## Instructions & Cover Page
RFI NUMBER RT13-2022
REQUEST FOR INFORMATION ON THE TECHNICAL SPECIFICATIONS
TRANSVERSAL CONTRACT DESCRIPTION: SUPPLY AND DELIVERY OF RESPIRATORY AIDS AND ANAESTHETIC ACCESSORIES TO THE STATE
CLOSING DATE AND TIME FOR SUBMISSION OF INFORMATION ON THE TECHNICAL SPECIFICATIONS IS 21 OCTOBER 2022 AT 16H00
In responding to the RFI, potential bidders are requested to complete ROW 14 to 17 on the following table as a cover page.
COVER PAGE TO BE COMPLETED BY POTENTIAL BIDDERS
To | Department of National Treasury (Transversal Contracting) | Date | 2022-10-21
Attention | Ms Masea Ramaloko
RFI no | RFI RT13-2022
Tel no | Insert the contact details of the company | E-mail address | Insert the email address of company
From | Insert the registered full legal name of the company | Address | Insert registered address of the company
Address | Insert the physical address of the company
Sender | Insert the full name of the sender at the company
Description of Transversal Contract | Supply and delivery of respiratory aid and anaesthetic accessories to the state
No of copies for submission | 1 x soft copy excel sheet submitted electronically
1. BACKGROUND
1.1 National Treasury, Office of the Chief Procurement Office (OCPO): Transversal Contracting Unit facilitates the arrangement and management of Transversal Term Contracts on behalf of State Institutions. With collaboration with National Department of Health this office is in the process to renew a Transversal Contract RT13: Supply and delivery of respiratory aids and anaesthetic accessories to the state.
1.2 To address the issue of non-compliance on Technical Specifications, National Treasury, Transversal Contracting Unit (TCU) is embarking on consulting with the industry for the review of the Technical Specifications.
2. EXPECTATION FROM POTENTIAL BIDDERS ON THE TECHNICAL SPECIFICATIONS AND SUBMISSION OF INFORMATION
2.1 Potential bidders are required to complete the instructions and cover page.
2.2 Potential bidders are required to review the technical specifications, where the potential bidders believe the technical specifications are not correct, you are advised to write your INPUTS on column E and highlight them for clear visibility.
2.3 If there is any additional information on the inputs and items on the Technical Specifications, it can submit and clearly marked with the category number and line-item number.
2.4 The information should be submitted on email be in an XLSX excel sheet format to Demand.Acquisition2@treasury.gov.za and NOT at the Tender Information Centre or to any other email address other than the one provided.
3. CONSULTATION PROCESS ON THE INPUTS TO BE RECEIVED
3.1 Upon the closing date of the inputs on the technical specifications, TCU will schedule a 2 days industry consultation meeting to discuss the inputs received on the Technical Specifications.
4. ENQUIRIES
4.1 All enquiries should be in writing to Demand.Acquisition2@treasury.gov.za

## Category 1
BID NUMBER | RT13
BID DESCRIPTION | SUPPLY AND DELIVERY OF RESPIRATORY AID AND ANAESTHETIC ACCESSORIES TO THE STATE
CATEGORY 1
OXYGEN DELIVERY DEVICES
| Item Number | Technical Specification | Product Name and Size | Insert inputs on this column
NASAL CANNULA
CATEGORY 1A | RT13-01-01A | NASAL CANNULA Purpose: For low flow nasal oxygen delivery in a spontaneously breathing patient Clear or colour tinted non-kinking, non-crush PVC tubing > 2m long Over-the-ear style with under-chin toggle or sleeve. Nasal cannula tips must be soft, anatomically shaped (curved) and tapered at the ends Any portion of the cannula which comes in contact with facial skin must be soft and uniformly contoured without protruding edges or ridges or flaps All components must be latex-free clinically clean and individually packed in peel pouch that is easy to open - For SINGLE use only | Nasal Cannula for oxygen delivery for use in an Extra Small Neonatal patient's Prong diameter: 1.5mm Prong length: 5.5mm Distance between medial surfaces of the prongs: 4mm
| RT13-01-02A |  | Nasal Cannula for oxygen delivery for use in Neonatal patients Prong diameter: 3mm Prong length: 8.0mm Distance between medial surfaces of the prongs: 5mm
| RT13-01-03A |  | Nasal Cannula for oxygen delivery for use in Infants Prong diameter: 4mm Prong length: 10mm Distance between medial surfaces of the prongs: 7mm
| RT13-01-04A |  | Nasal Cannula for oxygen delivery for use in Paediatric/ Child Prong diameter: 4mm Prong length:10mm Distance between medial surfaces of the prongs: 10mm
| RT13-01-05A |  | Nasal Cannula for oxygen delivery for use in an Adult Prong diameter: 4mm Prong length: 10mm ± 2mm Distance between medial surfaces of the prongs: 12mm
NASAL CANNULA WITH INTEGRATED C02 SAMPLING |  | NASAL CANNULA WITH INTEGRATED C02 SAMPLING
CATEGORY 1B | RT13-01-01B | NASAL CANNULA WITH INTEGRATED C02 SAMPLING Purpose: For low flow nasal oxygen delivery in a spontaneously breathing patient, with simultaneous C02 sampling Clear or colour tinted non-kinking, non-crush PVC tubing at least 2m long Over-the-ear style with under-chin toggle or sleeve. Nasal cannula tips must be soft, anatomically shaped (curved) and tapered at the ends Any portion of the cannula which comes in contact with facial skin must be soft and uniformly contoured without protruding edges or ridges or flaps Additional C02 sampling line with a male luer connector/ attached to the PVC tubing All components must be latex-free clinically clean and individually packed in peel pouch that is easy to open - For SINGLE use only | Nasal Cannula with integrated C02 sampling for use in an Extra Small Neonatal patients Prong diameter: 1.5mm Prong length: 5.5mm Distance between medial surfaces of the prongs: 4mm
| RT13-01-02B |  | Nasal Cannula with integrated C02 sampling for use in Neonatal patients Prong diameter: 3mm Prong length: 8.0mm Distance between medial surfaces of the prongs: 5mm
| RT13-01-03B |  | Nasal Cannula with integrated C02 sampling for use in Infants Prong diameter: 4mm Prong length: 10mm Distance between medial surfaces of the prongs: 7mm
| RT13-01-04B |  | Nasal Cannula with integrated C02 sampling for use in Paediatric/ Child Prong diameter: 4mm Prong length:10mm Distance between medial surfaces of the prongs: 10mm
| RT13-01-05B |  | Nasal Cannula with integrated C02 sampling for use in an Adult Prong diameter: 4mm Prong length:10mm ± 2mm Distance between medial surfaces of the prongs: 12mm
NASAL CANNULA - HYPOALLERGENIC
CATEGORY 1C | RT13-01-01C | NASAL CANNULA - HYPOALLERGENIC Purpose: For low flow nasal oxygen delivery in a spontaneously breathing patient with allergies Clear or colour tinted non-kinking, non-crush PVC tubing at least 2m long Over-the-ear style with under-chin toggle or sleeve. Nasal cannula tips must be soft, anatomically shaped (curved) and tapered at the ends Any portion of the cannula which comes in contact with facial skin must be soft and uniformly contoured without protruding edges or ridges or flaps All components must be latex-free, DEHP free and safe for use in hypoallergenic patients Clinically clean and individually packed in peel pouch that is easy to open - For SINGLE use only | Nasal Cannula for oxygen delivery to an Extra Small Neonatal patient's - Hypoallergenic Prong diameter: 1.5mm Prong length: 5.5mm Distance between medial surfaces of the prongs: 4mm
| RT13-01-02C |  | Nasal Cannula for oxygen delivery to Neonatal patient's- Hypoallergenic Prong diameter: 3mm Prong length: 8.0mm Distance between medial surfaces of the prongs: 5mm
| RT13-01-03C |  | Nasal Cannula for oxygen delivery to Infants - Hypoallergenic Prong diameter: 4mm Prong length: 10mm Distance between medial surfaces of the prongs: 7mm
| RT13-01-04C |  | Nasal Cannula for oxygen delivery to Paediatric/ Child - Hypoallergenic Prong diameter: 4mm Prong length:10mm Distance between medial surfaces of the prongs: 10mm
| RT13-01-05C |  | Nasal Cannula for oxygen delivery to an Adult - Hypoallergenic Prong diameter: 4mm Prong length: 10mm ± 2mm Distance between medial surfaces of the prongs: 12mm
NASAL CANNULA WITH INTEGRATED C02 SAMPLING - HYPOALLERGENIC
CATEGORY 1D | RT13-01-01D | NASAL CANNULA WITH INTEGRATED C02 SAMPLING - HYPOALLERGENIC Purpose: For low flow nasal oxygen delivery in a spontaneously breathing patient, with simultaneous C02 sampling Clear or colour tinted non-kinking, non-crush PVC tubing at least 2m long Over-the-ear style with under-chin toggle or sleeve. Nasal cannula tips must be soft, anatomically shaped (curved) and tapered at the ends Any portion of the cannula which comes in contact with facial skin must be soft and uniformly contoured without protruding edges or ridges or flaps Additional C02 sampling line with a male luer connector/ attached to the PVC tubing All components must be latex-free, DEHP free and safe for use in hypoallergenic patients Clinically clean and individually packed in peel pouch that is easy to open - For SINGLE use only | Nasal Cannula with integrated C02 sampling for use in an Extra Small Neonatal patients - Hypoallergenic Prong diameter: 1.5mm Prong length: 5.5mm Distance between medial surfaces of the prongs: 4mm
| RT13-01-02D |  | Nasal Cannula with integrated C02 sampling for use in Neonatal patients - Hypoallergenic Prong diameter: 3mm Prong length: 8.0mm Distance between medial surfaces of the prongs: 5mm
| RT13-01-03D |  | Nasal Cannula with integrated C02 sampling for use in Infants - Hypoallergenic Prong diameter: 4mm Prong length: 10mm Distance between medial surfaces of the prongs: 7mm
| RT13-01-04D |  | Nasal Cannula with integrated C02 sampling for use in Paediatric/ Child - Hypoallergenic Prong diameter: 4mm Prong length:10mm Distance between medial surfaces of the prongs: 10mm
| RT13-01-05D |  | Nasal Cannula with integrated C02 sampling for use in an Adult - Hypoallergenic Prong diameter: 4mm Prong length:10mm ± 2mm Distance between medial surfaces of the prongs: 12mm
VENTURI OXYGEN MASK
CATEGORY 1E | RT13-01-01E | VENTURI OXYGEN MASK Purpose: Delivery of oxygen and entrained room air to a spontaneously breathing patient. Modular design correctly sized and comfortable anatomically shaped mask to fit over the patient’s nose and mouth Under the chin cover for a better seal. Atraumatic to the nose. 2 exhalation holes in the side of the mask, a secure retaining strap and a malleable nose plate Flexible clear kink resistant corrugated tube: Length 10 - 20cm with a 22mm connector Crush proof delivery tubing > 2 m Colour coded Venturi device the Oxygen concentration in % and required flow rate must be moulded or clearly indicated on the device Components securely attached to each other and not separate accidentally in normal use. Latex free medical grade plastic - must comply with BS EN 13544-3:2001 standards Individually packed, easy to open without tubing getting tangled, clinically clean - SINGLE use Masks will be considered in series | Venturi Oxygen Mask 24% - Paediatric Colour coded venturi: Blue
| RT13-01-02E |  | Venturi Oxygen Mask 28% - Paediatric Colour coded venturi: Yellow
| RT13-01-03E |  | Venturi Oxygen Mask 40% - Paediatric Colour coded venturi: Pink
| RT13-01-04E |  | Venturi Oxygen Mask 60% - Paediatric Colour coded venturi: Red
| RT13-01-05E |  | Venturi Oxygen Mask 28% - Adult Colour coded venturi: Yellow
| RT13-01-06E |  | Venturi Oxygen Mask 40% - Adult Colour coded venturi: Pink
| RT13-01-07E |  | Venturi Oxygen Mask 60% - Adult Colour coded venturi: Red
VENTURI OXYGEN MASK: ELONGATED
CATEGORY 1F | RT13-01-01F | VENTURI OXYGEN MASK: ELONGATED Purpose: Delivery of oxygen and entrained room air to a spontaneously breathing patient. Modular design correctly sized and comfortable elongated anatomically shaped mask to fit over the patient’s nose and mouth Under the chin cover for a better seal. Atraumatic to the nose. 2 exhalation holes in the side of the mask, a secure retaining strap and a malleable nose plate Flexible clear kink resistant corrugated tube: Length 10 - 20cm with a 22mm connector Crush proof delivery tubing > 2 m Colour coded Venturi device the Oxygen concentration in % and required flow rate must be moulded or clearly indicated on the device Components securely attached to each other and not separate accidentally in normal use. Latex free medical grade plastic - must comply with BS EN 13544-3:2001 standards Individually packed, easy to open without tubing getting tangled, clinically clean - SINGLE use Masks will be considered in series | Venturi Oxygen Mask 28% - Elongated Adult Colour coded venturi: Yellow
| RT13-01-02F |  | Venturi Oxygen Mask 40% - Elongated Adult Colour coded venturi: Pink
| RT13-01-03F |  | Venturi Oxygen Mask 60% - Elongated Adult Colour coded venturi: Red
VENTURI OXYGEN MASK - MRI COMPATIBLE
CATEGORY 1G | RT13-01-01G | VENTURI OXYGEN MASK - MRI COMPATIBLE Purpose: Delivery of oxygen and entrained room air to a spontaneously breathing patient. Suitable for the MRI Suite Modular design correctly sized and comfortable anatomically shaped preformed mask shaped to fit over the patient’s nose and mouth with under the chin cover for a better seal Atraumatic to the nose. 2 exhalation holes in the side of the mask, a secure retaining strap. All components to be MRI compatible. Flexible corrugated tube: Length 10 - 20cm with a 22mm connector Must have clear kink resistant, crush proof delivery tubing > 2 m Colour coded Venturi device with the Oxygen concentration in % and required flow rate must be moulded or clearly indicated on the device The components must be securely attached to each other and not separate accidentally in normal use. Manufactured from latex-free medical grade plastic - must comply with BS EN 13544-3:2001 standards Individually packed, easy to open without tubing getting tangled, clinically clean - SINGLE use | Venturi Oxygen Mask MRI COMPATIBLE 24% - Paediatric Colour coded venturi: Blue
| RT13-01-02G |  | Venturi Oxygen Mask MRI COMPATIBLE 28% - Paediatric Colour coded venturi: Yellow
| RT13-01-03G |  | Venturi Oxygen Mask MRI COMPATIBLE 40% - Paediatric Colour coded venturi: Pink
| RT13-01-04G |  | Venturi Oxygen Mask MRI COMPATIBLE 60% - Paediatric Colour coded venturi: Red
| RT13-01-05G |  | Venturi Oxygen Mask MRI COMPATIBLE 28% - Adult Colour coded venturi: Yellow
| RT13-01-06G |  | Venturi Oxygen Mask MRI COMPATIBLE 40% - Adult Colour coded venturi: Pink
| RT13-01-07G |  | Venturi Oxygen Mask MRI COMPATIBLE 60% - Adult Colour coded venturi: Red
VENTURI OXYGEN MASK ELONGATED- MRI COMPATIBLE
CATEGORY 1H | RT13-01-01H | VENTURI OXYGEN MASK ELONGATED- MRI COMPATIBLE Purpose: Delivery of oxygen and entrained room air to a spontaneously breathing patient. Suitable for the MRI Suite Modular design correctly sized and comfortable anatomically shaped elongated preformed mask shaped to fit over the patient’s nose and mouth with under the chin cover for a better seal Atraumatic to the nose. 2 exhalation holes in the side of the mask, a secure retaining strap. All components to be MRI compatible. Flexible corrugated tube: Length 10 - 20cm with a 22mm connector Must have clear kink resistant, crush proof delivery tubing > 2 m Colour coded Venturi device with the Oxygen concentration in % and required flow rate must be moulded or clearly indicated on the device The components must be securely attached to each other and not separate accidentally in normal use. Manufactured from latex-free medical grade plastic - must comply with BS EN 13544-3:2001 standards Individually packed, easy to open without tubing getting tangled, clinically clean - SINGLE use | Venturi Oxygen Mask MRI COMPATIBLE 28% - Elongated Adult Colour coded venturi: Yellow
| RT13-01-02H |  | Venturi Oxygen Mask MRI COMPATIBLE 40% - Elongated Adult Colour coded venturi: Pink
| RT13-01-03H |  | Venturi Oxygen Mask MRI COMPATIBLE 60% - Elongated Adult Colour coded venturi: Red
VENTURI OXYGEN MASK REGULATING PACK
CATEGORY 1I | RT13-01-01I | VENTURI OXYGEN MASK REGULATING PACK Purpose: Delivery of oxygen and entrained room air to a spontaneously breathing patient with the option of 4 venturi concentrators Modular design correctly sized and comfortable anatomically shaped mask shaped to fit over the patient’s nose and mouth with under the chin cover for a better seal Atraumatic to the nose. Must have 2 exhalation holes in the side of the mask, a secure retaining strap and a malleable nose plate Flexible corrugated tube: Length 10 - 20cm with a 22mm connector Must have clear kink resistant, crush proof delivery tubing > 2 m 4 Colour coded Venturi devices 28% Yellow, 35% Green, 40% Pink, 60% Red with the Oxygen concentration in % and required flow rate must be moulded or clearly indicated on the device The components must be securely attached to each other and not separate accidentally in normal use. Manufactured from latex-free medical grade plastic - must comply with BS EN 13544-3:2001 standards Individually packed, easy to open without tubing getting tangled, clinically clean - SINGLE use | Venturi Oxygen Regulating pack of 4 concentrations – 28% Yellow, 35% Green, 40% Pink, 60% Red - Paediatric mask
| RT13-01-02I |  | Venturi Oxygen Regulating pack of 4 concentrations – 28% Yellow, 35% Green, 40% Pink, 60% Red - Adult mask
VENTURI OXYGEN MASK REGULATING PACK
CATEGORY 1J | RT13-01-01J | VENTURI OXYGEN MASK and ADJUSTABLE DIAL Purpose: Delivery of oxygen and entrained room air to a spontaneously breathing patient with the option of a dial that adjusts the oxygen concentration from 21% to 50% Modular design correctly sized and comfortable anatomically shaped mask shaped to fit over the patient’s nose and mouth with under the chine cover for a better seal Atraumatic to the nose. 2 exhalation holes in the side of the mask, a secure retaining strap and a malleable nose plate Flexible corrugated tube: Length 10 - 20cm with a 22mm connector Must have clear kink resistant, crush proof delivery tubing > 2 m The adjustable dial must turn easily to set the concentration - but not be change the concentration with standard movement The components must be securely attached to each other and not separate accidentally in normal use. Manufactured from latex-free medical grade plastic - must comply with BS EN 13544-3:2001 standards Individually packed, easy to open without tubing getting tangled, clinically clean - SINGLE use | Venturi Oxygen Mask Adjustable % - Paediatric
| RT13-01-02J |  | Venturi Oxygen Mask Adjustable % - Adult
VENTURI OXYGEN MASK REGULATING PACK
CATEGORY 1K | RT13-01-01K | VENTURI T-PIECE Purpose: Delivery of oxygen and entrained room air to a spontaneously breathing patient via an endotracheal tube or tracheostomy Modular design consisting of a T-piece connector. Base of ‘T’ must fit a standard 15mm connector. Side limbs of ‘T’ must be connected to a venturi device via flexible corrugated tube Flexible corrugated tube: Length 10 - 15cm Colour coded Venturi device with the Oxygen concentration in % and required flow rate must be moulded or clearly indicated on the device The components must be securely attached to each other and not separate accidentally in normal use. Latex free medical grade plastic Individually packed, easy to open without tubing getting tangled, clinically clean - SINGLE use To be considered in series | Venturi Oxygen T-piece 28% Colour coded venturi: Yellow
| RT13-01-02K |  | Venturi Oxygen T-piece 35% Colour coded venturi:Green
| RT13-01-03K |  | Venturi Oxygen T-piece 40% Colour coded venturi: Pink
| RT13-01-04K |  | Venturi Oxygen T-piece 60% Colour coded venturi: Red
FACE MASK WITH INTEGRATED C02 SAMPLING
CATEGORY 1L | RT13-01-01L | Face mask with integrated C02 sampling: Adult Purpose: Delivery of oxygen with simultaneous C02 sampling in sedated patient's Modular design consisting of an adult sized mask shaped to fit over the patient’s nose and mouth with a secure retaining strap and malleable nose plate – 2 exhalation holes on the side of the mask. The device should deliver different Oxygen concentrations based on the flow rate: Oxygen flow rates at 5 l/min should deliver 35% Fi02; at 6l/min – 40% Fi02; at 8l/min – 50% Fi02 A C02 sampling line: 2m, with a male luer lock connection at the end must be part of the system Manufactured from latex -free medical grade plastic. For SINGLE use; Individually packed , Clinically clean | Face mask with integrated C02 sampling: Adult
REBREATHING MASKS
CATEGORY 1M | RT13-01-01M | REBREATHING MASKS Purpose: Delivery of high concentrations of oxygen to a spontaneously breathing patient – safety feature in case of interruption of oxygen supply Modular designed, correctly sized and comfortable mask shaped to fit over the patient’s nose and mouth and chin. Atraumatic to the nose. The mask must have 2 exhalation holes in the side of the mask that can entrain room air, a secure retaining strap and a malleable nose plate Soft, strong, vinyl Oxygen reservoir bag attached to the base of the mask, with a two-way valve located between the mask and the reservoir bag Reservoir bag must only partially collapse on inspiration, must always contain ⅓CO2 and ⅔ O2 Kink resistant delivery tubing > 2 m The components must be securely attached to each other and not separate accidentally in normal use. Latex free, soft odourless polyvinyl chloride Individually packed, easy to open without tubing getting tangled, clinically clean - SINGLE use | Rebreather Oxygen Mask - Paediatric
| RT13-01-02M |  | Rebreather Oxygen Mask - Adult
| RT13-01-03M |  | Rebreather Oxygen Mask - Elongated Adult
NON-REBREATHING MASKS
CATEGORY 1N | RT13-01-01N | NON-REBREATHING MASKS Purpose: Delivery of high concentrations of oxygen to a spontaneously breathing patient – safety feature in case of interruption of oxygen supply Modular designed, correctly sized and comfortable mask shaped to fit over the patient’s nose and mouth and chin. The mask must have 2 exhalation holes in the side of the mask, a secure retaining strap and a malleable nose plate There must be two, one-way expiration valves manufacture of soft odourless polyvinyl chloride present on the exhalation holes. Removal of the valves must allow room air to be entrained - decreasing the Fi02 from 91% to 60 % A soft and strong Oxygen reservoir bag must be attached to base of mask, with a one-way valve located between the mask and the reservoir bag to prevent exhaled air from entering the reservoir bag. The reservoir bag must only partially collapse on inspiration Must have kink resistant delivery tubing > 2m The required oxygen flow rate must be stated for preferred Oxygen requirements. The components must be securely attached to each other and not separate accidentally in normal use. Manufactured from latex-free medical grade plastic Individually packed, easy to open without tubing getting tangled, clinically clean SINGLE use | Non-rebreather Oxygen Mask - Paediatric
| RT13-01-02N |  | Non-rebreather Oxygen Mask - Adult
| RT13-01-03N |  | Non-rebreather Oxygen Mask - Elongated Adult
NEBULISER SETS
CATEGORY 1O | RT13-01-01O | NEBULISER SETS Purpose: To deliver a fine mist of droplets for inhalational therapy Modular designed correctly sized and comfortable mask shaped to fit over the patient’s nose and mouth that is secured with a retaining strap and has a malleable nose plate and 2 exhalation holes in the side of the mask. Atraumatic to the nose Must have a kink resistant delivery tubing: > 2 m The Jet nebuliser chamber must have a capacity of > 5ml and < 10ml The top must be easy to screw and unscrew and must provide a good seal Must produce a particle size of 1-5 µ at an oxygen flow rate of 6 - 8 l/min but must not come apart from the tubing at oxygen flow rates of 8 l/min Must have a swivel connector that can rotate around 360° Must operate at vertical and horizontal positions All components must be securely attached to each other and must not separate accidentally in normal use. Latex-free, soft, medical grade PVC Clinically Clean - For SINGLE use only | Infant nebuliser set consisting of an infant mask, medication chamber and oxygen tubing and swivel connector
| RT13-01-02O |  | Child nebuliser set consisting of a child mask, medication chamber and oxygen tubing and swivel connectors
| RT13-01-03O |  | Adult nebuliser set consisting of an adult mask, medication chamber and oxygen tubing and swivel connector
| RT13-01-04O |  | Adult nebuliser set consisting of an elongated adult mask, medication chamber and oxygen tubing and swivel connector
HAND-HELD NEBULISER SET
CATEGORY 1P | RT13-01-01P | HAND-HELD NEBULISER SET with mouthpiece, T-connector, medication chamber and oxygen tubing Purpose: To deliver a fine mist of droplets for inhalational therapy Handheld mouthpiece - integral or seperate. Attached securely to a T-connector with flexible corrugated tube: Length 10 - 15cm. Base of ‘T’ must fit a standard 15mm connector. Jet nebuliser chamber must have a capacity of > 5 and < 10ml and must produce a particle size of 1-5 µ at an oxygen flow rate of 6 - 8 l/min The components must not come apart from the tubing at oxygen flow rates of 8 litres per minute. Must operate at vertical and horizontal position The top must be easy to screw and unscrew and must provide a good seal Kink resistant Oxygen tubing > 2 m in length, and must fit securely to the nebuliser Clinically Clean. For SINGLE use only | HAND-HELD NEBULISER SET with mouthpiece, T-connector, medication chamber and oxygen tubing
IN-LINE NEBULISER SET
CATEGORY 1Q | RT13-01-01Q | IN-LINE NEBULISER SET with self-sealing T-Piece, medication chamber and oxygen tubing Purpose: To deliver a fine mist of droplets for inhalational therapy Consists of a T-Piece with standard 22mm connector on either side. The T-piece must attach to the ventilator system without causing leaks or adding to the weight of the circuit. The Jet nebuliser chamber must have a capacity of > 5 and < 10ml and must produce a particle size of 1-5 µ at an oxygen flow rate of 6 - 8 l/min The components must not come apart from the tubing at oxygen flow rates of 8 litres per minute The top must be easy to screw and unscrew and must provide a good seal Kink resistant Oxygen tubing > 2 m in length, and must fit securely to the nebuliser Clinically Clean. For SINGLE use only | IN-LINE NEBULISER SET with self-sealing T-Piece, medication chamber and oxygen tubing
OROPHARYNGEAL AIRWAYS
CATEGORY 1R | RT13-01-01R | OROPHARYNGEAL AIRWAYS Purpose: Used to maintain oral airway patency Anatomically shaped PVA or PVC airway with an appropriate contour that has a flanged buccal end with the size clearly moulded onto the plastic. Must have a robust bite block to prevent occlusion, smooth rounded distal end to minimise tissue trauma. The edges or corners that come into contact with patient tissue, must be smooth with a minimal radius of curvature as per ISO 5364:2016 One-piece design or reinforcement components for that do not separate easily or could be an aspiration risk Distance measured from buccal flange to tip of OPA with tolerance as per ISO 5364:2016 Validation of measurements must be provided. Brochures will be acceptable. All components must be latex-free, clinically clean and individually packed in peel pouch that is easy to open - For SINGLE use only For safety and standardisation purposes, offers for items sized 30mm and/or 35mm to be awarded individually Items sized 40mm - 70mm will be considered in series.. Items sized 80mm - 100mm will be considered in series. Item sized 110mm and 120mm will be awarded individually | Oropharyngeal airway (guedels) 30mm Size 000 (Lilac/Clear) Minimum inside dimension: 2.5mm
| RT13-01-02R |  | Oropharyngeal airway (guedels) 35mm Size 000 (Light Green) Minimum inside dimension: 3.0mm
| RT13-01-03R |  | Oropharyngeal airway (guedels) 40-45mm Size 00 (Pink/Cream) Minimum inside dimension: 3.0mm
| RT13-01-04R |  | Oropharyngeal airway (guedels) 50-55mm Size 0 (Light Blue/Grey) Minimum inside dimension: 3.5mm
| RT13-01-05R |  | Oropharyngeal airway (guedels) 60 - 65mm 6.0 Size 1 (Black/Brown) Minimum inside dimension: 4.0mm
| RT13-01-06R |  | Oropharyngeal airway (guedels) 70mm Size 2 (White) Minimum inside dimension: 4.0mm
| RT13-01-07R |  | Oropharyngeal airway (guedels) 80mm ID: 3 (Green) Minimum inside dimension: 4.5mm
| RT13-01-08R |  | Oropharyngeal airway (guedels) 90mm Size 4 (Yellow) Minimum inside dimension: 4.5mm
| RT13-01-09R |  | Oropharyngeal airway (guedels) 100mm Size 5 (Red) Minimum inside dimension: 5.0mm
| RT13-01-10R |  | Oropharyngeal airway (guedels) 110mm (Orange) Minimum inside dimension: 5.5mm
| RT13-01-11R |  | Oropharyngeal airway (guedels) 120mm (Purple) Minimum inside dimension: 5.5mm
ORAL FIBRE OPTIC GUIDES - Berman type
CATEGORY 1S | RT13-01-01S | ORAL FIBRE OPTIC GUIDES - Berman type Purpose: Modified oral airway specifically designed for oral fibre-optic intubation. Used to guide the tracheal tube and scope around the tongue and protect the scope should the patient bite down Straight with groove to allow removal of scope upon intubation. Latex free medical grade plastic SINGLE use, clinically clean. Individually packed in peel pouch | Oral Fibre-Optic Guide – Berman type Size 000 Length: 40mm
| RT13-01-02S |  | Oral Fibre-Optic Guide – Berman type Size 00 Length: 50mm
| RT13-01-03S |  | Oral Fibre-Optic Guide – Berman type Size 0 Length: 60mm
| RT13-01-04S |  | Oral Fibre-Optic Guide – Berman type Size 1 Length: 70mm
| RT13-01-05S |  | Oral Fibre-Optic Guide – Berman type Size 2 Length: 80mm
| RT13-01-06S |  | Oral Fibre-Optic Guide – Berman type Size 3 Length: 90mm
| RT13-01-07S |  | Oral Fibre-Optic Guide – Berman type Size 4 Length: 100mm
| RT13-01-08S |  | Oral Fibre-Optic Guide – Berman type Size 5 Length: 110mm
ORAL FIBRE OPTIC GUIDES - Ovassapian type Adult
CATEGORY 1T | RT13-01-01T | ORAL FIBRE OPTIC GUIDES - Ovassapian type Adult Purpose: Modified oral airways specifically designed for oral fibre-optic intubation. Anatomically shaped EVA or PVC airway with an appropriate atraumatic contour The edges or corners that come into contact with patient tissue, must be smooth One-piece design with a strong bite-block to protect the fibre-optic should the patient bite down Must split longitudinally to facilitate the placing of an endotracheal tube over the fibre optic scope All components must be latex-free, clinically clean and individually packed in peel pouch that is easy to open - For SINGLE use only | Oral Fibre-optic Guide - Ovassapian type Size 2
| RT13-01-02T |  | Oral Fibre-optic Guide - Ovassapian type Size 3
| RT13-01-03T |  | Oral Fibre-optic Guide - Ovassapian type Size 4
NASOPHARYNGEAL AIRWAYS with OXYGEN PORTS
CATEGORY 1U | RT13-01-01U | NASOPHARYNGEAL AIRWAYS with OXYGEN PORTS Purpose: Nasally placed airway management Anatomically shaped soft silicone airway with an appropriate contour that maintains the pharyngeal curvature without collapse of the lumen Proximal end: Trumpeted end with integrated 15mm oxygen adaptor/ flanged end with a15mm oxygen adaptor. Separate oxygen elbow to allow for oxygen delivery Distal end: Soft atraumatic bevelled tip Must have sizing and dimension clearly marked on the device All components must be MRI compatible, latex-free, clinically clean and individually packed in peel pouch that is easy to open -For SINGLE use only | Nasopharyngeal airway with Oxygen Port – 18Fr ID: 4.0
| RT13-01-02U |  | Nasopharyngeal airway with Oxygen Port – 22Fr ID: 5.0
| RT13-01-03U |  | Nasopharyngeal airway with Oxygen Port – 25Fr ID: 6.0
| RT13-01-04U |  | Nasopharyngeal airway with Oxygen Port – 29Fr ID: 7.0
| RT13-01-05U |  | Nasopharyngeal airway with Oxygen Port – 32Fr ID: 8.0
NASOPHARYNGEAL AIRWAYS with OXYGEN AND CAPNOGRAPHY PORTS
CATEGORY 1V | RT13-01-01V | NASOPHARYNGEAL AIRWAYS with OXYGEN AND CAPNOGRAPHY PORTS Purpose: Nasally placed airway management Anatomically shaped soft silicone airway with an appropriate contour that maintains the pharyngeal curvature without collapse of the lumen Proximal end: Trumpeted end with integrated 15mm oxygen adaptor/ flanged end with a15mm oxygen adaptor. Separate oxygen elbow with direct side stream sampling line to allow for oxygen delivery and ETCO2 monitoring. Distal end: Soft atraumatic bevelled tip Must have sizing and dimension clearly marked on the device All components must be MRI compatible, latex-free, clinically clean and individually packed in peel pouch that is easy to open - For SINGLE use only | Nasopharyngeal airway with Oxygen Port and Capnography Line – 18Fr ID: 4.0
| RT13-01-02V |  | Nasopharyngeal airway with Oxygen Port and Capnography Line – 22Fr ID: 5.0
| RT13-01-03V |  | Nasopharyngeal airway with Oxygen Port and Capnography Line – 25Fr ID: 6.0
| RT13-01-04V |  | Nasopharyngeal airway with Oxygen Port and Capnography Line – 29Fr ID: 7.0
| RT13-01-05V |  | Nasopharyngeal airway with Oxygen Port and Capnography Line – 32Fr ID: 8.0
ANAESTHETIC MASKS: Round REUSABLE.
CATEGORY 1X | RT13-01-01X | ANAESTHETIC MASKS: Round REUSABLE. Purpose: Used during bag mask ventilation Latex-free round, clear silicone mask with a 15mm/22mm connector Malleable, non-collapsible with an easy user interface Must have a comfortable patient fit, be atraumatic to the nasal bridge and create a good seal between the mask and patient's’ face. Must have a smooth surface without ridges - must not cause skin trauma There must be a good fit between the mask connector and breathing circuit. The size of the mask must clearly be indicated on the mask Clinically clean and individually packed. Autoclavable | Anaesthetic mask. Round Reusable. For use in premature babies Size: 000
| RT13-01-02X |  | Anaesthetic mask. Round Reusable.For use in neonates Size: 00
| RT13-01-03X |  | Anaesthetic mask. Round Reusable. For use in small infants Size: 0
| RT13-01-04X |  | Anaesthetic mask. Round Reusable. For use in large infants Size: 1
| RT13-01-05X |  | Anaesthetic mask. Round Reusable. For use in small child Size: 2
ANAESTHETIC MASKS: Round SINGLE use.
CATEGORY 1Y | RT13-01-01Y | ANAESTHETIC MASKS: Round SINGLE use. Purpose: Used during bag mask ventilation Consist of a latex-free round, clear silicone mask with a 15mm/22mm connector Must be malleable but non-collapsible with an easy user interface Must have a comfortable patient fit, be atraumatic to the nasal bridge and create a good seal between the mask and patient's’ face. Must have a smooth surface without ridges - must not cause skin trauma There must be a good fit between the mask connector and breathing circuit. The size of the mask must clearly be indicated on the mask Clinically clean and individually packed in easy to open packaging - SINGLE use Must comply with SANS 1866-1:2018 or equivalent | Anaesthetic mask. Round SINGLE use. For use in premature babies Size: 000
| RT13-01-02Y |  | Anaesthetic mask. Round SINGLE use. For use in neonates Size: 00
| RT13-01-03Y |  | Anaesthetic mask. Round SINGLE use. For use in small infants Size: 0
| RT13-01-04Y |  | Anaesthetic mask. Round SINGLE use. For use in large infants Size: 1
| RT13-01-05Y |  | Anaesthetic mask. Round SINGLE use. For use in a child Size: 2
ANAESTHETIC MASKS: Anatomically correct - REUSABLE.
CATEGORY 1Z | RT13-01-01Z | ANAESTHETIC MASKS: Anatomically correct - REUSABLE. Purpose: Used during bag mask ventilation Latex free anatomically correct, soft rim, clear silicone mask Must be a SINGLE component design, malleable but non-collapsible Must have a comfortable patient fit, be atraumatic to the nasal bridge Must create a good seal between the mask and patient's’ face with an easy user interface Must have a smooth surface without ridges - must not cause skin trauma There must be a good fit between the mask connector and breathing circuit. The size of the mask must clearly be indicated on the mask Clinically clean and individually packed in easy to open packaging Must be autoclavable | Anaesthetic mask. Anatomically correct – Reusable. For use in small infants Size: 0
| RT13-01-02Z |  | Anaesthetic mask. Anatomically correct – Reusable. For use in large infants Size: 1
| RT13-01-03Z |  | Anaesthetic mask. Anatomically correct – Reusable. For use in a Child Size: 2
| RT13-01-04Z |  | Anaesthetic mask. Anatomically correct – Reusable. For use in a Small Adult Size: 3
| RT13-01-05Z |  | Anaesthetic mask. Anatomically correct – Reusable. For use in a Medium Adult Size: 4
| RT13-01-06Z |  | Anaesthetic mask. Anatomically correct – Reusable. For use in a Large Adult . Size 5
| RT13-01-07Z |  | Anaesthetic mask. Anatomically correct – Reusable. For use in a XL Adult Size: 6
ANAESTHETIC MASKS: Anatomically correct - SINGLE USE
CATEGORY 1AA | RT13-01-01AA | ANAESTHETIC MASKS: Anatomically correct - SINGLE USE Purpose: Used during bag mask ventilation Latex free anatomically correct, soft rim, clear silicone mask Must be a SINGLE component design, malleable but non-collapsible Comfortable patient fit, atraumatic to the nasal bridge Must create a good seal between the mask and patient's’ face with an easy user interface Must have a smooth surface without ridges - must not cause skin trauma There must be a good fit between the mask connector and breathing circuit. The size of the mask must clearly be indicated on the mask Clinically clean and individually packed in easy to open packaging - SINGLE use Must comply with SANS 1866-1:2018 or equivalent | Anaesthetic mask. Anatomically correct – SINGLE use .For use in small infants Size: 0
| RT13-01-02AA |  | Anaesthetic mask. Anatomically correct – SINGLE use. For use in large infants Size: 1
| RT13-01-03AA |  | Anaesthetic mask. Anatomically correct – SINGLE use. For use in a Child Size: 2
| RT13-01-04AA |  | Anaesthetic mask. Anatomically correct – SINGLE use. For use in a Small Adult Size: 3
| RT13-01-05AA |  | Anaesthetic mask. Anatomically correct – SINGLE use. For use in a Medium Adult Size: 4
| RT13-01-06AA |  | Anaesthetic mask. Anatomically correct – SINGLE use. For use in a Large Adult . Size 5
| RT13-01-07AA |  | Anaesthetic mask. Anatomically correct – SINGLE use. For use in a XL Adult Size: 6
ANAESTHETIC MASKS WITH INFLATABLE RIM. Anatomically correct –SINGLE use.
CATEGORY 1BB | RT13-01-01BB | ANAESTHETIC MASKS WITH INFLATABLE RIM. Anatomically correct –SINGLE use. Purpose: Used during bag mask ventilation Latex-free anatomically moulded mask with inflatable rim that is connected to a harder shell. Malleable but not collapsible The soft profiled facial cushion must be pre-inflated to 7- 10mmHg with valve for air volume adjustment Must create a good seal between the mask and patient's’ face without causing damaging facial pressure and must be atraumatic to nasal bridge and must not cover the patient’s eyes Must have a smooth surface without ridges - must not cause skin trauma There must be a good fit between the mask connector and breathing circuit. The size of the mask must clearly be indicated on the mask and the mask must be transparent for optimum visual clarity. Manufactured from soft malleable medical grade plastic Clinically clean and individually packed in easy to open packaging - SINGLE use Must comply with SANS 1866-1:2018 or equivalent | Anaesthetic mask with inflatable rim, Anatomically Correct –SINGLE use.For use in an Infant Size 0
| RT13-01-02BB |  | Anaesthetic mask with inflatable rim, Anatomically Correct –SINGLE use. For use in a Small Child Size 1
| RT13-01-03BB |  | Anaesthetic mask with inflatable rim, Anatomically correct –SINGLE use. For use in a Large Child Size 2
| RT13-01-04BB |  | Anaesthetic mask with inflatable rim,. Anatomically correct –SINGLE use. For use in a small Adult Size 3
| RT13-01-05BB |  | Anaesthetic mask with inflatable rim,. Anatomically correct –SINGLE use. For use in a Medium Adult Size 4
| RT13-01-06BB |  | Anaesthetic mask with inflatable rim, Anatomically correct –SINGLE use. For use in a Large Adult Size 5
| RT13-01-07BB |  | Anaesthetic mask with inflatable rim, Anatomically correct –SINGLE use. For use in a XL Adult Size 6
ANAESTHETIC MASKS WITH INFLATABLE RIM, SCENTED. Anatomically correct –SINGLE use.
CATEGORY 1CC | RT13-01-01CC | ANAESTHETIC MASKS WITH INFLATABLE RIM, SCENTED. Anatomically correct – SINGLE use. Purpose: Used during bag mask ventilation Latex-free anatomically moulded mask with inflatable rim that is connected to a harder shell. Malleable but not collapsible The soft profiled facial cushion must be pre-inflated to 7- 10mmHg with valve for air volume adjustment Must create a good seal between the mask and patient's’ face without causing damaging facial pressure and must be atraumatic to nasal bridge and must not cover the patient’s eyes Must have a smooth surface without ridges - must not cause skin trauma There must be a good fit between the mask connector and breathing circuit. The size of the mask must clearly be indicated on the mask and the mask must be transparent for optimum visual clarity. Manufactured from soft malleable medical grade plastic Size: 0 - Size: 2 will be considered in series will be considered in series Clinically clean and individually packed in easy to open packaging - SINGLE use Must comply with SANS 1866-1:2018 or equivalent | Anaesthetic mask with inflatable rim, Scented, Anatomically Correct –SINGLE use.For use in an Infant Size 0
| RT13-01-02CC |  | Anaesthetic mask with inflatable rim, Scented, Anatomically Correct –SINGLE use. For use in a Small Child Size 1
| RT13-01-03CC |  | Anaesthetic mask with inflatable rim, Scented, Anatomically correct –SINGLE use. For use in a Large Child Size 2
CPAP MASK
CATEGORY 1DD | RT13-01-01DD | CPAP MASK Purpose: For non-invasive continuous positive airway pressure delivery The double-ported mask must connect to the PEEP valve and the oxygen source. The mask must have a hard clear durable anatomically moulded face plate with inflatable rim with hooks to secure harness. The system must be able to create a good seal between the mask and patient's face, must be atraumatic to the nasal bridge and must not cause skin trauma. There must be hook rings for the attachment of the harness Latex-free, single use, individually packed and surgical clean.
| RT13-01-02DD |  | CPAP Mask - Medium
| RT13-01-03DD |  | CPAP Mask - Large
| RT13-01-04DD |  | CPAP Mask - X-Large
| RT13-01-05DD |  | CPAP System - X-Large
CONSTANT POSITIVE AIRWAY PRESSURE SYSTEM (CPAP)
CATEGORY 1EE | RT13-01-01EE | CONSTANT POSITIVE AIRWAY PRESSURE SYSTEM (CPAP) Purpose: To administer fresh gas to the patient through the entire respiratory cycle. The system must include a mask with a pressure relief valve, harness, oxygen tubing and a PEEP valve. The double-ported mask must connect to the PEEP valve and the oxygen source. The mask must have a hard clear durable anatomically moulded face plate with inflatable rim with hooks to secure harness. The system must be able to create a good seal between the mask and patient's face, must be atraumatic to the nasal bridge and must not cause skin trauma. There must be hook rings for the attachment of the harness The harness must be comfortable, adjustable, easy to release and non-irritant. An adjustable PEEP valve or different sized PEEP valves must be included. Must be supplied with kink resistant O2 delivery tubing ± 2m. A standard flow meter should achieve 5-20cmH2O of CPAP. Additional adapters must be included to fit PEEP valves and oxygen tubing Latex-free, single use, individually packed and surgical clean. | CPAP System - Paediatric
| RT13-01-02EE |  | CPAP System - Small
| RT13-01-03EE |  | CPAP System - Medium
| RT13-01-04EE |  | CPAP System - Large
| RT13-01-05EE |  | CPAP System - X-Large
PEEP MANOMETER
CATEGORY 1FF | RT13-01-01FF | PEEP manometer Purpose: Dedicated manometer to measure PEEP values when using CPAP non-invasive systems. Must be suitable for use with manual resuscitators or CPAP masks. Visible pressure ranges to give enhanced visual support during ventilation ranging from 0-40cmH2O Must have an universal adaptor to connect standard resuscitation bags. Single use, individually packed and surgical clean | PEEP manometer
BAG VALVE MASK RESUSCITATION BAG WITH PEEP VALVE - AUTOCLAVABLE
CATEGORY 1GG | RT13-01-01GG | BAG VALVE MASK RESUSCITATION BAG WITH PEEP VALVE - AUTOCLAVABLE Purpose: Used to provide positive pressure ventilation to a patient who is not breathing or who is breathing inadequately Ventilation bag: Clear, robust silicone. Self-inflating with a fast recoil time, and thin-walled bag to allow lung compliance assessment. Must have an ergonomic grip, Oxygen reservoir bag: Reusable PVC One-way Intake / Reservoir valve between ventilation bag and reservoir bag that only allows one way flow into the ventilation bag Oxygen from reservoir bag and room air must enter the ventilation bag if fresh gas flow is inadequate. Secure connector for oxygen tubing. Oxygen tubing: >1 m Patient valves between ventilation bag and mask. Comprising of a lip valve (duck-billed valve), expiratory valve to prevent entrainment of room air, pressure release valve – 35cmH20 (must be printed on valve) Standard connector 15mm ID, 22mm OD. Must have a PEEP valve Masks with comfortable patient fit and user interface. Malleable, non-collapsible with smooth surface or ridges. 1 piece modular face mask Able to create a tight seal between the mask and patient's’ face and must be transparent. There must be a secure fit between mask connector and breathing circuit All connections between components must be easily assembled and disassembled with tight fit and no leak. Must not disassemble during normal use Manufactured from latex-free, clear silicone, PVC, and medical grade plastic - All components Autoclavable | Bag valve mask resuscitation bag with PEEP valve. Preterm Patient size: Up to 2.5kg Ventilation bag: Volume: 240ml , Oxygen reservoir bag: Volume: 600ml Peep valve adjustable to ± 1.5-10 cmH2O Must be supplied with 2 round masks – size 00 and 0/1
| RT13-01-02GG |  | Bag valve mask resuscitation bag with PEEP valve. Paediatric Patient size: 2.5kg – 20kg Ventilation bag: Volume: 500ml Oxygen reservoir bag: Volume: 600ml Peep valve adjustable to ± 1.5-10 cmH2O Must be supplied with 2 round masks – size 0/1 and 2
| RT13-01-03GG |  | Bag valve mask resuscitation bag with PEEP valve.Adult Patient size: > 40kg Ventilation bag: Volume: 1000- 1600ml Oxygen reservoir bag: Volume: 2600ml Peep valve adjustable to 5-20 cm/H20 Must be supplied with 2 anatomically correct silicone masks – size 4 and 5
VOLUMETRIC INCENTIVE SPIROMETRY
CATEGORY 1HH | RT13-01-01HH | VOLUMETRIC INCENTIVE SPIROMETRY Purpose: Handheld device that helps achieve optimum lung capacity and restoring disrupted breathing patterns. Consists of a single/multiple ball volumetric exerciser with volume scale indicator. Must have a connecting tube with 22mm OD connector and mouthpiece. Made from compact and sturdy plastic. Single use, individually packed and surgical clean | Volumetric Incentive Spirometer - Paediatric 2500ml
| RT13-01-02HH |  | Volumetric Incentive Spirometer - Adult 5000ml
MOUTHPIECE FOR IPPB MACHINE
CATEGORY 1II | RT13-01-01II | MOUTHPIECE FOR IPPB MACHINE ID: 15mm OD: 22mm Consists of a mouthpiece made of translucent molded-plastic . Must have a comfortable fit and be compatible with IPPB machines. Single use, individually packed and surgical clean | MOUTHPIECE FOR IPPB MACHINE ID: 15mm OD: 22mm
HUMIDIFIER PRE-FILLED WATER PACKS
CATEGORY 1JJ | RT13-01-01JJ | HUMIDIFIER PRE-FILLED WATER PACKSPurpose: A humidifier that supplies effective humidity for the inhalation of dry medical air via an oxygen mask or nasal cannula. Consists of a 325ml - 350ml sterile water prefilled humidifier with a wing-nut to fit a standard DISS connection on oxygen flow meter. The connector must not break when screwed onto the flowmeter The universal male adaptor/outlet port must connect securely to oxygen tubing. Must allow for a flow rate 0-15 L/min. The tubing must not pop off under high flows and the humidifier should not make a whistling noise when in use. Surgical clean and individually pack | HUMIDIFIER PRE-FILLED WATER PACKSPurpose: A humidifier that supplies effective humidity for the inhalation of dry medical air via an oxygen mask or nasal cannula.
ENDOTRACHEAL CUFF PRESSURE MANOMETER
CATEGORY 1KK | RT13-01-01KK | Endotracheal Cuff Pressure Manometer Purpose: To inflate and monitor the pressure of laryngeal masks and intubation tubes with low-pressure cuffs so as to reduce the risk of pressure necrosis and mucosal ischemia as well as preventing the risk of aspiration. Consists of a large dial with easy to read display graded in cmH2O The display must clearly indicate the ideal ranges for either tracheal tubes or laryngeal tubes. 22–32cmH2O for tracheal tubes 32–60cmH2O For Laryngeal Tubes and Laryngeal Masks A luer-connection must allow the gauge to be connected to laryngeal masks or intubation tubes, either directly or via the connecting tube. Must have a release valve to allow air release in the case of over-inflation The cuff must be inflated using a sensitive and easy to compress inflation bulb Must have a suspension hook for easy rail connection The manometer must be ergonomic, durable and easy to clean | Endotracheal Cuff Pressure Manometer

## Category 2
BID NUMBER | RT13
BID DESCRIPTION | SUPPLY AND DELIVERY OF RESPIRATORY AID AND ANAESTHETIC ACCESSORIES TO THE STATE
CATEGORY 2
ENDOTRACHEAL TUBES
| Item Number | Technical Specification | Product Name and Size | Insert inputs on this column
UNCUFFED ORAL/NASAL ENDOTRACHEAL TUBE FOR SHORT-TERM INTUBATION
CATEGORY 2A | RT13-02-01A | UNCUFFED ORAL/NASAL ENDOTRACHEAL TUBE FOR SHORT-TERM INTUBATION Purpose: Short term oral or nasal intubation Medically safe, thermosensitive firm and kink resistant, clear PVC with a full Magill curvature. Must have a radiopaque line, marked with numbered depth markings in 1cm intervals visible on the inner curvature The tube markings must include the size, length, ID, OD, trade name. Markings must be clear and must not rub off when wiped with alcohol. Must have a precision band mark to aid with depth of intubation and guide proper placement. Solid or circumferential precision band markings accepted Preference will be given to tubes with a second circumferential band (not more than 1cm from each other) to guide proper placement. Proximal end - Secure 15mm OD circuit connector made from medical grade plastic Distal end - Atraumatic left-facing bevelled tip with 1 or NO Murphy eye on the opposite side. All components latex-free, sterile and individually packed in peel pouch that is easy to open - For SINGLE use only To comply with the latest ISO 5361 Size 2.0 and 2.5 will be considered individually; Size 3.0 - 6.0 will be considered in series | Uncuffed Oral/Nasal endotracheal tube for short-term intubation Size 2.0
| RT13-02-02A |  | Uncuffed Oral/Nasal endotracheal tube for short-term intubation Size 2.5
| RT13-02-03A |  | Uncuffed Oral/Nasal endotracheal tube for short-term intubation Size 3.0
| RT13-02-04A |  | Uncuffed Oral/Nasal endotracheal tube for short-term intubation Size 3.5
| RT13-02-05A |  | Uncuffed Oral/Nasal endotracheal tube for short-term intubation Size 4.0
| RT13-02-06A |  | Uncuffed Oral/Nasal endotracheal tube for short-term intubation Size 4.5
| RT13-02-07A |  | Uncuffed Oral/Nasal endotracheal tube for short-term intubation Size 5.0
| RT13-02-08A |  | Uncuffed Oral/Nasal endotracheal tube for short-term intubation Size 5.5
| RT13-02-09A |  | Uncuffed Oral/Nasal endotracheal tube for short-term intubation Size 6.0
UNCUFFED SILICONIZED ORAL/NASAL ENDOTRACHEAL TUBE FOR LONG -TERM INTUBATION
CATEGORY 2B | RT13-02-01B | UNCUFFED SILICONIZED ORAL/NASAL ENDOTRACHEAL TUBE FOR LONG-TERM INTUBATION Purpose: Short term oral or nasal intubation Medically safe, thermosensitive, soft, pliable and kink resistant, SILICONIZED PVC with a full Magill curvature. Must have a radiopaque line, marked with numbered depth markings in 1cm intervals visible on the inner curvature Tube markings size, length, ID, OD, trade name must be clear and must not rub off when wiped with alcohol. Must have a precision band mark to aid with depth of intubation and guide proper placement. Solid or circumferential precision band markings accepted Proximal end - secure 15mm OD circuit connector made from medical grade plastic Distal end - Atraumatic left-facing bevelled tip with Murphy eye on the opposite side All components latex-free, sterile and individually packed in peel pouch that is easy to open - For SINGLE use only To comply with the latest ISO 5361 | Uncuffed Siliconized Oral/Nasal Endotracheal Tube for long-term intubation ID: 2.0mm OD: 2.9 ± 0.15mm Precision band marks at 20mm and 30mm from tip with depth markings in 1cm intervals starting at 4cm
| RT13-02-02B |  | Uncuffed Siliconized Oral/Nasal Endotracheal Tube for short-term intubation: ID: 2.5mm OD: 3.5 mm ± 0.15mm Precision band marks at 20mm and 30mm from tip with depth markings in 1cm intervals starting at 4cm
| RT13-02-03B |  | Uncuffed Siliconized Oral/Nasal Endotracheal Tube for long-term intubation ID: 3.0mm OD: 4.2mm ± 0.15mm Precision band marks at 25mm and 35mm from tip with depth markings in 1cm intervals starting at 4cm
| RT13-02-04B |  | Uncuffed Siliconized Oral/Nasal Endotracheal Tube for long-term intubation ID: 3.5mm OD: 4.9mm ± 0.15mm Precision band marks at 30mm and 40mm and 50mm from tip with depth markings in 1cm intervals starting at 6cm
| RT13-02-05B |  | Uncuffed Siliconized Oral/Nasal Endotracheal Tube for long-term intubation ID: 4.0mm OD: 5.5mm ± 0.15mm Precision band marks at 30mm and 40mm and 50mm from tip with depth markings in 1cm intervals starting at 6cm
| RT13-02-06B |  | Uncuffed Siliconized Oral/Nasal Endotracheal Tube for long-term intubation ID: 4.5mm OD: 6.2mm ± 0.15mm Precision band marks at 30mm and 40mm and 50mm from tip with depth markings in 1cm intervals starting at 7cm
| RT13-02-07B |  | Uncuffed Siliconized Oral/Nasal Endotracheal Tube for long-term intubation ID: 5.0mm OD: 6.8mm ± 0.15mm Precision band marks at 35mm and 45mm and 55mm from tip with depth markings in 1cm intervals starting at 7cm
| RT13-02-08B |  | Uncuffed Siliconized Oral/Nasal Endotracheal Tube for long-term intubation ID: 5.5mm OD: 7.5mm ± 0.15mm Precision band marks at 35mm and 45mm and 55mm from tip with depth markings in 1cm intervals starting at 7cm
| RT13-02-09B |  | Uncuffed Siliconized Oral/Nasal Endotracheal Tube for long-term intubation ID: 6.0mm OD: 8.5mm ± 0.20mm Precision band marks at 35mm and 45mm and 55mm from tip with depth markings in 1cm intervals starting at 7cm
UNCUFFED EXTRA-SOFT PLACTICIZED PVC ORAL/NASAL TRACHEAL TUBE FOR LONG TERM INTUBATION
CATEGORY 2C | RT13-02-01C | UNCUFFED EXTRA-SOFT PLASTICIZED PVC ORAL/NASAL TRACHEAL TUBE FOR LONG TERM INTUBATION Purpose: Long term oral or nasal intubation in paediatric patient's Consists of a firm and kink resistant extra-soft plasticized PVC with a full Magill curvature that softens at body temperature to conform to anatomy Must have a radiopaque line and must be marked with numbered depth markings in 50mm or 1cm intervals and the side of the tube with tube size, length, ID, OD, trade name. The markings on tube must be clear and must not rub off Must have a precision band mark to aid with depth of intubation and guide proper placement. Solid or circumferential precision band markings accepted and must be no more than 20mm from the tip Proximal end - secure 15mm OD circuit connector made from medical grade plastic Distal end - Atraumatic left-facing bevelled tip without Murphy eye All components must be latex and DEHP free, sterile and individually packed in peel pouch that is easy to open - For SINGLE use only To comply with the latest ISO 5361 Considered in series | Uncuffed Extra-soft Plasticized PVC Oral/Nasal tracheal tube for long-term intubation. ID: 2.0mm OD: 3.4mm Length: 165mm Depth markings in 0.5cm intervals, starting at 3cm.
| RT13-02-02C |  | Uncuffed Extra-soft Plasticized PVC Oral/Nasal tracheal tube for long-term intubation. ID: 2.5mm OD: 4.1mm Length: 165mm Depth markings in 0.5cm intervals, starting at 3cm.
| RT13-02-03C |  | Uncuffed Extra-soft Plasticized PVC Oral/Nasal tracheal tube for long-term intubation. ID: 3.0mm OD: 4.6mm Length: 165mm Depth markings in 0.5cm intervals, starting at 3cm.
| RT13-02-04C |  | Uncuffed Extra-soft Plasticized PVC Oral/Nasal tracheal tube for long-term intubation. ID: 3.5mm OD: 5.2mm Length: 165mm Depth markings in 0.5cm intervals starting at 3cm.
| RT13-02-05C |  | Uncuffed Extra-soft Plasticized PVC Oral/Nasal tracheal tube for long-term intubation. ID: 4.0mm OD: 5.7mm Length: 230mm Depth markings in 1cm intervals from 9 - 21cm.
| RT13-02-06C |  | Uncuffed Extra-soft Plasticized PVC Oral/Nasal tracheal tube for long-term intubation. ID: 4.5mm OD: 6.2mm Length: 230mm Depth markings in 1cm intervals from 9 - 21cm
| RT13-02-07C |  | Uncuffed Extra-soft Plasticized PVC Oral/Nasal tracheal tube for long-term intubation. ID: 5.0mm OD: 7.0mm Length: 230mm Depth markings in 1cm intervals from 13 - 21cm
| RT13-02-08C |  | Uncuffed Extra-soft Plasticized PVC Oral/Nasal tracheal tube for long-term intubation. ID: 5.5mm OD: 8.0mm Length: 270mm Depth markings in 1cm intervals from 15 - 25cm
| RT13-02-09 |  | Uncuffed Extra-soft Plasticized PVC Oral/Nasal tracheal tube for long-term intubation. ID: 6.0mm OD: 8.5mm Length: 270mm Depth markings in 1cm intervals from 15 - 25cm
CUFFED ORAL/NASAL TRACHEAL TUBE WITH/ WITHOUT A MURPHY EYE - NEWBORN
CATEGORY 2D | RT13-02-01D | CUFFED ORAL/NASAL TRACHEAL TUBE WITH/ WITHOUT A MURPHY EYE - NEWBORN Purpose: Short term oral or nasal intubation Consists of a clear thermo-sensitive PVC tube that is kink resistant, atraumatic to airways and has a full Magill curvature. Must be radiopaque or have a radiopaque line Must be marked with numbered depth markings in 1cm intervals, tube size, length, ID, OD, trade name. The markings on tube must be clear and must not rub off when wiped with alcohol. Must have a precision band mark just above the cuff to aid with depth of intubation and guide proper placement. Preference will be given to tubes with a second circumferential band (not more than 1cm from each other) to guide proper placement. Proximal end - secure 15mm OD circuit connector made from medical grade plastic Distal end - Atraumatic left-facing bevelled tip without or 1 Murphy eye on the opposite side Must have a high-volume-low-pressure barrel-shaped polyvinyl chloride cuff with a pilot cuff. Cuff thickness ± 50 microns The pilot cuff must have a one-way valve and must indicate the degree of inflation accurately. The cuff must inflate and deflate easily, and when deflated must conform closely to tube to minimise the risk of trauma during intubation, and extubation and not increase the OD of the tube. Must be appropriately sized in relation to the tube size and appropriately placed with regards to the distance from the distal tip. The cuff must be resistant to Nitrous Oxide effects All components must be latex-free sterile and individually packed in peel pouch that is easy to open; For SINGLE use only To comply with the latest ISO 5361 Size 2.0mm and 2.5mm to be awarded individually | Cuffed Oral/Nasal tracheal tube with a Murphy eye ID: 2.0 Age: Newborn (< 2kg)
| RT13-02-02D |  | Cuffed Oral/Nasal tracheal tube with a Murphy eye ID: 2.5 Age: Newborn (< 2.5kg)
CUFFED ORAL/NASAL TRACHEAL TUBE WITH MURPHY EYES - PAEDIATRIC
CATEGORY 2E | RT13-02-01E | CUFFED ORAL/NASAL TRACHEAL TUBE WITH A MURPHY EYE - PAEDIATRIC Purpose: Short term oral or nasal intubation Consists of a clear thermo-sensitive PVC tube that is kink resistant, atraumatic to airways and has a full Magill curvature. Must be radiopaque or have a radiopaque line Must be marked with numbered depth markings in 1cm intervals, tube size, length, ID, OD, trade name. The markings on tube must be clear and must not rub off when wiped with alcohol. Must be marked with numbered depth markings in 1cm intervals, tube size, length, ID, OD, trade name. The markings on tube must be clear and must not rub off when wiped with alcohol. Must have a precision band mark just above the cuff to aid with depth of intubation and guide proper placement. Preference will be given to tubes with a second circumferential band (not more than 1cm from each other) to guide proper placement. Proximal end - secure 15mm OD circuit connector made from medical grade plastic Distal end - Atraumatic left-facing bevelled tip with 1 or 2 Murphy eyes on the opposite side Must have a high-volume-low-pressure barrel-shaped polyvinyl chloride cuff with a pilot cuff. Cuff thickness ± 50 microns The pilot cuff must have a one-way valve and must indicate the degree of inflation accurately. The cuff must inflate and deflate easily, and when deflated must conform closely to tube to minimise the risk of trauma during intubation, and extubation and not increase the OD of the tube. Must be appropriately sized in relation to the tube size and appropriately placed with regards to the distance from the distal tip. The cuff must be resistant to Nitrous Oxide effects All components must be latex-free sterile and individually packed in peel pouch that is easy to open; For SINGLE use only To comply with the latest ISO 5361 Sizes 3mm - 6mm considered in series | Cuffed Oral/Nasal tracheal tube with a Murphy eye ID: 3.0mm Age: Neonates (≥3 kg) to <1yr
| RT13-02-02E |  | Cuffed Oral/Nasal tracheal tube with a Murphy eye ID: 3.5mm Age:1 to < 2yr
| RT13-02-03E |  | Cuffed Oral/Nasal tracheal tube with a Murphy eye ID: 4.0mm Age: 2 to < 4yr
| RT13-02-04E |  | Cuffed Oral/Nasal tracheal tube with a Murphy eye ID: 4.5mm Age: 2 to < 4yr
| RT13-02-05E |  | Cuffed Oral/Nasal tracheal tube with a Murphy eye ID: 5.0mm Age: 6 to < 8yr
| RT13-02-06E |  | Cuffed Oral/Nasal tracheal tube with a Murphy eye ID: 5.5mm Age: 8 to < 10yr
| RT13-02-07E |  | Cuffed Oral/Nasal tracheal tube with a Murphy eye ID: 6.0mm Age: > 10yr
CUFFED ORAL/NASAL TRACHEAL TUBE WITH A MURPHY EYES - ADULT
CATEGORY 2F | RT13-02-01F | CUFFED ORAL/NASAL TRACHEAL TUBE WITH A MURPHY EYES - ADULT Purpose: Short term oral or nasal intubation Consists of a clear thermo-sensitive PVC tube that is kink resistant, atraumatic to airways and has a full Magill curvature. Must be radiopaque or have a radiopaque line Must be marked with numbered depth markings in 1cm intervals, tube size, length, ID, OD, trade name. The markings on tube must be clear and must not rub off when wiped with alcohol. Must have a precision band mark just above the cuff to aid with depth of intubation and guide proper placement. Preference will be given to tubes with a second circumferential band (not more than 2cm from each other) to guide proper placement. Proximal end - secure 15mm OD circuit connector made from medical grade plastic Distal end - Atraumatic left-facing bevelled tip with 1 or 2 Murphy eyes on the opposite side Must have a high-volume-low-pressure barrel-shaped polyvinyl chloride cuff with a pilot cuff. Cuff thickness ± 50 microns The pilot cuff must have a one-way valve and must indicate the degree of inflation accurately. The cuff must inflate and deflate easily, and when deflated must conform closely to tube to minimise the risk of trauma during intubation, and extubation and not increase the OD of the tube. Must be appropriately sized in relation to the tube size and appropriately placed with regards to the distance from the distal tip. The cuff must be resistant to Nitrous Oxide effects All components must be latex-free sterile and individually packed in peel pouch that is easy to open; For SINGLE use only To comply with the latest ISO 5361 Sizes 6.5mm - 8mm considered in series; Size 8.5mm and 9.0mm to be awarded individually | Cuffed Oral/Nasal tracheal tube with a Murphy eye Adult ID: 6.5mm
| RT13-02-01F |  | Cuffed Oral/Nasal tracheal tube with a Murphy eye Adult ID: 7.0mm
| RT13-02-01F |  | Cuffed Oral/Nasal tracheal tube with a Murphy eye Adult ID: 7.5mm
| RT13-02-01F |  | Cuffed Oral/Nasal tracheal tube with a Murphy eye Adult ID: 8.0mm
| RT13-02-01F |  | Cuffed Oral/Nasal tracheal tube with a Murphy eye Adult ID: 8.5mm
| RT13-02-01F |  | Cuffed Oral/Nasal tracheal tube with a Murphy eye Adult ID: 9.0mm
CUFFED SILICONIZED ENDOTRACHEAL TUBE WITH A MURPHY EYE - PAEDIATRIC
CATEGORY 2G | RT13-02-01G | CUFFED SILICONIZED ENDOTRACHEAL TUBE WITH A MURPHY EYE - PAEDIATRIC Purpose: Short term oral or nasal intubation Medically safe, thermosensitive, soft, pliable and kink resistant, SILICONIZED PVC tube with a radiopaque line and a full Magill curvature. Must be radiopaque or have a radiopaque line Must be marked with numbered depth markings in 1cm intervals, tube size, length, ID, OD, trade name. The markings on tube must be clear and must not rub off when wiped with alcohol. Must have a precision band mark just above the cuff to aid with depth of intubation and guide proper placement. Preference will be given to tubes with a second circumferential band (not more than 1cm from each other) to guide proper placement. Proximal end - secure 15mm OD circuit connector made from medical grade plastic Distal end - Atraumatic left-facing bevelled tip with 1 or 2 Murphy eyes on the opposite side Must have a high-volume-low-pressure barrel-shaped polyvinyl chloride cuff with a pilot cuff. Cuff thickness ± 50 microns The pilot cuff must have a one-way valve and must indicate the degree of inflation accurately. The cuff must inflate and deflate easily, and when deflated must conform closely to tube to minimise the risk of trauma during intubation, and extubation and not increase the OD of the tube. Must be appropriately sized in relation to the tube size and appropriately placed with regards to the distance from the distal tip. The cuff must be resistant to Nitrous Oxide effects All components must be latex-free; DEHP free; sterile and individually packed in peel pouch that is easy to open; For SINGLE use only To comply with the latest ISO 5361 Size 3.0mm - 6.0mm considered in series | Cuffed Siliconized Oral/Nasal endotracheal tube with Murphy Eye: Size ID: 3.0mm
| RT13-02-02G |  | Cuffed Siliconized Oral/Nasal endotracheal tube with Murphy Eye: Size ID: 3.5mm
| RT13-02-03G |  | Cuffed Siliconized Oral/Nasal endotracheal tube with Murphy Eye: Size ID: 4.0mm
| RT13-02-04G |  | Cuffed Siliconized Oral/Nasal endotracheal tube with Murphy Eye: Size ID: 4.5mm
| RT13-02-05G |  | Cuffed Siliconized Oral/Nasal endotracheal tube with Murphy Eye: Size ID: 5.0mm
| RT13-02-06G |  | Cuffed Siliconized Oral/Nasal endotracheal tube with Murphy Eye: Size ID: 5.5mm
| RT13-02-07G |  | Cuffed Siliconized Oral/Nasal endotracheal tube with Murphy Eye: Size ID: 6.0mm
CUFFED SILICONIZED ENDOTRACHEAL TUBE WITH A MURPHY EYE - ADULT
CATEGORY 2H | RT13-02-01H | CUFFED SILICONIZED ENDOTRACHEAL TUBE WITH A MURPHY EYE - ADULT Purpose: Short term oral or nasal intubation Medically safe, thermosensitive, soft, pliable and kink resistant, SILICONIZED PVC tube with a radiopaque line and a full Magill curvature. Must be radiopaque or have a radiopaque line Must be marked with numbered depth markings in 1cm intervals, tube size, length, ID, OD, trade name. The markings on tube must be clear and must not rub off when wiped with alcohol. Must have a precision band mark just above the cuff to aid with depth of intubation and guide proper placement. Preference will be given to tubes with a second circumferential band (not more than 2cm from each other) to guide proper placement. Proximal end - secure 15mm OD circuit connector made from medical grade plastic Distal end - Atraumatic left-facing bevelled tip with 1 or 2 Murphy eyes on the opposite side Must have a high-volume-low-pressure barrel-shaped polyvinyl chloride cuff with a pilot cuff. Cuff thickness ± 50 microns The pilot cuff must have a one-way valve and must indicate the degree of inflation accurately. The cuff must inflate and deflate easily, and when deflated must conform closely to tube to minimise the risk of trauma during intubation, and extubation and not increase the OD of the tube. Must be appropriately sized in relation to the tube size and appropriately placed with regards to the distance from the distal tip. The cuff must be resistant to Nitrous Oxide effects All components must be latex-free; DEHP free; sterile and individually packed in peel pouch that is easy to open; For SINGLE use only To comply with the latest ISO 5361 Sizes 6.5mm - 8mm considered in series; Size 8.5mm and 9.0mm to be awarded individually | Cuffed Siliconized Oral/Nasal endotracheal tube with a Murphy Eye Size ID: 6.5mm
| RT13-02-02H |  | Cuffed Siliconized Oral/Nasal endotracheal tube with a Murphy Eye Size ID: 7.0mm
| RT13-02-03H |  | Cuffed Siliconized Oral/Nasal endotracheal tube with a Murphy Eye Size ID: 7.5mm
| RT13-02-04H |  | Cuffed Siliconized Oral/Nasal endotracheal tube with a Murphy Eye Size ID: 8.0mm
| RT13-02-05H |  | Cuffed Siliconized Oral/Nasal endotracheal tube with a Murphy Eye Size ID: 8.5mm
| RT13-02-06H |  | Cuffed Siliconized Oral/Nasal endotracheal tube with a Murphy Eye Size ID: 9.0mm
SHORT-CUFFED ENDOTRACHEAL TUBE WITH SUBGLOTTIC PRESSURE-FREE ZONE, MAGILL TIP WITHOUT A MURPHY EYE
CATEGORY 2I | RT13-02-01I | SHORT-CUFFED ENDOTRACHEAL TUBE WITH SUBGLOTTIC PRESSURE-FREE ZONE, MAGILL TIP WITHOUT A MURPHY EYE Purpose: Short term oral or nasal intubation without a murphy’s eye Consists of a clear, kink-resistant, thermo-sensitive PVC tube that is atraumatic to airways and has a full Magill curvature without Murphy eye. Must be radiopaque or have a radiopaque line Must be marked with numbered depth markings in 1cm intervals, tube size, length, ID, OD, trade name. The markings on tube must be clear and must not rub off when wiped with alcohol. Must have a precision band mark just above the cuff to aid with depth of intubation and guide proper placement. Preference will be given to tubes with further markings to guide proper placement. Proximal end - secure 15mm OD circuit connector made from medical grade plastic Distal end - short atraumatic tip distal to a shortened cuff without Murphy eye The high-volume-low-pressure polyvinyl chloride cuff with pilot cuff must be short and cylindrical and made of micro-thin (10µ) polyurethane Must be placed close to tip of tube and allow for a subglottic pressure-free zone. The cuff must inflate and deflate easily, and when deflated must conform closely to tube to minimise the risk of trauma during intubation, and extubation and not increase the OD of the tube. The cuff must be resistant to the effects of Nitrous Oxide The pilot cuff must have a one-way valve and must indicate the degree of inflation accurately. All components must be latex-free, sterile and individually packed in peel pouch that is easy to open - For SINGLE use only To comply with the latest ISO 5361 Considered in series | Short-cuffed Endotracheal Tube with Subglottic Pressure-Free Zone, Magill Tip without a Murphy Eye ID: 3.0mm OD: 4.3 - 5.0mm Length: 160 - 190mm Age: Newborns (≥3 kg) to <1yr
| RT13-02-02I |  | Short-cuffed Endotracheal Tube with Subglottic Pressure-Free Zone, Magill Tip without a Murphy Eye ID: 3.5mm OD: 4.3 - 5.3mm Length: 190 - 220mm Age:1 to <2yr
| RT13-02-03I |  | Short-cuffed Endotracheal Tube with Subglottic Pressure-Free Zone, Magill Tip without a Murphy Eye ID: 4mm OD: 5.6 - 6.0mm Length: 220 - 240mm Age: 2 to <4yr
| RT13-02-04I |  | Short-cuffed Endotracheal Tube with Subglottic Pressure-Free Zone, Magill Tip without a Murphy Eye ID: 4.5mm OD: 6.2 - 6.3mm Length: 230 - 260mm Age: 4 to <6yr
| RT13-02-05I |  | Short-cuffed Endotracheal Tube with Subglottic Pressure-Free Zone, Magill Tip without a Murphy Eye ID: 5mm OD: 6.7 - 6.9mm Age: 6 to < 8yr
| RT13-02-06I |  | Short-cuffed Endotracheal Tube with Subglottic Pressure-Free Zone, Magill Tip without a Murphy Eye ID: 5.5mm OD: 7.3 - 7.5mm Age: 8 to <10yr
| RT13-02-07I |  | Short-cuffed Endotracheal Tube with Subglottic Pressure-Free Zone, Magill Tip without a Murphy Eye ID: 6.0 mm Age: 10 to <12yr
SHORT-CUFFED ORAL SOUTH-FACING PREFORMED (RAE) ENDOTRACHEAL TUBE WITH A SUBGLOTTIC PRESSURE-FREE ZONE, MAGILL TIP WITHOUT A MURPHY EYE
CATEGORY 2J | RT13-02-01J | SHORT-CUFFED ORAL SOUTH-FACING PREFORMED (RAE) ENDOTRACHEAL TUBE WITH A SUBGLOTTIC PRESSURE-FREE ZONE, MAGILL TIP WITHOUT A MURPHY EYE Purpose: Short term oral or nasal intubation without a murphy’s eye Consists of a clear, kink-resistant, thermo-sensitive PVC tube with a south-facing preformed curve without Murphy eye. Must be radiopaque or have a radiopaque line Must be atraumatic to airways and able to temporarily straightened to allow the passage of a suction catheter. Must be marked with numbered depth markings in 1cm intervals, tube size, length, ID, OD, trade name. The markings on tube must be clear and must not rub off when wiped with alcohol. Must have a precision band mark just above the cuff to aid with depth of intubation and guide proper placement. Preference will be given to tubes with further markings to guide proper placement. Must also have a distinct marker at the tube bend Proximal end - secure 15mm OD circuit connector made from medical grade plastic Distal end - short atraumatic tip distal to a shortened cuff without Murphy eye The high-volume-low-pressure polyvinyl chloride cuff with pilot cuff must be short and cylindrical and made of micro-thin (10µ) polyurethane Must be placed close to tip of tube and allow for a subglottic pressure-free zone. The cuff must inflate and deflate easily, and when deflated must conform closely to tube to minimise the risk of trauma during intubation, and extubation and not increase the OD of the tube. The cuff must be resistant to the effects of Nitrous Oxide The pilot cuff must have a one-way valve and must indicate the degree of inflation accurately. All components must be latex-free, sterile and individually packed in peel pouch that is easy to open - For SINGLE use only To comply with the latest ISO 5361 Considered in series | Short-Cuffed Oral South-Facing Preformed (RAE) Endotracheal Tube with Subglottic Pressure-Free Zone, Magill Tip without a Murphy Eye ID: 3.0 Age: Newborns (≥3 kg) to <1yr
| RT13-02-02J |  | Short-Cuffed Oral South-Facing Preformed (RAE) Endotracheal Tube with Subglottic Pressure-Free Zone, Magill Tip without a Murphy Eye ID: 3.5mm Age:1 to <2yr
| RT13-02-03J |  | Short-Cuffed Oral South-Facing Preformed (RAE) Endotracheal Tube with Subglottic Pressure-Free Zone, Magill Tip without a Murphy Eye ID: 4mm Age: 2 to <4yr
| RT13-02-04J |  | Short-Cuffed Oral South-Facing Preformed (RAE) Endotracheal Tube with Subglottic Pressure-Free Zone, Magill Tip without a Murphy Eye ID: 4.5mm Age: 4 to <6yr
| RT13-02-05J |  | Short-Cuffed Oral South-Facing Preformed (RAE) Endotracheal Tube with Subglottic Pressure-Free Zone, Magill Tip without a Murphy Eye ID: 5mm Age: 6 to < 8yr
| RT13-02-06J |  | Short-Cuffed Oral South-Facing Preformed (RAE) Endotracheal Tube with Subglottic Pressure-Free Zone, Magill Tip without a Murphy Eye ID: 5.5mm Age: 8 to <10yr
UNCUFFED ORAL/NASAL TRACHEAL TUBE WITH REDUCED TRAUMA TIP
CATEGORY 2K | RT13-02-01K | UNCUFFED ORAL/NASAL TRACHEAL TUBE WITH REDUCED TRAUMA TIP Purpose: Short term oral or nasal intubation with tip designed to reduce airway trauma Consists of a clear thermo-sensitive PVC tube that is kink resistant, atraumatic to airways and has a full Magill curvature and two Murphy eyes. Must be radiopaque or have a radiopaque line Must be marked with numbered depth markings in 1cm intervals, tube size, length, ID, OD, trade name. The markings on tube must be clear and must not rub off when wiped with alcohol. Must have a precision band mark to aid with depth of intubation and guide proper placement. Solid or circumferential precision band markings accepted Preference will be given to tubes with a second circumferential band (not more than 1cm from each other) to guide proper placement. Proximal end - secure 15mm OD circuit connector made from medical grade plastic Distal end - Tapered, soft flexible pointed tip with an upturned leading edge , configured with the tip point in the midline so as to pass easily between the vocal cords without touching them. The tube must be able to demonstrate a reduction in intubation induced airway trauma. All components must be latex-free sterile and individually packed in peel pouch that is easy to open - For SINGLE use only To comply with the latest ISO 5361 Considered in series | Uncuffed Oral/Nasal tracheal tube with Reduced Trauma Tip: Size 2.5
| RT13-02-02K |  | Uncuffed Oral/Nasal tracheal tube with Reduced Trauma Tip: Size 3.0
| RT13-02-03K |  | Uncuffed Oral/Nasal tracheal tube with Reduced Trauma Tip: Size 3.5
| RT13-02-04K |  | Uncuffed Oral/Nasal tracheal tube with Reduced Trauma Tip: Size 4.0
| RT13-02-05K |  | Uncuffed Oral/Nasal tracheal tube with Reduced Trauma Tip: Size 4.5
| RT13-02-06K |  | Uncuffed Oral/Nasal tracheal tube with Reduced Trauma Tip: Size 5.0
| RT13-02-07K |  | Uncuffed Oral/Nasal tracheal tube with Reduced Trauma Tip: Size 5.5
| RT13-02-08K |  | Uncuffed Oral/Nasal tracheal tube with Reduced Trauma Tip: Size 6.0
CUFFED ORAL/NASAL ENDOTRACHEAL TUBE WITH REDUCED TRAUMA TIP - PAEDIATRIC
CATEGORY 2L | RT13-02-01L | CUFFED ORAL/NASAL ENDOTRACHEAL TUBE WITH REDUCED TRAUMA TIP - PAEDIATRIC Purpose: Short term oral or nasal intubation with tip designed to reduce airway trauma Consists of a clear thermo-sensitive PVC tube that is kink resistant, atraumatic to airways and has a full Magill curvature and two Murphy eyes Must be radiopaque or have a radiopaque line Must be marked with numbered depth markings in 1cm intervals, tube size, length, ID, OD, trade name. The markings on tube must be clear and must not rub off when wiped with alcohol. Must have a precision band mark just above the cuff to aid with depth of intubation and guide proper placement. Preference will be given to tubes with a second circumferential band (not more than 1cm from each other) to guide proper placement. Proximal end - secure 15mm OD circuit connector made from medical grade plastic Distal end - Tapered, soft flexible pointed tip with an upturned leading edge , configured with the tip point in the midline so as to pass easily between the vocal cords without touching them. The tube must be able to demonstrate a reduction in intubation induced airway trauma. Must have a high-volume-low-pressure polyurethane barrel shaped cuff with pilot cuff. Cuff thickness ± 50 microns a shore hardness of no more than 86 microns The pilot cuff must have a one-way valve and must indicate the degree of inflation accurately. The cuff must inflate and deflate easily and when deflated must conform closely to tube to minimise the risk of trauma during intubation and extubation and not increase the OD of the tube. Cuff must be barrel shaped to facilitate optimal pressure dispersion against the tracheal wall. The cuff must be resistant to Nitrous Oxide effects All components must be latex-free sterile and individually packed in peel pouch that is easy to open - For SINGLE use only To comply with the latest ISO 5361 Size 3.0mm - 6.0mm considered in series | Cuffed Oral/Nasal endotracheal tube with Reduced Trauma Tip: ID: 3.0mm
| RT13-02-02L |  | Cuffed Oral/Nasal endotracheal tube with Reduced Trauma Tip: ID: 3.5mm
| RT13-02-03L |  | Cuffed Oral/Nasal endotracheal tube with Reduced Trauma Tip: ID: 4.0mm
| RT13-02-04L |  | Cuffed Oral/Nasal endotracheal tube with Reduced Trauma Tip: ID: 4.5mm
| RT13-02-05L |  | Cuffed Oral/Nasal endotracheal tube with Reduced Trauma Tip: ID: 5.0mm
| RT13-02-06L |  | Cuffed Oral/Nasal endotracheal tube with Reduced Trauma Tip: ID: 5.5mm
| RT13-02-07L |  | Cuffed Oral/Nasal endotracheal tube with Reduced Trauma Tip: ID: 6.0mm
CUFFED ORAL/NASAL ENDOTRACHEAL TUBE WITH REDUCED TRAUMA TIP - ADULT
CATEGORY 2M | RT13-02-01M | CUFFED ORAL/NASAL ENDOTRACHEAL TUBE WITH REDUCED TRAUMA TIP - ADULT Purpose: Short term oral or nasal intubation with tip designed to reduce airway trauma Consists of a clear thermo-sensitive PVC tube that is kink resistant, atraumatic to airways and has a full Magill curvature and two Murphy eyes Must be radiopaque or have a radiopaque line Must be marked with numbered depth markings in 1cm intervals, tube size, length, ID, OD, trade name. The markings on tube must be clear and must not rub off when wiped with alcohol. Must have a precision band mark just above the cuff to aid with depth of intubation and guide proper placement. Preference will be given to tubes with a second circumferential band (not more than 2cm from each other) to guide proper placement. Proximal end - secure 15mm OD circuit connector made from medical grade plastic Distal end - Tapered, soft flexible pointed tip with an upturned leading edge , configured with the tip point in the midline so as to pass easily between the vocal cords without touching them. The tube must be able to demonstrate a reduction in intubation induced airway trauma. Must have a high-volume-low-pressure polyurethane barrel shaped cuff with pilot cuff. Cuff thickness ± 50 microns a shore hardness of no more than 86 microns The pilot cuff must have a one-way valve and must indicate the degree of inflation accurately. The cuff must inflate and deflate easily and when deflated must conform closely to tube to minimise the risk of trauma during intubation and extubation and not increase the OD of the tube. Cuff must be barrel shaped to facilitate optimal pressure dispersion against the tracheal wall. The cuff must be resistant to Nitrous Oxide effects All components must be latex-free sterile and individually packed in peel pouch that is easy to open - For SINGLE use only To comply with the latest ISO 5361 Size 6.5mm - 8.5mm considered in series | Cuffed Oral/Nasal endotracheal tube with Reduced Trauma Tip: ID: 6.5mm
| RT13-02-02M |  | Cuffed Oral/Nasal endotracheal tube with Reduced Trauma Tip: ID: 7.0mm
| RT13-02-03M |  | Cuffed Oral/Nasal endotracheal tube with Reduced Trauma Tip: ID: 7.5mm
| RT13-02-04M |  | Cuffed Oral/Nasal endotracheal tube with Reduced Trauma Tip: ID: 8.0mm
| RT13-02-05M |  | Cuffed Oral/Nasal endotracheal tube with Reduced Trauma Tip: ID: 8.5mm
UNCUFFED ORAL SOUTH-FACING PREFORMED (RAE) ENDOTRACHEAL TUBE WITH REDUCED TRAUMA TIP
CATEGORY 2N | RT13-02-01N | UNCUFFED ORAL SOUTH-FACING PREFORMED (RAE) ENDOTRACHEAL TUBE WITH REDUCED TRAUMA TIP Purpose: Short term oral intubation where tube is South facing and removed from surgical field Consists of a clear thermo-sensitive PVC tube that is kink resistant, with a South polar preformed curve with a soft flexible pointed tip and two Murphy eyes Must be able to temporarily straightened the tube to allow the passage of a suction catheter Must have a radiopaque line and must be marked with numbered depth markings in 1cm intervals, tube size, length, ID, OD, trade name. The markings on tube must be clear and must not rub off. Must also have marks/reference lines proximal to cuff to aid with depth of intubation and guide proper placement. Preference will be given to tubes with a second circumferential band (not more than 1cm from each other) to guide proper placement. Must also have a distinct mark at preformed curve to aid positioning Proximal end - secure 15mm OD circuit connector made from medical grade plastic Distal end - soft flexible pointed tip with an upturned leading edge , configured with the tip point in the midline so as to pass easily between the vocal cords without touching them. The tube must be able to demonstrate a reduction in intubation induced airway trauma. All components must be latex-free sterile and individually packed in peel pouch that is easy to open - For SINGLE use only To comply with the latest ISO 5361 Size 3.0mm - 6.5mm considered in series | Cuffed Oral South-Facing Preformed (RAE) Endotracheal Tube with Reduced Trauma Tip. ID: 3.0mm OD 4.0mm Length ± 105mm
| RT13-02-02N |  | Cuffed Oral South-Facing Preformed (RAE) Endotracheal Tube with Reduced Trauma Tip. ID: 3.0mm OD 4.0mm Length ± 105mm
| RT13-02-03N |  | Cuffed Oral South-Facing Preformed (RAE) Endotracheal Tube with Reduced Trauma Tip. ID: 4.0mm OD 5.3mm Length ± 130mm
| RT13-02-04N |  | Cuffed Oral South-Facing Preformed (RAE) Endotracheal Tube with Reduced Trauma Tip ID: 4.5mm OD 6.0mm Length ± 145mm
| RT13-02-05N |  | Cuffed Oral South-Facing Preformed (RAE) Endotracheal Tube with Reduced Trauma Tip. ID: 5.0mm OD 6.7mm Length ± 155mm
| RT13-02-06N |  | Cuffed Oral South-Facing Preformed (RAE) Endotracheal Tube with Reduced Trauma Tip ID: 5.5mm OD 7.3mm Length ± 165mm
| RT13-02-07N |  | Cuffed Oral South-Facing Preformed (RAE) Endotracheal Tube with Reduced Trauma Tip ID: 6.0mm OD 8.0mm Length ± 180mm
| RT13-02-08N |  | Cuffed Oral South-Facing Preformed (RAE) Endotracheal Tube with Reduced Trauma Tip ID: 6.5mm OD 8.2mm Length ± 180mm
CUFFED ORAL SOUTH-FACING PREFORMED (RAE) ENDOTRACHEAL TUBE WITH REDUCED TRAUMA TIP
CATEGORY 2O | RT13-02-01O | UNCUFFED NASAL NORTH-FACING PREFORMED (RAE) ENDOTRACHEAL TUBE WITH REDUCED TRAUMA TIP Purpose: Short term oral intubation where tube is North facing and removed from surgical field Consists of a clear thermo-sensitive PVC tube that is kink resistant, with a North polar preformed curve with a soft flexible pointed tip and two Murphy eyes Must be able to temporarily straightened the tube to allow the passage of a suction catheter Must have a radiopaque line and must be marked with numbered depth markings in 1cm intervals, tube size, length, ID, OD, trade name. The markings on tube must be clear and must not rub off. Must also have marks/reference lines proximal to cuff to aid with depth of intubation and guide proper placement. Preference will be given to tubes with a second circumferential band (not more than 1cm from each other) to guide proper placement. Must also have a distinct mark at preformed curve to aid positioning Proximal end - secure 15mm OD circuit connector made from medical grade plastic Distal end - Tapered soft flexible pointed tip with an upturned leading edge , configured with the tip point in the midline so as to pass easily between the vocal cords without touching them. The tube must be able to demonstrate a reduction in intubation induced airway trauma. All components must be latex-free sterile and individually packed in peel pouch that is easy to open - For SINGLE use only To comply with the latest ISO 5361 Size 3.0mm - 6.5mm considered in series | Uncuffed Nasal North-Facing Preformed (RAE) Endotracheal Tube with Reduced Trauma Tip ID: 3.0mm OD 4.0mm Length at nasal curve ± 155mm
| RT13-02-02O |  | Uncuffed Nasal North-Facing Preformed (RAE) Endotracheal Tube with Reduced Trauma Tip ID: 3.5mm OD 5.3mm Length at nasal curve ± 160mm
| RT13-02-03O |  | Uncuffed Nasal North-Facing Preformed (RAE) Endotracheal Tube with Reduced Trauma Tip ID: 4.0mm OD 6.0mm Length at nasal curve ± 185mm
| RT13-02-04O |  | Uncuffed Nasal North-Facing Preformed (RAE) Endotracheal Tube with Reduced Trauma Tip ID: 4.5mm OD 6.3mm Length at nasal curve ± 220mm
| RT13-02-05O |  | Uncuffed Nasal North-Facing Preformed (RAE) Endotracheal Tube with Reduced Trauma Tip ID: 5.0mm OD 6.7mm Length at nasal curve ± 235mm
| RT13-02-06O |  | Uncuffed Nasal North-Facing Preformed (RAE) Endotracheal Tube with Reduced Trauma Tip ID: 5.5mm OD 7.3mm Length at nasal curve ± 245mm
| RT13-02-07O |  | Uncuffed Nasal North-Facing Preformed (RAE) Endotracheal Tube with Reduced Trauma Tip ID: 6.0mm OD 8.0mm Length at nasal curve ± 260mm
| RT13-02-08O |  | Uncuffed Nasal North-Facing Preformed (RAE) Endotracheal Tube with Reduced Trauma Tip ID: 6.5mm OD 8.7mm Length at nasal curve ± 270mm
UNCUFFED NASAL NORTH-FACING PREFORMED (RAE) ENDOTRACHEAL TUBE WITH REDUCED TRAUMA TIP
CATEGORY 2P | RT13-02-01P | UNCUFFED REINFORCED ORAL/NASAL ENDOTRACHEAL TUBE Purpose: Oral or nasal intubation where tube patency and flexibility is essential Consists of a soft flexible PVC tube with spiral-wound medical grade reinforcing wire that is sealed tightly against the bonded 15mm connector. Must be kink resistant and atraumatic to airways. Must be radiopaque or have a radiopaque line Must be marked with numbered depth markings in 1cm intervals, tube size, length, ID, OD, trade name. The markings on tube must be clear and must not rub off when wiped with alcohol. Must have a precision band mark to aid with depth of intubation and guide proper placement. Solid or circumferential precision band markings accepted Preference will be given to tubes with a second circumferential band (not more than 1cm from each other) to guide proper placement. Proximal end - secure 15mm OD circuit connector made from medical grade plastic Distal end - Atraumatic left-facing bevelled tip with Murphy eye on the opposite side All components must be latex-free, sterile and individually packed in peel pouch that is easy to open - For SINGLE use only To comply with the latest ISO 5361 Size 3mm - 6mm considered in series | Uncuffed Reinforced Oral/Nasal endotracheal tube: Size ID: 3.0mm OD 4.8mm Length ±184mm
| RT13-02-02P |  | Uncuffed Reinforced Oral/Nasal endotracheal tube: Size ID: 3.5mm OD 5.5mm Length ±185mm
| RT13-02-03P |  | Uncuffed Reinforced Oral/Nasal endotracheal tube: Size ID: 4.0mm OD 6.0mm Length ±195mm
| RT13-02-04P |  | Uncuffed Reinforced Oral/Nasal endotracheal tube: Size ID: 4.5mm OD 6.5mm Length ±198mm
| RT13-02-05P |  | Uncuffed Reinforced Oral/Nasal endotracheal tube: Size ID: 5.0mm OD 7.2mm Length ±208mm
| RT13-02-06P |  | Uncuffed Reinforced Oral/Nasal endotracheal tube: Size ID: 5.5mm OD 7.7mm Length ±208mm
| RT13-02-07P |  | Uncuffed Reinforced Oral/Nasal endotracheal tube: Size ID: 6.0mm OD 8.4mm Length ± 329mm
CUFFED REINFORCED ORAL/NASAL ENDOTRACHEAL TUBE PAEDIATRIC Size 3mm - 6mm considered in series
CATEGORY 2Q | RT13-02-01Q | CUFFED REINFORCED ORAL/NASAL ENDOTRACHEAL TUBE - PAEDIATRIC Purpose: Oral or nasal intubation where tube patency and flexibility is essential Consists of a soft flexible PVC tube with spiral-wound medical grade reinforcing wire that is sealed tightly against the bonded 15mm connector. Must be kink resistant and atraumatic to airways. Must be radiopaque or have a radiopaque line Must be marked with numbered depth markings in 1cm intervals, tube size, length, ID, OD, trade name. The markings on tube must be clear and must not rub off when wiped with alcohol. Must have a precision band mark to aid with depth of intubation and guide proper placement. Solid or circumferential precision band markings accepted Preference will be given to tubes with a second circumferential band (not more than 1cm from each other) to guide proper placement. Proximal end - secure 15mm OD circuit connector made from medical grade plastic Distal end - Atraumatic left-facing bevelled tip with Murphy eye on the opposite side Must have a high-volume-low-pressure polyurethane barrel shaped cuff with pilot cuff. Cuff thickness ± 50 microns. The pilot cuff must have a one-way valve and must indicate the degree of inflation accurately. The cuff must inflate and deflate easily and when deflated must conform closely to tube to minimise the risk of trauma during intubation and extubation and not increase the OD of the tube. The cuff must be resistant to Nitrous Oxide effects All components must be latex-free, sterile and individually packed in peel pouch that is easy to open - For SINGLE use only To comply with the latest ISO 5361 Size 3mm - 6mm considered in series | Cuffed Reinforced Oral/Nasal endotracheal tube: ID: 3.0mm OD 4.8mm Length ±184mm
| RT13-02-02Q |  | Cuffed Reinforced Oral/Nasal endotracheal tube: ID: 3.5mm OD 5.5mm Length ±185mm
| RT13-02-03Q |  | Cuffed Reinforced Oral/Nasal endotracheal tube: ID: 4.0mm OD 6.0mm Length ±195mm
| RT13-02-04Q |  | Cuffed Reinforced Oral/Nasal endotracheal tube: ID: 4.5mm OD 6.5mm Length ±198mm
| RT13-02-05Q |  | Cuffed Reinforced Oral/Nasal endotracheal tube: ID: 5.0mm OD 7.3mm Length ± 321mm
| RT13-02-06Q |  | Cuffed Reinforced Oral/Nasal endotracheal tube: ID: 5.5mm OD 7.8mm Length ± 322mm
| RT13-02-07Q |  | Cuffed Reinforced Oral/Nasal endotracheal tube: ID: 6.0mm OD 8.4mm Length ± 329mm
CUFFED REINFORCED ORAL/NASAL ENDOTRACHEAL TUBE - ADULT Size 6.5mm - 8mm considered in series; Size 8.5mm and 9.0mm to be awarded individually
CATEGORY 2R | RT13-02-01R | CUFFED REINFORCED ORAL/NASAL ENDOTRACHEAL TUBE - Adult Purpose: Oral or nasal intubation where tube patency and flexibility is essential Consists of a soft flexible PVC tube with spiral-wound medical grade reinforcing wire that is sealed tightly against the bonded connector. Must be kink resistant and atraumatic to airways Must have a radiopaque line and must be marked with numbered depth markings in 1cm intervals, tube size, length, ID, OD, trade name. The markings on the tube must be clear and must not rub off. Must have a precision band mark to aid with depth of intubation and guide proper placement. Solid or circumferential precision band markings accepted Preference will be given to tubes with a second circumferential band (not more than 2cm from each other) to guide proper placement. Proximal end - secure 15mm OD: circuit connector made from medical grade plastic Distal end - Atraumatic left-facing bevelled tip with Murphy eye on the opposite side The high-volume-low-pressure barrel-shaped polyvinyl chloride cuff with pilot cuff . The pilot cuff must have a one-way valve and must indicate the degree of inflation accurately. The cuff must inflate and deflate easily, and when deflated must conform closely to tube to minimise the risk of trauma during intubation, and extubation and not increase the OD of the tube. Must be appropriately sized in relation to the tube size and appropriately placed with regards to the distance from the distal tip. The cuff must be resistant to Nitrous Oxide effects Cuff thickness ± 50 microns All components must be latex-free sterile and individually packed in peel pouch that is easy to open; For SINGLE use only To comply with the latest ISO 5361 | Cuffed Reinforced Oral/Nasal endotracheal tube: ID: 6.5mm OD: 9.0mm Length ± 334mm
| RT13-02-02R |  | Cuffed Reinforced Oral/Nasal endotracheal tube: ID: 7.0mm OD: 9.6mm Length ± 352mm
| RT13-02-03R |  | Cuffed Reinforced Oral/Nasal endotracheal tube: ID: 7.5mm OD: 10.2mm Length ± 356mm
| RT13-02-04R |  | Cuffed Reinforced Oral/Nasal endotracheal tube: ID: 8.0mm OD: 10.8mm Length ± 364mm
| RT13-02-05R |  | Cuffed Reinforced Oral/Nasal endotracheal tube: ID: 8.5mm OD: 11.3mm Length ± 388mm
| RT13-02-06R |  | Cuffed Reinforced Oral/Nasal endotracheal tube: ID: 9.0mm OD: 11.9mm Length ± 398mm
TAPER-CUFFED ENDOTRACHEAL TUBE
CATEGORY 2S | RT13-02-01S | TAPER-CUFFED ENDOTRACHEAL TUBE Purpose: endotracheal tube with a tapered cuff for long stay oral/nasal intubation aimed to reduce micro-aspiration Consists of a clear thermo-sensitive PVC tube that is kink resistant, atraumatic to airways and have a full Magill curvature. Must have a radiopaque line and must be marked with numbered depth markings in 1cm intervals, tube size, length, ID, OD, trade name. The markings on tube must be clear and must not rub off. Must have a precision band mark just above the cuff to aid with depth of intubation and guide proper placement. Preference will be given to tubes with a second circumferential band (not more than 2cm from each other) to guide proper placement. Proximal end - secure 15mm OD circuit connector made from medical grade plastic Distal end - Atraumatic left-facing bevelled tip with Murphy eye on the opposite side Must have a high-volume-low-pressure polyurethane taper-shaped cuff with pilot cuff. Cuff thickness ± 50 microns The cuff must inflate and deflate easily and when deflated must conform closely to tube to minimise the risk of trauma during intubation and extubation. Cuff must be barrel shaped to facilitate optimal pressure dispersion against the tracheal wall The cuff must be resistant to Nitrous Oxide effects The pilot cuff must have a one-way valve and must indicate the degree of inflation accurately. All components must be latex-free sterile and individually packed in peel pouch that is easy to open - For SINGLE use only To comply with the latest ISO 5361:2016 Size 5mm - 8mm considered in series Size 8.0mm, 8.5mm and 9.0mm considered individually | Taper-Cuffed Oral/Nasal endotracheal tube: ID: 5.0mm OD: 6.9mm
| RT13-02-02S |  | Taper-Cuffed Oral/Nasal endotracheal tube: ID: 5.5mm OD: 7.5mm
| RT13-02-03S |  | Taper-Cuffed Oral/Nasal endotracheal tube: ID: 6.0mm OD: 8.2mm
| RT13-02-04S |  | Taper-Cuffed Oral/Nasal endotracheal tube: ID: 6.5mm OD: 8.9mm
| RT13-02-05S |  | Taper-Cuffed Oral/Nasal endotracheal tube: ID: 7.0mm OD: 9.5mm
| RT13-02-06S |  | Taper-Cuffed Oral/Nasal endotracheal tube: ID: 7.5mm OD: 10.2mm
| RT13-02-07S |  | Taper-Cuffed Oral/Nasal endotracheal tube: ID: 8.0mm OD: 10.8mm
| RT13-02-08S |  | Taper-Cuffed Oral/Nasal endotracheal tube: ID: 8.5mm OD: 11.4mm
| RT13-02-09S |  | Taper-Cuffed Oral/Nasal endotracheal tube: ID: 9.0mm OD: 12.1mm
TAPER-CUFFED ORAL/NASAL ENDOTRACHEAL TUBE WITH EVACUATION LUMEN AT THE SUBGLOTTIC SPACE
CATEGORY 2T | RT13-02-01T | TAPER-CUFFED ORAL/NASAL ENDOTRACHEAL TUBE WITH EVACUATION LUMEN AT THE SUBGLOTTIC SPACE Purpose: endotracheal tube with a tapered cuff for long stay oral/nasal intubation aimed to reduce micro-aspiration Consists of a clear thermo-sensitive PVC tube that is kink resistant, atraumatic to airways and have a full Magill curvature. Must have a radiopaque line and must be marked with numbered depth markings in 1cm intervals, tube size, length, ID, OD, trade name. The markings on tube must be clear and must not rub off. Must have a precision band mark just above the cuff to aid with depth of intubation and guide proper placement. Preference will be given to tubes with a second circumferential band (not more than 2cm from each other) to guide proper placement. Proximal end - secure 15mm OD circuit connector made from medical grade plastic Distal end - Atraumatic left-facing bevelled tip with Murphy eye on the opposite side Must have a high-volume-low-pressure polyurethane taper-shaped cuff with pilot cuff. Cuff thickness ± 50 microns The integrated evacuation lumen above the tapered cuff should be on the posterior aspect of the tube for the suction of subglottic secretions. The suction lumen must have a toggled capped The cuff must inflate and deflate easily and when deflated must conform closely to tube to minimise the risk of trauma during intubation and extubation. Cuff must be barrel shaped to facilitate optimal pressure dispersion against the tracheal wall The cuff must be resistant to Nitrous Oxide effects The pilot cuff must have a one-way valve and must indicate the degree of inflation accurately. All components must be latex-free sterile and individually packed in peel pouch that is easy to open - For SINGLE use only To comply with the latest ISO 5361:2016 Size 6mm - 8mm considered in series Size 8.0mm, 8.5mm and 9.0mm considered individually | Taper-Cuffed Oral/Nasal endotracheal tube with Evacuation Lumen: ID: 6.0mm OD: 9.0mm
| RT13-02-02T |  | Taper-Cuffed Oral/Nasal endotracheal tube with Evacuation Lumen: ID: 6.5mm OD: 9.8mm
| RT13-02-03T |  | Taper-Cuffed Oral/Nasal endotracheal tube with Evacuation Lumen: ID: 7.0mm OD: 10.4mm
| RT13-02-04T |  | Taper-Cuffed Oral/Nasal endotracheal tube with Evacuation Lumen: ID: 7.5mm OD: 11.2mm
| RT13-02-05T |  | Taper-Cuffed Oral/Nasal endotracheal tube with Evacuation Lumen: ID: 8.0mm OD: 11.8mm
| RT13-02-06T |  | Taper-Cuffed Oral/Nasal endotracheal tube with Evacuation Lumen: ID: 8.5mm OD: 12.6mm
| RT13-02-07T |  | Taper-Cuffed Oral/Nasal endotracheal tube with Evacuation Lumen: ID: 9.0mm OD: 13.1mm
UNCUFFED NASAL NORTH-FACING PREFORMED (RAE) ENDOTRACHEAL TUBE - NEWBORN
CATEGORY 2U | RT13-02-01U | UNCUFFED NASAL NORTH-FACING PREFORMED (RAE) ENDOTRACHEAL TUBE - NEWBORN Purpose: Short term oral intubation where tube is North facing and removed from surgical field Consists of a clear thermo-sensitive PVC tube that is kink resistant, with a North polar preformed curve with a soft flexible pointed tip and two Murphy eyes Must be able to temporarily straightened the tube to allow the passage of a suction catheter Must have a radiopaque line and must be marked with numbered depth markings in 1cm intervals, tube size, length, ID, OD, trade name. The markings on tube must be clear and must not rub off. Must also have marks/reference lines proximal to cuff to aid with depth of intubation and guide proper placement. Solid or circumferential precision band markings accepted Preference will be given to tubes with a second circumferential band (not more than 1cm from each other) to guide proper placement. Must also have a distinct mark at preformed curve to aid positioning Proximal end - secure 15mm OD: circuit connector made from medical grade plastic Distal end - Atraumatic left-facing bevelled tip NO or 1 Murphy eyes on the opposite side All components must be latex-free sterile and individually packed in peel pouch that is easy to open - For SINGLE use only To comply with the latest ISO 5361 Size 2.0mm and 2.5mm considered individually | Uncuffed Nasal North-Facing Preformed (RAE) Endotracheal Tube - Newborn ID: 2.0mm OD 4.0mm Length ± 115mm
| RT13-02-02U |  | Uncuffed Nasal North-Facing Preformed (RAE) Endotracheal Tube - Newborn ID: 2.5mm OD 4.0mm Length ± 115mm
UNCUFFED NASAL NORTH-FACING PREFORMED (RAE) ENDOTRACHEAL TUBE - PAEDIATRIC
CATEGORY 2V | RT13-02-01V | UNCUFFED NASAL NORTH-FACING PREFORMED (RAE) ENDOTRACHEAL TUBE - PAEDIATRIC Purpose: Short term oral intubation where tube is North facing and removed from surgical field Consists of a clear thermo-sensitive PVC tube that is kink resistant, with a North polar preformed curve with a soft flexible pointed tip and two Murphy eyes Must be able to temporarily straightened the tube to allow the passage of a suction catheter Must have a radiopaque line and must be marked with numbered depth markings in 1cm intervals, tube size, length, ID, OD, trade name. The markings on tube must be clear and must not rub off. Must also have marks/reference lines proximal to cuff to aid with depth of intubation and guide proper placement. Solid or circumferential precision band markings accepted Preference will be given to tubes with a second circumferential band (not more than 1cm from each other) to guide proper placement. Must also have a distinct mark at preformed curve to aid positioning Proximal end - secure 15mm OD: circuit connector made from medical grade plastic Distal end - Atraumatic left-facing bevelled tip 1 Murphy eye on the opposite side All components must be latex-free sterile and individually packed in peel pouch that is easy to open - For SINGLE use only To comply with the latest ISO 5361 Size 3.0mm - 6.0mm considered in series | Uncuffed Nasal North-Facing Preformed (RAE) Endotracheal Tube - Paediatric ID: 3.0mm OD 4.0mm Length ± 115mm
| RT13-02-02V |  | Uncuffed Nasal North-Facing Preformed (RAE) Endotracheal Tube - Paediatric ID: 3.5mm OD 5.3mm Length ± 155mm
| RT13-02-03V |  | Uncuffed Nasal North-Facing Preformed (RAE) Endotracheal Tube - Paediatric ID: 4.0mm OD 5.3mm Length ± 155mm
| RT13-02-04V |  | Uncuffed Nasal North-Facing Preformed (RAE) Endotracheal Tube - Paediatric ID: 4.5mm OD 6.0mm Length ± 180mm
| RT13-02-05V |  | Uncuffed Nasal North-Facing Preformed (RAE) Endotracheal Tube - Paediatric ID: 5.0mm OD 6.7mm Length ± 190mm
| RT13-02-06V |  | Uncuffed Nasal North-Facing Preformed (RAE) Endotracheal Tube - Paediatric ID: 5.5mm OD 7.3mm Length ± 195mm
| RT13-02-07V |  | Uncuffed Nasal North-Facing Preformed (RAE) Endotracheal Tube - Paediatric ID: 6.0mm OD 8.0mm Length ± 205mm
CUFFED NASAL NORTH-FACING PREFORMED (RAE) ENDOTRACHEAL TUBE - NEWBORN
CATEGORY 2X | RT13-02-01X | CUFFED NASAL NORTH-FACING PREFORMED (RAE) ENDOTRACHEAL TUBE - PAEDIATRIC Purpose: Short term oral intubation where tube is North facing and removed from surgical field Consists of a clear thermo-sensitive PVC tube that is kink resistant, with a North polar preformed curve with a soft flexible pointed tip and two Murphy eyes Must be able to temporarily straightened the tube to allow the passage of a suction catheter Must have a radiopaque line and must be marked with numbered depth markings in 1cm intervals, tube size, length, ID, OD, trade name. The markings on tube must be clear and must not rub off. Must also have marks/reference lines proximal to cuff to aid with depth of intubation and guide proper placement. Solid or circumferential precision band markings accepted Preference will be given to tubes with a second circumferential band (not more than 1cm from each other) to guide proper placement. Must also have a distinct mark at preformed curve to aid positioning Proximal end - secure 15mm OD: circuit connector made from medical grade plastic Distal end - Atraumatic left-facing bevelled tip with 1 Murphy eye on the opposite side Must have a high-volume-low-pressure polyurethane barrel shaped cuff with pilot cuff. Cuff thickness ± 50 microns. The pilot cuff must have a one-way valve and must indicate the degree of inflation accurately. The cuff must inflate and deflate easily and when deflated must conform closely to tube to minimise the risk of trauma during intubation and extubation and not increase the OD of the tube. The cuff must be resistant to Nitrous Oxide effects All components must be latex-free sterile and individually packed in peel pouch that is easy to open - For SINGLE use only To comply with the latest ISO 5361 Size 3.0mm and 6.0mm considered in series. | Cuffed Nasal North-Facing Preformed (RAE) Endotracheal Tube - Newborn ID: 2.0mm
| RT13-02-02X |  | Cuffed Nasal North-Facing Preformed (RAE) Endotracheal Tube - Newborn ID: 2.5mm
CUFFED NASAL NORTH-FACING PREFORMED (RAE) ENDOTRACHEAL TUBE - PAEDIATRIC
CATEGORY 2Y | RT13-02-01Y | CUFFED NASAL NORTH-FACING PREFORMED (RAE) ENDOTRACHEAL TUBE - PAEDIATRIC Purpose: Short term oral intubation where tube is North facing and removed from surgical field Consists of a clear thermo-sensitive PVC tube that is kink resistant, with a North polar preformed curve with a soft flexible pointed tip and two Murphy eyes Must be able to temporarily straightened the tube to allow the passage of a suction catheter Must have a radiopaque line and must be marked with numbered depth markings in 1cm intervals, tube size, length, ID, OD, trade name. The markings on tube must be clear and must not rub off. Must also have marks/reference lines proximal to cuff to aid with depth of intubation and guide proper placement. Solid or circumferential precision band markings accepted Preference will be given to tubes with a second circumferential band (not more than 1cm from each other) to guide proper placement. Must also have a distinct mark at preformed curve to aid positioning Proximal end - secure 15mm OD: circuit connector made from medical grade plastic Distal end - Atraumatic left-facing bevelled tip with 1 Murphy eye on the opposite side Must have a high-volume-low-pressure polyurethane barrel shaped cuff with pilot cuff. Cuff thickness ± 50 microns. The pilot cuff must have a one-way valve and must indicate the degree of inflation accurately. The cuff must inflate and deflate easily and when deflated must conform closely to tube to minimise the risk of trauma during intubation and extubation and not increase the OD of the tube. The cuff must be resistant to Nitrous Oxide effects All components must be latex-free sterile and individually packed in peel pouch that is easy to open - For SINGLE use only To comply with the latest ISO 5361 Size 3.0mm and 6.0mm considered in series. | Cuffed Nasal North-Facing Preformed (RAE) Endotracheal Tube - Paediatric ID: 3.0mm OD 4.0mm Length ± 115mm
| RT13-02-02Y |  | Cuffed Nasal North-Facing Preformed (RAE) Endotracheal Tube - Paediatric ID: 3.5mm OD 5.3mm Length ± 155mm
| RT13-02-03Y |  | Cuffed Nasal North-Facing Preformed (RAE) Endotracheal Tube - Paediatric ID: 4.0mm OD 5.3mm Length ± 155mm
| RT13-02-04Y |  | Cuffed Nasal North-Facing Preformed (RAE) Endotracheal Tube - Paediatric ID: 4.5mm OD 6.0mm Length ± 180mm
| RT13-02-05Y |  | Cuffed Nasal North-Facing Preformed (RAE) Endotracheal Tube - Paediatric ID: 5.0mm OD 6.7mm Length ± 190mm
| RT13-02-06Y |  | Cuffed Nasal North-Facing Preformed (RAE) Endotracheal Tube - Paediatric ID: 5.5mm OD 7.3mm Length ± 195mm
| RT13-02-07Y |  | Cuffed Nasal North-Facing Preformed (RAE) Endotracheal Tube - Paediatric ID: 6.0mm OD 8.0mm Length ± 205mm
CUFFED NASAL NORTH-FACING PREFORMED (RAE) ENDOTRACHEAL TUBE - ADULT
CATEGORY 2Z | RT13-02-01Z | CUFFED NASAL NORTH-FACING PREFORMED (RAE) ENDOTRACHEAL TUBE - ADULT Purpose: Short term oral intubation where tube is North facing and removed from surgical field Consists of a clear thermo-sensitive PVC tube that is kink resistant, with a North polar preformed curve with a soft flexible pointed tip and two Murphy eyes Must be able to temporarily straightened the tube to allow the passage of a suction catheter Must have a radiopaque line and must be marked with numbered depth markings in 1cm intervals, tube size, length, ID, OD, trade name. The markings on tube must be clear and must not rub off. Must also have marks/reference lines proximal to cuff to aid with depth of intubation and guide proper placement. Solid or circumferential precision band markings accepted Preference will be given to tubes with a second circumferential band (not more than 2cm from each other) to guide proper placement. Must also have a distinct mark at preformed curve to aid positioning Proximal end - secure 15mm OD: circuit connector made from medical grade plastic Distal end - Atraumatic left-facing bevelled tip with 1 Murphy eye on the opposite side Must have a high-volume-low-pressure polyurethane barrel shaped cuff with pilot cuff. Cuff thickness ± 50 microns. The pilot cuff must have a one-way valve and must indicate the degree of inflation accurately. The cuff must inflate and deflate easily and when deflated must conform closely to tube to minimise the risk of trauma during intubation and extubation and not increase the OD of the tube. The cuff must be resistant to Nitrous Oxide effects All components must be latex-free sterile and individually packed in peel pouch that is easy to open - For SINGLE use only To comply with the latest ISO 5361 Size 6.5mm and 8.0mm considered in series. | Cuffed Nasal North-Facing Preformed (RAE) Endotracheal Tube - Adult ID: 6.5mm
| RT13-02-02Z |  | Cuffed Nasal North-Facing Preformed (RAE) Endotracheal Tube - Adult ID: 7.0mm
| RT13-02-03Z |  | Cuffed Nasal North-Facing Preformed (RAE) Endotracheal Tube - Adult ID: 7.5mm
| RT13-02-04Z |  | Cuffed Nasal North-Facing Preformed (RAE) Endotracheal Tube - Adult ID: 8.0mm
UNCUFFED ORAL SOUTH-FACING PREFORMED (RAE) ENDOTRACHEAL TUBE WITH MURPHY EYE - PAEDIATRIC
CATEGORY 2AA | RT13-02-01AA | UNCUFFED ORAL SOUTH-FACING PREFORMED (RAE) ENDOTRACHEAL TUBE WITH MURPHY EYE Purpose: Short term oral intubation where tube is South facing and removed from surgical field Consists of a clear thermo-sensitive PVC tube that is kink resistant, with a South polar preformed curve with two Murphy eyes Must be able to temporarily straightened the tube to allow the passage of a suction catheter Must have a radiopaque line and must be marked with numbered depth markings in 1cm intervals, tube size, length, ID, OD, trade name. The markings on tube must be clear and must not rub off. Must also have marks/reference lines proximal to cuff to aid with depth of intubation and guide proper placement. Preference will be given to tubes with a second circumferential band (not more than 1cm from each other) to guide proper placement. Must also have a distinct mark at preformed curve to aid positioning Proximal end - secure 15mm OD circuit connector made from medical grade plastic Distal end - Atraumatic left-facing bevelled tip with Murphy eye on the opposite side All components must be latex-free sterile and individually packed in peel pouch that is easy to open - For SINGLE use only To comply with the latest ISO 5361 Size 2.0mm and 2.5mm will be considered individually Size 3.0mm - 6.0mm considered in series | Uncuffed Oral South-Facing Preformed (RAE) Endotracheal Tube with Murphy eye. ID: 2.0mm
| RT13-02-02AA |  | Uncuffed Oral South-Facing Preformed (RAE) Endotracheal Tube with Murphy eye. ID: 2.5mm
| RT13-02-03AA |  | Uncuffed Oral South-Facing Preformed (RAE) Endotracheal Tube with Murphy eye. ID: 3.0mm OD 4.0mm Length ± 105mm
| RT13-02-04AA |  | Uncuffed Oral South-Facing Preformed (RAE) Endotracheal Tube with Murphy eye. ID: 3.5mm OD 4.7mm Length ± 120mm
| RT13-02-05AA |  | Uncuffed Oral South-Facing Preformed (RAE) Endotracheal Tube with Murphy eye. ID: 4.0mm OD 5.3mm Length ± 130mm
| RT13-02-06AA |  | Uncuffed Oral South-Facing Preformed (RAE) Endotracheal Tube with Murphy eye. ID: 4.5mm OD 6.0mm Length ± 145mm
| RT13-02-07AA |  | Uncuffed Oral South-Facing Preformed (RAE) Endotracheal Tube with Murphy eye. ID: 5.0mm OD 6.7mm Length ± 155mm
| RT13-02-08AA |  | Uncuffed Oral South-Facing Preformed (RAE) Endotracheal Tube with Murphy eye. ID: 5.5mm OD 7.3mm Length ± 165mm
| RT13-02-09AA |  | Uncuffed Oral South-Facing Preformed (RAE) Endotracheal Tube with Murphy eye. ID: 6.0mm OD 8.0mm Length ± 180mm
CUFFED ORAL SOUTH-FACING PREFORMED (RAE) ENDOTRACHEAL TUBE WITH MURPHY EYE - PAEDIATRIC
CATEGORY 2BB | RT13-02-01BB | CUFFED ORAL SOUTH-FACING PREFORMED (RAE) ENDOTRACHEAL TUBE WITH MURPHY EYE - PAEDIATRIC Purpose: Short term oral intubation where tube is South facing and removed from surgical field Consists of a clear thermo-sensitive PVC tube that is kink resistant, with a South polar preformed curve with two Murphy eyes Must be able to temporarily straightened the tube to allow the passage of a suction catheter Must have a radiopaque line and must be marked with numbered depth markings in 1cm intervals, tube size, length, ID, OD, trade name. The markings on tube must be clear and must not rub off. Must also have marks/reference lines proximal to cuff to aid with depth of intubation and guide proper placement. Preference will be given to tubes with a second circumferential band (not more than 1cm from each other) to guide proper placement. Must also have a distinct mark at preformed curve to aid positioning Proximal end - secure 15mm OD circuit connector made from medical grade plastic Distal end - Atraumatic left-facing bevelled tip with Murphy eye on the opposite side Must have a high-volume-low-pressure polyurethane barrel shaped cuff with pilot cuff. Cuff thickness ± 50 microns. The pilot cuff must have a one-way valve and must indicate the degree of inflation accurately. The cuff must inflate and deflate easily and when deflated must conform closely to tube to minimise the risk of trauma during intubation and extubation and not increase the OD of the tube. The cuff must be resistant to Nitrous Oxide effects All components must be latex-free sterile and individually packed in peel pouch that is easy to open - For SINGLE use only To comply with the latest ISO 5361 Size 2.0mm and 2.5mm will be considered individually, with or without murphy eye. Size 3.0mm - 6.0mm considered in series | Cuffed Oral South-Facing Preformed (RAE) Endotracheal Tube with Murphy eye - Paediatric. ID: 2.0mm
| RT13-02-02BB |  | Cuffed Oral South-Facing Preformed (RAE) Endotracheal Tube with Murphy eye - Paediatric. ID: 2.5mm
| RT13-02-03BB |  | Cuffed Oral South-Facing Preformed (RAE) Endotracheal Tube with Murphy eye - Paediatric. ID: 3.0mm OD 4.0mm Length ± 105mm
| RT13-02-04BB |  | Cuffed Oral South-Facing Preformed (RAE) Endotracheal Tube with Murphy eye - Paediatric. ID: 3.5mm OD 4.7mm Length ± 120mm
| RT13-02-05BB |  | Cuffed Oral South-Facing Preformed (RAE) Endotracheal Tube with Murphy eye - Paediatric. ID: 4.0mm OD 5.3mm Length ± 130mm
| RT13-02-06BB |  | Cuffed Oral South-Facing Preformed (RAE) Endotracheal Tube with Murphy eye - Paediatric. ID: 4.5mm OD 6.0mm Length ± 145mm
| RT13-02-07BB |  | Cuffed Oral South-Facing Preformed (RAE) Endotracheal Tube with Murphy eye - Paediatric. ID: 5.0mm OD 6.7mm Length ± 155mm
| RT13-02-08BB |  | Cuffed Oral South-Facing Preformed (RAE) Endotracheal Tube with Murphy eye - Paediatric. ID: 5.5mm OD 7.3mm Length ± 165mm
| RT13-02-09BB |  | Cuffed Oral South-Facing Preformed (RAE) Endotracheal Tube with Murphy eye - Paediatric. ID: 6.0mm OD 8.0mm Length ± 180mm
CUFFED ORAL SOUTH-FACING PREFORMED (RAE) ENDOTRACHEAL TUBE WITH MURPHY EYE - ADULT
CATEGORY2CC | RT13-02-01CC | CUFFED ORAL SOUTH-FACING PREFORMED (RAE) ENDOTRACHEAL TUBE WITH MURPHY EYE - ADULT Purpose: Short term oral intubation where tube is South facing and removed from surgical field Consists of a clear thermo-sensitive PVC tube that is kink resistant, with a South polar preformed curve with two Murphy eyes Must be able to temporarily straightened the tube to allow the passage of a suction catheter Must have a radiopaque line and must be marked with numbered depth markings in 1cm intervals, tube size, length, ID, OD, trade name. The markings on tube must be clear and must not rub off. Must also have marks/reference lines proximal to cuff to aid with depth of intubation and guide proper placement. Preference will be given to tubes with a second circumferential band (not more than 2cm from each other) to guide proper placement. Must also have a distinct mark at preformed curve to aid positioning Proximal end - secure 15mm OD circuit connector made from medical grade plastic Distal end - Atraumatic left-facing bevelled tip with Murphy eye on the opposite side Must have a high-volume-low-pressure polyurethane barrel shaped cuff with pilot cuff. Cuff thickness ± 50 microns. The pilot cuff must have a one-way valve and must indicate the degree of inflation accurately. The cuff must inflate and deflate easily and when deflated must conform closely to tube to minimise the risk of trauma during intubation and extubation and not increase the OD of the tube. The cuff must be resistant to Nitrous Oxide effects All components must be latex-free sterile and individually packed in peel pouch that is easy to open - For SINGLE use only To comply with the latest ISO 5361 Size 6.5mm to 8.0mm considered in series | Cuffed Oral South-Facing Preformed (RAE) Endotracheal Tube with Murphy eye - Adult. ID: 6.5mm OD 8.7mm Length ± 190mm
| RT13-02-02CC |  | Cuffed Oral South-Facing Preformed (RAE) Endotracheal Tube with Murphy eye - Adult ID:7.0mm OD 9.3mm Length ± 200mm
| RT13-02-03CC |  | Cuffed Oral South-Facing Preformed (RAE) Endotracheal Tube with Murphy eye - Adult ID:7.5mm OD 10mm Length ± 210mm
| RT13-02-04CC |  | Cuffed Oral South-Facing Preformed (RAE) Endotracheal Tube with Murphy eye - Adult ID:8.0mm OD 10.7mm Length ± 220mm
DOUBLE LUMEN ENDOBRONCHIAL TUBE- LEFT
CATEGORY2DD | RT13-02-01DD | DOUBLE LUMEN ENDOBRONCHIAL TUBE- LEFT Purpose: For isolation of one lung to prevent contamination, or for ventilation of one lung in thoracic surgery Consists of a double Lumen clear PVC tube that is thermo-sensitive with radiopaque lines and markers at key features. Markings starting at 26cm from distal end marked every 2cm up to and about 32cm Tubes to be flexible enough to allow manipulation through distorted airways, but without luminal collapse. Must have a smooth transition of the inner lumen without narrowing, to allow easy passage of flexible fibre-optic-bronchoscope. Tracheal and Bronchial extensions on the main tube must be colour coded - blue for bronchial, clear for tracheal Stylet: Malleable shape forming - packaged in place down the bronchial lumen Supplied with 2 proximal right-angled connectors with a Bronchoscopy/Aspiration Port that has a diaphragm which hugs fibre-optic bronchoscopes The port must have a toggled cap and allow for easy passage of either a flexible fibre-optic-bronchoscope, or a suction catheter – both without a leak. The proximal angled connectors to connect to the bronchial and tracheal tubes and the distal connectors to connect to a standard Y-connector that is compatible with a standard breathing circuit. 2 - 4 Suction catheters: with suction control, to fit the total length of the tube Colourless tracheal high-volume-low-pressure polyvinyl chloride cuff – with colourless pilot balloon, clearly labelled “tracheal” Blue bronchial low pressure Polyurethane cuff - with blue pilot balloon, clearly labelled “bronchial”. Bronchial cuff to be short and cylindrical. One-way valves in pilot balloons. The cuffs must inflate and deflate easily, and the pilot cuffs must indicate the degree of inflation accurately. Cuffs must be resistant to the effects of nitrous oxide. When deflated, cuffs must conform closely to the tube to minimise the risk of trauma during intubation and extubation. All components must be latex-free, sterile and individually packed in peel pouch - SINGLE use only To comply with ISO 16628:2008 | Double Lumen Endobronchial tube: Left Size: 26Fr Shaft Length: > mm ID: > (bilaterally) OD: <
| RT13-02-02DD |  | Double Lumen Endobronchial tube: Left Size: 28Fr Shaft Length: > 420mm ID: > 3.1mm (bilaterally) OD: < 10mm ± 2mm
| RT13-02-03DD |  | Double Lumen Endobronchial tube: Left Size: 32Fr Shaft Length: > 420mm ID: > 3.3mm (bilaterally) OD: < 10.7mm ± 2mm
| RT13-02-04DD |  | Double Lumen Endobronchial tube: Left Size: 35Fr Shaft Length: > 420mm ID: > 4.7mm (bilaterally) OD: < 11.7mm ± 2mm
| RT13-02-05DD |  | Double Lumen Endobronchial tube: Left Size: 37Fr Shaft Length: > 420mm ID: > 5.0mm (bilaterally) OD: < 12.3mm ± 2mm
| RT13-02-06DD |  | Double Lumen Endobronchial tube: Left Size: 39Fr Shaft Length: > 420mm ID: > 5.2mm (bilaterally) OD: < 13mm ± 2mm
| RT13-02-07DD |  | Double Lumen Endobronchial tube: Left Size: 41Fr Shaft Length:> 420mm ID: > 5.3mm (bilaterally) OD: < 13.7mm ± 2mm
DOUBLE LUMEN ENDOBRONCHIAL TUBE- RIGHT
CATEGORY2EE | RT13-02-01EE | DOUBLE LUMEN ENDOBRONCHIAL TUBE- RIGHT Purpose: For isolation of one lung to prevent contamination, or for ventilation of one lung in thoracic surgery Consists of a double Lumen clear PVC tube that is thermo-sensitive with radiopaque lines and markers at key features. Markings starting at 26cm from distal end marked every 2cm up to and about 32cm Tubes to be flexible enough to allow manipulation through distorted airways, but without luminal collapse. Must have a smooth transition of the inner lumen without narrowing, to allow easy passage of flexible fibre-optic-bronchoscope. Tracheal and Bronchial extensions on the main tube must be colour coded - blue for bronchial, clear for tracheal Stylet: Malleable shape forming - packaged in place down the bronchial lumen Supplied with 2 proximal right-angled connectors with a Bronchoscopy/Aspiration Port that has a diaphragm which hugs fibre-optic bronchoscopes The port must have a toggled cap and allow for easy passage of either a flexible fibre-optic-bronchoscope, or a suction catheter – both without a leak. The proximal angled connectors to connect to the bronchial and tracheal tubes and the distal connectors to connect to a standard Y-connector that is compatible with a standard breathing circuit. 2 - 4 Suction catheters: with suction control, to fit the total length of the tube Colourless tracheal high-volume-low-pressure polyvinyl chloride cuff – with colourless pilot balloon, clearly labelled “tracheal” Blue bronchial low pressure Polyurethane cuff - with blue pilot balloon, clearly labelled “bronchial”. Bronchial cuff to be saddle-shaped One-way valves in pilot balloons. The cuffs must inflate and deflate easily, and the pilot cuffs must indicate the degree of inflation accurately. Cuffs must be resistant to the effects of nitrous oxide. When deflated, cuffs must conform closely to the tube to minimise the risk of trauma during intubation and extubation. All components must be latex-free, sterile and individually packed in peel pouch - SINGLE use only To comply with ISO 16628:2008 | Double Lumen Endobronchial tube: Right Size: 26Fr Shaft Length: 420mm ID: > mm (bilaterally) OD: <
| RT13-02-02EE |  | Double Lumen Endobronchial tube: Right Size: 28Fr Shaft Length: 420mm ID: > 3.2mm (bilaterally) OD: < 9.3mm ± 2mm
| RT13-02-03EE |  | Double Lumen Endobronchial tube: Right Size: 32Fr Shaft Length: 420mm ID: > 3.4mm (bilaterally) OD: < 10.7mm ± 2mm
| RT13-02-04EE |  | Double Lumen Endobronchial tube: Right Size: 35Fr Shaft Length: 420mm ID: > 4.7mm (bilaterally) OD: < 11.7mm ± 2mm
| RT13-02-05EE |  | Double Lumen Endobronchial tube: Right Size: 37Fr Shaft Length: 420mm ID: > 5.0mm (bilaterally) OD: < 12.3mm ± 2mm
| RT13-02-06EE |  | Double Lumen Endobronchial tube: Right Size: 39Fr Shaft Length: 420mm ID: > 5.2mm (bilaterally) OD: < 13mm ± 2mm
| RT13-02-07EE |  | Double Lumen Endobronchial tube: Right Size: 41Fr Shaft Length: 420mm ID: > 5.3mm (bilaterally) OD: < 13.7mm ± 2mm
COMBINED ENDOTRACHEAL TUBE AND ENDOBRONCHIAL BLOCKER
CATEGORY 2FF | RT13-02-01FF | COMBINED ENDOTRACHEAL TUBE AND ENDOBRONCHIAL BLOCKER Purpose: For isolation of one lung to prevent contamination, or for ventilation of one lung in thoracic surgery Consists of a single main lumen endotracheal tube for ventilation and a secondary integrated lumen provided to guide a bi-directional, torque controlled bronchus blocker. The tube should function as a conventional endotracheal tube once the bronchus blocker is removed The blocker must allow occlusion of either the right or the left bronchus or for selective lobar blockade The endobronchial blocker tube should be hollow and allow CPAP or high frequency jet ventilation, irrigation, suctioning and/or gas sampling from the occluded lung. The endobronchial blocker should have a malleable shaft with high torque control for manipulation, a soft distal tip with a high volume low pressure cuff, and should extend 8-10cm beyond the tip. All components must be latex-free, sterile and individually packed in peel pouch - SINGLE use only | Combined Endotracheal Tube and Endobronchial Blocker Size 6.0mm
| RT13-02-02FF |  | Combined Endotracheal Tube and Endobronchial Blocker Size 7.0mm
| RT13-02-03FF |  | Combined Endotracheal Tube and Endobronchial Blocker Size 8.0mm
| RT13-02-04FF |  | Combined Endotracheal Tube and Endobronchial Blocker Size 9.0mm
ENDOBRONCHIAL BLOCKER with BIFURCATED DISTAL EXTENSIONS
CATEGORY 2GG | RT13-02-01GG | ENDOBRONCHIAL BLOCKER with BIFURCATED DISTAL EXTENSIONS Size: 7Fr Length: 75cm Purpose: For one lung isolation via an existing single lumen endotracheal tube Consists of a radiopaque shaft with depth markings and bifurcated colour-coded distal extensions The distal extensions must have have colour-coded inflatable polyurethane cuffs that match the colour-coded pilot cuffs on both proximal ends and allows occlusion of either the right or the left bronchus. The extensions must have a central lumen to allow for oxygen to be administered to the isolated lung during procedure. The proximal end must be able to allow CO2 flow check to ensure cuff is providing a full seal. The set must include a transparent multi-port adapter that will fascilitate ventilation, allow insertion of the endobronchial blocker and a fibre-optic endoscope or suction cathater. Must allow for CPAP and must have 2 free caps to close ports All components must be latex-free, sterile and individually packed in peel pouch - SINGLE use only | ENDOBRONCHIAL BLOCKER with BIFURCATED DISTAL EXTENSIONS Size: 7Fr Length: 75cm
MICROLARYNGEAL TUBES
CATEGORY 2HH | RT13-02-01HH | MICROLARYNGEAL TUBES Purpose: Longer tube for use in narrowed airways or laryngeal surgery Consists of a clear thermo-sensitive extra-long PVC tube that is kink resistant, atraumatic to airways and have a full Magill curvature. Must have a radiopaque line and must be marked with numbered depth markings in 1 cm intervals, tube size, length, ID, OD, trade name. The markings on the tube must be clear and must not rub off. Proximal end - secure 15mm OD circuit connector made from medical grade plastic Distal end - bevelled atraumatic tip with a murphy eye on opposite side to bevel. The high-volume-low-pressure polyurethane cuff with pilot cuff . The pilot cuff must have a one-way valve and must indicate the degree of inflation accurately. The cuff must inflate and deflate easily, and when deflated must conform closely to tube to minimise the risk of trauma during intubation, and extubation and not increase the OD of the tube. Must be appropriately sized in relation to the tube size and appropriately placed with regards to the distance from the distal tip. The cuff must be resistant to Nitrous Oxide effects Cuff thickness ± 50 micron All components must be latex-free sterile and individually packed in peel pouch that is easy to open - For SINGLE use only To comply with the latest ISO 5361 | Cuffed MicroLaryngeal Oral/Nasal tracheal tube with Murphy eye Size: ID: 4.0mm OD: 5.5 – 5.8mm Tube length ≥ 320mm
| RT13-02-02HH |  | Cuffed MicroLaryngeal Oral/Nasal tracheal tube with Murphy eye Size: ID: 5.0mm OD: 6.9 - 7.2mm Tube length ≥ 320mm
| RT13-02-03HH |  | Cuffed MicroLaryngeal Oral/Nasal tracheal tube with Murphy eye Size: ID: 6.0mm OD: 8.2mm Tube length ≥ 320mm
UNCUFFED LASER ENDOTRACHEAL TUBE WITHOUT A MURPHY EYE
CATEGORY 2II | RT13-02-01II | UNCUFFED LASER ENDOTRACHEAL TUBES WITHOUT A MURPHY EYE Purpose: Endotracheal tube for use during laser airway surgery. Consists of an airtight stainless steel endotracheal tube that ends in a permanently attached 15mm connector proximally. The tube must have a Magill curvature, be flexible but kink-resistant to allow it to be positioned outside of the surgical field. The stainless steel design must have a matte finish to defocus the reflected laser beams. The tube must be resistant to C02 and KTP lasers and non-flammable. Company must produce evidence of laser power settings at which ignition does not occur. Distal end must end in a medical grade plastic atraumatic left-facing bevelled tip without a Murphy eye. Must have a precision band mark to aid with depth of intubation and guide proper placement. Solid or circumferential precision band markings accepted Packaging must be marked with tube size, length, ID, OD, trade name. . All components must be sterile and individually packed in peel pouch that is easy to open - For SINGLE use only To comply with the latest ISO xxxx | Uncuffed LASER Endotracheal Tubes without a Murphy Eye ID: 3.0mm OD: 5.2mm
| RT13-02-02II |  | Uncuffed LASER Endotracheal Tubes without a Murphy Eye ID: 3.5mm OD: 5.7mm
| RT13-02-03II |  | Uncuffed LASER Endotracheal Tubes without a Murphy Eye ID: 4.0mm OD: 6.1mm
DOUBLE-CUFFED LASER ENDOTRACHEAL TUBES
CATEGORY 2JJ | RT13-02-01JJ | DOUBLE-CUFFED STAINLESS STEEL LASER ENDOTRACHEAL TUBES -FOR VENTILATION DURING CO2 AND KTP LASER Purpose: Endotracheal tube for use during laser airway surgery. Consists of an airtight stainless steel endotracheal tube that ends in a permanently attached 15mm connector proximally. The tube must have a Magill curvature, be flexible but kink-resistant to allow it to be positioned outside of the surgical field. The stainless steel design must have a matte finish to defocus the reflected laser beams. Must have two cuffs that can be filled with saline to seal the airway and maintain ventilation should one cuff be punctured. The pilot balloons for the cuffs, must be colour-coded, have a self sealing valve and must indicate proximal or distal. The tube must be resistant to C02 and KTP lasers and non-flammable. Company must produce evidence of laser power settings at which ignition does not occur. Distal end must end in a medical grade plastic atraumatic left-facing bevelled tip with Murphy eye on the opposite side Must have a precision band mark to aid with depth of intubation and guide proper placement. Solid or circumferential precision band markings accepted Packaging must be marked with tube size, length, ID, OD, trade name. . All components must be sterile and individually packed in peel pouch that is easy to open - For SINGLE use only To comply with the latest ISO xxxx | Double-Cuffed Stainless Steel LASER Endotracheal Tubes ID: 4.5mm OD: 7.0mm
| RT13-02-02JJ |  | Double-Cuffed Stainless Steel LASER Endotracheal Tubes ID: 5.0mm OD: 7.5mm
| RT13-02-03JJ |  | Double-Cuffed Stainless Steel LASER Endotracheal Tubes ID: 5.5mm OD: 7.9mm
| RT13-02-04JJ |  | Double-Cuffed Stainless Steel LASER Endotracheal Tubes ID: 6.0mm OD: 8.5mm
DOUBLE-CUFFED LASER ENDOTRACHEAL TUBES
CATEGORY 2KK | RT13-02-01KK | DOUBLE-CUFFED RUBBER LASER ENDOTRACHEAL TUBES -FOR VENTILATION DURING ARGON, ND/YAG AND CO2 TYPE LASER Purpose: Endotracheal tube for use during laser airway surgery. Consists of a soft natural rubber tube with micro-corrugated silver foil that ends in a 15mm connector. The tube must have a Magill curvature, be flexible but kink-resistant to allow it to be positioned outside of the surgical field. The tube must be marked with tube size, length, ID, OD, trade name. . Proximal end - secure 15mm OD circuit connector made from medical grade plastic Distal end - Atraumatic left-facing bevelled tip with Murphy eye on the opposite side Must have two cuffs that must be able to be saline filled to seal the airway and maintain ventilation should one cuff be punctured. The cuffs must be resistant to the diffusion of anesthetic gas. The tube must be resistant to Argon, Nd/YAG and CO2 type laser in the Wave range of 0.488mm to 10.6mm. Company must produce evidence of laser power settings at which ignition does not occur. All components must be sterile and individually packed in peel pouch that is easy to open - For SINGLE use only To comply with the latest ISO xxxx | Double-Cuffed Rubber LASER Endotracheal Tubes ID: 4.0mm OD: 8.0mm Length: 400mm
| RT13-02-02KK |  | Double-Cuffed Rubber LASER Endotracheal Tubes ID: 5.0mm OD: 9.0mm Length: 400mm
| RT13-02-03KK |  | Double-Cuffed Rubber LASER Endotracheal Tubes ID: 6.0mm OD: 10.8mm Length: 400mm
| RT13-02-04KK |  | Double-Cuffed Rubber LASER Endotracheal Tubes ID: 7.0mm OD: 12.3mm Length: 400mm
| RT13-02-05KK |  | Double-Cuffed Rubber LASER Endotracheal Tubes ID: 8.0mm OD: 13.4mm Length: 400mm

## Category 3
BID NUMBER | RT13
BID DESCRIPTION | SUPPLY AND DELIVERY OF RESPIRATORY AID AND ANAESTHETIC ACCESSORIES TO THE STATE
CATEGORY 3
SUPRAGLOTTIC AIRWAY DEVICES
| Item Number | Technical Specification | Product Name and Size | Insert inputs on this column
1st GENERATION SUPRAGLOTTIC AIRWAY DEVICE WITH APERTURE BARS - SINGLE USE
CATEGORY 3A | RT13-03-01A | 1st GENERATION SUPRAGLOTTIC AIRWAY DEVICE WITH APERTURE BARS - SINGLE USE Purpose: Ist Generation Supraglottic airway device for a spontaneously breathing or ventilated patient. Consists of a kink resistant airway tube ending in a soft silicone elliptical pharyngeal cuff with a one-way valve and bulb for inflation and deflation of the cuff The cuff must be suitable for positive pressure ventilation up to 20cm/H20 and designed to minimize epiglottic obstruction The tip should resist folding over during insertion. Must have aperture bars to prevent obstruction and protect the epiglottis The airway tube must have a smooth transition to the posterior aspect of the boot of the cuff and a 15mm connector proximally Must have clear markings that do not rub off to indicate Trade name, Size, Recommended patient weight There must also be a guiding line on the posterior aspect to assist with orientation Manufactured from latex-free medical grade PVC /Silicone/ Styrene-ethylene-butylene-styrene (SEBS) Must be sterile and individually packed in peel pouch that is easy to open - For SINGLE use only Must comply with CE 0086 or equivalent Size 1 and 1.5 and 6 considered individually, Size 2 - 5 will be considered in series | Ist Generation Supraglottic Airway Device with Aperture Bars - Single Use Size: 1 For weight: up to 5 kg
| RT13-03-02A |  | Ist Generation Supraglottic Airway Device with Aperture Bars - Single Use Size: 1.5 For weight: 5-10kg
| RT13-03-03A |  | Ist Generation Supraglottic Airway Device with Aperture Bars - Single Use Size: 2 For weight: 10-20kg
| RT13-03-04A |  | Ist Generation Supraglottic Airway Device with Aperture Bars - Single Use Size: 2.5 For weight: 20-30kg
| RT13-03-05A |  | Ist Generation Supraglottic Airway Device with Aperture Bars - Single Use Size: 3 For weight: 30-50kg
| RT13-03-06A |  | Ist Generation Supraglottic Airway Device with Aperture Bars - Single Use Size: 4 For weight: 50-70kg
| RT13-03-07A |  | Ist Generation Supraglottic Airway Device with Aperture Bars - Single Use Size: 5 For weight: 70-100kg
| RT13-03-08A |  | Ist Generation Supraglottic Airway Device with Aperture Bars - Single Use Size: 6 For weight: > 100kg
1st GENERATION SUPRAGLOTTIC AIRWAY DEVICE WITHOUT APERTURE BARS - SINGLE USE
CATEGORY 3B | RT13-03-01B | 1st GENERATION SUPRAGLOTTIC AIRWAY DEVICE WITHOUT APERTURE BARS - SINGLE USE Purpose: Ist Generation Supraglottic airway device for a spontaneously breathing or ventilated patient. Consists of a kink resistant airway tube ending in a soft silicone elliptical pharyngeal cuff with a one-way valve and bulb for inflation and deflation of the cuff The cuff must be suitable for positive pressure ventilation up to 20cm/H20 and designed to minimize epiglottic obstruction The tip should resist folding over during insertion. Must not have aperture bars The airway tube must have a smooth transition to the posterior aspect of the boot of the cuff and a 15mm connector proximally Must have clear markings that do not rub off to indicate Trade name, Size, Recommended patient weight There must also be a guiding line on the posterior aspect to assist with orientation Manufactured from latex-free medical grade PVC /Silicone/ Styrene-ethylene-butylene-styrene (SEBS) Must be sterile and individually packed in peel pouch that is easy to open - For SINGLE use only Must comply with CE 0086 or equivalent Size 1 and 1.5 and 6 considered individually, Size 2 - 5 will be considered in series | Ist Generation Supraglottic Airway Device without Aperture Bars - Single Use Size: 1 For weight: up to 5 kg
| RT13-03-02B |  | Ist Generation Supraglottic Airway Device without Aperture Bars - Single Use Size: 1.5 For weight: 5-10kg
| RT13-03-03B |  | Ist Generation Supraglottic Airway Device without Aperture Bars - Single Use Size: 2 For weight: 10-20kg
| RT13-03-04B |  | Ist Generation Supraglottic Airway Device without Aperture Bars - Single Use Size: 2.5 For weight: 20-30kg
| RT13-03-05B |  | Ist Generation Supraglottic Airway Device without Aperture Bars - Single Use Size: 3 For weight: 30-50kg
| RT13-03-06B |  | Ist Generation Supraglottic Airway Device without Aperture Bars - Single Use Size: 4 For weight: 50-70kg
| RT13-03-07B |  | Ist Generation Supraglottic Airway Device without Aperture Bars - Single Use Size: 5 For weight: 70-100kg
| RT13-03-08B |  | Ist Generation Supraglottic Airway Device without Aperture Bars - Single Use Size: 6 For weight: > 100kg
Ist GENERATION SUPRAGLOTTIC AIRWAY DEVICE WITH PREFORMED CURVE - SINGLE USE
CATEGORY 3C | RT13-03-01C | Ist GENERATION SUPRAGLOTTIC AIRWAY DEVICE WITH PREFORMED CURVE - SINGLE USE Purpose: Ist Generation Supraglottic airway device in a spontaneously breathing or ventilated patient. Consists of an preformed anatomically curved, kink resistant airway tube ending in a soft elliptical PVC pharyngeal cuff with a one-way valve and bulb for inflation and deflation of the cuff The airway tube must allow for easy one hand insertion, have a 15mm connector proximally and a smooth transition to the posterior aspect of the boot distally The cuff must create a high-pressure seal and must be suitable for positive pressure ventilation of at least 20cm/H20 The tip should resist folding over during insertion. Must not have aperture bars The cuff must have a one-way valve and bulb that indicates the degree of inflation and deflation of pharyngeal cuff. The device must have clear markings that do not rub off to indicate Trade name; Size; Recommended patient weight Guideline on posterior aspect to assist with orientation Manufactured from latex-free medical grade PVC /Silicone/ Styrene-ethylene-butylene-styrene (SEBS) Must be sterile and individually packed in peel pouch that is easy to open - For SINGLE use only Must comply with CE 0086 or equivalent Size 1 and 1.5 and 6 considered individually, Size 2 - 5 will be considered in series | Ist Generation Supraglottic Airway Device with Preformed Curve - Single use Size: 1 For weight: up to 5 kg
| RT13-03-02C |  | Ist Generation Supraglottic Airway Device with Preformed Curve - Single use Size: 1.5 For weight: 5-10kg
| RT13-03-03C |  | Ist Generation Supraglottic Airway Device with Preformed Curve - Single use Size: 2 For weight: 10-20kg
| RT13-03-04C |  | Ist Generation Supraglottic Airway Device with Preformed Curve - Single use Size: 2.5 For weight: 20-30kg
| RT13-03-05 |  | Ist Generation Supraglottic Airway Device with Preformed Curve - Single use Size: 3 For weight: 30-50kg
| RT13-03-06 |  | Ist Generation Supraglottic Airway Device with Preformed Curve - Single use Size: 4 For weight: 50-70kg
| RT13-03-07 |  | Ist Generation Supraglottic Airway Device with Preformed Curve - Single use Size: 5 For weight: 70-100kg
| RT13-03-08 |  | Ist Generation Supraglottic Airway Device with Preformed Curve - Single use Size: 6 For weight: > 100kg
1st GENERATION SUPRAGLOTTIC AIRWAY DEVICE WITH APERTURE BARS AND REINFORCED TUBING - SINGLE USE
CATEGORY 3D | RT13-03-01D | 1st GENERATION SUPRAGLOTTIC AIRWAY DEVICE WITH APERTURE BARS AND REINFORCED TUBING - SINGLE USE Purpose: 1st Generation Supraglottic Airway Device control in a spontaneously breathing or ventilated patient. Consists of a kink resistant wire reinforced flexible airway tube ending in a soft silicone elliptical pharyngeal cuff with a one-way valve and bulb for inflation and deflation of the cuff. The cuff must be suitable for positive pressure ventilation up to 20cm/H20. Must have aperture bars to prevent obstruction and protect the epiglottis. The tip should resist folding over during insertion. The airway tube must have a smooth transition to the posterior aspect of the boot of the cuff and a 15mm connector proximally Must have clear markings that do not rub off, to indicate Trade name, Size, Recommended patient weight There must also be a guiding line on the posterior aspect to assist with orientation Manufactured from latex-free medical grade PVC /Silicone/ Styrene-ethylene-butylene-styrene (SEBS) Must be sterile and individually packed in peel pouch that is easy to open - For SINGLE use only Must comply with CE 0086 or equivalent Size 1 and 1.5 and 6 considered individually, Size 2 - 5 will be considered in series | 1st Generation Supraglottic Airway Device with Aperture Bars and Reinforced Tubing - Single Use Size: 1 For weight: up to 5kg
| RT13-03-02D |  | 1st Generation Supraglottic Airway Device with Aperture Bars and Reinforced Tubing - Single Use Size: 1.5 For weight: 5 - 10kg
| RT13-03-03D |  | 1st Generation Supraglottic Airway Device with Aperture Bars and Reinforced Tubing - Single Use Size: 2 For weight: 10-20kg
| RT13-03-04D |  | 1st Generation Supraglottic Airway Device with Aperture Bars and Reinforced Tubing - Single Use Size: 2.5 For weight: 20-30kg
| RT13-03-05D |  | 1st Generation Supraglottic Airway Device with Aperture Bars and Reinforced Tubing - Single Use Size: 3 For weight: 30-50kg
| RT13-03-06D |  | 1st Generation Supraglottic Airway Device with Aperture Bars and Reinforced Tubing - Single Use Size: 4 For weight: 50 - 70 kg
| RT13-03-07D |  | 1st Generation Supraglottic Airway Device with Aperture Bars and Reinforced Tubing - Single Use Size: 5 For weight: 70-100kg
| RT13-03-08D |  | 1st Generation Supraglottic Airway Device with Aperture Bars and Reinforced Tubing - Single Use Size: 6 For weight: > 100kg
1st GENERATION SUPRAGLOTTIC AIRWAY DEVICE WITH REINFORCED TUBING AND WITHOUT APERTURE BARS - SINGLE USE
CATEGORY 3E | RT13-03-01E | 1st GENERATION SUPRAGLOTTIC AIRWAY DEVICE WITH REINFORCED TUBING AND WITHOUT APERTURE BARS - SINGLE USE Purpose: 1st Generation Supraglottic Airway Device control in a spontaneously breathing or ventilated patient. Consists of a kink resistant wire reinforced flexible airway tube ending in a soft silicone elliptical pharyngeal cuff with a one-way valve and bulb for inflation and deflation of the cuff. The cuff must be suitable for positive pressure ventilation up to 20cm/H20. Must NOT have aperture bars The tip should resist folding over during insertion. The airway tube must have a smooth transition to the posterior aspect of the boot of the cuff and a 15mm connector proximally Must have clear markings that do not rub off, to indicate Trade name, Size, Recommended patient weight There must also be a guiding line on the posterior aspect to assist with orientation Manufactured from latex-free medical grade PVC /Silicone/ Styrene-ethylene-butylene-styrene (SEBS) Must be sterile and individually packed in peel pouch that is easy to open - For SINGLE use only Must comply with CE 0086 or equivalent Size 1 and 1.5 and 6 considered individually, Size 2 - 5 will be considered in series | 1st Generation Supraglottic Airway Device with Reinforced Tubing and without Aperture Bars - Single Use Size: 1 For weight: up to 5kg
| RT13-03-02E |  | 1st Generation Supraglottic Airway Device with Reinforced Tubing and without Aperture Bars - Single Use Size: 1.5 For weight: 5 - 10kg
| RT13-03-03E |  | 1st Generation Supraglottic Airway Device with Reinforced Tubing and without Aperture Bars - Single Use Size: 2 For weight: 10-20kg
| RT13-03-04E |  | 1st Generation Supraglottic Airway Device with Reinforced Tubing and without Aperture Bars - Single Use Size: 2.5 For weight: 20-30kg
| RT13-03-05E |  | 1st Generation Supraglottic Airway Device with Reinforced Tubing and without Aperture Bars - Single Use Size: 3 For weight: 30-50kg
| RT13-03-06E |  | 1st Generation Supraglottic Airway Device with Reinforced Tubing and without Aperture Bars - Single Use Size: 4 For weight: 50 - 70 kg
| RT13-03-07E |  | 1st Generation Supraglottic Airway Device with Reinforced Tubing and without Aperture Bars - Single Use Size: 5 For weight: 70-100kg
| RT13-03-08E |  | 1st Generation Supraglottic Airway Device with Reinforced Tubing and without Aperture Bars - Single Use Size: 6 For weight: > 100kg
2nd GENERATION SUPRAGLOTTIC AIRWAY DEVICE WITH INFLATABLE CUFF AND A GASTRIC OUTLET - SINGLE USE
CATEGORY 3F | RT13-03-01F | 2nd GENERATION SUPRAGLOTTIC AIRWAY DEVICE WITH INFLATABLE CUFF AND A GASTRIC OUTLET - SINGLE USE Purpose: 2nd Generation Supraglottic airway device in a spontaneously breathing or ventilated patient. Consists of an preformed anatomically curved, kink resistant soft airway tube ending in a soft elliptical silicone pharyngeal cuff. The airway tube must allow for easy one hand insertion, must have an integral bite block, have a 15mm connector proximally and a smooth transition to the posterior aspect of the boot distally The cuff must create a high-pressure seal and must be suitable for positive pressure ventilation of at least 30cm/H20 The tip should resist folding over during insertion. Must NOT have aperture bars. The proximal end of the gastric outlet tube must be able to connect directly to a suction source or allow a suction catheter to be passed for gastric deflation as required. This gastric outlet must be transparent. The gastric drain outlet must open at the tip of the elliptical cuff distally. The cuff must have a one-way valve and bulb that indicates the degree of inflation and deflation of pharyngeal cuff. The device must have clear markings that do not rub off to indicate Trade name; Size; Recommended patient weight Guideline on posterior aspect to assist with orientation Manufactured from latex-free medical grade PVC /Silicone/ Styrene-ethylene-butylene-styrene (SEBS) Must be sterile and individually packed in peel pouch that is easy to open - For SINGLE use only Must comply with CE 0086 or equivalent Size 1 and 1.5 and 6 considered individually, Size 2 - 5 will be considered in series | 2nd Generation Supraglottic Airway Device with Inflatable Cuff and a Gastric Outlet - Single Use Size: 1 For weight: 1 - 5kg Gastric drain to allow 6Fr catheter
| RT13-03-02F |  | 2nd Generation Supraglottic Airway Device with Inflatable Cuff and a Gastric Outlet - Single Use Size: 1.5 For weight: 5-10 kg Gastric drain to allow 8Fr catheter
| RT13-03-03F |  | 2nd Generation Supraglottic Airway Device with Inflatable Cuff and a Gastric Outlet - Single Use Size: 2 For weight: 10-20kg Gastric drain to allow 10Fr catheter
| RT13-03-04F |  | 2nd Generation Supraglottic Airway Device with Inflatable Cuff and a Gastric Outlet - Single Use Size:: 2.5 For weight: 20-30 kg Gastric drain to allow 12Fr catheter
| RT13-03-05F |  | 2nd Generation Supraglottic Airway Device with Inflatable Cuff and a Gastric Outlet - Single Use Size: 3 For weight: 30-50kg Gastric drain to allow 14Fr catheter
| RT13-03-06F |  | 2nd Generation Supraglottic Airway Device with Inflatable Cuff and a Gastric Outlet - Single Use Size: 4 For weight: 50-70kg Gastric drain to allow 14Fr catheter
| RT13-03-07F |  | 2nd Generation Supraglottic Airway Device with Inflatable Cuff and a Gastric Outlet - Single Use Size: 5 For weight: 70-100kg Gastric drain to allow 14Fr catheter
| RT13-03-08F |  | 2nd Generation Supraglottic Airway Device with Inflatable Cuff and a Gastric Outlet - Single Use Size: 6 For weight: > 100kg Gastric drain to allow 14Fr catheter
2nd GENERATION GEL-LIKE SUPRAGLOTTIC AIRWAY DEVICE WITH NON-INFLATABLE CUFF AND A GASTRIC OUTLET - SINGLE USE
CATEGORY 3G | RT13-03-01G | 2nd GENERATION GEL-LIKE SUPRAGLOTTIC AIRWAY DEVICE WITH NON-INFLATABLE CUFF AND A GASTRIC OUTLET- SINGLE USE Purpose: 2nd Generation Supraglottic airway Device control in a spontaneously breathing or ventilated patient. Consists of a preformed anatomically curved, kink resistant soft gel-like airway tube ending in a non-inflatable elliiptical bowl allowing easy and atraumatic insertion. The distal gel cuff should warm to body temperatures and mold to ensure good seating and allow for positive pressure ventilation of at least 30cm/H20 Must have an integral bite block incorporated into the airway tube and a buccal cavity stabiliser to aid with insertion and prevent rotation. Must have clear markings that do not rub off to indicate Trade name; Size; Recommended patient weight Manufactured from latex-free medical grade thermoplastic elastomer. Must be sterile and individually packed in peel pouch that is easy to open, For SINGLE use only Must comply with ISO 11712 Size 1 and 1.5 considered individually, Size 2 - 5 will be considered in series | 2nd Generation Supraglottic Airway Device with Non-Inflatable Cuff and NO Gastric Outlet - Single Use. Size: 1 For weight: 2-5kg
| RT13-03-02G |  | 2nd Generation Supraglottic Airway Device with Non-Inflatable Cuff and a Gastric Outlet - Single Use. Size: 1.5 For weight: 5-12kg Gastric drain to allow 10Fr catheter
| RT13-03-03G |  | 2nd Generation Supraglottic Airway Device with Non-Inflatable Cuff and a Gastric Outlet - Single Use. Size: 2 For weight: 10-25kg Gastric drain to allow 12Fr catheter
| RT13-03-04G |  | 2nd Generation Supraglottic Airway Device with Non-Inflatable Cuff and a Gastric Outlet - Single Use. Size: 2.5 For weight: 25-35kg Gastric drain to allow 12Fr catheter
| RT13-03-05G |  | 2nd Generation Supraglottic Airway Device with Non-Inflatable Cuff and a Gastric Outlet - Single Use. Size: 3 For weight: 30–60kg Gastric drain to allow 12Fr catheter
| RT13-03-06G |  | 2nd Generation Supraglottic Airway Device with Non-Inflatable Cuff and a Gastric Outlet - Single Use. Size: 4 For weight: 50–90kg Gastric drain to allow 12Fr catheter
| RT13-03-07G |  | 2nd Generation Supraglottic Airway Device with Non-Inflatable Cuff and a Gastric Outlet - Single Use. Size: 5 For weight: > 90kg Gastric drain to allow 14Fr catheter
1st GENERATION SUPRAGLOTTIC AIRWAY DEVICE - REUSABLE
CATEGORY 3H | RT13-03-01H | 1st GENERATION SUPRAGLOTTIC AIRWAY DEVICE WITH APERTURE BARS - REUSABLE Purpose: 1st Generation Supraglottic Airway Device control in a spontaneously breathing or ventilated patient. Consists of a kink resistant airway tube ending in a soft silicone elliptical pharyngeal cuff with a one-way valve and bulb for inflation and deflation of the cuff. The cuff must be suitable for positive pressure ventilation up to 20cm/H20. Must have aperture bars to prevent obstruction and protect the epiglottis. The tip should resist folding over during insertion. The airway tube must have a smooth transition to the posterior aspect of the boot of the cuff and a 15mm connector proximally Must have clear markings that do not rub off, to indicate Trade name, Size, Recommended patient weight There must also be a guiding line on the posterior aspect to assist with orientation Must be manufactured from latex-free Medical Grade silicone. Must comply with ISO 11712 Must be autoclavable – at least 40 times. Recording card must be included Size 1 and 1.5 considered individually, Size 2 - 5 will be considered in series | 1st Generation Supraglottic Airway Device with Aperture Bars - REUSABLE Size: 1 For weight: up to 5 kg
| RT13-03-02H |  | 1st Generation Supraglottic Airway Device with Aperture Bars - REUSABLE Size: 1.5 For weight: 5-10kg
| RT13-03-03H |  | 1st Generation Supraglottic Airway Device with Aperture Bars - REUSABLE Size: 2 For weight: 10 – 20 kg
| RT13-03-04H |  | 1st Generation Supraglottic Airway Device with Aperture Bars - REUSABLE Size: 2.5 For weight: 20 – 30 kg
| RT13-03-05H |  | 1st Generation Supraglottic Airway Device with Aperture Bars - REUSABLE Size: 3 For weight: 30 – 50 kg
| RT13-03-06H |  | 1st Generation Supraglottic Airway Device with Aperture Bars - REUSABLE Size: 4 For weight: 50 – 70 kg
| RT13-03-07H |  | 1st Generation Supraglottic Airway Device with Aperture Bars - REUSABLE Size: 5 For weight: 70 – 100 kg
2nd GENERATION SUPRAGLOTTIC AIRWAY DEVICE - REUSABLE
CATEGORY 3I | RT13-03-01I | 2nd GENERATION SUPRAGLOTTIC AIRWAY DEVICE WITH REINFORCED AIRWAY TUBING AND A GASTRIC OUTLET - REUSABLE Purpose: 2nd generation Supraglottic airway Device for control in a spontaneously breathing or ventilated patient. Consists of a kink resistant wire reinforced flexible airway ending in a soft silicone elliptical pharyngeal cuff with a one-way valve and bulb for inflation and deflation of the cuff The cuff must be suitable for positive pressure ventilation up to 30cm/H20. The tip should resist folding over during insertion. The boot of the elliptical cuff must have a pouch to accomodate an introducer. Must be packed with a compatible introducer. There must be a gastric drain tube to allow a suction catheter to be passed for gastric deflation as required. The gastric drain tube should open at the tip of the elliptical cuff. The airway tube must have a bite block with a smooth transition to the posterior aspect of the boot of the cuff and a 15mm connector proximally Must have clear markings that do not rub off to indicate Trade name, Size, Recommended patient weight There must also be a guiding line on the posterior aspect to assist with orientation Must be manufactured from latex-free, medical grade silicone. Must comply with ISO 11712 Must be autoclavable – at least 40 times. Recording card must be included. Individually packed and surgical clean. | 2nd Generation Supraglottic Airway Device with Reinforced Airway Tubing and a Gastric Outlet - Reusable Size: 1 For weight: 1-5kg Gastric drain to allow 8Fr catheter
| RT13-03-02I |  | 2nd Generation Supraglottic Airway Device with Reinforced Airway Tubing and a Gastric Outlet - Reusable Size: 1.5 For weight: 5-10kg Gastric drain to allow 8Fr catheter
| RT13-03-03I |  | 2nd Generation Supraglottic Airway Device with Reinforced Airway Tubing and a Gastric Outlet - Reusable Size: 2 For weight: 10-20kg Gastric drain to allow 8Fr catheter
| RT13-03-04I |  | 2nd Generation Supraglottic Airway Device with Reinforced Airway Tubing and a Gastric Outlet - Reusable Size: 2.5 For weight: 20-30kg Gastric drain to allow 12Fr catheter
| RT13-03-05I |  | 2nd Generation Supraglottic Airway Device with Reinforced Airway Tubing and a Gastric Outlet - Reusable Size: 3 For weight: 30-50kg Gastric drain to allow 14Fr catheter
| RT13-03-06I |  | 2nd Generation Supraglottic Airway Device with Reinforced Airway Tubing and a Gastric Outlet - Reusable Size: 4 For weight: 50-70kg Gastric drain to allow 14Fr catheter
| RT13-03-07I |  | 2nd Generation Supraglottic Airway Device with Reinforced Airway Tubing and a Gastric Outlet - Reusable Size: 5 For weight: 70-100kg Gastric drain to allow 16Fr catheter
INTUBATING SUPRAGLOTTIC AIRWAY DEVICE (SGAD) - REUSABLE
CATEGORY 3J | RT13-03-01J | INTUBATING SUPRAGLOTTIC AIRWAY DEVICE (SGAD) - REUSABLE Purpose: Preformed Supraglottic airway device that allows intubation in difficult airways Consists of a rigid anatomically curved airway tube with an inflatable elliptical cuff, fitted with a rigid stainless steel handle to facilitate one-handed insertion, removal, and adjustment of the device's position. Must have an epiglottic bar in the mask aperture which elevates as the ETT is passed through, and a ramp which directs the tube centrally and anteriorly to reduce the risk of arytenoid trauma or oesophageal placement. The tube must terminate in a standard 15mm OD connector. Intubating SGAD Size: 3 compatible with ISGAD ETT size 6 - 6.5 Intubating SGAD Size: 4 compatible with ISGAD ETT size 6.5 - 7 Intubating SGAD Size: 5 compatible with ISGAD ETT size 7.5 - 8 Must be manufactured from medical grade silicone and metal. All components must be latex-free .Must comply with ISO 11712 Must be autoclavable – at least 40 times. Recording card to be included | Intubating Supraglottic Airway Device ID: 3 Compatible with Intubating SGAD ETT Size 6 - 6.5mm and Flexible Silicone Stabilizer rod
| RT13-03-02J |  | Intubating Supraglottic Airway Device ID: 4 Compatible with Intubating SGAD ETT Size 6.5-7mm and Flexible Silicone Stabilizer rod
| RT13-03-03J |  | Intubating Supraglottic Airway Device ID: 5 Compatible with Intubating SGAD ETT Size 7.5mm and Flexible Silicone Stabilizer rod
CUFFED REINFORCED ENDOTRACHEAL TUBE FOR AN INTUBATING SGAD - SINGLE USE
CATEGORY 3K | RT13-03-01K | CUFFED REINFORCED ENDOTRACHEAL TUBE FOR AN INTUBATING SGAD - SINGLE USE Purpose: To be used through an Intubating SGAD for airway securement Consists of a 300mm, soft flexible silicone tube with spiral-wound medical grade reinforcing wire. Must be kink resistant and atraumatic to airways Must be marked with depth markings of 1cm intervals, starting from the tip of the tube, tube size, length, ID, OD, trade name. The markings on the tube must be clear and must not rub off. Proximal end: Removable but secure 15mm OD circuit connector made from medical grade plastic Distal end: Bevelled atraumatic tip with a Murphy eye on the opposite side to bevel Must have high-volume-low-pressure polyvinyl chloride cuff with pilot cuff Cuff thickness ± 50 microns. The cuff must inflate and deflate easily and when deflated must conform closely to the tube to minimise the risk of trauma during intubation and extubation. The cuff must be resistant to Nitrous Oxide effects The pilot cuff must have a one-way valve and must indicate the degree of inflation accurately. All components must be latex-free, sterile and individually packed in peel pouch that is easy to open - For SINGLE use only To comply with the latest ISO 5361 | Cuffed Reinforced Endotracheal Tube for an Intubating SGAD - Single Use ID: 6.0mm Compatible with Intubating SGAD Size: 3; 4; 5
| RT13-03-02K |  | Cuffed Reinforced Endotracheal Tube for an Intubating SGAD - Single Use ID: 6.5mm Compatible with Intubating SGAD Size: 3; 4; 5
| RT13-03-03K |  | Cuffed Reinforced Endotracheal Tube for an Intubating SGAD - Single Use ID: 7.0mm Compatible with Intubating SGAD Size: 4; 5
| RT13-03-04K |  | Cuffed Reinforced Endotracheal Tube for an Intubating SGAD - Single Use ID: 7.5mm Compatible with Intubating SGAD Size: 5
| RT13-03-05K |  | Cuffed Reinforced Endotracheal Tube for an Intubating SGAD - Single Use ID: 8mm Compatible with Intubating SGAD Size: 5
STABILIZER ROD FOR INTUBATING SGAD
CATEGORY 3L | RT13-03-01L | Stabilizer Rod for Intubating SGAD Flexible silicone stabilizer rod that is Reusable and autoclavable and compatible with Intubating SGAD Size: 3; 4; 5 The rod must be calibrated markings down the side Individually packed and surgically clean. | Stabilizer Rod for Intubating SGAD Flexible silicone stabilizer rod that is Reusable and autoclavable and compatible with Intubating SGAD Size: 3; 4; 5

## Category 4
BID NUMBER | RT13
BID DESCRIPTION | SUPPLY AND DELIVERY OF RESPIRATORY AID AND ANAESTHETIC ACCESSORIES TO THE STATE
CATEGORY 4
RESCUE AIRWAY DEVICES
| Item Number | Technical Specification | Product Name and Size | Insert inputs on this column
BLIND INSERTION AIRWAY DEVICE - SINGLE-LUMEN DOUBLE-CUFFED:
CATEGORY 4A | RT13-04-01A | BLIND INSERTION AIRWAY DEVICE - SINGLE-LUMEN DOUBLE-CUFFED: Purpose: SINGLE lumen ventilatory tube with an integral drain tube designed for difficult or emergency airway intubation and ventilation without the use of a laryngoscope Consists of a DOUBLE-cuffed SINGLE-lumen tube. Proximal cuff - High volume cuff to seal the oro and nasopharynx and hold the device firmly in place Distal cuff - High volume cuff to occlude the oesophageal inlet and minimises the possibility for gastric ventilation. Must have a precalibrated high-volume syringe to inflate the a SINGLE pilot balloon that inflates both cuffs and acts as a pressure indicator Proximal end must have a 15mm connector There must be ventilation holes situated between the two cuffs for ventilation and that allow for the use of a fibre-optic device or a tube exchanger into the trachea. There must be a separate drain tube for the insertion of a gastric catheter The tube must have a radiopaque line down its full length of the tube. The distal end must sit above oesophageal opening and must be soft and non-traumatic Manufactured from medical grade plastic All components must be latex-free, sterile and individually packed in peel pouch that is easy to open For SINGLE use only Must comply with ISO 11712 or equivalent | Blind Insertion Airway Device: Single-Lumen Double-Cuffed Size: 3 Adult patient's: < 155cm tall
| RT13-04-02A |  | Blind Insertion Airway Device: Single-Lumen Double-Cuffed Size: 4 Adult patient's: 155 - 180cm tall
| RT13-04-03A |  | Blind Insertion Airway Device: Single-Lumen Double-Cuffed Size: 5 Adult patient's: > 180cm tall
BLIND INSERTION AIRWAY DEVICES: DOUBLE LUMEN DOUBLE-CUFFED
CATEGORY 4B | RT13-04-01B | BLIND INSERTION AIRWAY DEVICES: DOUBLE LUMEN DOUBLE-CUFFED Purpose: Dual lumen tube designed for difficult or emergency airway intubation and ventilation without the use of a laryngoscope Consists of a cuffed, DOUBLE-lumen tube, DOUBLE -cuffed with a soft non-traumatic tip. High-volume proximal cuff for oro and nasopharyngeal occlusion. The tube must be held firmly in place by the proximal cuff and it must prevent the escape of gases through the nose or mouth Low-volume distal cuff for oesophageal occlusion Colour coded pilot balloons and 2 syringes prefilled with the correct amount of air for the inflation and deflation of the proximal and distal cuffs Bilateral ventilation ports between the two cuffs The tube must allow for rapid intubation of difficult airways and should be a viable airway whether it is placed in the trachea or in the oesophagus It should also allow for a suction catheter into the oesophagus when the distal cuff is placed in the oesophagus All components must be latex-free sterile and individually packed in peel pouch that is easy to open - For SINGLE use only Must comply with ISO 11712 or equivalent | DOUBLE Lumen Blind insertion airway device – Oesophageal-tracheal DOUBLE lumen airway - Medium 37Fr Patient: 120-160cm tall
| RT13-04-02B |  | DOUBLE Lumen Blind insertion airway device – Oesophageal-tracheal DOUBLE lumen airway - Large 41Fr Patient: >120cm tall
EMERGENCY CRICOTHYROIDOTOMY KIT - WITH VARESS NEEDLE
CATEGORY 4C | RT13-04-01C | Emergency Cricothyrotomy Kit – with Varess needle Purpose: Used for emergency airway access when conventional endotracheal intubation or ventilation can’t be performed Consists of a radiopaque, cuffed, 6mm Cricothyroidotomy tube with a 15mm connector pre-assembled on a varess needle Varess needle to confirm entry into the trachea and indicate any subsequent contact with the posterior tracheal wall. To include 10ml Syringe, Scalpel Size 15, Lubricating Gel, Nylon sutures, HME filter Manufactured from medical grade metal and plastic. All components must be latex-free sterile and individually packed in peel pouch that is easy to open - For SINGLE use only | Emergency Cricothyrotomy Kit – with Varess needle Consists of a radiopaque, cuffed, 6mm Cricothyroidotomy tube with a 15mm connector pre-assembled on a varess needle
EMERGENCY CRICOTHYROTOMY KIT - SELDINGER DILATION TECHNIQUE
CATEGORY 4D | RT13-04-01D | Emergency Cricothyrotomy Kit –Seldinger dilation technique via the cricothyroid membrane Purpose: Used for emergency airway access when conventional endotracheal intubation or ventilation can’t be performed Consists of a radiopaque, cuffed 5mm Cricothyroidotomy tube, 9cm long with 15mm connector Catheter to sit securely on tapered dilator and must be easy to advance over the dilator Short handled scalpel blade; Syringe with introducer needle; Guide-wire and curved dilator with a tapered end Manufactured from medical grade metal and plastic. All components must be latex-free sterile and individually packed in peel pouch that is easy to open - For SINGLE use only | Emergency Cricothyrotomy Kit –Seldinger dilation technique via the cricothyroid membrane Consists of a radiopaque, cuffed 5mm Cricothyroidotomy tube, 9cm long with 15mm connector
EMERGENCY CRICOTHYROTOMY KIT - SCALPEL AND BOUGIE GUIDED TECHNIQUE
CATEGORY 4E | RT13-04-01E | Emergency Cricothyrotomy Kit –Scalpel and Bougie Guided technique Purpose: Used for emergency airway access when conventional endotracheal intubation or ventilation can’t be performed Consists of a radiopaque, cuffed 6mm Cricothyroidotomy tube with 15mm connector To include a short handled scalpel blade Size:10; Syringe 10ml; Catheter mount with angled bronchoscopy port and neck tape Must have a 14Fr Bougie 40cm - with a "hockey stick" 35 - 40 degrees angled tip Manufactured from medical grade metal and plastic. All components must be latex-free sterile and individually packed in peel pouch that is easy to open - For SINGLE use only | Emergency Cricothyrotomy Kit –Scalpel and Bougie Guided technique Consists of a radiopaque, cuffed 6mm Cricothyroidotomy tube with 15mm connector
BLIND INSERTION AIRWAY DEVICE: SINGLE LUMEN DOUBLE-CUFFED
CATEGORY 4F | RT13-04-01F | BLIND INSERTION AIRWAY DEVICE - SINGLE-LUMEN DOUBLE-CUFFED: Purpose: SINGLE lumen ventilatory tube with an integral drain tube designed for difficult or emergency airway intubation and ventilation without the use of a laryngoscope Consists of a DOUBLE-cuffed SINGLE-lumen tube. Proximal cuff - High volume cuff to seal the oro and nasopharynx and hold the device firmly in place Distal cuff - High volume cuff to occlude the oesophageal inlet and minimises the possibility for gastric ventilation. Must have a precalibrated high-volume syringe to inflate the a SINGLE pilot balloon that inflates both cuffs and acts as a pressure indicator Proximal end must have a 15mm connector There must be ventilation holes situated between the two cuffs for ventilation and that allow for the use of a fibre-optic device or a tube exchanger into the trachea. There must be a separate drain tube for the insertion of a gastric catheter The tube must have a radiopaque line down its full length of the tube. The distal end must sit above oesophageal opening and must be soft and non-traumatic Manufactured from medical grade plastic All components must be latex-free, sterile and individually packed in peel pouch that is easy to open For SINGLE use only Must comply with ISO 11712 or equivalent | Blind Insertion Airway Device: SINGLE Lumen Double-Cuffed Size: 3 Adult patient's: < 155cm tall
| RT13-04-02F |  | Blind Insertion Airway Device: SINGLE Lumen Double-Cuffed Size: 4 Adult patient's: 155 - 180cm tall
| RT13-04-03F |  | Blind Insertion Airway Device: SINGLE Lumen Double-Cuffed Size: 5 Adult patient's: > 180cm tall
BLIND INSERTION AIRWAY DEVICE: DOUBLE LUMEN DOUBLE-CUFFED
CATEGORY 4G | RT13-04-01G | BLIND INSERTION AIRWAY DEVICES: DOUBLE LUMEN DOUBLE-CUFFED Purpose: Dual lumen tube designed for difficult or emergency airway intubation and ventilation without the use of a laryngoscope Consists of a cuffed, DOUBLE-lumen tube, DOUBLE -cuffed with a soft non-traumatic tip. High-volume proximal cuff for oro and nasopharyngeal occlusion. The tube must be held firmly in place by the proximal cuff and it must prevent the escape of gases through the nose or mouth Low-volume distal cuff for oesophageal occlusion Colour coded pilot balloons and 2 syringes prefilled with the correct amount of air for the inflation and deflation of the proximal and distal cuffs Bilateral ventilation ports between the two cuffs The tube must allow for rapid intubation of difficult airways and should be a viable airway whether it is placed in the trachea or in the oesophagus It should also allow for a suction catheter into the oesophagus when the distal cuff is placed in the oesophagus All components must be latex-free sterile and individually packed in peel pouch that is easy to open - For SINGLE use only Must comply with ISO 11712 or equivalent | DOUBLE Lumen Blind insertion airway device – Oesophageal-tracheal DOUBLE lumen airway - Medium 37Fr Patient: 120-160cm tall
| RT13-04-02G |  | DOUBLE Lumen Blind insertion airway device – Oesophageal-tracheal DOUBLE lumen airway - Large 41Fr Patient: >120cm tall
EMERGENCY CRICOTHYROTOMY KIT - WITH VARESS NEEDLE
CATEGORY 4H | RT13-04-01H | Emergency Cricothyrotomy Kit – with Varess needle Purpose: Used for emergency airway access when conventional endotracheal intubation or ventilation can’t be performed Consists of a radiopaque, cuffed, 6mm Cricothyroidotomy tube with a 15mm connector pre-assembled on a varess needle Varess needle to confirm entry into the trachea and indicate any subsequent contact with the posterior tracheal wall. To include 10ml Syringe, Scalpel Size 15, Lubricating Gel, Nylon sutures, HME filter Manufactured from medical grade metal and plastic. All components must be latex-free sterile and individually packed in peel pouch that is easy to open - For SINGLE use only | Emergency Cricothyrotomy Kit – with Varess needle Consists of a radiopaque, cuffed, 6mm Cricothyroidotomy tube with a 15mm connector pre-assembled on a varess needle
EMERGENCY CRICOTHYROTOMY KIT - SELDINGER DILATION TECHNIQUE
CATEGORY 4I | RT13-04-01I | Emergency Cricothyrotomy Kit –Seldinger dilation technique via the cricothyroid membrane Purpose: Used for emergency airway access when conventional endotracheal intubation or ventilation can’t be performed Consists of a radiopaque, cuffed 5mm Cricothyroidotomy tube, 9cm long with 15mm connector Catheter to sit securely on tapered dilator and must be easy to advance over the dilator Short handled scalpel blade; Syringe with introducer needle; Guide-wire and curved dilator with a tapered end Manufactured from medical grade metal and plastic. All components must be latex-free sterile and individually packed in peel pouch that is easy to open - For SINGLE use only | Emergency Cricothyrotomy Kit –Seldinger dilation technique via the cricothyroid membrane Consists of a radiopaque, cuffed 5mm Cricothyroidotomy tube, 9cm long with 15mm connector
EMERGENCY CRICOTHYROTOMY KIT - SCALPEL AND BOUGIE GUIDED TECHNIQUE
CATEGORY 4J | RT13-04-01J | Emergency Cricothyrotomy Kit –Scalpel and Bougie Guided technique Purpose: Used for emergency airway access when conventional endotracheal intubation or ventilation can’t be performed Consists of a radiopaque, cuffed 6mm Cricothyroidotomy tube with 15mm connector To include a short handled scalpel blade Size:10; Syringe 10ml; Catheter mount with angled bronchoscopy port and neck tape Must have a 14Fr Bougie 40cm - with a "hockey stick" 35 - 40 degrees angled tip Manufactured from medical grade metal and plastic. All components must be latex-free sterile and individually packed in peel pouch that is easy to open - For SINGLE use only

## Category 5
BID NUMBER | RT13
BID DESCRIPTION | SUPPLY AND DELIVERY OF RESPIRATORY AID AND ANAESTHETIC ACCESSORIES TO THE STATE
CATEGORY 5
INTUBATION AIDS
| Item Number | Technical Specification | Product Name and Size | Insert inputs on this column
INTUBATION STYLET
CATEGORY 5A | RT13-05-01A | INTUBATION STYLET Purpose: To pre-form an endotracheal tube to a desired shape to aid intubation Consists of a malleable core coated and tightly bonded with high density polyurethane. Must have a soft atraumatic tip Must have a U-bend proximally or other mechanism at proximal end to limit depth Latex-free clinically clean and individually packed in peel pouch that is easy to open - For SINGLE use only | Intubation Stylet, Neonatal OD: 2.0mm Length: 225mm To fit ETT's 2.5 - 3.5mm
| RT13-05-02A |  | Intubation Stylet, Paediatric OD: 3.3mm Length: 275mm To fit ETT's 4.0 - 5.5mm
| RT13-05-03A |  | Intubation Stylet, Adult OD: 4.7mm Total Length: 335 - 345mm To fit ETT's > 6.0mm
| RT13-05-04A |  | Intubation Stylet, Adult extra length OD: 4.7mm Total Length: 365mm To fit ETT's > 6.0mm
INTUBATION STYLET - WITH METAL CORE
CATEGORY 5B | RT13-05-01B | INTUBATION STYLET with metal core Purpose: To pre-form an endotracheal tube to a desired shape to aid intubation Consists of a malleable aluminium core coated and tightly bonded with high density polyurethane/ PVC / medical grade plastic/ polyethylene Must have a soft atraumatic tip Must have a U-bend proximally or other mechanism at proximal end to limit depth Latex-free clinically clean and individually packed in peel pouch that is easy to open - For SINGLE use only | Intubation Stylet, Neonatal OD: 2.0mm Length: 225mmTo fit ETT's 2.5 - 3.5mm
| RT13-05-02B |  | Intubation Stylet, Paediatric OD: 3.3mm Length: 275mmTo fit ETT's 4.0 - 5.5mm
| RT13-05-03B |  | Intubation Stylet, Adult OD: 4.7mm Length: 335 - 345mmTo fit ETT's > 6.0mm
| RT13-05-04B |  | Intubation Stylet, Extra-Length Adult OD: 5.0mm Length: 365mm To fit ETT's > 7.5mm
ARTICULATED FLEXIBLE STYLET - REUSABLE
CATEGORY 5C | RT13-05-01C | ARTICULATED FLEXIBLE STYLET - REUSABLE Purpose: To increase the angle of an endotracheal tube during difficult intubations or the use of a D Blade Consists of a stylet with articulating arm, that manoeuvres the tip of an endotracheal tube in a single vertical plane to enable the proper placement of the tube in a difficult airway. Single thumb movement must increase the tube curvature Individually packed and surgical clean in a peel pouch Reusable. | Articulated flexible stylet. Size: 5mm
| RT13-05-02C |  | Articulated flexible stylet. Size: 5.5 - 6.0mm
| RT13-05-03C |  | Articulated flexible stylet. Size: 6.5 - 7.0mm
| RT13-05-04C |  | Articulated flexible stylet. Size: 7.5 - 8.0mm
SOLID INTUBATING BOUGIE - STRAIGHT TIP SINGLE USE
CATEGORY 5C | RT13-05-01C | SOLID INTUBATING BOUGIE - STRAIGHT TIP - Single use Purpose: To assist with endotracheal intubation where visualisation of the glottis is impossible or incomplete and for railroad exchange of endotracheal tubes. Consists of a solid tube manufactured from medical grade plastic that has a non-malleable core that is flexible but kink resistant and will aid intubation Must have an atraumatic rounded tip. Must be graduated in cm marking. Packaged in straight packaging Manufactured from latex-free polyurethane - clinically clean and individually packed in peel pouch that is easy to open | Solid Intubating bougie Straight Tip Single use Neonatal OD: 1.7mm Total Length: 500mm To fit ETT's 2.5 – 4.0mm
| RT13-05-02C |  | Solid Intubating bougie Straight Tip Single use Paediatric OD: 3.3mm Total Length: 600 – 700mm To fit ETT's 4.0 – 6.0mm
| RT13-05-03C |  | Solid Intubating bougie Straight Tip Single use Adult OD: 5.0mm Total Length: 600 – 700mm To fit ETT's > 5.0mm
SOLID INTUBATING BOUGIE - CURVED TIP- SINGLE USE
CATEGORY 5D | RT13-05-01D | SOLID INTUBATING BOUGIE - CURVED-TIP- SINGLE USE Purpose: To assist with endotracheal intubation where visualisation of the glottis is impossible or incomplete and for railroad exchange of endotracheal tubes. Consists of a solid tube manufactured from medical grade PVC that has a non-malleable core that is flexible but kink resistant and will aid intubation. Must have an atraumatic "hockey stick" 35 - 40 degrees angled tip. The tip must clicks on tracheal rings as the bougie is advanced Packaged in straight packaging Manufactured from latex-free polyurethane - clinically clean and individually packed in peel pouch that is easy to open | Solid Intubating bougie Curved tip SINGLE use Neonatal OD: 1.7mm Total Length: 500mm To fit ETT's 2.5 – 4.0mm
| RT13-05-02D |  | Solid Intubating bougie Curved tip SINGLE use Paediatric OD: 3.3mm Total Length: 600 - 700mm To fit ETT's 4.0 – 6.0mm
| RT13-05-03D |  | Solid Intubating bougie Curved tip SINGLE use Adult OD: 5.0mm Total Length: 600 – 700mm To fit ETT's > 5.0mm
SOLID INTUBATING BOUGIE - CURVED TIP- REUSABLE
CATEGORY 5E | RT13-05-01E | SOLID INTUBATING BOUGIE - CURVED-TIP- REUSABLE Purpose: To assist with endotracheal intubation where visualisation of the glottis is impossible or incomplete and for railroad exchange of endotracheal tubes. Consists of a solid tube manufactured from medical grade PVC that has a non-malleable core that is flexible but kink resistant and will aid intubation Must have an atraumatic "hockey stick" 35 - 40 degrees angled tip. The tip must clicks on tracheal rings as the bougie is advanced All components must be latex-free clinically clean and individually packed in peel pouch that is easy to open. Must be autoclavable – at least 40 times | Solid Intubating bougie Curved tip Reusable Neonatal OD: 1.7mm Total Length: 500mm To fit ETT's 2.5 – 4.0mm
| RT13-05-02E |  | Solid Intubating bougie Curved tip Reusable Paediatric OD: 3.3mm Total Length: 600 - 700mm To fit ETT's 4.0 – 6.0mm
| RT13-05-03E |  | Solid Intubating bougie Curved tip Reusable Adult OD: 5.0mm Total Length: 600 – 700mm To fit ETT's > 5.0mm
AIRWAY EXCHANGE CATHETER
CATEGORY 5F | RT13-05-01F | AIRWAY EXCHANGE CATHETER Purpose: For simple, atraumatic endotracheal tube exchange whilst providing oxygenation Consists of a catheter manufactured from radiopaque medical grade plastic Must have a hollow lumen for constant oxygenation of the patient, with distal side ports to ensure adequate airflow Must be supplied with connection and adaptor for attachment of a ventilatory device. Must have a blunt tip for atraumatic intubation, centimetre markings for accurate tube placement. Latex-free, sterile and individually packed in peel pouch that is easy to open - For SINGLE use only | Airway Exchange Catheter. Infant 8Fr Length: ± 450mm To fit ETT's Size: 3.0mm
| RT13-05-02F |  | Airway Exchange Catheter. Paediatric 11Fr Length: ± 830mm To fit ETT's Size: 4.0mm
| RT13-05-03F |  | Airway Exchange Catheter. Adult 14Fr Length: ± 830mm To fit ETT's Size: 5.0mm
| RT13-05-04F |  | Airway Exchange Catheter. Adult 19Fr Length: ± 830mm To fit ETT's Size: 7.0mm
AIRWAY EXCHANGE CATHETER FOR FIBRE-OPTIC GUIDED INTUBATION
CATEGORY 5G | RT13-05-01G | AIRWAY EXCHANGE CATHETER FOR FIBRE-OPTIC GUIDED INTUBATION Purpose: For simple, atraumatic fibre-optic guided intubation while providing oxygenation. Consists of a large diameter hollow lumen catheter that fits over a Size 4.2 fibre-optic flexible bronchoscope Manufactured from radiopaque medical grade plastic, with distal side ports to ensure adequate airflow Must be supplied with connection and adaptor for attachment of a ventilatory device. Must have a blunt tip slightly tapered tip to fit snugly on the bronchoscope. Centimetre markings from 5cm in 1cm increments Preference will be given to items with a distal side hole for ventilation Latex-free, sterile and individually packed in peel pouch that is easy to open - For SINGLE use only | Airway Exchange Catheter for Fibre-Optic Guided Intubation Size: 19Fr ID: 4.7mm OD: < 5mm Length: 560mm To fit ETT's 7.0mm
AIRWAY EXCHANGE CATHETER FOR FIBRE-OPTIC GUIDED INTUBATION
CATEGORY 5H | RT13-05-01H | Airway Exchange Catheter Kit with Guidewire for staged extubation. Purpose: For staged or difficult extubation of patient's with cardiorespiratory dysfunction and potentially difficult reintubation Consists of an exchange catheter, extra stiff guidewire, Bronchoscopic port, 2 Adaptors 1 x luer lock, 15mm connector Catheter 14Fr ID: 2.3mm OD: < 4mm Catheter Length: 700 - 900mm Guidewire: Diameter: 0.1cm Length: 160cm The wire guide to fit into the working channel of a bronchoscope. The catheter must have distal side ports to allow for airflow. Must have an airway adaptor to allow for ventilation To fit an endotracheal tube of Size 5 or larger Latex-free medical grade metal and plastic, sterile and individually packed in peel pouch that is easy to open - For SINGLE use only | Airway Exchange Catheter Kit with Guidewire for staged extubation.Catheter 14Fr ID: 2.3mm OD: < 4mm Catheter Length: 700 - 900mm
RETROGRADE INTUBATION KIT
CATEGORY 5I | RT13-05-01I | RETROGRADE INTUBATION KIT 11Fr Length: 70cm for ETT's ID: 4mm or larger Purpose: Used to assist in the placement of an endotracheal tube using retrograde techniques during a difficult airway access procedure. Consists of an Introducer needle/ Catheter introducer needle:18G and Guide Wire: Size: 0.038mm Length 110cm Guidewire made from stainless steel with two positioning marks 20cm from each end; atraumatic tips; one end straight and the other a flexible J tip. radiopaque catheter with distal side ports and a 15mm connector proximally for ventilatory devices. Luer slip syringe to be included. Seldinger technique access to the airway via the cricothyroid membrane. Retrograde (cephalad) advancement of wire towards the oral/ nasal cavities. Antegrade advancement of a hollow catheter with central & peripheral side ports for optimal ventilation. Latex-free, sterile and individually packed in peel pouch that is easy to open - For SINGLE use only | RETROGRADE INTUBATION KIT 11Fr Length: 70cm for ETT's ID: 4mm or larger
LARYNGO-TRACHEAL ATOMISER
CATEGORY 5J | RT13-05-01J | LARYNGO-TRACHEAL ATOMISER STYLET Purpose: For the delivery of local anaesthetic agents by atomisation to anaesthetise upper and lower airway Consists of a long narrow malleable stylet with a tip that atomises drugs into a fine mist of particles 30-100microns The distal end must have multiple orifices around the tube (360°) and the stylet must be able to conform to the patient's anatomy. The proximal end must be compatible with a luer lock or luer slip syringe and must allow for high pressure to be applied through the stylet. All components must be latex-free, clinically clean and individually packed in peel pouch that is easy to open - For SINGLE use only | Laryngo-tracheal Atomiser Stylet Size: Paediatric
Laryngo-tracheal Atomiser Stylet Size: Adult
LARYNGO-TRACHEAL ATOMISER KIT
CATEGORY 5K | RT13-05-01K | LARYNGO-TRACHEAL ATOMISER KIT Purpose: For the delivery of local anaesthetic agents by atomisation to anaesthetise upper and lower airway Consists of a long narrow malleable stylet with a tip that atomises drugs into a fine mist of particles 30-100microns The distal end must have multiple orifices around the tube (360°) and the stylet must be able to conform to the patient's anatomy. The proximal end must be compatible with a luer lock or luer slip syringe and must allow for high pressure to be applied through the stylet. Must contain compatible 2ml syringe All components must be latex-free, clinically clean and individually packed in peel pouch that is easy to open - For SINGLE use only | Laryngo-tracheal Atomiser Kit Size: Paediatric with 2ml syringe
| RT13-05-02K |  | Laryngo-tracheal Atomiser Size: Adult with 2ml syringe
INTRANASAL ATOMISER
CATEGORY 5J | RT13-05-01J | INTRANASAL ATOMISER Purpose: For the delivery of atomised drugs to the nasal mucosa. Consists of a sturdy adaptor connected to a 3ml luer lock syringe with a soft atraumatic conical nasal plug. The soft conical plug must prevent leakage out of the nostril and must deliver atomised drugs into a fine mist with particles 30-100microns in size 3ml Luer Lock syringe must connect tightly and allow for high pressures. All components must be latex-free, clinically clean and individually packed in peel pouch that is easy to open - For SINGLE use only | INTRANASAL ATOMISER
OROPHARYNGEAL AIRWAY ATOMISER
CATEGORY 5K | RT13-05-01K | OROPHARYNGEAL AIRWAY ATOMISER Purpose: An adult oropharyngeal airway with integral atomiser to topicalize airways while facilitating oxygenation. Must have two ports connected to the distal end of the oropharyngeal airway. One port for atomisation - must connect to a 2ml or 5ml syringe and allow for atomization with a particle size of 30-100microns The 2nd port must have a universal connector that connects to standard oxygen tubing. Manufactured from medical grade plastic, and must be latex-free and DEHP-free. Single use, individually packed and surgical clean. For airways requiring a fiberoptic scope Delivers topical anesthetic and oxygenation | OROPHARYNGEAL AIRWAY ATOMISER

## Category 6
BID NUMBER | RT13
BID DESCRIPTION | SUPPLY AND DELIVERY OF RESPIRATORY AID AND ANAESTHETIC ACCESSORIES TO THE STATE
CATEGORY 6
CONNECTORS AND CATHETER MOUNTS
| Item Number | Technical Specification | Product Name and Size | Insert inputs on this column
Anaesthetic Circuit Elbow (angled) Connector
CATEGORY 6A | RT13-06-01A | ANAESTHETIC ELBOW (angled) CONNECTORS Consists of a connector made from latex-free, non-pyrogenic, hypoallergenic, DEHP free, transparent, non-conductive medical grade plastic. All components latex-free, clinically clean and individually packed in peel pouch that is easy to open - For SINGLE use only To comply with ISO 5367 | Anaesthetic circuit elbow (angled) connector 22mm male/ 15mm female (Patient end) 15mm male (Circuit end)
| RT13-06-02A |  | Anaesthetic circuit elbow (angled) connector 22mm male/ 15mm female (Patient end) 22mm female (Circuit end)
| RT13-06-03A |  | Anaesthetic circuit elbow (angled) connector, with toggled, capped luer lock sample port for Capnography: 22mm male/ 15mm female (Patient end) 15mm male (Circuit end)
| RT13-06-04A |  | Anaesthetic circuit elbow (angled) connector, with toggled, capped luer lock sample port for Capnography: 22mm male/ 15mm female (Patient end) 22mm female (Circuit end)
| RT13-06-05A |  | Anaesthetic DOUBLE swivel elbow connector with 360° rotation, with suction port and toggled cap: 22mm male/ 15mm female (Patient end) 15mm male (Circuit end)
| RT13-06-06A |  | Anaesthetic DOUBLE swivel elbow connector with 360° rotation, and a sealable bronchoscopy port and toggled cap Must have a diaphragm which hugs fibre-optic bronchoscope 22mm male/ 15mm female (Patient end) 15mm male (Circuit end)
ANAESTHETIC STRAIGHT CONNECTORS
CATEGORY 6B | RT13-06-01B | ANAESTHETIC STRAIGHT CONNECTORS Consists of a connector made from latex-free, non-pyrogenic, hypoallergenic, DEHP free, transparent, non-conductive medical grade plastic. All components latex-free, clinically clean and individually packed in peel pouch that is easy to open - For SINGLE use only To comply with ISO 5367 or equivalent | Anaesthetic circuit straight connector 15mm male to 15mm male
| RT13-06-02B |  | Anaesthetic circuit straight connector 15mm male to 15mm female
| RT13-06-03B |  | Anaesthetic circuit straight connector 15mm female/22mm male to 22mm male
| RT13-06-04B |  | Anaesthetic circuit straight connector 15mm female/22mm male to 22mm female
| RT13-06-05B |  | Anaesthetic circuit straight connector 22mm male to 22mm male
| RT13-06-06B |  | Anaesthetic circuit straight connector 22mm male to 22mm female
ANAESTHETIC T-PIECE CONNECTORS
CATEGORY 6C | RT13-06-01C | ANAESTHETIC T-PIECE CONNECTORS Consists of a connector made from latex-free, non-pyrogenic, hypoallergenic, DEHP free, transparent, non-conductive medical grade plastic. All components latex-free, clinically clean and individually packed in peel pouch that is easy to open - For SINGLE use only To comply with ISO 5367 or equivalent | Anaesthetic T-Piece connectors: T - opening (90°) 15mm Female; Straight opening: 15mm Male to 15mm Female
| RT13-06-02C |  | Anaesthetic T-Piece connectors: T - opening (90°) 22mm Male/15 Female; Straight opening: 22mm Male to 22mm Male with toggled 7.6mm fibreoptic/suction port on straight part
| RT13-06-03C |  | Anaesthetic T-Piece connectors: T - opening (90°) 22mm Male/15 Female; Straight opening: 22mm Male to 22mm Male with toggled 7.6mm fibreoptic/suction port on straight part
SCAVENGING CONNECTORS
CATEGORY 6D | RT13-06-01D | CONNECTORS and OXYGEN DELIVERY ACCESSORIES Consists of a connector made from latex-free, non-pyrogenic, hypoallergenic, DEHP free, transparent, non-conductive medical grade plastic. All components latex-free, clinically clean and individually packed in peel pouch that is easy to open - For SINGLE use only To comply with ISO 5367 or equivalent | Scavenging connector straight connector 19mm male/22mm male Pendant scavenger side
| RT13-06-02D |  | Scavenging connector straight connector 22mm male to 30mm female
| RT13-06-03D |  | Scavenging connector straight connector 22mm male to 30mm male
| RT13-06-04D |  | Scavenging connector straight connector 30mm male to 30mm male
| RT13-06-05D |  | Scavenging connector straight connector 30mm male to 30mm female
AEROSOL STRAIGHT CONNECTOR
CATEGORY 6E | RT13-06-01E | AEROSOL STRAIGHT CONNECTOR with toggled port to fit nebulising canister: 15mm Male to 15mm Female Consists of a connector made from latex-free, non-pyrogenic, hypoallergenic, DEHP free, transparent, non-conductive medical grade plastic. All components latex-free, clinically clean and individually packed in peel pouch that is easy to open - For SINGLE use only To comply with ISO 5367 or equivalent | AEROSOL STRAIGHT CONNECTOR with toggled port to fit nebulising canister: 15mm Male to 15mm Female
NIPPLE NUT-AND-STEM “CHRISTMAS TREE” OXYGEN ADAPTER
CATEGORY 6F | RT13-06-01F | NIPPLE NUT-AND-STEM “CHRISTMAS TREE” OXYGEN ADAPTER Purpose: Connector between standard flow meter DISS connection and oxygen tubing Wing-nut to fit standard DISS connection on oxygen flow meter. Thread to be secure without leaking. Tapered stem to allow easy connection of standard bore oxygen tubing, with deep grooved barbs on the stem to prevent disconnection of tubing Manufactured from latex free medical grade plastic/ medical grade metal To comply with ISO 5367 or equivalent Clinically clean packaging | Nipple Nut-and-stem “Christmas tree” oxygen adapter - Plastic
| RT13-06-02F |  | Nipple Nut-and-stem “Christmas tree” oxygen adapter - Metal
CATHETER MOUNT - STRAIGHT
CATEGORY 6G | RT13-06-01G | CATHETER MOUNT - STRAIGHT CONNECTOR Purpose: Connector to join an endotracheal or tracheostomy tube to a breathing circuit Consists of flexible, low-flow smooth-bore tubing with an external coil to prevent kinking.. All components latex-free, sterile and individually packed in peel pouch that is easy to open - For SINGLE use only | Catheter mount with Straight Connector- Neonatal Length ± 90mm Connectors: 22mm Male/ 15mm Female (patient end); 15mm Male (machine end) Volume: < 15ml
| RT13-06-02G |  | Catheter mount (15mm) with Straight Connector- Paediatric Length ± 90mm Connectors: 22mm Male/ 15mm Female (patient end); 15mm Male (machine end) Volume: < 20ml
| RT13-06-03G |  | Catheter mount (22mm) with Straight Connector- Adult Length ± 150mm Connectors: 22mm Male/ 15mm Female (patient end); 22mm Female (machine end)
CATHETER MOUNT WITH ELBOW CONNECTOR, LUER LOCK SAMPLE PORT AND TOGGLED CAP
CATEGORY 6H | RT13-06-01H | CATHETER MOUNT WITH ELBOW CONNECTOR, LUER LOCK SAMPLE PORT AND TOGGLED CAP Purpose: Connector to join an endotracheal or tracheostomy tube to a breathing circuit with sample port for Capnography Consists of flexible, low-flow smooth-bore tubing with an external coil to prevent kinks. Elbow connector with a Luer lock sample port and toggled cap All components latex-free, sterile and individually packed in peel pouch that is easy to open - For SINGLE use only | Catheter mount with Elbow connector, C02 sampling port and toggled cap - Neonatal Volume: ± 10-16ml Length ± 90mm Connector: 22mm Male/ 15mm Female (patient end); 15mm Male (machine end)
| RT13-06-02H |  | Catheter mount with Elbow connector, C02 sampling port and toggled cap - Paediatric Volume: ± 15-30ml Length ± 90mm Connectors: 22mm Male/ 15mm Female (patient end); 22mm Female (machine end)
| RT13-06-03H |  | Catheter mount with Elbow connector, C02 sampling port and toggled cap - Adult Length ± 120-150mm Connectors: 22mm Male/ 15mm Female (patient end); 22mm Female (machine end)
CATHETER MOUNT WITH DOUBLE SWIVEL ELBOW CONNECTOR, BRONCHOSCOPY/ASPIRATION PORT AND TOGGLED CAP
CATEGORY 6I | RT13-06-01I | CATHETER MOUNT WITH DOUBLE SWIVEL ELBOW CONNECTOR, BRONCHOSCOPY/ASPIRATION PORT AND TOGGLED CAP Purpose: Connector to join an endotracheal or tracheostomy tube to a breathing circuit with a bronchoscopy/aspiration port. Consists of flexible, low-flow smooth-bore tubing with an external coil to prevent kinks. Double swivel Elbow connector with a Bronchoscopy/Aspiration Port. Must have a diaphragm which hugs fibre-optic bronchoscope. All components latex-free, sterile and individually packed in peel pouch that is easy to open - For SINGLE use only | Catheter mount with DOUBLE swivel elbow connector, Bronchoscopy/Aspiration Port and toggled cap - Neonatal Volume: ± 10-16ml Length ± 90mm Connectors: 22mm Male/ 15mm Female (patient end); 15mm Male (machine end)
| RT13-06-02I |  | Catheter mount with DOUBLE swivel elbow connector, Bronchoscopy/Aspiration Port and toggled cap - Paediatric Volume: ± 15-30ml Length ± 90mm Connectors: 22mm Male/ 15mm Female (patient end); 22mm Female (Machine end).
| RT13-06-03I |  | Catheter mount with DOUBLE swivel elbow connector, Bronchoscopy/Aspiration Port and toggled cap - Adult Length: 120-150mm Connectors: 22mm Male/ 15mm Female to patient and 22mm Female to machine
CATHETER MOUNT - EXTENDIBLE WITH ELBOW CONNECTOR, LUER LOCK SAMPLE PORT AND TOGGLED CAP
CATEGORY 6J | RT13-06-01J | CATHETER MOUNT - Extendible with Elbow Connector, Luer Lock Sample Port and Toggled Cap Purpose: Connector to join an endotracheal or tracheostomy tube to a breathing circuit with sample port for Capnography. Consists of flexible, extendible tubing. The internal diameter must remain constant in every position. Must have an Elbow connector with a Luer lock sample port and toggled cap All components latex-free, sterile and individually packed in peel pouch that is easy to open - For SINGLE use only | Catheter mount: Extendible with elbow connector, Luer Lock Sample Port and Toggled Cap - Premature Volume: ± 10-16ml Length ± 45mm - 100mm Connectors: 8.5mm Female (patient end); 15mm Male (machine end)
| RT13-06-02J |  | Catheter mount: Extendible with connector, Luer Lock Sample Port and Toggled Cap - Neonatal Volume: ± 10-16ml Length ± 45mm - 100mm Connectors: 15mm Female (patient end) ; 15mm Male (machine end)
| RT13-06-03J |  | Catheter mount: Extendible connector, Luer Lock Sample Port and Toggled Cap - Paediatric Volume: ± 15-30ml Length ± 45mm - 100mm Connectors: 22mm Male/ 15mm Female (patient end) 22mm Female (machine end)
CATHETER MOUNT - EXTENDIBLE WITH DOUBLE SWIVEL ELBOW CONNECTOR, BRONCHOSCOPY/ASPIRATION PORT AND TOGGLED CAP
CATEGORY 6K | RT13-06-01K | CATHETER MOUNT - Extendible with DOUBLE swivel elbow connector, Bronchoscopy/Aspiration Port and toggled cap Purpose: Connector to join an endotracheal or tracheostomy tube to a breathing circuit with with a bronchoscopy/aspiration port. Consists of flexible, extendible tubing. The internal diameter must remain constant in every position. Double swivel Elbow connector with a Bronchoscopy/Aspiration Port. Must have a diaphragm which hugs fibre-optic bronchoscope All components latex-free, sterile and individually packed in peel pouch that is easy to open - For SINGLE use only | Catheter mount: Extendible with DOUBLE swivel elbow connector, Bronchoscopy/Aspiration Port and toggled cap - Premature Volume: ± 10-16ml Length ± 45mm - 100mm Connectors: 8.5mm Female (patient end); 15mm Male (machine end)
| RT13-06-02K |  | Catheter mount: Extendible with DOUBLE swivel elbow connector, Bronchoscopy/Aspiration Port and toggled cap - Neonatal Volume: ± 10-16ml Length ± 45mm - 100mm Connectors: 15mm Female (patient end) ; 15mm Male (machine end)
| RT13-06-03K |  | Catheter mount: Extendible with DOUBLE swivel elbow connector, Bronchoscopy/Aspiration Port and toggled cap - Paediatric Volume: ± 15-30ml Length ± 45mm - 100mm Connectors: 22mm Male/ 15mm Female (patient end) 22mm Female (machine end)
| RT13-06-04K |  | Catheter mount: Extendible with DOUBLE swivel elbow connector, Bronchoscopy/Aspiration Port and toggled cap - Adult Length ± 70mm - 150mm Connectors: 22mm Male/ 15mm Female (patient end) 22mm Female (machine end)

## Category 7
BID NUMBER | RT13
BID DESCRIPTION | SUPPLY AND DELIVERY OF RESPIRATORY AID AND ANAESTHETIC ACCESSORIES TO THE STATE
CATEGORY 7
CLOSED CATHETER SUCTION SYSTEMS
| Item Number | Technical Specification | Product Name and Size | Insert inputs on this column
CLOSED SUCTION CATHETER - PAEDIATRIC ENDOTRACHEAL Y-TYPE 48HR
CATEGORY 7A | RT13-07-01A | CLOSED SUCTION CATHETER - PAEDIATRIC ENDOTRACHEAL Y-TYPE 48HR Purpose: Closed system for respiratory suctioning through endotracheal tubes without disconnecting or interrupting ventilation. Consists of a closed suction catheter with an atraumatic tip. The catheter must have markings along the length to indicate depth of suctioning and a radiopaque line on the catheter tip. While connected but not in suction mode the catheter tip must not extend beyond the sheath or impede airflow. The suction catheter must be enclosed within a durable, flexible, plastic sheath for storage to avoid contamination. The transparent sheath must allow visualisation of markings on the suction catheter and must connect to the endotracheal tube via a clear, firmly-fitting, plastic connector or adaptor. The sheath and connector must allow for smooth advancement of catheter. The system must seal the airway and catheter from atmospheric contamination and allow uninterrupted positive pressure ventilation throughout the suction process. An effective seal around the catheter must prevent ventilation gas entering the lumen of the sheath. The system must protect the user from contamination by secretions and suctioned material. Proximal end of suction catheter must have clear chamber, so that suctioned contents can be visualised and must have lockable suction valve with toggled capped suction tubing connector. Paediatric Catheters must have a Y- connector (without locking mechanism) with an appropriate range of diameters for each size catheter. The connector must allow variations in patient positions without causing torque, twisting, or obstruction of the tracheal tube, and must not cause tissue or skin pressure. The system must allow for irrigation via a lateral irrigation port which must be a capped with a one-way valve that connects to a standard luer-slip syringe. The irrigation port must be situated at the patient end of the system for catheters and at the suction end of the catheter for catheters.made from medically safe PVC The suction end of the system must be fitted with capped, lockable thumb valve to control suctioning. It must have a tapered connection to fit a range of suction tubing. The system must be as light in weight as possible. The system is for single patient use for a specified 48 or 72 hours. Must have coloured day sticker labels to facilitate regular changing of the catheter system. Made from medically safe PVC. Sterile and individually packed in easy to open pouch with one transparent side; For SINGLE use only The system must conform to ISO 8836:2019. | Closed Suction Catheter - Paediatric Endotracheal Y-Type 48Hr 5Fr OD: 1.66mm For ETT ID: 2.5mm Y connectors: 2.0mm, 2.5mm
| RT13-07-02A |  | Closed Suction Catheter - Paediatric Endotracheal Y-Type 48Hr 6Fr OD: 2.0mm For ETT ID: 3mm Y connectors: 2.5mm, 3.0mm, 3.5mm
| RT13-07-03A |  | Closed Suction Catheter - Paediatric Endotracheal Y-Type 48Hr 7Fr OD: 2.3mm For ETT ID: 3.5mm Y connectors: 3.0mm, 3.5mm, 4mm
| RT13-07-04A |  | Closed Suction Catheter - Paediatric Endotracheal Y-Type 48Hr 8Fr OD: 2.6mm For ETT with ID: 4mm Y connectors: 3.5mm, 4mm, 4.5mm
| RT13-07-05A |  | Closed Suction Catheter - Paediatric Endotracheal Y-Type 48Hr 10Fr OD: 3.3mm For ETT with ID: 5.0mm Y connectors: 4.5mm, 5mm, 5.5mm
| RT13-07-06A |  | Closed Suction Catheter - Paediatric Endotracheal Y-Type 48Hr 12Fr OD: 4.0mm For ETT with ID: 6.0mm Y connectors: 5.5mm, 6mm, 6.5mm
CLOSED SUCTION CATHETER - PAEDIATRIC ENDOTRACHEAL Y-TYPE 72HR
CATEGORY 7B | RT13-07-01B | CLOSED SUCTION CATHETER - PAEDIATRIC ENDOTRACHEAL Y-TYPE 72HR Purpose: Closed system for respiratory suctioning through endotracheal tubes without disconnecting or interrupting ventilation. Consists of a closed suction catheter with an atraumatic tip. The catheter must have markings along the length to indicate depth of suctioning and a radiopaque line on the catheter tip. While connected but not in suction mode the catheter tip must not extend beyond the sheath or impede airflow. The suction catheter must be enclosed within a durable, flexible, plastic sheath for storage to avoid contamination. The transparent sheath must allow visualisation of markings on the suction catheter and must connect to the endotracheal tube via a clear, firmly-fitting, plastic connector or adaptor. The sheath and connector must allow for smooth advancement of catheter. The system must seal the airway and catheter from atmospheric contamination and allow uninterrupted positive pressure ventilation throughout the suction process. An effective seal around the catheter must prevent ventilation gas entering the lumen of the sheath. The system must protect the user from contamination by secretions and suctioned material. Proximal end of suction catheter must have clear chamber, so that suctioned contents can be visualised and must have lockable suction valve with toggled capped suction tubing connector. Paediatric Catheters must have a Y- connector (without locking mechanism) with an appropriate range of diameters for each size catheter. The connector must allow variations in patient positions without causing torque, twisting, or obstruction of the tracheal tube, and must not cause tissue or skin pressure. The system must allow for irrigation via a lateral irrigation port which must be a capped with a one-way valve that connects to a standard luer-slip syringe. The irrigation port must be situated at the patient end of the system for catheters and at the suction end of the catheter for catheters.made from medically safe PVC The suction end of the system must be fitted with capped, lockable thumb valve to control suctioning. It must have a tapered connection to fit a range of suction tubing. The system must be as light in weight as possible. The system is for single patient use for a specified 48 or 72 hours. Must have coloured day sticker labels to facilitate regular changing of the catheter system. Made from medically safe PVC. Sterile and individually packed in easy to open pouch with one transparent side; For SINGLE use only The system must conform to ISO 8836:2019. | Closed Suction Catheter - Paediatric Endotracheal Y-Type 72Hr 5Fr OD: 1.66mm For ETT ID: 2.5mm Y connectors: 2.0mm, 2.5mm
| RT13-07-02B |  | Closed Suction Catheter - Paediatric Endotracheal Y-Type 72Hr 6Fr OD: 2.0mm For ETT ID: 3mm Y connectors: 2.5mm, 3.0mm, 3.5mm
| RT13-07-03B |  | Closed Suction Catheter - Paediatric Endotracheal Y-Type 72Hr 7Fr OD: 2.3mm For ETT ID: 3.5mm Y connectors: 3.0mm, 3.5mm, 4mm
| RT13-07-04B |  | Closed Suction Catheter - Paediatric Endotracheal Y-Type 72Hr 8Fr OD: 2.6mm For ETT with ID: 4mm Y connectors: 3.5mm, 4mm, 4.5mm
| RT13-07-05B |  | Closed Suction Catheter - Paediatric Endotracheal Y-Type 72Hr 10Fr OD: 3.3mm For ETT with ID: 5.0mm Y connectors: 4.5mm, 5mm, 5.5mm
| RT13-07-06B |  | Closed Suction Catheter - Paediatric Endotracheal Y-Type 72Hr 12Fr OD: 4.0mm For ETT with ID: 6.0mm Y connectors: 5.5mm, 6mm, 6.5mm
CLOSED SUCTION CATHETER ADULT ENDOTRACHEAL - 48HRS
CATEGORY 7C | RT13-07-01C | CLOSED SUCTION CATHETER ADULT ENDOTRACHEAL - 48HRS Purpose: Closed system for respiratory suctioning through endotracheal tubes without disconnecting or interrupting ventilation. Consists of a closed suction catheter with an atraumatic tip. The catheter must have markings along the length to indicate depth of suctioning and a radiopaque line on the catheter tip. While connected but not in suction mode the catheter tip must not extend beyond the sheath or impede airflow. The suction catheter must be enclosed within a durable, flexible, plastic sheath for storage to avoid contamination. The transparent sheath must allow visualisation of markings on the suction catheter and must connect to the endotracheal tube via a clear, firmly-fitting, plastic connector or adaptor. The sheath and connector must allow for smooth advancement of catheter. The system must seal the airway and catheter from atmospheric contamination and allow uninterrupted positive pressure ventilation throughout the suction process. An effective seal around the catheter must prevent ventilation gas entering the lumen of the sheath. The system must protect the user from contamination by secretions and suctioned material. Proximal end of suction catheter must have clear chamber, so that suctioned contents can be visualised and must have lockable suction valve with capped suction tubing connector. The suction catheter must connect onto the tube with a DOUBLE swivel connector or a double-swivel Y-connector with standard 15 mm OD taper fittings. This connection must be secure and accidental disconnection must be minimal The whole device must be as lightweight as possible. Must include a light, straight, flexible, clear, concertina catheter-mount (15x22 mm connections) tube for connection to the breathing circuit The connector must allow variations in patient positions without causing torque, twisting, or obstruction of the tracheal tube, and must not cause tissue or skin pressure. The system must allow for irrigation via a lateral irrigation port which must be a capped with a one-way valve that connects to a standard luer-slip syringe. The irrigation port must be situated at the patient end of the system for catheters and at the suction end of the catheter for catheters.made from medically safe PVC The suction end of the system must be fitted with capped, lockable thumb valve to control suctioning. It must have a tapered connection to fit a range of suction tubing. The system is for single patient use for a specified 48 or 72 hours. Must have coloured day sticker labels to facilitate regular changing of the catheter system. Made from medically safe PVC. Sterile and individually packed in easy to open pouch with one transparent side; For SINGLE use only The system must conform to ISO 8836:2019. | Closed Suction Catheter - Adult Endotracheal 48Hr 10Fr OD: 3.3mm Length range: 520 - 600mm
| RT13-07-02C |  | Closed Suction Catheter - Adult Endotracheal 48Hr 12Fr OD: 4.0mm Length range: 520 - 600mm
| RT13-07-03C |  | Closed Suction Catheter - Adult Endotracheal 48Hr 14Fr OD: 4.7mm Length range: 520 - 600mm
| RT13-07-04C |  | Closed Suction Catheter - Adult Endotracheal 48Hr 16Fr OD: 5.3mm Length range: 520 - 600mm
CLOSED SUCTION CATHETER ADULT ENDOTRACHEAL - 72HRS
CATEGORY 7D | RT13-07-01D | CLOSED SUCTION CATHETER ADULT ENDOTRACHEAL - 72HRS Purpose: Closed system for respiratory suctioning through endotracheal tubes without disconnecting or interrupting ventilation. Consists of a closed suction catheter with an atraumatic tip. The catheter must have markings along the length to indicate depth of suctioning and a radiopaque line on the catheter tip. While connected but not in suction mode the catheter tip must not extend beyond the sheath or impede airflow. The suction catheter must be enclosed within a durable, flexible, plastic sheath for storage to avoid contamination. The transparent sheath must allow visualisation of markings on the suction catheter and must connect to the endotracheal tube via a clear, firmly-fitting, plastic connector or adaptor. The sheath and connector must allow for smooth advancement of catheter. The system must seal the airway and catheter from atmospheric contamination and allow uninterrupted positive pressure ventilation throughout the suction process. An effective seal around the catheter must prevent ventilation gas entering the lumen of the sheath. The system must protect the user from contamination by secretions and suctioned material. Proximal end of suction catheter must have clear chamber, so that suctioned contents can be visualised and must have lockable suction valve with capped suction tubing connector. The suction catheter must connect onto the tube with a DOUBLE swivel connector or a double-swivel Y-connector with standard 15 mm OD taper fittings. This connection must be secure and accidental disconnection must be minimal The whole device must be as lightweight as possible. Must include a light, straight, flexible, clear, concertina catheter-mount (15x22 mm connections) tube for connection to the breathing circuit The connector must allow variations in patient positions without causing torque, twisting, or obstruction of the tracheal tube, and must not cause tissue or skin pressure. The system must allow for irrigation via a lateral irrigation port which must be a capped with a one-way valve that connects to a standard luer-slip syringe. The irrigation port must be situated at the patient end of the system for catheters and at the suction end of the catheter for catheters.made from medically safe PVC The suction end of the system must be fitted with capped, lockable thumb valve to control suctioning. It must have a tapered connection to fit a range of suction tubing. The system is for single patient use for a specified 48 or 72 hours. Must have coloured day sticker labels to facilitate regular changing of the catheter system. Made from medically safe PVC. Sterile and individually packed in easy to open pouch with one transparent side; For SINGLE use only The system must conform to ISO 8836:2019. | Closed Suction Catheter - Adult Endotracheal 72Hr 10Fr OD: 3.3mm Length range: 520 - 600mm
| RT13-07-02D |  | Closed Suction Catheter - Adult Endotracheal 72Hr 12Fr OD: 4.0mm Length range: 520 - 600mm
| RT13-07-03D |  | Closed Suction Catheter - Adult Endotracheal 72Hr 14Fr OD: 4.7mm Length range: 520 - 600mm
| RT13-07-04D |  | Closed Suction Catheter - Adult Endotracheal 72Hr 16Fr OD: 5.3mm Length range: 520 - 600mm
CLOSED SUCTION CATHETER - PAEDIATRIC TRACHEOSTOMY Y-TYPE 48HR
CATEGORY 7E | RT13-07-01E | CLOSED SUCTION CATHETER - PAEDIATRIC TRACHEOSTOMY Y-TYPE 48HR Purpose: Closed system for respiratory suctioning through tracheostomy tubes without disconnecting or interrupting ventilation. Consists of a closed suction catheter with an atraumatic tip. The catheter must have markings along the length to indicate depth of suctioning and a radiopaque line on the catheter tip. While connected but not in suction mode the catheter tip must not extend beyond the sheath or impede airflow. The suction catheter must be enclosed within a durable, flexible, plastic sheath for storage to avoid contamination. The transparent sheath must allow visualisation of markings on the suction catheter and must connect to the tracheal tube via a clear, firmly-fitting, plastic connector or adaptor. The sheath and connector must allow for smooth advancement of catheter. The system must seal the airway and catheter from atmospheric contamination and allow uninterrupted positive pressure ventilation throughout the suction process. An effective seal around the catheter must prevent ventilation gas entering the lumen of the sheath. The system must protect the user from contamination by secretions and suctioned material. Proximal end of suction catheter must have clear chamber, so that suctioned contents can be visualised and must have lockable suction valve with capped suction tubing connector. Paediatric Catheters must have a Y- connector (without locking mechanism) with an appropriate range of diameters for each size catheter. The connector must allow variations in patient positions without causing torque, twisting, or obstruction of the tracheal tube, and must not cause tissue or skin pressure. The system must allow for irrigation via a lateral irrigation port which must be a capped with a one-way valve that connects to a standard luer-slip syringe. The irrigation port must be situated at the patient end of the system for catheters and at the suction end of the catheter for catheters.made from medically safe PVC The suction end of the system must be fitted with capped, lockable thumb valve to control suctioning. It must have a tapered connection to fit a range of suction tubing. The system must be as light in weight as possible. The system is for single patient use for a specified 48 or 72 hours. Must have coloured day sticker labels to facilitate regular changing of the catheter system. Made from medically safe PVC. Sterile and individually packed in easy to open pouch with one transparent side; For SINGLE use only The system must conform to ISO 8836:2019. | Closed Suction Catheter - Paediatric Tracheostomy Y-Type 48Hr 6Fr OD: 2.0mm For ETT
| RT13-07-02E |  | Closed Suction Catheter - Paediatric Tracheostomy Y-Type 48Hr 7Fr OD: 2.3mm For ETT
| RT13-07-03E |  | Closed Suction Catheter - Paediatric Tracheostomy Y-Type 48Hr 8Fr Suction Catheter OD: 2.6mm For ETT with ID: 4mm Length: 30.5cm Y connectors: 3.5mm, 4mm, 4.5mm
| RT13-07-04E |  | Closed Suction Catheter - Paediatric Tracheostomy Y-Type 48Hr 10Fr OD: 3.3mm For ET
| RT13-07-05E |  | Closed Suction Catheter - Paediatric Tracheostomy Y-Type 48Hr 12Fr OD: 4.0mm For ETT with ID: 6.0mm Length: 40.5cm Y connectors: 5.5mm, 6mm, 6.5mm
CLOSED SUCTION CATHETER - PAEDIATRIC TRACHEOSTOMY Y-TYPE 72HR
CATEGORY 7F | RT13-07-01F | CLOSED SUCTION CATHETER PAEDIATRIC TRACHEOSTOMY - 72HRS Purpose: Closed system for respiratory suctioning through tracheostomy tubes without disconnecting or interrupting ventilation. Consists of a closed suction catheter with an atraumatic tip. The catheter must have markings along the length to indicate depth of suctioning and a radiopaque line on the catheter tip. While connected but not in suction mode the catheter tip must not extend beyond the sheath or impede airflow. The suction catheter must be enclosed within a durable, flexible, plastic sheath for storage to avoid contamination. The transparent sheath must allow visualisation of markings on the suction catheter and must connect to the tracheal tube via a clear, firmly-fitting, plastic connector or adaptor. The sheath and connector must allow for smooth advancement of catheter. The system must seal the airway and catheter from atmospheric contamination and allow uninterrupted positive pressure ventilation throughout the suction process. An effective seal around the catheter must prevent ventilation gas entering the lumen of the sheath. The system must protect the user from contamination by secretions and suctioned material. Proximal end of suction catheter must have clear chamber, so that suctioned contents can be visualised and must have lockable suction valve with capped suction tubing connector. Paediatric Catheters must have a Y- connector (without locking mechanism) with an appropriate range of diameters for each size catheter. The connector must allow variations in patient positions without causing torque, twisting, or obstruction of the tracheal tube, and must not cause tissue or skin pressure. The system must allow for irrigation via a lateral irrigation port which must be a capped with a one-way valve that connects to a standard luer-slip syringe. The irrigation port must be situated at the patient end of the system for catheters and at the suction end of the catheter for catheters.made from medically safe PVC The suction end of the system must be fitted with capped, lockable thumb valve to control suctioning. It must have a tapered connection to fit a range of suction tubing. The system must be as light in weight as possible. The system is for single patient use for a specified 48 or 72 hours. Must have coloured day sticker labels to facilitate regular changing of the catheter system. Made from medically safe PVC. Sterile and individually packed in easy to open pouch with one transparent side; For SINGLE use only The system must conform to ISO 8836:2019. | Closed Suction Catheter - Paediatric Tracheostomy Y-Type 72Hr 6Fr OD: 2.0mm For ETT ID: 3mm Length: 30.5cm Y connectors: 2.5mm, 3.0mm, 3.5mm
| RT13-07-02F |  | Closed Suction Catheter - Paediatric Tracheostomy Y-Type 72Hr 7Fr OD: 2.3mm For ETT ID: 3.5mm Length: 30.5cm Y connectors: 3.0mm, 3.5mm, 4mm
| RT13-07-03F |  | Closed Suction Catheter - Paediatric Tracheostomy Y-Type 72Hr 8Fr OD: 2.6mm For ETT with ID: 4mm Length: 30.5cm Y connectors: 3.5mm, 4mm, 4.5mm
| RT13-07-04F |  | Closed Suction Catheter - Paediatric Tracheostomy Y-Type 72Hr 10Fr OD: 3.3mm For ETT with ID: 5.0mm Length: 40.5cm Y connectors: 4.5mm, 5mm, 5.5mm
| RT13-07-05F |  | Closed Suction Catheter - Paediatric Tracheostomy Y-Type 72Hr 12Fr OD: 4.0mm For ETT with ID: 6.0mm Length: 40.5cm Y connectors: 5.5mm, 6mm, 6.5mm
CLOSED SUCTION CATHETER ADULT TRACHEOSTOMY - 48HRS
CATEGORY 7G | RT13-07-01G | CLOSED SUCTION CATHETER ADULT TRACHEOSTOMY - 48HRS Purpose: Closed system for respiratory suctioning through tracheostomy tubes without disconnecting or interrupting ventilation. Consists of a closed suction catheter with an atraumatic tip. The catheter must have markings along the length to indicate depth of suctioning and a radiopaque line on the catheter tip. While connected but not in suction mode the catheter tip must not extend beyond the sheath or impede airflow. The suction catheter must be enclosed within a durable, flexible, plastic sheath for storage to avoID:contamination. The transparent sheath must allow visualisation of markings on the suction catheter and must connect to the tracheal tube via a clear, firmly-fitting, plastic connector or adaptor. The sheath and connector must allow for smooth advancement of catheter. The system must seal the airway and catheter from atmospheric contamination and allow uninterrupted positive pressure ventilation throughout the suction process. An effective seal around the catheter must prevent ventilation gas entering the lumen of the sheath. The system must protect the user from contamination by secretions and suctioned material. Proximal end of suction catheter must have clear chamber, so that suctioned contents can be visualised and must have lockable suction valve with capped suction tubing connector. The suction catheter must connect onto the tube with a DOUBLE swivel connector or a double-swivel Y-connector with standard 15 mm OD taper fittings. This connection must be secure and accidental disconnection must be minimal The whole device must be as lightweight as possible. Must include a light, straight, flexible, clear, concertina catheter-mount (15x22 mm connections) tube for connection to the breathing circuit The connector must allow variations in patient positions without causing torque, twisting, or obstruction of the tracheal tube, and must not cause tissue or skin pressure. The system must allow for irrigation via a lateral irrigation port which must be a capped with a one-way valve that connects to a standard luer-slip syringe. The irrigation port must be situated at the patient end of the system for catheters and at the suction end of the catheter for catheters.made from medically safe PVC The suction end of the system must be fitted with capped, lockable thumb valve to control suctioning. It must have a tapered connection to fit a range of suction tubing. The system is for single patient use for a specified 48 or 72 hours. Must have coloured day sticker labels to facilitate regular changing of the catheter system. Made from medically safe PVC. Sterile and individually packed in easy to open pouch with one transparent side; For SINGLE use only The system must conform to ISO 8836:2019. | Closed Suction Catheter - Adult Tracheostomy 48Hr 10Fr OD: 3.3mm Length range: 300 - 400mm
| RT13-07-02G |  | Closed Suction Catheter - Adult Tracheostomy 48Hr 12Fr OD: 4.0mm Length range: 300 - 400mm
| RT13-07-03G |  | Closed Suction Catheter - Adult Tracheostomy 48Hr 14Fr OD: 4.7mm Length range: 300 - 400mm
| RT13-07-04G |  | Closed Suction Catheter - Adult Tracheostomy 48Hr 16Fr OD: 5.3mm Length range: 300 - 400mm
CLOSED SUCTION CATHETER ADULT TRACHEOSTOMY - 72HRS
CATEGORY 7G | RT13-07-01G | CLOSED SUCTION CATHETER ADULT TRACHEOSTOMY - 72HRS Purpose: Closed system for respiratory suctioning through tracheostomy tubes without disconnecting or interrupting ventilation. Consists of a closed suction catheter with an atraumatic tip. The catheter must have markings along the length to indicate depth of suctioning and a radiopaque line on the catheter tip. While connected but not in suction mode the catheter tip must not extend beyond the sheath or impede airflow. The suction catheter must be enclosed within a durable, flexible, plastic sheath for storage to avoID:contamination. The transparent sheath must allow visualisation of markings on the suction catheter and must connect to the tracheal tube via a clear, firmly-fitting, plastic connector or adaptor. The sheath and connector must allow for smooth advancement of catheter. The system must seal the airway and catheter from atmospheric contamination and allow uninterrupted positive pressure ventilation throughout the suction process. An effective seal around the catheter must prevent ventilation gas entering the lumen of the sheath. The system must protect the user from contamination by secretions and suctioned material. Proximal end of suction catheter must have clear chamber, so that suctioned contents can be visualised and must have lockable suction valve with capped suction tubing connector. The suction catheter must connect onto the tube with a DOUBLE swivel connector or a double-swivel Y-connector with standard 15 mm OD taper fittings. This connection must be secure and accidental disconnection must be minimal The whole device must be as lightweight as possible. Must include a light, straight, flexible, clear, concertina catheter-mount (15x22 mm connections) tube for connection to the breathing circuit The connector must allow variations in patient positions without causing torque, twisting, or obstruction of the tracheal tube, and must not cause tissue or skin pressure. The system must allow for irrigation via a lateral irrigation port which must be a capped with a one-way valve that connects to a standard luer-slip syringe. The irrigation port must be situated at the patient end of the system for catheters and at the suction end of the catheter for catheters.made from medically safe PVC The suction end of the system must be fitted with capped, lockable thumb valve to control suctioning. It must have a tapered connection to fit a range of suction tubing. The system is for single patient use for a specified 48 or 72 hours. Must have coloured day sticker labels to facilitate regular changing of the catheter system. Made from medically safe PVC. Sterile and individually packed in easy to open pouch with one transparent side; For SINGLE use only The system must conform to ISO 8836:2019. | Closed Suction Catheter - Adult Tracheostomy 72Hr 10Fr OD: 3.3mm Length range: 300 - 400mm
| RT13-07-02G |  | Closed Suction Catheter - Adult Tracheostomy 72Hr 12Fr OD: 4.0mm Length range: 300 - 400mm
| RT13-07-03G |  | Closed Suction Catheter - Adult Tracheostomy 72Hr 14Fr OD: 4.7mm Length range: 300 - 400mm
| RT13-07-04G |  | Closed Suction Catheter - Adult Tracheostomy 72Hr 16Fr OD: 5.3mm Length range: 300 - 400mm

## Category 8
BID NUMBER | RT13
BID DESCRIPTION | SUPPLY AND DELIVERY OF RESPIRATORY AID AND ANAESTHETIC ACCESSORIES TO THE STATE
CATEGORY 8
FILTERS
| Item Number | Technical Specification | Product Name and Size | Insert inputs on this column
HEAT MOISTURE EXCHANGE FILTER (HMEF)- INFANT
CATEGORY 8A | RT13-08-01A | HEAT MOISTURE EXCHANGE FILTER (HMEF) - INFANT Purpose: Heat and moisture exchange and filtration device to provide passive humidification and patient protection against contamination - used on patient end Bidirectional use filter must be hydrophobic with low resistance throughout clinical use and must remain effective when wet. Must be hygroscopic to humidify the patient's airway Maximum use: 24 hrs Liquid and airborne bacterial and viral filtration: 99,99% efficiency to protect patient and ventilator/ anaesthetic machine. Preference will be given to filters with higher efficiency that do not increase resistance. Microbiologically validated against mycobacterium tuberculosis, Pseudomonas species, Serratia marcescens, MS2-virus, Staphylococcus aureus, Hepatitis C-virus,human immunodeficiency virus, and COVID 19 virus. Validation must be provided. To comply with ISO 9360 or equivalent Consists of an ergonomic, non-conductive smooth housing profile. Latex-free non-pyrogenic, hypoallergenic and DEHP-free transparent medical grade plastic Connections to fit standard breathing circuits and ventilator outlets: Patient ends 15F/22M ISO connections and circuit end to be compatible with commonly used anaesthesia breathing circuits and ventilator outlets Must have a luer monitoring port with a toggled cap for integral monitoring of end tidal CO2 Sterile and individually packed in peel pouch with view paper; For SINGLE use only | Heat Moisture Exchange Filter (HMEF) – Infant, Straight with Tidal Volumes: 30 -100ml (maximum: 200ml ) Maximum filter volume: < 10ml; Filter weight: 10G; Resistance: 1.0cmH2O at 11L/min Gas flow. Humidification > 25mg H2O /litre. Viral and Bacterial Filtration > 99.99%
| RT13-08-02A |  | Heat Moisture Exchange Filter (HMEF) – Infant, Angled with Tidal Volumes: 30 -100ml (maximum: 200ml ) Integrated 45 - 90° angled elbow design Maximum filter volume: < 10ml; Filter weight: 10G; Resistance: 1.0cmH2O at 11L/min Gas flow. Humidification: > 25mg H2O /litre. Viral and Bacterial Filtration > 99.99%
HEAT MOISTURE EXCHANGE FILTER (HMEF)- PAEDIATRIC |  | CORRUGATED BREATHING CIRCUITS WITH WATER TRAPS AND EXTRA LIMB AND BAG - SINGLE USE
CATEGORY 8B | RT13-08-01B | HEAT MOISTURE EXCHANGE FILTER (HMEF) - PAEDIATRIC Purpose: Heat and moisture exchange and filtration device to provide passive humidification and patient protection against contamination - used on patient end Bidirectional use filter must be hydrophobic with low resistance throughout clinical use and must remain effective when wet. Must be hygroscopic to humidify the patient's airway Maximum use: 24 hrs Liquid and airborne bacterial and viral filtration: 99,999% efficiency to protect patient and ventilator/ anaesthetic machine. Microbiologically validated against mycobacterium tuberculosis, Pseudomonas species, Serratia marcescens, MS2-virus, Staphylococcus aureus, Hepatitis C-virus,human immunodeficiency virus, and COVID 19 virus. Validation must be provided. To comply with ISO 9360 or equivalent Consists of an ergonomic, non-conductive smooth housing profile. Latex-free non-pyrogenic, hypoallergenic and DEHP-free transparent medical grade plastic Connections to fit standard breathing circuits and ventilator outlets: Patient ends 15F/22M ISO connections and circuit end to be compatible with commonly used anaesthesia breathing circuits and ventilator outlets Must have a luer monitoring port with a toggled cap for integral monitoring of end tidal CO2 Sterile and individually packed in peel pouch with view paper; For SINGLE use only | Heat Moisture Exchange Filter (HMEF) – Paediatric, Straight with Tidal Volumes > 75ml Maximum filter volume: < 30ml Maximum filter weight: ≤ 20g; Resistance: < 2.5 cmH2O at 30 l/minGas flow. Filtration ≥ 99.999 Humidification > 30mg H2O/litre.
| RT13-08-02B |  | Heat Moisture Exchange Filter (HMEF) – Paediatric, Angled with Tidal Volumes > 75ml Integrated 45 - 90° angled elbow design Maximum filter volume: < 30ml Maximum filter weight: ≤ 20g; Resistance: < 2.5 cmH2O at 30 l/minGas flow. Filtration ≥ 99.999 Humidification > 30mg H2O/litre.
HIGH-EFFICIENCY PARTICULATE AIR FILTER (HEPA)
CATEGORY 8C | RT13-08-01C | HIGH-EFFICIENCY PARTICULATE AIR FILTER (HEPA) Purpose: Heat and moisture exchange and filtration device to provide passive humidification and patient protection against contamination - used on patient end Adult, Straight, straight with Tidal Volumes: 300ml – 1500ml Maximum filter volume: 55ml; Resistance: < 2.0 mbar at 30 l/min Gas flow. Humidification: ≥ 34.2 mg H2O /litre. Liquid and airborne bacterial and viral filtration: 99,9999% Bidirectional use filter must be hydrophobic with low resistance throughout clinical use and must remain effective when wet. Must be hygroscopic to humidify the patient's airway Liquid and airborne bacterial and viral filtration: 99,9999% efficiency to protect patient and ventilator/ anaesthetic machine. Preference will be given to filters with higher efficiency. Microbiologically validated against mycobacterium tuberculosis, Pseudomonas species, Serratia marcescens, MS2-virus, Staphylococcus aureus, Hepatitis C-virus,human immunodeficiency virus, and COVID 19 virus. Validation must be provided. To comply with ISO 9360 or equivalent Consists of an ergonomic, non-conductive smooth housing profile. Latex-free non-pyrogenic, hypoallergenic transparent medical grade plastic Connections to fit standard breathing circuits and ventilator outlets: Patient ends 15F/22M ISO connections and circuit end to be compatible with commonly used anaesthesia breathing circuits and ventilator outlets Must have a luer monitoring port with a toggled cap for integral monitoring of end tidal CO2 Preference will be given to items that are DEHP free Sterile and individually packed in peel pouch with view paper; For SINGLE use only | High-efficiency Particulate Air Filter (HEPA) - Adult, Straight Tidal Volumes: 300ml – 1500ml Maximum use: 24 hrs Maximum filter volume: 55ml; Resistance: < 2.0 mbar at 30l/min Gas flow. Humidification: ≥ 34.2mg H2O/litre. Liquid and airborne bacterial and viral filtration: 99,9999%
| RT13-08-02C |  | High-efficiency Particulate Air Filter (HEPA) - Adult, Straight Tidal Volumes: 300ml – 1500ml Maximum use: 48 hrs Maximum filter volume: 55ml; Resistance: < 2.0 mbar at 30l/min Gas flow. Humidification: ≥ 34.2mg H2O/litre. Liquid and airborne bacterial and viral filtration: 99,9999%
HEAT AND MOISTURE EXCHANGER (HME) - NEONATAL HME REPLACEMENT CONNECTOR
CATEGORY 8D | RT13-08-01D | HEAT AND MOISTURE EXCHANGER (HME) - NEONATAL HME REPLACEMENT CONNECTOR Purpose: To warm and moisten air breathed through a endotracheal tube during anaesthesia Consists of a bidirectional hygroscopic bacteriostatic paper filter, with low resistance throughout clinical use Must remain effective when wet. For 24 hourly use The HME to be used as a replacement for size 3.0 and/or 3.5 ETT circuit connectors. The connectors must fit standard breathing circuits and ventilator outlets. Must have a side port allowing for access for suctioning and C02 sampling. The patient end must be tapered to fit securely into the endotracheal tube. Must have a non-conductive smooth housing profile. Latex-free non-pyrogenic, hypoallergenic and DEHP-free transparent medical grade plastic Sterile and individually packed in peel pouch with view paper; For SINGLE use only Must comply with ISO 9360 | HEAT AND MOISTURE EXCHANGER (HME) -NEONATAL HME REPLACEMENT CONNECTOR Size 3.0 for VT 10-50ml Dead space: < 3ml; Maximum filter weight: < 5g; Resistance: 1.8 cmH2O at 5l/min Gas flow. Humidification: ± 30mg H2O /litre for VT 20ml
| RT13-08-02D |  | HEAT AND MOISTURE EXCHANGER (HME) - NEONATAL HME REPLACEMENT CONNECTOR Size 3.5 for VT 10-50ml Dead space: < 3ml; Maximum filter weight: < 5g; Resistance: 1.1cmH2O at 5l/min Gas flow. Humidification: ± 30mg H2O /litre for VT 10ml
HEAT AND MOISTURE EXCHANGER (HME) -NEONATAL Straight
CATEGORY 8E | RT13-08-01E | HEAT AND MOISTURE EXCHANGER (HME) -NEONATAL Straight Purpose: To warm and moisten air breathed through a tracheostomy. Bidirectional use filter must be hygroscopic with low resistance throughout clinical use. Foam/ Pleated paper acceptable Must remain effective when wet. For 24 hourly use Must have a non-conductive smooth housing profile. Latex-free non-pyrogenic, hypoallergenic and DEHP-free transparent medical grade plastic Where applicable connections to fit standard breathing circuits and ventilator outlets: Patient ends 15F/22M ISO connections and circuit end to be compatible with commonly used anaesthesia breathing circuits and ventilator outlets Sterile and individually packed in peel pouch with view paper; For SINGLE use only Must comply with ISO 9360 | HEAT AND MOISTURE EXCHANGER (HME) -NEONATAL Straight For Tidal Volumes > 8ml Dead space: < 3ml; Maximum filter weight: < 5g; Resistance: < 0.9cmH2O at 10l/min Gas flow. Humidification: ± 30mg H2O /litre.
HEAT AND MOISTURE EXCHANGER INFANT
CATEGORY 8F | RT13-08-01F | HEAT AND MOISTURE EXCHANGER INFANT Purpose: Heat and moisture exchange device to provide passive humidification to patient's airways - for direct connection to tracheostomies Bidirectional use filter must be hydrophobic with low resistance throughout clinical use and must remain effective when wet. For short term use - transport or short procedures Consists of an ergonomic, non-conductive smooth housing profile. Latex-free non-pyrogenic, hypoallergenic and DEHP-free transparent medical grade plastic Connections to fit standard breathing circuits and ventilator outlets: Patient ends 15F/22M ISO connections and circuit end to be compatible with commonly used anaesthesia breathing circuits and ventilator outlets Sterile and individually packed in peel pouch with view paper; For SINGLE use only | Heat and Moisture Exchanger (HME) - Infant straight For Tidal Volumes: > 50ml Dead space: < 11ml; Maximum filter weight: < 12g; Resistance: < 0.6cmH2O at 5 l/min Gas flow. Humidification: > 30 mg H2O /litre.
| RT13-08-02F |  | Tracheostomy Heat and Moisture Exchanger (HME) - Infant straight with toggled Capnography port For Tidal Volumes: > 50ml Dead space: < 11ml; Maximum filter weight: < 12g; Resistance: < 0.6cmH2O at 5 l/min Gas flow. Humidification: > 30 mg H2O /litre.
TRACHEOSTOMY HEAT AND MOISTURE EXCHANGER ADULT
CATEGORY 8G | RT13-08-01G | TRACHEOSTOMY HEAT AND MOISTURE EXCHANGER ADULT Purpose: Heat and moisture exchange device to provide passive humidification to patient's airways - for direct connection to tracheostomies Bidirectional use filter must be hydrophobic with low resistance throughout clinical use and must remain effective when wet. For short term use Consists of a low profile T-shaped air bridge that connects directly to the endotracheal/tracheostomy tube. Paper/sponge filter. Latex-free non-pyrogenic, hypoallergenic and DEHP-free transparent medical grade plastic Connections to fit standard breathing circuits and ventilator outlets: Patient ends 15F/22M ISO connections and circuit end to be compatible with commonly used anaesthesia breathing circuits and ventilator outlets Sterile and individually packed in peel pouch with view paper; For SINGLE use only | Tracheostomy Heat and Moisture Exchanger (HME) - Adult, Single Sponge, with swivelled Oxygen connector and clipped Suction Port For Tidal Volumes: > 60ml Maximum filter weight: < 12g; Resistance: < 0.6cmH2O at 5 l/min Gas flow. Humidification: > 25mg H2O /litre.
| RT13-08-02G |  | Tracheostomy Heat and Moisture Exchanger (HME) Adult, Double paper, with swivelled Oxygen connector and clipped Suction Port For Tidal Volumes: > 50ml Maximum filter weight: 10g Resistance: 0.3cmH2O at 30l/min Gas flow. Humidification: > 25 mg H2O /litre.
MECHANICAL FILTER FOR MACHINE PROTECTION
CATEGORY 8H | RT13-08-01H | MECHANICAL FILTER FOR MACHINE PROTECTION WITH VIRAL AND BACTERIAL FILTRATION - Adult, Straight, straight with Tidal Volumes: 300ml – 1500ml Maximum filter volume: 55ml; Resistance: < 2.0 mbar at 30 l/min Gas flow. Purpose: Mechanical filtration device to protect ventilators. To be used on machine-end Bidirectional use filter must be hydrophobic with low resistance throughout clinical use and must remain effective when wet. Maximum use: 24 hrs Liquid and airborne bacterial and viral filtration: 99,9999% efficiency to protect ventilator/ anaesthetic machine. Microbiologically validated against mycobacterium tuberculosis, Pseudomonas species, Serratia marcescens, MS2-virus, Staphylococcus aureus, Hepatitis C-virus and human immunodeficiency virus and Covid-19. Validation must be provided. To comply with ISO 9360 or equivalent Consists of an ergonomic, non-conductive smooth housing profile. Latex-free non-pyrogenic, hypoallergenic transparent medical grade plastic Connections to fit standard breathing circuits and ventilator outlets: Patient ends 15F/22M ISO connections and circuit end to be compatible with commonly used anaesthesia breathing circuits and ventilator outlets Preference will be given to items that are DEHP free Sterile and individually packed in peel pouch with view paper; For SINGLE use only | MECHANICAL FILTER FOR MACHINE PROTECTION WITH VIRAL AND BACTERIAL FILTRATION - Adult, Straight, straight with Tidal Volumes: 300ml – 1500ml Maximum filter volume: 55ml; Resistance: < 2.0 mbar at 30 l/min Gas flow.

## Category 9
BID NUMBER | RT13
BID DESCRIPTION | SUPPLY AND DELIVERY OF RESPIRATORY AID AND ANAESTHETIC ACCESSORIES TO THE STATE
CATEGORY 9
BREATHING CIRCUITS and REBREATHING BAGS
| Item Number | Technical Specification | Product Name and Size | Insert inputs on this column
CORRUGATED BREATHING CIRCUITS WITH EXTRA LIMB AND BAG - SINGLE USE
CATEGORY 9A | RT13-09-01A | CORRUGATED BREATHING CIRCUITS WITH EXTRA LIMB AND BAG - SINGLE USE Purpose: Used to deliver oxygen and Volatile Gases to patient's and to eliminate C02 during anaesthesia Transparent circuit with low resistance, non-extendable, corrugated tubing that is non-kinking, lightweight and crush resistant. The tubing must be pliable and must not retain a shape. Uninterrupted tubing (not joined). Shallow corrugations to minimise flow resistance and turbulence. Must have an extra limb with an anti-static reservoir bag that gives good feedback regarding patient lung compliance. The reservoir bag must not twist on itself and must have a mechanism to prevent this. Connectors Malleable 22mm female connectors bonded to the ends of inspiratory and expiratory limbs – must fit securely to the machine. Spare 22mm M/M connector. Y-piece at patient's end. Separate elbow-piece with a toggled cap on a luer lock port. Good quality to ensure no leakage or easy disconnections Latex-free, non-pyrogenic, hypoallergenic, non-toxic and non-conductive medical grade plastic No variation in Circuit limb length or Extra limb will be accepted Clinically clean and individually packed, For SINGLE use only To comply with ISO 5367 or equivalent | Corrugated breathing circuit with extra limb and breathing bag - Single Use - Neonatal 15mm Circuit: 1.6m; 15mm Extra limb: 1.2m Reservoir bag: 500ml
| RT13-09-02A |  | Corrugated breathing circuit with extra limb and breathing bag - Single Use - Neonatal 15mm Circuit: 1.6m; 15mm Extra limb: 1.8m Reservoir bag: 500ml
| RT13-09-03A |  | Corrugated breathing circuit with extra limb and breathing bag - Single Use - Paediatric > 3 yrs 15mm Circuit: 1.6m; 15mm Extra limb: 1.2m Reservoir bag: 1l
| RT13-09-04A |  | Corrugated breathing circuit with extra limb and breathing bag - Single Use - Paediatric > 3 yrs 15mm Circuit: 1.6m; 15mm Extra limb: 1.8m Reservoir bag: 1l
| RT13-09-05A |  | Corrugated breathing circuit with extra limb and breathing bag - Single Use - Adult 22mm Circuit: 1.8m; 22mm Extra limb: 1.2m OD: 22mm Reservoir bag: 2l
| RT13-09-06A |  | Corrugated breathing circuit with extra limb and breathing bag - Single Use - Adult 22mm Circuit: 1.8m; 22mm Extra limb: 1.5m OD: 22mm Reservoir bag: 2l
| RT13-09-07A |  | Corrugated breathing circuit with extra limb and breathing bag - Single Use - Adult 22mm Circuit: 1.8m; 22mm Extra limb: 1.8m OD: 22mm Reservoir bag: 2l
| RT13-09-08A |  | Corrugated breathing circuit with extra limb and breathing bag - Single Use - Adult Long 22mm Circuit: 2.5m; 22mm Extra limb: 1.2m OD: 22mm Reservoir bag: 2l
| RT13-09-09A |  | Corrugated breathing circuit with extra limb and breathing bag - Single Use - Adult Long 22mm Circuit: 2.5m; 22mm Extra limb: 1.5m OD: 22mm Reservoir bag: 2l
| RT13-09-10A |  | Corrugated breathing circuit with extra limb and breathing bag - Single Use - Adult Long 22mm Circuit: 2.5m; 22mm Extra limb: 1.8m OD: 22mm Reservoir bag: 2l
CORRUGATED BREATHING CIRCUITS WITH WATER TRAPS AND EXTRA LIMB AND BAG - SINGLE USE |  | CORRUGATED BREATHING CIRCUITS WITH WATER TRAPS AND EXTRA LIMB AND BAG - SINGLE USE
CATEGORY 9B | RT13-09-01B | CORRUGATED BREATHING CIRCUITS WITH WATER TRAPS AND EXTRA LIMB AND BAG - SINGLE USE Purpose: Used to deliver oxygen and Volatile Gases to patient's and to eliminate C02 during anaesthesia Transparent circuit with low resistance, non-extendable, corrugated tubing that is non-kinking, lightweight and crush resistant. The tubing must be pliable and must not retain a shape Shallow corrugations to minimise flow resistance and turbulence. Uninterrupted tubing (not joined). Lightweight water traps centrally on inspiratory and expiratory limb Must have an extra limb with an anti-static reservoir bag that gives good feedback regarding patient lung compliance. The reservoir bag must not twist on itself and must have a mechanism to prevent this. Connectors Malleable 22mm female connectors bonded to the ends of inspiratory and expiratory limbs – must fit securely to the machine. Spare 22mm M/M connector. Y-piece at patient's end. Separate angle-piece with luer lock port Good quality to ensure no leakage or easy disconnections Latex-free, non-pyrogenic, hypoallergenic, non-toxic and non-conductive medical grade plastic No variation in Circuit limb length or Extra limb will be accepted Clinically clean and individually packed, For SINGLE use only To comply with ISO 5367 or equivalent | Corrugated breathing circuit with water traps and extra limb with breathing bag - Single Use - Paediatric > 3 yrs 15mm Circuit: 1.6m; 15mm Extra limb: 1.2m Reservoir bag: 1l
| RT13-09-02B |  | Corrugated breathing circuit with water traps and extra limb with breathing bag - Single use Paediatric > 3 yrs 15mm Circuit: 1.6m; 15mm Extra limb: 1.8m Reservoir bag: 1l
| RT13-09-03B |  | Corrugated breathing circuit with water traps and extra limb with breathing bag - Single Use Adult 22mm Circuit: 1.8m; 22mm Extra limb: 1.2m OD: 22mm Reservoir bag: 2l
| RT13-09-04B |  | Corrugated breathing circuit with water traps and extra limb with breathing bag - Single Use Adult 22mm Circuit: 1.8m; 22mm Extra limb: 1.5m OD: 22mm Reservoir bag: 2l
| RT13-09-05B |  | Corrugated breathing circuit with water traps and extra limb with breathing bag - Single UseAdult 22mm Circuit: 1.8m; 22mm Extra limb: 1.5m OD: 22mm Reservoir bag: 2l
| RT13-09-06B |  | Corrugated breathing circuit with water traps and extra limb with breathing bag - Single Use Adult Long 22mm Circuit: 2.5m; 22mm Extra limb: 1.2m OD: 22mm Reservoir bag: 2l
| RT13-09-07B |  | Corrugated breathing circuit with water traps and extra limb with breathing bag - Single Use Adult Long 22mm Circuit: 2.5m; 22mm Extra limb: 1.5m OD: 22mm Reservoir bag: 2l
| RT13-09-08B |  | Corrugated breathing circuit with water traps and extra limb with breathing bag - Single Use Adult Long 22mm Circuit: 2.5m; 22mm Extra limb: 1.8m OD: 22mm Reservoir bag: 2l
SINGLE CORRUGATED BREATHING TUBING - SINGLE USE |  | SINGLE CORRUGATED BREATHING TUBING - SINGLE USE
CATEGORY 9C | RT13-09-01C | SINGLE CORRUGATED BREATHING TUBING - SINGLE USE Purpose: Used as extension or extra limb in breathing circuits Transparent circuit with low resistance, non-extendable, corrugated tubing that is non-kinking, lightweight and crush resistant. The tubing must be pliable and must not retain a shape Shallow corrugations to minimise flow resistance and turbulence. Connectors Malleable 22mm female connectors bonded to the end of the tubing Spare connectors to be provided Latex-free, non-pyrogenic, hypoallergenic, non-toxic and non-conductive medical grade plastic No variation in Circuit limb length or Extra limb will be accepted Clinically clean and individually packed, For SINGLE use only To comply with ISO 5367 or equivalent | Single Corrugated Breathing Tubing - Single Use Paediatric 15mm Length: 1.2m Connector: 15mm M/F; 15mm M/M; 15mm F/F
| RT13-09-02C |  | Single Corrugated Breathing Tubing - Single Use Paediatric 15mm Length: 1.8m Connector: 15mm M/F; 15mm M/M; 15mm F/F
| RT13-09-03C |  | Single Corrugated Breathing Tubing - Single Use Adult 22mm Length: 1.2m Connector: 22mm M/F; 22mm M/M; 22mm F/F
| RT13-09-04C |  | Single Corrugated Breathing Tubing - Single Use Adult 22mm Length: 1.8m Connector: 22mm M/F; 22mm M/M; 22mm F/F
MAPLESON F PAEDIATRIC BREATHING CIRCUIT - (Modified Ayre’s T piece)
CATEGORY 9D | RT13-09-01D | MAPLESON F PAEDIATRIC BREATHING CIRCUIT - (Modified Ayre’s T piece) for children < 20-25 kg Purpose: Used to deliver oxygen and volatile gases to patient's and to eliminate C02 during anaesthesia Transparent circuit with low resistance, non-extendable, corrugated tubing that is non-kinking, lightweight and crush resistant. The tubing must be pliable and must not retain a shape with shallow corrugations to minimise flow resistance and turbulence. Tubing must be connected to a straight T-Piece with inspiratory port, expiratory port and patient connection port 22M/15F. Inspiratory tubing: 10mm ± 1.8 - 2.1m Bonded 15mm male connector at the end of the inspiratory limb– must fit securely to the machine. . Expiratory tubing 15mm diameter 0.5m with a 0.5l open-ended reservoir bag to allow gas to escape via the tail of the bag The reservoir bag must give good feedback regarding patient lung compliance and must not twist on itself and must have a mechanism to prevent this. Connectors: Elbow-piece with toggled cap on a luer lock port, capped connector at patient’s end. Good quality to ensure no leakage or easy disconnections Manufactured from transparent non-conductive medical grade plastic that is latex-free non-pyrogenic, hypoallergenic and non-toxic. | Mapleson F breathing circuit - (Modified Ayre’s T piece) for children < 20-25 kg 15mm Circuit ± 1.8 - 2.1m Expiratory Limb 0.4m Reservoir bag: 0.5l open ended
MAPLESON C BREATHING CIRCUIT
CATEGORY 9E | RT13-09-01E | MAPLESON C BREATHING CIRCUIT Purpose: A bagging system for the induction of anaesthesia, resuscitation and patient transport Consists of transparent low resistance, non-extendable, 15mm corrugated tubing that is non-kinking, lightweight and crush resistant. ID: 15mm diameter Length: 1.8m The tubing must be pliable and must not retain a shape with shallow corrugations to minimise flow resistance and turbulence. The tubing must be connected to a straight T-Piece via an inspiratory port. The straight T-Piece must have an angled elbow for patient connection 22M/15F, and an APL valve on the T of the T-Piece and a rebreathing bag Connectors: Elbow-piece with toggled cap on a luer lock port, capped connector at patient’s end. Good quality to ensure no leakage or easy disconnections Manufactured from transparent non-conductive medical grade plastic that is latex-free non-pyrogenic, hypoallergenic and non-toxic. | Mapleson F Breathing Circuit - Paediatric with 1.0 litre Rebreather Bag
| RT13-09-02E |  | Mapleson F Breathing Circuit - Adult with 2.0 litre Rebreather Bag
SMOOTH BORE BREATHING CIRCUITS WITH EXTRA LIMB AND BREATHING BAG - SINGLE USE
CATEGORY 9F | RT13-09-01F | SMOOTH BORE BREATHING CIRCUITS WITH EXTRA LIMB AND BREATHING BAG - SINGLE USE Purpose: Used to deliver oxygen and Volatile Gases to patient's and to eliminate C02 during anaesthesia Transparent circuit with low resistance, non-extendable, smooth-bore tubing that is non-kinking, lightweight and crush resistant. The tubing must be pliable and must not retain a shape Spiral outer casing that supports the inner lumen to maintain tube integrity. Must have an extra limb with an anti-static reservoir bag that gives good feedback regarding patient lung compliance. The reservoir bag must not twist on itself and must have a mechanism to prevent this. Connectors Malleable 22mm female connectors bonded to the ends of inspiratory and expiratory limbs – must fit securely to the machine. Spare 22mm M/M connector. Y-piece at patient's end. Separate elbow-piece with capped luer lock port. Good quality to ensure no leakage or easy disconnections Latex-free, non-pyrogenic, hypoallergenic, non-toxic and non-conductive medical grade plastic No variation in Circuit limb length or Extra limb will be accepted Clinically clean and individually packed, For SINGLE use only To comply with ISO 5367 or equivalent | Smooth bore breathing circuit with extra limb and breathing bag - Single Use Paediatric > 3yrs 15mm Circuit: 1.6m; Extra limb: 1.2m Reservoir bag: 1l
| RT13-09-02F |  | Smooth bore breathing circuit with extra limb and breathing bag - Single Use Paediatric > 3yrs 15mm Circuit: 1.6m; Extra limb: 1.8m Reservoir bag: 1l
| RT13-09-03F |  | Smooth bore breathing circuit with extra limb and breathing bag - Single Use Adult 22mm Circuit: 1.8m; Extra limb: 1.2m OD: Reservoir bag: 2l
| RT13-09-04F |  | Smooth bore breathing circuit with extra limb and breathing bag - Single Use Adult 22mm Circuit: 1.8m; Extra limb: 1.5m OD: Reservoir bag: 2l
| RT13-09-05F |  | Smooth bore breathing circuit with extra limb and breathing bag - Single Use Adult 22mm Circuit: 1.8m; Extra limb: 1.8m OD: Reservoir bag: 2l
| RT13-09-06F |  | Smooth bore breathing circuit with extra limb and breathing bag - Single Use Adult Long 22mm Circuit: 2.6m; Extra limb: 1.2m Reservoir bag: 2l
| RT13-09-07F |  | Smooth bore breathing circuit with extra limb and breathing bag - Single Use Adult Long 22mm Circuit: 2.6m; Extra limb: 1.5m Reservoir bag: 2l
| RT13-09-08F |  | Smooth bore breathing circuit with extra limb and breathing bag - Single Use Adult Long 22mm Circuit: 2.6m; Extra limb: 1.8m Reservoir bag: 2l
SMOOTH BORE BREATHING CIRCUITS WITH WATER TRAPS AND EXTRA LIMB AND BAG - SINGLE USE
CATEGORY 9G | RT13-09-01G | SMOOTH BORE BREATHING CIRCUITS WITH WATER TRAPS AND EXTRA LIMB AND BAG - SINGLE USE Purpose: Used to deliver oxygen and Volatile Gases to patient's and to eliminate C02 during anaesthesia Transparent circuit with low resistance, non-extendable, smooth-bore tubing that is non-kinking, lightweight and crush resistant. The tubing must be pliable and must not retain a shape Spiral outer casing that supports the inner lumen to maintain tube integrity. Lightweight water traps centrally on inspiratory and expiratory limbs Must have an extra limb with an anti-static reservoir bag that gives good feedback regarding patient lung compliance. The reservoir bag must not twist on itself and must have a mechanism to prevent this. Connectors Malleable 22mm female connectors bonded to the ends of inspiratory and expiratory limbs – must fit securely to the machine. Spare 22mm M/M connector. Y-piece at patient's end. Separate angle-piece with luer lock port Good quality to ensure no leakage or easy disconnections Latex-free, non-pyrogenic, hypoallergenic, non-toxic and non-conductive medical grade plastic No variation in Circuit limb length or Extra limb will be accepted Clinically clean and individually packed, For SINGLE use only To comply with ISO 5367 or equivalent | Smooth bore breathing circuit with with water traps, extra limb and breathing bag - Single Use Adult Circuit: 1.8m; Extra limb: 1.2m OD: 22mm Reservoir bag: 2l
| RT13-09-02G |  | Smooth bore breathing circuit with with water traps, extra limb and breathing bag - Single Use Adult Circuit: 1.8m; Extra limb: 1.5m OD: 22mm Reservoir bag: 2l
| RT13-09-03G |  | Smooth bore breathing circuit with with water traps, extra limb and breathing bag - Single Use Adult Circuit: 1.8m; Extra limb: 1.8m OD: 22mm Reservoir bag: 2l
| RT13-09-04G |  | Smooth bore breathing circuit with with water traps, extra limb and breathing bag - Single Use Adult Long Circuit: 2.6m; Extra limb: 1.2m OD: 22mm Reservoir bag: 2l
| RT13-09-05G |  | Smooth bore breathing circuit with with water traps, extra limb and breathing bag - Single Use Adult Long Circuit: 2.6m; Extra limb: 1.5m OD: 22mm Reservoir bag: 2l
| RT13-09-06G |  | Smooth bore breathing circuit with with water traps, extra limb and breathing bag - Single Use Adult Long Circuit: 2.6m; Extra limb: 1.8m OD: 22mm Reservoir bag: 2l
SINGLE SMOOTH-BORE BREATHING TUBING - SINGLE USE
CATEGORY 9H | RT13-09-01H | SINGLE SMOOTH-BORE BREATHING TUBING - SINGLE USE Purpose: Used as extension or extra limb in breathing circuits Transparent circuit with low resistance, non-extendable, smooth-bore tubing that is non-kinking, lightweight and crush resistant. The tubing must be pliable and must not retain a shape Spiral outer casing that supports the inner lumen to maintain tube integrity. Connectors Malleable 22mm female connectors bonded to the end of the tubing Spare connectors to be provided Latex-free, non-pyrogenic, hypoallergenic, non-toxic and non-conductive medical grade plastic Clinically clean and individually packed, For SINGLE use only To comply with ISO 5367 or equivalent | Single Smooth-bore Breathing Tubing - Single Use Adult 22mm Length: 1.2m Connector: 22mm M/F; 22mm M/M; 22mm F/F
| RT13-09-02H |  | Single Smooth-bore Breathing Tubing - Single Use Adult 22mm Length: 1.8m Connector: 22mm M/F; 22mm M/M; 22mm F/F
| RT13-09-03H |  | Single Smooth-bore Breathing Tubing - Single Use Paediatric 15mm Length: 1.2m Connector: 15mm M/F; 15mm M/M; 15mm F/F
| RT13-09-04H |  | Single Smooth-bore Breathing Tubing - Single Use Paediatric 15mm Length: 1.8m Connector: 15mm M/F; 15mm M/M; 15mm F/F
CORRUGATED BREATHING CIRCUITS - REUSABLE
CATEGORY 9I | RT13-09-01I | CORRUGATED BREATHING CIRCUITS WITH EXTRA LIMB AND BAG - REUSABLE Purpose: Used to deliver oxygen and Volatile Gases to patient's and to eliminate C02 during anaesthesia Consists of a circuit with low resistance, non-extendable, corrugated tubing that is non-kinking, clear and lightweight and crush resistant. The tubing must be pliable and must not retain a shape The corrugations must be shallow to minimise flow resistance and turbulence. Must have an extra limb with an anti-static Reservoir bag that gives good feedback regarding patient lung compliance and does not twist on itself. Must have a mechanism to prevent this. Connectors Malleable 22mm female connectors at the end of inspiratory and expiratory limbs – must fit securely to the machine. Spare 22mm M/M connector. Y-piece at patient's end. Good quality to ensure no leakage or easy disconnections No variation in Circuit limb length or Extra limb will be accepted Manufactured from latex- free, non-pyrogenic hypoallergenic and non-toxic Silicone/ Hydro-tubing . Reusable, autoclavable, individually packed and surgical clean. | Corrugated breathing circuit with extra limb and breathing bag - Reusable- Neonatal > 3 yrs Circuit: 1.6m; Extra limb: 1.2m OD: 15mm Reservoir bag: 500ml
| RT13-09-02I |  | Corrugated breathing circuit with extra limb and breathing bag - Reusable- Neonatal > 3 yrs Circuit: 1.6m; Extra limb: 1.8m OD: 15mm Reservoir bag: 500ml
| RT13-09-03I |  | Corrugated breathing circuit with extra limb and breathing bag - Reusable- Paediatric > 3 yrs Circuit: 1.6m; Extra limb: 1.2m OD: 15mm Reservoir bag: 1l
| RT13-09-04I |  | Corrugated breathing circuit with extra limb and breathing bag - Reusable- Paediatric > 3 yrs Circuit: 1.6m; Extra limb: 1.8m OD: 15mm Reservoir bag: 1l
| RT13-09-05I |  | Corrugated breathing circuit with extra limb and breathing bag - Reusable - Adult Circuit: 1.8m; Extra limb: 1.2m OD: 22mm Reservoir bag: 2l
| RT13-09-06I |  | Corrugated breathing circuit with extra limb and breathing bag - Reusable - Adult Circuit: 1.8m; Extra limb: 1.5m OD: 22mm Reservoir bag: 2l
| RT13-09-07I |  | Corrugated breathing circuit with extra limb and breathing bag - Reusable - Adult Circuit: 1.8m; Extra limb: 1.8m OD: 22mm Reservoir bag: 2l
SCAVENGING CORRUGATED TUBING
CATEGORY 9J | RT13-09-01J | SCAVENGING CORRUGATED TUBING Consists of transparent corrugated segmented tubing with cuffed repeating segments To fit 22mm M connectors Roll ± 50m. Single use, individually packed and surgical clean. | SCAVENGING Corrugated Tubing - segments repeating every ± 0.45m
| RT13-09-02J |  | SCAVENGING Corrugated Tubing - segments repeating every ± 0.15m
REBREATHING BAGS SINGLE USE
CATEGORY 9K | RT13-09-01K | REBREATHING BAGS SINGLE use Purpose: Stores gas between respiratory cycles. Provides information to anaesthetist in manual ventilation regarding pulmonary compliance and airway patency. Consists of alatex-free antistatic, non-conductive reservoir bag with a tapered neck and collar Compliance: 30-60cmH20 when filled to 4 X nominal volume Must hold at least 10 X nominal volume before it will burst Must retain original volume and shape (within 10%) at deflation Must must not adhere to itself and must not twist or kink on itself –must have a mechanism to prevent this. SINGLE patient use, clinically clean, latex-free To comply with ISO 5362:2006 or equivalent | Rebreathing reservoir bag 0.5 litre with 15mm F Connector - Single use
| RT13-09-02K |  | Rebreathing reservoir bag 1 litre with 15mm F connector - Single use
| RT13-09-03K |  | Rebreathing reservoir bag 2 litre with 22mm F connector- Single use
REBREATHING BAGS REUSABLE
CATEGORY 9L | RT13-09-01L | REBREATHING BAGS REUSABLE Purpose: Stores gas between respiratory cycles. Provides information to anaesthetist in manual ventilation regarding pulmonary compliance and airway patency. Consists of alatex-free antistatic, non-conductive reservoir bag with a tapered neck and collar Compliance: 30-60cmH20 when filled to 4 X nominal volume Must hold at least 10 X nominal volume before it will burst Must retain original volume and shape (within 10%) at deflation Must must not adhere to itself and must not twist or kink on itself –must have a mechanism to prevent this. Must clearly indicate that it is Autoclavable on the bag To comply with ISO 5362:2006 or equivalent | Rebreathing reservoir bag 0.5 litre with 15mm F connector - Reusable
| RT13-09-02L |  | Rebreathing reservoir bag 1 litre with 15mm F connector - Reusable
| RT13-09-03L |  | Rebreathing reservoir bag 2 litre with 22mm F connector- Reusable
| RT13-09-04L |  | Rebreathing reservoir bag 2 litre with 22mm F connector- Reusable
CORRUGATED BREATHING CIRCUITS WITH WATER TRAPS AND EXTRA LIMB AND BAG - SINGLE USE
CATEGORY 9M | RT13-09-01M | CORRUGATED BREATHING CIRCUITS WITH WATER TRAPS AND EXTRA LIMB AND BAG - SINGLE USE Purpose: Used to deliver oxygen and Volatile Gases to patient's and to eliminate C02 during anaesthesia Transparent circuit with low resistance, non-extendable, corrugated tubing that is non-kinking, lightweight and crush resistant. The tubing must be pliable and must not retain a shape Shallow corrugations to minimise flow resistance and turbulence. Uninterrupted tubing (not joined). Lightweight water traps centrally on inspiratory and expiratory limb Must have an extra limb with an anti-static reservoir bag that gives good feedback regarding patient lung compliance. The reservoir bag must not twist on itself and must have a mechanism to prevent this. Connectors Malleable 22mm female connectors bonded to the ends of inspiratory and expiratory limbs – must fit securely to the machine. Spare 22mm M/M connector. Y-piece at patient's end. Separate angle-piece with luer lock port Good quality to ensure no leakage or easy disconnections Latex-free, non-pyrogenic, hypoallergenic, non-toxic and non-conductive medical grade plastic No variation in Circuit limb length or Extra limb will be accepted Clinically clean and individually packed, For SINGLE use only To comply with ISO 5367 or equivalent | Corrugated breathing circuit with water traps and extra limb with breathing bag - Single Use - Paediatric > 3 yrs 15mm Circuit: 1.6m; 15mm Extra limb: 1.2m Reservoir bag: 1l
| RT13-09-02M |  | Corrugated breathing circuit with water traps and extra limb with breathing bag - Single use Paediatric > 3 yrs 15mm Circuit: 1.6m; 15mm Extra limb: 1.8m Reservoir bag: 1l
| RT13-09-03M |  | Corrugated breathing circuit with water traps and extra limb with breathing bag - Single Use Adult 22mm Circuit: 1.8m; 22mm Extra limb: 1.2m OD: 22mm Reservoir bag: 2l
| RT13-09-04M |  | Corrugated breathing circuit with water traps and extra limb with breathing bag - Single Use Adult 22mm Circuit: 1.8m; 22mm Extra limb: 1.5m OD: 22mm Reservoir bag: 2l
| RT13-09-05M |  | Corrugated breathing circuit with water traps and extra limb with breathing bag - Single Use Adult 22mm Circuit: 1.8m; 22mm Extra limb: 1.8m OD: 22mm Reservoir bag: 2l
| RT13-09-06M |  | Corrugated breathing circuit with water traps and extra limb with breathing bag - Single Use Adult Long 22mm Circuit: 2.5m; 22mm Extra limb: 1.2m OD: 22mm Reservoir bag: 2l
| RT13-09-07M |  | Corrugated breathing circuit with water traps and extra limb with breathing bag - Single Use Adult Long 22mm Circuit: 2.5m; 22mm Extra limb: 1.5m OD: 22mm Reservoir bag: 2l
| RT13-09-08M |  | Corrugated breathing circuit with water traps and extra limb with breathing bag - Single Use Adult Long 22mm Circuit: 2.5m; 22mm Extra limb: 1.8m OD: 22mm Reservoir bag: 2l
SINGLE CORRUGATED BREATHING TUBING - SINGLE USE
CATEGORY 9N | RT13-09-01N | SINGLE CORRUGATED BREATHING TUBING - SINGLE USE Purpose: Used as extension or extra limb in breathing circuits Transparent circuit with low resistance, non-extendable, corrugated tubing that is non-kinking, lightweight and crush resistant. The tubing must be pliable and must not retain a shape Shallow corrugations to minimise flow resistance and turbulence. Connectors Malleable 22mm female connectors bonded to the end of the tubing Spare connectors to be provided Latex-free, non-pyrogenic, hypoallergenic, non-toxic and non-conductive medical grade plastic No variation in Circuit limb length or Extra limb will be accepted Clinically clean and individually packed, For SINGLE use only To comply with ISO 5367 or equivalent | Single Corrugated Breathing Tubing - Single Use Paediatric 15mm Length: 1.2m Connector: 15mm M/F; 15mm M/M; 15mm F/F
| RT13-09-02N |  | Single Corrugated Breathing Tubing - Single Use Paediatric 15mm Length: 1.8m Connector: 15mm M/F; 15mm M/M; 15mm F/F
| RT13-09-03N |  | Single Corrugated Breathing Tubing - Single Use Adult 22mm Length: 1.2m Connector: 22mm M/F; 22mm M/M; 22mm F/F
| RT13-09-04N |  | Single Corrugated Breathing Tubing - Single Use Adult 22mm Length: 1.8m Connector: 22mm M/F; 22mm M/M; 22mm F/F
MAPLESON F PAEDIATRIC BREATHING CIRCUIT - (Modified Ayre’s T piece)
CATEGORY 9O | RT13-09-01O | MAPLESON F PAEDIATRIC BREATHING CIRCUIT - (Modified Ayre’s T piece) for children < 20-25 kg Purpose: Used to deliver oxygen and volatile gases to patient's and to eliminate C02 during anaesthesia Transparent circuit with low resistance, non-extendable, corrugated tubing that is non-kinking, lightweight and crush resistant. The tubing must be pliable and must not retain a shape with shallow corrugations to minimise flow resistance and turbulence. Tubing must be connected to a straight T-Piece with inspiratory port, expiratory port and patient connection port 22M/15F. Inspiratory tubing: 10mm ± 1.8 - 2.1m Bonded 15mm male connector at the end of the inspiratory limb– must fit securely to the machine. . Expiratory tubing 15mm diameter 0.5m with a 0.5l open-ended reservoir bag to allow gas to escape via the tail of the bag The reservoir bag must give good feedback regarding patient lung compliance and must not twist on itself and must have a mechanism to prevent this. Connectors: Elbow-piece with toggled cap on a luer lock port, capped connector at patient’s end. Good quality to ensure no leakage or easy disconnections Manufactured from transparent non-conductive medical grade plastic that is latex-free non-pyrogenic, hypoallergenic and non-toxic. | Mapleson F breathing circuit - (Modified Ayre’s T piece) for children < 20-25 kg 15mm Circuit ± 1.8 - 2.1m Expiratory Limb 0.4m Reservoir bag: 0.5l open ended
MAPLESON C BREATHING CIRCUIT
CATEGORY 9P | RT13-09-01P | MAPLESON C BREATHING CIRCUIT Purpose: A bagging system for the induction of anaesthesia, resuscitation and patient transport Consists of transparent low resistance, non-extendable, 15mm corrugated tubing that is non-kinking, lightweight and crush resistant. ID: 15mm diameter Length: 1.8m The tubing must be pliable and must not retain a shape with shallow corrugations to minimise flow resistance and turbulence. The tubing must be connected to a straight T-Piece via an inspiratory port. The straight T-Piece must have an angled elbow for patient connection 22M/15F, and an APL valve on the T of the T-Piece and a rebreathing bag Connectors: Elbow-piece with toggled cap on a luer lock port, capped connector at patient’s end. Good quality to ensure no leakage or easy disconnections Manufactured from transparent non-conductive medical grade plastic that is latex-free non-pyrogenic, hypoallergenic and non-toxic. | Mapleson F Breathing Circuit - Paediatric with 1.0 litre Rebreather Bag
| RT13-09-02P |  | Mapleson F Breathing Circuit - Adult with 2.0 litre Rebreather Bag
SMOOTH BORE BREATHING CIRCUITS WITH EXTRA LIMB AND BREATHING BAG - SINGLE USE
CATEGORY 9Q | RT13-09-01Q | SMOOTH BORE BREATHING CIRCUITS WITH EXTRA LIMB AND BREATHING BAG - SINGLE USE Purpose: Used to deliver oxygen and Volatile Gases to patient's and to eliminate C02 during anaesthesia Transparent circuit with low resistance, non-extendable, smooth-bore tubing that is non-kinking, lightweight and crush resistant. The tubing must be pliable and must not retain a shape Spiral outer casing that supports the inner lumen to maintain tube integrity. Must have an extra limb with an anti-static reservoir bag that gives good feedback regarding patient lung compliance. The reservoir bag must not twist on itself and must have a mechanism to prevent this. Connectors Malleable 22mm female connectors bonded to the ends of inspiratory and expiratory limbs – must fit securely to the machine. Spare 22mm M/M connector. Y-piece at patient's end. Separate elbow-piece with capped luer lock port. Good quality to ensure no leakage or easy disconnections Latex-free, non-pyrogenic, hypoallergenic, non-toxic and non-conductive medical grade plastic Clinically clean and individually packed, For SINGLE use only To comply with ISO 5367 or equivalent | Smooth bore breathing circuit with extra limb and breathing bag - Single Use Paediatric > 3yrs 15mm Circuit: ± 1.6m; Extra limb: ± 1.2m Reservoir bag: 1l
| RT13-09-02Q |  | Smooth bore breathing circuit with extra limb and breathing bag - Single Use Paediatric > 3yrs 15mm Circuit: ± 1.6m; Extra limb: ± 1.8m Reservoir bag: 1l
| RT13-09-03Q |  | Smooth bore breathing circuit with extra limb and breathing bag - Single Use Adult 22mm Circuit: ± 1.8m; Extra limb: ± 1.2m OD: Reservoir bag: 2l
| RT13-09-04Q |  | Smooth bore breathing circuit with extra limb and breathing bag - Single Use Adult 22mm Circuit: ± 1.8m; Extra limb: ± 1.5m OD: Reservoir bag: 2l
| RT13-09-05Q |  | Smooth bore breathing circuit with extra limb and breathing bag - Single Use Adult 22mm Circuit: ± 1.8m; Extra limb: ± 1.8m OD: Reservoir bag: 2l
| RT13-09-06Q |  | Smooth bore breathing circuit with extra limb and breathing bag - Single Use Adult Long 22mm Circuit: ± 2.6m; Extra limb: ± 1.2m Reservoir bag: 2l
| RT13-09-07Q |  | Smooth bore breathing circuit with extra limb and breathing bag - Single Use Adult Long 22mm Circuit: ± 2.6m; Extra limb: ± 1.5m Reservoir bag: 2l
| RT13-09-08Q |  | Smooth bore breathing circuit with extra limb and breathing bag - Single Use Adult Long 22mm Circuit: ± 2.6m; Extra limb: ± 1.8m Reservoir bag: 2l
SMOOTH BORE BREATHING CIRCUITS WITH WATER TRAPS AND EXTRA LIMB AND BAG - SINGLE USE
CATEGORY 9R | RT13-09-01R | SMOOTH BORE BREATHING CIRCUITS WITH WATER TRAPS AND EXTRA LIMB AND BAG - SINGLE USE Purpose: Used to deliver oxygen and Volatile Gases to patient's and to eliminate C02 during anaesthesia Transparent circuit with low resistance, non-extendable, smooth-bore tubing that is non-kinking, lightweight and crush resistant. The tubing must be pliable and must not retain a shape Spiral outer casing that supports the inner lumen to maintain tube integrity. Lightweight water traps centrally on inspiratory and expiratory limbs Must have an extra limb with an anti-static reservoir bag that gives good feedback regarding patient lung compliance. The reservoir bag must not twist on itself and must have a mechanism to prevent this. Connectors Malleable 22mm female connectors bonded to the ends of inspiratory and expiratory limbs – must fit securely to the machine. Spare 22mm M/M connector. Y-piece at patient's end. Separate angle-piece with luer lock port Good quality to ensure no leakage or easy disconnections Latex-free, non-pyrogenic, hypoallergenic, non-toxic and non-conductive medical grade plastic Clinically clean and individually packed, For SINGLE use only To comply with ISO 5367 or equivalent | Smooth bore breathing circuit with with water traps, extra limb and breathing bag - Single Use Adult Circuit: ± 1.8m; Extra limb: ± 1.2m OD: 22mm Reservoir bag: 2l
| RT13-09-02R |  | Smooth bore breathing circuit with with water traps, extra limb and breathing bag - Single Use Adult Circuit: ± 1.8m; Extra limb: ± 1.5m OD: 22mm Reservoir bag: 2l
| RT13-09-03R |  | Smooth bore breathing circuit with with water traps, extra limb and breathing bag - Single Use Adult Circuit: ± 1.8m; Extra limb: ± 1.8m OD: 22mm Reservoir bag: 2l
| RT13-09-04R |  | Smooth bore breathing circuit with with water traps, extra limb and breathing bag - Single Use Adult Long Circuit: ± 2.6m; Extra limb: ± 1.2m OD: 22mm Reservoir bag: 2l
| RT13-09-05R |  | Smooth bore breathing circuit with with water traps, extra limb and breathing bag - Single Use Adult Long Circuit: ± 2.6m; Extra limb: ± 1.5m OD: 22mm Reservoir bag: 2l
| RT13-09-06R |  | Smooth bore breathing circuit with with water traps, extra limb and breathing bag - Single Use Adult Long Circuit: ± 2.6m; Extra limb: ± 1.8m OD: 22mm Reservoir bag: 2l
SINGLE SMOOTH-BORE BREATHING TUBING - SINGLE USE
CATEGORY 9S | RT13-09-01S | SINGLE SMOOTH-BORE BREATHING TUBING - SINGLE USE Purpose: Used as extension or extra limb in breathing circuits Transparent circuit with low resistance, non-extendable, smooth-bore tubing that is non-kinking, lightweight and crush resistant. The tubing must be pliable and must not retain a shape Spiral outer casing that supports the inner lumen to maintain tube integrity. Connectors Malleable 22mm female connectors bonded to the end of the tubing Spare connectors to be provided Latex-free, non-pyrogenic, hypoallergenic, non-toxic and non-conductive medical grade plastic Clinically clean and individually packed, For SINGLE use only To comply with ISO 5367 or equivalent | Single Smooth-bore Breathing Tubing - Single Use Adult 22mm Length: ± 1.2m Connector: 22mm M/F; 22mm M/M; 22mm F/F
| RT13-09-02S |  | Single Smooth-bore Breathing Tubing - Single Use Adult 22mm Length: ± 1.8m Connector: 22mm M/F; 22mm M/M; 22mm F/F
| RT13-09-03S |  | Single Smooth-bore Breathing Tubing - Single Use Paediatric 15mm Length: ± 1.2m Connector: 15mm M/F; 15mm M/M; 15mm F/F
| RT13-09-04S |  | Single Smooth-bore Breathing Tubing - Single Use Paediatric 15mm Length: ± 1.8m Connector: 15mm M/F; 15mm M/M; 15mm F/F
CORRUGATED BREATHING CIRCUITS WITH EXTRA LIMB AND BAG - REUSABLE
CATEGORY 9T | RT13-09-01T | CORRUGATED BREATHING CIRCUITS WITH EXTRA LIMB AND BAG - REUSABLE Purpose: Used to deliver oxygen and Volatile Gases to patient's and to eliminate C02 during anaesthesia Consists of a circuit with low resistance, non-extendable, corrugated tubing that is non-kinking, clear and lightweight and crush resistant. The tubing must be pliable and must not retain a shape The corrugations must be shallow to minimise flow resistance and turbulence. Must have an extra limb with an anti-static Reservoir bag that gives good feedback regarding patient lung compliance and does not twist on itself. Must have a mechanism to prevent this. Connectors Malleable 22mm female connectors at the end of inspiratory and expiratory limbs – must fit securely to the machine. Spare 22mm M/M connector. Y-piece at patient's end. Good quality to ensure no leakage or easy disconnections No variation in Circuit limb length or Extra limb will be accepted Manufactured from latex- free, non-pyrogenic hypoallergenic and non-toxic Silicone/ Hydro-tubing . Reusable, autoclavable, individually packed and surgical clean. | Corrugated breathing circuit with extra limb and breathing bag - Reusable- Paediatric > 3 yrs Circuit: ± 1.6m; Extra limb: ± 1.2m OD: 15mm Reservoir bag: 500ml
| RT13-09-02T |  | Corrugated breathing circuit with extra limb and breathing bag - Reusable- Paediatric > 3 yrs Circuit: ± 1.6m; Extra limb: ± 1.8m OD: 15mm Reservoir bag: 500ml
| RT13-09-03T |  | Corrugated breathing circuit with extra limb and breathing bag - Reusable- Paediatric > 3 yrs Circuit: ± 1.6m; Extra limb: ± 1.2m OD: 15mm Reservoir bag: 1l
| RT13-09-04T |  | Corrugated breathing circuit with extra limb and breathing bag - Reusable- Paediatric > 3 yrs Circuit: ± 1.6m; Extra limb: ± 1.8m OD: 15mm Reservoir bag: 1l
| RT13-09-05T |  | Corrugated breathing circuit with extra limb and breathing bag - Reusable - Adult Circuit: ± 1.8m; Extra limb: ± 1.2m OD: 22mm Reservoir bag: 2l
| RT13-09-06T |  | Corrugated breathing circuit with extra limb and breathing bag - Reusable - Adult Circuit: ± 1.8m; Extra limb: ± 1.5m OD: 22mm Reservoir bag: 2l
| RT13-09-07T |  | Corrugated breathing circuit with extra limb and breathing bag - Reusable - Adult Circuit: ± 1.8m; Extra limb: ± 1.8m OD: 22mm Reservoir bag: 2l
SCAVENGING CORRUGATED TUBING
CATEGORY 9U | RT13-09-01U | SCAVENGING CORRUGATED TUBING Consists of transparent corrugated segmented tubing with cuffed repeating segments To fit 22mm M connectors Roll ± 50m. Single use, individually packed and surgical clean. | SCAVENGING Corrugated Tubing - segments repeating every ± 0.45m
| RT13-09-02U |  | SCAVENGING Corrugated Tubing - segments repeating every ± 0.15m
REBREATHING BAGS REUSABLE
CATEGORY 9X | RT13-09-01X | REBREATHING BAGS REUSABLE Purpose: Stores gas between respiratory cycles. Provides information to anaesthetist in manual ventilation regarding pulmonary compliance and airway patency. Consists of alatex-free antistatic, non-conductive reservoir bag with a tapered neck and collar Compliance: 30-60cmH20 when filled to 4 X nominal volume Must hold at least 10 X nominal volume before it will burst Must retain original volume and shape (within 10%) at deflation Must must not adhere to itself and must not twist or kink on itself –must have a mechanism to prevent this. Must clearly indicate that it is Autoclavable on the bag To comply with ISO 5362:2006 or equivalent | Rebreathing reservoir bag 0.5 litre with 15mm F connector - Reusable
| RT13-09-02X |  | Rebreathing reservoir bag 1 litre with 15mm F connector - Reusable
| RT13-09-03X |  | Rebreathing reservoir bag 2 litre with 22mm F connector- Reusable

## Category 10
BID NUMBER | RT13
BID DESCRIPTION | SUPPLY AND DELIVERY OF RESPIRATORY AID AND ANAESTHETIC ACCESSORIES TO THE STATE
CATEGORY 10
ENDOTRACHEAL TUBES
| Item Number | Technical Specification | Product Name and Size | Insert inputs on this column
PERCUTANEOUS TRACHEOSTOMY KITS
CATEGORY 10A | RT13-10-01A | PERCUTANEOUS TRACHEOSTOMY KITS - WITH DILATOR Purpose: Tracheostomy insertion at the bedside Consists of a, radiopaque, cuffed Cricothyroidotomy tube with an inner cannula and a 15mm connector. Must include a double-swivel angled connector with a bronchoscopy port and a hypoallergenic neck tape. The set must have a short-handled scalpel blade; a curved dissecting forceps; a 14-15G puncture needle with a catheter; a kink-resistant guidewire with a guidewire inserter; a short 14-15FR dilator, an angled tracheostomy dilator on a guiding catheter with a safety stop and a loading dilator. The guiding catheter must have clear marking lines to prevent the dilator from being inserted too deeply The tracheostomy tube must be loaded onto the loading dilator. The tube must sit snugly on the loading dilator and must not flange Manufactured from latex-free medical grade metal and plastic Individually packed in peel pouch that is easy to open, SINGLE use, Sterile To comply with latest ISO 5366 | Percutaneous Tracheostomy Kit with Dilator Size 7.0
| RT13-10-02A |  | Percutaneous Tracheostomy Kit with Dilator Size 7.5
| RT13-10-03A |  | Percutaneous Tracheostomy Kit with Dilator Size 8.0
| RT13-10-04A |  | Percutaneous Tracheostomy Kit with Dilator Size 8.5
PERCUTANEOUS TRACHEOSTOMY KITS - WITH DILATING FORCEPS
CATEGORY 10B | RT13-10-01B | PERCUTANEOUS TRACHEOSTOMY KITS - WITH DILATING FORCEPS Purpose: Tracheostomy insertion at the bedside Consists of a, radiopaque, cuffed Cricothyroidotomy tube with an inner cannula and a 15mm connector. Must include a double-swivel angled connector with a bronchoscopy port and a hypoallergenic neck tape. The set must have a short-handled scalpel blade; a curved dissecting forceps; a 14-15G puncture needle with a catheter; a kink-resistant guidewire with a guidewire inserter; a short 14-15FR dilator, angled dilating forceps with a specialised notch for the guidewire. The angled dilating forceps must have an ergonomic handle and must pass through the skin and tissue easily. . The forceps must lock over the guidewire catheter and allow for single step dilation. Manufactured from latex-free medical grade metal and plastic Individually packed in peel pouch that is easy to open, SINGLE use, Sterile To comply with latest ISO 5366 | Percutaneous Tracheostomy kits with dilating forceps Size 7.0
| RT13-10-02B |  | Percutaneous Tracheostomy kits with dilating forceps Size 8.0
| RT13-10-03B |  | Percutaneous Tracheostomy kits with dilating forceps Size 9.0