## 1 Title
ANNEXURE 3A TECHNICAL SPECIFICATIONS AND PRICING SCHEDULE
BID NUMBER: RT233-2023
SUPPLY AND DELIVERY OF WHEELCHAIRS, COMMODES, SEAT CUSHIONS, CUSHION COVERS, SEATING SYSTEMS AND POSITIONERS TO THE STATE
CLOSING DATE: 23 JANUARY 2023

## 2 Category Profiles
SUPPLY AND DELIVERY OF WHEELCHAIRS, SEATING SYSTEMS, POSITIONERS AND COMMODES TO THE STATE FOR A PERIOD OF 60 MONTHS
SUMMARY OF CATEGORIES |  |  | Quantity (Different Sizes)
| Item No. | WHEELCHAIRS AND COMMODES | 177
Category 1 | RT233-1.1 | INDOOR / LOW-ACTIVE BASIC FOLDING FRAME WHEELCHAIRS | 26
| RT233-1.2 | BASIC FOLDING FRAME, STANDARD BACK HEIGHT, LOW SLUNG SEAT | 9
| RT233-1.3 | BASIC FOLDING FRAME, SEMI-RECLINER | 7
| RT233-1.4 | BARIATRIC FOLDING FRAME WHEELCHAIRS | 3
| RT233-1.5 | URBAN ACTIVE WHEELCHAIRS | 36
| RT233-1.6 | HYBRID ACTIVE WHEELCHAIRS | 13
| RT233-1.7 | RURAL ACTIVE WHEELCHAIRS | 4
| RT233-1.8 | ADJUSTABLE POSTURE SUPPORT WHEELCHAIRS | 14
| RT233-1.9 | MOTORISED FOLDING FRAME WHEELCHAIRS | 36
| RT233-1.10 | MOTORISED HEAVY-DUTY RIGID FRAME WHEELCHAIR | 5
| RT233-1.11 | BUGGIES | 10
| RT233-1.12 | PORTER WHEELCHAIRS | 2
| RT233-1.13 | COMMODES | 4
| RT233-1.14 | TRANSPORTERS | 8
SUMMARY OF CATEGORIES |  |  | Quantity (Different Sizes)
Category 2 | Item No. | WHEELCHAIR CUSHIONS AND FOAM | 169
| RT233-2.1 | BASIC FOAM CUSHIONS | 122
| RT233-2.2 | FLUID FILLED CUSHION | 10
| RT233-2.3 | AIR FILLED CUSHION | 17
| RT233-2.4 | COMPRESSED FOAM WEDGE | 2
| RT233-2.5 | COMPRESSED FOAM SHEET | 8
| RT233-2.6 | HIGH DENSITY FOAM SHEET | 10
SUMMARY OF CATEGORIES |  |  | Quantity (Different Sizes)
Category 3 | Item No. | WHEELCHAIR CUSHIONS COVERS | 253
| RT233-3.1 | CUSHION COVER FOR CONTOURED BASED, BASIC FOAM | 80
| RT233-3.2 | CUSHION COVER FOR CONTOURED BASED, BASIC FOAM EXTENDED | 64
| RT233-3.3 | CUSHION COVER FOR STRAIGHT BASE, BASIC FOAM | 104
| RT233-3.4 | CUSHION COVER FABRIC | 2
| RT233-3.5 | FLUID FILLED CUSHION COVER | 2
| RT233-3.6 | AIR FILLED CUSHION COVER | 1
SUMMARY OF CATEGORIES |  |  | Quantity (Different Sizes)
Category 4 | Item No. | WHEELCHAIR CUSHIONS WITH COVERS | 242
| RT233-4.1 | BASIC FOAM CUSHIONS WITH NYLON COVERS | 121
| RT233-4.2 | BASIC FOAM CUSHIONS WITH WATERPROOF COVERS | 121
SUMMARY OF CATEGORIES |  |  | Quantity (Different Sizes)
Category 5 | Item No. | BACKREST AND SEATING SYSTEMS | 194
| RT233-5.1 | TENSION ADJUSTABLE BACK | 108
| RT233-5.2 | TENSION ADJUSTABLE BACKREST WITH RIGID HEIGHT ADJUSTABLE FULL-LENGTH LATERALS | 18
| RT233-5.3 | RIGID ADJUSTABLE BACK WITH FULL LENGTH LATERAL SUPPORT | 23
| RT233-5.4 | RIGID ADJUSTABLE BACK | 12
| RT233-5.5 | ADJUSTABLE MODULAR BACK | 10
| RT233-5.6 | LOW COST MODULAR SEATING SYSTEM | 8
| RT233-5.7 | TRUCK SIDE SUPPORT | 7
| RT233-5.8 | BASE SEATING SYSTEM | 4
| RT233-5.9 | POSTURE SUPPORT SEAT SYSTEM | 4
SUMMARY OF CATEGORIES |  |  | Quantity (Different Sizes)
Category 6 | Item No. | TRAY TABLES | 18
| RT233-6.1 | TRAY TABLE, BASIC WOOD | 9
| RT233-6.2 | TRAY TABLE, BASIC, PLASTIC | 9
SUMMARY OF CATEGORIES |  |  | Quantity (Different Sizes)
Category 7 | Item No. | POSITIONERS | 17
| RT233-7.1 | SIDELYERS | 4
| RT233-7.2 | PRONE WEDGES | 4
| RT233-7.3 | STANDING FRAMES | 9
|  | Total Items | 1070

## 3 Main Pricing Schedule
SUPPLY AND DELIVERY OF WHEELCHAIRS, SEATING SYSTEMS, POSITIONERS AND COMMODES TO THE STATE FOR A PERIOD OF 60 MONTHS
All items being tendered for must comply with the following technical specification, failure to comply with any of the conditions set out below may result in immediate disqualification. Please note that where the technical specification calls for certification, this certification must accompany the bid, failure to provide such certification will result in immediate disqualification. Bidders must give details of items being tendered for in the COMMENTS section where there are variation for further deliberation. Complete relevant columns for each line item
1. WHEELCHAIRS AND COMMODES
Item No. | Item Code | Description | Quantity | Are you bidding for this item?: Yes or No | Minimum order Quantities (where applicable) | Import Percentage (where applicable) | Local Percentage (where applicable) | Country of Manufacture (Local or Imported) | Foreign Currency Name | Brand Name | Model Number | Delivery Period (Days) | Unit Of Measure (UOM) | Equipment Expected Lifespan (where applicable) | Unit Price | Comments
| RT233-1.1 | INDOOR / LOW-ACTIVE BASIC FOLDING FRAME WHEELCHAIRS | 26
| RT233-1.1.1 | Basic folding frame, standard back height | 9
1 | RT233-1.1.1a | Size: 51 cm (20") wide | 1
2 | RT233-1.1.1b | Size: 46 cm (18") wide | 1
3 | RT233-1.1.1c | Size: 43 cm (17”) wide | 1
4 | RT233-1.1.1d | Size: 41 cm (16") wide | 1
5 | RT233-1.1.1e | Size: 38 cm (15") wide | 1
6 | RT233-1.1.1f | Size: 36 cm (14") wide | 1
7 | RT233-1.1.1g | Size: 30 cm (12") wide with 40 cm seat frame length | 1
8 | RT233-1.1.1h | Size: 30 cm (12") wide with 30 cm seat frame length | 1
9 | RT233-1.1.1i | Size: 25 cm (10”) wide with 30 cm seat frame length | 1
| RT233-1.1.2 | Basic folding frame, tall back height | 8
10 | RT233-1.1.2a | Size: 51 cm (20") wide | 1
11 | RT233-1.1.2b | Size: 46 cm (18") wide | 1
12 | RT233-1.1.2c | Size: 43 cm (17") wide | 1
13 | RT233-1.1.2d | Size: 41 cm (16") wide | 1
14 | RT233-1.1.2e | Size: 38 cm (15") wide | 1
15 | RT233-1.1.2f | Size: 36 cm (14") wide | 1
16 | RT233-1.1.2g | Size: 30 cm (12") wide with 40 cm seat frame length | 1
17 | RT233-1.1.2h | Size: 30 cm (12") wide with 30 cm seat frame length | 1
| RT233-1.1.3 | Basic folding frame, short back height | 9
18 | RT233-1.1.3a | Size: 51 cm (20") wide | 1
19 | RT233-1.1.3b | Size: 46 cm (18") wide | 1
20 | RT233-1.1.3c | Size: 43 cm (17”) wide | 1
21 | RT233-1.1.3d | Size: 41 cm (16") wide | 1
22 | RT233-1.1.3e | Size: 38 cm (15") wide | 1
23 | RT233-1.1.3f | Size: 36 cm (14") wide | 1
24 | RT233-1.1.3g | Size: 30 cm (12") wide with 40 cm seat frame length | 1
25 | RT233-1.1.3h | Size: 30 cm (12”) wide with 30 cm seat frame length | 1
26 | RT233-1.1.3i | Size: 25 cm (10”) wide with 30 cm seat frame length | 1
| RT233-1.2 | BASIC FOLDING FRAME, STANDARD BACK HEIGHT, LOW SLUNG SEAT | 9
| RT233-1.2.1 | Basic folding frame, standard back height, low slung seat
27 | RT233-1.2.1a | Size: 51 cm (20") wide | 1
28 | RT233-1.2.1a | Size: 46 cm (18") wide | 1
29 | RT233-1.2.1a | Size: 43 cm (17”) wide | 1
30 | RT233-1.2.1a | Size: 41 cm (16") wide | 1
31 | RT233-1.2.1a | Size: 38 cm (15") wide | 1
32 | RT233-1.2.1a | Size: 36 cm (14") wide | 1
33 | RT233-1.2.1a | Size: 30 cm (12") wide with 40 cm seat frame length | 1
34 | RT233-1.2.1a | Size: 30 cm (12”) wide with 30 cm seat frame length | 1
35 | RT233-1.2.1a | Size: 25 cm (10”) wide with 30 cm seat frame length | 1
| RT233-1.3 | BASIC FOLDING FRAME, SEMI-RECLINER | 7
| RT233-1.3.1 | Basic folding frame, semi-recliner
36 | RT233-1.3.1a | Size: 51 cm (20") wide | 1
37 | RT233-1.3.1b | Size: 46 cm (18") wide | 1
38 | RT233-1.3.1c | Size: 43 cm (17”) wide | 1
39 | RT233-1.3.1d | Size: 41 cm (16") wide | 1
40 | RT233-1.3.1e | Size: 38 cm (15”) wide | 1
41 | RT233-1.3.1f | Size: 36 cm (14”) wide | 1
42 | RT233-1.3.1g | Size: 30 cm (12”) wide | 1
| RT233-1.4 | BARIATRIC FOLDING FRAME WHEELCHAIRS | 3
| RT233-1.4.1 | Bariatric folding frame chairs
43 | RT233-1.4.1a | Size: 51cm (20”) | 1
44 | RT233-1.4.1b | Size: 56cm ( 22”) | 1
45 | RT233-1.4.1c | Size: 61cm ( 24”) | 1
| RT233-1.5 | URBAN ACTIVE WHEELCHAIRS | 36
| RT233-1.5.1 | Adjustable folding frame urban active wheelchair | 9
46 | RT233-1.5.1a | Size: 51 cm (20") wide with 43 + 1 cm seat frame length | 1
47 | RT233-1.5.1b | Size: 46 cm (18") wide with 43 + 1 cm seat frame length | 1
48 | RT233-1.5.1c | Size: 47 cm (17") wide with 43 + 1 cm seat frame length | 1
49 | RT233-1.5.1d | Size: 41 cm (16") wide with 43 + 1 cm seat frame length | 1
50 | RT233-1.5.1e | Size: 38 cm (15") wide with 43 + 1 cm seat frame length | 1
51 | RT233-1.5.1f | Size: 36 cm (14") wide with 43 + 1 cm seat frame length | 1
52 | RT233-1.5.1g | Size: 30 cm (12") wide with 43 + 1 cm seat frame length | 1
53 | RT233-1.5.1h | Size: 30 cm (12") wide with 33 + 1 cm seat frame length | 1
54 | RT233-1.5.1i | Size: 25 cm (10") wide with 33 + 1 cm seat frame length | 1
| RT233-1.5.2 | Adjustable folding frame urban active wheelchair - light weight | 7
55 | RT233-1.5.2a | Size: 51 cm (20”) wide with 43 + 1 cm seat frame length | 1
56 | RT233-1.5.2b | Size: 46 cm (18") wide with 43 + 1 cm seat frame length | 1
57 | RT233-1.5.2c | Size: 41 cm (16") wide with 43 + 1 cm seat frame length | 1
58 | RT233-1.5.2d | Size: 36 cm (14") wide with 43 + 1 cm seat frame length | 1
59 | RT233-1.5.2e | Size: 30 cm (12”) wide with 43 + 1 cm seat frame length | 1
60 | RT233-1.5.2f | Size: 30 cm (12”) wide with 33 + 1 cm seat frame length | 1
61 | RT233-1.5.2g | Size: 25 cm (10”) wide with 33 + 1 cm seat frame length | 1
| RT233-1.5.3 | Adjustable rigit frame wheelchair, 4-wheeler urban active chair with standard or long frame | 14
| RT233-1.5.3.1 | Adjustable rigit frame wheelchair, 4-wheeler urban active chair with standard frame | 8
62 | RT233-1.5.3.1a | Size: 46 cm (18") wide | 1
63 | RT233-1.5.3.1b | Size: 43 cm (17") wide | 1
64 | RT233-1.5.3.1c | Size: 41 cm (16") wide | 1
65 | RT233-1.5.3.1d | Size: 38 cm (15") wide | 1
66 | RT233-1.5.3.1e | Size: 35 cm (14") wide | 1
67 | RT233-1.5.3.1f | Size: 30 cm (12") wide with 40 cm seat frame length | 1
68 | RT233-1.5.3.1g | Size: 30 cm (12") wide with 30 cm seat frame length | 1
69 | RT233-1.5.3.1h | Size: 25 cm (10”) wide with 30 cm seat frame length but 25 cm seat | 1
| RT233-1.5.3.2 | Adjustable rigit frame wheelchair, 4-wheeler urban active chair with long frame | 6
70 | RT233-1.5.3.2a | Size: 46 cm (18") wide | 1
71 | RT233-1.5.3.2b | Size: 43 cm (17") wide | 1
72 | RT233-1.5.3.2c | Size: 41 cm (16") wide | 1
73 | RT233-1.5.3.2d | Size: 38cm (15") wide | 1
74 | RT233-1.5.3.2e | Size: 35 cm (14") wide | 1
75 | RT233-1.5.3.2f | Size: 30 cm (12") wide with 40 cm seat frame length | 1
| RT233-1.6 | HYBRID ACTIVE WHEELCHAIRS | 13
| RT233-1.6.1 | Adjustable, 3-wheeler hybrid | 9
76 | RT233-1.6.1a | Size: 46 cm (18") wide | 1
77 | RT233-1.6.1b | Size: 43 cm (17") wide | 1
78 | RT233-1.6.1c | Size: 41 cm (16") wide | 1
79 | RT233-1.6.1d | Size: 38 cm (15") wide | 1
80 | RT233-1.6.1e | Size: 35 cm (14") wide | 1
81 | RT233-1.6.1f | Size: 34 cm (13”) wide | 1
82 | RT233-1.6.1g | Size: 30 cm (12") wide with 40 cm seat frame length | 1
83 | RT233-1.6.1h | Size: 30 cm (12") wide with 30cm seat frame length | 1
84 | RT233-1.6.1i | Size: 25 cm (10”) wide with 25 cm seat frame length | 1
| RT233-1.6.2 | Adjustable folding frame, hybrid with long wheel base | 4
85 | RT233-1.6.2a | Size: 51 cm (20”) | 1
86 | RT233-1.6.2b | Size: 46 cm (18”) | 1
87 | RT233-1.6.2c | Size: 41 cm (16”) | 1
88 | RT233-1.6.2d | Size: 36 cm (14”) | 1
| RT233-1.7 | RURAL ACTIVE WHEELCHAIRS | 4
| RT233-1.7.1 | Adjustable 3-wheeler
89 | RT233-1.7.1a | Size: 48 cm (18.9") wide | 1
90 | RT233-1.7.1b | Size: 44 cm (17.3") wide | 1
91 | RT233-1.7.1c | Size: 40 cm (15.7") wide | 1
92 | RT233-1.7.1d | Size: 36 cm (14.2") wide | 1
| RT233-1.8 | ADJUSTABLE POSTURE SUPPORT WHEELCHAIRS | 14
| RT233-1.8.1 | Adjustable posture support wheelchair, self-propelling with modular detachable seating system (Buggy type) | 6
93 | RT233-1.8.1a | Size: 12” short Seat 30 cm wide x 25 cm - 32 cm adjustable depth | 1
94 | RT233-1.8.1b | Size:12” standard Seat 30 cm wide x 30 cm - 40 cm adjustable depth | 1
95 | RT233-1.8.1c | Size: 14” short Seat 35 cm wide x 25 cm - 35 cm adjustable depth | 1
96 | RT233-1.8.1d | Size: 14” standard Seat 35 cm wide x 35 cm - 45 cm adjustable depth | 1
97 | RT233-1.8.1e | Size: 16” standard Seat 40 cm wide x 35 cm - 45 cm adjustable depth | 1
98 | RT233-1.8.1f | Size: 18” standard Seat 45 cm wide x 35 cm - 45 cm adjustable depth | 1
| RT233-1.8.2 | Peadiatric seating system: self propelling paediatric wheelchair with adjustable posture support backrest | 8
99 | RT233-1.8.2a | Size: 45cm (18”) wide with adjustable seat frame length 35 – 45cm | 1
100 | RT233-1.8.2b | Size: 41cm (16”) wide with adjustable seat frame length 35 – 45cm | 1
101 | RT233-1.8.2c | Size: 36 cm (14”) wide with adjustable seat frame length 35 – 45cm | 1
102 | RT233-1.8.2d | Size: 36cm (14”) wide with adjustable seat frame length 25 – 35cm | 1
103 | RT233-1.8.2e | Size: 30cm (12”) wide with adjustable seat frame length 30 – 40cm | 1
104 | RT233-1.8.2f | Size: 30cm (12”) wide with adjustable seat frame length 25 – 35cm | 1
105 | RT233-1.8.2g | Size: 25cm (10”) wide with adjustable seat frame length 22 – 35cm | 1
106 | RT233-1.8.2h | Size: 20cm (8”) wide with adjustable seat frame length 17 – 25cm | 1
| RT233-1.9 | MOTORISED FOLDING FRAME WHEELCHAIRS | 36
| RT233-1.9.1 | Motorised wheelchair, folding frame, standard back height with hand control | 9
107 | RT233-1.9.1a | Size: 51 cm (20") wide | 1
108 | RT233-1.9.1b | Size: 46 cm (18") wide | 1
109 | RT233-1.9.1c | Size: 43 cm (17”) wide | 1
110 | RT233-1.9.1d | Size: 41 cm (16") wide | 1
111 | RT233-1.9.1e | Size: 38 cm (15") wide | 1
112 | RT233-1.9.1f | Size: 36 cm (14") wide | 1
113 | RT233-1.9.1g | Size: 34 cm (13") wide with 40 cm seat frame length | 1
114 | RT233-1.9.1h | Size: 34 cm (13") wide with 30 cm seat frame length | 1
115 | RT233-1.9.1i | Size: 30 cm (12”) wide with 30 cm seat frame length | 1
| RT233-1.9.2 | Motorised wheelchair, folding frame, tall back height with hand control | 9
116 | RT233-1.9.2a | Size: 51 cm (20") wide | 1
117 | RT233-1.9.2b | Size: 46 cm (18") wide | 1
118 | RT233-1.9.2c | Size: 43 cm (17”) wide | 1
119 | RT233-1.9.2d | Size: 41 cm (16") wide | 1
120 | RT233-1.9.2e | Size: 38 cm (15") wide | 1
121 | RT233-1.9.2f | Size: 36 cm (14") wide | 1
122 | RT233-1.9.2g | Size: 34 cm (13") wide with 40 cm seat frame length | 1
123 | RT233-1.9.2h | Size: 34 cm (13") wide with 30 cm seat frame length | 1
124 | RT233-1.9.2i | Size: 30 cm (12”) wide with 30 cm seat frame length | 1
| RT233-1.9.3 | Motorised wheelchair, adjustable folding frame with adjustable tilt in space with hand control | 9
125 | RT233-1.9.3a | Size: 51 cm (20") wide | 1
126 | RT233-1.9.3b | Size: 46 cm (18") wide | 1
127 | RT233-1.9.3c | Size: 43 cm (17”) wide | 1
128 | RT233-1.9.3d | Size: 41 cm (16") wide | 1
129 | RT233-1.9.3e | Size: 38 cm (15") wide | 1
130 | RT233-1.9.3f | Size: 36 cm (14") wide | 1
131 | RT233-1.9.3g | Size: 34 cm (13") wide with 40 cm seat frame length | 1
132 | RT233-1.9.3h | Size: 34 cm (13") wide with 30 cm seat frame length | 1
133 | RT233-1.9.3i | Size: 30 cm (12”) wide with 30 cm seat frame length | 1
| RT233-1.9.4 | Motorised wheelchair, adjustable folding frame with adjustable tilt in space and adjustable backrest recline | 9
134 | RT233-1.9.4a | Size: 51 cm (20") wide with 43 + 1 cm seat frame length | 1
135 | RT233-1.9.4b | Size: 46 cm (18") wide with 43 + 1 cm seat frame length | 1
136 | RT233-1.9.4c | Size: 43 cm (17”) wide with 43 + 1 cm seat frame length | 1
137 | RT233-1.9.4d | Size: 41 cm (16") wide with 43 + 1 cm seat frame length | 1
138 | RT233-1.9.4e | Size: 38 cm (15") wide with 43 + 1 cm seat frame length | 1
139 | RT233-1.9.4f | Size: 36 cm (14") wide with 43 + 1 cm seat frame length | 1
140 | RT233-1.9.4g | Size: 30 cm (12") wide with 43 + 1 cm seat frame length | 1
141 | RT233-1.9.4h | Size: 30 cm (12") wide with 33 + 1 cm seat frame length | 1
142 | RT233-1.9.4i | Size: 25 cm (10") wide with 33 + 1 cm seat frame length | 1
| RT233-1.10 | MOTORISED HEAVY-DUTY RIGID FRAME WHEELCHAIR | 5
| RT233-1.10.1 | Heavy duty rigid frame
143 | RT233-1.10.1a | Size: 51 cm (20") standard, Seat 51 cm wide x 40 cm – 50 cm deep | 1
144 | RT233-1.10.1a | Size: 46 cm (18") standard, Seat 46 cm wide x 40 cm – 50 cm deep | 1
145 | RT233-1.10.1a | Size: 43 cm (17") standard, Seat 41 cm wide x 40 cm – 50 cm deep | 1
146 | RT233-1.10.1a | Size: 41 cm (16") standard, Seat 41 cm wide x 40 cm – 50 cm deep | 1
147 | RT233-1.10.1a | Size: 36 cm (14") standard, Seat 36 cm wide x 40 cm – 50 cm deep | 1
| RT233-1.11 | BUGGIES | 10
| RT233-1.11.1 | Modular, rigid buck, foam adjustable posture support buggy with rigid frame | 6
148 | RT233-1.11.1a | Size: Seat width 25cm, depth 24cm, height 40cm | 1
149 | RT233-1.11.1b | Size: Seat width 27cm, depth 24cm, height 40cm | 1
150 | RT233-1.11.1c | Size: Seat width 29cm, depth 30cm, height 40cm | 1
151 | RT233-1.11.1d | Size: Seat width 32cm, depth 36cm, height 49cm | 1
152 | RT233-1.11.1e | Size: Seat width 32cm, depth 45cm, height 54cm | 1
153 | RT233-1.11.1f | Size: Seat width 39cm, depth 49cm, height 55cm | 1
| RT233-1.11.2 | Modular, rigid adjustable posture support buggy with folding frame (positioning buggy with recline option) | 4
154 | RT233-1.11.2a | Size: Back height 31cm-41cm; Seat width 18cm; Seat lengths 17cm-26cm; Trunk side support height range 17cm-28cm. Trunk side support width range 16cm-21 cm | 1
155 | RT233-1.11.2b | Size: Back height 28cm-38cm; Seat width 23cm; Seat lengths 24cm-32cm. Trunk side support height range 21cm-28cm. Trunk side support width range 18cm-24cm. | 1
156 | RT233-1.11.2c | Size: Back height 37cm-44cm; Seat width 25cm; Seat lengths 29cm-36cm. Trunk side support height range 23cm-33cm. Trunk side support width range 18cm-24cm | 1
157 | RT233-1.11.2d | Size: Back height 40cm-50cm; Seat width 30cm; Seat lengths 37cm-47cm. Trunk side support height range 28cm-35cm. Trunk side support width range 23cm-29cm | 1
| RT233-1.12 | PORTER WHEELCHAIRS | 2
Porter chair for institutional use
158 | RT233-1.12.1 | Size: 51 cm (20”) wide | 1
Porter chair for institutional use with drip stand holder and T-rod
159 | RT233-1.12.2 | Size: 51 cm (20”) wide | 1
| RT233-1.13 | COMMODES | 4
Commode, basic
160 | RT233-1.13.1 | Size: Width range 46cm – 51cm, seat depth 40 cm | 1
Aluminium toilet and shower commode (wheelchair type)
161 | RT233-1.13.2 | Size: Width range 46cm – 51cm, seat depth 40 cm | 1
162 | RT233-1.13.3 | Commode, wheelchair type | 1
Commode, foldable
163 | RT233-1.13.4 | Size: Width 46cm | 1
| RT233-1.14 | TRANSPORTERS | 8
| RT233-1.14.1 | Transporter buggy - folding / collapsible type | 4
164 | RT233-1.14.1a | Size: Seat width: 10” and Approximately 20-40cm | 1
165 | RT233-1.14.1b | Size: Seat width: 12” and Approximately 20-40cm | 1
166 | RT233-1.14.1c | Size: Seat width: 14” and Approximately 20-40cm | 1
167 | RT233-1.14.1d | Size: Seat width: 15” and Approximately 20-40cm | 1
| RT233-1.14.2 | Transporter with side frame covers | 3
168 | RT233-1.14.2a | Size: Seat width: 12” and Approximately 40cm | 1
169 | RT233-1.14.2b | Size: Seat width: 14” and Approximately 40cm | 1
170 | RT233-1.14.2c | Size: Seat width: 15” and Approximately 40cm | 1
|  | Adult Transporter | 1
171 | RT233-1.14.3 | Size: Seat width of 17” to 18”. | 1
2. WHEELCHAIR CUSHIONS AND FOAM
Item No. | Item Code | Description | Quantity | Are you bidding for this item?: Yes or No | Minimum order Quantities (where applicable) | Import Percentage (where applicable) | Local Percentage (where applicable) | Country of Manufacture (Local or Imported) | Foreign Currency Name | Brand Name | Model Number | Delivery Period (Days) | Unit Of Measure (UOM) | Equipment Expected Lifespan (where applicable) | Unit Price | Comments
| RT233-2.1 | BASIC FOAM CUSHIONS | 122
| RT233-2.1.1 | Contoured Base, Thick pressure care cushion | 11
172 | RT233-2.1.1a | Size: 61 cm wide; 43 cm long; 9 cm high in front | 1
173 | RT233-2.1.1b | Size: 56 cm wide; 43 cm long; 9 cm high in front | 1
174 | RT233-2.1.1c | Size: 51 cm wide; 43 cm long; 9 cm high in front | 1
175 | RT233-2.1.1d | Size: 46 cm wide; 43 cm long; 9 cm high in front | 1
176 | RT233-2.1.1e | Size: 43 cm wide; 43 cm long; 9 cm high in front | 1
177 | RT233-2.1.1f | Size: 41 cm wide; 43 cm long; 9 cm high in front | 1
178 | RT233-2.1.1g | Size: 38 cm wide; 43 cm long; 9 cm high in front | 1
179 | RT233-2.1.1h | Size: 36 cm wide; 43 cm long; 9 cm high in front | 1
180 | RT233-2.1.1i | Size: 30 cm wide; 43 cm long; 9 cm high in front | 1
181 | RT233-2.1.1j | Size: 30 cm wide; 30 cm long; 6.5 cm high in front | 1
182 | RT233-2.1.1k | Size: 25 cm wide; 30 cm long; 6 cm high in front | 1
| RT233-2.1.2 | Contoured Base, Wedged, thick pressure care cushion | 10
183 | RT233-2.1.2a | Size: 61 cm wide; 43 cm long; 6 cm high at rear, 11 cm high in front | 1
184 | RT233-2.1.2b | Size: 56 cm wide; 43 cm long; 6 cm high at rear, 11 cm high in front | 1
185 | RT233-2.1.2c | Size: 51 cm wide; 43 cm long; 6 cm high at rear, 11 cm high in front | 1
186 | RT233-2.1.2d | Size: 46 cm wide; 43 cm long; 6 cm high at rear, 11 cm high in front | 1
187 | RT233-2.1.2e | Size: 41 cm wide; 43 cm long; 6 cm high at rear, 11 cm high in front | 1
188 | RT233-2.1.2f | Size: 38 cm wide; 43 cm long; 6 cm high at rear, 11 cm high in front | 1
189 | RT233-2.1.2g | Size: 36 cm wide; 43 cm long; 6 cm high at rear, 11 cm high in front | 1
190 | RT233-2.1.2h | Size: 30 cm wide; 43 cm long; 6 cm high at rear, 11 cm high in front | 1
191 | RT233-2.1.2i | Size: 30 cm wide; 30 cm long; 4 cm high at rear, 8.5 cm high in front | 1
192 | RT233-2.1.2j | Size: 25 cm wide; 30 cm long; 4 cm high at rear, 7.5 cm high in front | 1
| RT233-2.1.3 | Contoured Base, Basic Thin Foam, positioner cushion | 9
193 | RT233-2.1.3a | Size: 51 cm wide; 43 cm long; 7 cm high | 1
194 | RT233-2.1.3b | Size: 46 cm wide; 43 cm long; 7 cm high | 1
195 | RT233-2.1.3c | Size: 43 cm wide; 43 cm long; 7 cm high | 1
196 | RT233-2.1.3d | Size: 41 cm wide; 43 cm long; 7 cm high | 1
197 | RT233-2.1.3e | Size: 38 cm wide; 43 cm long; 7 cm high | 1
198 | RT233-2.1.3f | Size: 36 cm wide; 43 cm long; 7 cm high | 1
199 | RT233-2.1.3g | Size: 30 cm wide; 43 cm long; 6.5 cm high | 1
200 | RT233-2.1.3h | Size: 30 cm wide; 33 cm long; 6.5 cm high | 1
201 | RT233-2.1.3i | Size: 25 cm wide; 28 cm long; 6 cm high | 1
| RT233-2.1.4 | Contoured Base, Wedged, thin foam positioner | 9
202 | RT233-2.1.4a | Size: 51 cm wide; 43 cm long; 4 cm high at rear, 9 cm high in front | 1
203 | RT233-2.1.4b | Size: 46 cm wide; 43 cm long; 4 cm high at rear, 9 cm high in front | 1
204 | RT233-2.1.4c | Size: 43 cm wide; 43 cm long; 4 cm high at rear, 9 cm high in front | 1
205 | RT233-2.1.4d | Size: 41 cm wide; 43 cm long; 4 cm high at rear, 9 cm high in front | 1
206 | RT233-2.1.4e | Size: 38 cm wide; 43 cm long; 4 cm high at rear, 9 cm high in front | 1
207 | RT233-2.1.4f | Size: 36 cm wide; 43 cm long; 4 cm high at rear, 9 cm high in front | 1
208 | RT233-2.1.4g | Size: 30 cm wide; 43 cm long; 4 cm high at rear, 9 cm high in front | 1
209 | RT233-2.1.4h | Size: 30 cm wide; 33 cm long; 4 cm high at rear, 8 cm high in front | 1
210 | RT233-2.1.4i | Size: 25 cm wide; 28 cm long; 4 cm high at rear, 7.5cm high in front | 1
| RT233-2.1.5 | Contoured Base, Thick pressure care foam cushion, extended length | 9
211 | RT233-2.1.5a | Size: 61 cm wide; 50 cm long; 9 cm high | 1
212 | RT233-2.1.5b | Size: 56 cm wide; 50 cm long; 9 cm high | 1
213 | RT233-2.1.5c | Size: 51 cm wide; 50 cm long; 9 cm high | 1
214 | RT233-2.1.5d | Size: 46 cm wide; 50 cm long; 9 cm high | 1
215 | RT233-2.1.5e | Size: 43 cm wide; 50 cm long; 9 cm high | 1
216 | RT233-2.1.5f | Size: 41 cm wide; 50 cm long; 9 cm high | 1
217 | RT233-2.1.5g | Size: 38 cm wide; 50 cm long; 9 cm high | 1
218 | RT233-2.1.5h | Size: 36 cm wide; 50 cm long; 8.5 cm high | 1
219 | RT233-2.1.5i | Size: 30 cm wide; 50 cm long; 8.5 cm high | 1
| RT233-2.1.6 | Contoured Base, Wedged, thick pressure care foam cushion, extended length | 9
220 | RT233-2.1.6a | Size: 61 cm wide; 50 cm long; 6 cm high at rear, 11 cm high in front | 1
221 | RT233-2.1.6b | Size: 56 cm wide; 50 cm long; 6 cm high at rear, 11 cm high in front | 1
222 | RT233-2.1.6c | Size: 51 cm wide; 50 cm long; 6 cm high at rear, 11 cm high in front | 1
223 | RT233-2.1.6d | Size: 46 cm wide; 50 cm long; 6 cm high at rear, 11 cm high in front | 1
224 | RT233-2.1.6e | Size: 43 cm wide; 50 cm long; 6 cm high at rear, 11 cm high in front | 1
225 | RT233-2.1.6f | Size: 41 cm wide; 50 cm long; 6 cm high at rear, 11 cm high in front | 1
226 | RT233-2.1.6g | Size: 38 cm wide; 50 cm long; 6 cm high at rear, 11cm high in front | 1
227 | RT233-2.1.6h | Size: 36 cm wide; 50 cm long; 6 cm high at rear, 11 cm high in front | 1
228 | RT233-2.1.6i | Size: 30 cm wide; 50 cm long; 6 cm high at rear, 10.5 cm high in front | 1
| RT233-2.1.7 | Contoured Base, Thin Basic positioner foam cushion, extended length | 7
229 | RT233-2.1.7a | Size: 51 cm wide; 50 cm long; 7 cm high in front | 1
230 | RT233-2.1.7b | Size: 46 cm wide; 50 cm long; 7 cm high in front | 1
231 | RT233-2.1.7c | Size: 43 cm wide; 50 cm long; 7 cm high in front | 1
232 | RT233-2.1.7d | Size: 41 cm wide; 50 cm long; 7 cm high in front | 1
233 | RT233-2.1.7e | Size: 38 cm wide; 50 cm long; 7 cm high in front | 1
234 | RT233-2.1.7f | Size: 36 cm wide; 50 cm long; 7 cm high in front | 1
235 | RT233-2.1.7g | Size: 30 cm wide; 50 cm long; 7 cm high in front | 1
| RT233-2.1.8 | Contoured Base, Wedged, thin foam positioner, extended length | 7
236 | RT233-2.1.8a | Size: 51 cm wide; 50 cm long; 4 cm high at rear, 9 cm in front | 1
237 | RT233-2.1.8b | Size: 46 cm wide; 50 cm long; 4 cm high at rear, 9 cm in front | 1
238 | RT233-2.1.8c | Size: 43 cm wide; 50 cm long; 4 cm high at rear, 9 cm in front | 1
239 | RT233-2.1.8d | Size: 41 cm wide; 50 cm long; 4 cm high at rear, 9 cm in front | 1
240 | RT233-2.1.8e | Size: 38 cm wide; 50 cm long; 4 cm high at rear, 9 cm in front | 1
241 | RT233-2.1.8f | Size: 36 cm wide; 50 cm long; 4 cm high at rear, 9 cm in front | 1
242 | RT233-2.1.8g | Size: 30 cm wide; 50 cm long; 4 cm high at rear, 8.5 cm in front | 1
| RT233-2.1.9 | Non-contoured Base, straight base, basic foam, thick pressure care cushion with no abduction block | 13
243 | RT233-2.1.9a | Size: 20 cm wide, 25 cm long, 9cm high | 1
244 | RT233-2.1.9b | Size: 22.5 cm wide, 25 cm long, 9cm high | 1
245 | RT233-2.1.9c | Size: 25 cm wide, 35 cm long, 9cm high | 1
246 | RT233-2.1.9d | Size: 27.5 cm wide, 40 cm long, 9cm high | 1
247 | RT233-2.1.9e | Size: 30 cm wide, 35 cm long, 9cm high | 1
248 | RT233-2.1.9f | Size: 30 cm wide, 40 cm long, 9cm high | 1
249 | RT233-2.1.9g | Size: 36 cm wide, 40 cm long, 9cm high | 1
250 | RT233-2.1.9h | Size: 36 cm wide , 50 cm long, 9cm high | 1
251 | RT233-2.1.9i | Size: 38cm wide, 50cm long, 9cm high | 1
252 | RT233-2.1.9j | Size: 40 cm wide, 50 cm long, 9cm high | 1
253 | RT233-2.1.9k | Size: 43 cm wide, 50 cm long, 9cm high | 1
254 | RT233-2.1.9l | Size: 45 cm wide, 50 cm long, 9cm high | 1
255 | RT233-2.1.9m | Size: 51 cm wide, 50 cm long, 9cm high | 1
| RT233-2.1.10 | Non-contoured Base, straight base, thick pressure care cushion with abduction block | 13
256 | RT233-2.1.10a | Size: 20 cm wide, 25 cm long, 9cm high with extra-small abduction block | 1
257 | RT233-2.1.10b | Size: 22.5 cm wide, 25 cm long, 9cm high with extra-small abduction block | 1
258 | RT233-2.1.10c | Size: 25 cm wide, 35 cm long, 9cm high with extra-small abduction block | 1
259 | RT233-2.1.10d | Size: 27.5 cm wide, 40 cm long, 9cm high with extra-small abduction block | 1
260 | RT233-2.1.10e | Size: 30 cm wide, 35 cm long, 9cm high, with small abduction block | 1
261 | RT233-2.1.10f | Size: 30 cm wide, 40 cm long, 9cm high, with small abduction block | 1
262 | RT233-2.1.10g | Size: 35 cm wide, 40 cm long, 9cm high, with small abduction block | 1
263 | RT233-2.1.10h | Size: 36 cm wide, 50 cm long, 9cm high, with small abduction block | 1
264 | RT233-2.1.10i | Size: 38cm wide, 50cm long, 9cm high, with small abduction block | 1
265 | RT233-2.1.10j | Size: 41 cm wide, 50 cm long, 9cm high with large abduction block | 1
266 | RT233-2.1.10k | Size: 43 cm wide, 50 cm long, 9cm high with large abduction block | 1
267 | RT233-2.1.10l | Size: 46 cm wide, 50 cm long, 9cm high with large abduction block | 1
268 | RT233-2.1.10m | Size: 51 cm wide, 50 cm long, 9cm high with large abduction block | 1
| RT233-2.1.11 | Non-contoured Base, straight base, thin foam positioner with no abduction block | 13
269 | RT233-2.1.11a | Size: 20 cm wide, 25 cm long, 7cm high | 1
270 | RT233-2.1.11b | Size: 22.5 cm wide, 25 cm long, 7cm high | 1
271 | RT233-2.1.11c | Size: 25 cm wide, 35 cm long, 7cm high | 1
272 | RT233-2.1.11d | Size: 27.5 cm wide, 40 cm long, 7cm high | 1
273 | RT233-2.1.11e | Size: 30 cm wide, 35 cm long, 7cm high | 1
274 | RT233-2.1.11f | Size: 30 cm wide, 40 cm long, 7cm high | 1
275 | RT233-2.1.11g | Size: 36 cm wide, 40 cm long, 7cm high | 1
276 | RT233-2.1.11h | Size: 36 cm wide, 50 cm long, 7cm high | 1
277 | RT233-2.1.11i | Size: 38cm wide, 50cm long, 7 cm high | 1
278 | RT233-2.1.11j | Size: 40 cm wide, 50 cm long, 7cm high | 1
279 | RT233-2.1.11k | Size: 43 cm wide, 50 cm long, 7cm high | 1
280 | RT233-2.1.11l | Size: 45 cm wide, 50 cm long, 7cm high | 1
281 | RT233-2.1.11m | Size: 51 cm wide, 50 cm long, 7cm high | 1
| RT233-2.1.12 | Non-contoured Base, straight base, thin foam positioner with abduction block | 12
282 | RT233-2.1.12a | Size: 20 cm wide, 25 cm long, 7cm high with extra-small abduction block | 1
283 | RT233-2.1.12b | Size: 22.5 cm wide, 25 cm long, 7cm high with extra-small abduction block | 1
284 | RT233-2.1.12c | Size: 25 cm wide, 35 cm long, 7cm high with extra-small abduction block | 1
285 | RT233-2.1.12d | Size: 27.5 cm wide, 40 cm long, 7cm high with extra-small abduction block | 1
286 | RT233-2.1.12e | Size: 30 cm wide, 35 cm long, 7cm high with small abduction block | 1
287 | RT233-2.1.12f | Size: 30 cm wide, 40 cm long, 7cm high with small abduction block | 1
288 | RT233-2.1.12g | Size: 35 cm wide, 40 cm long, 7cm high with small abduction block | 1
289 | RT233-2.1.12h | Size: 36 cm wide, 50 cm long, 7cm high with small abduction block | 1
290 | RT233-2.1.12i | Size: 41 cm wide, 50 cm long, 7cm high with large abduction block | 1
291 | RT233-2.1.12j | Size: 43 cm wide, 50 cm long, 7cm high with large abduction block | 1
292 | RT233-2.1.12k | Size: 46 cm wide, 50 cm long, 7cm high with large abduction block | 1
293 | RT233-2.1.12l | Size: 51 cm wide, 50 cm long, 7cm high with large abduction block | 1
| RT233-2.2 | FLUID FILLED CUSHION | 10
| RT233-2.2.1 | Fluid filled cushion with contoured base and rea fluid pad
294 | RT233-2.2.1a | Size: 51 cm wide, 45cm long, no thicker than 8 cm | 1
295 | RT233-2.2.1b | Size: 46 cm wide, 43 cm long, no thicker than 8 cm | 1
296 | RT233-2.2.1c | Size: 46 cm wide, 45 cm long, no thicker than 8 cm | 1
297 | RT233-2.2.1d | Size: 43 cm ) wide, 45 cm long, no thicker than 8 cm | 1
298 | RT233-2.2.1e | Size: 43 cm wide, 43 cm long, no thicker than 8 cm | 1
299 | RT233-2.2.1f | Size: 41 cm wide, 45 cm long, no thicker than 8 cm | 1
300 | RT233-2.2.1g | Size: 41 cm wide, 43 cm long, no thicker than 8 cm | 1
301 | RT233-2.2.1h | Size: 38 cm wide, 43 cm long, no thicker than 8 cm | 1
302 | RT233-2.2.1i | Size: 36 cm wide, 45 cm long, no thicker than 8 cm | 1
303 | RT233-2.2.1j | Size: 36 cm wide, 41 cm long, no thicker than 8 cm | 1
| RT233-2.3 | AIR FILLED CUSHION | 17
| RT233-2.3.1 | Air filled cushion with flexible base | 8
304 | RT233-2.3.1a | Size: 51 cm wide, 46 cm long, no higher than 10 cm | 1
305 | RT233-2.3.1b | Size: 51 cm wide, 41 cm long, no higher than 10 cm | 1
306 | RT233-2.3.1c | Size: 46 cm wide, 46 cm long, no higher than 10 cm | 1
307 | RT233-2.3.1d | Size: 46 cm wide, 41 cm long, no higher than 10 cm | 1
308 | RT233-2.3.1e | Size: 41 cm wide, 46 cm long, no higher than 10 cm | 1
309 | RT233-2.3.1f | Size: 41 cm wide, 41 cm long, no higher than 10 cm | 1
310 | RT233-2.3.1g | Size: 38 cm wide, 46 cm long, no higher than 10 cm | 1
311 | RT233-2.3.1h | Size: 36 cm wide, 41 cm long, no higher than 10 cm | 1
| RT233-2.3.2 | Air filled cushion with individual removable air filled pockets or cells | 9
312 | RT233-2.3.2a | Size: 51 cm wide, 51 cm long, no higher than 10 cm | 1
313 | RT233-2.3.2b | Size: 51 cm wide, 46 cm long, no higher than 10 cm | 1
314 | RT233-2.3.2c | Size: 46 cm wide, 51 cm long, no higher than 10 cm | 1
315 | RT233-2.3.2d | Size: 46 cm wide, 46 cm long, no higher than 10 cm | 1
316 | RT233-2.3.2e | Size: 46 cm wide, 41 cm long, no higher than 10 cm | 1
317 | RT233-2.3.2f | Size: 43 cm wide, 43 cm long, no higher than 10 cm | 1
318 | RT233-2.3.2g | Size: 41 cm wide, 46 cm long, no higher than 10 cm | 1
319 | RT233-2.3.2h | Size: 41 cm wide, 41 cm long, no higher than 10 cm | 1
320 | RT233-2.3.2i | Size: 36 cm wide, 41 cm long, no higher than 10 cm | 1
| RT233-2.4 | COMPRESSED FOAM WEDGE | 2
| RT233-2.4.1 | Compressed foam wedge
321 | RT233-2.4.1a | Size: 45 cm wide; 45 cm long; 1 cm high at rear, 3 cm high in front | 1
322 | RT233-2.4.1b | Size: 45 cm wide; 45 cm long;1 cm high at rear, 5 cm high in front | 1
| RT233-2.5 | COMPRESSED FOAM SHEET | 8
| RT233-2.5.1 | Compressed foam sheet
323 | RT233-2.5.1a | Size: 122 cm x 180 cm x 5 cm | 1
324 | RT233-2.5.1b | Size: 122 cm x 180 cm x 3 cm | 1
325 | RT233-2.5.1c | Size: 122 cm x 180 cm x 2 cm | 1
326 | RT233-2.5.1d | Size: 120 cm x 180 cm x 1.2 cm | 1
327 | RT233-2.5.1e | Size: 137 cm x 94 cm x 5cm | 1
328 | RT233-2.5.1f | Size: 137 cm x 94 cm x 3cm | 1
329 | RT233-2.5.1g | Size: 137 cm x 94 cm x 2 cm | 1
330 | RT233-2.5.1h | Size: 137 cmx 94 cm x 1.2 cm | 1
| RT233-2.6 | HIGH DENSITY FOAM SHEET | 10
| RT233-2.6.1 | High density foam sheet
331 | RT233-2.6.1a | Size: 188 cm x 183 cm x 1.2 cm | 1
332 | RT233-2.6.1b | Size: 188 cm x 183 cm x 3 cm | 1
333 | RT233-2.6.1c | Size: 188 cm x 183 cm x 5 cm | 1
334 | RT233-2.6.1d | Size: 188 cm x 183 cm x 10 cm | 1
335 | RT233-2.6.1e | Size: 188 cm x 183 cm x 15 cm | 1
336 | RT233-2.6.1f | Size: 190 cm x 115 cm x 1.2 cm | 1
337 | RT233-2.6.1g | Size: 190 cm x 115 cm x 3 cm | 1
338 | RT233-2.6.1h | Size: 190 cm x 115 cm x 5 cm | 1
339 | RT233-2.6.1i | Size: 190 cm x 115 cm x 10 cm | 1
340 | RT233-2.6.1j | Size: 190 cm x 115 cm x 15 cm | 1
3. WHEELCHAIR CUSHION COVERS
Item No. | Item Code | Description | Quantity | Are you bidding for this item?: Yes or No | Minimum order Quantities (where applicable) | Import Percentage (where applicable) | Local Percentage (where applicable) | Country of Manufacture (Local or Imported) | Foreign Currency Name | Brand Name | Model Number | Delivery Period (Days) | Unit Of Measure (UOM) | Equipment Expected Lifespan (where applicable) | Unit Price | Comments
| RT233-3.1 | WHEELCHAIR CUSHION COVERS FOR CONTOURED BASED, BASIC FOAM | 80
| RT233-3.1.1 | Nylon wheelchair cushion cover for thick pressure care cushion with contoured base | 11
341 | RT233-3.1.1a | Size: 61 cm wide; 43 cm long; 9 cm high front | 1
342 | RT233-3.1.1b | Size: 56 cm wide; 43 cm long; 9 cm high front | 1
343 | RT233-3.1.1c | Size: 51 cm wide; 43 cm long; 9 cm high front | 1
344 | RT233-3.1.1d | Size: 43 cm wide; 43 cm long; 9 cm high front | 1
345 | RT233-3.1.1e | Size: 46 cm wide; 43 cm long; 9 cm high front | 1
346 | RT233-3.1.1f | Size: 41 cm wide; 43 cm long; 9 cm high front | 1
347 | RT233-3.1.1g | Size: 38 cm wide; 43 cm long; 9 cm high front | 1
348 | RT233-3.1.1h | Size: 36 cm wide; 43 cm long; 9 cm high front | 1
349 | RT233-3.1.1i | Size: 30 cm wide; 43 cm long; 9 cm high front | 1
350 | RT233-3.1.1j | Size: 30 cm wide; 30 cm long; 6.5 cm high front | 1
351 | RT233-3.1.1k | Size: 25 cm wide; 30 cm long; 6 cm high front | 1
| RT233-3.1.2 | Waterproof wheelchair cushion cover for thick pressure care cushion with contoured base | 11
352 | RT233-3.1.2a | Size: 61 cm wide; 43 cm long; 9 cm high front | 1
353 | RT233-3.1.2b | Size: 56 cm wide; 43 cm long; 9 cm high front | 1
354 | RT233-3.1.2c | Size: 51 cm wide; 43 cm long; 9 cm high front | 1
355 | RT233-3.1.2d | Size: 43 cm wide; 43 cm long; 9 cm high front | 1
356 | RT233-3.1.2e | Size: 46 cm wide; 43 cm long; 9 cm high front | 1
357 | RT233-3.1.2f | Size: 41 cm wide; 43 cm long; 9 cm high front | 1
358 | RT233-3.1.2g | Size: 38 cm wide; 43 cm long; 9 cm high front | 1
359 | RT233-3.1.2h | Size: 36 cm wide; 43 cm long; 9 cm high front | 1
360 | RT233-3.1.2i | Size: 30 cm wide; 43 cm long; 9 cm high front | 1
361 | RT233-3.1.2j | Size: 30 cm wide; 30 cm long; 6.5 cm high front | 1
362 | RT233-3.1.2k | Size: 25 cm wide; 30 cm long; 6 cm high front | 1
| RT233-3.1.3 | Nylon wheelchair cushion cover to fit wedged, thick pressure care cushion with contoured base | 11
363 | RT233-3.1.3a | Size: 61 cm wide; 43 cm long; 6 cm high at rear, 11 cm in front | 1
364 | RT233-3.1.3b | Size: 56 cm wide; 43 cm long; 6 cm high at rear, 11 cm in front | 1
365 | RT233-3.1.3c | Size: 51 cm wide; 43 cm long; 6 cm high at rear, 11 cm in front | 1
366 | RT233-3.1.3d | Size: 43 cm wide; 43 cm long; 6 cm high at rear, 11 cm in front | 1
367 | RT233-3.1.3e | Size: 46 cm wide; 43 cm long; 6 cm high at rear, 11 cm in front | 1
368 | RT233-3.1.3f | Size: 41 cm wide; 43 cm long; 6 cm high at rear, 11 cm in front | 1
369 | RT233-3.1.3g | Size: 38 cm wide; 43 cm long; 6 cm high at rear, 11 cm in front | 1
370 | RT233-3.1.3h | Size: 36 cm wide; 43 cm long; 6 cm high at rear, 11 cm in front | 1
371 | RT233-3.1.3i | Size: 30 cm wide; 43 cm long; 6 cm high at rear, 11 cm in front | 1
372 | RT233-3.1.3j | Size: 30 cm wide; 30 cm long; 4 cm high at rear, 8.5 cm in front | 1
373 | RT233-3.1.3k | Size: 25 cm wide; 30 cm long; 4 cm high at rear, 7.5 cm in front | 1
| RT233-3.1.4 | Waterproof wheelchair cushion cover to fit wedged, thick pressure care cushion with contoured base | 11
374 | RT233-3.1.4a | Size: 61 cm wide; 43 cm long; 6 cm high at rear, 11 cm in front | 1
375 | RT233-3.1.4b | Size: 56 cm wide; 43 cm long; 6 cm high at rear, 11 cm in front | 1
376 | RT233-3.1.4c | Size: 51 cm wide; 43 cm long; 6 cm high at rear, 11 cm in front | 1
377 | RT233-3.1.4d | Size: 43 cm wide; 43 cm long; 6 cm high at rear, 11 cm in front | 1
378 | RT233-3.1.4e | Size: 46 cm wide; 43 cm long; 6 cm high at rear, 11 cm in front | 1
379 | RT233-3.1.4f | Size: 41 cm wide; 43 cm long; 6 cm high at rear, 11 cm in front | 1
380 | RT233-3.1.4g | Size: 38 cm wide; 43 cm long; 6 cm high at rear, 11 cm in front | 1
381 | RT233-3.1.4h | Size: 36 cm wide; 43 cm long; 6 cm high at rear, 11 cm in front | 1
382 | RT233-3.1.4i | Size: 30 cm wide; 43 cm long; 6 cm high at rear, 11 cm in front | 1
383 | RT233-3.1.4j | Size: 30 cm wide; 30 cm long; 4 cm high at rear, 8.5 cm in front | 1
384 | RT233-3.1.4j | Size: 25 cm wide; 30 cm long; 4 cm high at rear, 7.5 cm in front | 1
| RT233-3.1.5 | Nylon wheelchair cushion cover to fit thin foam positioner cushion with contoured base | 9
385 | RT233-3.1.5a | Size: 51 cm wide; 43 cm long; 7 cm high | 1
386 | RT233-3.1.5b | Size: 46 cm wide; 43 cm long; 7 cm high | 1
387 | RT233-3.1.5c | Size: 43 cm wide; 43 cm long; 7 cm high | 1
388 | RT233-3.1.5d | Size: 41 cm wide; 43 cm long; 7 cm high | 1
389 | RT233-3.1.5e | Size: 38 cm wide; 43 cm long; 7 cm high | 1
390 | RT233-3.1.5f | Size: 36 cm wide; 43 cm long; 7 cm high | 1
391 | RT233-3.1.5g | Size: 30 cm wide; 43 cm long; 6.5 cm high | 1
392 | RT233-3.1.5h | Size: 30 cm wide; 33 cm long; 6.5 cm high | 1
393 | RT233-3.1.5i | Size: 25 cm wide; 28 cm long; 6 cm high | 1
| RT233-3.1.6 | Waterproof wheelchair cushion cover to fit thin foam positioner cushion with contoured base | 9
394 | RT233-3.1.6a | Size: 51 cm wide; 43 cm long; 7 cm high | 1
395 | RT233-3.1.6b | Size: 46 cm wide; 43 cm long; 7 cm high | 1
396 | RT233-3.1.6c | Size: 43 cm wide; 43 cm long; 7 cm high | 1
397 | RT233-3.1.6d | Size: 41 cm wide; 43 cm long; 7 cm high | 1
398 | RT233-3.1.6e | Size: 38 cm wide; 43 cm long; 7 cm high | 1
399 | RT233-3.1.6f | Size: 36 cm wide; 43 cm long; 7 cm high | 1
400 | RT233-3.1.6g | Size: 30 cm wide; 43 cm long; 6.5 cm high | 1
401 | RT233-3.1.6h | Size: 30 cm wide; 33 cm long; 6.5 cm high | 1
402 | RT233-3.1.6i | Size: 25 cm wide; 28 cm long; 6 cm high | 1
| RT233-3.1.7 | Nylon wheelchair cushion cover to fit wedged, thin foam positioner with contoured base | 9
403 | RT233-3.1.7a | Size: 51 cm wide; 43 cm long; 4 cm high at rear, 9 cm in front | 1
404 | RT233-3.1.7a | Size: 46 cm wide; 43 cm long; 4 cm high at rear, 9 cm in front | 1
405 | RT233-3.1.7a | Size: 43 cm wide; 43 cm long; 4 cm high at rear, 9 cm in front | 1
406 | RT233-3.1.7a | Size: 41 cm wide; 43 cm long; 4 cm high at rear, 9 cm in front | 1
407 | RT233-3.1.7a | Size: 38 cm wide; 43 cm long; 4 cm high at rear, 9 cm in front | 1
408 | RT233-3.1.7a | Size: 36 cm wide; 43 cm long; 4 cm high at rear, 9 cm in front | 1
409 | RT233-3.1.7a | Size: 30 cm wide; 43 cm long; 4 cm high at rear, 9 cm in front | 1
410 | RT233-3.1.7a | Size: 30 cm wide; 33 cm long; 4 cm high at rear, 8.5 cm in front | 1
411 | RT233-3.1.7a | Size: 25 cm wide; 28 cm long; 4 cm high at rear, 7.5cm in front | 1
| RT233-3.1.8 | Waterproof cushion cover to fit wedged, thin foam positioner with contoured base | 9
412 | RT233-3.1.8a | Size: 51 cm wide; 43 cm long; 4 cm high at rear, 9 cm in front | 1
413 | RT233-3.1.8b | Size: 46 cm wide; 43 cm long; 4 cm high at rear, 9 cm in front | 1
414 | RT233-3.1.8c | Size: 43 cm wide; 43 cm long; 4 cm high at rear, 9 cm in front | 1
415 | RT233-3.1.8d | Size: 41 cm wide; 43 cm long; 4 cm high at rear, 9 cm in front | 1
416 | RT233-3.1.8e | Size: 38 cm wide; 43 cm long; 4 cm high at rear, 9 cm in front | 1
417 | RT233-3.1.8f | Size: 36 cm wide; 43 cm long; 4 cm high at rear, 9 cm in front | 1
418 | RT233-3.1.8g | Size: 30 cm wide; 43 cm long; 4 cm high at rear, 9 cm in front | 1
419 | RT233-3.1.8h | Size: 30 cm wide; 33 cm long; 4 cm high at rear, 8.5 cm in front | 1
420 | RT233-3.1.8i | Size: 25 cm wide; 28 cm long; 4 cm high at rear, 7.5cm in front | 1
| RT233-3.2 | WHEELCHAIR CUSHION COVERS FOR CONTOURED BASED, BASIC FOAM EXTENDED | 64
| RT233-3.2.1 | Nylon wheelchair cushion cover to fit thick pressure care foam cushion with contoured base, extended | 9
421 | RT233-3.2.1a | Size: 61 cm wide; 50 cm long; 9 cm high | 1
422 | RT233-3.2.1b | Size: 56 cm wide; 50 cm long; 9 cm high | 1
423 | RT233-3.2.1c | Size: 51 cm wide; 50 cm long; 9 cm high | 1
424 | RT233-3.2.1d | Size: 46 cm wide; 50 cm long; 9 cm high | 1
425 | RT233-3.2.1e | Size: 43 cm wide; 50 cm long; 9 cm high | 1
426 | RT233-3.2.1f | Size: 41 cm wide; 50 cm long; 9 cm high | 1
427 | RT233-3.2.1g | Size: 38 cm wide; 50 cm long; 9 cm high | 1
428 | RT233-3.2.1h | Size: 36 cm wide; 50 cm long; 9 cm high | 1
429 | RT233-3.2.1i | Size: 30 cm wide; 50 cm long; 6.5 cm high | 1
| RT233-3.2.2 | Waterproof wheelchair cushion cover to fit thick pressure care foam cushion with contoured base, extended | 9
430 | RT233-3.2.2a | Size: 61 cm wide; 50 cm long; 9 cm high | 1
431 | RT233-3.2.2b | Size: 56 cm wide; 50 cm long; 9 cm high | 1
432 | RT233-3.2.2c | Size: 51 cm wide; 50 cm long; 9 cm high | 1
433 | RT233-3.2.2d | Size: 46 cm wide; 50 cm long; 9 cm high | 1
434 | RT233-3.2.2e | Size: 43 cm wide; 50 cm long; 9 cm high | 1
435 | RT233-3.2.2f | Size: 41 cm wide; 50 cm long; 9 cm high | 1
436 | RT233-3.2.2g | Size: 38 cm wide; 50 cm long; 9 cm high | 1
437 | RT233-3.2.2h | Size: 36 cm wide; 50 cm long; 9 cm high | 1
438 | RT233-3.2.2i | Size: 30 cm wide; 50 cm long; 6.5 cm high | 1
| RT233-3.2.3 | Nylon wheelchair cushion cover to fit wedged, thick pressure care foam cushion with contoured base, extended length | 9
439 | RT233-3.2.3a | Size: 61 cm wide; 50 cm long; 6 cm high at rear, 11 cm in front | 1
440 | RT233-3.2.3b | Size: 56 cm wide; 50 cm long; 6 cm high at rear, 11 cm in front | 1
441 | RT233-3.2.3c | Size: 51 cm wide; 50 cm long; 6 cm high at rear, 11 cm in front | 1
442 | RT233-3.2.3d | Size: 46 cm wide; 50 cm long; 6 cm high at rear, 11 cm in front | 1
443 | RT233-3.2.3e | Size: 43 cm wide; 50 cm long; 6 cm high at rear, 11 cm in front | 1
444 | RT233-3.2.3f | Size: 41 cm wide; 50 cm long; 6 cm high at rear, 11 cm in front | 1
445 | RT233-3.2.3g | Size: 38 cm wide; 50 cm long; 6 cm high at rear, 11cm in front | 1
446 | RT233-3.2.3h | Size: 36 cm wide; 50 cm long; 6 cm high at rear, 11 cm in front | 1
447 | RT233-3.2.3i | Size: 30 cm wide; 50 cm long; 4 cm high at rear, 8.5 cm in front | 1
| RT233-3.2.4 | Waterproof wheelchair cushion cover to fit wedged, thick pressure care foam cushion with contoured base, extended length | 9
448 | RT233-3.2.4a | Size: 61 cm wide; 50 cm long; 6 cm high at rear, 11 cm in front | 1
449 | RT233-3.2.4b | Size: 56 cm wide; 50 cm long; 6 cm high at rear, 11 cm in front | 1
450 | RT233-3.2.4c | Size: 51 cm wide; 50 cm long; 6 cm high at rear, 11 cm in front | 1
451 | RT233-3.2.4d | Size: 46 cm wide; 50 cm long; 6 cm high at rear, 11 cm in front | 1
452 | RT233-3.2.4e | Size: 43 cm wide; 50 cm long; 6 cm high at rear, 11 cm in front | 1
453 | RT233-3.2.4f | Size: 41 cm wide; 50 cm long; 6 cm high at rear, 11 cm in front | 1
454 | RT233-3.2.4g | Size: 38 cm wide; 50 cm long; 6 cm high at rear, 11cm in front | 1
455 | RT233-3.2.4h | Size: 36 cm wide; 50 cm long; 6 cm high at rear, 11 cm in front | 1
456 | RT233-3.2.4i | Size: 30 cm wide; 50 cm long; 64 cm high at rear, 8.5 cm in front | 1
| RT233-3.2.5 | Nylon wheelchair cushion cover to fit thin foam, basic positioner cushion with contoured base, extended length | 7
457 | RT233-3.2.5a | Size: 51 cm wide; 50 cm long; 7 cm high in front | 1
458 | RT233-3.2.5b | Size: 46 cm wide; 50 cm long; 7 cm high in front | 1
459 | RT233-3.2.5c | Size: 43 cm wide; 50 cm long; 7 cm high in front | 1
460 | RT233-3.2.5d | Size: 41 cm wide; 50 cm long; 7 cm high in front | 1
461 | RT233-3.2.5e | Size: 38 cm wide; 50 cm long; 7 cm high in front | 1
462 | RT233-3.2.5f | Size: 36 cm wide; 50 cm long; 7 cm high in front | 1
463 | RT233-3.2.5g | Size: 30 cm wide; 50 cm long; 6.5 cm high in front | 1
| RT233-3.2.6 | Waterproof wheelchair cushion cover to fit thin foam, basic positioner cushion with contoured base, extended length | 7
464 | RT233-3.2.6a | Size: 51 cm wide; 50 cm long; 7 cm high in front | 1
465 | RT233-3.2.6b | Size: 46 cm wide; 50 cm long; 7 cm high in front | 1
466 | RT233-3.2.6c | Size: 43 cm wide; 50 cm long; 7 cm high in front | 1
467 | RT233-3.2.6d | Size: 41 cm wide; 50 cm long; 7 cm high in front | 1
468 | RT233-3.2.6e | Size: 38 cm wide; 50 cm long; 7 cm high in front | 1
469 | RT233-3.2.6f | Size: 36 cm wide; 50 cm long; 7 cm high in front | 1
470 | RT233-3.2.6g | Size: 30 cm wide; 50 cm long; 6.5 cm high in front | 1
| RT233-3.2.7 | Nylon wheelchair cushion cover to fit wedged, thin forma positioner with contoured base, extended length | 7
471 | RT233-3.2.7a | Size: 51 cm wide; 50 cm long; 4 cm high at rear, 9 cm in front | 1
472 | RT233-3.2.7b | Size: 46 cm wide; 50 cm long; 4 cm high at rear, 9 cm in front | 1
473 | RT233-3.2.7c | Size: 43 cm wide; 50 cm long; 4 cm high at rear, 9 cm in front | 1
474 | RT233-3.2.7d | Size: 41 cm wide; 50 cm long; 4 cm high at rear, 9 cm in front | 1
475 | RT233-3.2.7e | Size: 38 cm wide; 50 cm long; 4 cm high at rear, 9 cm in front | 1
476 | RT233-3.2.7f | Size: 36 cm wide; 50 cm long; 4 cm high at rear, 9 cm in front | 1
477 | RT233-3.2.7g | Size: 30 cm wide; 50 cm long; 4 cm high at rear, 8.5 cm in front | 1
| RT233-3.2.8 | Waterproof wheelchair cushion cover to fit wedged, thin form positioner with contoured base, extended length | 7
478 | RT233-3.2.8a | Size: 51 cm wide; 50 cm long; 4 cm high at rear, 9 cm in front | 1
479 | RT233-3.2.8b | Size: 46 cm wide; 50 cm long; 4 cm high at rear, 9 cm in front | 1
480 | RT233-3.2.8c | Size: 43 cm wide; 50 cm long; 4 cm high at rear, 9 cm in front | 1
481 | RT233-3.2.8d | Size: 41 cm wide; 50 cm long; 4 cm high at rear, 9 cm in front | 1
482 | RT233-3.2.8e | Size: 38 cm wide; 50 cm long; 4 cm high at rear, 9 cm in front | 1
483 | RT233-3.2.8f | Size: 36 cm wide; 50 cm long; 4 cm high at rear, 9 cm in front | 1
484 | RT233-3.2.8g | Size: 30 cm wide; 50 cm long; 4 cm high at rear, 8.5 cm in front | 1
| RT233-3.3 | WHEELCHAIR CUSHION COVER FOR STRAIGHT BASE, BASIC FOAM | 104
| RT233-3.3.1 | Nylon wheelchair cushion cover to fit straight base, thick pressure care cushion with no adbuction block | 13
485 | RT233-3.3.1a | Size: 20 cm wide, 25 cm long, 9cm high | 1
486 | RT233-3.3.1b | Size: 22.5 cm wide, 25 cm long, 9cm high | 1
487 | RT233-3.3.1c | Size: 25 cm wide, 35 cm long, 9cm high | 1
488 | RT233-3.3.1d | Size: 27.5 cm wide, 40 cm long, 9cm high | 1
489 | RT233-3.3.1e | Size: 30 cm wide, 35 cm long, 9cm high | 1
490 | RT233-3.3.1f | Size: 30 cm wide, 40 cm long, , 9cm high | 1
491 | RT233-3.3.1g | Size: 36 cm wide, 40 cm long, , 9cm high | 1
492 | RT233-3.3.1h | Size: 36 cm wide , 50 cm long, 9cm high | 1
493 | RT233-3.3.1i | Size: 38cm wide, 50cm long, 9cm high | 1
494 | RT233-3.3.1j | Size: 40 cm wide, 50 cm long, 9cm high | 1
495 | RT233-3.3.1k | Size: 43 cm wide, 50 cm long, 9cm high | 1
496 | RT233-3.3.1l | Size: 45 cm wide, 50 cm long, 9cm high | 1
497 | RT233-3.3.1m | Size: 51 cm wide, 50 cm long, 9cm high | 1
| RT233-3.3.2 | Nylon wheelchair cushion cover to fit straight base, thick pressure care cushion with adbuction block | 13
498 | RT233-3.3.2a | Size: 20 cm wide, 25 cm long, 9cm high with extra-small abduction block | 1
499 | RT233-3.3.2b | Size: 22.5 cm wide, 25 cm long, 9cm high with extra-small abduction block | 1
500 | RT233-3.3.2c | Size: 25 cm wide, 35 cm long, 9cm high with extra-small abduction block | 1
501 | RT233-3.3.2d | Size: 27.5 cm wide, 40 cm long, 9cm high with extra-small abduction block | 1
502 | RT233-3.3.2e | Size: 30 cm wide, 35 cm long, 9cm high, with small abduction block | 1
503 | RT233-3.3.2f | Size: 30 cm wide, 40 cm long, 9cm high, with small abduction block | 1
504 | RT233-3.3.2g | Size: 35 cm wide, 40 cm long, 9cm high, with small abduction block | 1
505 | RT233-3.3.2h | Size: 36 cm wide, 50 cm long, 9cm high, with small abduction block | 1
506 | RT233-3.3.2i | Size: 38cm wide, 50cm long, 9cm high, with small abduction block | 1
507 | RT233-3.3.2j | Size: 41 cm wide, 50 cm long, 9cm high with large abduction block | 1
508 | RT233-3.3.2k | Size: 43 cm wide, 50 cm long, 9cm high with large abduction block | 1
509 | RT233-3.3.2l | Size: 46 cm wide, 50 cm long, 9cm high with large abduction block | 1
510 | RT233-3.3.2m | Size: 51 cm wide, 50 cm long, 9cm high with large abduction block | 1
| RT233-3.3.3 | Waterproof cushion cover to fit straight base, thick pressure care cushion with no adbuction block | 13
511 | RT233-3.3.3a | Size: 20 cm wide, 25 cm long, 9cm high | 1
512 | RT233-3.3.3b | Size: 22.5 cm wide, 25 cm long, 9cm hig | 1
513 | RT233-3.3.3c | Size: 25 cm wide, 35 cm long, 9cm high | 1
514 | RT233-3.3.3d | Size: 27.5 cm wide, 40 cm long, 9cm high | 1
515 | RT233-3.3.3e | Size: 30 cm wide, 35 cm long, 9cm high | 1
516 | RT233-3.3.3f | Size: 30 cm wide, 40 cm long, 9cm high | 1
517 | RT233-3.3.3g | Size: 36 cm wide, 40 cm long, 9cm high | 1
518 | RT233-3.3.3h | Size: 36 cm wide , 50 cm long, 9cm high | 1
519 | RT233-3.3.3i | Size: 38cm wide, 50cm long, 9cm high | 1
520 | RT233-3.3.3j | Size: 40 cm wide, 50 cm long, 9cm high | 1
521 | RT233-3.3.3k | Size: 43 cm wide, 50 cm long, 9cm high | 1
522 | RT233-3.3.3l | Size: 45 cm wide, 50 cm long, 9cm high | 1
523 | RT233-3.3.3m | Size: 51 cm wide, 50 cm long, 9cm high | 1
| RT233-3.3.4 | Waterproof wheelchair cushion cover to fit straight base, thick pressure care cushion with adbuction block | 13
| RT233-3.3.4a | Size: 20 cm wide, 25 cm long, 9cm high with extra-small abduction block | 1
524 | RT233-3.3.4b | Size: 22.5 cm wide, 25 cm long, 9cm high with extra-small abduction block | 1
525 | RT233-3.3.4c | Size: 25 cm wide, 35 cm long, 9cm high with extra-small abduction block | 1
526 | RT233-3.3.4d | Size: 27.5 cm wide, 40 cm long, 9cm high with extra-small abduction block | 1
527 | RT233-3.3.4e | Size: 30 cm wide, 35 cm long, 9cm high, with small abduction block | 1
528 | RT233-3.3.4f | Size: 30 cm wide, 40 cm long, 9cm high, with small abduction block | 1
529 | RT233-3.3.4g | Size: 35 cm wide, 40 cm long, 9cm high, with small abduction block | 1
530 | RT233-3.3.4h | Size: 36 cm wide, 50 cm long, 9cm high, with small abduction block | 1
531 | RT233-3.3.4i | Size: 38 cm wide, 50cm long, 9cm high, with small abduction block | 1
532 | RT233-3.3.4j | Size: 41 cm wide, 50 cm long, 9cm high with large abduction block | 1
533 | RT233-3.3.4k | Size: 43 cm wide, 50 cm long, 9cm high with large abduction block | 1
534 | RT233-3.3.4l | Size: 46 cm wide, 50 cm long, 9cm high with large abduction block | 1
535 | RT233-3.3.4m | Size: 51 cm wide, 50 cm long, 9cm high with large abduction block | 1
| RT233-3.3.5 | Nylon wheelchair cushion cover for straight base, thin foam positioner with no adbuction block | 13
536 | RT233-3.3.5a | Size: 20 cm wide, 25 cm long, 7cm high | 1
537 | RT233-3.3.5b | Size: 22.5 cm wide, 25 cm long, 7cm high | 1
538 | RT233-3.3.5c | Size: 25 cm wide, 35 cm long, 7cm high | 1
539 | RT233-3.3.5d | Size: 27.5 cm wide, 40 cm long, 7cm high | 1
540 | RT233-3.3.5e | Size: 30 cm wide, 35 cm long, 7cm high | 1
541 | RT233-3.3.5f | Size: 30 cm wide, 40 cm long, 7cm high | 1
542 | RT233-3.3.5g | Size: 35 cm wide, 40 cm long, 7cm high | 1
543 | RT233-3.3.5h | Size: 36 cm wide, 50 cm long, 7cm high | 1
544 | RT233-3.3.5i | Size: 38 cm wide, 50 cm long, 7 cm high | 1
545 | RT233-3.3.5j | Size: 41 cm wide, 50 cm long, 7cm high | 1
546 | RT233-3.3.5k | Size: 43 cm wide, 50 cm long, 7cm high | 1
547 | RT233-3.3.5l | Size: 46 cm wide, 50 cm long, 7cm high | 1
548 | RT233-3.3.5m | Size: 51 cm wide, 50 cm long, 7cm high | 1
| RT233-3.3.6 | Nylon wheelchair cushion cover to fit straight base, thin foam positioner with adbuction block | 13
549 | RT233-3.3.6a | Size: 20 cm wide, 25 cm long, 7cm high with extra-small abduction block | 1
550 | RT233-3.3.6b | Size: 22.5 cm wide, 25 cm long, 7cm high with extra-small abduction block | 1
551 | RT233-3.3.6c | Size: 25 cm wide, 35 cm long, 7cm high with extra-small abduction block | 1
552 | RT233-3.3.6d | Size: 27.5 cm wide, 40 cm long, 7cm high with extra-small abduction block | 1
553 | RT233-3.3.6e | Size: 30 cm wide, 35 cm long, 7cm high with small abduction block | 1
554 | RT233-3.3.6f | Size: 30 cm wide, 40 cm long, 7cm high with small abduction block | 1
555 | RT233-3.3.6g | Size: 35 cm wide, 40 cm long, 7cm high with small abduction block | 1
556 | RT233-3.3.6h | Size: 36 cm wide, 50 cm long, 7cm high with small abduction block | 1
557 | RT233-3.3.6i | Size: 38 cm wide, 50 cm long, 7 cm high with small abduction block | 1
558 | RT233-3.3.6j | Size: 41 cm wide, 50 cm long, 7cm high with large abduction block | 1
559 | RT233-3.3.6k | Size: 43 cm wide, 50 cm long, 7cm high with large abduction block | 1
560 | RT233-3.3.6l | Size: 46 cm wide, 50 cm long, 7cm high with large abduction block | 1
561 | RT233-3.3.6m | Size: 51 cm wide, 50 cm long, 7cm high with large abduction block | 1
| RT233-3.3.7 | Waterproof wheelchair cushion cover to fit straight base, thin foam positioner with no adbuction block | 13
562 | RT233-3.3.7a | Size: 20 cm wide, 25 cm long, 7cm high | 1
563 | RT233-3.3.7b | Size: 22.5 cm wide, 25 cm long, 7cm high | 1
564 | RT233-3.3.7c | Size: 25 cm wide, 35 cm long, 7cm high | 1
565 | RT233-3.3.7d | Size: 27.5 cm wide, 40 cm long, 7cm high | 1
566 | RT233-3.3.7e | Size: 30 cm wide, 35 cm long, 7cm high | 1
567 | RT233-3.3.7f | Size: 30 cm wide, 40 cm long, 7cm high | 1
568 | RT233-3.3.7g | Size: 35 cm wide, 40 cm long, 7cm high | 1
569 | RT233-3.3.7h | Size: 36 cm wide, 50 cm long, 7cm high | 1
570 | RT233-3.3.7i | Size: 38 cm wide, 50 cm long, 7 cm high | 1
571 | RT233-3.3.7j | Size: 41 cm wide, 50 cm long, 7cm high | 1
572 | RT233-3.3.7k | Size: 43 cm wide, 50 cm long, 7cm high | 1
573 | RT233-3.3.7l | Size: 46 cm wide, 50 cm long, 7cm high | 1
574 | RT233-3.3.7m | Size: 51 cm wide, 50 cm long, 7cm high | 1
| RT233-3.3.8 | Waterproof wheelchair cushion cover to fit straight base, thin foam positioner with adbuction block | 13
575 | RT233-3.3.8a | Size: 20 cm wide, 25 cm long, 7cm high with extra-small abduction block | 1
576 | RT233-3.3.8b | Size: 22.5 cm wide, 25 cm long, 7cm high with extra-small abduction block | 1
577 | RT233-3.3.8c | Size: 25 cm wide, 35 cm long, 7cm high with extra-small abduction block | 1
578 | RT233-3.3.8d | Size: 27.5 cm wide, 40 cm long, 7cm high with extra-small abduction block | 1
579 | RT233-3.3.8e | Size: 30 cm wide, 35 cm long, 7cm high with small abduction block | 1
580 | RT233-3.3.8f | Size: 30 cm wide, 40 cm long, 7cm high with small abduction block | 1
581 | RT233-3.3.8g | Size: 35 cm wide, 40 cm long, 7cm high with small abduction block | 1
582 | RT233-3.3.8h | Size: 36 cm wide, 50 cm long, 7cm high with small abduction block | 1
583 | RT233-3.3.8i | Size: 38 cm wide, 50 cm long, 7 cm high with small abduction block | 1
584 | RT233-3.3.8j | Size: 41 cm wide, 50 cm long, 7cm high with large abduction block | 1
585 | RT233-3.3.8k | Size: 43 cm wide, 50 cm long, 7cm high with large abduction block | 1
586 | RT233-3.3.8l | Size: 46 cm wide, 50 cm long, 7cm high with large abduction block | 1
587 | RT233-3.3.8m | Size: 51 cm wide, 50 cm long, 7cm high with large abduction block | 1
| RT233-3.4 | WHEELCHAIR CUSHION COVER FABRIC | 2
588 | RT233-3.4.1 | Wheelchair cushion cover fabric, basic nylon stretchable | 1
589 | RT233-3.4.2 | Wheelchair cushion cover fabric, waterproof, stretchable | 1
| RT233-3.5 | FLUID FILLED WHEELCHAIR CUSHION COVER | 2
590 | RT233-3.5.1 | Fluid filled wheelchair cushion with removable base and fluid pad | 1
3.6 | RT233-3.6 | AIR FILLED WHEELCHAIR CUSHION COVER | 1
591 | RT233-3.6.1 | Air filled wheelchair cushion with removable cover | 1
4. WHEELCHAIR CUSHIONS WITH COVERS
| RT233-4.1 | BASIC FOAM WHEELCHAIR CUSHIONS WITH NYLON COVERS | 121
| RT233-4.1.1 | Contoured base thick pressure care wheelchair cushion with nylon cover | 10
592 | RT233-4.1.1a | Size: 61 cm wide, 43 cm long, 9 cm high in front | 1
593 | RT233-4.1.1b | Size: 56 cm wide, 43 cm long, 9 cm high in front | 1
594 | RT233-4.1.1c | Size: 51 cm wide, 43 cm long, 9 cm high in front | 1
595 | RT233-4.1.1d | Size: 46 cm wide, 43 cm long, 9 cm high in front | 1
596 | RT233-4.1.1e | Size: 41 cm wide, 43 cm long, 9 cm high in front | 1
597 | RT233-4.1.1f | Size: 38 cm wide, 43 cm long, 9 cm high in front | 1
598 | RT233-4.1.1g | Size: 36 cm wide, 43 cm long, 9 cm high in front | 1
599 | RT233-4.1.1h | Size: 30 cm wide, 43 cm long, 9 cm high in front | 1
600 | RT233-4.1.1i | Size: 30 cm wide, 30 cm long, 6.5cm high in front | 1
601 | RT233-4.1.1j | Size: 25 cm wide, 30 cm long, 6 cm high in front | 1
| RT233-4.1.2 | Contoured base wedged thick pressure care wheelchair cushion with nylon cover | 10
602 | RT233-4.1.2a | Size: 61 cm wide; 43 cm long; 6 cm high at rear, 11 cm high in front | 1
603 | RT233-4.1.2b | Size: 56 cm wide; 43 cm long; 6 cm high at rear, 11 cm high in front | 1
604 | RT233-4.1.2c | Size: 51 cm wide; 43 cm long; 6 cm high at rear, 11 cm high in front | 1
605 | RT233-4.1.2d | Size: 46 cm wide; 43 cm long; 6 cm high at rear, 11 cm high in front | 1
606 | RT233-4.1.2e | Size: 41 cm wide; 43 cm long; 6 cm high at rear, 11 cm high in front | 1
607 | RT233-4.1.2f | Size: 38 cm wide; 43 cm long; 6 cm high at rear, 11 cm high in front | 1
608 | RT233-4.1.2g | Size: 36 cm wide; 43 cm long; 6 cm high at rear, 11 cm high in front | 1
609 | RT233-4.1.2h | Size: 30 cm wide; 43 cm long; 6 cm high at rear, 11 cm high in front | 1
610 | RT233-4.1.2i | Size: 30 cm wide; 30 cm long; 4 cm high at rear, 8.5 cm high in front | 1
611 | RT233-4.1.2j | Size: 25 cm wide; 30 cm long; 4 cm high at rear, 7.5 cm high in front | 1
| RT233-4.1.3 | Contoured base thin foam basic positioner wheelchair cushion with nylon cover | 9
612 | RT233-4.1.3a | Size: 51 cm wide; 43 cm long; 7 cm high | 1
613 | RT233-4.1.3b | Size: 46 cm wide; 43 cm long; 7 cm high | 1
614 | RT233-4.1.3c | Size: 43 cm wide; 43 cm long; 7 cm high | 1
615 | RT233-4.1.3d | Size: 41 cm wide; 43 cm long; 7 cm high | 1
616 | RT233-4.1.3e | Size: 38 cm wide; 43 cm long; 7 cm high | 1
617 | RT233-4.1.3f | Size: 36 cm wide; 43 cm long; 7 cm high | 1
618 | RT233-4.1.3g | Size: 30 cm wide; 43 cm long; 6.5 cm high | 1
619 | RT233-4.1.3h | Size: 30 cm wide; 33 cm long; 6.5 cm high | 1
620 | RT233-4.1.3i | Size: 25 cm wide; 28 cm long; 6 cm high | 1
| RT233-4.1.4 | Contoured base wedged thin foam positioner with nylon wheelchair cover | 9
621 | RT233-4.1.4a | Size: 51 cm wide; 43 cm long; 4 cm high at rear, 9 cm high in front | 1
622 | RT233-4.1.4b | Size: 46 cm wide; 43 cm long; 4 cm high at rear, 9 cm high in front | 1
623 | RT233-4.1.4c | Size: 43 cm wide; 43 cm long; 4 cm high at rear, 9 cm high in front | 1
624 | RT233-4.1.4d | Size: 41 cm wide; 43 cm long; 4 cm high at rear, 9 cm high in front | 1
625 | RT233-4.1.4e | Size: 38 cm wide; 43 cm long; 4 cm high at rear, 9 cm high in front | 1
626 | RT233-4.1.4f | Size: 36 cm wide; 43 cm long; 4 cm high at rear, 9 cm high in front | 1
627 | RT233-4.1.4g | Size: 30 cm wide; 43 cm long; 4 cm high at rear, 9 cm high in front | 1
628 | RT233-4.1.4h | Size: 30 cm wide; 33 cm long; 4 cm high at rear, 8 cm high in front | 1
629 | RT233-4.1.4i | Size: 25 cm wide; 28 cm long; 4 cm high at rear, 7.5cm high in front | 1
| RT233-4.1.5 | Contoured base thick pressure care wheelchair cushion, extended length with nylon cover | 9
630 | RT233-4.1.5a | Size: 61 cm wide; 50 cm long; 9 cm high | 1
631 | RT233-4.1.5b | Size: 56 cm wide; 50 cm long; 9 cm high | 1
632 | RT233-4.1.5c | Size: 51 cm wide; 50 cm long; 9 cm high | 1
633 | RT233-4.1.5d | Size: 46 cm wide; 50 cm long; 9 cm high | 1
634 | RT233-4.1.5e | Size: 43 cm wide; 50 cm long; 9 cm high | 1
635 | RT233-4.1.5f | Size: 41 cm wide; 50 cm long; 9 cm high | 1
636 | RT233-4.1.5g | Size: 38 cm wide; 50 cm long; 9 cm high | 1
637 | RT233-4.1.5h | Size: 36 cm wide; 50 cm long; 8.5 cm high | 1
638 | RT233-4.1.5i | Size: 30 cm wide; 50 cm long; 8.5 cm high | 1
| RT233-4.1.6 | Contoured base wedged thick pressure care wheelchair cushion extended length with nylon cover | 9
639 | RT233-4.1.6a | Size: 61 cm wide; 50 cm long; 6 cm high at rear, 11 cm high in front | 1
640 | RT233-4.1.6b | Size: 56 cm wide; 50 cm long; 6 cm high at rear, 11 cm high in front | 1
641 | RT233-4.1.6c | Size: 51 cm wide; 50 cm long; 6 cm high at rear, 11 cm high in front | 1
642 | RT233-4.1.6d | Size: 46 cm wide; 50 cm long; 6 cm high at rear, 11 cm high in front | 1
643 | RT233-4.1.6e | Size: 43 cm wide; 50 cm long; 6 cm high at rear, 11 cm high in front | 1
644 | RT233-4.1.6f | Size: 41 cm wide; 50 cm long; 6 cm high at rear, 11 cm high in front | 1
645 | RT233-4.1.6g | Size: 38 cm wide; 50 cm long; 6 cm high at rear, 11cm high in front | 1
646 | RT233-4.1.6h | Size: 36 cm wide; 50 cm long; 6 cm high at rear, 11 cm high in front | 1
647 | RT233-4.1.6i | Size: 30 cm wide; 50 cm long; 6 cm high at rear, 10.5 cm high in front | 1
| RT233-4.1.7 | Contoured base thin foam basic positioner wheelchair cushion extended length with nylon cover | 7
648 | RT233-4.1.7a | Size: 51 cm wide; 50 cm long; 7 cm high in front | 1
649 | RT233-4.1.7b | Size: 46 cm wide; 50 cm long; 7 cm high in front | 1
650 | RT233-4.1.7c | Size: 43 cm wide; 50 cm long; 7 cm high in front | 1
651 | RT233-4.1.7d | Size: 41 cm wide; 50 cm long; 7 cm high in front | 1
652 | RT233-4.1.7e | Size: 38 cm wide; 50 cm long; 7 cm high in front | 1
653 | RT233-4.1.7f | Size: 36 cm wide; 50 cm long; 7 cm high in front | 1
654 | RT233-4.1.7g | Size: 30 cm wide; 50 cm long; 7 cm high in front | 1
| RT233-4.1.8 | Contoured base wedged thin foam positioner, extended length with nylon wheelchair cover | 7
655 | RT233-4.1.8a | Size: 51 cm wide; 50 cm long; 4 cm high at rear, 9 cm in front | 1
656 | RT233-4.1.8b | Size: 46 cm wide; 50 cm long; 4 cm high at rear, 9 cm in front | 1
657 | RT233-4.1.8c | Size: 43 cm wide; 50 cm long; 4 cm high at rear, 9 cm in front | 1
658 | RT233-4.1.8d | Size: 41 cm wide; 50 cm long; 4 cm high at rear, 9 cm in front | 1
659 | RT233-4.1.8e | Size: 38 cm wide; 50 cm long; 4 cm high at rear, 9 cm in front | 1
660 | RT233-4.1.8f | Size: 36 cm wide; 50 cm long; 4 cm high at rear, 9 cm in front | 1
661 | RT233-4.1.8g | Size: 30 cm wide; 50 cm long; 4 cm high at rear, 8.5 cm in front | 1
| RT233-4.1.9 | Non-contoured base, straight base, basic foam, thick pressure care wheelchair cushion with no abduction block with nylon cover | 13
662 | RT233-4.1.9a | Size: 20 cm wide, 25 cm long, 9cm high | 1
663 | RT233-4.1.9b | Size: 22.5 cm wide, 25 cm long, 9cm high | 1
664 | RT233-4.1.9c | Size: 25 cm wide, 35 cm long, 9cm high | 1
665 | RT233-4.1.9d | Size: 27.5 cm wide, 40 cm long, 9cm high | 1
666 | RT233-4.1.9e | Size: 30 cm wide, 35 cm long, 9cm high | 1
667 | RT233-4.1.9f | Size: 30 cm wide, 40 cm long, 9cm high | 1
668 | RT233-4.1.9g | Size: 36 cm wide, 40 cm long, 9cm high | 1
669 | RT233-4.1.9h | Size: 36 cm wide , 50 cm long, 9cm high | 1
670 | RT233-4.1.9i | Size: 38cm wide, 50cm long, 9cm high | 1
671 | RT233-4.1.9j | Size: 40 cm wide, 50 cm long, 9cm high | 1
672 | RT233-4.1.9k | Size: 43 cm wide, 50 cm long, 9cm high | 1
673 | RT233-4.1.9l | Size: 45 cm wide, 50 cm long, 9cm high | 1
674 | RT233-4.1.9m | Size: 51 cm wide, 50 cm long, 9cm high | 1
| RT233-4.1.10 | Non-contoured base, straight base, thick pressure care wheelchair cushion with abduction block and nylon cover | 13
675 | RT233-4.1.10a | Size: 20 cm wide, 25 cm long, 9cm high with extra-small abduction block | 1
676 | RT233-4.1.10b | Size: 22.5 cm wide, 25 cm long, 9cm high with extra-small abduction block | 1
677 | RT233-4.1.10c | Size: 25 cm wide, 35 cm long, 9cm high with extra-small abduction block | 1
678 | RT233-4.1.10d | Size: 27.5 cm wide, 40 cm long, 9cm high with extra-small abduction block | 1
679 | RT233-4.1.10e | Size: 30 cm wide, 35 cm long, 9cm high, with small abduction block | 1
680 | RT233-4.1.10f | Size: 30 cm wide, 40 cm long, 9cm high, with small abduction block | 1
681 | RT233-4.1.10g | Size: 35 cm wide, 40 cm long, 9cm high, with small abduction block | 1
682 | RT233-4.1.10h | Size: 36 cm wide, 50 cm long, 9cm high, with small abduction block | 1
683 | RT233-4.1.10i | Size: 38cm wide, 50cm long, 9cm high, with small abduction block | 1
684 | RT233-4.1.10j | Size: 41 cm wide, 50 cm long, 9cm high with large abduction block | 1
685 | RT233-4.1.10k | Size: 43 cm wide, 50 cm long, 9cm high with large abduction block | 1
686 | RT233-4.1.10l | Size: 46 cm wide, 50 cm long, 9cm high with large abduction block | 1
687 | RT233-4.1.10m | Size: 51 cm wide, 50 cm long, 9cm high with large abduction block | 1
| RT233-4.1.11 | Non-contoured base, straight base, positioner with no abduction block with nylon wheelchair cover | 13
688 | RT233-4.1.11a | Size: 20 cm wide, 25 cm long, 7cm high | 1
689 | RT233-4.1.11b | Size: 22.5 cm wide, 25 cm long, 7cm high | 1
690 | RT233-4.1.11c | Size: 25 cm wide, 35 cm long, 7cm high | 1
691 | RT233-4.1.11d | Size: 27.5 cm wide, 40 cm long, 7cm high | 1
692 | RT233-4.1.11e | Size: 30 cm wide, 35 cm long, 7cm high | 1
693 | RT233-4.1.11f | Size: 30 cm wide, 40 cm long, 7cm high | 1
694 | RT233-4.1.11g | Size: 36 cm wide, 40 cm long, 7cm high | 1
695 | RT233-4.1.11h | Size: 36 cm wide, 50 cm long, 7cm high | 1
696 | RT233-4.1.11i | Size: 38cm wide, 50cm long, 7 cm high | 1
697 | RT233-4.1.11j | Size: 40 cm wide, 50 cm long, 7cm high | 1
698 | RT233-4.1.11k | Size: 43 cm wide, 50 cm long, 7cm high | 1
699 | RT233-4.1.11l | Size: 45 cm wide, 50 cm long, 7cm high | 1
700 | RT233-4.1.11m | Size: 51 cm wide, 50 cm long, 7cm high | 1
| RT233-4.1.12 | Non-contoured base, straight base, positioner with abduction block with nylon wheelchair cover | 12
701 | RT233-4.1.12a | Size: 20 cm wide, 25 cm long, 7cm high with extra-small abduction block | 1
702 | RT233-4.1.12b | Size: 22.5 cm wide, 25 cm long, 7cm high with extra-small abduction block | 1
703 | RT233-4.1.12c | Size: 25 cm wide, 35 cm long, 7cm high with extra-small abduction block | 1
704 | RT233-4.1.12d | Size: 27.5 cm wide, 40 cm long, 7cm high with extra-small abduction block | 1
705 | RT233-4.1.12e | Size: 30 cm wide, 35 cm long, 7cm high with small abduction block | 1
706 | RT233-4.1.12f | Size: 30 cm wide, 40 cm long, 7cm high with small abduction block | 1
707 | RT233-4.1.12g | Size: 35 cm wide, 40 cm long, 7cm high with small abduction block | 1
708 | RT233-4.1.12h | Size:36 cm wide, 50 cm long, 7cm high with small abduction block | 1
709 | RT233-4.1.12i | Size: 41 cm wide, 50 cm long, 7cm high with large abduction block | 1
710 | RT233-4.1.12j | Size: 43 cm wide, 50 cm long, 7cm high with large abduction block | 1
711 | RT233-4.1.12k | Size: 46 cm wide, 50 cm long, 7cm high with large abduction block | 1
712 | RT233-4.1.12l | Size: 51 cm wide, 50 cm long, 7cm high with large abduction block | 1
| RT233-4.2 | BASIC FOAM WHEELCHAIR CUSHIONS WITH WATERPROOF COVERS | 121
| RT233-4.2.1 | Contoured base thick pressure care wheelchair cushion with waterproof cover | 10
713 | RT233-4.2.1a | Size: 61 cm wide, 43 cm long, 9 cm high in front | 1
714 | RT233-4.2.1b | Size: 56 cm wide, 43 cm long, 9 cm high in front | 1
715 | RT233-4.2.1c | Size: 51 cm wide, 43 cm long, 9 cm high in front | 1
716 | RT233-4.2.1d | Size: 46 cm wide, 43 cm long, 9 cm high in front | 1
717 | RT233-4.2.1e | Size: 41 cm wide, 43 cm long, 9 cm high in front | 1
718 | RT233-4.2.1f | Size: 38 cm wide, 43 cm long, 9 cm high in front | 1
719 | RT233-4.2.1g | Size: 36 cm wide, 43 cm long, 9 cm high in front | 1
720 | RT233-4.2.1h | Size: 30 cm wide, 43 cm long, 9 cm high in front | 1
721 | RT233-4.2.1i | Size: 30 cm wide, 30 cm long, 6.5cm high in front | 1
722 | RT233-4.2.1j | Size: 25 cm wide, 30 cm long, 6 cm high in front | 1
| RT233-4.2.2 | Contoured base, wedged, thick pressure care wheelchair cushion with waterproof cover | 10
723 | RT233-4.2.2a | Size: 61 cm wide; 43 cm long; 6 cm high at rear, 11 cm high in front | 1
724 | RT233-4.2.2b | Size: 56 cm wide; 43 cm long; 6 cm high at rear, 11 cm high in front | 1
725 | RT233-4.2.2c | Size: 51 cm wide; 43 cm long; 6 cm high at rear, 11 cm high in front | 1
726 | RT233-4.2.2d | Size: 46 cm wide; 43 cm long; 6 cm high at rear, 11 cm high in front | 1
727 | RT233-4.2.2e | Size: 41 cm wide; 43 cm long; 6 cm high at rear, 11 cm high in front | 1
728 | RT233-4.2.2f | Size: 38 cm wide; 43 cm long; 6 cm high at rear, 11 cm high in front | 1
729 | RT233-4.2.2g | Size: 36 cm wide; 43 cm long; 6 cm high at rear, 11 cm high in front | 1
730 | RT233-4.2.2h | Size: 30 cm wide; 43 cm long; 6 cm high at rear, 11 cm high in front | 1
731 | RT233-4.2.2i | Size: 30 cm wide; 30 cm long; 4 cm high at rear, 8.5 cm high in front | 1
732 | RT233-4.2.2j | Size: 25 cm wide; 30 cm long; 4 cm high at rear, 7.5 cm high in front | 1
| RT233-4.2.3 | Contoured base, thin foam, basic positioner wheelchair cushion with waterproof cover | 9
733 | RT233-4.2.3a | Size: 51 cm wide; 43 cm long; 7 cm high | 1
734 | RT233-4.2.3b | Size: 46 cm wide; 43 cm long; 7 cm high | 1
735 | RT233-4.2.3c | Size: 43 cm wide; 43 cm long; 7 cm high | 1
736 | RT233-4.2.3d | Size: 41 cm wide; 43 cm long; 7 cm high | 1
737 | RT233-4.2.3e | Size: 38 cm wide; 43 cm long; 7 cm high | 1
738 | RT233-4.2.3f | Size: 36 cm wide; 43 cm long; 7 cm high | 1
739 | RT233-4.2.3g | Size: 30 cm wide; 43 cm long; 6.5 cm high | 1
740 | RT233-4.2.3h | Size: 30 cm wide; 33 cm long; 6.5 cm high | 1
741 | RT233-4.2.3i | Size: 25 cm wide; 28 cm long; 6 cm high | 1
| RT233-4.2.4 | Contoured base, wedged, thin foam positioner with waterproof wheelchair cover | 9
742 | RT233-4.2.4a | Size: 51 cm wide; 43 cm long; 4 cm high at rear, 9 cm high in front | 1
743 | RT233-4.2.4b | Size: 46 cm wide; 43 cm long; 4 cm high at rear, 9 cm high in front | 1
744 | RT233-4.2.4c | Size: 43 cm wide; 43 cm long; 4 cm high at rear, 9 cm high in front | 1
745 | RT233-4.2.4d | Size: 41 cm wide; 43 cm long; 4 cm high at rear, 9 cm high in front | 1
746 | RT233-4.2.4e | Size: 38 cm wide; 43 cm long; 4 cm high at rear, 9 cm high in front | 1
747 | RT233-4.2.4f | Size: 36 cm wide; 43 cm long; 4 cm high at rear, 9 cm high in front | 1
748 | RT233-4.2.4g | Size: 30 cm wide; 43 cm long; 4 cm high at rear, 9 cm high in front | 1
749 | RT233-4.2.4h | Size: 30 cm wide; 33 cm long; 4 cm high at rear, 8 cm high in front | 1
750 | RT233-4.2.4i | Size: 25 cm wide; 28 cm long; 4 cm high at rear, 7.5cm high in front | 1
| RT233-4.2.5 | Contoured base, thick pressure care wheelchair cushion, extended length with waterproof cover | 9
751 | RT233-4.2.5a | Size: 61 cm wide; 50 cm long; 9 cm high | 1
752 | RT233-4.2.5b | Size: 56 cm wide; 50 cm long; 9 cm high | 1
753 | RT233-4.2.5c | Size: 51 cm wide; 50 cm long; 9 cm high | 1
754 | RT233-4.2.5d | Size: 46 cm wide; 50 cm long; 9 cm high | 1
755 | RT233-4.2.5e | Size: 43 cm wide; 50 cm long; 9 cm high | 1
756 | RT233-4.2.5f | Size: 41 cm wide; 50 cm long; 9 cm high | 1
757 | RT233-4.2.5g | Size: 38 cm wide; 50 cm long; 9 cm high | 1
758 | RT233-4.2.5h | Size: 36 cm wide; 50 cm long; 8.5 cm high | 1
759 | RT233-4.2.5i | Size: 30 cm wide; 50 cm long; 8.5 cm high | 1
| RT233-4.2.6 | Contoured base, wedged, thick pressure care wheelchair cushion, extended length with waterproof cover | 9
760 | RT233-4.2.6a | Size: 61 cm wide; 50 cm long; 6 cm high at rear, 11 cm high in front | 1
761 | RT233-4.2.6b | Size: 56 cm wide; 50 cm long; 6 cm high at rear, 11 cm high in front | 1
762 | RT233-4.2.6c | Size: 51 cm wide; 50 cm long; 6 cm high at rear, 11 cm high in front | 1
763 | RT233-4.2.6d | Size: 46 cm wide; 50 cm long; 6 cm high at rear, 11 cm high in front | 1
764 | RT233-4.2.6e | Size: 43 cm wide; 50 cm long; 6 cm high at rear, 11 cm high in front | 1
765 | RT233-4.2.6f | Size: 41 cm wide; 50 cm long; 6 cm high at rear, 11 cm high in front | 1
766 | RT233-4.2.6g | Size: 38 cm wide; 50 cm long; 6 cm high at rear, 11cm high in front | 1
767 | RT233-4.2.6h | Size: 36 cm wide; 50 cm long; 6 cm high at rear, 11 cm high in front | 1
768 | RT233-4.2.6i | Size: 30 cm wide; 50 cm long; 6 cm high at rear, 10.5 cm high in front | 1
| RT233-4.2.7 | Contoured base, thin foam, basic positioner wheelchair cushion, extended length with waterproof cover | 7
769 | RT233-4.2.7a | Size: 51 cm wide; 50 cm long; 7 cm high in front | 1
770 | RT233-4.2.7b | Size: 46 cm wide; 50 cm long; 7 cm high in front | 1
771 | RT233-4.2.7c | Size: 43 cm wide; 50 cm long; 7 cm high in front | 1
772 | RT233-4.2.7d | Size: 41 cm wide; 50 cm long; 7 cm high in front | 1
773 | RT233-4.2.7e | Size: 38 cm wide; 50 cm long; 7 cm high in front | 1
774 | RT233-4.2.7f | Size: 36 cm wide; 50 cm long; 7 cm high in front | 1
775 | RT233-4.2.7g | Size: 30 cm wide; 50 cm long; 7 cm high in front | 1
| RT233-4.2.8 | Contoured base, wedged, thin foam positioner, extended length with waterproof wheelchair cover | 7
776 | RT233-4.2.8a | Size: 51 cm wide; 50 cm long; 4 cm high at rear, 9 cm in front | 1
777 | RT233-4.2.8b | Size: 46 cm wide; 50 cm long; 4 cm high at rear, 9 cm in front | 1
778 | RT233-4.2.8c | Size: 43 cm wide; 50 cm long; 4 cm high at rear, 9 cm in front | 1
779 | RT233-4.2.8d | Size: 41 cm wide; 50 cm long; 4 cm high at rear, 9 cm in front | 1
780 | RT233-4.2.8e | Size: 38 cm wide; 50 cm long; 4 cm high at rear, 9 cm in front | 1
781 | RT233-4.2.8f | Size: 36 cm wide; 50 cm long; 4 cm high at rear, 9 cm in front | 1
782 | RT233-4.2.8g | Size: 30 cm wide; 50 cm long; 4 cm high at rear, 8.5 cm in front | 1
| RT233-4.2.9 | Non-contoured base, straight base, basic foam, thick pressure care wheelchair cushion with no abduction block with waterproof cover | 13
783 | RT233-4.2.9a | Size: 20 cm wide, 25 cm long, 9cm high | 1
784 | RT233-4.2.9b | Size: 22.5 cm wide, 25 cm long, 9cm high | 1
785 | RT233-4.2.9c | Size: 25 cm wide, 35 cm long, 9cm high | 1
786 | RT233-4.2.9d | Size: 27.5 cm wide, 40 cm long, 9cm high | 1
787 | RT233-4.2.9e | Size: 30 cm wide, 35 cm long, 9cm high | 1
788 | RT233-4.2.9f | Size: 30 cm wide, 40 cm long, 9cm high | 1
789 | RT233-4.2.9g | Size: 36 cm wide, 40 cm long, 9cm high | 1
790 | RT233-4.2.9h | Size: 36 cm wide , 50 cm long, 9cm high | 1
791 | RT233-4.2.9i | Size: 38cm wide, 50cm long, 9cm high | 1
792 | RT233-4.2.9j | Size: 40 cm wide, 50 cm long, 9cm high | 1
793 | RT233-4.2.9k | Size: 43 cm wide, 50 cm long, 9cm high | 1
794 | RT233-4.2.9l | Size: 45 cm wide, 50 cm long, 9cm high | 1
795 | RT233-4.2.9m | Size: 51 cm wide, 50 cm long, 9cm high | 1
| RT233-4.2.10 | Non-contoured base, straight base, thick pressure care wheelchair cushion with abduction block and waterproof cover | 13
796 | RT233-4.2.10a | Size: 20 cm wide, 25 cm long, 9cm high with extra-small abduction block | 1
797 | RT233-4.2.10b | Size: 22.5 cm wide, 25 cm long, 9cm high with extra-small abduction block | 1
798 | RT233-4.2.10c | Size: 25 cm wide, 35 cm long, 9cm high with extra-small abduction block | 1
799 | RT233-4.2.10d | Size: 27.5 cm wide, 40 cm long, 9cm high with extra-small abduction block | 1
800 | RT233-4.2.10e | Size: 30 cm wide, 35 cm long, 9cm high, with small abduction block | 1
801 | RT233-4.2.10f | Size: 30 cm wide, 40 cm long, 9cm high, with small abduction block | 1
802 | RT233-4.2.10g | Size: 35 cm wide, 40 cm long, 9cm high, with small abduction block | 1
803 | RT233-4.2.10h | Size: 36 cm wide, 50 cm long, 9cm high, with small abduction block | 1
804 | RT233-4.2.10i | Size: 38cm wide, 50cm long, 9cm high, with small abduction block | 1
805 | RT233-4.2.10j | Size: 41 cm wide, 50 cm long, 9cm high with large abduction block | 1
806 | RT233-4.2.10k | Size: 43 cm wide, 50 cm long, 9cm high with large abduction block | 1
807 | RT233-4.2.10l | Size: 46 cm wide, 50 cm long, 9cm high with large abduction block | 1
808 | RT233-4.2.10m | Size: 51 cm wide, 50 cm long, 9cm high with large abduction block | 1
| RT233-4.2.11 | Non-contoured base, straight base, positioner with no abduction block with waterproof wheelchair cover | 13
809 | RT233-4.2.11a | Size: 20 cm wide, 25 cm long, 7cm high | 1
810 | RT233-4.2.11b | Size: 22.5 cm wide, 25 cm long, 7cm high | 1
811 | RT233-4.2.11c | Size: 25 cm wide, 35 cm long, 7cm high | 1
812 | RT233-4.2.11d | Size: 27.5 cm wide, 40 cm long, 7cm high | 1
813 | RT233-4.2.11e | Size: 30 cm wide, 35 cm long, 7cm high | 1
814 | RT233-4.2.11f | Size: 30 cm wide, 40 cm long, 7cm high | 1
815 | RT233-4.2.11g | Size: 36 cm wide, 40 cm long, 7cm high | 1
816 | RT233-4.2.11h | Size: 36 cm wide, 50 cm long, 7cm high | 1
817 | RT233-4.2.11i | Size: 38cm wide, 50cm long, 7 cm high | 1
818 | RT233-4.2.11j | Size: 40 cm wide, 50 cm long, 7cm high | 1
819 | RT233-4.2.11k | Size: 43 cm wide, 50 cm long, 7cm high | 1
820 | RT233-4.2.11l | Size: 45 cm wide, 50 cm long, 7cm high | 1
821 | RT233-4.2.11m | Size: 51 cm wide, 50 cm long, 7cm high | 1
| RT233-4.2.12 | Non-contoured base, straight base, positioner with abduction block with waterproof cover | 12
822 | RT233-4.2.12a | Size: 20 cm wide, 25 cm long, 7cm high with extra-small abduction block | 1
823 | RT233-4.2.12b | Size: 22.5 cm wide, 25 cm long, 7cm high with extra-small abduction block | 1
824 | RT233-4.2.12c | Size: 25 cm wide, 35 cm long, 7cm high with extra-small abduction block | 1
825 | RT233-4.2.12d | Size: 27.5 cm wide, 40 cm long, 7cm high with extra-small abduction block | 1
826 | RT233-4.2.12e | Size: 30 cm wide, 35 cm long, 7cm high with small abduction block | 1
827 | RT233-4.2.12f | Size: 30 cm wide, 40 cm long, 7cm high with small abduction block | 1
828 | RT233-4.2.12g | Size: 35 cm wide, 40 cm long, 7cm high with small abduction block | 1
829 | RT233-4.2.12h | Size: 36 cm wide, 50 cm long, 7cm high with small abduction block | 1
830 | RT233-4.2.12i | Size: 41 cm wide, 50 cm long, 7cm high with large abduction block | 1
831 | RT233-4.2.12j | Size: 43 cm wide, 50 cm long, 7cm high with large abduction block | 1
832 | RT233-4.2.12k | Size: 46 cm wide, 50 cm long, 7cm high with large abduction block | 1
833 | RT233-4.2.12l | Size: 51 cm wide, 50 cm long, 7cm high with large abduction block | 1
5. BACKREST SYSTEMS
Item No. | Item Code | Description | Quantity |  | Minimum order Quantities (where applicable) | Import Percentage (where applicable) | Local Percentage (where applicable) | Country of Manufacture (Local or Imported) | Foreign Currency Name | Brand Name | Model Number | Delivery Period (Days) | Unit Of Measure (UOM) | Equipment Expected Lifespan (where applicable) | Unit Price | Comments
| RT233-5.1 | TENSION ADJUSTABLE BACK | 108
| RT233-5.1.1 | Tension adjustable back, standard | 10
834 | RT233-5.1.1a | Size (width) 61 cm (24”), Standard back height 47.5 cm | 1
835 | RT233-5.1.1b | Size (width) 56 cm (22”), Standard back height 47.5 cm | 1
836 | RT233-5.1.1c | Size (width) 51 cm (20"), Standard back height 45 cm | 1
837 | RT233-5.1.1d | Size (width) 46 cm (18") (Cover range 43 cm (17”) to 46 cm (18")), Standard back height 45 cm | 1
838 | RT233-5.1.1e | Size (width) 41 cm (16") (cover range 38 cm (15") to 41 cm (16")), Standard back height 45 cm | 1
839 | RT233-5.1.1f | Size (width) 41 cm (16") (cover range 38 cm (15") to 41 cm (16")), Standard back height 45 cm | 1
840 | RT233-5.1.1g | Size (width) 36 cm (14"), Standard back height 45 cm | 1
841 | RT233-5.1.1h | Size (width) 30 cm (12") wide with 40 cm seat frame length, Standard back height 45 cm | 1
842 | RT233-5.1.1i | Size (width) 30 cm (12") wide with 30 cm seat frame length, Standard back height 35 cm | 1
843 | RT233-5.1.1j | Size (width) 25 cm (10") wide with 30 cm seat frame length, Standard back height 32-33 cm | 1
| RT233-5.1.2 | Tension adjustable back, short | 10
844 | RT233-5.1.2a | Size (width) 61 cm (24”), Tall back height 51 cm | 1
845 | RT233-5.1.2b | Size (width) 56 cm (22”), Tall back height 51 cm | 1
846 | RT233-5.1.2c | Size (width) 51 cm (20"), Tall back height 55 cm | 1
847 | RT233-5.1.2d | Size (width) 46 cm (18") (Cover range 43 cm (17”) to 46 cm (18")), Tall back height 55 cm | 1
848 | RT233-5.1.2e | Size (width) 41 cm (16") (cover range 38 cm (15") to 41 cm (16")), Tall back height 55 cm | 1
849 | RT233-5.1.2f | Size (width) 41 cm (16") (cover range 38 cm (15") to 41 cm (16")), Tall back height 55 cm | 1
850 | RT233-5.1.2g | Size (width) 36 cm (14"), Tall back height 55 cm | 1
851 | RT233-5.1.2h | Size (width) 30 cm (12") wide with 40 cm seat frame length, Tall back height 55 cm | 1
852 | RT233-5.1.2i | Size (width) 30 cm (12") wide with 30 cm seat frame length, Tall back height 45 cm | 1
853 | RT233-5.1.2j | Size (width) 25 cm (10") wide with 30 cm seat frame length, Not Applicable | 1
| RT233-5.1.3 | Tension adjustable back, tall | 10
854 | RT233-5.1.3a | Size (width) 61 cm (24”) | 1
855 | RT233-5.1.3b | Size (width) 56 cm (22”) | 1
856 | RT233-5.1.3c | Size (width) 51 cm (20"), Short back height 35 cm | 1
857 | RT233-5.1.3d | Size (width) 46 cm (18") (Cover range 43 cm (17”) to 46 cm (18")), Short back height 35 cm | 1
858 | RT233-5.1.3e | Size (width) 41 cm (16") (cover range 38 cm (15") to 41 cm (16")), Short back height 35 cm | 1
859 | RT233-5.1.3f | Size (width) 41 cm (16") (cover range 38 cm (15") to 41 cm (16")), Short back height 35 cm | 1
860 | RT233-5.1.3g | Size (width) 36 cm (14"), Short back height 35 cm | 1
861 | RT233-5.1.3h | Size (width) 30 cm (12") wide with 40 cm seat frame length, Short back height 35 cm | 1
862 | RT233-5.1.3i | Size (width) 30 cm (12") wide with 30 cm seat frame length, Not Applicable | 1
863 | RT233-5.1.3j | Size (width) 25 cm (10") wide with 30 cm seat frame length, Not Applicable | 1
5.2 | RT233-5.2 | TENSION ADJUSTABLE BACKREST WITH RIGID HEIGHT ADJUSTABLE FULL-LENGTH LATERALS | 18
| RT233-5.2.1 | Tension adjustable back with rigid, height adjustable full length laterals, short | 6
864 | RT233-5.2.1a | Size (width) 51 cm (20") (cover range 48 cm (19”) to 51 cm (20")), Short back height 30 cm | 1
865 | RT233-5.2.1b | Size (width) 46 cm (18") (cover range 43 cm (17”) to 46 cm (18")), Short back height 30 cm | 1
866 | RT233-5.2.1c | Size (width) 41 cm (16") (cover range 38 cm (15") to 41 cm (16")), Short back height 30 cm | 1
867 | RT233-5.2.1d | Size (width) 36 cm (14"), Short back height 30 cm | 1
868 | RT233-5.2.1e | Size (width) 30 cm (12”) 40cm seat length, Short back height 30 cm | 1
869 | RT233-5.2.1f | Size (width) 30 cm (12”) 30 cm seat length, Nil | 1
| RT233-5.2.2 | Tension adjustable back with rigid, height adjustable full length laterals, standard | 6
870 | RT233-5.2.2a | Size (width) 51 cm (20") (cover range 48 cm (19”) to 51 cm (20")), Standard back height 42 cm | 1
871 | RT233-5.2.2b | Size (width) 46 cm (18") (cover range 43 cm (17”) to 46 cm (18")), Standard back height 42 cm | 1
872 | RT233-5.2.2c | Size (width) 41 cm (16") (cover range 38 cm (15") to 41 cm (16")), Standard back height 42 cm | 1
873 | RT233-5.2.2d | Size (width) 36 cm (14"), Standard back height 42 cm | 1
874 | RT233-5.2.2e | Size (width) 30 cm (12”) 40cm seat length, Standard back height 42 cm | 1
875 | RT233-5.2.2f | Size (width) 30 cm (12”) 30 cm seat length, Standard back height 30 cm | 1
| RT233-5.2.3 | Tension adjustable back with rigid, height adjustable full length laterals, tall | 6
876 | RT233-5.2.3a | Size (width) 51 cm (20") (cover range 48 cm (19”) to 51 cm (20")), Tall back height 50 cm | 1
877 | RT233-5.2.3b | Size (width) 46 cm (18") (cover range 43 cm (17”) to 46 cm (18")), Tall back height 50 cm | 1
878 | RT233-5.2.3c | Size (width) 41 cm (16") (cover range 38 cm (15") to 41 cm (16")), Tall back height 50 cm | 1
879 | RT233-5.2.3d | Size (width) 36 cm (14"), Tall back height 50 cm | 1
880 | RT233-5.2.3e | Size (width) 30 cm (12”) 40cm seat length, Tall back height 50 cm | 1
881 | RT233-5.2.3f | Size (width) 30 cm (12”) 30 cm seat length, Tall back height 42 cm | 1
| RT233-5.3 | RIGID ADJUSTABLE BACK WITH FULL LENGTH LATERAL SUPPORT | 23
| RT233-5.3.1 | Rigid adjustable back with full length integrated trunk and pelvis side support, mid thoracic height | 7
882 | RT233-5.3.1a | See Table 9: Rigid Adjustable Back with full length laterals, Mid-Thoracic Height Measurements of Annexure 3B | 1
883 | RT233-5.3.1b | See Table 9: Rigid Adjustable Back with full length laterals, Mid-Thoracic Height Measurements of Annexure 3B | 1
884 | RT233-5.3.1c | See Table 9: Rigid Adjustable Back with full length laterals, Mid-Thoracic Height Measurements of Annexure 3B | 1
885 | RT233-5.3.1d | See Table 9: Rigid Adjustable Back with full length laterals, Mid-Thoracic Height Measurements of Annexure 3B | 1
886 | RT233-5.3.1e | See Table 9: Rigid Adjustable Back with full length laterals, Mid-Thoracic Height Measurements of Annexure 3B | 1
887 | RT233-5.3.1f | See Table 9: Rigid Adjustable Back with full length laterals, Mid-Thoracic Height Measurements of Annexure 3B | 1
888 | RT233-5.3.1g | See Table 9: Rigid Adjustable Back with full length laterals, Mid-Thoracic Height Measurements of Annexure 3B | 1
| RT233-5.3.2 | Rigid adjustable back with full length integrated trunk and pelvis side support, upper thoracic height | 8
889 | RT233-5.3.2a | See Table 10: Rigid Adjustable Back with full length laterals, Upper Thoracic Height Measurements on Annexure 3B | 1
890 | RT233-5.3.2b | See Table 10: Rigid Adjustable Back with full length laterals, Upper Thoracic Height Measurements on Annexure 3B | 1
891 | RT233-5.3.2c | See Table 10: Rigid Adjustable Back with full length laterals, Upper Thoracic Height Measurements on Annexure 3B | 1
892 | RT233-5.3.2d | See Table 10: Rigid Adjustable Back with full length laterals, Upper Thoracic Height Measurements on Annexure 3B | 1
893 | RT233-5.3.2e | See Table 10: Rigid Adjustable Back with full length laterals, Upper Thoracic Height Measurements on Annexure 3B | 1
894 | RT233-5.3.2f | See Table 10: Rigid Adjustable Back with full length laterals, Upper Thoracic Height Measurements on Annexure 3B | 1
895 | RT233-5.3.2g | See Table 10: Rigid Adjustable Back with full length laterals, Upper Thoracic Height Measurements on Annexure 3B | 1
896 | RT233-5.3.2h | See Table 10: Rigid Adjustable Back with full length laterals, Upper Thoracic Height Measurements on Annexure 3B | 1
| RT233-5.3.3 | Rigid adjustable back with full length integrated trunk and pelvis side support, shoulder thoracic height | 8
897 | RT233-5.3.3a | See Table 11: Rigid Adjustable Back with full length laterals, Shoulder Height Measurements on Annexure 3B | 1
898 | RT233-5.3.3b | See Table 11: Rigid Adjustable Back with full length laterals, Shoulder Height Measurements on Annexure 3B | 1
899 | RT233-5.3.3c | See Table 11: Rigid Adjustable Back with full length laterals, Shoulder Height Measurements on Annexure 3B | 1
900 | RT233-5.3.3d | See Table 11: Rigid Adjustable Back with full length laterals, Shoulder Height Measurements on Annexure 3B | 1
901 | RT233-5.3.3e | See Table 11: Rigid Adjustable Back with full length laterals, Shoulder Height Measurements on Annexure 3B | 1
902 | RT233-5.3.3f | See Table 11: Rigid Adjustable Back with full length laterals, Shoulder Height Measurements on Annexure 3B | 1
903 | RT233-5.3.3g | See Table 11: Rigid Adjustable Back with full length laterals, Shoulder Height Measurements on Annexure 3B | 1
904 | RT233-5.3.3h | See Table 11: Rigid Adjustable Back with full length laterals, Shoulder Height Measurements on Annexure 3B | 1
| RT233-5.4 | RIGID ADJUSTABLE BACK | 12
| RT233-5.4.1 | Rigid adjustable back, mid thoracis height | 4
905 | RT233-5.4.1a | Size of back (measured from outside to outside of trunk side supports): 12” | 1
906 | RT233-5.4.1b | Size of back (measured from outside to outside of trunk side supports): 14” | 1
907 | RT233-5.4.1c | Size of back (measured from outside to outside of trunk side supports): 16” | 1
908 | RT233-5.4.1d | Size of back (measured from outside to outside of trunk side supports): 18” | 1
| RT233-5.4.2 | Rigid adjustable back, upper thoracis height | 4
909 | RT233-5.4.2a | Size of back (measured from outside to outside of trunk side supports): 12” | 1
910 | RT233-5.4.2b | Size of back (measured from outside to outside of trunk side supports): 14” | 1
911 | RT233-5.4.2c | Size of back (measured from outside to outside of trunk side supports): 16” | 1
912 | RT233-5.4.2d | Size of back (measured from outside to outside of trunk side supports): 18” | 1
| RT233-5.4.3 | Rigid adjustable back, shoulder height | 4
913 | RT233-5.4.3a | Size of back (measured from outside to outside of trunk side supports): 12” | 1
914 | RT233-5.4.3b | Size of back (measured from outside to outside of trunk side supports): 14” | 1
915 | RT233-5.4.3c | Size of back (measured from outside to outside of trunk side supports): 16” | 1
916 | RT233-5.4.3d | Size of back (measured from outside to outside of trunk side supports): 18” | 1
| RT233-5.5 | ADJUSTABLE MODULAR BACK | 10
| RT233-5.5.1 | Rigid modular back, standard height | 5
917 | RT233-5.5.1a | See Table 13: Adjustable Modular Back Size Ranges on Annexure 3B | 1
918 | RT233-5.5.1b | See Table 13: Adjustable Modular Back Size Ranges on Annexure 3B | 1
919 | RT233-5.5.1c | See Table 13: Adjustable Modular Back Size Ranges on Annexure 3B | 1
920 | RT233-5.5.1d | See Table 13: Adjustable Modular Back Size Ranges on Annexure 3B | 1
921 | RT233-5.5.1e | See Table 13: Adjustable Modular Back Size Ranges on Annexure 3B | 1
| RT233-5.5.2 | Rigid modular back, tall height | 5
922 | RT233-5.5.2a | See Table 13: Adjustable Modular Back Size Ranges on Annexure 3B | 1
923 | RT233-5.5.2b | See Table 13: Adjustable Modular Back Size Ranges on Annexure 3B | 1
924 | RT233-5.5.2c | See Table 13: Adjustable Modular Back Size Ranges on Annexure 3B | 1
925 | RT233-5.5.2d | See Table 13: Adjustable Modular Back Size Ranges on Annexure 3B | 1
926 | RT233-5.5.2e | See Table 13: Adjustable Modular Back Size Ranges on Annexure 3B | 1
| RT233-5.6 | LOW COST MODULAR SEATING SYSTEM | 8
927 | RT233-5.6.1 | Low-cost modular seating unit - seat board | 4
928 | RT233-5.6.1a | See Table 14: Low-Cost Modular Seating System Size Ranges on Annexure 3B | 1
929 | RT233-5.6.1b | See Table 14: Low-Cost Modular Seating System Size Ranges on Annexure 3B | 1
930 | RT233-5.6.1c | See Table 14: Low-Cost Modular Seating System Size Ranges on Annexure 3B | 1
931 | RT233-5.6.1d | See Table 14: Low-Cost Modular Seating System Size Ranges on Annexure 3B | 1
932 | RT233-5.6.2 | Low-cost modular seating unit - backrest board | 4
933 | RT233-5.6.2a | See Table 14: Low-Cost Modular Seating System Size Ranges on Annexure 3B | 1
934 | RT233-5.6.2b | See Table 14: Low-Cost Modular Seating System Size Ranges on Annexure 3B | 1
935 | RT233-5.6.2c | See Table 14: Low-Cost Modular Seating System Size Ranges on Annexure 3B | 1
936 | RT233-5.6.2d | See Table 14: Low-Cost Modular Seating System Size Ranges on Annexure 3B | 1
| RT233-5.7 | TRUCK SIDE SUPPORT | 7
| RT233-5.7.1 | Truck side support
937 | RT233-5.7.1a | Size: 10 cm x10 cm + 2 cm | 1
938 | RT233-5.7.1b | Size: 10 cm x12.5 cm +2 cm | 1
939 | RT233-5.7.1c | Size: 10 cm x15 cm + 2 cm | 1
940 | RT233-5.7.1d | Size: 15 cm x10 cm + 2 cm | 1
941 | RT233-5.7.1e | Size: 15 cm x12.5 cm +2 cm | 1
942 | RT233-5.7.1f | Size: 18cm x 10cm ± 2cm | 1
943 | RT233-5.7.1g | Size:18 cm x 12 cm ± 2cm | 1
| RT233-5.8 | BASE SEATING SYSTEM | 4
| RT233-5.8.1 | Pelvic Straps for all Wheelchairs (excluding posture chairs and buggies)
944 | RT233-5.8.1.1 | See Section 5.8 on Annexure 3B: Pelvic straps, basic adjustable, with D-ring fitting (5 cm) | 1
945 | RT233-5.8.1.2 | See Section 5.8 on Annexure 3B: Pelvic straps, basic adjustable, with D-ring fitting (3 cm) | 1
946 | RT233-5.8.1.3 | See Section 5.8 on Annexure 3B: Pelvic straps, basic adjustable, with male-female plastic buckle system | 1
947 | RT233-5.8.1.4 | See Section 5.8 on Annexure 3B: Pelvic straps, padded adjustable | 1
| RT233-5.9 | POSTURE SUPPORT SEAT SYSTEM | 4
| RT233-5.9.1 | Posture support seat system for use in transporting child in vehicle
948 | RT233-5.9.1a | Size: Seat width 30 cm, depth 31 cm, height 50 cm | 1
949 | RT233-5.9.1b | Size: Seat width 35 cm, depth 36 cm, height 60 cm | 1
950 | RT233-5.9.1c | Size: Seat width 40 cm, depth 41 cm, height 70 cm | 1
951 | RT233-5.9.1d | Size: Seat width 40 cm, depth 45 cm, height 80 cm | 1
6. TRAY TABLES
Item No. | Item Code | Description | Quantity |  | Minimum order Quantities (where applicable) | Import Percentage (where applicable) | Local Percentage (where applicable) | Country of Manufacture (Local or Imported) | Foreign Currency Name | Brand Name | Model Number | Delivery Period (Days) | Unit Of Measure (UOM) | Equipment Expected Lifespan (where applicable) | Unit Price | Comments
| RT233-6.1 | TRAY TABLE, BASIC WOOD | 18
| RT233-6.1 | Tray table to fit both basic and adjustable folding frame wheelchair | 9
952 | RT233-6.1a | To fit wheelchair 61 cm (24") wide (Total dimensions = 70+3 cm x 70+3 cm) | 1
953 | RT233-6.1b | To fit wheelchair 56 cm (22") wide (Total dimensions = 65+3 cm x 65+3 cm) | 1
954 | RT233-6.1c | To fit wheelchair 51 cm (20") wide (Total dimensions = 60+3 cm x 60+3 cm) | 1
955 | RT233-6.1d | To fit wheelchair 46 cm (18") wide (Total dimensions = 55+3 cm x 55+3 cm) | 1
956 | RT233-6.1e | To fit wheelchair 41 cm (16") wide (Total dimensions = 50+3 cm x 50+3 cm) | 1
957 | RT233-6.1f | To fit wheelchair 36 cm (14") wide (Total dimensions = 45+3 cm x 45+3 cm) | 1
958 | RT233-6.1g | To fit wheelchair 30 cm (12") wide with 40 cm seat frame length (Total dimensions = 40+3 cm x 40+3 cm) | 1
959 | RT233-6.1h | To fit wheelchair 30 cm (12") wide with 30 cm seat frame length (Total dimensions = 35+3 cm x 35+3 cm) | 1
960 | RT233-6.1i | To fit wheelchair 25 cm (10") wide with 30 cm seat frame length (Total dimensions = 35+3 cm x 35+3 cm) | 1
| RT233-6.2 | TRAY TABLE, BASIC PLASTIC
| RT233-6.2.1 | Tray table to fit both basic and adjustable folding frame wheelchair | 9
961 | RT233-6.2.1a | To fit wheelchair 61 cm (24") wide (Total dimensions = 70+3 cm x 70+3 cm) | 1
962 | RT233-6.2.1b | To fit wheelchair 56 cm (22") wide (Total dimensions = 65+3 cm x 65+3 cm) | 1
963 | RT233-6.2.1c | To fit wheelchair 51 cm (20") wide (Total dimensions = 60+3 cm x 60+3 cm) | 1
964 | RT233-6.2.1d | To fit wheelchair 46 cm (18") wide (Total dimensions = 55+3 cm x 55+3 cm) | 1
965 | RT233-6.2.1e | To fit wheelchair 41 cm (16") wide (Total dimensions = 50+3 cm x 50+3 cm) | 1
966 | RT233-6.2.1f | To fit wheelchair 36 cm (14") wide (Total dimensions = 45+3 cm x 45+3 cm) | 1
967 | RT233-6.2.1g | To fit wheelchair 30 cm (12") wide with 40 cm seat frame length (Total dimensions = 40+3 cm x 40+3 cm) | 1
968 | RT233-6.2.1h | To fit wheelchair 30 cm (12") wide with 30 cm seat frame length (Total dimensions = 35+3 cm x 35+3 cm) | 1
969 | RT233-6.2.1i | To fit wheelchair 25 cm (10") wide with 30 cm seat frame length (Total dimensions = 35+3 cm x 35+3 cm) | 1
7. POSITIONERS
Item No. | Item Code | Description | Quantity |  | Minimum order Quantities (where applicable) | Import Percentage (where applicable) | Local Percentage (where applicable) | Country of Manufacture (Local or Imported) | Foreign Currency Name | Brand Name | Model Number | Delivery Period (Days) | Unit Of Measure (UOM) | Equipment Expected Lifespan (where applicable) | Unit Price | Comments
| RT233-7.1 | SIDELYERS | 4
| RT233-7.1.1 | Positioner - sidelyer | 4
970 | RT233-7.1.1a | Size: Extra small 100 cm long x 60 cm wide, Back height 35 cm (Back height measured from top of lying support) | 1
971 | RT233-7.1.1a | Size: Small 120 long cm x 60 wide cm, Back height 35cm (Back height measured from top of lying support) | 1
972 | RT233-7.1.1a | Size: Medium 150 long cm x 60 wide cm, Back height 45cm (Back height measured from top of lying support) | 1
973 | RT233-7.1.1a | Size: Large 180 long cm x 70 wide cm, Back height 50cm (Back height measured from top of lying support) | 1
| RT233-7.2 | PRONE WEDGES | 4
| RT233-7.2.1 | Positioner - prone wedge
974 | RT233-7.2.1a | Size: Extra Small: length 50 cm x width 60 cm x 10 cm front height x 4 cm back height | 1
975 | RT233-7.2.1b | Size: Small: length 70 cm x width 60 cm x 15 cm front height x 6 cm back height | 1
976 | RT233-7.2.1c | Size: Medium: length 90 cm x width 60 cm x 20 cm front height x 8 cm back height | 1
977 | RT233-7.2.1d | Size: Large: length 110 cm x width 70 cm x 25 cm front height x 10 cm back height | 1
| RT233-7.3 | STANDING FRAMES | 9
| RT233-7.3.1 | Positioner - upright standing frame | 4
978 | RT233-7.3.1a | Size: Extra small: Heel to Head 65cm-100cm ;Hip width maximum 30cm | 1
979 | RT233-7.3.1b | Size: Small: Heel to head 85cm-120cm ;Hip width maximum 30cm | 1
980 | RT233-7.3.1c | Size: Medium: Heel to head 100cm-135cm;Hip width maximum 35cm | 1
981 | RT233-7.3.1d | Size: Large: Heel to head 135cm-160cm;Hip width maximum 40cm | 1
982 | RT233-7.3.2 | Positioner - prone standing frame, adult - Size: 78cm width, 84cm depth, height adjustable from 102 -122 cm | 1
| RT233-7.3.3 | Positioner - prone standing frame, adjustable forward tilt to prone 30 degrees | 4
983 | RT233-7.3.3a | Size: Extra small: Heel to Head 65cm-100cm ;Hip width maximum 30cm | 1
984 | RT233-7.3.3b | Size: Small: Heel to head 85cm-120cm ;Hip width maximum 30cm | 1
985 | RT233-7.3.3c | Size: Medium: Heel to head 100cm-135cm;Hip width maximum 35cm | 1
986 | RT233-7.3.3d | Size: Large: Heel to head 135cm-160cm;Hip width maximum 40cm | 1
1070

## 4 Category 1
CATEGORY 1
All items being tendered for must comply with the following technical specification, failure to comply with any of the conditions set out below may result in immediate disqualification. Please note that where the technical specification calls for certification, this certification must accompany the bid, failure to provide such certification will result in immediate disqualification. Bidders must give details of items being tendered for in the COMMENTS section where there are variation for further deliberation
WHEELCHAIRS AND COMMODES
RT233-1.1 | INDOOR / LOW-ACTIVE BASIC FOLDING FRAME WHEELCHAIRS | Indicate the item bidding for below bby stating "Yes bidding for this item". Item left blank will be assumed it is not bidded for | Complies Yes/No | Comments (Technical Specification Deviation)
Technical Specifications
1.1.1 | BASIC FOLDING FRAME, STANDARD BACK HEIGHT
Size range:
51 cm (20") wide
46 cm (18") wide
43 cm (17”) wide
41 cm (16") wide
38 cm (15") wide
36 cm (14") wide
30 cm (12") wide with 40 cm seat frame length
30 cm (12") wide with 30 cm seat frame length
25 cm (10”) wide with 30 cm seat frame length
1.1.2 | BASIC FOLDING FRAME, TALL BACK HEIGHT
Size range:
51 cm (20") wide
46 cm (18") wide
43 cm (17") wide
41 cm (16") wide
38 cm (15") wide
36 cm (14") wide
30 cm (12") wide with 40 cm seat frame length
30 cm (12") wide with 30 cm seat frame length
1.1.3 | BASIC FOLDING FRAME, SHORT BACK HEIGHT
Size range:
51 cm (20") wide
46 cm (18") wide
43 cm (17”) wide
41 cm (16") wide
38 cm (15") wide
36 cm (14") wide
30 cm (12") wide with 40 cm seat frame length
30 cm (12”) wide with 30 cm seat frame length
25 cm (10”) wide with 30 cm seat frame length
Backrest / Seat configuration: for Section 1.1.1 to 1.1.3 above • Seat tilt: 5 degrees minimum and 8 degrees maximum (approximate minimum of 3.5cm and maximum of 5.7cm front and rear seat height differential, on 40cm seat length). • Back to seat angle not to exceed 90 degrees. • 8-10 degrees bend in back post, except for short back and low-slung short back subcategories. • Back height: see Table 3 for measurements. • Seat length:  Standard = 40cm (allow + 1cm) long on all wheelchairs with 40cm seat frame. Exception: 30cm (allow + 1cm) long on 30cm (12") wide with 30cm seat. Frame length: 25cm (10”) wide with 30cm seat. Frame length; 25cm (10”) wide with 25cm seat frame length. Shorter length may be specified by end user.
See Table 3: Size Range for Basic Folding Frame Wheelchairs on Annexure ??
Upholstery: for Section 1.1.1 to 1.1.3 above • Back upholstery: Nylon (600 D or stronger) or vinyl: Exception - Reinforced on all 20” or reinforced 18” and 16” wheelchairs. Reinforcement by means of a single nylon strip, at least 4cm wide stitched to outside of back upholstery. • Seat upholstery: Nylon (600 D or stronger) or vinyl (40cm long or any shorter length as specified). • All back and seat upholstery to be reinforced with an inner Polyvinyl Chloride (PVC) minimum technical specification of 500g/m².The inner coated PVC to have a minimum tensile strength of 200kg/50mm (WMD) and 160kg/50mm (AMD) and minimum tear strength of 25kg both WMD and AMD. (Reinforced wheelchairs as specified will therefore have a second, outer layer of PVC for additional reinforcement. (Reinforcement by means of strips not acceptable).
Armrests: • Removable or fold-away, sport, full or desk with metal or durable plastic skirt guard. Exception: Height adjustable full or desk, with metal or durable plastic skirt guard only on 30cm (12") wide with 30cm seat frame length and 10” wheelchairs. • Arm rest sockets to be on side of seat rail only. • Shape of arm rests may in no way impede the fitting of commercial rigid adjustable back systems onto back posts.
Castors and forks: • 8x2" (20cm x 5cm) castors with plastic hub, outer, shore hardness of 65-70 and inner, shore hardness of 33-70, with the inner value not exceeding the outer value. Exception: 20” wheelchair - 8x2" (20cm x 5cm) castors with aluminium hub, and outer, shore hardness of 65-70 and inner, shore hardness of 33-70, with the inner value not exceeding the outer value. Alternative, 8x2” (20cm x 5cm) rubber castors. • 8x1” (20cm x 2.5cm) castors with plastic hub to be available as an option on request. • Forks to be minimum 0.4cm thick mild steel. Stem to be fastened to the fork with a thread and nut system to allow for re-fastening.
Rear wheels and axles: • 24" x 1 3/8 mag (standard) or spoke (optional) wheel with one-part semi-solid tyre as specified by end user. Exception, 12” wheelchair with 30cm seat frame length as well as 10” wheelchair, to use 22” x 1 3/8 mag (standard) or spoke (optional) wheel complete with semi-solid one-piece tyre. • Axle 1.2cm thick.
Footrest hangers: with 2cm variance • Extra-short (approximately 15cm), short (20cm), long (25cm) or fully detachable 90 degrees hangers with height, angle adjustable and reversible footplates as specified by end user.
Footrest: • Footplate to be hard heavy-duty plastic, aluminium, or mild steel • Height adjustable.
Leg strap: • Nylon strap (4cm to 5cm wide) with hook and loop fasteners. The leg strap should have hook and loop fasteners fitted to its full length to allow maximum adjustability.
Frame: • Basic, non-adjustable folding frame. • Finish: Epoxy coated.
Guarantee: • Frame and crossbars guaranteed for a minimum of 2 years from date of issue to end user.
Weight: • Not to exceed 23kg (complete). With rubber castors, allow +1 kg. • Not to exceed 16kg for 8” & 10” wheelchair. With rubber castors allow +1 kg.
Weight of user: • Visibly marked on each wheelchair to indicate maximum weight load of end user.
Repair / tool kit: • 4 rear wheels: bearings • 4 castors: axle bearings • 4 castor stem bearings • 4 seat guides (where applicable) • Adjustment tool
General requirements: • Bidder must be able to provide wheelchair with all abovementioned options. The end user will specify the desired option at the time of procuring. The different options available within each technical specification description must be fully interchangeable and available at the same price. Where the end user does not specify options, the wheelchair will be issued as in the option given first for the categories with more than one option. • No part or component may be changed deviate from the technical specifications set out in this document. • Bidder to submit accredited test report to ensure device durability
1.1 C | CUSTOMISING OPTIONS FOR INDOOR / LOW ACTIVE BASIC FOLDING FRAME WHEELCHAIRS
1.1.C.1 | One arm drive unit To be fitted onto the rear axle of the frame and should allow for propulsion of the wheelchair with either left or right hand only. The mounting frame or block should allow for positioning of the rear wheel axle to be in front of the original axle position. The connecting unit between the rear wheels should be a solid bar and have a quick release action to allow removal before folding the wheelchair. The concertina type connecting bar will not be accepted. This option must fit onto items 1.1 above
1.1.C.2 | Anti-tip levers (pair) Removable, length adjustable extensions that fit onto the tipping levers of the side frame to prevent backwards tipping of wheelchair; quick release mechanism; must be able to be mounted with anti-tip castor facing either down or up. This option must fit onto items 1.1 above. As specified by end user
1.1 C.3 | Amputation extension (pair) Removable extension brackets that can be fitted onto side frames and tipping levers to allow the rear wheels to be mounted 5 cm (+/- 1 cm) back from original setting. This option must fit onto items 1.1 above
1.1 C.4 | 45cm and 47.5cm seat extension 45cm and 47.5cm seats, with seat frame extension kit and full-length extended seat upholstery (short strip of upholstery not allowed) to lengthen original seat frames and seat. Upholstery screws to be provided. This option must fit onto items 1.1 above and should be available in the full wheelchair width range.
1.1 C.5 | Height adjustable desk armrest Height adjustable desk armrests, removable/foldaway, complete with mild steel or durable plastic skirt guard to fit basic folding frame manual wheelchair. Length 40 cm, height adjustment range 160-320mm (Range is measured from seat). This option must fit onto items 1.1 above. As specified by end user
1.1 C.6 | Nylon skirt guard (pair) Nylon skirt guards (15cm high x 25cm long) to be fitted onto seat and back of wheelchair to protect clothes from dirt on wheels when no armrests are used. This option must fit onto items 1.1 above. As specified by end user
1.1 C.7 | Bolt on footrest (pair) Small footplate (15 x 13cm) without hanger, with attachment to be fitted directly onto front post of side frame of wheelchair. This option must fit onto items 1.1 above
1.1 C.8 | Elevating footrest hanger (pair) Angle of footrest hanger to be adjustable between 0 degrees and 70 degrees of knee flexion; Height adjustable calf support; fitted with adult, narrow or juvenile footplates as specified by end user. To be ordered as a pair. This option must fit onto items 1.1 above.
1.1 C.9 | Elevating footrest hanger (individual) Angle of footrest hanger to be adjustable between 0 degrees and 70 degrees of knee flexion; Height adjustable calf support; Fitted with adult, narrow or juvenile footplates as specified by end user. To be ordered as an individual item specified Left or Right. This option must fit onto items 1.1 above.
1.1 C.10 | Footrest with 90 degrees hangers Footplate of footrests with 90 degrees hangers only in aluminium/mild steel or hard heavy duty plastic, must be height, angle and forward/backward adjustable.
1.1 C.11 | Rear wheels
1.1 C.11a | 24 x 1 3/8 spoke wheel (pair) 24 x 1 3/8 rear wheel rim with spokes, mild steel, complete with semi-solid one-piece tyre. This option must fit onto items 1.1 above; excluding 12” wheelchairs with 30cm seat frame length and 10” wheelchairs.
1.1 C.11b | 24 x 1¾ spoke wheel (pair) 24 x 1¾ rear wheel rim with spokes, mild steel, complete with semi-solid one-piece tyre. This option must fit onto items 1.1 above; excluding 12” wheelchairs with 30cm seat frame length and 10” wheelchairs.
1.1 C.11c | 22 x 1 3/8 spoke wheel (pair) 22 x 1 3/8 rear wheel rim with spokes, mild steel, complete with semi-solid one-piece tyre. This option must fit onto items 1.1 above; to fit only 12” wheelchairs with 30cm seat frame length and 10” wheelchairs.
1.1 C.12 | Front castors
1.1 C.12a | 20 x 5 cm (8 x 2”) castor with Aluminium hub (pair) 8x2" (20 x 5cm) castors with plastic hub, outer shore hardness of 65-70 and inner,, shore hardness of 33-70, with the inner value not exceeding the outer value. This option must fit onto items 1.1 above.
1.1 C.12b | 20 x 2.5 cm (8 x 1”) castor with plastic hub (pair) 20 x 2.5cm (8 x 1”) castor with plastic hub and rubber tyre. This option must fit onto items 1.1 above.
1.1 C.12c | 20 x 5 cm (8 x 2”) castor with plastic hub (pair) 8x2" (20 x 5cm) castors with plastic hub, outer shore hardness of 65-70 and inner shore hardness of 33-70, with the inner value not exceeding the outer value. This option must fit onto items 1.1 above.
1.1 C.13 | Quick-release axle for rear wheel 12 mm or 12.7 mm axle with release mechanism to allow for easy removal of the rear wheel. When in place must secure the rear wheel to allow for normal propulsion without limiting wheel rotation. This option must fit onto items 1.1 above.
1.1 C.14 | Head support, adjustable, basic occipital support only Easily removable Height and depth adjustable Kidney shaped head support flexible and rake adjustable Padded and covered in heavy gauge, removable, washable fabric. This option must be compatible with option 1.1.C.16 and also fit on items 1.1 above
1.1 C.15 | Head support, complex adjustable (hand shaped) Easily removable Height and depth adjustable 4-point, hand shaped head support, flexible and rake adjustable Moulded and padded Single Vehicle Approval (EVA) This option must be compatible with option 1.1.C.16 and also fit on items 1.1 above
1.1 C.16 | Removable universal head rest with horizontal mounting bar Universal bar that attaches in a horizontal position onto the back post/handles of the wheelchair to be fitted with a removable head rest. Mounting to the wheelchair: quick release mechanism to allow folding of wheelchair. Secure attachment to ensure safety when used as push handle. Mount for head rest: Adjustable mount to allow 4cm +- 1cm offset left and right from central position. Should be adjustable. This option must fit onto items 1.1 above.
1.1 C.17 | Sport armrest (pair) Removable/foldaway sport armrest, with durable plastic skirt guard. This option must fit onto items 1.1 above and should be available in each size.
1.1 C.18 | Full armrest (pair) Removable/foldaway full armrest with durable plastic skirt guard. This option must fit onto items 1.1 above and should be available in each size
1.1. C.19 | Desk armrest (pair) Removable/foldaway desk armrest with durable plastic skirt guard. This option must fit onto items 1.1 above and should be available in each size
1.1. C.20 | Standard Backrest, Standard height Back upholstery: Nylon (600 D or stronger) or vinyl: Exception: Reinforced on all 20” or reinforced 18” and 16” wheelchairs. Reinforcement by means of a single nylon strip, at least 4cm wide stitched to outside of back upholstery. Back upholstery to be reinforced with an inner PVC minimum technical specification of 500 g/m². The inner coated PVC to have a minimum tensile strength of 200 kg/50mm WMD and 160kg/50mm AMD and a minimum tear strength of 25kg both WMD and AMD. (Reinforced wheelchairs as specified will therefore have a second, outer layer of PVC for additional reinforcement. Reinforcement by means of strips not acceptable.)
1.1. C.21 | Standard Backrest, Tall backrest height Back upholstery: Nylon (600 D or stronger) or vinyl: Exception: Reinforced on all 20” or reinforced 18” and 16” wheelchairs. Reinforcement by means of a single nylon strip, at least 4cm wide stitched to outside of back upholstery. Back upholstery to be reinforced with an inner PVC minimum technical specification of 500 g/m². The inner coated PVC to have a minimum tensile strength of 200 kg/50mm WMD and 160kg/50mm AMD and a minimum tear strength of 25kg both WMD and AMD. (Reinforced wheelchairs as specified will therefore have a second, outer layer of PVC for additional reinforcement. Reinforcement by means of strips not acceptable.)
1.1. C.22 | Standard Backrest, Short backrest height Back upholstery: Nylon (600 D or stronger) or vinyl: Exception: Reinforced on all 20” or reinforced 18” and 16” wheelchairs. Reinforcement by means of a single nylon strip, at least 4cm wide stitched to outside of back upholstery. Back upholstery to be reinforced with an inner PVC minimum technical specification of 500 g/m². The inner coated PVC to have a minimum tensile strength of 200 kg/50mm WMD and 160kg/50mm AMD and a minimum tear strength of 25kg both WMD and AMD. (Reinforced wheelchairs as specified will therefore have a second, outer layer of PVC for additional reinforcement. Reinforcement by means of strips not acceptable.)
RT233-1.2 | BASIC FOLDING FRAME, STANDARD BACK HEIGHT, LOW SLUNG SEAT | Indicate the item bidding for below bby stating "Yes bidding for this item". Item left blank will be assumed it is not bidded for | Complies Yes/No | Comments (Technical Specification Deviation)
Size range:
51 cm (20") wide
46 cm (18") wide
43 cm (17”) wide
41 cm (16") wide
38 cm (15") wide
36 cm (14") wide
30 cm (12") wide with 40 cm seat frame length
30 cm (12”) wide with 30 cm seat frame length
25 cm (10”) wide with 30 cm seat frame length
Backrest / Seat configuration: for Section 1.2 above • Seat tilt: 5 degrees minimum and 8 degrees maximum (approximate minimum of 3.5cm and maximum of 5.7cm front and rear seat height differential on 40cm seat length). • Back to seat angle not to exceed 90 degrees. • 8-10 degrees bend in back post, except for short back and low-slung short back categories. • Back height: see Table 4 for measurements. • Seat length:  Standard = 40cm (allow + 1cm) long on all wheelchairs with 40cm seat frame.  Exception: 30cm (allow + 1cm) long on 30cm (12") wide with 30cm seat frame length.  Shorter length may be specified by end user
See Table 4: Range of Sizes for Low Slung Seat on Annexure ??
Upholstery: • Back upholstery: Nylon (600 D or stronger) or vinyl: Exception - Reinforced on all 20” or reinforced 18” and 16” wheelchairs. Reinforcement by means of a single nylon strip, at least 4cm wide stitched to outside of back upholstery. • Seat upholstery: Nylon (600 D or stronger) or vinyl (40cm long or any shorter length as specified). • All back and seat upholstery to be reinforced with an inner PVC minimum technical specification of 500 g/m². The inner coated PVC to have a minimum tensile strength of 200 kg/50mm WMD and 160kg/50mm AMD and a minimum tear strength of 25kg both WMD and AMD. (Reinforced wheelchairs as specified will therefore have a second, outer layer of PVC for additional reinforcement. Reinforcement by means of strips not acceptable).
Armrests: • Removable or fold away sport, full or desk with metal or durable plastic skirt guard. Exception: Height adjustable full or desk with metal or durable plastic skirt guard only on 30cm (12") wide with 30cm seat frame length and 10” wheelchairs. • Armrest sockets to be on side of seat rail only. • Shape of arm rests may in no way impede the fitting of commercial rigid adjustable back systems onto back posts. As specified by end user
Castors and forks: • 8x2" (20 x 5cm) castors with plastic hub, outer shore hardness of 65-70 and inner shore hardness of 33-70, with the inner value not exceeding the outer value. Exception: 20” wheelchair: 8x2" (20 x 5cm) castors with aluminium hub, and outer shore hardness of 65-70 and inner shore hardness of 33-70, with the inner value not exceeding the outer value. Alternative, 8x2” (20 x 5cm) rubber castors. • 8x1” (20 x 2.5cm) castors with plastic hub to be available as an option on request. • Forks to be minimum 4mm thick mild steel forks. Stem to be fastened to the fork with a thread and nut system to allow for re-fastening.
Rear wheels and axles: • 24" x 1 3/8 mag (standard) or spoke (optional) wheel with one-part semi-solid tyre as specified by end user. Exception, 12” wheelchair with 30cm seat frame length as well as 10” wheelchair, to use 22” x 1 3/8 mag (standard) or spoke (optional) wheel complete with semi-solid one-piece tyre. • Axle 12mm thick.
Footrest hangers: • Extra-short (approximately 15cm), short (20cm), long (25cm) or fully detachable 90-degree hangers with angle, height adjustable, reversible footplates as specified by end user.
Footrest: • Footplate to be hard heavy-duty plastic, or aluminium, or mild steel, and height adjustable.
Leg strap: • Nylon strap (4cm to 5cm wide) with hook and loop fasteners. The leg strap should have hook and loop fasteners fitted to its full length to allow maximum adjustability.
Frame: • Basic, non-adjustable folding frame. • Finish: Epoxy coated.
Guarantee: • Frame and crossbars guaranteed for a minimum of 2 years from date of issue to end user.
Weight: • Not to exceed 23kg (complete). With rubber castors, allow +1 kg.
Weight of user: • Visibly marked on each wheelchair to indicate maximum weight load of end user.
Repair / tool kit: • 4 rear wheels: bearings • 4 castors: axle bearings • 4 castor stem bearings • 4 seat guides (where applicable) • Adjustment tool
General requirement: • Bidder must be able to provide wheelchair with all abovementioned options. The end user will specify the desired option at the time of procuring. The different options available within each technical specification description must be fully interchangeable and available at the same price. Where the end user does not specify options the wheelchair will be issued as in the option given first for the categories with more than one option. • No part or component may be changed / deviate from the technical specifications set out in this document. • Bidder to submit accredited test report to ensure device durability
1.2 C | CUSTOMISING OPTIONS FOR INDOOR / LOW ACTIVE BASIC FOLDING FRAME WHEELCHAIRS, LOW SLUNG SEAT
1.2.C.1 | One arm drive unit To be fitted onto the rear axle of the frame and should allow for propulsion of the wheelchair with either left or right hand only. The mounting frame or block should allow for positioning of the rear wheel axle to be in front of the original axle position. The connecting unit between the rear wheels should be a solid bar and have a quick release action to allow removal before folding the wheelchair. The concertina type connecting bar will not be accepted. This option must fit onto items 1.1-1.3.
1.2. C.2 | Anti-tip levers (pair) Removable, length adjustable extensions that fit onto the tipping levers of the side frame to prevent backwards tipping of wheelchair; quick release mechanism; must be able to be mounted with anti-tip castor facing either down or up. This option must fit onto items 1.1-1.3. As specified by end user
1.2. C.3 | Amputation extension (pair) Removable extension brackets that can be fitted onto side frames and tipping levers to allow the rear wheels to be mounted 5 cm (+/- 1 cm) back from original setting. This option must fit onto items 1.1-1.3.
1.2. C.4 | 45cm and 47.5cm seat extension 45cm and 47.5cm seats, with seat frame extension kit and full-length extended seat upholstery (short strip of upholstery not allowed) to lengthen original seat frames and seat. Upholstery screws to be provided. This option must fit onto items 1.1-1.3. and should be available in the full wheelchair width range.
1.2 C.5 | Height adjustable desk armrest Height adjustable desk armrests, removable/foldaway, complete with mild steel or durable plastic skirt guard to fit basic folding frame manual wheelchair. Length 40 cm, height adjustment range 160-320mm (Range is measured from seat). This option must fit onto items 1.1-1.3. As specified by end user
1.2. C.6 | Nylon skirt guard (pair) Nylon skirt guards (15cm high x 25cm long) to be fitted onto seat and back of wheelchair to protect clothes from dirt on wheels when no armrests are used. This option must fit onto items 1.1-1.3. As specified by end user
1.2. C.7 | Bolt on footrest (pair) Small footplate (15 x 13cm) without hanger, with attachment to be fitted directly onto front post of side frame of wheelchair. This option must fit onto items 1.1-1.3.
1.2. C.8 | Elevating footrest hanger (pair) Angle of footrest hanger to be adjustable between 0 degrees and 70 degrees of knee flexion; Height adjustable calf support; fitted with adult, narrow or juvenile footplates as specified by end user. To be ordered as a pair. This option must fit onto items 1.1-1.3.
1.2. C.9 | Elevating footrest hanger (individual) Angle of footrest hanger to be adjustable between 0 degrees and 70 degrees of knee flexion; Height adjustable calf support; Fitted with adult, narrow or juvenile footplates as specified by end user. To be ordered as an individual item specified Left or Right. This option must fit onto items 1.1-1.3.
1.2 C.10 | Footrest with 90 degrees hangers Footplate of footrests with 90 degrees hangers only in aluminium/mild steel or hard heavy duty plastic, must be height, angle and forward/backward adjustable.
1.2. C.11 | Rear wheels
1.2. C.11.a | 24 x 1 3/8 spoke wheel (pair) 24 x 1 3/8 rear wheel rim with spokes, mild steel, complete with semi-solid one-piece tyre. This option must fit onto items 1.1-1.3; excluding 12” wheelchairs with 30cm seat frame length and 10” wheelchairs.
1.2. C.11.b | 24 x 1¾ spoke wheel (pair) 24 x 1¾ rear wheel rim with spokes, mild steel, complete with semi-solid one-piece tyre. This option must fit onto items 1.1-1.3; excluding 12” wheelchairs with 30cm seat frame length and 10” wheelchairs.
1.2. C.11.c | 22 x 1 3/8 spoke wheel (pair) 22 x 1 3/8 rear wheel rim with spokes, mild steel, complete with semi-solid one-piece tyre. This option must fit onto items 1.1-1.3; to fit only 12” wheelchairs with 30cm seat frame length and 10” wheelchairs.
1.2. C.12 | Front castors
1.2. C.12.a | 20 x 5 cm (8 x 2”) castor with Aluminium hub (pair) 8x2" (20 x 5cm) castors with plastic hub, outer shore hardness of 65-70 and inner,, shore hardness of 33-70, with the inner value not exceeding the outer value. This option must fit onto items 1.1-1.3.
1.2. C.12.b | 20 x 2.5 cm (8 x 1”) castor with plastic hub (pair) 20 x 2.5cm (8 x 1”) castor with plastic hub and rubber tyre. This option must fit onto items 1.1-1.3.
1.2. C.12.c | 20 x 5 cm (8 x 2”) castor with plastic hub (pair) 8x2" (20 x 5cm) castors with plastic hub, outer shore hardness of 65-70 and inner shore hardness of 33-70, with the inner value not exceeding the outer value. This option must fit onto items 1.1-1.3.
1.2. C.13 | Quick-release axle for rear wheel 12 mm or 12.7 mm axle with release mechanism to allow for easy removal of the rear wheel. When in place must secure the rear wheel to allow for normal propulsion without limiting wheel rotation. This option must fit onto items 1.1-1.3.
1.2. C.14 | Head support, adjustable, basic occipital support only Easily removable Height and depth adjustable Kidney shaped head support flexible and rake adjustable Padded and covered in heavy gauge, removable, washable fabric. This option must be compatible with option 1.1.C.16 and also fit on items 1.1-1.3
1.2. C.15 | Head support, complex adjustable (hand shaped) Easily removable Height and depth adjustable 4-point, hand shaped head support, flexible and rake adjustable Moulded and padded Single Vehicle Approval (EVA) This option must be compatible with option 1.1.C.16 and also fit on items 1.1-1.3
1.2. C.16 | Removable universal head rest with horizontal mounting bar Universal bar that attaches in a horizontal position onto the back post/handles of the wheelchair to be fitted with a removable head rest. Mounting to the wheelchair: quick release mechanism to allow folding of wheelchair. Secure attachment to ensure safety when used as push handle. Mount for head rest: Adjustable mount to allow 4cm +- 1cm offset left and right from central position. Should be adjustable. This option must fit onto items 1.1-1.3.
RT233-1.3 | BASIC FOLDING FRAME, SEMI-RECLINER | Indicate the item bidding for below bby stating "Yes bidding for this item". Item left blank will be assumed it is not bidded for | Complies Yes/No | Comments (Technical Specification Deviation)
Size range:
• 51 cm (20") wide
• 46 cm (18") wide
• 43 cm (17”) wide
• 41 cm (16") wide
• 38 cm (15”) wide
• 36 cm (14”) wide
• 30 cm (12”) wide
Backrest / Seat configuration:
• Seat tilt: 5 degrees minimum and 8 degrees maximum (approximate minimum of 3.5cm and maximum of 5.7cm front and rear seat height differential on 40cm seat length). • Back to seat angle : Maximum 120 degrees. • 8-10 degrees bend in back post. • Back height: see table 2 row V • Seat length:  Standard = 40cm (allow + 1 cm) long on all wheelchairs with 40cm seat frame.  Shorter length may be specified by end user.
Upholstery:
• Back upholstery: Nylon (600 D or stronger) or vinyl: Exception - Reinforced on all 20” or reinforced 18” and 16” wheelchairs. Reinforcement by means of a single nylon strip, at least 4cm wide stitched to outside of back upholstery. • Seat upholstery: Nylon (600 D or stronger) or vinyl (40 cm long or any shorter length as specified). • All back and seat upholstery to be reinforced with an inner PVC minimum technical specification of 500 g/m². The inner coated PVC to have a minimum tensile strength of 200 kg/50mm WMD and 160kg/50mm AMD and a minimum tear strength of 25kg both WMD and AMD. (Reinforced wheelchairs as specified will therefore have a second, outer layer of PVC for additional reinforcement. Reinforcement by means of strips not acceptable).
Armrests:
• Removable or fold-away, sport, full or desk with metal or durable plastic skirt guard. • Armrest sockets to be on side of seat rail only. • Shape of arm rests may in no way impede the fitting of commercial rigid adjustable back systems onto back posts.
Castors and forks:
• 8x2" (20 x 5cm) castors with plastic hub, outer shore hardness of 65-70 and inner shore hardness of 33-70, with the inner value not exceeding the outer value. Exception: 20” wheelchair: 8x2" (20 x 5cm) castors with aluminium hub, and outer shore hardness of 65-70 and inner shore hardness of 33-70, with the inner value not exceeding the outer value. Alternative, 8x2” (20 x 5cm) rubber castors. • 8x1” (20 x 2.5cm) castors with plastic hub to be available as an option on request. • Forks to be minimum 4mm thick mild steel forks. Stem to be fastened to the fork with a thread and nut system to allow for re-fastening.
Rear wheels and axles:
• 24" x 1 3/8 mag (standard) or spoke (optional) wheel with one-part semi-solid tyre as specified by end user. • Axle 12mm thick.
Footrest hangers:
• Extra-short (15cm), short (20cm), long (25cm) or fully detachable 90-degree hangers with angle, height adjustable, reversible footplates as specified by end user.
Footrest:
• Footplate to be hard heavy-duty plastic, or aluminium, or mild steel, and height adjustable.
Leg strap:
• Nylon strap (4cm to 5cm wide) with hook and loop fasteners. The leg strap should have hook and loop fasteners fitted to its full length to allow maximum adjustability.
Frame:
• Basic, non-adjustable folding frame. • Finish: Epoxy coated.
Guarantee:
• Frame and crossbars guaranteed for a minimum of 2 years from date of issue to end user.
Weight:
• Not to exceed 23kg (complete). With rubber castors, allow +1 kg.
Weight of user:
• Visibly marked on each wheelchair to indicate maximum weight load of end user.
Repair / tool kit:
• 4 rear wheels: bearings • 4 castors: axle bearings • 4 castor stem bearings • 4 seat guides (where applicable) • Adjustment tool
General requirement:
• Bidder must be able to provide wheelchair with all abovementioned options. The end user will specify the desired option at the time of procuring. The different options available within each technical specification description must be fully interchangeable and available at the same price. Where the end user does not specify options the wheelchair will be issued as in the option given first for the categories with more than one option. • No part or component may be changed / deviate from the specifications set out in this document. • Bidder to submit accredited test report to ensure device durability
1.3. C | CUSTOMISING OPTIONS FOR INDOOR / LOW ACTIVE BASIC FOLDING FRAME CHAIRS, SEMI-RECLINER WHEELCHAIRS
1.3. C.1 | Anti-tip levers (pair)
Removable, length adjustable extensions that fit onto the tipping levers of the side frame to prevent backwards tipping of wheelchair; Quick release mechanism; Must be able to be mounted with anti-tip castor facing either down or up. This option must fit onto item 1.3.
1.3. C.2 | 45cm and 47.5cm seats extension
45cm and 47.5cm seats, with seat frame extension kit and full-length extended seat upholstery (short strip of upholstery not allowed) to lengthen original seat frames and seats. Upholstery screws to be provided. This option must fit onto item 1.3 and should be available in each size.
1.3. C.3 | Nylon skirt guard (pair)
Nylon skirt guards (15cm high x 25cm long) that can be fitted onto seat and back of wheelchair to protect clothes from dirt on wheels when no armrests are used. This option must fit onto item 1.3.
1.3. C.4 | Bolt on footrest (pair)
Small footplate (15 x 13cm) without hanger, with attachment to be fitted directly onto front post of side frame of wheelchair. This option must fit onto item 1.3.
1.3. C.5 | Elevating footrest hanger (pair)
Angle of footrest hanger to be adjustable between 0 degrees and 70 degrees of knee flexion; Height adjustable calf support; Fitted with adult, narrow or juvenile footplates as specified by end user. To be ordered as a pair. This option must fit onto item 1.3.
1.3. C.6 | Elevating footrest hanger (individual)
Angle of footrest hanger to be adjustable between 0 degrees and 70 degrees of knee flexion; Height adjustable calf support; Fitted with adult, narrow or juvenile footplates as specified by end user. To be ordered as an individual item. This option must fit onto items 1.1-1.3.
1.3 C.7 | Footrest with 90 degrees hangers
Footplate of footrests with 90 degrees hangers only in aluminium/mild steel or hard heavy-duty plastic, must be height/ angle/ forward/ backward adjustable
1.3. C.8 | Rear wheels
1.3. C.8.1 | 24 x 1 3/8 spoke wheel
24 x 1 3/8 rear wheel rim with spokes, mild steel, complete with semi-solid one-piece tyre. This option must fit onto item 1.3.
1.3. C.8.2 | 24 x 1¾ spoke wheel
24 x 1¾ rear wheel rim with spokes, mild steel, complete with semi-solid one-piece tyre. This option must fit onto item 1.3.
1.3. C.9 | Front castors
1.3. C.9.1 20 x 5cm (8 x 2”) castor with aluminium hub
8x2" (20 x 5cm) castors with plastic hub, outer shore hardness of 65-70 and inner shore hardness of 33-70, with the inner value not exceeding the outer value. This option must fit onto item 1.3.
1.3. C.9.2 | 20 x 2.5cm (8 x 1”) castor with plastic hub
20 x 2.5cm (8 x 1”) castor with plastic hub and rubber tyre. This option must fit onto item 1.3.
1.3. C.9.3 | 20 x 5cm (8 x 2”) castor with plastic hub
8x2" (20 x 5cm) castors with plastic hub, outer shore hardness of 65-70 and inner shore hardness of 33-70, with the inner value not exceeding the outer value. This option must fit onto item 1.3.
1.3. C.10 | Quick-release axle for rear wheel
This option must fit onto item 1.3.
1.3. C.11 | Removable universal head rest with horizontal mounting bar
Universal bar that attaches in a horizontal position onto the back post/handles of the wheelchair to be fitted with a removable head rest. Mounting to the wheelchair: Quick release mechanism to allow folding of wheelchair. Secure attachment to ensure safety when used as push handle. Mount for head rest: Adjustable mount to allow 4cm +- 1cm offset left and right from central position. Should be adjustable. This option must fit on item 1.3.
RT233-1.4 | BARIATRIC FOLDING FRAME WHEELCHAIRS | Indicate the item bidding for below bby stating "Yes bidding for this item". Item left blank will be assumed it is not bidded for | Complies Yes/No | Comments (Technical Specification Deviation)
Size range:
• 51cm (20”)
• 56cm ( 22”)
• 61cm ( 24”)
Backrest / Seat configuration:
• Seat tilt: 5 degrees minimum and 8 degrees maximum (approximate minimum of 3.5cm and maximum of 5.7cm front and rear seat height differential on 40cm seat length). • Back to seat angle not to exceed 90 degrees. • 8-10 degrees bend in back post. • Back height: adjustable 40cm-50cm (allow + 1 cm). • Seat length:  Standard = 45cm (allow + 1 cm) long.  Shorter length may be specified by end user.
Upholstery:
• Back upholstery: Nylon (600 D or stronger) or vinyl: Reinforcement by means of a single nylon strip, at least 4cm wide stitched to outside of back upholstery. • Seat upholstery: Nylon (600 D or stronger) or vinyl (45cm long or any shorter length as specified). Reinforced wheelchairs as specified will therefore have a second, outer layer of PVC for additional reinforcement. Reinforcement by means of strips not acceptable. • All back and seat upholstery to be reinforced with an inner PVC minimum technical specification of 500 g/m². The inner coated PVC to have a minimum tensile strength of 200 kg/50mm WMD and 160kg/50mm AMD and a minimum tear strength of 25kg both WMD and AMD.
Armrests:
• Removable or fold away full or desk with aluminium or mild steel skirt guard. • Arm rest sockets to be on side of seat rail only. • Shape of arm rests may in no way impede the fitting of commercial rigid adjustable back systems onto back posts.
Castors and forks:
• 8x2" (20 x 5cm) castors with aluminium hub, outer, shore hardness of 65-70 and inner, shore hardness of 33-70, with the inner value not exceeding the outer value. Alternative, 8x2” (20 x 5cm) rubber castors. • 8x1” (20 x 2.5cm) castors with aluminium hub to be available as an option on request. • Forks to be minimum 4mm thick mild steel forks. Stem to be fastened to the fork with a thread and nut system to allow for re-fastening.
Rear wheels and axles:
• 24" x 1 3/8 mag (standard) or spoke (optional) wheel with one-part semi-solid tyre as specified by end user. • Axle 14mm thick.
Footrest hangers:
• Extra-short (approximately 15cm), short (20cm) or long (25cm) fully detachable hangers as specified by end user.
Footrest:
• Footplate to be hard heavy-duty plastic, or aluminium, or mild steel, and height adjustable.
Leg strap:
• Nylon strap (4cm to 5cm wide) with hook and loop fasteners. The leg strap should have hook and loop fasteners fitted to its full length to allow maximum adjustability.
Frame:
• Double cross bar. • Stabiliser bar to limit flexion / distortion of the frame. • Finish: Epoxy coated.
Guarantee:
• Frame and crossbars guaranteed for a minimum of 2 years from date of issue to end user.
Weight:
• Not to exceed 25kg (complete). With rubber castors, allow +1 kg.
Weight of user:
• Maximum user weight of 180kgs – 200kgs. To ensure the device fulfils the capacity requirement of Bariatric wheelchair. Indication of maximum weight load visibly marked on each wheelchair
Repair / tool kit:
• 4 rear wheels: bearings • 4 castors: axle bearings • 4 castor stem bearings • 4 seat guides (where applicable) • Adjustment tool
General requirement:
• Bidder must be able to provide wheelchair with all abovementioned options. The end user will specify the desired option at the time of procuring. The different options available within each technical specification description must be fully interchangeable and available at the same price. Where the end user does not specify options the wheelchair will be issued as in the option given first for the categories with more than one option. • No part or component may be changed / deviate from the technical specifications set out in this document. • Bidder to submit accredited test report to ensure device durability
1.4. C | CUSTOMISING OPTIONS FOR BARIATRIC WHEELCHAIRS
1.4. C.1 | Anti-tip levers (pair)
Removable, length adjustable extensions that fit onto the tipping levers of the side frame to prevent backwards tipping of wheelchair; Quick release mechanism; Must be able to be mounted with anti-tip castor facing either down or up. As specified by end user.
1.4. C.2 | 47.5cm seat extension
47.5cm seat, with seat frame extension kit and full-length extended seat upholstery (short strip of upholstery not allowed) to lengthen original seat frames and seat. Upholstery screws to be provided.
1.4. C.3 | Height adjustable desk or full arms (pair)
Height adjustable desk arms, removable, complete with mild steel or durable plastic skirt guard to fit basic folding frame manual wheelchair. Height adjustable range 160-320mm measured from seat.
1.4. C.4 | Elevating footrest hanger (pair)
Angle of footrest hanger to be adjustable between 0 degrees and 70 degrees of knee flexion; Height adjustable calf support; Fitted with adult, narrow or juvenile footplates as specified by end user. To be ordered as a pair.
1.4. C.5 | Elevating footrest hanger (individual)
Angle of footrest hanger to be adjustable between 0 degrees and 70 degrees of knee flexion; Height adjustable calf support; Fitted with adult, narrow or juvenile footplates as specified by end user. To be ordered as an individual item.
1.4. C.6 | Rear wheels
1.4. C.6.1 | 24 x 1 3/8” spoke wheel (pair)
24 x 1 3/8” rear wheel rim with spokes, mild steel, complete with semi-solid one-piece tyre.
1.4. C.6.2 | 24 x 1¾” spoke wheel (pair)
24 x 1¾ rear wheel rim with spokes, mild steel, complete with semi-solid one-piece tyre.
1.4. C.7 | Front castors
1.4. C.7.1 | 20 x 5cm (8 x 2”) castor with aluminium hub (pair)
8x2" (20cm x 5cm) castors with plastic hub, outer shore hardness of 65-70 and inner shore hardness of 33-70, with the inner value not exceeding the outer value.
1.4. C.7.2 | 20 x 2.5cm (8 x 1”) castor with aluminium hub (pair)
20cm x 2.5cm (8 x 1”) castor with aluminium hub and rubber tyre.
RT233-1.5 | URBAN ACTIVE WHEELCHAIRS | Indicate the item bidding for below bby stating "Yes bidding for this item". Item left blank will be assumed it is not bidded for | Complies Yes/No | Comments (Technical Specification Deviation)
RT233-1.5.1 | ADJUSTABLE FOLDING FRAME URBAN ACTIVE WHEELCHAIRS
Size range:
• 51 cm (20") wide with 43 + 1 cm seat frame length
• 46 cm (18") wide with 43 + 1 cm seat frame length
• 47 cm (17") wide with 43 + 1 cm seat frame length
• 41 cm (16") wide with 43 + 1 cm seat frame length
• 38 cm (15") wide with 43 + 1 cm seat frame length
• 36 cm (14") wide with 43 + 1 cm seat frame length
• 30 cm (12") wide with 43 + 1 cm seat frame length
• 30 cm (12") wide with 33 + 1 cm seat frame length
• 25 cm (10") wide with 33 + 1 cm seat frame length
Backrest / Seat configuration:
• Seat tilt: Adjustable, secondary to centre of gravity adjustment of wheelchair. • Standard Frame: Back to seat angle 90 degrees. For recline backrest options – see customising options • Angle of bend in back post to be 8-10 degrees. • Standard back height: Height adjustable: 40cm-50cm. For additional ranges- see customising options. • Standard seat depth = 40cm + 1cm long for additional ranges - see customising options. For the 12” with 33cm seat rail, the standard = 30 + 1cm long. For the 10“ with 33cm seat rail, the standard = 25 + 1cm long.
Centre of gravity adjustability:
• Adjustable by adjusting rear wheel - at least 3 settings for vertical and horizontal. • The adjustment range of the front castor should ensure that the castor stem always remains perpendicular to the ground within the available tilt in space range.
Upholstery:
• Back upholstery: Tension adjustable back upholstery. 3-piece tension adjustable with major components and quality as specified under tension adjustable backrests in this technical specification. • Seat upholstery: Nylon (600 D or stronger) (40cm long or any shorter length as specified). All seat upholstery to be reinforced with an inner PVC minimum technical specification of 500 g/m².
Armrests:
• Armrest sockets to be on side of seat rail only. • Shape of arm rests may in no way impede the fitting of commercial rigid adjustable back systems onto back posts.
Castors and forks:
• 8x1" (20 x 2,5cm) or 8x2" (5 x 20cm) castors as specified by end user. • 8x2" (20 x 5cm) castors with plastic hub, outer shore hardness of 65-70 and inner shore hardness of 33-70, with the inner value not exceeding the outer value. • Matching adjustments of the front castor to equal any adjustments done to the rear wheel position. • Forks to be minimum 0.4cm thick mild steel forks. Stem to be fastened to the fork with a thread and nut system to allow for re-fastening.
Rear wheels and axles:
• 24" x 1 3/8 mag wheel with one-part semi-solid tyre. Exception: 22” x 1 3/8 wheel on 30cm (12") wide with 33cm + 1cm seat frame length and 25cm (10") wide with 33 cm+ 1cm seat frame length.
Rear Wheel adjustments and axles:
• Vertical and horizontal adjustments of rear wheel (at least 3 settings each way). • Axle 12mm thick
Footrest hangers:
• Extra-short (approximately 15cm), short (20cm), long (25cm) or fully detachable 90-degree hangers with angle adjustable footplates as specified by end user.
Footrest:
• Footplate to be hard heavy-duty plastic, or aluminium, or mild steel, and height adjustable.
Leg strap:
• Nylon strap (4cm-5cm wide) with hook and loop fasteners. The leg strap should have hook and loop fasteners fitted to its full length to allow maximum adjustability.
Guarantee:
• Frame and crossbars guaranteed for a minimum of 2 years from date of issue to end user.
Frame:
• Adjustable, folding frame • Aluminium / mild steel • 30cm (12") wide with 33 + 1cm seat frame length and 25cm (10") wide with 33 + 1cm seat frame length has shorter wheelbase and overall length as the other sizes in this range.
Weight:
• Not to exceed 14kg, excluding rear wheels.
Weight of user:
• Visibly marked on each wheelchair to indicate maximum weight load of end user.
Repair / tool kit:
• 4 rear wheels: bearings • 4 castors: axle bearings • 4 castor stem bearings • 4 seat guides (where applicable) • Adjustment tool
General requirement:
• Bidder must be able to provide wheelchair with all abovementioned options. The end user will specify the desired option at the time of ordering. The different options available within each technical specification description must be fully interchangeable and available at the same price. Where the end user does not specify options the wheelchair will be issued as in the option given first for the categories with more than one option. • No part or component may be changed / deviate from the specifications set out in this document. • Bidder to submit accredited test report to ensure device durability
1.5.1 C | CUSTOMISING OPTIONS FOR ADJUSTABLE FOLDING FRAME URBAN ACTIVE WHEELCHAIRS
1.5.1 C.1 | One arm drive unit
A unit that can be fitted onto the rear axle of the frame and allows for propulsion of the wheelchair with either left or right hand only. The mounting frame or block should allow for positioning of the rear wheel axle to be in front of the original axle position. The connecting unit between the rear wheels should be a solid bar and have a quick release action to allow removal before folding the wheelchair. The concertina type connecting bar will not be accepted. This option must fit onto item 1.5.1.
1.5.1 C.2 | Anti-tip levers (pair)
Removable, length adjustable extensions that fit onto the tipping levers of the side frame to prevent backwards tipping of wheelchair; Quick release mechanism; Must be able to be mounted with anti-tip castor facing either down or up. This option must fit onto item 1.5.1.
1.5.1 C.3 | Option to accommodate extra thigh length using any one the following:
1.5.1 C.3.1 | For sizes 12” with standard frame and 14” and larger:
Medium Frame length and seat depth: 40cm +1cm; 43cm +1cm; or 45cm +1 cm
For sizes: 12” with 33cm seat frame: 33cm +1cm seat ext For sizes: 10” with 33cm seat frame: 30cm +1cm seat ext
1.5.1 C.3.2 | Long frame length and seat depth:
43cm +1cm 45cm +1cm or 47cm +1cm
1.5.1 C.3 3 | Frame with seat depth adjustability range. Minimum 40cm-50 cm. End user to specify option and seat depth and frame length.
1.5.1 C.4 | Back height
Adjustable back height ranges: 30 – 40cm (no bend), 35 – 45 cm (7 bend), 40 – 50 cm (7 bend), 45 – 55 cm (7 bend)
1.5.1 C.5 | Nylon skirt guard (pair)
Nylon skirt guards (15cm high x 25cm long) that can be fitted onto seat and back of wheelchair to protect clothes from dirt on wheels when no armrests are used. This option must fit onto item 1.5.1.
1.5.1 C.6 | Bolt on footrest (pair)
Small footplate (15 x 13cm) without hanger, with attachment to be fitted directly onto front post of side frame of wheelchair. This option must fit onto item 1.5.1.
1.5.1 C.7 | Elevating footrest hanger (pair)
Angle of footrest hanger to be adjustable between 0 degrees and 70 degrees of knee flexion; Height adjustable calf support; Fitted with adult, narrow or juvenile footplates as specified by end user. To be ordered as a pair. This option must fit onto item 1.5.1.
1.1. C.8 | Elevating footrest hanger (individual)
Angle of footrest hanger to be adjustable between 0 degrees and 70 degrees of knee flexion; Height adjustable calf support; Fitted with adult, narrow or juvenile footplates as specified by end user. To be ordered as an individual item.
1.5.1 C.9 | Footrest with 90 degrees hangers
Footplate of footrests with 90 degrees hangers only in aluminium/mild steel or hard heavy-duty plastic, must be height/ angle/ forward/ backward adjustable.
1.5.1 C10 | . Rear wheels
1.5.1 C.10.1 | Spoke wheel: 24 x 1 3/8” (pair)
24 x 1 3/8” rear wheel rim with spokes, mild steel, complete with semi-solid one-piece tyre. This option must fit onto item 1.5.1.
1.5.1 C.10.2 | Spoke wheel: 24 x 1¾” (pair)
24 x 1¾ rear wheel rim with spokes, mild steel, complete with semi-solid one-piece tyre. This option must fit onto item 1.5.1.
1.5.1 C.11 | Front castors
1.5.1 C.11.1 | Castor with aluminium hub: 20 x 5cm (8 x 2”) (pair)
8x2" (20 x 5cm) castors with plastic hub, outer, shore hardness of 65-70 and inner, shore hardness of 33-70, with the inner value not exceeding the outer value. This option must fit onto item 1.5.1.
1.5.1 C.11.2 | Castor with plastic hub: 20 x 2.5cm (8 x 1”) (pair)
20 x 2.5cm (8 x 1”) castor with plastic hub and rubber tyre. This option must fit onto item 1.5.1.
1.5.1 C.11.3 | Castor with plastic hub: 20 x 5cm (8 x 2”) (pair)
8x2" (20 x 5cm) castors with plastic hub, outer, shore hardness of 65-70 and inner, shore hardness of 33-70, with the inner value not exceeding the outer value. This option must fit onto item 1.5.1.
1.5.1 C.12 | Head support, adjustable, basic occipital support only
Easily removable Height and depth adjustable Kidney shaped head support flexible and rake adjustable Padded and covered in heavy gauge, removable, washable fabric. This option must be compatible with option 1.5.1.C.14 and also fit on item 1.5.1
1.5.1. C.13 | Head support, complex adjustable (hand shaped)
Easily removable Height and depth adjustable 4-point, hand shaped head support, flexible and rake adjustable Moulded and padded Single Vehicle Approval (EVA) This option must be compatible with option 1.5.1.C.14 and also fit on item 1.5.1
1.5.1 C.14 | Removable universal head rest with horizontal mounting bar
Universal bar that attaches in a horizontal position onto the back post/handles of the wheelchair to be fitted with a removable head rest. Mounting to the wheelchair: Quick release mechanism to allow folding of wheelchair. Secure attachment to ensure safety when used as push handle. Mount for head rest: Adjustable mount to allow 4cm +- 1cm offset left and right from central position. Should be adjustable. This option must fit on item 1.5.1.
1.5.1 C.15 | Adjustable folding frame with a tool adjustable recline back post
Adjustable back to seat angle 90-105 degrees. Adjustment mechanism tool adjustable with minimum 3 adjustment increments.
1.5.1 C.16 | Adjustable backrest height ranges:
35 – 45 cm (straight back posts), 40 – 50 cm (straight back posts), 45 – 55 cm (straight back posts)
RT233-1.5.2 | ADJUSTABLE FOLDING FRAME URBAN ACTIVE WHEELCHAIR - LIGHT WEIGHT
Size range:
• 51 cm (20”) wide with 43 + 1 cm seat frame length • 46 cm (18") wide with 43 + 1 cm seat frame length • 41 cm (16") wide with 43 + 1 cm seat frame length • 36 cm (14") wide with 43 + 1 cm seat frame length • 30 cm (12”) wide with 43 + 1 cm seat frame length • 30 cm (12”) wide with 33 + 1 cm seat frame length • 25 cm (10”) wide with 33 + 1 cm seat frame length
Backrest / Seat configuration:
• Seat tilt: Adjustable, secondary to centre of gravity adjustment of wheelchair. • Standard frame: Back to seat angle: 90 degrees. For recline backrest options – see below customising options • Standard back height: Height adjustable 40cm-50cm. For additional ranges see customising options. • Standard seat depth = 40cm + 1cm long for additional ranges see customising options. For the 12” with 33cm seat rail, the standard = 30 + 1cm long. For the 10“ with 33cm seat rail, the standard = 25 + 1cm long
Centre of gravity adjustability:
• Adjustable by adjusting rear wheel - at least 3 settings for vertical and horizontal. • The adjustment range of the front castor should ensure that the castor stem always remains perpendicular to the ground within the available tilt in space range.
Upholstery:
• Back upholstery: Tension adjustable back upholstery. 3-piece tension adjustable with major components and quality as specified under tension adjustable backrests in this technical specification. • Seat upholstery: Nylon (600 D or stronger) (40cm long or any shorter length as specified). All seat upholstery to be reinforced with an inner PVC minimum technical specification of 500 g/m².
Height adjustable Armrests:
• Flip back and removable, full or desk, with durable plastic skirt guard. • Shape of arm rests may in no way impede the fitting of commercial rigid adjustable back systems onto back posts.
Castors and forks:
• Castors: 6 x 1" (15 x 2.5cm). • Stem to be fastened to the fork with a thread and nut system to allow for re-fastening. • Matching adjustments of the front castor to equal any adjustments done to the rear wheel position. • Forks to be minimum 4mm thick mild steel or aluminium. Stem to be fastened to the fork with a thread and nut system to allow for re-fastening.
Rear wheels and axles:
• 24" x 1 3/8 mag wheel with one-part semi-solid tyre. Exception: 22” x 1 3/8 wheel on 30cm (12") wide with 33 + 1cm seat frame length and 25cm (10") wide with 33 + 1cm seat frame length.
Rear wheel adjustment and axles:
• Vertical and horizontal adjustments of rear wheel (at least 3 settings each way). • Axle 12mm thick.
Footrest hangers:
• Fixed or swing away. As specified by end user
Footrest:
• Footplate to be hard heavy-duty plastic, or aluminium, or mild steel, and height adjustable and flip up
Leg strap:
• Nylon strap (4-5cm wide) with hook and loop fasteners. The leg strap should have hook and loop fasteners fitted to its full length to allow maximum adjustability
Guarantee:
• Frame and crossbars guaranteed for a minimum of 2 years from date of issue to end user.
Frame:
• Adjustable, folding frame • Aluminium / mild steel • 30cm (12") wide with 33 + 1cm seat frame length and 25cm (10") wide with 33 + 1cm seat frame length has shorter wheelbase and overall length as the other sizes in this range.
Weight:
• Not to exceed 12kg, excluding rear wheels
Weight of user:
• Visibly marked on each wheelchair to indicate maximum weight load of end user.
Repair / tool kit:
• 4 rear wheels: bearings • 4 castors: axle bearings • 4 castor stem bearings • 4 seat guides (where applicable) • Adjustment tool
General requirement:
• Bidder must be able to provide wheelchair with all abovementioned options. The end user will specify the desired option at the time of ordering. The different options available within each technical specification description must be fully interchangeable and available at the same price. Where the end user does not specify options the wheelchair will be issued as in the option given first for the categories with more than one option. • No part or component may be changed / deviate from the specifications set out in this document. • Bidder to submit accredited test report to ensure device durability
1.5.2. C | CUSTOMISING OPTIONS FOR ADJUSTABLE FOLDING FRAME URBAN ACTIVE CHAIRS – LIGHT WEIGHT
1.5.2. C.1 Anti-tip levers (pair)
Removable, length adjustable extensions that fit onto the tipping levers of the side frame to prevent backwards tipping of wheelchair; Quick release mechanism; Must be able to be mounted with anti-tip castor facing either down or up. This option must fit onto item 1.5.2.
1.5.2.C.2. Option to accommodate extra thigh length using any one the following:
1.5.2.C.2.1 For sizes 12” with standard frame and 14” and larger:
Medium Frame length and seat depth: 40cm +1cm; 43cm +1cm ; 45cm +1 cm For sizes: 12” with 33cm seat frame: 33cm +1cm seat ext For sizes: 10” with 33cm seat frame: 30cm +1cm seat ext
1.5.1 C.2.2 Long frame length and seat depth:
43cm +1 45cm +1cm or 47cm +1cm
1.5.2.C.2 3 Frame with seat depth adjustability range. Minimum 40cm-50cm. End user to specify option and seat depth and frame length
1.5.2.C.3 Back height
Adjustable back heigh ranges: 30 – 40cm (no bend), 35 – 45 cm (7 bend), 40 – 50 cm (7 bend), 45 – 55 cm (7 bend)
1.5.2.C.4 Adjustable folding frame with an adjustable recline back post
Adjustable back to seat angle 90⁰-105⁰ degrees. Adjustment mechanism tool adjustable with minimum 3 adjustment increments of 5⁰ (0⁰, -5⁰, -10⁰, -15⁰)
1.5.2. C.5 Rear wheels
1.5.2. C.5.1 Spoke wheel: 24 x 1 3/8” (pair)
24 x 1 3/8” rear wheel rim with spokes, mild steel, complete with semi-solid one-piece tyre. This option must fit onto item 1.5.2.
1.5.2. C.5.2 Spoke wheel: 24 x 1¾” (pair)
24 x 1¾ rear wheel rim with spokes, mild steel, complete with semi-solid one-piece tyre. This option must fit onto item 1.5.2.
1.5.2. C.6 Front castors
1.5.2. C.6.1 Castors: 4 x1" (10cm x 2.5cm)
This option must fit onto item 1.5.2.
1.5.2. C.6.2 Castors: 5"x1” (12cm x 2.4cm)
This option must fit onto item 1.5.2.
1.5.2. C.67 Adjustable back height ranges:
35 – 45 cm (straight back posts), 40 – 50 cm (straight back posts), 45 – 55 cm (straight back posts)
RT233-1.5.3 | ADJUSTABLE RIGID FRAME WHEELCHAIRS, 4-WHEELER URBAN ACTIVE WHEELCHAIR WITH STANDARD OR LONG FRAME)
A | Wheelbase (Allow + 1cm)
B | Overall length (maximum, with footrest set fully extended, allowing 5cm clearance from the ground).
C | Back upholstery height (measured from top of seat rail to top of back rest upholstery).
D | Seat width measured from outside of one seat rail to outside of opposite seat rail.
E | Seat length measured from front of back post to front of seat upholstery.
See Table 5: Adjustable rigid frame 4-wheeler urban active wheelchair on Annexure 3B
Size range:
RT233-1.5.3.1 | ADJUSTABLE RIGID FRAME, STANDARD FRAME LENGTH
• 46 cm (18") wide • 43 cm (17") wide • 41 cm (16") wide • 38 cm (15") wide • 35 cm (14") wide • 30 cm (12") wide with 40 cm seat frame length • 30 cm (12") wide with 30 cm seat frame length • 25 cm (10”) wide with 30 cm seat frame length but 25 cm seat
RT233-1.5.3.2 | ADJUSTABLE RIGID FRAME, LONG FRAME LENGTH
• 46 cm (18") wide • 43 cm (17") wide • 41 cm (16") wide • 38cm (15") wide • 35 cm (14") wide • 30 cm (12") wide with 40 cm seat frame length
Backrest / Seat configuration:
• Back to seat angle adjustable and not to exceed 85-100 degrees range, adjustment tool adjustable with additional quick release adjustment, nonfraying. • Seat angle adjustable, must be able to adjust to level (horizontal) position. • No bend in back post. • Back height: Low back = 28-38 + 1cm, standard back = 33-43 + 1cm. • Push handles attached to back posts. • Extended push handle on 10” and 12” wheelchair. • Seat depth:  + 1cm long on all wheelchairs with standard frame.  43 + 1cm on all wheelchairs with long frame.  Exception: 30 + 1cm long on 30cm (12") wide with 30cm seat frame length and 25 + 1cm on.  25cm (10”) wide with 30cm seat frame length.  Shorter length may be specified by end user.
Centre of gravity adjustability:
• Seat tilt : 5degrees – 10degrees • Adjustability of Centre of Gravity by changing both seat angle and either seat-back system horizontal and vertical position or changing wheel position both vertically and horizontally. • Minimum of 5 horizontal settings, 1 negative, 1 neutral and 3 forward. (Seat or rear wheel). If infinitely adjustable, should have minimum adjustment range of 3cm.
Upholstery:
• Back upholstery: Tension adjustable back upholstery. 3-piece tension adjustable with major components and quality as specified under tension adjustable backrests in this technical specification. • Seat upholstery: 40 + 1cm long on all wheelchairs with standard frame. 43 + 1cm on all wheelchairs with long frame. Shorter length may be specified by end user. Exception: 30 + 1cm long on 30cm (12") wide with 30cm seat frame length or any shorter length as specified. Fabric to be Nylon (600 D or stronger). All seat upholstery to be reinforced with an inner PVC minimum technical specification of 500 g/m².
Skirt guards:
• Skirt guard to cover wheel rim arch protruding above seat. • No arm rests.
Castors and forks:
• Castors: 6 x 1" (15 x 2.5cm). • Stem to be fastened to the fork with a thread and nut system to allow for re-fastening.
Rear wheels and axles:
• 24" x 1 3/8 mag wheel with one-part semi-solid tyre. Exception, 12” wheelchair with 30cm seat rail length and 25cm (10") wide wheelchair to use 22” x 1 3/8 mag (standard) or spoke (optional) wheel complete with semi-solid one-piece tyre
Quick release axles:
• 12 mm or 12.7 mm axle.
Footrest hangers:
• Integral part of frame.
Footrest:
• Fixed angle adjustable, height adjustable without modification with broad weight-bearing surface. On 12” and 10” wheelchairs, the footplate to be on 15cm elevations.
Leg strap:
• Nylon strap (4-5cm wide) with hook and loop fasteners. The leg strap should have hook and loop fasteners fitted to its full length to allow maximum adjustability.
Frame:
• Adjustable rigid frame. • Mild steel or aluminium. • Long frame is 5cm longer than standard frame on all sizes except 30cm with 30cm seat rail, where difference is between 4 and 5cm. The wheelbase is measured with castor fork pointing back to rear wheel from centre of castor to centre of rear wheel along the horizontal plane (as per diagram for basic folding frame wheelchair). • Standard frame wheelbase is 38 + 1cm. Long frame is 43 + 1cm. For the 30cm wheelchair with 30cm seat rail, the standard frame is 33 + 1cm. • 30cm (12") wide with 33 + 1cm seat frame length and 25cm (10") wide with 33 + 1cm seat frame length has shorter wheelbase and overall length as the other sizes in this range.
Guarantee:
• Frame and crossbars guaranteed for a minimum of 2 years from date of issue to end user.
Weight:
• Not to exceed 15kg excluding rear wheels
Weight of user:
• Visibly marked on each wheelchair to indicate maximum weight load of end user.
Repair / tool kit:
• 4 rear wheels: bearings • 4 castors: axle bearings • 4 castor stem bearings • 4 seat guides (where applicable) • Adjustment tool • Tension adjustable filler straps to allow for adjustment of back height.
General requirement:
• Bidder must be able to provide wheelchair with all abovementioned options. The end user will specify the desired option at the time of ordering. The different options available within each technical specification description must be fully interchangeable and available at the same price. Where the end user does not specify options the wheelchair will be issued as in the option given first for the categories with more than one option. • No part or component may be changed / deviate from the specifications set out in this document. • Bidder to submit accredited test report to ensure device durability
1.5.3. C | CUSTOMISING OPTIONS FOR ADJUSTABLE RIGID FRAME 4-WHEELER URBAN ACTIVE WHEELCHAIRS WITH STANDARD OR LONG FRAME
1.5.3.C.1. Option to accommodate extra thigh length using any one the following:
1.5.3.C.1.1 Seat depth (extension): Standard frame
38cm +1cm 41cm +1cm 43cm +1cm;
1.5.3.C.1.2 Long wheelchair frame:
45cm +1cm; 47cm+1cm 49 cm +1.
1.5.3.C.1 3 Frame with seat depth adjustability range. Minimum 40cm-50cm. End user to specify option and seat depth and frame length.
1.5.3.C.2 Short footplate hanger inserts
Special insert to increase the range of the standard footplate height in all wheelchairs in 1.5.3.
1.5.3.C.3 Extended push handles
Extended T-handle style push handles with mounting mechanism to mount onto rear axle. Quick-lease mechanism to remove handles from wheelchair and mounting mechanism.
1.5.3. C.4 Quick-release clip-on castor wheel unit
A large castor wheel unit that can be fitted onto the frame of the wheelchair to turn the wheelchair into a 3-wheeler and allows for the front castors to clear the ground. Should have a quick-lease mechanism to attach and remove device from wheelchair. Secure attachment to ensure safety when used. When detached, the unit should fit securely into an assembly onto the back of the wheelchair.
1.5.3 C 5 Backrest height:
Adjustable between 36cm – 46 +1cm
RT233-1.6 | HYBRID ACTIVE WHEELCHAIRS | Indicate the item bidding for below bby stating "Yes bidding for this item". Item left blank will be assumed it is not bidded for | Complies Yes/No | Comments (Technical Specification Deviation)
A | Wheelbase (Allow + 1cm).
B | Overall length (Maximum, with footrest set fully extended, allowing 5cm clearance from the ground).
C | Back upholstery height (measured from top of seat rail to top of back rest upholstery).
D | Seat width measured from outside of one seat rail to outside of opposite seat rail
E | Seat length measured from front of back post to front of seat upholstery.
See Table 6: Hybrid Active Wheelchairs on Annexure 3B
RT233-1.6.1 | ADJUSTABLE, 3-WHEELER HYBRID
Size range:
• 46 cm (18") wide • 43 cm (17") wide • 41 cm (16") wide • 38 cm (15") wide • 35 cm (14") wide • 34 cm (13”) wide • 30 cm (12") wide with 40 cm seat frame length • 30 cm (12") wide with 30cm seat frame length • 25 cm (10”) wide with 25 cm seat frame length
Backrest / Seat configuration:
• Seat tilt: Adjustable. • Back to seat angle not to exceed 70-98 degrees. • No bend in back post. • Back height: Low back = 28-38 + 1cm, standard back = 33-43 + 1cm. • Push handles attached to back posts. • Seat length:  Standard = 40cm + 1cm long on all wheelchairs with 40cm seat frame.  Exception: 30cm + 1cm long on 30cm (12") wide with 30cm seat frame length and 25cm + 1cm long on 25cm (10”) wide with 25cm seat frame length.  Shorter length may be specified by end user.
Centre of gravity adjustability:
• Adjustability of Centre of Gravity by positioning rear wheel and tilting seat angle.
Upholstery:
• Back upholstery: Tension adjustable back upholstery. 3-piece tension adjustable with major components and quality as specified under tension adjustable backrests in this technical specification. • Seat length: Standard = 40cm + 1cm long on all wheelchairs with 40cm seat frame. Shorter length may be specified by end user. Exception: 30 + 1cm long on 30cm (12") wide with 30cm seat frame length and 25cm on 25cm (10”) wide or any shorter length as specified. Fabric to be Nylon (600 D or stronger). All seat upholstery to be reinforced with an inner PVC minimum technical specification of 500 g/m².
Skirt guards:
• Skirt guard to cover wheel rim arch protruding above seat. • No arm rests.
Front Castors and forks:
• 8x2" (20 x 5cm) castors with plastic hub, outer shore hardness of 65-70 and inner shore hardness of 33-70, with the inner value not exceeding the outer value. Alternative 8x2” (20 x 5cm) rubber castors. • Forks to be minimum 4mm thick mild steel forks. Stem to be fastened to the fork with a thread and nut system to allow for re-fastening.
Rear wheels and axles:
• 24" x 1 3/8 mag wheel with one-part semi-solid tyre. Exception, 12” wheelchair with 30cm seat rail length and 25cm+ 1cm long on 25cm (10”) wide with 25cm seat frame length to use 22” x 1 3/8 mag (standard) or spoke (optional) wheel complete with semi-solid one-piece tyre.
Quick release axles:
• 2 degrees camber. • At least 3 vertical and 4 horizontal adjustments on rear wheel axle position. • Axle 12 mm thick.
Footrest hangers:
• Integral part of frame.
Footrest:
• Fixed or flip-up, angle adjusted and height adjustable without modification.
Heel strap:
• Nylon heel strap (4cm-5cm wide).
Guarantee:
• Frame and crossbars guaranteed for a minimum of 2 years from date of issue to end user.
Weight:
• Not to exceed 16kg (complete). With rubber castor, allow +0.5kg. • 12” not to exceed 14kg.
Weight of user:
• Visibly marked on each wheelchair to indicate maximum weight load of end user.
Repair / tool kit:
• 4 rear wheels: bearings • 4 castors: axle bearings • 4 castor stem bearings • 4 seat guides (where applicable) • Adjustment tool • Tension adjustable filler straps to allow for adjustment of back height.
General requirement:
• Bidder must be able to provide wheelchair with all abovementioned options. The end user will specify the desired option at the time of procuring. The different options available within each technical specification description must be fully interchangeable and available at the same price. Where the end user does not specify options the wheelchair will be issued as in the option given first for the categories with more than one option. • No part or component may be changed / deviate from the technical specifications set out in this document. • Bidder to submit accredited test report to ensure device durability
Frame:
• Adjustable three-wheeler. • Rigid; mild steel or aluminium. • The wheelbase is measured with castor fork pointing back to rear wheel from centre of castor to centre of rear wheel along the horizontal plane (as per diagram for basic folding frame wheelchair). Short wheelbase (active position) is 68cm + 1cm. Long wheelbase (least active / safe) is 76cm + 1cm. For the 30cm wheelchair with 30cm seat rail, the range is 59cm + 1cm to 67 + 1cm. • Bearing spacer to be used in all bearing applications, i.e. rear wheel, front castor, and castor stem.
1.6.1. C | CUSTOMISING OPTIONS FOR ADJUSTABLE 3-WHEELER HYBRID ACTIVE WHEELCHAIRS
1.6.1. C.1 43cm and 45cm extended seat length
43cm and 45cm extended seats upholstery, with seat frames and screws complete. The extended seat is a complete unit, and a short strip extension of upholstery is not allowed. This option must fit onto item 1.6.1 and should be available in each size.
1.6.1. C.2 Extended push handles
Extended T-handle style push handles with mounting mechanism to mount onto rear axle. Quick-lease mechanism to remove handles from wheelchair and mounting mechanism.
1.6.1 C 3 Backrest height:
Adjustable between 36cm – 46 +1cm
1.6.1 C 4 Seat depth:
Option to accommodate extra thigh length using any one the following: Seat depth (extension): Standard frame 38cm +1cm 41cm +1cm 43cm +1cm; Long wheelchair frame: 45cm +1cm; 47cm+1cm 49 cm +1cm
RT233-1.6.2 | ADJUSTABLE FOLDING FRAME, HYBRID WITH LONG WHEELBASE
Size range:
• 51 cm (20”) • 46 cm (18”) • 41 cm (16”) • 36 cm (14”)
Backrest / Seat configuration:
• Seat tilt: 8 degrees (approximately 8cm front and rear seat height differential on 40cm seat length). • Back to seat angle not to exceed 90 degrees. • No bend in back post. • Back height: 35cm-45cm range adjustable. • Seat length:  Standard = 40cm + 1cm long on all wheelchairs with 40cm seat frame.  Shorter length may be specified by end user.
Centre of gravity adjustability:
• Adjustable by adjusting rear wheel position horizontally at least three (3) positions.
Upholstery:
• Back upholstery: Tension adjustable back upholstery. 3-piece tension adjustable with major components and quality as specified under tension adjustable backrests in this technical specification. • Seat upholstery: Nylon (600 D or stronger) (40cm long or any shorter length as specified). All seat upholstery to be reinforced with an inner PVC minimum technical specification of 500 g/m².
Skirt guards:
• Optional or integrated (bidder to specify)
Castors and forks:
• 10cm x 6.5cm complete rubber unit. • Fitted with 37 x 12mm bearings. • Stem to be fastened to the fork with a thread and nut system to allow for re-fastening.
Rear wheels and axles:
• 24 x 1 ¾ “mag or spoke wheel with one-part semi-solid tyre or pneumatic. • Minimum of 6 rear wheel positions – safe, neutral, and active. • Axle 0.12cm thick.
Footrest hangers:
• Footrest hanger fitted to frame / removable with tools.
Footrest:
• Flip-up footplate to be hard heavy-duty plastic, or aluminium, or mild steel, and height adjustable.
Leg strap:
• Nylon strap (4cm wide) with hook and loop fasteners. The leg strap should have hook and loop fasteners fitted to its full length to allow maximum adjustability.
Guarantee:
• Frame and crossbars guaranteed for a minimum of 2 years from date of issue to end user.
Frame:
• Folding frame • Mild steel
Weight:
• Not to exceed 23kg (complete)
Weight of user:
• Visibly marked on each wheelchair to indicate maximum weight load of end user.
Repair / tool kit:
• 4 rear wheels: bearings • 4 castors: axle bearings • 4 castor stem bearings • 4 seat guides (where applicable) • Adjustment tool • Tension adjustable filler straps to allow for adjustment of back height.
General requirement:
• Bidder must be able to provide wheelchair with all abovementioned options. The end user will specify the desired option at the time of procuring. The different options available within each specification description must be fully interchangeable and available at the same price. Where the end user does not specify options the wheelchair will be issued as in the option given first for the categories with more than one option. • No part or component may be changed / deviate from the technical specifications set out in this document. • Bidder to submit accredited test report to ensure device durability
1.6.2. C | CUSTOMISING OPTIONS FOR ADJUSTABLE FOLDING FRAME HYBRID ACTIVE WHEELCHAIRS WITH LONG WHEELBASE
1.6.2. C.1 43cm and 45cm extended seat length
43cm and 45cm extended seat upholstery, with seat frames and screws complete. The extended seat is a complete unit, and a short strip extension of upholstery is not allowed. This option must fit onto item 1.6.2 and should be available in each size.
RT233-1.7 | RURAL ACTIVE CHAIRS – ADJUSTABLE 3-WHEELER | Indicate the item bidding for below bby stating "Yes bidding for this item". Item left blank will be assumed it is not bidded for | Complies Yes/No | Comments (Technical Specification Deviation)
Size range:
• 48 cm (18.9") wide • 44 cm (17.3") wide • 40 cm (15.7") wide • 36 cm (14.2") wide
Backrest / Seat configuration:
• Back rest angle adjustable: settings 86, 90 and 94 degrees (Backrest angle relative to seat angle of 10 degrees from ground). • No bend in back post. • Back height: Height adjustable backrest: to measure from solid seat Low back 38cm, Tall back 47cm. • Push handles attached to back posts. • Seat tilt: 10 degrees. • Seat length: To be specified by end user. 48cm (18.9") wide x 40, 44, 48 or 50cm long 44cm (17.3") wide x 40, 44, 48 or 50cm long 40cm (15.7") wide x 36, 40, 44, 48 or 50cm long 36cm (14.2") wide x 36, 40, 44, 48 or 50cm long
Centre of gravity adjustability:
• Adjustability of Centre of Gravity by positioning rear wheel in 2 positions within a 6,5cm range, i.e. safe / active.
Cushion:
• Pressure relief cushion contoured polyurethane foam with pre-ischial shelf. Cover fabric, waterproof, terry cloth cotton covered, 2-way stretch, breathable, anti-microbial shield treatment, and machine washable and can be dried in tumble drier and water resistant/ fire retardant nylon upholstery.
Upholstery:
• Back rest: Solid aluminium shell with shallow lateral walls, padded with open-cell, medium density foam of density 40 kg/m³, angular/height adjustment and adjustable sacral pad. • Backrest cover removable either a combination of heavy-duty lycra (220g/m² and yield 3.4m/kg) and 600D nylon or 200D nylon. • Seat: solid plywood or polyethylene (6/8/10mm thick appropriate to size) 36cm-48cm cut and setup to end-user technical specifications.
Arm rests:
• Removable arm rests.
Castors and forks:
• One rubber castor, size of 22 x 6cm.
Rear wheels and axles:
• 26 x 1 3/8" spoke wheel with one-part semi-solid tyre. • Quick release axles. • 2 degrees of camber
Footrest hangers:
• Integral part of frame.
Footrest:
• Fixed or flip-up, angle adjusted and height adjustable range of 14cm without modification. • Leg position and foot angle adjustment.
Leg strap:
• Heel, calf, and toe straps.
Frame:
• Rigid 3-wheeler. Minimum overall length with rear wheel in active position 120cm + 2cm. • Backrest - fold down. • The wheelbase is measured with castor fork pointing back to rear wheel from centre of castor to centre of rear wheel along the horizontal plane (as per diagram for basic folding frame wheelchair). Short wheelbase (active position) is 75cm+ 2cm. Long wheelbase (least active / safe) is 81cm+ 2cm. • Frame material – aluminium to be grade 6061 T4 with minimum tensile strength of 150 MPa and minimum yield of 115 MPa. Mild steel with minimum yield of 215 MPa and tensile strength of 340 MPa.
Guarantee:
• Frame and crossbars guaranteed for a minimum of 2 years from date of issue to end user.
Weight:
• Not to exceed 25kg, complete.
Weight of user:
• Visibly marked on each wheelchair to indicate maximum weight load of end user.
Repair / tool kit:
• 4 rear wheels: bearings • 4 castors: axle bearings • 4 castor stem bearings • 4 seat guides (where applicable) • Adjustment tool
General requirement:
• In the technical specifications where more than 1 option is provided: Bidder must be able to provide wheelchair with all of options as specified. The end user will specify the desired option at the time of ordering. The different options stated within each technical specification description must be fully interchangeable and available at the same price. Where the end user does not specify options the wheelchair will be issued as in the option listed first. • No part or component may be changed / deviate from the technical specifications set out in this document. • Bidder to submit accredited test report to ensure device durability
RT233-1.8 | ADJUSTABLE POSTURE SUPPORT WHEELCHAIRS | Indicate the item bidding for below bby stating "Yes bidding for this item". Item left blank will be assumed it is not bidded for | Complies Yes/No | Comments (Technical Specification Deviation)
RT233-1.8.1 | ADJUSTABLE POSTURE SUPPORT CHAIRS, SELF-PROPELLING, WITH MODULAR DETACHABLE SEATING SYSTEM (BUGGY TYPE)
Size range:
• 12” short Seat 30 cm wide x 25 cm - 32 cm adjustable depth • 12” standard Seat 30 cm wide x 30 cm - 40 cm adjustable depth • 14” short Seat 35 cm wide x 25 cm - 35 cm adjustable depth • 14” standard Seat 35 cm wide x 35 cm - 45 cm adjustable depth • 16” standard Seat 40 cm wide x 35 cm - 45 cm adjustable depth • 18” standard Seat 45 cm wide x 35 cm - 45 cm adjustable depth
Size Range:
The sizes below are maximum internal usable measurements of seat system and exclude postural seating supports but include the seat cushion.
• Seat width 17 cm – 23 cm, depth 18 cm – 28 cm, height 28 cm – 39 cm • Seat width 17 cm – 2,7 cm, depth 23 cm – 34 cm, height 35 cm – 45 cm • Seat width 22 cm – 28 cm, depth 29 cm – 38 cm, height 36 cm – 46 cm
Seat System:
• Removable modular adjustable seating system • Seating system to be removable/attachable without tools • Aluminium, wood, or plastic • Trunk support: Minimum length 100mm for smallest size / 180mm for largest size. Trunk support independently adjustable vertically and horizontally. Padded with a minimum of 6mm nonporous/ sealed cell foam. • Lined/padded with minimum of 6mm nonporous/ sealed cell foam • Removable washable waterproof covers. • Adjustable tilt in space up to 60 degrees • Recline - Seat to back angle adjustable in range 90 to 110 degrees. • Pelvic strap length adjustable, padded and at a 60–90-degree angle at the hips • Width of pelvic straps for 12” and smaller – 3 cm • Width of straps for 14” and bigger – 5 cm • Seating systems allows for adaptation, build-up and customising • Set of modular positioning support wedges / inserts • Seat cushion with / without abduction block as specified by end user: Top layer high density foam 2.5 - 5 cm thick; 36/20with hardness factor 18 – 21.6; Density 32.1-40 kg/m3). Base layer 2.5 cm high density compressed (reconstituted flexible PU) foam 1300; Hardness factor 52-63; Density 130 +10% kg/m3). • (SANS 642 – 1976). Waterproof 2-way stretch, breathable cover with anti-microbial shield (AEGIF) treatment • Seat length adjustable as specified • Back height adjustable as specified • Back support to accommodate thoracic spine or designed in such a way to accommodate user without causing any restriction / interference with self-propelling / posture
Seat and frame configuration:
• The chair will be modular. The seat system (combined seat and backrest with optional accessories such as headrest) will be separate from the base frame • The base frame to be rigid or folding • Infinite adjustable seat system tilt between 0 and minimum 30 degrees • Push handles minimum 93.5 cm from ground with 10 cm height adjustability. For smaller sizes – adult should comfortably reach push handle. Push handle may in no way restrict self-propelling / affect posture / set-up of back support system. • The design should allow for COG adjustment either as a specific wheel/seat adjustment or by means of a mechanism which will always result in COG being safely within base of support.
Front castors and forks:
• Seat width over 30 cm: Front castor wheel diameter range: 10 cm and 20 cm (+ 0,2 cm) • Seat width under 30 cm: Front castor wheel width range: 2,5 cm and 8 cm • End user to specify castor with and diameter • Stem to be fastened to the fork with thread and nut system to allow for re-fastening and adjusting.
Rear wheels and axles:
• Size 12” standard chair and larger: Minimum 24” x 1 3/8 diameter with one-part semi-solid tyre • Size 12” short chair and smaller: Minimum 20” x 1 3/8 diameter with one-part semi-solid tyre or foam insert • Optional pneumatic tyre • To be fitted with quick-release axle
Anti-tip wheels:
• Must be removable • Should allow for adjustability to match the level of independence and balance of the user. Range of adjustability to allow for stability in static position and to adjust for the required degree of stability in propulsion and if applicable, when front castors are raised for obstacle clearance.
Footrest:
• Height adjustable hanger with minimum 15 cm range (10 cm on 12” chair and smaller) to allow for variations in user profile and seat length adjustments • Must flip-up / fold up or must be detachable • Should allow for age-appropriate hip abduction • Will be constructed of rigid moisture and corrosion resistant material and must have a non-slip surface (e g. Rubber).
Lap Tray:
• Removable • Height adjustable • Basic, not padded. Vinyl covered • To support elbows
Guarantee:
• Frame and seat system guaranteed for a minimum of 2 years from date of issue to end user.
General requirement:
• In the technical specifications where more than 1 option is provided: Bidder must be able to provide wheelchair with all of options as specified. The end user will specify the desired option at the time of ordering. The different options stated within each technical specification description must be fully interchangeable and available at the same price. Where the end user does not specify options the wheelchair will be issued as in the option listed first. • No part or component may be changed / deviate from the technical specifications set out in this document. • Bidder to submit accredited test report to ensure device durability
1.8.1 C | CUSTOMISING OPTIONS FOR ADJUSTABLE POSTURE SUPPORT CHAIRS, SELF-PROPELLING, WITH MODULAR DETACHABLE SEATING SYSTEM
1.8.1 C.1 Tray Table Basic
Size Range: To fit posture chair in the above category
Description: • Easily removable multipurpose tray • Provides stable non-slip surface • When attached, the surface should provide adequate full resting support for the elbows • When attached, the curves around the body should be cut in such a way as to prevent a large opening between the body and the tray where the elbow, forearms or hands can slip through.
Thickness: Between 0.6 cm and 1.5cm Material: Wood, plastic with washable vinyl top Quality: Smooth, washable, splinter free surface
1.8.1 C.2 Tray Table Padded
Size Range: To fit posture chair in the above category Description: • Easily removable multipurpose tray • Provides stable, dynamic, non-slip surface • When attached, the surface should provide full resting support for the elbows • When attached, the curves around the body should be cut in such a way as to prevent a large opening between the body and the tray where the elbow, forearms or hands can slip through. Thickness: Between 2.0cm and 3.0cm Material: Wood or plastic inner, 2.5mm high density compressed (reconstituted flexible PU) foam with washable stretch vinyl top – fully upholstered
1.8.1 C.3 Head support, adjustable, basic occipital support only
• Minimum 2 sizes, small and standard • Easily removable • Height and depth adjustable • Kidney shaped head support flexible and rake adjustable • Moulded and padded • Washable
1.8.1 C.4 Head support, adjustable, occipital and side support
• Minimum 2 sizes, small and standard • Easily removable • Height and depth adjustable • Flexible and rake adjustable • Moulded and padded • Washable
1.8.1 C.5 Chest Strap
• Poly-cotton webbing, 10 cm • Two separate sections, adjustable, with 2 x 50mm D-ring (chromium plated, galvanised steel) male and female Velcro Loop Velcro immediately adjacent to hook Velcro • Padding) attached to D-ring section, , extending beyond the D-ring and width of strap
1.8.1 C 6 Knee straps with attachments for modular seating system
• Poly-cotton webbing, • One section • Length adjustable. • 2 x 5cm D-ring steel, epoxy coated to be attached to the seat mould • Adjustable webbing sections with female hook and loop fasteners to be 20cm long. • Hook 30cm long on both ends, with loop 30cm long immediately adjacent to hook loop. Padding attached to contact areas of webbing.
1.8.1 C 7 Pelvic straps with attachments for modular seating system
• Poly-cotton webbing, 5cm wide. • Two separate sections, one non-adjustable, other section adjustable. • Non-adjustable section fitted with 5cm D-ring (chromium plated, galvanise steel) and other with male and female hook and loop fasteners. • Size specific (12” and smaller = 3cm,14” and larger 5cm) - to fit seating systems above hook on one end, with loop immediately adjacent to hook. Padding attached to D-ring section, 6cm wide, 50cm long of which half extending beyond the D-ring
1.8.1 C 8 Footbox raiser for modular seating system.
• To clip onto the existing footrest of the above seating system. Raise the back and side height by approximately 10cm, preventing the feet kicking out of the footrest. Provide firm scuff resistant padding on back and side surfaces.
1.8.1 C 9 Adduction boards for modular seating system (specify size)
1.8.1 C.10 Mobile (wheeled) compact base for seating system from range 1.8.2
• Incorporating wheels and/or castors on the four corners. Castors should be lockable or a minimum of 2 wheels will be supplied with locks or brakes. • COG adjustable to safely accommodate recline and tilt options
1.8.1 C.11 Motorised base frame only - Front wheel or rear wheel drive with electronic hand control system for use with removable posture support system
Size options: To fit chairs:
• Seat width 17 cm – 23 cm, depth 18 cm – 28 cm, height 28 cm – 39 cm • Seat width 17 cm – 2,7 cm, depth 23 cm – 34 cm, height 35 cm – 45 cm • Seat width 22 cm – 28 cm, depth 29 cm – 38 cm, height 36 cm – 46 cm
Frame:
• Accommodate COG adjustment • Anti-tip levers • Frame material - mild steel to be of a minimum yield of 375 MPA and a minimum tensile strength of 400 MPA. • Finish: Epoxy coated
Castor and forks:
• 8” x 2” (20 cm x 5 cm) castor with plastic hub. • Cast Aluminium fork, with top 1 cm and sides tapering to 0.4 cm. • Forks to be attached to the castor stem by means of bolt system to allow for re-fastening of the stem.
Drive wheels and axles:
• Minimum 12”x2” diameter tyres with deep treads. • Can be changed manually from drive to free wheel through Dog clutch or Gear box release mechanism
Size range:
• To fit posture chair in the above category
Description:
• Easily removable multi-purpose tray • When attached, the extensions should provide full resting support for the elbows. • When attached, the curve around the body should be cut to prevent a large opening between the body and the tray where the elbows, forearms or hands can slip through. Material: 0.8cm-1.5cm see through polycarbonate or acrylic type sheet Quality: Smooth, washable, non-breakable, transparent
Controller:
• Electronic sensitivity – programmable speed and sensitivity control system • Position adjustable vertical and horizontal • Left- or right-hand control as specified by end user
Batteries:
• 2 x 12-volt sealed lead acid batteries • Deep cycle 38 Amp hour sealed batteries and dedicated charger
Motors:
• DC motors 24 volt, 240- 320 peak Watt, • Gear ratio 1:32
Speed:
• Top speed 6 km/h
Range:
• Average range 13 – 16 km
Guarantee:
• Frame and crossbars guaranteed for a minimum of 2 years from date of issue to end user. • Electronics and motors guaranteed for a minimum of 1 year from issue during appropriate use as stated per category
Repair/ tool kit:
• 4 Rear wheel: bearings • 4 Castors axle bearings • 4 castor stem bearings • Adjustment tools
Tray table –See through type- Supplied standard with the powerbase
• Tray tables Removable- wood with vinyl (not padded), height and depth adjustable, to fit seating support system • Tray tables Removable- wood with vinyl (padded), height and depth adjustable, to fit seating support system • Tray tables Removable- perspex , height and depth adjustable, to fit seating support system
RT233-1.8.2 | PEDIATRIC SEATING SYSTEM: SELF PROPELLING SELF-PROPELLING PAEDIATRIC WHEELCHAIR WITH ADJUSTABLE POSTURE SUPPORT BACKREST
Size range:
• 45cm (18”) wide with adjustable seat frame length 35 – 45cm • 41cm (16”) wide with adjustable seat frame length 35 – 45cm • 36 cm (14”) wide with adjustable seat frame length 35 – 45cm • 36cm (14”) wide with adjustable seat frame length 25 – 35cm • 30cm (12”) wide with adjustable seat frame length 30 – 40cm • 30cm (12”) wide with adjustable seat frame length 25 – 35cm • 25cm (10”) wide with adjustable seat frame length 22 – 35cm • 20cm (8”) wide with adjustable seat frame length 17 – 25cm
Frame:
• Rigid, mild steel or aluminium
Backrest/seat configuration:
• Seat tilt: Adjustable range up to 20 degrees • Back post recline: Adjustable range up to 35 degrees • Seat/backrest configuration is detachable from the wheels/frame to allow for transportability
Upholstery:
• All upholstery should be washable and removable
Seat cushion:
• Pressure relief cushion with Dual layer foam construction, high density foam top layer of 5cm thickness and high density compressed foam bottom layer of 5cm thickness
Armrests:
• Removable or detachable • Shape of armrests should in no way impede the fitment of commercial rigid adjustable backrest systems • Optionally integrated into tray table support
Castors and forks:
• 5cm minimum width of castors
Rear wheels and axles:
• 24” x 1 3/8 spoke with one-part semi-solid tyre on 18” wide chair up to 12” 30-40cm wide chairs • 22” x 1 /38 spoke wheels with one part semi-solid tyre on 8” wide chair up to 12” 25-35cm wide chairs
Footrest:
• Integrated into the backrest seat configuration or frame • Height adjustable appropriate to width of chair • Flip-up or detachable • Footplate to be hard heavy-duty plastic, or aluminium, or mild steel
Pelvic strap:
• Pelvic strap should be 2.5-4cm in width with padding, D-ring loop and hook and loop fasteners • Adjustable for length
Guarantee:
• Frame and seat system guaranteed for a minimum of 2 years from date of issue to end user.
Weight of user:
• Maximum user weight should not exceed 70kg, appropriate to the size of the chair
Optional Extras
• Head Support • Complex head support height and depth and rake adjustable, to fit modular seating support system, flexible, moulded and padded
Tray tables
• Tray tables Removable- wood with vinyl (not padded), height and depth adjustable, to fit seating support system • Tray tables Removable- wood with vinyl (padded), height and depth adjustable, to fit seating support system • Tray tables Removable- perspex , height and depth adjustable, to fit seating support system
General requirement:
• In the technical specifications where more than 1 option is provided: Bidder must be able to provide wheelchair with all of options as specified. The end user will specify the desired option at the time of ordering. The different options stated within each technical specification description must be fully interchangeable and available at the same price. Where the end user does not specify options the wheelchair will be issued as in the option listed first. • No part or component may be changed / deviate from the technical specifications set out in this document. • Bidder to submit accredited test report to ensure device durability
RT233-1.9 | MOTORISED FOLDING FRAME WHEELCHAIRS | Indicate the item bidding for below bby stating "Yes bidding for this item". Item left blank will be assumed it is not bidded for | Complies Yes/No | Comments (Technical Specification Deviation)
RT233-1.9.1 | MOTORISED WHEELCHAIRS, FOLDING FRAME, STANDARD BACK HEIGHT WITH HAND CONTROL
RT233-1.9.2 | MOTORISED WHEELCHAIRS, FOLDING FRAME, TALL BACK HEIGHT WITH HAND CONTROL
RT233-1.9.3 | MOTORISED WHEELCHAIRS, ADJUSTABLE FOLDING FRAME, WITH ADJUSTABLE TILT IN SPACE WITH HAND CONTROL
Size range for standard (1.10.1) and tall (1.10.2) back height motorised wheelchairs
• 51 cm (20") wide • 46 cm (18") wide • 43 cm (17”) wide • 41 cm (16") wide • 38 cm (15") wide • 36 cm (14") wide • 34 cm (13") wide with 40 cm seat frame length • 34 cm (13") wide with 30 cm seat frame length • 30 cm (12”) wide with 30 cm seat frame length
Backrest / Seat configuration:
• Seat tilt: 5 degrees minimum and 8 degrees maximum (approximate minimum of 3.5cm and maximum of 5.7cm front and rear seat height differential on 40cm seat length). • Back to seat angle not to exceed 90 degrees. • 8-10 degrees bend in back post. • Standard back height: 42.5cm ± 2.5cm or max 45cm on all. Exception - 30cm (12") wide with 30cm seat frame length which is 32.5cm ± 2.5cm or maximum 35cm. • Tall back height: 52.5cm ± 2.5cm or maximum 55cm on all. Exception - 30cm (12") wide with 30cm seat frame length which is 42.5cm ± 2.5cm or maximum 45cm. • Seat length: standard = 40cm + 1cm long on all wheelchairs with 40cm seat frame. Shorter length may be specified by end user. Exception - 30cm + 1cm long on 30cm (12") wide with 30cm seat frame length.
Upholstery:
• Back upholstery: Tension adjustable back upholstery. 3-piece tension adjustable as specified under tension adjustable backrests in this technical specification. • Seat upholstery: Nylon (600 D or stronger) or vinyl (40cm long or any shorter length as specified). All back and seat upholstery to be reinforced with an inner PVC minimum technical specification of 500 g/m². (Re-enforced wheelchairs as specified will therefore have a second, outer layer of PVC for additional reinforcement. Reinforcement by means of strips not acceptable).
Armrests:
• Removable, height adjustable desk or full armrest mild steel/ durable plastic skirt guard. • Arm rest sockets to be on side of seat rail only. • Shape of arm rests may in no way impede the fitting of commercial rigid adjustable back systems onto back posts.
Castors and forks:
• 8x2" (20cm x 5cm) castors with aluminium hub, and outer shore hardness of 65-70 and inner shore hardness of 33-70, with the inner value not exceeding the outer value. Alternative 8x2” (20cm x 5cm) rubber castors. • Forks to be minimum 0.4cm thick mild steel forks. Stem to be fastened to the fork with a thread and nut system to allow for re-fastening.
Rear wheels and axles:
• 31cm x 5cm (12" x 2") Semi solid tyres with deep treads. Can be changed manually from drive to free wheel through dog clutch or gear box release mechanism.
Footrest hangers:
• Extra-short (15cm), short (20cm), long (25cm) or fully detachable 90-degree hangers with angle adjustable footplates as specified by end user.
Footrest:
• Footplate to be hard heavy-duty plastic, or aluminium, or mild steel, and height adjustable. Swing away
Leg strap:
• Nylon strap (4cm-5cm wide) with hook and loop fasteners. The leg strap should have hook and loop fasteners fitted to its full length to allow maximum adjustability.
Controller:
• Adjustable vertical and horizontal position. • Chin left or right-hand control as specified by end user (the chin control in this instance will be a hand controller positioned at the chin).
Battery tray:
• Removable. • Batteries x 12 volt sealed lead acid batteries x 36 AH batteries (minimum requirement). • Motor strength 2 x 280 – 350 Watt. • Charger included standard with surge protector. • Brakes. • Fully automatic magnetic safety brakes.
Speed:
• Top speed 6 km/h.
Range:
• Average range 13-16 km. • Grade climbing up to12 degrees.
Guarantee:
• Frame and crossbars guaranteed for a minimum 2 year guarantee from date of issue to end user. • Electronics and motors guaranteed for a minimum of 1 year from issue during appropriate use as stated per category
Weight:
Not to exceed 70kg (complete). Allow +1kg with rubber castors.
Weight of user:
• Visibly marked on each wheelchair to indicate maximum weight load of end user.
Repair / tool kit:
• 4 Rear wheel: bearings • 4 Castors: axle bearings • 4 castor stem bearings • 4 seat guides (where applicable) • Adjustment tools
General requirement:
• Bidder must be able to provide wheelchair with all abovementioned options. The end user will specify the desired option at the time of procuring. The different options available within each technical specification description must be fully interchangeable and available at the same price. Where the end user does not specify options the wheelchair will be issued as in the option given first for the categories with more than one option. • No part or component may be changed / deviate from the technical specifications set out in this document. • Bidder to submit accredited test report to ensure device durability
Frame:
• Basic, non-adjustable folding frame. • Anti-tip levers. • Frame material - mild steel to be of a minimum yield of 375 MPa and a minimum tensile strength of 400 MPa. • Frame to be zinc phosphate rust protected prior to powder coating. • Finish: Epoxy coated. • Back rest metal tubing wall thickness to be 0.16cm. • X-bars to be constructed of tubing with a minimum of 2.5cm diameter and 0.2cm wall thickness.
RT233-1.9.4 | MOTORISED WHEELCHAIRS, ADJUSTABLE FOLDING FRAME WITH ADJUSTABLE TILT IN SPACE AND ADJUSTABLE BACKREST HEIGHT
Size range:
• 51 cm (20") wide with 43 + 1 cm seat frame length • 46 cm (18") wide with 43 + 1 cm seat frame length • 43 cm (17”) wide with 43 + 1 cm seat frame length • 41 cm (16") wide with 43 + 1 cm seat frame length • 38 cm (15") wide with 43 + 1 cm seat frame length • 36 cm (14") wide with 43 + 1 cm seat frame length • 30 cm (12") wide with 43 + 1 cm seat frame length • 30 cm (12") wide with 33 + 1 cm seat frame length • 25 cm (10") wide with 33 + 1 cm seat frame length
Backrest / Seat configuration:
• Seat tilt: Adjustable, secondary to centre of gravity adjustment of wheelchair. • Back to seat angle not to exceed 90 degrees. • Angle in back post to be 8-10 degrees. • Back height: Adjustable 35cm-45cm, 40cm-50cm and 45cm-55cm. The back range of the 2 smallest sizes is 35cm-45cm only and has the 8-10 degrees bend in the back post. • Seat length:  Standard = 40cm + 1cm long as standard, with option to be fitted with 44cm+ 1cm seat on 43cm seat rail.  For the 12” with 33cm seat rail, the standard = 30cm + 1cm long as standard, with option to be fitted with 33cm + 1cm seat.  For the 10“ with 33cm seat rail, the standard = 25cm + 1cm long as standard. Shorter length may be specified by end user.
Centre of gravity adjustability:
• Adjustable by adjusting rear wheel position horizontally and vertical (at least 3 settings each way). • Adjustable by adjusting rear wheel - at least 3 settings for vertical and horizontal. • The adjustment range of the front castor should ensure that the castor stem always remains perpendicular to the ground within the available tilt in space range.
Upholstery:
• Back upholstery: Tension adjustable back upholstery. 3-piece tension adjustable with major components and quality as specified under tension adjustable backrests in this technical specification. • Seat upholstery: Nylon (600 D or stronger) (40cm long or any shorter length as specified). All seat upholstery to be reinforced with an inner PVC minimum technical specification of 500 g/m².
Armrests:
• Removable height adjustable desk or full armrests • mild steel/ durable plastic skirt guard. • Arm rest sockets to be on side of seat rail only. • Shape of arm rests may in no way impede the fitting of commercial rigid adjustable back systems onto back posts.
Castors and forks:
• 8x1" (20cm x 2.5cm) or 8x2" (5cm x 20cm) castors as specified by end user. • 8x2" (20cm x 5cm) castors with plastic hub, outer shore hardness of 65-70 and inner shore hardness of 33-70, with the inner value not exceeding the outer value. • Matching adjustments of the front castor to equal any adjustments done to the rear wheel position. • Forks to be minimum 0.4cm thick mild steel forks. Stem to be fastened to the fork with a thread and nut system to allow for re-fastening.
Rear wheels and axles:
• 31cm x 5cm (12" x 2") Semi solid tyres with deep treads. Can be changed manually from drive to free wheel through dog clutch or gear box release mechanism. • Vertical and horizontal adjustments of rear wheel (at least 3 settings each way).
Footrest hangers:
• Extra-short (15cm), short (20cm), long (25cm) or fully detachable 90-degree hangers with angle adjustable footplates as specified by end user. Swing away
Footrest:
• Footplate to be hard heavy-duty plastic, or aluminium, or mild steel, and height adjustable. Flip up
Leg strap:
• Nylon strap (4cm-5cm wide) with hook and loop fasteners. The leg strap should have hook and loop fasteners fitted to its full length to allow maximum adjustability.
Controller:
• Adjustable vertical and horizontal position. • Chin left or right-hand control as specified by end user (the chin control in this instance will be a hand controller positioned at the chin).
Battery tray:
• Removable. • Batteries x 12 volt sealed lead acid batteries x 36 AH batteries (minimum requirement). • Motor strength 2 x 280-350 Watts. • Charger included standard with surge protector.
Brakes:
• Fully automatic magnetic safety brakes.
Speed:
• Top speed 6 km/h.
Range:
• Average range 13-16 km. • Grade climbing up to12 degrees.
Guarantee:
• Frame and crossbars guaranteed for a minimum 2 year guarantee from date of issue to end user. • Electronics and motors guaranteed for a minimum of 1 year from issue during appropriate use as stated per category
Frame:
• Adjustable, folding frame. • Aluminium / mild steel. • 30cm (12") wide with 33 + 1cm seat frame length and 25cm (10") wide with 33cm + 1cm seat frame length has shorter wheelbase and overall length as the other sizes in this range.
Weight:
• Not to exceed 70kg (complete), Allow +1kg with rubber castors.
Weight of user:
• Visibly marked on each wheelchair to indicate maximum weight load of end user.
Repair / tool kit:
• 4 rear wheels: bearings • 4 castors: axle bearings • 4 castor stem bearings • 4 seat guides (where applicable) • Adjustment tool
General requirement:
• Bidder must be able to provide wheelchair with all abovementioned options. The end user will specify the desired option at the time of ordering. The different options available within each technical specification description must be fully interchangeable and available at the same price. Where the end user does not specify options the wheelchair will be issued as in the option given first for the categories with more than one option. • No part or component may be changed / deviate from the technical specifications set out in this document. • Bidder to submit accredited test report to ensure device durability
1.9. C. | CUSTOMISING OPTIONS FOR MOTORISED WHEELCHAIRS
1.9. C.1 Dedicated chin controller:
Dedicated chin controller. On chin mount with adjustable vertical and horizontal position. This option must fit onto item of 1.10.
1.9.C.2. Option to accommodate extra thigh length using any one the following:
1.9.C.2.1 For sizes 12” with standard frame and 14” and larger: Seat depth (extension):
43cm +1cm; 45cm +1cm ; 47cm+1cm ;49 cm +1. For sizes: 12” with 33cm seat frame: 33cm +1cm seat extension For sizes: 10” with 33cm seat frame : 30cm +1cm seat extension
1.9.C.2.2 Long wheelchair frame
1.9.C.2.3 Frame with seat depth adjustability range. Minimum 40cm-50cm. End user to specify option and seat depth and frame length.
1.9. C.3 Bolt on footrest (pair)
Small footplate (15 x 13cm) without hanger, with attachment to be fitted directly onto side frame of wheelchair. This option must fit onto items 1.9 and 1.10.
1.9.C.4 Adjustable folding frame with an adjustable recline back post
Adjustable back to seat angle 90⁰-105⁰. Adjustment mechanism adjustable with minimum 3 adjustment increments of 5⁰((0⁰, -5⁰, -10⁰, -15⁰)
1.9. C.5 Elevating footrest hanger (pair)
Angle of footrest hanger to be adjustable between 0 degrees and 70 degrees of flexion; Height adjustable calf support; Fitted with adult, narrow or juvenile footplates as specified by end user. To be ordered as a pair. This option must fit onto items 1.9 and 1.10.
1.9. C.6 Elevating footrest hanger (individual)
Angle of footrest hanger to be adjustable between 0 degrees and 70 degrees of knee flexion; Height adjustable calf support; Fitted with adult, narrow or juvenile footplates as specified by end user. To be ordered as an individual item. This option must fit onto items 1.9 and 1.10.
1.9. C.7 Head support, adjustable, basic occipital support only
Easily removable Height and depth adjustable Kidney shaped head support flexible and rake adjustable Padded and covered in heavy gauge, removable, washable fabric. This option must be compatible with option 1.9.3.C.9 and also fit on item 1.10
1.9.3. C.8 Head support, complex adjustable (hand shaped)
Easily removable Height and depth adjustable 4-point, hand shaped head support, flexible and rake adjustable Moulded and padded (EVA) This option must be compatible with option 1.9.3.C.9 and also fit on item 1.10
1.9.3. C.9 Removable universal head rest with horizontal mounting bar
Universal bar that attaches in a horizontal position onto the back post/handles of the wheelchair to be fitted with a removable head rest. Mounting to the wheelchair: Quick release mechanism to allow folding of wheelchair. Secure attachment to ensure safety when used as push handle. Mount for head rest: Adjustable mount to allow 4cm +- 1cm offset left and right from central position. Should be adjustable. This option must fit on item of 1.10.
1.9.3 C.10 Footrest with 90 degrees hangers
Footplate of footrests with 90 degrees hangers only in aluminium/mild steel or hard heavy-duty plastic, must be height/ angle/ forward/ backward adjustable.
1.9.3 C.11 Adjustable backrest height ranges
35 – 45 cm (straight back posts), 40 – 50 cm (straight back posts), 45 – 55 cm (straight back posts)
RT233-1.10 | MOTORISED HEAVY-DUTY RIGID FRAME WHEELCHAIRS | Indicate the item bidding for below bby stating "Yes bidding for this item". Item left blank will be assumed it is not bidded for | Complies Yes/No | Comments (Technical Specification Deviation)
Size range:
• 51 cm (20") standard, Seat 51 cm wide x 40 cm – 50 cm deep • 46 cm (18") standard, Seat 46 cm wide x 40 cm – 50 cm deep • 43 cm (17") standard, Seat 41 cm wide x 40 cm – 50 cm deep • 41 cm (16") standard, Seat 41 cm wide x 40 cm – 50 cm deep • 36 cm (14") standard, Seat 36 cm wide x 40 cm – 50 cm deep
Backrest / Seat configuration:
• Seat system tilt (back and seat unit combined): Manual, adjustable tilt in space. • Seat system tilt in space infinitely adjustable between 0 and minimum 20 degrees. • Centre of gravity settings a range of 4x2 adjustments (8 total settings). • Seat back angle at pelvis not to exceed 90 degrees. • Adjustable seat height ranges from 45cm-50cm, measured at rear of seat at 0-degree tilt. • Backrest height, measured from rear of seat, 45cm-65cm adjustable range. Push handle may in no way affect posture / set-up of back support system. • Minimum 10-degree bend in back post to accommodate thoracic spine or designed in such a way to accommodate the back system at the required degree of recline without causing any restriction / interference with posture. • Seat length adjustable forward by a minimum of 15cm from minimum length. • Strap: Pelvic, angle at 45 degrees over hips
Upholstery:
• Back upholstery: Tension adjustable back upholstery. 3-piece tension adjustable with major components and quality as specified under tension adjustable backrests in this technical specification. • Seat upholstery:  Seat base material to be moisture and corrosion resistant and must provide rigid support.  Length adjustable as specified under size range above.
Armrests:
• Removable full or desk with mild steel skirt guard. • Shape of arm rests may in no way impede the fitting of commercial rigid adjustable back systems onto back posts.
Castors and forks
• 9” x3" (23cm x 7.5cm) or 8x2" (20cm x 5cm) castors with aluminium hub, and outer shore hardness of 65-70 and inner shore hardness of 33-70, with the inner value not exceeding the outer value. Castors as specified by end user. • Matching adjustments of the front castor to equal any adjustments done to the rear wheel position. • Forks to be minimum 0.4cm thick mild steel forks. Castor stems need to be between 1.5cm-1.7cm. Stem to be fastened to the fork with a thread and nut system to allow for re-fastening.
Rear wheels and axles:
• 11”-14.5” diameter x 3.5” width, solid or pneumatic tyres. Can be changed manually from drive to free wheel through dog clutch or gear box release mechanism.
Anti-tip levers:
• Removable or flip up and adjustable in length 10cm x 3.5cm (4”x 1.4”).
Footrest hangers:
• Extra-short (15cm), short (20cm), long (25cm) or fully detachable 90-degree hangers with angle adjustable footplates as specified by end user.
Footrest:
• Footplate to be hard heavy-duty plastic, or aluminium, or mild steel, and height adjustable.
Leg strap:
• Nylon strap (4cm-5cm wide) with hook and loop fasteners. The leg strap should have hook and loop fasteners fitted to its full length to allow maximum adjustability.
Controller:
• Adjustable vertical and horizontal position. • Chin left or right-hand control as specified by end user.
Battery tray:
• Removable
Batteries:
• x 12 volt sealed lead acid batteries • Minimum capacity 2 x 48 AH batteries (minimum requirement) • Minimum Motor strength 2 x 450-700 Watts • Minimum amperage 50+ Amp • Charger included 6 Amp standard with surge protector
Brakes:
• Fully automatic magnetic safety brakes.
Speed:
• Average speed 6+ km/h
Range:
• Average range 20+ km • Grade climbing up to 15 degrees
Guarantee:
• Frame and crossbars guaranteed for a minimum 2 year guarantee from date of issue to end user. • Electronics and motors guaranteed for a minimum of 1 year from issue during appropriate use as stated per category
Frame:
• Rigid frame. • Seat system tilt in space infinitely adjustable between 0 and minimum 20 degrees (Seat system to tilt as unit). • Tilt and recline mechanisms may not interfere with the fitting of postural support systems. • Aluminium / mild steel. • Wheelbase is measured with castor fork pointing back to rear wheel from centre of castor to centre of rear wheel along the horizontal plane 46cm.
Weight:
• With battery: not to exceed 120kg. • Without battery: not to exceed 90kg.
Weight of user:
• Visibly marked on each wheelchair to indicate maximum weight load of end user.
Repair / tool kit:
• 4 rear wheels: bearings • 4 castors: axle bearings • 4 castor stem bearings • 4 seat guides (where applicable) • Adjustment tool
General requirement:
• Bidder must be able to provide wheelchair with all abovementioned options. The end user will specify the desired option at the time of ordering. The different options available within each technical specification description must be fully interchangeable and available at the same price. Where the end user does not specify options the wheelchair will be issued as in the option given first for the categories with more than one option. • No part or component may be changed / deviate from the technical specifications set out in this document. • Bidder to submit accredited test report to ensure device durability
1.10. C | CUSTOMISING OPTIONS MOTORISED HEAVY-DUTY RIGID FRAME WHEELCHAIRS
1.10. C.1 Dedicated chin controller:
Dedicated chin controller. On chin mount with adjustable vertical and horizontal position. This option must fit onto item 1.10.
1.10. C.2 45cm and 47.5cm seats extension
45cm and 47.5cm seats, with seat frame extension kit and full-length extended seat upholstery (short strip of upholstery not allowed) to lengthen original seat frames and seat. Upholstery screws to be provided. This option must fit onto item 1.10 and should be available in each size.
1.10. C.3 Bolt on footrest (pair)
Small footplate (15cm x 13cm) without hanger, with attachment to be fitted directly onto side frame of wheelchair. This option must fit onto item 1.10.
1.10. C.4 Elevating footrest hanger (pair)
Angle of footrest hanger to be adjustable between 0 degrees and 70 degrees of flexion; Height adjustable calf support; Fitted with adult, narrow or juvenile footplates as specified by end user. To be ordered as a pair. This option must fit onto item 1.10.
1.10. C.5 Elevating footrest hanger (individual)
Angle of footrest hanger to be adjustable between 0 degrees and 70 degrees of knee flexion; Height adjustable calf support; Fitted with adult, narrow or juvenile footplates as specified by end user. To be ordered as an individual item. This option must fit onto item 1.10.
1.10. C.6 Footrest with 90 degrees hangers
Footplate of footrests with 90 degrees hangers only in aluminium/mild steel or hard heavy-duty plastic, must be height/ angle/ forward/ backward adjustable.
1.10. C.7 Head support, adjustable, basic occipital support only
Easily removable Height and depth adjustable Kidney shaped head support flexible and rake adjustable Padded and covered in heavy gauge, removable, washable fabric. This option must fit onto item 1.10.
1.10. C.8 Electric tilt with range from 0-45 degrees
This option must fit onto item 1.10.
RT233-1.11 | BUGGIES | Indicate the item bidding for below bby stating "Yes bidding for this item". Item left blank will be assumed it is not bidded for | Complies Yes/No | Comments (Technical Specification Deviation)
RT233-1.11.1 | MODULAR, RIGID BUCK, FOAM ADJUSTABLE POSTURE SUPPORT BUGGY WITH RIGID FRAME
Size range:
The sizes below are maximum internal usable measurements of seat mould and exclude postural seating supports but include seat cushion.
• Seat width 25cm, depth 24cm, height 40cm • Seat width 27cm, depth 24cm, height 40cm • Seat width 29cm, depth 30cm, height 40cm • Seat width 32cm, depth 36cm, height 49cm • Seat width 32cm, depth 45cm, height 54cm • Seat width 39cm, depth 49cm, height 55cm
All heights increase when optional head extension is utilised.
Seat system:
• Removable fibreglass seating system, padded with 1.2cm high density foam, 1 seat system cover, removable, heavy duty denim twill, machine washable and can be dried in tumble drier. • Seat back angle not to exceed 98 degrees. • Tilt in space adjustable up to 30 degrees. • Pelvic strap length adjustable (both origins and middle), padded. Widths of straps are size appropriate (12” and smaller= 3cm,14” and larger 5cm). • Modular seating support inserts covered with removable, heavy duty denim twill, machine washable and can be dried in tumble drier. High density (26/20) with hardness factor 18.0-21.6 and density range 24.1-27. • Seating cushion cover to be waterproof fabric, terry cloth cotton covered, 2-way stretch, breathable, anti-microbial shield treated • Seat cushion with / without abduction block as specified by end user: Top layer high density foam 5cm thick; 36/20 with hardness factor 18-21.6; Density 32.1-40 kg/m³). Base layer 2.5cm high density compressed (reconstituted flexible Poly Utherane (PU)) foam 1300; Hardness factor 52-63; Density 130+10% kg/m³). (SANS 642-1976). Waterproof, terry cloth cotton covered, 2-way stretch, breathable, anti-microbial shield treatment
Seating support:
• 1 x 5cm back support (exception: baby buggy nil 5cm back supports) • 2 x 2,5cm back support • 1 x seat cushion with / without abduction block as specified by end user • 1 x head support cushion • 2 x trunk side supports • 1 seat system cover, removable and washable
Lap tray:
• Easily removable multi-purpose tray and provides stable surface. When attached, the surface extensions should provide support for the elbows. • When attached, the curve around the body should be cut in such a way to prevent a large opening between the body and the tray where the elbows, forearms or hands can slip through. • Basic, not padded, vinyl covered. • To support elbows. • Height, depth, and angle adjustable independently from tilt in space.
Material:
• Wood (Supawood covered/protected with sealer and wood varnish)/ plastic with washable vinyl top.
Quality:
• Smooth, washable, splinter free surface, splinter proof, rounded edges, easily wipe-able and cleanable.
Castors and forks:
• 8” x 2” (20cm x 5cm) castors with plastic hub. • Cast Aluminium fork, with top 1cm and sides tapering to 0.4cm. • Forks to be attached to the castor stem by means of a thread and bolt system to allow for re-fastening of the stem.
Rear wheels and axles:
• Minimum 20cm x 8cm
Footrests:
• Height adjustable, detachable • Solid with lateral and side supports • Footplate to be hard heavy-duty plastic, or aluminium, or mild steel, and height adjustable
Frame:
• Mild steel, epoxy coated • Size folded: Maximum 60cm x 45cm x 50cm
Weight:
• Not to exceed 25kg for largest size
Weight of user:
• Visibly marked on each wheelchair to indicate maximum weight load of end user.
Guarantee:
• Frame and seat system guaranteed for a minimum 2 year guarantee from date of issue to end user.
General requirement:
• In the technical specifications where more than 1 option is provided: Bidder must be able to provide wheelchair with all of options as specified. The end user will specify the desired option at the time of ordering. The different options stated within each technical specification description must be fully interchangeable and available at the same price. Where the end user does not specify options the wheelchair will be issued as in the option listed first. • No part or component may be changed / deviate from the technical specifications set out in this document. • Bidder to submit accredited test report to ensure device durability
1.11.1 C | CUSTOMISING OPTIONS FOR MODULAR, RIGID BUCK, FOAM ADJUSTABLE POSTURE SUPPORT BUGGY WITH RIGID FRAME
1.11.1. C.1 Tray table basic
Size range: To fit the modular, rigid buck, foam adjustable posture support buggy with rigid frame size range
Description: • Easily removable multi-purpose tray. • Provides stable non- slip surface. • When attached, the surface extensions should provide adequate full resting support for the elbows. • When attached, the curve around the body should be cut in such a way as to prevent a large opening between the body and the tray where the elbows, forearms or hands can slip through.
Thickness: Between 0.6cm and 1.5cm
Material: Wood (Supawood covered/protected with sealer and wood varnish)/ plastic with washable vinyl top.
Quality: Smooth, washable, splinter free surface, splinter proof, rounded edges, easily wipe-able and cleanable
1.11.1. C.2 Tray table padded
Size range: To fit the modular, rigid buck, foam adjustable posture support buggy with rigid frame size range.
Description: • Easily removable multi-purpose tray. • Provides stable, dynamic, non- slip surface. • When attached, the extensions should provide full resting support for the elbows. • When attached, the curve around the body should be cut to prevent a large opening between the body and the tray where the elbows, forearms or hands can slip through. • Height, depth, and angle adjustable independently from tilt in space.
Thickness: Between 2.0cm and 4.0cm
Material: Wood (Supawood covered/protected with sealer and wood varnish) or plastic inner, 2.5cm high density compressed (reconstituted flexible PU) foam with washable stretch vinyl top – fully upholstered, waterproof, material needs to be durable for regular cleaning, and flame retardant and heat resistant.
1.11.1. C.3 Head support, adjustable, basic occipital support only
Easily removable Height and depth adjustable Kidney shaped head support flexible and rake adjustable Padded and covered in heavy gauge, removable, washable fabric
1.11.1. C.4 Head support, adjustable, basic occipital, and side support
Easily removable Height and depth adjustable 4-point, hand shaped head support, flexible and rake adjustable Moulded and padded EVA
1.11.1. C.5 Chest strap
Poly-cotton webbing, 6cm or 10cm Two separate sections, adjustable, with 2 x 3cm or 5cm D-ring (chromium plated, galvanised steel) male and female loop immediately adjacent to hook Padding) attached to D-ring section, extending beyond the D-ring and width strap
1.11.1. C.6 Thigh strap
Poly-cotton webbing, One section Length adjustable. 2 x 5cm D-ring steel, epoxy coated to be attached to the seat mould Adjustable webbing sections with female hook and loop fasteners to be 20cm long. Hook 30cm long on both ends, with loop 30cm long immediately adjacent to hook loop. Padding attached to contact areas of webbing.
1.11.1. C.7 Pelvic straps, adjustable
Poly-cotton webbing, 5cm wide. Two separate sections, one non-adjustable, other section adjustable. Non-adjustable section fitted with 5cm D-ring (chromium plated, galvanise steel) and other with male and female hook and loop fasteners. Size specific (12” and smaller = 3cm,14” and larger 5cm) - to fit seating systems above hook on one end, with loop immediately adjacent to hook. Padding attached to D-ring section, 6cm wide, 50cm long of which half extending beyond the D-ring.
1.11.1. C.8 Foot box raisers
To clip onto the existing footrest of the above seating system. Raise the back and side height by approximately 10cm, preventing the feet kicking out of the footrest. Provide firm scuff resistant padding on back and side surfaces.
1.11.1. C.9 Indoor / School base unit only for removable seat
4 brake castors 4 legs – height and tilt adjustable Seat clips on and off school frame for easy transfer to self-propelling base Distance of castor feet must match or exceed the footprint of the appropriate size of push frame of seat
1.11.1. C.10
Pad support, postural seating, complete set, washable covers for x-small, small, medium or large positioner buggy.
RT233-1.11.2 | MODULAR, RIGID ADJUSTABLE POSTURE SUPPORT BUGGY WITH FOLDING FRAME (Positioning Buggy with recline option)
Size range:
• Back height 31cm-41cm; Seat width 18cm; Seat lengths 17cm-26cm; Trunk side support height range 17cm-28cm. Trunk side support width range 16cm-21 cm. • Back height 28cm-38cm; Seat width 23cm; Seat lengths 24cm-32cm. Trunk side support height range 21cm-28cm. Trunk side support width range 18cm-24cm. • Back height 37cm-44cm; Seat width 25cm; Seat lengths 29cm-36cm. Trunk side support height range 23cm-33cm. Trunk side support width range 18cm-24cm. • Back height 40cm-50cm; Seat width 30cm; Seat lengths 37cm-47cm. Trunk side support height range 28cm-35cm. Trunk side support width range 23cm-29cm.
Seat system:
• Removable modular adjustable seating system. • Aluminium, wood, or plastic padded / lined. • Trunk side support minimum length 10cm for smallest size / 18cm for largest size. • Trunk side support independently adjustable vertically and horizontally. • Padded with a minimum of 0.6cm nonporous/ sealed cell foam. • Lined/padded with minimum of 0.6cm nonporous/ sealed cell foam. • Pelvic side support lined/padding of minimum 2cm. • Removable washable waterproof covers. • Covers to accommodate customisation of back supports. • Tilt in space up to 30 degrees. • Recline - Seat to back angle adjustable in range 90 to 120 degrees. • Pelvic strap length adjustable, padded. • Seating systems allows for adaptation, build-up and customisation. • Seat cushion with / without abduction block as specified by end user: Top layer high density foam 2.5cm - 5cm thick; 36/20 with hardness factor 18-21.6; Density - 32.1-40kg/m³). Base layer 2.5cm high density compressed (reconstituted flexible PU) foam 1300; Hardness factor 52-63; Density130+10% kg/m³). (SANS 642-1976). Waterproof 2-way stretch, breathable cover with anti-microbial shield treatment. • Head extension – height and depth adjustable, offering minimum occipital side head support. • Footrest height and depth adjustable in relation to seat length.
Lap tray:
Size range: To fit posture support buggies in the above category
Description: • Easily removable (user should be able to fit and remove table by themselves), multi-purpose tray and provides stable surface. When attached, the surface extensions should provide support for the elbows. • When attached, the curve around the body should be cut in such a way to prevent a large opening between the body and the tray where the elbows, forearms or hands can slip through. Height, depth, and angle adjustable independently from tilt in space.
Thickness: Between 0.6cm and 1.5cm Material: Wood (Supawood covered/protected with sealer and wood varnish)/ plastic with washable vinyl top Quality: Smooth, washable, splinter free surface, splinter proof, rounded edges, easily wipe-able and cleanable
Footrest:
• Solid with side supports/ foot straps. • Height and depth adjustable.
Frame:
• Folding frame • Mild Steel, epoxy coated • Size folded
Castors and forks:
• 8” x 2” (20cm x 5cm) castor wheel width between 40cm and 8cm with plastic hub. • Cast aluminium fork, with top 1cm and sides tapering to 0.4cm. • Forks to be attached to the castor stem minimum 1.2cm thick bolt system to allow for re-fastening of the stem.
Rear wheels:
• Minimum 20cm x 8cm pneumatic or solid • Range from 8”x3” to 24”x1 3/8
Weight:
• Not to exceed 25kg for largest size
Weight of user:
• Visibly marked on each wheelchair to indicate maximum weight load of end user.
Guarantee:
• Frame and seat system guaranteed for a minimum 2 year from date of issue to end user.
General requirement:
• In the technical specifications where more than 1 option is provided: Bidder must be able to provide wheelchair with all of options as specified. The end user will specify the desired option at the time of ordering. The different options stated within each technical specification description must be fully interchangeable and available at the same price. Where the end user does not specify options the wheelchair will be issued as in the option listed first. • No part or component may be changed / deviate from the technical specifications set out in this document. • Bidder to submit accredited test report to ensure device durability
1.11.2 C | CUSTOMISING OPTIONS FOR MODULAR, RIGID ADJUSTABLE POSTURE SUPPORT BUGGY WITH FOLDING FRAME
1.11.2. C.1 Tray table basic
Size range: To fit modular, rigid adjustable posture support buggy with folding frame size range.
Description: • Easily removable multi-purpose tray. • Provides stable non- slip surface. • When attached, the surface extensions should provide adequate full resting support for the elbows. • When attached, the curve around the body should be cut in such a way as to prevent a large opening between the body and the tray where the elbows, forearms or hands can slip through.
Thickness: Between 0.6cm and 1.5cm Material: Wood (Supawood covered/protected with sealer and wood varnish)/ plastic with washable vinyl top. Quality: Smooth, washable, splinter free surface, splinter proof, rounded edges, easily wipe-able and cleanable
1.11.2. C.2 Tray table padded
Size range: To fit modular, rigid adjustable posture support buggy with folding frame size range.
Description: • Easily removable multi-purpose tray. • Provides stable, dynamic, non- slip surface. • When attached, the extensions should provide full resting support for the elbows. • When attached, the curve around the body should be cut to prevent a large opening between the body and the tray where the elbows, forearms or hands can slip through. • Height, depth, and angle adjustable independently from tilt in space.
Thickness: Between 2.0cm and 4.0cm Material: Wood (Supawood covered/protected with sealer and wood varnish) or plastic inner, 2.5cm high density compressed (reconstituted flexible PU) foam with washable stretch vinyl top – fully upholstered, waterproof, material needs to be durable for regular cleaning, and flame retardant and heat resistant.
1.11.2. C.3 Head support, adjustable, basic occipital support only
Easily removable Height and depth adjustable Kidney shaped head support flexible and rake adjustable Padded and covered in heavy gauge, removable, washable fabric
1.11.2. C.4 Head support, adjustable, basic occipital, and side support
Easily removable Height and depth adjustable 4-point, hand shaped head support, flexible and rake adjustable Moulded and padded EVA
1.11.2. C.5 Chest strap
Poly-cotton webbing, 6cm or 10cm Two separate sections, adjustable, with 2 x 3cm or 5cm D-ring (chromium plated, galvanised steel) male and female loop immediately adjacent to hook Padding) attached to D-ring section, extending beyond the D-ring and width strap
1.11.2. C.6 Thigh strap
Poly-cotton webbing, One section Length adjustable. 2 x 5cm D-ring steel, epoxy coated to be attached to the seat mould Adjustable webbing sections with female hook and loop fasteners to be 20cm long. Hook 30cm long on both ends, with loop 30cm long immediately adjacent to hook loop. Padding attached to contact areas of webbing.
1.11.2. C.7 Pelvic straps, adjustable
Poly-cotton webbing, 5cm wide. Two separate sections, one non-adjustable, other section adjustable. Non-adjustable section fitted with 5cm D-ring (chromium plated, galvanise steel) and other with male and female hook and loop fasteners. Size specific ( 12” and smaller = 3cm,14” and larger 5cm) - to fit seating systems above hook on one end, with loop immediately adjacent to hook. Padding attached to D-ring section, 6cm wide, 50cm long of which half extending beyond the D-ring.
1.11.2. C.8 Foot box raisers
To clip onto the existing footrest of the above seating system. Raise the back and side height by approximately 10cm, preventing the feet kicking out of the footrest. Provide firm scuff resistant padding on back and side surfaces.
1.11.2. C.9 Indoor / School base unit only for removable seat
4 brake castors 4 legs – height and tilt adjustable Seat clips on and off school frame for easy transfer to self-propelling base Distance of castor feet must match or exceed the footprint of the appropriate size of push frame of seat
1.11.2. C.10
Pad support, postural seating, complete set, washable covers for x-small, small, medium or large positioner buggy.
RT233-1.12 | PORTER WHEELCHAIRS | Indicate the item bidding for below bby stating "Yes bidding for this item". Item left blank will be assumed it is not bidded for | Complies Yes/No | Comments (Technical Specification Deviation)
RT233-1.12.1 | PORTER WHEELCHAIRS (INSTITUTIONAL USE)
Size:
• 51 cm (20”) wide
Backrest / Seat configuration:
• Seat tilt: 5 degrees • Back – seat angle to be 90 degrees • Seat: Solid, padded • Back height: 42.5cm +2.5cm or maximum 45cm • Seat length: 40cm + 1cm
Upholstery:
• Back upholstery: Vinyl • Seat upholstery: Vinyl solid padded seat (40cm long)
Armrests:
• Removable or flip up desk type with mild steel skirt guard.
Castors and forks:
• 8” x2" (20cm x 5cm) castors with plastic hub, outer shore hardness of 65-70 and inner shore hardness of 33-70, with the inner value not exceeding the outer value. • Forks to be minimum 0.4cm thick mild steel forks. Stem to be fastened to the fork with a thread and nut system to allow for re-fastening.
Rear wheel:
• 24”x 1 3/8
Footrest hangers:
• Long (25cm) or short (20cm); Fully detachable
Footrests:
• Aluminium; Height adjustable
Foot strap:
• Vinyl strap (4cm-6cm wide) with hook and loop fasteners
Frame:
• Non-folding with stabilizer bar to limit torsion within the frame
Weight of user:
• Minimum weight capacity of 160kg. • Visibly marked on each wheelchair to indicate maximum weight load of end user.
Guarantee:
• Frame and crossbars guaranteed for a minimum of 2 years from date of issue to end user.
Tool kit:
• Adjustment tool
General requirement:
• Bidder must be able to provide wheelchair with all abovementioned options. The end user will specify the desired option at the time of ordering. The different options available within each technical specification description must be fully interchangeable and available at the same price. Where the end user does not specify options the wheelchair will be issued as in the option given first for the categories with more than one option. • No part or component may be changed / deviate from the technical specifications set out in this document. • Bidder to submit accredited test report to ensure device durability
RT233-1.12.2 | PORTER WHEELCHAIRS (INSTITUTIONAL USE) WITH DRIP STAND HOLDER AND T ROD
Size:
• 51 cm (20”) wide
Backrest / Seat configuration:
• Seat tilt: 5 degrees • Back – seat angle to be 90 degrees • Seat: Solid, padded • Back height: 42.5cm +2.5cm or maximum 45cm • Seat length: 40cm + 1cm
Upholstery:
• Back upholstery: Vinyl • Seat upholstery: Vinyl solid padded seat (40cm long)
Armrests:
• Fixed desk type with mild steel skirt guard
Castors and forks:
• 8” x2" (20cm x 5cm) castors with plastic hub, outer shore hardness of 65-70 and inner shore hardness of 33-70, with the inner value not exceeding the outer value. • Forks to be minimum 0.4cm thick mild steel forks. Stem to be fastened to the fork with a thread and nut system to allow for re-fastening.
Rear wheel:
• 12.5” x 2.25" (30cm x 6.5cm) solid tyres
Footrest hangers:
• Long (25cm) or short (20cm); Fully detachable
Footrests:
• Footplate to be hard heavy-duty plastic, or aluminium, or mild steel, and height adjustable.
Foot strap:
• Vinyl strap (4cm-6cm wide) with hook and loop fasteners
Frame:
• Non-folding with stabilizer bar to limit torsion within the frame • Height adjustable drip stand with T-bar securely attached to frame
Weight of user:
• Minimum weight capacity of 160kg. • Visibly marked on each wheelchair to indicate maximum weight load of end user.
Guarantee:
• Frame and crossbars guaranteed for a minimum of 2 years from date of issue to end user.
Tool kit:
• Adjustment tool
General requirement:
• Bidder must be able to provide wheelchair with all abovementioned options. The end user will specify the desired option at the time of ordering. The different options available within each technical specification description must be fully interchangeable and available at the same price. Where the end user does not specify options the wheelchair will be issued as in the option given first for the categories with more than one option. • No part or component may be changed / deviate from the technical specifications set out in this document. • Bidder to submit accredited test report to ensure device durability
RT233-1.13 | COMMODES | Indicate the item bidding for below bby stating "Yes bidding for this item". Item left blank will be assumed it is not bidded for | Complies Yes/No | Comments (Technical Specification Deviation)
RT233-1.13.1 | BASIC
Frame:
• Rigid frame, Mild steel, Galvanised, Epoxy coated; Constructed in such a way that it can be reversed over a standard toilet bowl.
Size:
• Width range 46cm – 51cm, seat depth 40 cm
Seat:
• Moulded polyurethane seat; with a vinyl wheelchair seat cover; removable bucket system.
Backrest:
• Full length back support in vinyl or nylon; Height 40cm
Armrests:
• Completely removable, should not form part of the rigid frame.
Footrests:
• Removable
Castors:
• 4 x Lockable, Minimum 5cm diameter
Weight limit:
• Maximum of 130 kg (consider weight limitations for bigger sizes)
Guarantee:
• Frame and crossbars guaranteed for a minimum of 2 years from date of issue during appropriate use as stated per category
General requirement:
• Bidder must be able to provide wheelchair with all of options as specified. The end user will specify the desired option at the time of ordering. The different options stated within each technical specification description must be fully interchangeable and available at the same price. Where the end user does not specify options the wheelchair will be issued as in the option listed first. • No part or component may be changed / deviate from the technical specifications set out in this document. • Bidder to submit accredited test report to ensure device durability
RT233-1.13.2 | ALUMINIUM TOILET AND SHOWER COMMODE (WHEELCHAIR TYPE)
Frame:
• Rigid frame, aluminium frame, epoxy coated; constructed in such a way that it can be reversed over a standard toilet bowl
Size:
• Width range 46cm – 51cm, seat depth 40 cm
Seat:
• Moulded polyurethane toilet seat shaped; with a vinyl seat cover; removable bucket system.
Backrest:
• Full back length back support; Height 40cm.
Armrests:
• Completely removable
Footrests:
• Removable
Rear mag wheels and axles:
• Quick release, 24x 1 3/8 wheels
Castors:
• Standard 8 x 1” or 8 x 2”
Weight limit:
• Maximum 130kg
Guarantee:
• Frame and crossbars guaranteed for a minimum of 2 years from date of issue during appropriate use as stated per category
General requirement:
• Bidder must be able to provide wheelchair with all of options as specified. The end user will specify the desired option at the time of ordering. The different options stated within each technical specification description must be fully interchangeable and available at the same price. Where the end user does not specify options the wheelchair will be issued as in the option listed first. • No part or component may be changed / deviate from the technical specifications set out in this document. • Bidder to submit accredited test report to ensure device durability
RT233-1.13.3 | COMMODE, WHEELCHAIR TYPE
• Commode must be supplied with bucket attachment, including bucket. • To comply with the technical specifications in Appendix F attached.
Overall width:
• Over the wheel hubs must not exceed 660mm
Armrest:
• The support must be so hinged on vertical members of the backrest by means of tubular hinges that they can be swung aside. • It must have means to lock them in the normal position. • Armrest surface must be of 20mm hardwood with a 75mm width and length of 300mm.
Seat:
• Must be made of at least 15mm laminated wood.
Seat Cover:
• Must be made of laminated wood or plastic.
Frame:
• The frame and armrest support shall be made from mild steel tubes having a nominal outside diameter of at least 25mm with epoxy/nylon powder-coated finish to comply with SANS 778 paragraph 5.2.
Castors:
• Both front castors to be fitted with a locking device. • Castors must comply to the latest issue of SANS 621. • Where castors are fitted into steel tubular legs, the tube shall be of wall thickness not less than 2,0mm and castors shall be fixed to the tube by one of the following methods:  Solid plug (long) complying with SANS 621 subsection 3.5.6 or  Screwed into 35mm long sleeves welded into the tubular members and locked in an acceptable manner. • Rubber or plastic expanding sleeves for fitting castors are not acceptable. • Provision shall be made on the underside of the seat to hold the slop bucket and to allow its easy insertion and withdrawal.
Guarantee:
• Frame and crossbars guaranteed for a minimum of 2 years from date of issue during appropriate use as stated per category
Accessory:
• Must comply to above technical specification, be compatible and supplied with the main unit. • 14L Stainless steel slop bucket with hinged handle
General requirement:
• Bidder must be able to provide wheelchair with all of options as specified. The end user will specify the desired option at the time of ordering. The different options stated within each technical specification description must be fully interchangeable and available at the same price. Where the end user does not specify options the wheelchair will be issued as in the option listed first. • No part or component may be changed / deviate from the technical specifications set out in this document. • Bidder to submit accredited test report to ensure device durability
RT233-1.13.4 | FOLDABLE COMMODE
Size:
• Width 46cm,
Frame:
• Aluminium frame, foldable to compact dimensions for ease of transportation and storage • Height adjustable legs
Seat:
• Moulded polyurethane toilet seat shape for removable bucket system with lid.
Backrest:
• Aluminium frame
Armrests:
• Part of foldable frame with durable plastic arm support. • Legs to be supported by non-slip rubber base.
Weight limit:
• Maximum 130kg
Guarantee:
• Frame and crossbars guaranteed for a minimum of 2 years from date of issue during appropriate use as stated per category
General requirement:
• Bidder must be able to provide wheelchair with all of options as specified. The end user will specify the desired option at the time of ordering. The different options stated within each technical specification description must be fully interchangeable and available at the same price. Where the end user does not specify options the wheelchair will be issued as in the option listed first. • No part or component may be changed / deviate from the technical specifications set out in this document. • Bidder to submit accredited test report to ensure device durability
RT233-1.14 | TRANSPORTERS | Indicate the item bidding for below bby stating "Yes bidding for this item". Item left blank will be assumed it is not bidded for | Complies Yes/No | Comments (Technical Specification Deviation)
RT233-1.14.1 | TRANSPORTER BUGGY (FOLDING / COLLAPSIBLE TYPE)
Size range:
• Seat width: 10”, 12”, 14”, 15” • Approximately 20-40cm
Backrest / Seat configuration:
• Canvas sling type or padded with support structures • Back height: Approximately 50-120cm • Seat depth: Approximately 25-45cm
Castors and forks:
• Minimum 5"
Rear wheels:
• Minimum 5"
Brakes:
• Fitted to rear wheels or fitted to handle bar
Footrests:
• Height adjustable • Fitted to frame • Detachable/ collapsible for foldable frame
Size folded:
• Maximum 120cm x 48cm x 55cm
Weight:
• Not to exceed 13kg
Weight of user:
• Visibly marked on each device to indicate maximum weight load of end user.
Frame:
• Folding, aluminium • To include recline option (90 degrees – 160 degrees) • Collapsible / foldable for storage and transportation • Datable seat for storage and folding • Padded inner structures with support for positioning
Guarantee:
• Frame and crossbars guaranteed for a minimum of 2 years from date of issue to end user.
General requirement:
• Bidder must be able to provide wheelchair with all of options as specified. The end user will specify the desired option at the time of ordering. The different options stated within each technical specification description must be fully interchangeable and available at the same price. Where the end user does not specify options the wheelchair will be issued as in the option listed first. • No part or component may be changed / deviate from the technical specifications set out in this document. • Bidder to submit accredited test report to ensure device durability
Other:
• To include headrest / head positioner • To include straps / fasteners over chest or pelvis or hips • To include storage/carrier bag facility. • To include canopy cover/ thermal cover/ rain cover • To include seat belt • Net basket on backrest • Semi-rigid seat with inner aluminium supports • Breathable upholstery
RT233-1.14.2 | TRANSPORTER WITH SIDE FRAME COVERS
Size range:
• Seat width: 12”,14”,15” • Approximately 40cm
Backrest / Seat configuration:
• Canvas sling type • Back height: Approximately 54cm • Seat depth: Approximately 29cm • Removable, washable canvas covers on side frames.
Castors and forks:
• Minimum 5"
Rear wheels:
• Minimum 5"
Brakes:
• Fitted to rear wheels
Footrests:
• Height adjustable, removable or swing away
Size folded:
• Maximum 120cm x 48cm x 55cm
Weight:
• Not to exceed 8kg
Weight of user:
• Visibly marked on each wheelchair to indicate maximum weight load of end user.
Frame:
• Folding, aluminium
Guarantee:
• Frame and crossbars guaranteed for a minimum of 2 years from date of issue to end user.
General requirement:
• Bidder must be able to provide wheelchair with all of options as specified. The end user will specify the desired option at the time of ordering. The different options stated within each technical specification description must be fully interchangeable and available at the same price. Where the end user does not specify options the wheelchair will be issued as in the option listed first. • No part or component may be changed / deviate from the technical specifications set out in this document. • Bidder to submit accredited test report to ensure device durability
RT233-1.14.3 | ADULT TRANSPORTER
Size range:
• Seat width of 17” to 18”.
Footrests:
• Height adjustable, flip up or removable.
Armrests:
• Can be swing away, removable, or fixed.
Size folded:
• Must be compactly foldable to about 1/3 of its size, with or without a carrier bag or carrier handles
Weight:
• Approximately 10 to 15 kg
Weight of user:
• Must have a weight limit of at least 100kg
Frame:
• Made from a lightweight material such as aluminium.
Guarantee:
• Frame and crossbars guaranteed for a minimum of 2 years from date of issue to end user.
General requirement:
• Bidder must be able to provide wheelchair with all of options as specified. The end user will specify the desired option at the time of ordering. The different options stated within each technical specification description must be fully interchangeable and available at the same price. Where the end user does not specify options the wheelchair will be issued as in the option listed first. • No part or component may be changed / deviate from the technical specifications set out in this document. • Bidder to submit accredited test report to ensure device durability

## 5 Category 2
CATEGORY 2
All items being tendered for must comply with the following technical specification, failure to comply with any of the conditions set out below may result in immediate disqualification. Please note that where the technical specification calls for certification, this certification must accompany the bid, failure to provide such certification will result in immediate disqualification. Bidders must give details of items being tendered for in the COMMENTS section where there are variation for further deliberation
WHEELCHAIRS CUSHIONS AND FOAM
RT233-2.1 | BASIC FOAM CUSHIONS | Indicate the item bidding for below bby stating "Yes bidding for this item". Item left blank will be assumed it is not bidded for | Complies Yes/No | Comments (Technical Specification Deviation)
TECHNICAL SPECIFICATIONS
RT233-2.1.1 | CONTOURED BASE - THICK PRESSURE CARE CUSHION
Variance allowed on overall measurement for cushions: +1cm
Size range:
• 61 cm wide; 43 cm long; 9 cm high in front
• 56 cm wide; 43 cm long; 9 cm high in front
• 51 cm wide; 43 cm long; 9 cm high in front
• 46 cm wide; 43 cm long; 9 cm high in front
• 43 cm wide; 43 cm long; 9 cm high in front
• 41 cm wide; 43 cm long; 9 cm high in front
• 38 cm wide; 43 cm long; 9 cm high in front
• 36 cm wide; 43 cm long; 9 cm high in front
• 30 cm wide; 43 cm long; 9 cm high in front
• 30 cm wide; 30 cm long; 6.5 cm high in front
• 25 cm wide; 30 cm long; 6 cm high in front
Description:
• Wheelchair basic foam pressure relieving cushion. Bevelled contoured base to match sling seat of wheelchair. • Must have trochanteric loading pads to create a well. • All cushions not to be glued at the back. • The depth of the well for 14” (36cm) - 24” (61cm) should be 3cm, for 12”(30cm) should be 2,5cm and 10” (25cm) should be 2cm. • The width of the well = 50% of the width of the cushion. • Length of the well = the width of the well up to maximum 20cm. • Sizes 14”-24”: Top layer - 5cm thick high-density foam following contour of base through full length of cushion; only the front half of the cushion to be glued. • Size 10” (keep and check) -12”: Top layer - 3cm thick high-density foam following contour of base through full length of cushion; only the front half of the cushion to be glued. • Type, size, and direction of cushion use printed on bottom of base.
Material:
Top layer:
For sizes: 25cm - 46cm High-density foam 36/20; Hardness factor 18-21.6; Density range 32.1-40kg/m³
For sizes: 51cm-61cm High density foam 50/30 Hardness factor: 249-291 and Density: 37.1- 41kg/m³
Base layer:
High density compressed (reconstituted flexible PU) foam 1300; Hardness factor-52-63; Density 130+10% kg/m³ (SANS 642 - 1976)
RT233-2.1.2 | CONTOURED BASE - WEDGED, THICK PRESSURE CARE CUSHION
Size range:
• 61 cm wide; 43 cm long; 6 cm high at rear, 11 cm high in front
• 56 cm wide; 43 cm long; 6 cm high at rear, 11 cm high in front
• 51 cm wide; 43 cm long; 6 cm high at rear, 11 cm high in front
• 46 cm wide; 43 cm long; 6 cm high at rear, 11 cm high in front
• 41 cm wide; 43 cm long; 6 cm high at rear, 11 cm high in front
• 38 cm wide; 43 cm long; 6 cm high at rear, 11 cm high in front
• 36 cm wide; 43 cm long; 6 cm high at rear, 11 cm high in front
• 30 cm wide; 43 cm long; 6 cm high at rear, 11 cm high in front
• 30 cm wide; 30 cm long; 4 cm high at rear, 8.5 cm high in front
• 25 cm wide; 30 cm long; 4 cm high at rear, 7.5 cm high in front
Variance allowed on overall measurement for cushions: +1 cm
Description:
• Wheelchair basic foam pressure relieving cushion. Bevelled contoured base to match sling seat of wheelchair. • Must have trochanteric loading pads to create a well. • The depth of the well for 14” (36cm) - 24” (61cm) should be 3cm, 12” (30cm) should be 2,5cm and 10’ (2cm) . • The width of the well = 50% of the width of the cushion. • Length of the well = the width of the well up to maximum 20cm. • 2cm wedge in front of pre-ischial shelf extending to the front of the cushion. • Sizes 14”-24”: Top layer - 5cm thick high-density foam following contour of base through full length of cushion; only the front half of the cushion to be glued. • Size 10” -12”: Top - 3 cm thick high-density foam following contour of base through full length of cushion; only the front half of the cushion to be glued. • Type, size, and direction of cushion use printed on bottom of base.
Material:
Top layer:
For sizes: 25cm – 46 cm High-density foam 36/20; Hardness factor 18 – 21.6; Density range 32.1-40kg/m³
For sizes: 51cm -61cm High density foam 50/30 Hardness factor: 249 - 291 and Density: 37.1- 41kg/m³
Base layer:
High density compressed (reconstituted flexible PU) foam 1300; Hardness factor 52-63; Density 130+10% kg/m³ (SANS 642 - 1976)
RT233-2.1.3 | CONTOURED BASE – BASIC THIN FOAM, POSITIONER CUSHION
Size range:
• 51 cm wide; 43 cm long; 7 cm high
• 46 cm wide; 43 cm long; 7 cm high
• 43 cm wide; 43 cm long; 7 cm high
• 41 cm wide; 43 cm long; 7 cm high
• 38 cm wide; 43 cm long; 7 cm high
• 36 cm wide; 43 cm long; 7 cm high
• 30 cm wide; 43 cm long; 6.5 cm high
• 30 cm wide; 33 cm long; 6.5 cm high
• 25 cm wide; 28 cm long; 6 cm high
Variance allowed on overall measurement for cushions: +1cm
Description:
• Bevelled contoured base to match sling seat of wheelchair. Sitting surface with pre-ischial shelf. • The depth of the pre-ischial shelf for 14” (36cm) - 20 (51cm) should be 3cm minimum, for 12” (30cm) should be 2,5cm minimum, and for 10” (25cm) should be 2cm minimum. • The distance of the shelf from the rear of the cushion = 50% of the width of the cushion up to maximum of 20cm. • Top: 3cm thick high-density foam following contour of base through full length of cushion; only the front half of the cushion to be glued. • Type, size, and direction of cushion use printed on bottom of base.
Material:
Top layer:
For sizes: 25cm - 48cm High-density foam 36/20; Hardness factor 18 – 21.6; Density range 32.1-40kg/m³
For sizes: 51cm High density foam 50/30 Hardness factor: 249 - 291 and Density: 37.1- 41kg/m³
Base layer:
High density compressed (reconstituted flexible PU) foam 1300; Hardness factor 52-63; Density 130+10% kg/m³ (SANS 642 - 1976)
RT233-2.1.4 | CONTOURED BASE - WEDGED, THIN FOAM POSITIONER
Size range:
• 51 cm wide; 43 cm long; 4 cm high at rear, 9 cm high in front
• 46 cm wide; 43 cm long; 4 cm high at rear, 9 cm high in front
• 43 cm wide; 43 cm long; 4 cm high at rear, 9 cm high in front
• 41 cm wide; 43 cm long; 4 cm high at rear, 9 cm high in front
• 38 cm wide; 43 cm long; 4 cm high at rear, 9 cm high in front
• 36 cm wide; 43 cm long; 4 cm high at rear, 9 cm high in front
• 30 cm wide; 43 cm long; 4 cm high at rear, 9 cm high in front
• 30 cm wide; 33 cm long; 4 cm high at rear, 8 cm high in front
• 25 cm wide; 28 cm long; 4 cm high at rear, 7.5cm high in front
Variance allowed on overall measurement for cushions: width: +/-1cm
Variance in front to be +-0.5cm
Description:
• Bevelled contoured base to match sling seat of wheelchair. Sitting surface with pre-ischial shelf. • The depth of the pre-ischial shelf for 14” (36cm) - 20” (51cm) should be 3cm minimum, 12” (30cm) should be 2,5cm minimum, and 10” (25cm) should be 2cm minimum. • The distance of the shelf from the rear = 50% of the width of the cushion up to maximum of 20cm. • 2cm wedge in front of pre-ischial shelf extending to the front of the cushion. • Top: 3cm thick high-density foam following contour of base through full length of cushion; only the front half of the cushion to be glued. • Type, size, and direction of cushion use printed on bottom of base.
Material:
Top layer:
For sizes: 25cm-46cm High-density foam 36/20; Hardness factor 18 – 21.6; Density range 32.1-40kg/m³
For sizes: 51cm High density foam 50/30 Hardness factor: 249 - 291 and Density: 37.1- 41kg/m³
Base layer:
High density compressed (reconstituted flexible PU) foam 1300; Hardness factor 52-63; Density 130+10% kg/m³ (SANS 642 - 1976)
RT233-2.1.5 | CONTOURED BASE - THICK PRESSURE CARE FOAM CUSHION, EXTENDED LENGTH
Variance allowed on overall measurement for cushion: +1 cm
Size range:
• 61 cm wide; 50 cm long; 9 cm high
• 56 cm wide; 50 cm long; 9 cm high
• 51 cm wide; 50 cm long; 9 cm high
• 46 cm wide; 50 cm long; 9 cm high
• 43 cm wide; 50 cm long; 9 cm high
• 41 cm wide; 50 cm long; 9 cm high
• 38 cm wide; 50 cm long; 9 cm high
• 36 cm wide; 50 cm long; 8.5 cm high
• 30 cm wide; 50 cm long; 8.5 cm high
Description:
• Wheelchair basic foam pressure relieving cushion. Bevelled contoured base to match sling seat of wheelchair. • Must have trochanteric loading pads to create a well. • The depth of the well for 14” (36cm) - 24” (61cm) should be 3cm, for 12” (30cm) should be 2,5cm. • The width of the well = 50% of the width of the cushion. • Length of the well = the width of the well up to maximum 20cm. • Size 14”-24”, 5cm thick high-density foam top layer following contour of base through full length of cushion; only the front half of the cushion to be glued. • Size 12”, 3cm thick high-density foam top layer, following contour of base through full length of cushion; only the front half of the cushion to be glued. • Type, size, and direction of cushion use printed on bottom of base.
Material:
Top layer:
For sizes: 30cm-46cm High-density foam 36/20; Hardness factor 18 – 21.6; Density range 32.1-40kg/m³
For sizes: 51cm-61cm High density foam 50/30 Hardness factor: 249 - 291 and Density: 37.1- 41kg/m³
Base layer:
High density compressed (reconstituted flexible PU) foam 1300; Hardness factor 52-63 Density 130+10% kg/m³ (SANS 642 - 1976)
RT233-2.1.6 | CONTOURED BASE - WEDGED, THICK PRESSURE CARE FOAM CUSHION, EXTENDED LENGTH
Size range:
• 61 cm wide; 50 cm long; 6 cm high at rear, 11 cm high in front
• 56 cm wide; 50 cm long; 6 cm high at rear, 11 cm high in front
• 51 cm wide; 50 cm long; 6 cm high at rear, 11 cm high in front
• 46 cm wide; 50 cm long; 6 cm high at rear, 11 cm high in front
• 43 cm wide; 50 cm long; 6 cm high at rear, 11 cm high in front
• 41 cm wide; 50 cm long; 6 cm high at rear, 11 cm high in front
• 38 cm wide; 50 cm long; 6 cm high at rear, 11cm high in front
• 36 cm wide; 50 cm long; 6 cm high at rear, 11 cm high in front
• 30 cm wide; 50 cm long; 6 cm high at rear, 10.5 cm high in front
Variance allowed on overall measurement for cushions: +1 cm
Description:
• Wheelchair basic foam pressure relieving cushion. Bevelled contoured base to match sling seat of wheelchair. • Must have trochanteric loading pads to create a well. • The depth of the well for 14” (36cm) - 24” (61cm) should be 3cm, 12” (30cm) should be 2,5cm. • The width of the well = 50% of the width of the cushion. • Length of the well = the width of the well up to maximum 20cm. • Size 14”-24”, 5cm thick high-density foam top layer, following contour of base; only the front half of the cushion to be glued. • Size 12”, 3cm, thick high density foam top layer, following contour of base through full length of cushion, only the front half of the cushion to be glued. • Type, size, and direction of cushion use printed on bottom of base.
Material:
Top layer:
For sizes: 30cm-46cm High-density foam 36/20; Hardness factor 18 – 21.6; Density range 32.1-40kg/m³
For sizes: 51cm-61cm High density foam 50/30 Hardness factor: 249 - 291 and Density: 37.1- 41kg/m³
Base layer:
High density compressed (reconstituted flexible PU) foam 1300; Hardness factor 52-63 Density 130+10% kg/m³ (SANS 642 - 1976)
RT233-2.1.7 | CONTOURED BASE - THIN, BASIC POSITIONER FOAM CUSHION, EXTENDED LENGTH
Size range:
• 51 cm wide; 50 cm long; 7 cm high in front
• 46 cm wide; 50 cm long; 7 cm high in front
• 43 cm wide; 50 cm long; 7 cm high in front
• 41 cm wide; 50 cm long; 7 cm high in front
• 38 cm wide; 50 cm long; 7 cm high in front
• 36 cm wide; 50 cm long; 7 cm high in front
• 30 cm wide; 50 cm long; 7 cm high in front
Variance allowed on overall measurement for cushions: +1 cm
The entire size range will remain one series
Description:
• Bevelled contoured base to match sling seat of wheelchair. Sitting surface with pre-ischial shelf. • The depth of the pre-ischial shelf for 14” (36cm) - 20” (51cm) should be 3cm minimum; 12” (30cm) should be 2, 5cm minimum. • The distance of the shelf from the rear of the cushion = 50% of the width of the cushion up to maximum of 20cm. • Top: 3cm thick high-density foam following contour of base through full length of cushion; only the front half of the cushion to be glued. • Type, size, and direction of cushion use printed on bottom of base.
Material:
Top layer:
For Sizes: 30cm -46cm High-density foam 36/20; Hardness factor 18 – 21.6; Density range 32.1-40kg/m³
For sizes: 51cm High density foam 50/30 Hardness factor: 249 - 291 and Density: 37.1- 41kg/m³
Base layer:
High density compressed (reconstituted flexible PU) foam 1300; Hardness factor 52-63; Density 130+10% kg/m³ (SANS 642 - 1976)
RT233-2.1.8 | CONTOURED BASE - WEDGED, THIN FOAM, POSITIONER, EXTENDED LENGTH
Size range:
• 51 cm wide; 50 cm long; 4 cm high at rear, 9 cm in front
• 46 cm wide; 50 cm long; 4 cm high at rear, 9 cm in front
• 43 cm wide; 50 cm long; 4 cm high at rear, 9 cm in front
• 41 cm wide; 50 cm long; 4 cm high at rear, 9 cm in front
• 38 cm wide; 50 cm long; 4 cm high at rear, 9 cm in front
• 36 cm wide; 50 cm long; 4 cm high at rear, 9 cm in front
• 30 cm wide; 50 cm long; 4 cm high at rear, 8.5 cm in front
Variance allowed on overall measurement for cushions: +1 cm
The entire size range will remain one series
Description:
• Bevelled contoured base to match sling seat of wheelchair. Sitting surface with pre-ischial shelf. • The depth of the pre-ischial shelf for 14” (36cm) - 20” (51cm) should be 3cm minimum, 12” (30cm) should be 2, 5 cm minimum. • The distance of the shelf from the rear of the cushion = 50% of the width of the cushion up to maximum of 20cm. • 2cm wedge in front of pre-ischial shelf extending to the front of the cushion. • Top: 3cm thick high-density foam following contour of base through full length of cushion; only the front half of the cushion to be glued. • Type, size, and direction of cushion use printed on bottom of base
Material:
Top layer:
For sizes: 30cm -46cm High-density foam 36/20; Hardness factor 18 – 21.6; Density range 32.1-40kg/m³
For sizes: 51cm High density foam 50/30 Hardness factor: 249 - 291 and Density: 37.1- 41kg/m³
Base layer:
High density compressed (reconstituted flexible PU) foam 1300; Hardness factor 52-63; Density 130+10% kg/m³ (SANS 642 - 1976)
RT233-2.1.9 | NON-CONTOURED BASE - STRAIGHT BASE, BASIC FOAM, THICK PRESSURE CARE CUSHION WITH NO ABDUCTION BLOCK
Size Range:
• 20 cm wide, 25 cm long, 9cm high
• 22.5 cm wide, 25 cm long, 9cm high
• 25 cm wide, 35 cm long, 9cm high
• 27.5 cm wide, 40 cm long, 9cm high
• 30 cm wide, 35 cm long, 9cm high
• 30 cm wide, 40 cm long, , 9cm high
• 36 cm wide, 40 cm long, , 9cm high
• 36 cm wide , 50 cm long, 9cm high
• 38cm wide, 50cm long, 9cm high
• 40 cm wide, 50 cm long, 9cm high
• 43 cm wide, 50 cm long, 9cm high
• 45 cm wide, 50 cm long, 9cm high
• 51 cm wide, 50 cm long, 9cm high
Variance allowed on overall measurement for cushions: width: +1 cm
Description:
• Wheelchair basic foam pressure relieving cushion. Flat base to match solid wheelchair seat. • Must have trochanteric loading pads to create a well. • The depth of the well for 14” (36cm) - 20”(51cm) should be 3cm, for 8” (20cm) - 12”(30cm) should be 2,5cm. • The width of the well = 50% of the width of the cushion • The length of the well = the width of the well up to maximum 20cm. • Base layer: Minimum thickness 4 cm for sizes 14”-20” and minimum thickness 3cm for 12” and minimum 2,5cm for 10” and smaller. • The thickness of the top layer for 14” (36cm) -20” (51cm) should be 5cm minimum and for 8” (20cm) -12” (30cm ) should be 3cm; only the front half of the cushion to be glued. • Type, size, and direction of cushion use printed on bottom of base.
Material:
Top layer:
For sizes to fit: 8”-18” High-density foam 36/20; Hardness factor 18 – 21.6; Density range 32.1-40kg/m³
For sizes: 20” High density foam 50/30 Hardness factor: 249 - 291 and Density: 37.1- 41kg/m³
Base layer:
High density compressed (reconstituted flexible PU) foam 1300 Hardness factor 52-63; Density 130+10% kg/m³ (SANS 642 - 1976)
RT233-2.1.10 | NON-CONTOURED BASE - STRAIGHT BASE, THICK PRESSURE CARE WITH ABDUCTION BLOCK
Size Range:
• 20 cm wide, 25 cm long, 9cm high with extra-small abduction block
• 22.5 cm wide, 25 cm long, 9cm high with extra-small abduction block
• 25 cm wide, 35 cm long, 9cm high with extra-small abduction block
• 27.5 cm wide, 40 cm long, 9cm high with extra-small abduction block
• 30 cm wide, 35 cm long, 9cm high, with small abduction block
• 30 cm wide, 40 cm long, 9cm high, with small abduction block
• 35 cm wide, 40 cm long, 9cm high, with small abduction block
• 36 cm wide, 50 cm long, 9cm high, with small abduction block
• 38cm wide, 50cm long, 9cm high, with small abduction block
• 41 cm wide, 50 cm long, 9cm high with large abduction block
• 43 cm wide, 50 cm long, 9cm high with large abduction block
• 46 cm wide, 50 cm long, 9cm high with large abduction block
• 51 cm wide, 50 cm long, 9cm high with large abduction block
Variance allowed on overall measurement for cushions: width: +1 cm
Description:
| • Wheelchair basic foam pressure relieving cushion. Flat base to match solid wheelchair seat. • Must have trochanteric loading pads to create a well. • The depth of the well for 14” (36cm) - 20”(51cm) should be 3cm, for 8” (20cm) - 12”(30cm) should be 2,5cm. • The width of the well = 50% of the width of the cushion • The length of the well = the width of the well up to maximum 20cm. • Base layer: Minimum thickness 4 cm for sizes 14”-20” and minimum thickness 3cm for 12” and minimum 2,5cm for 10” and smaller. • The thickness of the top layer for 14” (36cm) -20” (51cm) should be 5cm minimum and for 8” (20cm) -12” (30cm ) should be 3cm; only the front half of the cushion to be glued. • The pommel should extend beyond the end of the cushion to reach the front of the user’s knees. • The height of the pommel = the full thickness of the cushion from the base and includes average height of user’s thigh in the front. • The length of the pommel should not be more than the distal third of the femur. • Type, size, and direction of cushion use printed on bottom of base. |  |  | No specific measurements of the pommel provided
Material:
Top layer:
For sizes to fit: 8”-18” High-density foam 36/20; Hardness factor 18 – 21.6; Density range 32.1-40kg/m³
For sizes: 20” High density foam 50/30 Hardness factor: 249 - 291 and Density: 37.1- 41kg/m³
Base layer:
High density compressed (reconstituted flexible PU) foam 1300 Hardness factor 52-63; Density 130+10% kg/m³ (SANS 642 - 1976)
RT233-2.1.11 | NON-CONTOURED BASE - STRAIGHT BASE, THIN FOAM POSITIONER WITH NO ABDUCTION BLOCK
Size Range:
• 20 cm wide, 25 cm long, 7cm high
• 22.5 cm wide, 25 cm long, 7cm high
• 25 cm wide, 35 cm long, 7cm high
• 27.5 cm wide, 40 cm long, 7cm high
• 30 cm wide, 35 cm long, 7cm high
• 30 cm wide, 40 cm long, 7cm high
• 36 cm wide, 50 cm long, 7cm high
• 38cm wide, 50cm long, 7 cm high
• 40 cm wide, 50 cm long, 7cm high
• 43 cm wide, 50 cm long, 7cm high
• 45 cm wide, 50 cm long, 7cm high
• 51 cm wide, 50 cm long, 7cm high
Variance allowed on overall measurement for cushions: width: +1 cm
The entire size range will remain one series
Description:
• Wheelchair basic foam positioner cushion. Flat base to match solid wheelchair seat • Sitting surface with pre-ischial shelf. • The depth of the pre-ischial shelf for 14” (36cm) - 20” (51cm) should be 3cm minimum, 12” (30cm) should be 2,5 cm minimum. • The distance of the shelf from the rear of the cushion = 50% of the width of the cushion up to maximum of 20cm. • Base layer: Minimum thickness 4cm for sizes 14”-20” and minimum thickness 3cm for 12” and minimum 2,5cm for 10” and smaller. • The thickness of the top layer for 14” (36cm) -20” (51cm) should be 3cm minimum and for 8” (20cm) -12” (30cm ) should be 2.5cm.; only the front half of the cushion to be glued. • Type, size, and direction of cushion use printed on bottom of base.
Material:
Top layer:
For sizes to fit: 8”-18” High-density foam 36/20; Hardness factor 18 – 21.6; Density range 32.1-40kg/m³
For sizes: 20” High density foam 50/30 Hardness factor: 249 - 291 and Density: 37.1- 41kg/m³
Base layer:
High density compressed (reconstituted flexible PU) foam 1300 Hardness factor 52-63: Density 130+10% kg/m³ (SANS 642 - 1976)
RT233-2.1.12 | NON-CONTOURED BASE - STRAIGHT BASE, THIN FOAM POSITIONER WITH ABDUCTION BLOCK
Size Range:
• 20 cm wide, 25 cm long, 7cm high with extra-small abduction block
• 22.5 cm wide, 25 cm long, 7cm high with extra-small abduction block
• 25 cm wide, 35 cm long, 7cm high with extra-small abduction block
• 27.5 cm wide, 40 cm long, 7cm high with extra-small abduction block
• 30 cm wide, 35 cm long, 7cm high with small abduction block
• 30 cm wide, 40 cm long, 7cm high with small abduction block
• 35 cm wide, 40 cm long, 7cm high with small abduction block
• 36 cm wide, 50 cm long, 7cm high with small abduction block
• 41 cm wide, 50 cm long, 7cm high with large abduction block
• 43 cm wide, 50 cm long, 7cm high with large abduction block
• 46 cm wide, 50 cm long, 7cm high with large abduction block
• 51 cm wide, 50 cm long, 7cm high with large abduction block
Variance allowed on overall measurement for cushions: width: +1 cm
The entire size range will remain one series
Description:
| • Wheelchair basic foam positioner cushion. Flat base to match solid wheelchair seat • Sitting surface with pre-ischial shelf. • The depth of the well for 14” (36cm) - 20”(51cm) should be 30mm, for 12”(30cm) should be 2,5cm. • The distance of the shelf from the rear of the cushion = 50% of the width of the cushion up to maximum of 20cm. • Base layer: Minimum thickness 4cm for sizes 14”-20” and minimum thickness 3cm for 12” and minimum 2,5cm for 10” and smaller. • The thickness of the top layer for 14” (36cm) -20” (51cm) should be 3cm minimum and for 8” (20cm) -12” (30cm ) should be 2.5cm.; only the front half of the cushion to be glued. • The pommel should extend beyond the end of the cushion to reach the front of the user’s knees. • The height of the pommel = the full thickness of the cushion from the base and includes average height of user’s thigh in the front. • The length of the pommel should not be more than the distal third of the femur. • Type, size, and direction of cushion use printed on bottom of base. |  |  | No specific measurements of the pommel provided
Material:
Top layer:
For sizes to fit: 8”-18” High-density foam 36/20; Hardness factor 18 – 21.6; Density range 32.1-40kg/m³
For sizes: 20” High density foam 50/30 Hardness factor: 249 - 291 and Density: 37.1- 41kg/m³
Base layer:
High density compressed (reconstituted flexible PU) foam 1300 Hardness factor 52-63: Density 130+10% kg/m³ (SANS 642 - 1976)
RT233-2.2 | FLUID FILLED CUSHION | Indicate the item bidding for below bby stating "Yes bidding for this item". Item left blank will be assumed it is not bidded for | Complies Yes/No | Comments (Technical Specification Deviation)
RT233-2.2.1 | FLUID FILLED CUSHIONS WITH CONTOURED BASE AND REAR FLUID PAD
Size range:
• 51 cm wide, 45cm long, no thicker than 8 cm
• 46 cm wide, 43 cm long, no thicker than 8 cm
• 46 cm wide, 45 cm long, no thicker than 8 cm
• 43 cm ) wide, 45 cm long, no thicker than 8 cm
• 43 cm wide, 43 cm long, no thicker than 8 cm
• 41 cm wide, 45 cm long, no thicker than 8 cm
• 41 cm wide, 43 cm long, no thicker than 8 cm
• 38 cm wide, 43 cm long, no thicker than 8 cm
• 36 cm wide, 45 cm long, no thicker than 8 cm
• 36 cm wide, 41 cm long, no thicker than 8 cm
Variance allowed on overall measurement for cushions: width: -1cm (minus 1cm); length: +1cm; thickness: +0.5cm
Description:
• Wheelchair contoured fluid filled medium to high-risk pressure relieving cushion. • Type, size, and direction of cushion use printed on base of cushion.
Base layer:
• Bottom of base flat. To be supplied with solid seat insert or sling filler for use on sling seats. • Sitting surface of base to be contoured with ischial well and trochanteric loading pads. Ischial well to be at least 4 + 1cm high and front edge straight or slightly curved. Fluid pad in ischial well. Flat area under thighs with central raise between thighs of 1.25cm- 2.5cm high. • Removable side inserts to position legs. These may integrate with the base cover or be applied directly to the base.
Fluid pad:
• Viscous fluid. Able to move away from pressure and mould to pressure area. • Viscous fluid of a non-evaporating, non-toxic nature. Fluid may not contain any irritants. Viscosity like petroleum jelly. Fluid should not be able to be compressed and volume may not change with temperature. Viscosity may not change properties under repeated sheer. Must remain fluid even in temperatures to 0 degrees. • Temperature range at which it still works (viscosity/nature of the gel will change in extreme temperatures. • Fluid pad minimum 3 sections, one each for left and right ischial tuberosities and one under the coccyx. • Adequate volume to completely fill ischial well when fitted correctly. • Removable. Secure attachment by means of minimum 5 hook and loop strips. • Design and attachment to minimise risk to bottom out by allowing for adequate volume of pad, volume of fluid, viscosity of fluid and attachment parameters. • Fluid pad at least 25% wider than underlying base when pad is laid down flat.
Covers:
• Separate removable base cover. • Separate removable base cover waterproof. • Separate removable external cover. External cover to allow moulding with fluid pad without interfering with pressure relieving properties. • External cover machine washable.
Weight of cushion:
• Maximum weight 5kg
User weight capacity:
• To accommodate users weighing up to 150kg
Guarantee:
• Minimum 2-year guarantee against defective manufacturing
RT233-2.3 | AIR FILLED CUSHIONS | Indicate the item bidding for below bby stating "Yes bidding for this item". Item left blank will be assumed it is not bidded for | Complies Yes/No | Comments (Technical Specification Deviation)
RT233-2.3.1 | AIR FILLED CUSHIONS WITH FLEXIBLE BASE
Size range:
• 51 cm wide, 46 cm long, no higher than 10 cm
• 51 cm wide, 41 cm long, no higher than 10 cm
• 46 cm wide, 46 cm long, no higher than 10 cm
• 46 cm wide, 41 cm long, no higher than 10 cm
• 41 cm wide, 46 cm long, no higher than 10 cm
• 41 cm wide, 41 cm long, no higher than 10 cm
• 38 cm wide, 46 cm long, no higher than 10 cm
• 36 cm wide, 41 cm long, no higher than 10 cm
Variance allowed on overall measurement for cushions: width: -1cm (minus 1cm); length: +1cm; thickness: +0.5cm
Description:
• Air -filled high risk pressure relieving cushion. • Air contained in multiple interlinked compartments, 10cm in height. • Able to regulate the volume of air in compartments to prevent bottoming out. • Dedicated dual purpose in- and outlet valve. • Multiple compartments arranged in 4 functional units which can be sealed off from each other to allow for customising of air pressures in each of the 4 compartments. • Dedicated air lock to control flow between the 4 compartments.
Base layer:
• Thin rubber, flexible, base integrated with air-filled interface layer.
Accessories:
• Dedicated pump • Repair kit • Cover
Weight of cushion:
• Up to 1,7 kg
Guarantee:
Minimum 2-year guarantee against defective manufacturing
RT233-2.3.2 | CUSHIONS WITH INDIVIDUAL REMOVABLE AIR-FILLED POCKETS/CELLS
Size range:
• 51 cm wide, 51 cm long, no higher than 10 cm
• 51 cm wide, 46 cm long, no higher than 10 cm
• 46 cm wide, 51 cm long, no higher than 10 cm
• 46 cm wide, 46 cm long, no higher than 10 cm
• 46 cm wide, 41 cm long, no higher than 10 cm
• 43 cm wide, 43 cm long, no higher than 10 cm
• 41 cm wide, 46 cm long, no higher than 10 cm
• 41 cm wide, 41 cm long, no higher than 10 cm
• 36 cm wide, 41 cm long, no higher than 10 cm
Variance allowed on overall measurement for cushions: width: -1cm (minus 1cm); length: +1cm; thickness: +0.5cm
Description:
• Air -filled high risk pressure relieving cushion. • Compartments/sectioned with clear ischial seat area and individual left /right sections. Minimum 6 pockets in total: Compartments filled with individual pre- filled airtight pockets/cells.
Cover:
• Separate stretchable, removable cover. Cover machine washable. • Non-slip bottom of cover with hook and loop strips to attach cushion to wheelchair. • Type, size, and direction of cushion use printed on cover.
Guarantee:
• Minimum 2-year guarantee against defective manufacturing
Accessories:
• Covers and • Individual air pockets
RT233-2.4 | COMPRESSED FOAM WEDGE | Indicate the item bidding for below bby stating "Yes bidding for this item". Item left blank will be assumed it is not bidded for | Complies Yes/No | Comments (Technical Specification Deviation)
Variance allowed on overall measurement for wedge: width: -1cm (minus 1cm); length: +1cm
Range:
• 45 cm wide; 45 cm long; 1 cm high at rear, 3 cm high in front • 45 cm wide; 45 cm long;1 cm high at rear, 5 cm high in front
Description:
• Compressed foam wedge to be used to customise wheelchair cushions (non-contoured base).
Material:
• High density compressed (reconstituted flexible PU) foam 1300; Hardness factor –52-63; Density 130+10% kg/m³ (SANS 642 - 1976).
RT233-2.5 | COMPRESSED FOAM SHEET | Indicate the item bidding for below bby stating "Yes bidding for this item". Item left blank will be assumed it is not bidded for | Complies Yes/No | Comments (Technical Specification Deviation)
Variance allowed on overall measurement for foam sheets: width: -1cm (minus 1cm); length: +1cm
Range:
• 122 cm x 180 cm x 5 cm
• 122 cm x 180 cm x 3 cm
• 122 cm x 180 cm x 2 cm
• 120 cm x 180 cm x 1.2 cm
• 137 cm x 94 cm x 5cm
• 137 cm x 94 cm x 3cm
• 137 cm x 94 cm x 2 cm
• 137 cmx 94 cm x 1.2 cm
Description:
• Compressed foam sheet
Material:
• High density compressed (reconstituted flexible PU) foam 1300; Hardness factor 52-63; Density 130+10% kg/m³ (SABS 642 - 1976)
Optional:
• No cost option to order 1 full sheet or 1 sheet cut into two at 68,5cm x 94cm x thickness
RT233-2.6 | HIGH DENSITY FOAM SHEET | Indicate the item bidding for below bby stating "Yes bidding for this item". Item left blank will be assumed it is not bidded for | Complies Yes/No | Comments (Technical Specification Deviation)
Variance allowed on overall measurement for foam sheets: width: -1cm (minus 1cm); length: +1cm
Range:
• 188 cm x 183 cm x 1.2 cm
• 188 cm x 183 cm x 3 cm
• 188 cm x 183 cm x 5 cm
• 188 cm x 183 cm x 10 cm
• 188 cm x 183 cm x 15 cm
• 190 cm x 115 cm x 1.2 cm
• 190 cm x 115 cm x 3 cm
• 190 cm x 115 cm x 5 cm
• 190 cm x 115 cm x 10 cm
• 190 cm x 115 cm x 15 cm
Description:
• High density foam sheet
Material:
• High-density foam 36/20; Hardness factor 18 – 21.6; Density range 32.1-40kg/m³
Optional:
• No cost option to order 1 full sheet or 1 sheet cut into two at 68,5cm x 94cm x thickness

## 6 Category 3
CATEGORY 3
All items being tendered for must comply with the following technical specification, failure to comply with any of the conditions set out below may result in immediate disqualification. Please note that where the technical specification calls for certification, this certification must accompany the bid, failure to provide such certification will result in immediate disqualification. Bidders must give details of items being tendered for in the COMMENTS section where there are variation for further deliberation
WHEELCHAIRS CUSHIONS WITH COVERS
RT233-3.1 | CUSHION COVERS FOR CONTOURED BASED, BASIC FOAM | Indicate the item bidding for below bby stating "Yes bidding for this item". Item left blank will be assumed it is not bidded for | Complies Yes/No | Comments (Technical Specification Deviation)
TECHNICAL SPECIFICATIONS
RT233-3.1.1 | NYLON COVER TO FIT THICK PRESSURE CARE CUSHION WITH CONTOURED BASE
Size range:
• 61 cm wide; 43 cm long; 9 cm high front
• 56 cm wide; 43 cm long; 9 cm high front
• 51 cm wide; 43 cm long; 9 cm high front
• 43 cm wide; 43 cm long; 9 cm high front
• 46 cm wide; 43 cm long; 9 cm high front
• 41 cm wide; 43 cm long; 9 cm high front
• 38 cm wide; 43 cm long; 9 cm high front
• 36 cm wide; 43 cm long; 9 cm high front
• 30 cm wide; 43 cm long; 9 cm high front
• 30 cm wide; 30 cm long; 6.5 cm high front
• 25 cm wide; 30 cm long; 6 cm high front
Variance allowed on overall measurement for cushions: +1cm.
Description of cushion cover:
• Heavy duty nylon (300g/m²), colour fast. • 2 ways stretch. • Breathable and machine washable. • Pull over elastic band with size label. • Colour: Preferably Black, alternately Navy Blue
RT233-3.1.2 | WATERPROOF COVER TO FIT THICK PRESSURE CARE CUSHION WITH CONTOURED BASE
Size range:
• 61 cm wide; 43 cm long; 9 cm high front
• 56 cm wide; 43 cm long; 9 cm high front
• 51 cm wide; 43 cm long; 9 cm high front
• 43 cm wide; 43 cm long; 9 cm high front
• 46 cm wide; 43 cm long; 9 cm high front
• 41 cm wide; 43 cm long; 9 cm high front
• 38 cm wide; 43 cm long; 9 cm high front
• 36 cm wide; 43 cm long; 9 cm high front
• 30 cm wide; 43 cm long; 9 cm high front
• 30 cm wide; 30 cm long; 6.5 cm high front
• 25 cm wide; 30 cm long; 6 cm high front
Variance allowed on overall measurement for cushions: +1cm.
Description of cushion cover:
• Fabric, waterproof, terry cloth covered. • 2 ways stretch, breathable, anti-microbial shield treatment. • Machine washable. • Colour: Preferably Black, alternately Navy Blue • Pull over elastic band with size label.
RT233-3.1.3 | NYLON COVER TO FIT WEDGED, THICK PRESSURE CARE CUSHION WITH CONTOURED BASE
Size range:
• 61 cm wide; 43 cm long; 6 cm high at rear, 11 cm in front
• 56 cm wide; 43 cm long; 6 cm high at rear, 11 cm in front
• 51 cm wide; 43 cm long; 6 cm high at rear, 11 cm in front
• 43 cm wide; 43 cm long; 6 cm high at rear, 11 cm in front
• 46 cm wide; 43 cm long; 6 cm high at rear, 11 cm in front
• 41 cm wide; 43 cm long; 6 cm high at rear, 11 cm in front
• 38 cm wide; 43 cm long; 6 cm high at rear, 11 cm in front
• 36 cm wide; 43 cm long; 6 cm high at rear, 11 cm in front
• 30 cm wide; 43 cm long; 6 cm high at rear, 11 cm in front
• 30 cm wide; 30 cm long; 4 cm high at rear, 8.5 cm in front
• 25 cm wide; 30 cm long; 4 cm high at rear, 7.5 cm in front
Variance allowed on overall measurement for cushions: +1 cm
Description of cushion cover:
• Heavy duty nylon (300g/m²), colour fast. • 2-way stretch. • Breathable and machine washable. • Pull over elastic band with size label. • Colour: Preferably Black, alternately Navy Blue
RT233-3.1.4 | WATERPROOF COVER TO FIT WEDGED, THICK PRESSURE CARE CUSHION WITH CONTOURED BASE
Size range:
• 61 cm wide; 43 cm long; 6 cm high at rear, 11 cm in front
• 56 cm wide; 43 cm long; 6 cm high at rear, 11 cm in front
• 51 cm wide; 43 cm long; 6 cm high at rear, 11 cm in front
• 43 cm wide; 43 cm long; 6 cm high at rear, 11 cm in front
• 46 cm wide; 43 cm long; 6 cm high at rear, 11 cm in front
• 41 cm wide; 43 cm long; 6 cm high at rear, 11 cm in front
• 38 cm wide; 43 cm long; 6 cm high at rear, 11 cm in front
• 36 cm wide; 43 cm long; 6 cm high at rear, 11 cm in front
• 30 cm wide; 43 cm long; 6 cm high at rear, 11 cm in front
• 30 cm wide; 30 cm long; 4 cm high at rear, 8.5 cm in front
• 25 cm wide; 30 cm long; 4 cm high at rear, 7.5 cm in front
Variance allowed on overall measurement for cushions: +1 cm.
Description of cushion cover:
• Fabric, waterproof, terry cloth covered. • 2 ways stretch, breathable, anti-microbial shield treatment. • Machine washable. • Pull over elastic band with size label. • Colour: Preferably Black, alternately Navy Blue.
RT233-3.1.5 | NYLON COVER TO FIT THIN FOAM POSITIONER CUSHION WITH CONTOURED BASE
Size range:
• 51 cm wide; 43 cm long; 7 cm high
• 46 cm wide; 43 cm long; 7 cm high
• 43 cm wide; 43 cm long; 7 cm high
• 41 cm wide; 43 cm long; 7 cm high
• 38 cm wide; 43 cm long; 7 cm high
• 36 cm wide; 43 cm long; 7 cm high
• 30 cm wide; 43 cm long; 6.5 cm high
• 30 cm wide; 33 cm long; 6.5 cm high
• 25 cm wide; 28 cm long; 6 cm high
Variance allowed on overall measurement for cushions: +1cm
Description of cushion cover:
• Heavy duty nylon (300g/m²), colour fast. • 2 ways stretch. • Breathable and machine washable. • Pull over elastic band with size label. • Colour: Preferably Black, alternately Navy Blue.
RT233-3.1.6 | WATERPROOF COVER TO FIT THIN FOAM POSITIONER CUSHION WITH CONTOURED BASE
Size range:
• 51 cm wide; 43 cm long; 7 cm high
• 46 cm wide; 43 cm long; 7 cm high
• 43 cm wide; 43 cm long; 7 cm high
• 41 cm wide; 43 cm long; 7 cm high
• 38 cm wide; 43 cm long; 7 cm high
• 36 cm wide; 43 cm long; 7 cm high
• 30 cm wide; 43 cm long; 6.5 cm high
• 30 cm wide; 33 cm long; 6.5 cm high
• 25 cm wide; 28 cm long; 6 cm high
Variance allowed on overall measurement for cushions: +1cm
Description of cushion cover:
• Fabric, waterproof, terry cloth covered. • 2 ways stretch, breathable, anti-microbial shield treatment. • Machine washable. • Colour: Preferably Black, alternately Navy Blue.
RT233-3.1.7 | NYLON CUSHION COVER TO FIT WEDGED, THIN FOAM POSITIONER WITH CONTOURED BASE
Size range:
• 51 cm wide; 43 cm long; 4 cm high at rear, 9 cm in front
• 46 cm wide; 43 cm long; 4 cm high at rear, 9 cm in front
• 43 cm wide; 43 cm long; 4 cm high at rear, 9 cm in front
• 41 cm wide; 43 cm long; 4 cm high at rear, 9 cm in front
• 38 cm wide; 43 cm long; 4 cm high at rear, 9 cm in front
• 36 cm wide; 43 cm long; 4 cm high at rear, 9 cm in front
• 30 cm wide; 43 cm long; 4 cm high at rear, 9 cm in front
• 30 cm wide; 33 cm long; 4 cm high at rear, 8.5 cm in front
• 25 cm wide; 28 cm long; 4 cm high at rear, 7.5cm in front
Variance allowed on overall measurement for cushions: width: +1cm
Description of cushion cover:
• Heavy duty nylon (300g/m²), colour fast. • 2 ways stretch. • Breathable and machine washable. • Pull over elastic band with size label. • Colour: Preferably Black, alternately Navy Blue.
RT233-3.1.8 | WATERPROOF CUSHION COVER TO FIT WEDGED, THIN FOAM POSITIONER WITH CONTOURED BASE
Size range:
• 51 cm wide; 43 cm long; 4 cm high at rear, 9 cm in front
• 46 cm wide; 43 cm long; 4 cm high at rear, 9 cm in front
• 43 cm wide; 43 cm long; 4 cm high at rear, 9 cm in front
• 41 cm wide; 43 cm long; 4 cm high at rear, 9 cm in front
• 38 cm wide; 43 cm long; 4 cm high at rear, 9 cm in front
• 36 cm wide; 43 cm long; 4 cm high at rear, 9 cm in front
• 30 cm wide; 43 cm long; 4 cm high at rear, 9 cm in front
• 30 cm wide; 33 cm long; 4 cm high at rear, 8.5 cm in front
• 25 cm wide; 28 cm long; 4 cm high at rear, 7.5cm in front
Variance allowed on overall measurement for cushions: width: +1cm
Description of cushion cover:
• Fabric, waterproof, terry cloth covered. • 2 ways stretch, breathable, anti-microbial shield treatment. • Machine washable. • Colour: Preferably Black, alternately Navy Blue.
RT233-3.2 | CUSHION COVERS FOR CONTOURED BASED, BASIC FOAM EXTENDED LENGTH | Indicate the item bidding for below bby stating "Yes bidding for this item". Item left blank will be assumed it is not bidded for | Complies Yes/No | Comments (Technical Specification Deviation)
RT233-3.2.1 | NYLON CUSHION COVER TO FIT THICK PRESSURE CARE FOAM CUSHION WITH CONTOURED BASE, EXTENDED LENGTH
Size range:
• 61 cm wide; 50 cm long; 9 cm high • 56 cm wide; 50 cm long; 9 cm high • 51 cm wide; 50 cm long; 9 cm high • 46 cm wide; 50 cm long; 9 cm high • 43 cm wide; 50 cm long; 9 cm high • 41 cm wide; 50 cm long; 9 cm high • 38 cm wide; 50 cm long; 9 cm high • 36 cm wide; 50 cm long; 9 cm high • 30 cm wide; 50 cm long; 6.5 cm high
Variance allowed on overall measurement for cushion: +1 cm
Description of cushion cover:
• Heavy duty nylon (300g/m²), colour fast. • 2 ways stretch. • Breathable and machine washable. • Pull over elastic band with size label. • Colour: Preferably Black, alternately Navy Blue.
RT233-3.2.2 | WATERPROOF CUSHION COVER TO FIT THICK PRESSURE CARE FOAM CUSHION WITH CONTOURED BASE, EXTENDED LENGTH
Size range:
• 61 cm wide; 50 cm long; 9 cm high
• 56 cm wide; 50 cm long; 9 cm high
• 51 cm wide; 50 cm long; 9 cm high
• 46 cm wide; 50 cm long; 9 cm high
• 43 cm wide; 50 cm long; 9 cm high
• 41 cm wide; 50 cm long; 9 cm high
• 38 cm wide; 50 cm long; 9 cm high
• 36 cm wide; 50 cm long; 9 cm high
• 30 cm wide; 50 cm long; 6.5 cm high
Variance allowed on overall measurement for cushion: +1 cm
Description of cushion cover:
• Fabric, waterproof, terry cloth covered. • 2 ways stretch, breathable, anti-microbial shield treatment. • Machine washable. • Colour: Preferably Black, alternately Navy Blue.
RT233-3.2.3 | NYLON CUSHION COVER TO FIT WEDGED, THICK PRESSURE CARE FOAM CUSHION WITH CONTOURED BASE, EXTENDED LENGTH
Size range:
• 61 cm wide; 50 cm long; 6 cm high at rear, 11 cm in front
• 56 cm wide; 50 cm long; 6 cm high at rear, 11 cm in front
• 51 cm wide; 50 cm long; 6 cm high at rear, 11 cm in front
• 46 cm wide; 50 cm long; 6 cm high at rear, 11 cm in front
• 43 cm wide; 50 cm long; 6 cm high at rear, 11 cm in front
• 41 cm wide; 50 cm long; 6 cm high at rear, 11 cm in front
• 38 cm wide; 50 cm long; 6 cm high at rear, 11cm in front
• 36 cm wide; 50 cm long; 6 cm high at rear, 11 cm in front
• 30 cm wide; 50 cm long; 4 cm high at rear, 8.5 cm in front
Variance allowed on overall measurement for cushions:+1 cm
Description of cushion cover:
• Heavy duty nylon (300g/m²), colour fast. • 2 ways stretch. • Breathable and machine washable. • Pull over elastic band with size label. • Colour: Preferably Black, alternately Navy Blue.
RT233-3.2.4 | WATERPROOF CUSHION COVER TO FIT WEDGED, THICK PRESSURE CARE FOAM CUSHION WITH CONTOURED BASE, EXTENDED LENGTH
Size range:
• 61 cm wide; 50 cm long; 6 cm high at rear, 11 cm in front
• 56 cm wide; 50 cm long; 6 cm high at rear, 11 cm in front
• 51 cm wide; 50 cm long; 6 cm high at rear, 11 cm in front
• 46 cm wide; 50 cm long; 6 cm high at rear, 11 cm in front
• 43 cm wide; 50 cm long; 6 cm high at rear, 11 cm in front
• 41 cm wide; 50 cm long; 6 cm high at rear, 11 cm in front
• 38 cm wide; 50 cm long; 6 cm high at rear, 11cm in front
• 36 cm wide; 50 cm long; 6 cm high at rear, 11 cm in front
• 30 cm wide; 50 cm long; 64 cm high at rear, 8.5 cm in front
Variance allowed on overall measurement for cushions: +1 cm
Description of cushion cover:
• Fabric, waterproof, terry cloth covered. • 2 ways stretch, breathable, anti-microbial shield treatment. • Machine washable. • Colour: Preferably Black, alternately Navy Blue.
RT233-3.2.5 | NYLON CUSHION COVER TO FIT THIN FOAM, BASIC POSITIONER CUSHION WITH CONTOURED BASE, EXTENDED LENGTH
Size range:
• 51 cm wide; 50 cm long; 7 cm high in front
• 46 cm wide; 50 cm long; 7 cm high in front
• 43 cm wide; 50 cm long; 7 cm high in front
• 41 cm wide; 50 cm long; 7 cm high in front
• 38 cm wide; 50 cm long; 7 cm high in front
• 36 cm wide; 50 cm long; 7 cm high in front
• 30 cm wide; 50 cm long; 6.5 cm high in front
Variance allowed on overall measurement for cushions: +1 cm
Description of cushion cover:
• Heavy duty nylon (300g/m²), colour fast. • 2 ways stretch. • Breathable and machine washable. • Pull over elastic band with size label. • Colour: Preferably Black, alternately Navy Blue.
RT233-3.2.6 | WATERPROOF CUSHION COVER TO FIT THIN FOAM, BASIC POSITIONER CUSHION WITH CONTOURED BASE, EXTENDED LENGTH
Size range:
• 51 cm wide; 50 cm long; 7 cm high in front
• 46 cm wide; 50 cm long; 7 cm high in front
• 43 cm wide; 50 cm long; 7 cm high in front
• 41 cm wide; 50 cm long; 7 cm high in front
• 38 cm wide; 50 cm long; 7 cm high in front
• 36 cm wide; 50 cm long; 7 cm high in front
• 30 cm wide; 50 cm long; 6.5 cm high in front
Variance allowed on overall measurement for cushions: +1 cm
Description of cushion cover:
• Fabric, waterproof, terry cloth covered. • 2 ways stretch, breathable, anti-microbial shield treatment. • Machine washable. • Colour: Preferably Black, alternately Navy Blue.
RT233-3.2.7 | NYLON CUSHION COVER TO FIT WEDGED, THIN FOAM, POSITIONER WITH CONTOURED BASE, EXTENDED LENGTH
Size range:
• 51 cm wide; 50 cm long; 4 cm high at rear, 9 cm in front
• 46 cm wide; 50 cm long; 4 cm high at rear, 9 cm in front
• 43 cm wide; 50 cm long; 4 cm high at rear, 9 cm in front
• 41 cm wide; 50 cm long; 4 cm high at rear, 9 cm in front
• 38 cm wide; 50 cm long; 4 cm high at rear, 9 cm in front
• 36 cm wide; 50 cm long; 4 cm high at rear, 9 cm in front
• 30 cm wide; 50 cm long; 4 cm high at rear, 8.5 cm in front
Variance allowed on overall measurement for cushions: +1 cm
Description of cushion cover:
• Heavy duty nylon (300g/m²), colour fast. • 2 ways stretch. • Breathable and machine washable. • Pull over elastic band with size label. • Colour: Preferably Black, alternately Navy Blue.
RT233-3.2.8 | WATERPROOF CUSHION COVER TO FIT WEDGED, THIN FOAM, POSITIONER WITH CONTOURED BASE, EXTENDED LENGTH
Size range:
• 51 cm wide; 50 cm long; 4 cm high at rear, 9 cm in front
• 46 cm wide; 50 cm long; 4 cm high at rear, 9 cm in front
• 43 cm wide; 50 cm long; 4 cm high at rear, 9 cm in front
• 41 cm wide; 50 cm long; 4 cm high at rear, 9 cm in front
• 38 cm wide; 50 cm long; 4 cm high at rear, 9 cm in front
• 36 cm wide; 50 cm long; 4 cm high at rear, 9 cm in front
• 30 cm wide; 50 cm long; 4 cm high at rear, 8.5 cm in front
Variance allowed on overall measurement for cushions: +1 cm
Description of cushion cover:
• Fabric, waterproof, terry cloth covered. • 2 ways stretch, breathable, anti-microbial shield treatment. • Machine washable. • Colour: Preferably Black, alternately Navy Blue.
RT233-3.3 | CUSHION COVERS FOR STRAIGHT BASE, BASIC FOAM | Indicate the item bidding for below bby stating "Yes bidding for this item". Item left blank will be assumed it is not bidded for | Complies Yes/No | Comments (Technical Specification Deviation)
RT233-3.3.1 | NYLON CUSHION COVER TO FIT STRAIGHT BASE, THICK PRESSURE CARE CUSHION WITH NO ABDUCTION BLOCK
Size Range:
• 20 cm wide, 25 cm long, 9cm high
• 22.5 cm wide, 25 cm long, 9cm high
• 25 cm wide, 35 cm long, 9cm high
• 27.5 cm wide, 40 cm long, 9cm high
• 30 cm wide, 35 cm long, 9cm high
• 30 cm wide, 40 cm long, , 9cm high
• 36 cm wide, 40 cm long, , 9cm high
• 36 cm wide , 50 cm long, 9cm high
• 38cm wide, 50cm long, 9cm high
• 40 cm wide, 50 cm long, 9cm high
• 43 cm wide, 50 cm long, 9cm high
• 45 cm wide, 50 cm long, 9cm high
• 51 cm wide, 50 cm long, 9cm high
Variance allowed on overall measurement for cushions: width: +1 cm
Description of cushion cover:
• Heavy duty nylon (300g/m²), colour fast. • 2 ways stretch. • Breathable and machine washable. • Pull over elastic band with size label. • Colour: Preferably Black, alternately Navy Blue.
RT233-3.3.2 | NYLON CUSHION COVER TO FIT STRAIGHT BASE, THICK PRESSURE CARE CUSHION WITH ABDUCTION BLOCK
Size Range:
• 20 cm wide, 25 cm long, 9cm high with extra-small abduction block
• 22.5 cm wide, 25 cm long, 9cm high with extra-small abduction block
• 25 cm wide, 35 cm long, 9cm high with extra-small abduction block
• 27.5 cm wide, 40 cm long, 9cm high with extra-small abduction block
• 30 cm wide, 35 cm long, 9cm high, with small abduction block
• 30 cm wide, 40 cm long, 9cm high, with small abduction block
• 35 cm wide, 40 cm long, 9cm high, with small abduction block
• 36 cm wide, 50 cm long, 9cm high, with small abduction block
• 38cm wide, 50cm long, 9cm high, with small abduction block
• 41 cm wide, 50 cm long, 9cm high with large abduction block
• 43 cm wide, 50 cm long, 9cm high with large abduction block
• 46 cm wide, 50 cm long, 9cm high with large abduction block
• 51 cm wide, 50 cm long, 9cm high with large abduction block
Variance allowed on overall measurement for cushions: width: +1 cm
Description of cushion cover:
• Heavy duty nylon (300g/m²), colour fast. • 2 ways stretch. • Breathable and machine washable. • Pull over elastic band with size label. • Colour: Preferably Black, alternately Navy Blue.
RT233-3.3.3 | WATERPROOF CUSHION COVER TO FIT STRAIGHT BASE, THICK PRESSURE CARE WITH NO ABDUCTION BLOCK
Size Range:
• 20 cm wide, 25 cm long, 9cm high
• 22.5 cm wide, 25 cm long, 9cm high
• 25 cm wide, 35 cm long, 9cm high
• 27.5 cm wide, 40 cm long, 9cm high
• 30 cm wide, 35 cm long, 9cm high
• 30 cm wide, 40 cm long, , 9cm high
• 36 cm wide, 40 cm long, , 9cm high
• 36 cm wide , 50 cm long, 9cm high
• 38cm wide, 50cm long, 9cm high
• 40 cm wide, 50 cm long, 9cm high
• 43 cm wide, 50 cm long, 9cm high
• 45 cm wide, 50 cm long, 9cm high
• 45 cm wide, 50 cm long, 9cm high
• 51 cm wide, 50 cm long, 9cm high
Variance allowed on overall measurement for cushions: width: +1 cm
Description of cushion cover:
• Fabric, waterproof, terry cloth covered. • 2 ways stretch, breathable, anti-microbial shield treatment. • Machine washable. • Colour: Preferably Black, alternately Navy Blue.
RT233-3.3.4 | WATERPROOF CUSHION COVER TO FIT STRAIGHT BASE, THICK PRESSURE CARE CUSHION WITH ABDUCTION BLOCK
Size Range:
• 20 cm wide, 25 cm long, 9cm high with extra-small abduction block
• 22.5 cm wide, 25 cm long, 9cm high with extra-small abduction block
• 25 cm wide, 35 cm long, 9cm high with extra-small abduction block
• 27.5 cm wide, 40 cm long, 9cm high with extra-small abduction block
• 30 cm wide, 35 cm long, 9cm high, with small abduction block
• 30 cm wide, 40 cm long, 9cm high, with small abduction block
• 35 cm wide, 40 cm long, 9cm high, with small abduction block
• 36 cm wide, 50 cm long, 9cm high, with small abduction block
• 38cm wide, 50cm long, 9cm high, with small abduction block
• 41 cm wide, 50 cm long, 9cm high with large abduction block
• 43 cm wide, 50 cm long, 9cm high with large abduction block
• 46 cm wide, 50 cm long, 9cm high with large abduction block
• 51 cm wide, 50 cm long, 9cm high with large abduction block
Variance allowed on overall measurement for cushions: width: +1 cm
Description of cushion cover:
• Fabric, waterproof, terry cloth covered. • 2 ways stretch, breathable, anti-microbial shield treatment. • Machine washable and can be dried in tumble drier. • Colour: Preferably Black, alternately Navy Blue.
RT233-3.3.5 | NYLON CUSHION COVER TO FIT STRAIGHT BASE, THIN FOAM POSITIONER WITH NO ABDUCTION BLOCK
Size Range:
• 20 cm wide, 25 cm long, 7cm high
• 22.5 cm wide, 25 cm long, 7cm high
• 25 cm wide, 35 cm long, 7cm high
• 27.5 cm wide, 40 cm long, 7cm high
• 30 cm wide, 35 cm long, 7cm high
• 30 cm wide, 40 cm long, 7cm high
• 35 cm wide, 40 cm long, 7cm high
• 36 cm wide, 50 cm long, 7cm high
• 38 cm wide, 50 cm long, 7 cm high
• 41 cm wide, 50 cm long, 7cm high
• 43 cm wide, 50 cm long, 7cm high
• 46 cm wide, 50 cm long, 7cm high
• 51 cm wide, 50 cm long, 7cm high
Variance allowed on overall measurement for cushions: width: +1 cm
Description of cushion cover:
• Heavy duty nylon (300g/m²), colour fast. • 2 ways stretch. • Breathable and machine washable. • Pull over elastic band with size label. • Colour: Preferably Black, alternately Navy Blue.
RT233-3.3.6 | NYLON CUSHION COVER TO FIT STRAIGHT BASE, THIN FOAM POSITIONER WITH ABDUCTION BLOCK
Size Range:
• 20 cm wide, 25 cm long, 7cm high with extra-small abduction block
• 22.5 cm wide, 25 cm long, 7cm high with extra-small abduction block
• 25 cm wide, 35 cm long, 7cm high with extra-small abduction block
• 27.5 cm wide, 40 cm long, 7cm high with extra-small abduction block
• 30 cm wide, 35 cm long, 7cm high with small abduction block
• 30 cm wide, 40 cm long, 7cm high with small abduction block
• 35 cm wide, 40 cm long, 7cm high with small abduction block
• 36 cm wide, 50 cm long, 7cm high with small abduction block
• 38 cm wide, 50 cm long, 7 cm high with small abduction block
• 41 cm wide, 50 cm long, 7cm high with large abduction block
• 43 cm wide, 50 cm long, 7cm high with large abduction block
• 46 cm wide, 50 cm long, 7cm high with large abduction block
• 51 cm wide, 50 cm long, 7cm high with large abduction block
Variance allowed on overall measurement for cushions: width: +1 cm
Description of cushion cover:
• Heavy duty nylon (300g/m²), colour fast. • 2 ways stretch. • Breathable and machine washable. • Pull over elastic band with size label. • Colour: Preferably Black, alternately Navy Blue.
RT233-3.3.7 | WATERPROOF CUSHION COVER TO FIT STRAIGHT BASE, THIN FOAM POSITIONER WITH NO ABDUCTION BLOCK
Size Range:
• 20 cm wide, 25 cm long, 7cm high
• 22.5 cm wide, 25 cm long, 7cm high
• 22.5 cm wide, 25 cm long, 7cm high
• 25 cm wide, 35 cm long, 7cm high
• 27.5 cm wide, 40 cm long, 7cm high
• 30 cm wide, 35 cm long, 7cm high
• 30 cm wide, 40 cm long, 7cm high
• 35 cm wide, 40 cm long, 7cm high
• 36 cm wide, 50 cm long, 7cm high
• 38 cm wide, 50 cm long, 7 cm high
• 41 cm wide, 50 cm long, 7cm high
• 43 cm wide, 50 cm long, 7cm high
• 46 cm wide, 50 cm long, 7cm high
• 51 cm wide, 50 cm long, 7cm high
• 51 cm wide, 50 cm long, 7cm high
• 51 cm wide, 50 cm long, 7cm high
Variance allowed on overall measurement for cushions: width: +1 cm
Description of cushion cover:
• Fabric, waterproof, terry cloth covered. • 2 ways stretch, breathable, anti-microbial shield treatment. • Machine washable. • Colour: Preferably Black, alternately Navy Blue
RT233-3.3.8 | WATERPROOF COVER TO FIT STRAIGHT BASE, THIN FOAM POSITIONER WITH ABDUCTION BLOCK
Size Range:
• 20 cm wide, 25 cm long, 7cm high with extra-small abduction block
• 22.5 cm wide, 25 cm long, 7cm high with extra-small abduction block
• 25 cm wide, 35 cm long, 7cm high with extra-small abduction block
• 27.5 cm wide, 40 cm long, 7cm high with extra-small abduction block
• 30 cm wide, 35 cm long, 7cm high with small abduction block
• 30 cm wide, 40 cm long, 7cm high with small abduction block
• 35 cm wide, 40 cm long, 7cm high with small abduction block
• 36 cm wide, 50 cm long, 7cm high with small abduction block
• 38 cm wide, 50 cm long, 7 cm high with small abduction block
• 41 cm wide, 50 cm long, 7cm high with large abduction block
• 43 cm wide, 50 cm long, 7cm high with large abduction block
• 46 cm wide, 50 cm long, 7cm high with large abduction block
• 51 cm wide, 50 cm long, 7cm high with large abduction block
Variance allowed on overall measurement for cushions: width: +1 cm
Description of cushion cover:
• Fabric, waterproof, terry cloth covered. • 2 ways stretch, breathable, anti-microbial shield treatment. • Machine washable. • Colour: Preferably Black, alternately Navy Blue
RT233-3.4 | CUSHION COVER FABRIC | Indicate the item bidding for below bby stating "Yes bidding for this item". Item left blank will be assumed it is not bidded for | Complies Yes/No | Comments (Technical Specification Deviation)
RT233-3.4.1 | CUSHION COVER FABRIC, BASIC NYLON STRETCHABLE
Fabric, heavy-duty nylon (300 g/m2), colour fast, 2-way stretch, breathable, and machine washable, tube at least 55 cm wide, Colour: Preferably in Black alternately in Navy Blue. Order per meter.
RT233-3.4.2 | CUSHION COVER FABRIC, WATERPROOF, STRETCHABLE
Fabric, waterproof, terry cloth cotton covered, 2-way stretch, breathable, anti-microbial shield treatment, machine washable, and 150 cm wide. Colour: Preferably in Black alternately in Navy Blue. Order per meter.
RT233-3.5 | FLUID FILLED CUSHION COVER | Indicate the item bidding for below bby stating "Yes bidding for this item". Item left blank will be assumed it is not bidded for | Complies Yes/No | Comments (Technical Specification Deviation)
• Separate removable base cover. • Separate removable base cover waterproof. • Separate removable external cover. External cover to allow moulding with fluid pad without interfering with pressure relieving properties.
RT233-3.6 | AIR FILLED CUSHION COVERS | Indicate the item bidding for below bby stating "Yes bidding for this item". Item left blank will be assumed it is not bidded for | Complies Yes/No | Comments (Technical Specification Deviation)
• Separate stretchable, removable cover. Cover machine washable. • Non-slip bottom of cover with hook and loop strips to attach cushion to wheelchair. • Type, size, and direction of cushion use printed on cover.

## 7 Category 4
CATEGORY 4
All items being tendered for must comply with the following technical specification, failure to comply with any of the conditions set out below may result in immediate disqualification. Please note that where the technical specification calls for certification, this certification must accompany the bid, failure to provide such certification will result in immediate disqualification. Bidders must give details of items being tendered for in the COMMENTS section where there are variation for further deliberation
WHEELCHAIR CUSHION COVERS
RT233-4.1 | BASIC FOAM CUSHIONS WITH NYLON COVERS | Indicate the item bidding for below bby stating "Yes bidding for this item". Item left blank will be assumed it is not bidded for | Complies Yes/No | Comments (Technical Specification Deviation)
TECHNICAL SPECIFICATIONS
RT233-4.1.1 | CONTOURED BASE THICK PRESSURE CARE CUSHION WITH NYLON COVER
Variance allowed on overall measurement +1 cm
Size range:
• 61 cm wide, 43 cm long, 9 cm high in front
• 56 cm wide, 43 cm long, 9 cm high in front
• 51 cm wide, 43 cm long, 9 cm high in front
• 46 cm wide, 43 cm long, 9 cm high in front
• 41 cm wide, 43 cm long, 9 cm high in front
• 38 cm wide, 43 cm long, 9 cm high in front
• 36 cm wide, 43 cm long, 9 cm high in front
• 30 cm wide, 43 cm long, 9 cm high in front
• 30 cm wide, 30 cm long, 6.5cm high in front
• 25 cm wide, 30 cm long, 6 cm high in front
Description of Cushion:
• Wheelchair basic foam pressure relieving cushion, Bevelled contoured base to match sling of seat of wheelchair • Must have trochanteric loading pads to create a well • All cushions not to be glued at the back • The depth of the well for 14”(36cm) – 24”(61cm) should be 3 cm, for 12”(30cm) should be 2.5cm, and 10”(25cm) should be 2cm. • The width of the well = 50% the width of the cushion • Length of the well = the width of the well up to a maximum of 20cm. • Sizes 14”-24”: Top layer – 5cm thick high density foam following the contour of the base through the full length of the cushion, only the front half of the cushion to be glued • Type, size and direction of use printed on bottom of base
Material:
Top layer:
For sizes: 25cm - 46cm High-density foam 36/20; Hardness factor 18-21.6; Density range 32.1-40kg/m³
For sizes: 51cm-61cm High density foam 50/30 Hardness factor: 249-291 and Density: 37.1- 41kg/m³
Base layer:
High density compressed (reconstituted flexible PU) foam 1300; Hardness factor-52-63; Density 130+10% kg/m³ (SANS 642 - 1976)
Description of cushion cover:
• Heavy duty nylon (300g/m²), colour fast. • 2 ways stretch. • Breathable and machine washable. • Pull over elastic band with size label. • Colour: Preferably Black, alternately Navy Blue
RT233-4.1.2 | CONTOURED BASE - WEDGED, THICK PRESSURE CARE CUSHION WITH NYLON COVER
Variance allowed on overall measurement: +1 cm Variance allowed on front height: +0.5cm
Size range:
• 61 cm wide; 43 cm long; 6 cm high at rear, 11 cm high in front
• 56 cm wide; 43 cm long; 6 cm high at rear, 11 cm high in front
• 51 cm wide; 43 cm long; 6 cm high at rear, 11 cm high in front
• 46 cm wide; 43 cm long; 6 cm high at rear, 11 cm high in front
• 41 cm wide; 43 cm long; 6 cm high at rear, 11 cm high in front
• 38 cm wide; 43 cm long; 6 cm high at rear, 11 cm high in front
• 36 cm wide; 43 cm long; 6 cm high at rear, 11 cm high in front
• 30 cm wide; 43 cm long; 6 cm high at rear, 11 cm high in front
• 30 cm wide; 30 cm long; 4 cm high at rear, 8.5 cm high in front
• 25 cm wide; 30 cm long; 4 cm high at rear, 7.5 cm high in front
Description of cushion:
• Wheelchair basic foam pressure relieving cushion. Bevelled contoured base to match sling seat of wheelchair. • Must have trochanteric loading pads to create a well. • The depth of the well for 14” (36cm) - 24” (61cm) should be 3cm, 12” (30cm) should be 2,5cm and 10’ (2cm) . • The width of the well = 50% of the width of the cushion. • Length of the well = the width of the well up to maximum 20cm. • 2cm wedge in front of pre-ischial shelf extending to the front of the cushion. • Sizes 14”-24”: Top layer - 5cm thick high-density foam following contour of base through full length of cushion; only the front half of the cushion to be glued. • Size 10” -12”: Top - 3 cm thick high-density foam following contour of base through full length of cushion; only the front half of the cushion to be glued. • Type, size, and direction of cushion use printed on bottom of base.
Material:
Top layer:
For sizes: 25cm – 46 cm High-density foam 36/20; Hardness factor 18 – 21.6; Density range 32.1-40kg/m³
For sizes: 51cm -61cm High density foam 50/30 Hardness factor: 249 - 291 and Density: 37.1- 41kg/m³
Base layer:
High density compressed (reconstituted flexible PU) foam 1300; Hardness factor 52-63; Density 130+10% kg/m³ (SANS 642 - 1976)
Description of cushion cover:
• Heavy duty nylon (300g/m²), colour fast. • 2 ways stretch. • Breathable and machine washable. • Pull over elastic band with size label. • Colour: Preferably Black, alternately Navy Blue
RT233-4.1.3 | CONTOURED BASE - THIN FOAM, BASIC POSITIONER CUSHION WITH NYLON COVER
Variance allowed on overall measurement : +1 cm
Size range:
• 51 cm wide; 43 cm long; 7 cm high
• 46 cm wide; 43 cm long; 7 cm high
• 43 cm wide; 43 cm long; 7 cm high
• 41 cm wide; 43 cm long; 7 cm high
• 38 cm wide; 43 cm long; 7 cm high
• 36 cm wide; 43 cm long; 7 cm high
• 30 cm wide; 43 cm long; 6.5 cm high
• 30 cm wide; 33 cm long; 6.5 cm high
• 25 cm wide; 28 cm long; 6 cm high
Description of cushion:
• Bevelled contoured base to match sling seat of wheelchair. Sitting surface with pre-ischial shelf. • The depth of the pre-ischial shelf for 14” (36cm) - 20 (51cm) should be 3cm minimum, for 12” (30cm) should be 2,5cm minimum, and for 10” (25cm) should be 2cm minimum. • The distance of the shelf from the rear of the cushion = 50% of the width of the cushion up to maximum of 20cm. • Top: 3cm thick high-density foam following contour of base through full length of cushion; only the front half of the cushion to be glued. • Type, size, and direction of cushion use printed on bottom of base.
Material:
Top layer:
For sizes: 25cm - 48cm High-density foam 36/20; Hardness factor 18 – 21.6; Density range 32.1-40kg/m³
For sizes: 51cm High density foam 50/30 Hardness factor: 249 - 291 and Density: 37.1- 41kg/m³
Base layer:
High density compressed (reconstituted flexible PU) foam 1300; Hardness factor 52-63; Density 130+10% kg/m³ (SANS 642 - 1976)
Description of cushion cover:
• Heavy duty nylon (300g/m²), colour fast. • 2 ways stretch. • Breathable and machine washable. • Pull over elastic band with size label. • Colour: Preferably Black, alternately Navy Blue
RT233-4.1.4 | CONTOURED BASE - WEDGED, THIN FOAM POSITIONER WITH NYLON COVER
Variance allowed on overall measurement: +1cm Variance allowed on front height:+0.5cm
Size range:
• 51 cm wide; 43 cm long; 4 cm high at rear, 9 cm high in front
• 46 cm wide; 43 cm long; 4 cm high at rear, 9 cm high in front
• 43 cm wide; 43 cm long; 4 cm high at rear, 9 cm high in front
• 41 cm wide; 43 cm long; 4 cm high at rear, 9 cm high in front
• 38 cm wide; 43 cm long; 4 cm high at rear, 9 cm high in front
• 36 cm wide; 43 cm long; 4 cm high at rear, 9 cm high in front
• 30 cm wide; 43 cm long; 4 cm high at rear, 9 cm high in front
• 30 cm wide; 33 cm long; 4 cm high at rear, 8 cm high in front
• 25 cm wide; 28 cm long; 4 cm high at rear, 7.5cm high in front
Description of cushion:
• Bevelled contoured base to match sling seat of wheelchair. Sitting surface with pre-ischial shelf. • The depth of the pre-ischial shelf for 14” (36cm) - 20” (51cm) should be 3cm minimum, 12” (30cm) should be 2,5cm minimum, and 10” (25cm) should be 2cm minimum. • The distance of the shelf from the rear = 50% of the width of the cushion up to maximum of 20cm. • 2cm wedge in front of pre-ischial shelf extending to the front of the cushion. • Top: 3cm thick high-density foam following contour of base through full length of cushion; only the front half of the cushion to be glued. • Type, size, and direction of cushion use printed on bottom of base
Material:
Top layer:
For sizes: 25cm-46cm High-density foam 36/20; Hardness factor 18 – 21.6; Density range 32.1-40kg/m³
For sizes: 51cm High density foam 50/30 Hardness factor: 249 - 291 and Density: 37.1- 41kg/m³
Base layer:
High density compressed (reconstituted flexible PU) foam 1300; Hardness factor 52-63; Density 130+10% kg/m³ (SANS 642 - 1976)
Description of cushion cover:
• Heavy duty nylon (300g/m²), colour fast. • 2 ways stretch. • Breathable and machine washable. • Pull over elastic band with size label. • Colour: Preferably Black, alternately Navy Blue
RT233-4.1.5 | CONTOURED BASE - THICK PRESSURE CARE CUSHION, EXTENDED LENGTH WITH NYLON COVER
Variance allowed on overall measurement: +1cm
Size range:
• 61 cm wide; 50 cm long; 9 cm high
• 56 cm wide; 50 cm long; 9 cm high
• 51 cm wide; 50 cm long; 9 cm high
• 46 cm wide; 50 cm long; 9 cm high
• 43 cm wide; 50 cm long; 9 cm high
• 41 cm wide; 50 cm long; 9 cm high
• 38 cm wide; 50 cm long; 9 cm high
• 36 cm wide; 50 cm long; 8.5 cm high
• 30 cm wide; 50 cm long; 8.5 cm high
Description of cushion:
• Wheelchair basic foam pressure relieving cushion. Bevelled contoured base to match sling seat of wheelchair. • Must have trochanteric loading pads to create a well. • The depth of the well for 14” (36cm) - 24” (61cm) should be 3cm, for 12” (30cm) should be 2,5cm. • The width of the well = 50% of the width of the cushion. • Length of the well = the width of the well up to maximum 20cm. • Size 14”-24”, 5cm thick high-density foam top layer following contour of base through full length of cushion; only the front half of the cushion to be glued. • Size 12”, 3cm thick high-density foam top layer, following contour of base through full length of cushion; only the front half of the cushion to be glued. • Type, size, and direction of cushion use printed on bottom of base.
Material:
Top layer:
For sizes: 30cm-46cm High-density foam 36/20; Hardness factor 18 – 21.6; Density range 32.1-40kg/m³
For sizes: 51cm-61cm High density foam 50/30 Hardness factor: 249 - 291 and Density: 37.1- 41kg/m³
Base layer:
High density compressed (reconstituted flexible PU) foam 1300; Hardness factor 52-63 Density 130+10% kg/m³ (SANS 642 - 1976)
Description of cushion cover:
• Heavy duty nylon (300g/m²), colour fast. • 2 ways stretch. • Breathable and machine washable. • Pull over elastic band with size label. • Colour: Preferably Black, alternately Navy Blue
RT233-4.1.6 | CONTOURED BASE - WEDGED, THICK PRESSURE CARE CUSHION, EXTENDED LENGTH WITH NYLON COVER
Variance allowed on overall measurement:+1cm
Size range:
• 61 cm wide; 50 cm long; 6 cm high at rear, 11 cm high in front
• 56 cm wide; 50 cm long; 6 cm high at rear, 11 cm high in front
• 51 cm wide; 50 cm long; 6 cm high at rear, 11 cm high in front
• 46 cm wide; 50 cm long; 6 cm high at rear, 11 cm high in front
• 43 cm wide; 50 cm long; 6 cm high at rear, 11 cm high in front
• 41 cm wide; 50 cm long; 6 cm high at rear, 11 cm high in front
• 38 cm wide; 50 cm long; 6 cm high at rear, 11cm high in front
• 36 cm wide; 50 cm long; 6 cm high at rear, 11 cm high in front
• 30 cm wide; 50 cm long; 6 cm high at rear, 10.5 cm high in front
Description of cushion:
• Wheelchair basic foam pressure relieving cushion. Bevelled contoured base to match sling seat of wheelchair. • Must have trochanteric loading pads to create a well. • The depth of the well for 14” (36cm) - 24” (61cm) should be 3cm, 12” (30cm) should be 2,5cm. • The width of the well = 50% of the width of the cushion. • Length of the well = the width of the well up to maximum 20cm. • Size 14”-24”, 5cm thick high-density foam top layer, following contour of base; only the front half of the cushion to be glued. • Size 12”, 3cm, thick high density foam top layer, following contour of base through full length of cushion, only the front half of the cushion to be glued. • Type, size, and direction of cushion use printed on bottom of base.
Material:
Top layer:
For sizes: 30cm-46cm High-density foam 36/20; Hardness factor 18 – 21.6; Density range 32.1-40kg/m³
For sizes: 51cm-61cm High density foam 50/30 Hardness factor: 249 - 291 and Density: 37.1- 41kg/m³
Base layer:
High density compressed (reconstituted flexible PU) foam 1300; Hardness factor 52-63 Density 130+10% kg/m³ (SANS 642 - 1976)
Description of cushion cover:
• Heavy duty nylon (300g/m²), colour fast. • 2 ways stretch. • Breathable and machine washable. • Pull over elastic band with size label. • Colour: Preferably Black, alternately Navy Blue
RT233-4.1.7 | CONTOURED BASED - THIN FOAM, BASIC POSITIONER CUSHION, EXTENDED LENGTH WITH NYLON COVER
Variance allowed on overall measurement: +1cm
Size range:
• 51 cm wide; 50 cm long; 7 cm high in front
• 46 cm wide; 50 cm long; 7 cm high in front
• 43 cm wide; 50 cm long; 7 cm high in front
• 41 cm wide; 50 cm long; 7 cm high in front
• 38 cm wide; 50 cm long; 7 cm high in front
• 36 cm wide; 50 cm long; 7 cm high in front
• 30 cm wide; 50 cm long; 7 cm high in front
Description of cushion:
• Bevelled contoured base to match sling seat of wheelchair. Sitting surface with pre-ischial shelf. • The depth of the pre-ischial shelf for 14” (36cm) - 20” (51cm) should be 3cm minimum; 12” (30cm) should be 2, 5cm minimum. • The distance of the shelf from the rear of the cushion = 50% of the width of the cushion up to maximum of 20cm. • Top: 3cm thick high-density foam following contour of base through full length of cushion; only the front half of the cushion to be glued. • Type, size, and direction of cushion use printed on bottom of base.
Material:
Top layer:
For Sizes: 30cm -46cm High-density foam 36/20; Hardness factor 18 – 21.6; Density range 32.1-40kg/m³
For sizes: 51cm High density foam 50/30 Hardness factor: 249 - 291 and Density: 37.1- 41kg/m³
Base layer:
High density compressed (reconstituted flexible PU) foam 1300; Hardness factor 52-63; Density 130+10% kg/m³ (SANS 642 - 1976)
Description of cushion cover:
• Heavy duty nylon (300g/m²), colour fast. • 2 ways stretch. • Breathable and machine washable. • Pull over elastic band with size label. • Colour: Preferably Black, alternately Navy Blue
RT233-4.1.8 | CONTOURED BASE – WEDGED THIN FOAM POSITIONER, EXTENDED LENGTH WITH NYLON COVER
Variance allowed on overall measurement: +1 cm
Size range:
• 51 cm wide; 50 cm long; 4 cm high at rear, 9 cm in front
• 46 cm wide; 50 cm long; 4 cm high at rear, 9 cm in front
• 43 cm wide; 50 cm long; 4 cm high at rear, 9 cm in front
• 41 cm wide; 50 cm long; 4 cm high at rear, 9 cm in front
• 38 cm wide; 50 cm long; 4 cm high at rear, 9 cm in front
• 36 cm wide; 50 cm long; 4 cm high at rear, 9 cm in front
• 30 cm wide; 50 cm long; 4 cm high at rear, 8.5 cm in front
Description of cushion:
• Bevelled contoured base to match sling seat of wheelchair. Sitting surface with pre-ischial shelf. • The depth of the pre-ischial shelf for 14” (36cm) - 20” (51cm) should be 3cm minimum, 12” (30cm) should be 2, 5 cm minimum. • The distance of the shelf from the rear of the cushion = 50% of the width of the cushion up to maximum of 20cm. • 2cm wedge in front of pre-ischial shelf extending to the front of the cushion. • Top: 3cm thick high-density foam following contour of base through full length of cushion; only the front half of the cushion to be glued. • Type, size, and direction of cushion use printed on bottom of base
Material:
Top layer:
For sizes: 30cm -46cm High-density foam 36/20; Hardness factor 18 – 21.6; Density range 32.1-40kg/m³
For sizes: 51cm High density foam 50/30 Hardness factor: 249 - 291 and Density: 37.1- 41kg/m³
Base layer:
High density compressed (reconstituted flexible PU) foam 1300; Hardness factor 52-63; Density 130+10% kg/m³ (SANS 642 - 1976)
Description of cushion cover:
• Heavy duty nylon (300g/m²), colour fast. • 2 ways stretch. • Breathable and machine washable. • Pull over elastic band with size label. • Colour: Preferably Black, alternately Navy Blue
RT233-4.1.9 | NON-CONTOURED BASE - STRAIGHT BASE, BASIC FOAM, THICK PRESSURE CARE CUSHION WITH NO ABDUCTION BLOCK WITH NYLON COVER
Variance allowed on overall measurement: +1cm
Size Range:
• 20 cm wide, 25 cm long, 9cm high
• 22.5 cm wide, 25 cm long, 9cm high
• 25 cm wide, 35 cm long, 9cm high
• 27.5 cm wide, 40 cm long, 9cm high
• 30 cm wide, 35 cm long, 9cm high
• 30 cm wide, 40 cm long, , 9cm high
• 36 cm wide, 40 cm long, , 9cm high
• 36 cm wide , 50 cm long, 9cm high
• 38cm wide, 50cm long, 9cm high
• 40 cm wide, 50 cm long, 9cm high
• 43 cm wide, 50 cm long, 9cm high
• 45 cm wide, 50 cm long, 9cm high
• 51 cm wide, 50 cm long, 9cm high
Description of cushion:
• Wheelchair basic foam pressure relieving cushion. Flat base to match solid wheelchair seat. • Must have trochanteric loading pads to create a well. • The depth of the well for 14” (36cm) - 20”(51cm) should be 3cm, for 8” (20cm) - 12”(30cm) should be 2,5cm. • The width of the well = 50% of the width of the cushion • The length of the well = the width of the well up to maximum 20cm. • Base layer: Minimum thickness 4 cm for sizes 14”-20” and minimum thickness 3cm for 12” and minimum 2,5cm for 10” and smaller. • The thickness of the top layer for 14” (36cm) -20” (51cm) should be 5cm minimum and for 8” (20cm) -12” (30cm ) should be 3cm; only the front half of the cushion to be glued. • Type, size, and direction of cushion use printed on bottom of base.
Material:
Top layer:
For sizes to fit: 8”-18” High-density foam 36/20; Hardness factor 18 – 21.6; Density range 32.1-40kg/m³
For sizes: 20” High density foam 50/30 Hardness factor: 249 - 291 and Density: 37.1- 41kg/m³
Base layer:
High density compressed (reconstituted flexible PU) foam 1300 Hardness factor 52-63; Density 130+10% kg/m³ (SANS 642 - 1976)
Description of cushion cover:
• Heavy duty nylon (300g/m²), colour fast. • 2 ways stretch. • Breathable and machine washable. • Pull over elastic band with size label. • Colour: Preferably Black, alternately Navy Blue
RT233-4.1.10 | NON-CONTOURED BASE - STRAIGHT BASE, THICK PRESSURE CARE CUSHION WITH ABDUCTION BLOCK AND NYLON COVER
Variance allowed on overall measurement: +1cm
Size Range:
· 20 cm wide, 25 cm long, 9cm high with extra-small abduction block
· 22.5 cm wide, 25 cm long, 9cm high with extra-small abduction block
· 25 cm wide, 35 cm long, 9cm high with extra-small abduction block
· 27.5 cm wide, 40 cm long, 9cm high with extra-small abduction block
· 30 cm wide, 35 cm long, 9cm high, with small abduction block
· 30 cm wide, 40 cm long, 9cm high, with small abduction block
· 35 cm wide, 40 cm long, 9cm high, with small abduction block
· 36 cm wide, 50 cm long, 9cm high, with small abduction block
· 38cm wide, 50cm long, 9cm high, with small abduction block
· 41 cm wide, 50 cm long, 9cm high with large abduction block
· 43 cm wide, 50 cm long, 9cm high with large abduction block
· 46 cm wide, 50 cm long, 9cm high with large abduction block
· 51 cm wide, 50 cm long, 9cm high with large abduction block
Description of cushion:
| • Wheelchair basic foam pressure relieving cushion. Flat base to match solid wheelchair seat. • Must have trochanteric loading pads to create a well. • The depth of the well for 14” (36cm) - 20”(51cm) should be 3cm, for 8” (20cm) - 12”(30cm) should be 2,5cm. • The width of the well = 50% of the width of the cushion • The length of the well = the width of the well up to maximum 20cm. • Base layer: Minimum thickness 4 cm for sizes 14”-20” and minimum thickness 3cm for 12” and minimum 2,5cm for 10” and smaller. • The thickness of the top layer for 14” (36cm) -20” (51cm) should be 5cm minimum and for 8” (20cm) -12” (30cm ) should be 3cm; only the front half of the cushion to be glued. • The pommel should extend beyond the end of the cushion to reach the front of the user’s knees. • The height of the pommel = the full thickness of the cushion from the base and includes average height of user’s thigh in the front. • The length of the pommel should not be more than the distal third of the femur. • Type, size, and direction of cushion use printed on bottom of base. |  |  | No specific measurements of the pommel provided
Material:
Top layer:
For sizes to fit: 8”-18” High-density foam 36/20; Hardness factor 18 – 21.6; Density range 32.1-40kg/m³
For sizes: 20” High density foam 50/30 Hardness factor: 249 - 291 and Density: 37.1- 41kg/m³
Base layer:
High density compressed (reconstituted flexible PU) foam 1300 Hardness factor 52-63; Density 130+10% kg/m³ (SANS 642 - 1976)
Description of cushion cover:
• Heavy duty nylon (300g/m²), colour fast. • 2 ways stretch. • Breathable and machine washable. • Pull over elastic band with size label. • Colour: Preferably Black, alternately Navy Blue
RT233-4.1.11 | NON-CONTOURED BASE - STRAIGHT BASE, POSITIONER WITH NO ABDUCTION BLOCK WITH NYLON COVER
Variance allowed on overall measurement: +1cm
Size Range:
20 cm wide, 25 cm long, 7cm high
22.5 cm wide, 25 cm long, 7cm high
25 cm wide, 35 cm long, 7cm high
27.5 cm wide, 40 cm long, 7cm high
30 cm wide, 35 cm long, 7cm high
30 cm wide, 40 cm long, 7cm high
36 cm wide, 40 cm long, 7cm high
36 cm wide, 50 cm long, 7cm high
38cm wide, 50cm long, 7 cm high
40 cm wide, 50 cm long, 7cm high
43 cm wide, 50 cm long, 7cm high
45 cm wide, 50 cm long, 7cm high
51 cm wide, 50 cm long, 7cm high
Description of cushion:
• Wheelchair basic foam positioner cushion. Flat base to match solid wheelchair seat • Sitting surface with pre-ischial shelf. • The depth of the pre-ischial shelf for 14” (36cm) - 20” (51cm) should be 3cm minimum, 12” (30cm) should be 2,5 cm minimum. • The distance of the shelf from the rear of the cushion = 50% of the width of the cushion up to maximum of 20cm. • Base layer: Minimum thickness 4cm for sizes 14”-20” and minimum thickness 3cm for 12” and minimum 2,5cm for 10” and smaller. • The thickness of the top layer for 14” (36cm) -20” (51cm) should be 3cm minimum and for 8” (20cm) -12” (30cm ) should be 2.5cm.; only the front half of the cushion to be glued. • Type, size, and direction of cushion use printed on bottom of base.
Material:
Top layer:
For sizes to fit: 8”-18” High-density foam 36/20; Hardness factor 18 – 21.6; Density range 32.1-40kg/m³
For sizes: 20” High density foam 50/30 Hardness factor: 249 - 291 and Density: 37.1- 41kg/m³
Base layer:
High density compressed (reconstituted flexible PU) foam 1300 Hardness factor 52-63: Density 130+10% kg/m³ (SANS 642 - 1976)
Description of cushion cover:
• Heavy duty nylon (300g/m²), colour fast. • 2 ways stretch. • Breathable and machine washable. • Pull over elastic band with size label. • Colour: Preferably Black, alternately Navy Blue
RT233-4.1.12 | NON-CONTOURED BASE - STRAIGHT BASE, POSITIONER WITH ABDUCTION BLOCK WITH NYLON COVER
Variance allowed on overall measurement: +1cm
Size Range:
• 20 cm wide, 25 cm long, 7cm high with extra-small abduction block
• 22.5 cm wide, 25 cm long, 7cm high with extra-small abduction block
• 25 cm wide, 35 cm long, 7cm high with extra-small abduction block
• 27.5 cm wide, 40 cm long, 7cm high with extra-small abduction block
• 30 cm wide, 35 cm long, 7cm high with small abduction block
• 30 cm wide, 40 cm long, 7cm high with small abduction block
• 35 cm wide, 40 cm long, 7cm high with small abduction block
• 36 cm wide, 50 cm long, 7cm high with small abduction bloc
• 41 cm wide, 50 cm long, 7cm high with large abduction block
• 43 cm wide, 50 cm long, 7cm high with large abduction block
• 46 cm wide, 50 cm long, 7cm high with large abduction block
• 51 cm wide, 50 cm long, 7cm high with large abduction block
Description:
| • Wheelchair basic foam positioner cushion. Flat base to match solid wheelchair seat • Sitting surface with pre-ischial shelf. • The depth of the well for 14” (36cm) - 20”(51cm) should be 30mm, for 12”(30cm) should be 2,5cm. • The distance of the shelf from the rear of the cushion = 50% of the width of the cushion up to maximum of 20cm. • Base layer: Minimum thickness 4cm for sizes 14”-20” and minimum thickness 3cm for 12” and minimum 2,5cm for 10” and smaller. • The thickness of the top layer for 14” (36cm) -20” (51cm) should be 3cm minimum and for 8” (20cm) -12” (30cm ) should be 2.5cm.; only the front half of the cushion to be glued. • The pommel should extend beyond the end of the cushion to reach the front of the user’s knees. • The height of the pommel = the full thickness of the cushion from the base and includes average height of user’s thigh in the front. • The length of the pommel should not be more than the distal third of the femur. • Type, size, and direction of cushion use printed on bottom of base. |  |  | No specific measurements of the pommel provided
Material:
Top layer:
For sizes to fit: 8”-18” High-density foam 36/20; Hardness factor 18 – 21.6; Density range 32.1-40kg/m³
For sizes: 20” High density foam 50/30 Hardness factor: 249 - 291 and Density: 37.1- 41kg/m³
Base layer:
High density compressed (reconstituted flexible PU) foam 1300 Hardness factor 52-63: Density 130+10% kg/m³ (SANS 642 - 1976)
Description of cushion cover:
• Heavy duty nylon (300g/m²), colour fast. • 2 ways stretch. • Breathable and machine washable. • Pull over elastic band with size label. • Colour: Preferably Black, alternately Navy Blue
RT233-4.2 | BASIC FOAM CUSHIONS WITH WATERPROOF COVERS | Indicate the item bidding for below bby stating "Yes bidding for this item". Item left blank will be assumed it is not bidded for | Complies Yes/No | Comments (Technical Specification Deviation)
RT233-4.2.1 | CONTOURED BASE THICK PRESSURE CARE CUSHION WITH WATERPROOF COVER
Variance allowed on overall measurement +1 cm
Size range:
• 61 cm wide, 43 cm long, 9 cm high in front
• 56 cm wide, 43 cm long, 9 cm high in front
• 51 cm wide, 43 cm long, 9 cm high in front
• 46 cm wide, 43 cm long, 9 cm high in front
• 41 cm wide, 43 cm long, 9 cm high in front
• 38 cm wide, 43 cm long, 9 cm high in front
• 36 cm wide, 43 cm long, 9 cm high in front
• 30 cm wide, 43 cm long, 9 cm high in front
• 30 cm wide, 30 cm long, 6.5cm high in front
• 25 cm wide, 30 cm long, 6 cm high in front
Description of Cushion:
• Wheelchair basic foam pressure relieving cushion, Bevelled contoured base to match sling of seat of wheelchair • Must have trochanteric loading pads to create a well • All cushions not to be glued at the back • The depth of the well for 14”(36cm) – 24”(61cm) should be 3 cm, for 12”(30cm) should be 2.5cm, and 10”(25cm) should be 2cm. • The width of the well = 50% the width of the cushion • Length of the well = the width of the well up to a maximum of 20cm. • Sizes 14”-24”: Top layer – 5cm thick high density foam following the contour of the base through the full length of the cushion, only the front half of the cushion to be glued • Type, size and direction of use printed on bottom of base
Material:
Top layer:
For sizes: 25cm - 46cm High-density foam 36/20; Hardness factor 18-21.6; Density range 32.1-40kg/m³
For sizes: 51cm-61cm High density foam 50/30 Hardness factor: 249-291 and Density: 37.1- 41kg/m³
Base layer:
High density compressed (reconstituted flexible PU) foam 1300; Hardness factor-52-63; Density 130+10% kg/m³ (SANS 642 - 1976)
Description of cushion cover:
• Fabric, waterproof, terry cloth covered. • 2 ways stretch, breathable, anti-microbial shield treatment. • Machine washable. • Pull over elastic band with size label. • Colour: Preferably Black, alternately Navy Blue.
RT233-4.2.2 | CONTOURED BASE - WEDGED, THICK PRESSURE CARE CUSHION WITH WATERPROOF COVER
Variance allowed on overall measurement: +1 cm Variance allowed on front height: +0.5cm
Size range:
• 61 cm wide; 43 cm long; 6 cm high at rear, 11 cm high in front
• 56 cm wide; 43 cm long; 6 cm high at rear, 11 cm high in front
• 51 cm wide; 43 cm long; 6 cm high at rear, 11 cm high in front
• 46 cm wide; 43 cm long; 6 cm high at rear, 11 cm high in front
• 41 cm wide; 43 cm long; 6 cm high at rear, 11 cm high in front
• 38 cm wide; 43 cm long; 6 cm high at rear, 11 cm high in front
• 36 cm wide; 43 cm long; 6 cm high at rear, 11 cm high in front
• 30 cm wide; 43 cm long; 6 cm high at rear, 11 cm high in front
• 30 cm wide; 30 cm long; 4 cm high at rear, 8.5 cm high in front
25 cm wide; 30 cm long; 4 cm high at rear, 7.5 cm high in front
Description of cushion:
• Wheelchair basic foam pressure relieving cushion. Bevelled contoured base to match sling seat of wheelchair. • Must have trochanteric loading pads to create a well. • The depth of the well for 14” (36cm) - 24” (61cm) should be 3cm, 12” (30cm) should be 2,5cm and 10’ (2cm) . • The width of the well = 50% of the width of the cushion. • Length of the well = the width of the well up to maximum 20cm. • 2cm wedge in front of pre-ischial shelf extending to the front of the cushion. • Sizes 14”-24”: Top layer - 5cm thick high-density foam following contour of base through full length of cushion; only the front half of the cushion to be glued. • Size 10” -12”: Top - 3 cm thick high-density foam following contour of base through full length of cushion; only the front half of the cushion to be glued. • Type, size, and direction of cushion use printed on bottom of base.
Material:
Top layer:
For sizes: 25cm – 46 cm High-density foam 36/20; Hardness factor 18 – 21.6; Density range 32.1-40kg/m³
For sizes: 51cm -61cm High density foam 50/30 Hardness factor: 249 - 291 and Density: 37.1- 41kg/m³
Base layer:
High density compressed (reconstituted flexible PU) foam 1300; Hardness factor 52-63; Density 130+10% kg/m³ (SANS 642 - 1976)
Description of cushion cover:
• Fabric, waterproof, terry cloth covered. • 2 ways stretch, breathable, anti-microbial shield treatment. • Machine washable. • Pull over elastic band with size label. • Colour: Preferably Black, alternately Navy Blue.
RT233-4.2.3 | CONTOURED BASE - THIN FOAM, BASIC POSITIONER CUSHION WITH WATERPROOF COVER
Variance allowed on overall measurement : +1 cm
Size range:
• 51 cm wide; 43 cm long; 7 cm high
• 46 cm wide; 43 cm long; 7 cm high
• 43 cm wide; 43 cm long; 7 cm high
• 41 cm wide; 43 cm long; 7 cm high
• 38 cm wide; 43 cm long; 7 cm high
• 36 cm wide; 43 cm long; 7 cm high
• 30 cm wide; 43 cm long; 6.5 cm high
• 30 cm wide; 33 cm long; 6.5 cm high
• 25 cm wide; 28 cm long; 6 cm high
Description of cushion:
• Bevelled contoured base to match sling seat of wheelchair. Sitting surface with pre-ischial shelf. • The depth of the pre-ischial shelf for 14” (36cm) - 20 (51cm) should be 3cm minimum, for 12” (30cm) should be 2,5cm minimum, and for 10” (25cm) should be 2cm minimum. • The distance of the shelf from the rear of the cushion = 50% of the width of the cushion up to maximum of 20cm. • Top: 3cm thick high-density foam following contour of base through full length of cushion; only the front half of the cushion to be glued. • Type, size, and direction of cushion use printed on bottom of base.
Material:
Top layer:
For sizes: 25cm - 48cm High-density foam 36/20; Hardness factor 18 – 21.6; Density range 32.1-40kg/m³
For sizes: 51cm High density foam 50/30 Hardness factor: 249 - 291 and Density: 37.1- 41kg/m³
Base layer:
High density compressed (reconstituted flexible PU) foam 1300; Hardness factor 52-63; Density 130+10% kg/m³ (SANS 642 - 1976)
Description of cushion cover:
• Fabric, waterproof, terry cloth covered. • 2 ways stretch, breathable, anti-microbial shield treatment. • Machine washable. • Pull over elastic band with size label. • Colour: Preferably Black, alternately Navy Blue.
RT233-4.2.4 | CONTOURED BASE - WEDGED, THIN FOAM POSITIONER WITH WATERPROOF COVER
Variance allowed on overall measurement: +1cm Variance allowed on front height:+0.5cm
Size range:
• 51 cm wide; 43 cm long; 4 cm high at rear, 9 cm high in front
• 46 cm wide; 43 cm long; 4 cm high at rear, 9 cm high in front
• 43 cm wide; 43 cm long; 4 cm high at rear, 9 cm high in front
• 41 cm wide; 43 cm long; 4 cm high at rear, 9 cm high in front
• 38 cm wide; 43 cm long; 4 cm high at rear, 9 cm high in front
• 36 cm wide; 43 cm long; 4 cm high at rear, 9 cm high in front
• 30 cm wide; 43 cm long; 4 cm high at rear, 9 cm high in front
• 30 cm wide; 33 cm long; 4 cm high at rear, 8 cm high in front
• 25 cm wide; 28 cm long; 4 cm high at rear, 7.5cm high in front
Description of cushion:
• Bevelled contoured base to match sling seat of wheelchair. Sitting surface with pre-ischial shelf. • The depth of the pre-ischial shelf for 14” (36cm) - 20” (51cm) should be 3cm minimum, 12” (30cm) should be 2,5cm minimum, and 10” (25cm) should be 2cm minimum. • The distance of the shelf from the rear = 50% of the width of the cushion up to maximum of 20cm. • 2cm wedge in front of pre-ischial shelf extending to the front of the cushion. • Top: 3cm thick high-density foam following contour of base through full length of cushion; only the front half of the cushion to be glued. • Type, size, and direction of cushion use printed on bottom of base.
Material:
Top layer:
For sizes: 25cm-46cm High-density foam 36/20; Hardness factor 18 – 21.6; Density range 32.1-40kg/m³
For sizes: 51cm High density foam 50/30 Hardness factor: 249 - 291 and Density: 37.1- 41kg/m³
Base layer:
High density compressed (reconstituted flexible PU) foam 1300; Hardness factor 52-63; Density 130+10% kg/m³ (SANS 642 - 1976)
Description of cushion cover:
• Fabric, waterproof, terry cloth covered. • 2 ways stretch, breathable, anti-microbial shield treatment. • Machine washable. • Pull over elastic band with size label. • Colour: Preferably Black, alternately Navy Blue.
RT233-4.2.5 | CONTOURED BASE - THICK PRESSURE CARE CUSHION, EXTENDED LENGTH WITH WATERPROOF COVER
Variance allowed on overall measurement: +1cm
Size range:
• 61 cm wide; 50 cm long; 9 cm high
• 56 cm wide; 50 cm long; 9 cm high
• 51 cm wide; 50 cm long; 9 cm high
• 46 cm wide; 50 cm long; 9 cm high
• 43 cm wide; 50 cm long; 9 cm high
• 41 cm wide; 50 cm long; 9 cm high
• 38 cm wide; 50 cm long; 9 cm high
• 36 cm wide; 50 cm long; 8.5 cm high
• 30 cm wide; 50 cm long; 8.5 cm high
Description of cushion:
• Wheelchair basic foam pressure relieving cushion. Bevelled contoured base to match sling seat of wheelchair. • Must have trochanteric loading pads to create a well. • The depth of the well for 14” (36cm) - 24” (61cm) should be 3cm, for 12” (30cm) should be 2,5cm. • The width of the well = 50% of the width of the cushion. • Length of the well = the width of the well up to maximum 20cm. • Size 14”-24”, 5cm thick high-density foam top layer following contour of base through full length of cushion; only the front half of the cushion to be glued. • Size 12”, 3cm thick high-density foam top layer, following contour of base through full length of cushion; only the front half of the cushion to be glued. • Type, size, and direction of cushion use printed on bottom of base.
Material:
Top layer:
For sizes: 30cm-46cm High-density foam 36/20; Hardness factor 18 – 21.6; Density range 32.1-40kg/m³
For sizes: 51cm-61cm High density foam 50/30 Hardness factor: 249 - 291 and Density: 37.1- 41kg/m³
Base layer:
High density compressed (reconstituted flexible PU) foam 1300; Hardness factor 52-63 Density 130+10% kg/m³ (SANS 642 - 1976)
Description of cushion cover:
• Fabric, waterproof, terry cloth covered. • 2 ways stretch, breathable, anti-microbial shield treatment. • Machine washable. • Pull over elastic band with size label. • Colour: Preferably Black, alternately Navy Blue.
RT233-4.2.6 | CONTOURED BASE - WEDGED, THICK PRESSURE CARE CUSHION, EXTENDED LENGTH WITH WATERPROOF COVER
Variance allowed on overall measurement:+1cm
Size range:
• 61 cm wide; 50 cm long; 6 cm high at rear, 11 cm high in front
• 56 cm wide; 50 cm long; 6 cm high at rear, 11 cm high in front
• 51 cm wide; 50 cm long; 6 cm high at rear, 11 cm high in front
• 46 cm wide; 50 cm long; 6 cm high at rear, 11 cm high in front
• 43 cm wide; 50 cm long; 6 cm high at rear, 11 cm high in front
• 41 cm wide; 50 cm long; 6 cm high at rear, 11 cm high in front
• 38 cm wide; 50 cm long; 6 cm high at rear, 11cm high in front
• 36 cm wide; 50 cm long; 6 cm high at rear, 11 cm high in front
• 30 cm wide; 50 cm long; 6 cm high at rear, 10.5 cm high in front
Description of cushion:
• Wheelchair basic foam pressure relieving cushion. Bevelled contoured base to match sling seat of wheelchair. • Must have trochanteric loading pads to create a well. • The depth of the well for 14” (36cm) - 24” (61cm) should be 3cm, 12” (30cm) should be 2,5cm. • The width of the well = 50% of the width of the cushion. • Length of the well = the width of the well up to maximum 20cm. • Size 14”-24”, 5cm thick high-density foam top layer, following contour of base; only the front half of the cushion to be glued. • Size 12”, 3cm, thick high density foam top layer, following contour of base through full length of cushion, only the front half of the cushion to be glued. • Type, size, and direction of cushion use printed on bottom of base.
Material:
Top layer:
For sizes: 30cm-46cm High-density foam 36/20; Hardness factor 18 – 21.6; Density range 32.1-40kg/m³
For sizes: 51cm-61cm High density foam 50/30 Hardness factor: 249 - 291 and Density: 37.1- 41kg/m³
Base layer:
High density compressed (reconstituted flexible PU) foam 1300; Hardness factor 52-63 Density 130+10% kg/m³ (SANS 642 - 1976)
Description of cushion cover:
• Fabric, waterproof, terry cloth covered. • 2 ways stretch, breathable, anti-microbial shield treatment. • Machine washable. • Pull over elastic band with size label. • Colour: Preferably Black, alternately Navy Blue.
RT233-4.2.7 | CONTOURED BASED - THIN FOAM, BASIC POSITIONER CUSHION, EXTENDED LENGTH WITH WATERPROOF COVER
Variance allowed on overall measurement: +1cm
Size range:
• 51 cm wide; 50 cm long; 7 cm high in front
• 46 cm wide; 50 cm long; 7 cm high in front
• 43 cm wide; 50 cm long; 7 cm high in front
• 41 cm wide; 50 cm long; 7 cm high in front
• 38 cm wide; 50 cm long; 7 cm high in front
• 36 cm wide; 50 cm long; 7 cm high in front
• 30 cm wide; 50 cm long; 7 cm high in front
Description of cushion:
• Bevelled contoured base to match sling seat of wheelchair. Sitting surface with pre-ischial shelf. • The depth of the pre-ischial shelf for 14” (36cm) - 20” (51cm) should be 3cm minimum; 12” (30cm) should be 2, 5cm minimum. • The distance of the shelf from the rear of the cushion = 50% of the width of the cushion up to maximum of 20cm. • Top: 3cm thick high-density foam following contour of base through full length of cushion; only the front half of the cushion to be glued. • Type, size, and direction of cushion use printed on bottom of base.
Material:
Top layer:
For Sizes: 30cm -46cm High-density foam 36/20; Hardness factor 18 – 21.6; Density range 32.1-40kg/m³
For sizes: 51cm High density foam 50/30 Hardness factor: 249 - 291 and Density: 37.1- 41kg/m³
Base layer:
High density compressed (reconstituted flexible PU) foam 1300; Hardness factor 52-63; Density 130+10% kg/m³ (SANS 642 - 1976)
Description of cushion cover:
• Fabric, waterproof, terry cloth covered. • 2 ways stretch, breathable, anti-microbial shield treatment. • Machine washable. • Pull over elastic band with size label. • Colour: Preferably Black, alternately Navy Blue.
RT233-4.2.8 | CONTOURED BASE – WEDGED THIN FOAM POSITIONER, EXTENDED LENGTH WITH WATERPROOF COVER
Variance allowed on overall measurement: +1 cm
Size range:
• 51 cm wide; 50 cm long; 4 cm high at rear, 9 cm in front
• 46 cm wide; 50 cm long; 4 cm high at rear, 9 cm in front
• 43 cm wide; 50 cm long; 4 cm high at rear, 9 cm in front
• 41 cm wide; 50 cm long; 4 cm high at rear, 9 cm in front
• 38 cm wide; 50 cm long; 4 cm high at rear, 9 cm in front
• 36 cm wide; 50 cm long; 4 cm high at rear, 9 cm in front
• 30 cm wide; 50 cm long; 4 cm high at rear, 8.5 cm in front
Description of cushion:
• Bevelled contoured base to match sling seat of wheelchair. Sitting surface with pre-ischial shelf. • The depth of the pre-ischial shelf for 14” (36cm) - 20” (51cm) should be 3cm minimum, 12” (30cm) should be 2, 5 cm minimum. • The distance of the shelf from the rear of the cushion = 50% of the width of the cushion up to maximum of 20cm. • 2cm wedge in front of pre-ischial shelf extending to the front of the cushion. • Top: 3cm thick high-density foam following contour of base through full length of cushion; only the front half of the cushion to be glued. • Type, size, and direction of cushion use printed on bottom of base
Material:
Top layer:
For sizes: 30cm -46cm High-density foam 36/20; Hardness factor 18 – 21.6; Density range 32.1-40kg/m³
For sizes: 51cm High density foam 50/30 Hardness factor: 249 - 291 and Density: 37.1- 41kg/m³
Base layer:
High density compressed (reconstituted flexible PU) foam 1300; Hardness factor 52-63; Density 130+10% kg/m³ (SANS 642 - 1976)
Description of cushion cover:
• Fabric, waterproof, terry cloth covered. • 2 ways stretch, breathable, anti-microbial shield treatment. • Machine washable. • Pull over elastic band with size label. • Colour: Preferably Black, alternately Navy Blue.
RT233-4.2.9 | NON-CONTOURED BASE - STRAIGHT BASE, BASIC FOAM, THICK PRESSURE CARE CUSHION WITH NO ABDUCTION BLOCK WITH WATERPROOF COVER
Variance allowed on overall measurement: +1cm
Size Range:
• 20 cm wide, 25 cm long, 9cm high
• 22.5 cm wide, 25 cm long, 9cm high
• 25 cm wide, 35 cm long, 9cm high
• 27.5 cm wide, 40 cm long, 9cm high
• 30 cm wide, 35 cm long, 9cm high
• 30 cm wide, 40 cm long, , 9cm high
• 36 cm wide, 40 cm long, , 9cm high
• 36 cm wide , 50 cm long, 9cm high
• 38cm wide, 50cm long, 9cm high
• 40 cm wide, 50 cm long, 9cm high
• 43 cm wide, 50 cm long, 9cm high
• 45 cm wide, 50 cm long, 9cm high
• 51 cm wide, 50 cm long, 9cm high
Description of cushion:
• Wheelchair basic foam pressure relieving cushion. Flat base to match solid wheelchair seat. • Must have trochanteric loading pads to create a well. • The depth of the well for 14” (36cm) - 20”(51cm) should be 3cm, for 8” (20cm) - 12”(30cm) should be 2,5cm. • The width of the well = 50% of the width of the cushion • The length of the well = the width of the well up to maximum 20cm. • Base layer: Minimum thickness 4 cm for sizes 14”-20” and minimum thickness 3cm for 12” and minimum 2,5cm for 10” and smaller. • The thickness of the top layer for 14” (36cm) -20” (51cm) should be 5cm minimum and for 8” (20cm) -12” (30cm ) should be 3cm; only the front half of the cushion to be glued. • Type, size, and direction of cushion use printed on bottom of base.
Material:
Top layer:
For sizes to fit: 8”-18” High-density foam 36/20; Hardness factor 18 – 21.6; Density range 32.1-40kg/m³
For sizes: 20” High density foam 50/30 Hardness factor: 249 - 291 and Density: 37.1- 41kg/m³
Base layer:
High density compressed (reconstituted flexible PU) foam 1300 Hardness factor 52-63; Density 130+10% kg/m³ (SANS 642 - 1976)
Description of cushion cover:
• Fabric, waterproof, terry cloth covered. • 2 ways stretch, breathable, anti-microbial shield treatment. • Machine washable. • Pull over elastic band with size label. • Colour: Preferably Black, alternately Navy Blue.
RT233-4.2.10 | NON-CONTOURED BASE - STRAIGHT BASE, THICK PRESSURE CARE CUSHION WITH ABDUCTION BLOCK AND WATERPROOF COVER
Variance allowed on overall measurement: +1cm
Size Range:
• 20 cm wide, 25 cm long, 9cm high with extra-small abduction block
• 22.5 cm wide, 25 cm long, 9cm high with extra-small abduction block
• 25 cm wide, 35 cm long, 9cm high with extra-small abduction block
• 27.5 cm wide, 40 cm long, 9cm high with extra-small abduction block
• 30 cm wide, 35 cm long, 9cm high, with small abduction block
• 30 cm wide, 40 cm long, 9cm high, with small abduction block
• 35 cm wide, 40 cm long, 9cm high, with small abduction block
• 36 cm wide, 50 cm long, 9cm high, with small abduction block
• 38cm wide, 50cm long, 9cm high, with small abduction block
• 41 cm wide, 50 cm long, 9cm high with large abduction block
• 43 cm wide, 50 cm long, 9cm high with large abduction block
• 46 cm wide, 50 cm long, 9cm high with large abduction block
• 51 cm wide, 50 cm long, 9cm high with large abduction block
Description of cushion:
| • Wheelchair basic foam pressure relieving cushion. Flat base to match solid wheelchair seat. • Must have trochanteric loading pads to create a well. • The depth of the well for 14” (36cm) - 20”(51cm) should be 3cm, for 8” (20cm) - 12”(30cm) should be 2,5cm. • The width of the well = 50% of the width of the cushion • The length of the well = the width of the well up to maximum 20cm. • Base layer: Minimum thickness 4 cm for sizes 14”-20” and minimum thickness 3cm for 12” and minimum 2,5cm for 10” and smaller. • The thickness of the top layer for 14” (36cm) -20” (51cm) should be 5cm minimum and for 8” (20cm) -12” (30cm ) should be 3cm; only the front half of the cushion to be glued. • The pommel should extend beyond the end of the cushion to reach the front of the user’s knees. • The height of the pommel = the full thickness of the cushion from the base and includes average height of user’s thigh in the front. • The length of the pommel should not be more than the distal third of the femur. • Type, size, and direction of cushion use printed on bottom of base. |  |  | No specific measurements of the pommel provided
Material:
Top layer:
For sizes to fit: 8”-18” High-density foam 36/20; Hardness factor 18 – 21.6; Density range 32.1-40kg/m³
For sizes: 20” High density foam 50/30 Hardness factor: 249 - 291 and Density: 37.1- 41kg/m³
Base layer:
High density compressed (reconstituted flexible PU) foam 1300 Hardness factor 52-63; Density 130+10% kg/m³ (SANS 642 - 1976)
Description of cushion cover:
• Fabric, waterproof, terry cloth covered. • 2 ways stretch, breathable, anti-microbial shield treatment. • Machine washable. • Pull over elastic band with size label. • Colour: Preferably Black, alternately Navy Blue.
RT233-4.2.11 | NON-CONTOURED BASE - STRAIGHT BASE, POSITIONER WITH NO ABDUCTION BLOCK WITH WATERPROOF COVER
Variance allowed on overall measurement: +1cm
Size Range:
20 cm wide, 25 cm long, 7cm high
22.5 cm wide, 25 cm long, 7cm high
25 cm wide, 35 cm long, 7cm high
27.5 cm wide, 40 cm long, 7cm high
30 cm wide, 35 cm long, 7cm high
30 cm wide, 40 cm long, 7cm high
36 cm wide, 40 cm long, 7cm high
36 cm wide, 50 cm long, 7cm high
38cm wide, 50cm long, 7 cm high
40 cm wide, 50 cm long, 7cm high
43 cm wide, 50 cm long, 7cm high
45 cm wide, 50 cm long, 7cm high
51 cm wide, 50 cm long, 7cm high
Description of cushion:
• Wheelchair basic foam positioner cushion. Flat base to match solid wheelchair seat • Sitting surface with pre-ischial shelf. • The depth of the pre-ischial shelf for 14” (36cm) - 20” (51cm) should be 3cm minimum, 12” (30cm) should be 2,5 cm minimum. • The distance of the shelf from the rear of the cushion = 50% of the width of the cushion up to maximum of 20cm. • Base layer: Minimum thickness 4cm for sizes 14”-20” and minimum thickness 3cm for 12” and minimum 2,5cm for 10” and smaller. • The thickness of the top layer for 14” (36cm) -20” (51cm) should be 3cm minimum and for 8” (20cm) -12” (30cm ) should be 2.5cm.; only the front half of the cushion to be glued. • Type, size, and direction of cushion use printed on bottom of base.
Material:
Top layer:
For sizes to fit: 8”-18” High-density foam 36/20; Hardness factor 18 – 21.6; Density range 32.1-40kg/m³
For sizes: 20” High density foam 50/30 Hardness factor: 249 - 291 and Density: 37.1- 41kg/m³
Base layer:
High density compressed (reconstituted flexible PU) foam 1300 Hardness factor 52-63: Density 130+10% kg/m³ (SANS 642 - 1976)
Description of cushion cover:
• Fabric, waterproof, terry cloth covered. • 2 ways stretch, breathable, anti-microbial shield treatment. • Machine washable. • Pull over elastic band with size label. • Colour: Preferably Black, alternately Navy Blue.
RT233-4.2.12 | NON-CONTOURED BASE - STRAIGHT BASE, POSITIONER WITH ABDUCTION BLOCK WITH WATERPROOF COVER
Variance allowed on overall measurement: +1cm
Size Range:
• 20 cm wide, 25 cm long, 7cm high with extra-small abduction block
• 22.5 cm wide, 25 cm long, 7cm high with extra-small abduction block
• 25 cm wide, 35 cm long, 7cm high with extra-small abduction block
• 27.5 cm wide, 40 cm long, 7cm high with extra-small abduction block
• 30 cm wide, 35 cm long, 7cm high with small abduction block
• 30 cm wide, 40 cm long, 7cm high with small abduction block
• 35 cm wide, 40 cm long, 7cm high with small abduction block
• 36 cm wide, 50 cm long, 7cm high with small abduction bloc
• 41 cm wide, 50 cm long, 7cm high with large abduction block
• 43 cm wide, 50 cm long, 7cm high with large abduction block
• 46 cm wide, 50 cm long, 7cm high with large abduction block
• 51 cm wide, 50 cm long, 7cm high with large abduction block
Description:
| • Wheelchair basic foam positioner cushion. Flat base to match solid wheelchair seat • Sitting surface with pre-ischial shelf. • The depth of the well for 14” (36cm) - 20”(51cm) should be 30mm, for 12”(30cm) should be 2,5cm. • The distance of the shelf from the rear of the cushion = 50% of the width of the cushion up to maximum of 20cm. • Base layer: Minimum thickness 4cm for sizes 14”-20” and minimum thickness 3cm for 12” and minimum 2,5cm for 10” and smaller. • The thickness of the top layer for 14” (36cm) -20” (51cm) should be 3cm minimum and for 8” (20cm) -12” (30cm ) should be 2.5cm.; only the front half of the cushion to be glued. • The pommel should extend beyond the end of the cushion to reach the front of the user’s knees. • The height of the pommel = the full thickness of the cushion from the base and includes average height of user’s thigh in the front. • The length of the pommel should not be more than the distal third of the femur. • Type, size, and direction of cushion use printed on bottom of base. |  |  | No specific measurements of the pommel provided
Material:
Top layer:
For sizes to fit: 8”-18” High-density foam 36/20; Hardness factor 18 – 21.6; Density range 32.1-40kg/m³
For sizes: 20” High density foam 50/30 Hardness factor: 249 - 291 and Density: 37.1- 41kg/m³
Base layer:
High density compressed (reconstituted flexible PU) foam 1300 Hardness factor 52-63: Density 130+10% kg/m³ (SANS 642 - 1976)
Description of cushion cover:
• Fabric, waterproof, terry cloth covered. • 2 ways stretch, breathable, anti-microbial shield treatment. • Machine washable. • Pull over elastic band with size label. • Colour: Preferably Black, alternately Navy Blue.

## 8 Category 5
CATEGORY 5
All items being tendered for must comply with the following technical specification, failure to comply with any of the conditions set out below may result in immediate disqualification. Please note that where the technical specification calls for certification, this certification must accompany the bid, failure to provide such certification will result in immediate disqualification. Bidders must give details of items being tendered for in the COMMENTS section where there are variation for further deliberation
BACKREST AND SEATING SYSTEMS
RT233-5.1 | TENSION ADJUSTABLE BACK | Indicate the item bidding for below bby stating "Yes bidding for this item". Item left blank will be assumed it is not bidded for | Complies Yes/No | Comments (Technical Specification Deviation)
TECHNICAL SPECIFICATIONS
To fit both basic and adjustable folding frame wheelchairs
5.1.1 | TENSION ADJUSTABLE BACK, STANDARD
5.1.2 | TENSION ADJUSTABLE BACK, SHORT
5.1.3 | TENSION ADJUSTABLE BACK, TALL
Size range:
See Table 7: Tension Adjustable Back Size Ranges on Annexure ??
Backrest components:
• 3-piece tension adjustable backrest. • 2x side panels with full length tubes for fitting to wheelchair by either using screw-in type. • Fixing (basic folding wheelchairs with seat guide) or slipping tube over back rest (adjustable folding frame wheelchairs). No pre-made screw holes are to be provided in the tubes. One panel to be fitted with straps the other with metal buckles. • Panels: full length +1cm from top and bottom: Hook and loop fasteners strip 4cm wide placed 5cm from edge. On panel fitted with straps, a second 4cm loop strap immediately adjacent to inside of first strip, thus providing 8cm wide loop section. • A separate cover is provided. The front panel is padded. Front panel of cover to be clear of stitching, except for 7cm from edges. Front part of cover to be open from inside (pocket).
Backrest material:
• Nylon; Black; 600D or stronger Nylon for panels and tubes of tension adjustable panels; and appropriate nylon for cover; Tubes reinforced with 400gsm PVC or stronger; Cover front padded with 12mm thick high-density foam (30/20 with hardness 18.0 - 21.6; Density 27.1 – 32 kg/m³ or 26/20 with hardness 18.0 - 21.6; Density 24.1 – 27 kg/m³ or other suitable foam).
Backrest straps:
• Maximum 2,5cm wide straps no wider than 2cm apart; hook and loop fastenings; Must be able to conform to body contours when adjusted.
• Minimum of 10 evenly spaced for standard length, 8 for short length and 12 straps for tall back.
RT233-5.2 | TENSION ADJUSTABLE BACKREST WITH RIGID HEIGHT ADJUSTABLE FULL-LENGTH LATERALS | Indicate the item bidding for below bby stating "Yes bidding for this item". Item left blank will be assumed it is not bidded for | Complies Yes/No | Comments (Technical Specification Deviation)
Height adjustable full-length laterals to fit both basic and adjustable folding frame manual and power wheelchairs
5.2.1 | TENSION ADJUSTABLE BACK WITH RIGID, HEIGHT ADJUSTABLE FULL-LENGTH LATERALS, SHORT
5.2.2 | TENSION ADJUSTABLE BACK WITH RIGID, HEIGHT ADJUSTABLE FULL-LENGTH LATERALS, STANDARD
5.2.3 | TENSION ADJUSTABLE BACK WITH RIGID, HEIGHT ADJUSTABLE FULL-LENGTH LATERALS, TALL
Size range:
See Table 8: Tension Adjustable Backrest with Rigid, Height adjustable full length laterals, Size Ranges on Annexure ??
Components and description:
• Include attachments to fit back system to wheelchair. If attachment is not size adjustable (universal bracket), is to be supplied with additional brackets or sleeve / filler to allow fitting onto narrower inner section of height adjustable back posts with range 19-22mm. This implies that 2 sets of each size should be provided in case if brackets which are not size adjustable. • Design of bracket to allow fitting below to bend in back post or push handle with no interference during removal and placement. • Quick-release mechanism allows for back to detach from wheelchair for transport. • 3-piece tension adjustable backrest. • 2x side panels with height adjustable full length rigid lateral supports. One panel to be fitted with straps the other with metal buckles. • Trunk side supports should be lightweight, easy to bend and mould and height adjustable. Allows for contouring to fit body shape
Material:
• Separate covers are removable, washable heavy-duty lycra (220 g/m² and yield 3.4m/kg) with hook and loop attachments.
Straps:
• Maximum 25mm wide straps close and evenly spaced and not further apart than 20mm with steel sliders and hook and loop fastenings; Must be able to conform to body contours when adjusted; Straps allow for minimum of 100mm adjustment.
Fitting brackets:
• Set of fitting brackets for 19mm back post • Set of fitting brackets for 22mm back post
RT233-5.3 | RIGID ADJUSTABLE BACK WITH FULL LENGTH LATERAL SUPPORT | Indicate the item bidding for below bby stating "Yes bidding for this item". Item left blank will be assumed it is not bidded for | Complies Yes/No | Comments (Technical Specification Deviation)
Integrated trunk and pelvis side support
5.3.1 | RIGID ADJUSTABLE BACK WITH FULL LENGTH INTEGRATED TRUNK AND PELVIS SIDE SUPPORT, MID THORACIC HEIGHT
5.3.2 | RIGID ADJUSTABLE BACK WITH FULL LENGTH INTEGRATED TRUNK AND PELVIS SIDE SUPPORT, UPPER THORACIC HEIGHT
5.3.3 | RIGID ADJUSTABLE BACK WITH FULL LENGTH INTEGRATED TRUNK AND PELVIS SIDE SUPPORT, SHOULDER HEIGHT
Size range (+- 1 cm):
See Table 9: Rigid Adjustable Back with full length laterals, Mid-Thoracic Height Measurements on Annexure ??
See Table 10: Rigid Adjustable Back with full length laterals, Upper Thoracic Height Measurements on Annexure ??
See Table 11: Rigid Adjustable Back with full length laterals, Shoulder Height Measurements on Annexure ??
Components and description:
• Include attachments to fit back system to wheelchair. If attachment is not size adjustable (universal bracket), is to be supplied with sleeve / filler to allow fitting onto narrower inner section of height adjustable back posts with range 1.9cm-2.2cm. • Quick-release mechanism allows for back to detach from wheelchair for transport. • Back system fitting on back posts to be height adjustable and recline. • Angle adjustable setting - maximum 15 degrees recline. • Trunk side supports to allow width adjustment as per specified range • Maximum thickness of trunk side supports = 1.5cm. • Should allow build up and modifications to customise fit. • Separate covers are removable, washable appropriate stretch fabric.
5.3 C | CUSTOMISING OPTIONS
5.3 C.1 | Set of fitting brackets for a second wheelchair
5.3 C.2 | Set of fitting brackets for 19mm back post
5.3 C.3 | Set of fitting brackets for 22mm back post
RT233-5.4 | RIGID ADJUSTABLE BACK | Indicate the item bidding for below bby stating "Yes bidding for this item". Item left blank will be assumed it is not bidded for | Complies Yes/No | Comments (Technical Specification Deviation)
5.4.1 | RIGID ADJUSTABLE BACK, MID THORACIC HEIGHT. RANGE: 30CM-38CM
5.4.2 | RIGID ADJUSTABLE BACK, UPPER THORACIC HEIGHT. RANGE: 42CM-50CM
5.4.3 | RIGID ADJUSTABLE BACK, SHOULDER HEIGHT. RANGE: 53CM-61CM
Minimum of 3 back heights either as individual backrest or with an adjustability option.
Size of back (measured from outside to outside of trunk side supports)
12”
14”
16”
18”
Components and description:
• Include attachments to fit back system to wheelchair. If attachment is not size adjustable (universal bracket), it is to be supplied with sleeve / filler to allow fitting onto narrower inner section of height adjustable back posts with range 1.9cm-2.5cm. • Quick-release mechanism allows for back to detach from wheelchair for transport. • Lightweight rigid adjustable back support with removable / swing-away / flip-up trunk side support. • Back system fitting on back posts to be height adjustable and recline. • Angle adjustable setting - maximum 15 degrees recline. • Trunk side supports to allow width adjustment as per specified range. • Maximum thickness of trunk side supports = 1.5cm. • Should allow build up and modifications to customise fit. • Separate covers are removable, washable appropriate stretch fabric.
5.4 C | CUSTOMISING OPTIONS
5.4 C.1 | Set of backrests mounting brackets to accommodate at least 19mm-22mm diameter back post for second wheelchair.
5.4 C.2 | Pairs of side supports to support trunk and/ or pelvis. Side support assembly allows for both adjustment in horizontal and vertical planes to adjust for depth and height. End user to specify if one or two pairs are needed. Side support pad size (depth height). Back rest with one set of side support only.
10x10 : + 1cm
10x12.5 : + 1cm
10x15 : + 1cm
15x10 : + 1cm
15x12.5 : + 1cm
15x15 : + 1cm
5.4 C.3 | Headrest with mounting bar to fit on backrest
• Easily removable • Height and depth adjustable
RT233-5.5 | ADJUSTABLE MODULAR BACK | Indicate the item bidding for below bby stating "Yes bidding for this item". Item left blank will be assumed it is not bidded for | Complies Yes/No | Comments (Technical Specification Deviation)
Adjustable modular /solid back with individual flip up or swing away rigid trunk and or pelvis side support
5.5.1 | RIGID MODULAR BACK, STANDARD HEIGHT
5.5.2 | RIGID MODULAR BACK, TALL HEIGHT
Size range:
See Table 13: Adjustable Modular Back Size Ranges on Annexure ??
If no technical specification received with order, device will be issued with standard length supports with no curve.
Components and description:
• Adjustable in width, height and recline. • Independent lumbar-sacral support adjustable in height and angle. • Independent thoracic support adjustable in height. • Independent lateral supports adjustable in height and width. • Attachments to fit back system to range of wheelchair back posts with range 1.9cm-2.5cm
Material:
• Back support sections: high density polyethylene (HDPE) moulded with sealed cell foam supports, which can be customised. • Lateral support: HDPE moulded plastic with sealed cell foam supports which can be build up and customised – available in standard or optional long to accommodate extra body depth. • Covers made from washable heavy-duty lycra (220 g/m² and yield 3.4m/kg) with hook and loop attachments. • Frame: Mild steel passivated or stainless steel. • Back system detachable for transport.
5.5 C | CUSTOMISING OPTIONS
5.5 C.1 | Head support, basic adjustable
• Minimum 2 sizes, small and standard • Easily removable • Height and depth adjustable • Kidney shaped head support flexible and rake adjustable • Padded and covered in heavy gauge, removable, washable fabric
5.5 C.2 | Head support, complex adjustable
• Minimum 2 sizes, small and standard • Easily removable • Height and depth adjustable point, hand shaped head support, flexible and rake adjustable • Moulded and padded EVA
5.5 C.3 | Set of backrests mounting brackets to accommodate at least 19mm-22mm diameter back post for second wheelchair
RT233-5.6 | LOW-COST MODULAR SEATING SYSTEM | Indicate the item bidding for below bby stating "Yes bidding for this item". Item left blank will be assumed it is not bidded for | Complies Yes/No | Comments (Technical Specification Deviation)
Low-cost modular seating unit
Sizes:
See Table 14: Low-Cost Modular Seating System Size Ranges on Annexure ??
Components:
• Seat assembly • Backrest assembly • Lateral pad assembly • Straps assembly • Tray table assembly • Brackets • Fasteners
A | Seat assembly
One seat panel with components for basic layered foam cushion and cover. Components include chip foam base, thick enough or in layers to create pre-ischial shelf, high density top layer.
B | Back assembly
One seat panel with back pad and pad cover
C | Lateral pad assembly
Six lateral support pads to be used as trunk supports, hip supports and abductor in any combination required. To fit child and adult wheelchair users.
D | Straps
1x lap strap assembly 1x chest strap assembly 2 x foot straps assembly
E | Tray table assembly
1x Tray table and 2 support assembly with attachment brackets to use on wheelchairs with no armrests
F | Brackets
Steel brackets (“L” brackets for the posture support device attachment to the back and seat boards.)
G | Fasteners
Bolts and washers, nuts etc. needed to secure lateral support to back and seat panels and seating system to wheelchair. Components for attachment to wheelchair by means of quick release system.
5.6 C | CUSTOMISING OPTIONS
5.6 C.1 | Attendant propelled wheeled base for low-cost seating system with:
• Minimum 90 degrees back to seat angle • Tilt in space option • Folding frame
5.6 C.2 | Self‐propelled wheeled base for low-cost seating system with:
• Minimum 90 degrees back to seat angle • Tilt in space option • Folding frame
5.6 C.3 | Rear wheels to fit attendant propelled wheeled base
5.6 C.4 | Rear wheels to fit self-propelled wheeled base
RT233-5.7 | TRUNK SIDE SUPPORT | Indicate the item bidding for below bby stating "Yes bidding for this item". Item left blank will be assumed it is not bidded for | Complies Yes/No | Comments (Technical Specification Deviation)
Sizes:
• 10 cm x10 cm + 2 cm
• 10 cm x12.5 cm +2 cm
• 10 cm x15 cm + 2 cm
• 15 cm x10 cm + 2 cm
• 15 cm x12.5 cm +2 cm
• 18cm x 10cm ± 2cm
• 18 cm x 12 cm ± 2cm
Side support pad size (depth height)
Description:
• Pairs of side supports to support trunk by attaching directly onto the back post of the wheelchair. • To be used in combination with a tension adjustable backrest. • The assembly should allow horizontal adjustment to bring trunk pad closer to or further away from user. • End user to specify size • Side support assembly allows for both adjustments in horizontal and vertical planes to adjust for depth and height. • End user to specify if one or two pairs needed.
Back rest mounting brackets:
• Set of backrests mounting brackets to accommodate at least 19mm -22mm diameter back post tubes.
RT233-5.8 | BASE SEATING SYSTEMS | Indicate the item bidding for below bby stating "Yes bidding for this item". Item left blank will be assumed it is not bidded for | Complies Yes/No | Comments (Technical Specification Deviation)
Pelvic straps for all wheelchairs (excluding posture wheelchairs and buggies)
5.8.1 | Pelvic straps, basic adjustable, with D-ring fitting
• Poly-cotton webbing, first preference is Black, and alternative is Navy Blue, 5 cm wide. Two separate sections, one non-adjustable, other section adjustable. Non-adjustable section fitted with 5 cm D-ring (chromium plated, galvanised steel) and other with male and female hook and loop straps. Non-adjustable section to be 60 cm long. Adjustable section with hook and loop straps. To be 80 cm long. Hook strap 30 cm long on one end, with loop strap 30 cm long immediately adjacent to hook strap. 20 cm of belt not covered in strap.
5.8.2 | Pelvic straps, basic adjustable, with D-ring fitting
• Poly-cotton webbing, first preference is Black, and alternative is Navy Blue, 3 cm wide. Two separate sections, one non-adjustable, other section adjustable. Non-adjustable section fitted with 3,8 cm D-ring (chromium plated, galvanised steel) and other with male and female hook and loop straps. Non-adjustable section to be 50 cm long. Adjustable section with hook and loop straps to be 56 cm long. Hook strap 8 cm long on one end, with loop strap 8 cm long immediately adjacent to hook strap. 40 cm of belt not covered in strap.
5.8.3 | Pelvic straps, basic adjustable, with male-female plastic buckle system
• Poly-cotton webbing, first preference is Black, and alternative is Navy Blue, 5 cm wide. Two separate sections, one non-adjustable, other section adjustable. Non-adjustable section fitted with female end of buckle. Non-adjustable section to be 60 cm long. Adjustable section to be fitted with male buckle and to be 80 cm long.
5.8.4 | Pelvic straps, padded adjustable
• Poly-cotton webbing, first preference is Black, and alternative is Navy Blue, 5 cm wide. Two separate sections, one non-adjustable, other section adjustable. Non-adjustable section fitted with 5 cm D-ring (chromium plated, galvanised steel) and other with male and female hook and loop straps. Non-adjustable section to be 60 cm long. Adjustable section with hook and loop straps to be 8 cm long. Hook strap 30 cm long on one end, with loop strap 30 cm long immediately adjacent to hook strap. 20 cm of belt not covered in hook and loop straps. Padding (EVA 45) attached to D-ring section, 6 cm wide, 50 cm long of which half extending beyond the D-ring. The padded section to be covered with sleeve to prevent damage / tearing of the EVA 45 pad.
RT233-5.9 | POSTURE SUPPORT SEAT SYSTEM | Indicate the item bidding for below bby stating "Yes bidding for this item". Item left blank will be assumed it is not bidded for | Complies Yes/No | Comments (Technical Specification Deviation)
For use in transporting child in vehicle
Size range of seat system:
The sizes below are maximum usable measurements of complete seating back supports and cushions. Variances of +- 2 cm allowed
• Seat width 30 cm, depth 31 cm, height 50 cm
• Seat width 35 cm, depth 36 cm, height 60 cm
• Seat width 40 cm, depth 41 cm, height 70 cm
• Seat width 40 cm, depth 45 cm, height 80 cm
General requirements:
• No part or component may be changed to deviate from the technical specifications set out in this document. • Bidder to submit accredited test report to ensure device durability • The seating system must be used in conjunction with a standard vehicle seat belt system. The design therefore must be such that it allows for safe and secure use of standard vehicle seat belt system. • System will be removable from vehicle by way of a strapping attachment system
Seat support frame technical specifications:
• Seat and back unit will be 2 separate units which can be set up independently of each other to fit the user’s individual technical specifications • Rigid outer made from fiberglass or similar strength material such as Polyethylene board • Both seat and back units should have rigid outer • Design must allow for supported tilt in space of 20 degrees while maintaining 90-degree hip flexion
Back unit:
• Foam contour moulded back support (see technical specifications under seat system foam) • Contoured lateral trunk and pelvic supports to ensure maximum postural support, but not restrict upper limb function and movement • Thoracic and pelvic posterior support moulded to allow for postural support in neutral alignment • Integrated chest strap – see postural support straps
Head:
• Head support unit • Position adjustable • Moulded to allow for lateral head support • Lateral head support should not restrict vision
Seat unit:
• Contoured sitting surface with build in pre-ischial, adduction and abduction support (see technical specifications under seat system foam). • High density foam pressure relieving cushion. • Integrated pelvic strap – see postural support straps
POSTURAL SUPPORT STRAPS
Chest strap
• 10 cm wide adjustable non-stretch strap • Release system should be made inaccessible to user by means of a protective sleeve
Lap belt Support
• 5 cm wide adjustable non-stretch strap • Release system should be made inaccessible to user by means of a protective sleeve
Seat system foam:
• At least 2 layers Top layer for postural support and pressure relief: • High-density foam 30/20; Hardness factor 18 – 21.6; Density range 27.1 - 32 kg/m³
Base layer:
• High density chip foam (reconstituted flexible PU) foam SC 960; Hardness factor 38 - 60; Density 96+10% kg/m³ (SANS 642 - 1976)
Attachment system to secure seat to vehicle:
• Attachment system should separately secure the back and seat units to the vehicle seat and prevent sliding or movement in any direction

## 9 Category 6
CATEGORY 6
All items being tendered for must comply with the following technical specification, failure to comply with any of the conditions set out below may result in immediate disqualification. Please note that where the technical specification calls for certification, this certification must accompany the bid, failure to provide such certification will result in immediate disqualification. Bidders must give details of items being tendered for in the COMMENTS section where there are variation for further deliberation
TRAY TABLES
RT233-6.1 | TRAY TABLE, BASIC WOOD | Indicate the item bidding for below bby stating "Yes bidding for this item". Item left blank will be assumed it is not bidded for | Complies Yes/No | Comments (Technical Specification Deviation)
TECHNICAL SPECIFICATIONS
Table to fit both basic and adjustable folding frame wheelchairs
Size range:
• To fit wheelchair 61 cm (24") wide (Total dimensions = 70+3 cm x 70+3 cm)
• To fit wheelchair 56 cm (22") wide (Total dimensions = 65+3 cm x 65+3 cm)
• To fit wheelchair 51 cm (20") wide (Total dimensions = 60+3 cm x 60+3 cm)
• To fit wheelchair 46 cm (18") wide (Total dimensions = 55+3 cm x 55+3 cm)
• To fit wheelchair 41 cm (16") wide (Total dimensions = 50+3 cm x 50+3 cm)
• To fit wheelchair 36 cm (14") wide (Total dimensions = 45+3 cm x 45+3 cm)
• To fit wheelchair 30 cm (12") wide with 40 cm seat frame length (Total dimensions = 40+3 cm x 40+3 cm)
• To fit wheelchair 30 cm (12") wide with 30 cm seat frame length (Total dimensions = 35+3 cm x 35+3 cm)
• To fit wheelchair 25 cm (10") wide with 30 cm seat frame length (Total dimensions = 35+3 cm x 35+3 cm)
Description:
• When attached, the extensions around the body should be long enough to reach to the back posts. This will ensure that the elbows are fully supported whilst in the upright sitting position. • When attached, the curve around the body should be such to prevent openings between the body and the tray where the elbows, forearms or hands can slip into.
Thickness:
• Between 0.9cm and 1.5 cm
Material:
• Wood with sealer and wood varnish
Quality:
• Smooth, washable, splinter free surface, splinter proof, rounded edges, easily wipe-able and cleanable
Attachment:
• To attach with hook and loop straps around the armrest of wheelchair • If the tray has a hook ad loop straps section for attachment, the hook ad loop straps should be of soft (loop) type.
6.1. C | CUSTOMISING OPTIONS FOR TRAY TABLE BASIC WOOD
6.1. C.1 | Non-slip vinyl top
6.1. C.2 | Padded top and with non-slip vinyl covering
RT233-6.2 | TRAY TABLE, BASIC PLASTIC | Indicate the item bidding for below bby stating "Yes bidding for this item". Item left blank will be assumed it is not bidded for | Complies Yes/No | Comments (Technical Specification Deviation)
Table to fit both basic and adjustable folding frame wheelchairs
Size range:
• To fit wheelchair 61 cm (24") wide (Total dimensions = 70+3 cm x 70+3 cm)
• To fit wheelchair 56 cm (22") wide (Total dimensions = 65+3 cm x 65+3 cm)
• To fit wheelchair 51 cm (20") wide (Total dimensions = 60+3 cm x 60+3 cm)
• To fit wheelchair 46 cm (18") wide (Total dimensions = 55+3 cm x 55+3 cm)
• To fit wheelchair 41 cm (16") wide (Total dimensions = 50+3 cm x 50+3 cm)
• To fit wheelchair 36 cm (14") wide (Total dimensions = 45+3 cm x 45+3 cm)
• To fit wheelchair 30 cm (12") wide with 40 cm seat frame length (Total dimensions = 40+3 cm x 40+3 cm)
• To fit wheelchair 30 cm (12") wide with 30 cm seat frame length (Total dimensions = 35+3 cm x 35+3 cm)
• To fit wheelchair 25 cm (10") wide with 30 cm seat frame length (Total dimensions = 35+3 cm x 35+3 cm)
Description:
• When attached, the extensions around the body should be long enough to reach to the back posts. This will ensure that the elbows are fully supported whilst in the upright sitting position. • When attached, the curve around the body should be such to prevent openings between the body and the tray where the elbows, forearms or hands can slip into.
Material:
• Clear, see-through durable plastic
Thickness:
• Between 0.9cm and 1.5 cm
Quality:
• Smooth, washable, splinter free surface, splinter proof, rounded edges, easily wipe-able and cleanable
Attachment:
• To attach with hook and loop straps around the armrest of wheelchair • If the tray has a hook ad loop straps section for attachment, the hook and loop straps should be of soft (loop) type.
6.2 C | CUSTOMISING OPTIONS FOR TRAY TABLE PLASTIC
6.2.C.1 | Non-slip vinyl top
6.2.C.2 | Padded top and with non-slip vinyl covering

## 10 Category 7
CATEGORY 7
All items being tendered for must comply with the following technical specification, failure to comply with any of the conditions set out below may result in immediate disqualification. Please note that where the technical specification calls for certification, this certification must accompany the bid, failure to provide such certification will result in immediate disqualification. Bidders must give details of items being tendered for in the COMMENTS section where there are variation for further deliberation
POSITIONERS
RT233-7.1 | SIDELYERS | Indicate the item bidding for below bby stating "Yes bidding for this item". Item left blank will be assumed it is not bidded for | Complies Yes/No | Comments (Technical Specification Deviation)
TECHNICAL SPECIFICATIONS
Type:
• Side positioner consisting of fully adjustable body support • To support body in alternate side lying position while stabilizing head, trunk, and pelvis. Legs to be accommodated in opposing flexor and extending postures. • Provide mattress with back support angles at 90 degrees to lying surface
Size range: (All measurements have a tolerance of + 10 cm)
• Extra small 100 cm long x 60 cm wide, Back height 35 cm (Back height measured from top of lying support)
• Small 120 long cm x 60 wide cm, Back height 35cm (Back height measured from top of lying support)
• Medium 150 long cm x 60 wide cm, Back height 45cm (Back height measured from top of lying support)
• Large 180 long cm x 70 wide cm, Back height 50cm (Back height measured from top of lying support)
Frame:
• Folding frame with webbing, hook, and loop straps attachments • Folding frame to be made from either metal or wood (these are options) • Mattress support should be made from metal strips or webbing or wooden slats
Foam supports:
• High density foam (26/20) Density range 24.1 – 27 Kg/m³ and harness factor 18.0 – 21.6, except head support cushion high density foam (30/9) with density range 27,1 – 32 kg/m³ and hardness factor 7.2 – 10.8. • Removable washable cotton twill • Lying surfaces covered with waterproof stretch terry cloth cotton covers, 2-way stretch, breathable, anti-microbial shield treated, • Both fabrics machine washable and can be dried in tumble drier, • Fabric with hook and loop straps for adjustments • 1 x integrated back and base support • 2 x leg supports (base and small positioning block) Block sizes appropriate to fully support thigh, lower leg, and foot of the user in lying • 1x trunk support with waterproof cover • 1 x head support with waterproof cover
Positioning Blocks Adjustment ranges:
• Chest Block with depth adjustment appropriate to the size (extra-small, small, medium and large) • Pelvic Block with depth adjustment appropriate to the size (extra-small, small, medium, and large) • Knee Block with depth adjustment appropriate to the size (extra-small, small, medium, and large)
Positioning Blocks sizes:
• 1x Chest Block Size for extra-small, small, medium, and large
• 1x Pelvic Block size for extra-small, small, medium, and large
• 1x Knee Block size for extra-small, small, medium, and large
• 1 x head support for extra-small, small, medium, and large
Covers:
• All covers removable for machine washing
RT233-7.2 | PRONE WEDGES | Indicate the item bidding for below bby stating "Yes bidding for this item". Item left blank will be assumed it is not bidded for | Complies Yes/No | Comments (Technical Specification Deviation)
Type:
• Prone wedge with pelvic support and abduction pommel • To support body in prone lying while stabilizing the pelvis. Legs to be accommodated in abduction.
Size range: (Tolerance + 1 cm)
• Extra Small: length 50 cm x width 60 cm x 10 cm front height x 4 cm back height
• Small: length 70 cm x width 60 cm x 15 cm front height x 6 cm back height
• Medium: length 90 cm x width 60 cm x 20 cm front height x 8 cm back height
• Large: length 110 cm x width 70 cm x 25 cm front height x 10 cm back height
Abduction blocks:
• 3 cm front x 8 cm back x 16 cm long x 15 cm high
• 6 cm front x 11 cm x 16 cm long x 15 cm high
• 7 cm front x 13 cm x 16 cm long x 15 cm high
• 7.5 cm front x 15 cm x 20 cm long x 15 cm high
Pelvic stabilisation:
• Pelvic strap 10 cm wide webbing strap with hook and loop straps attachments for extra small, small and medium
• Pelvic strap 15 cm wide webbing strap with hook and loop straps attachments for large
• Sandbag: 1.5 kg, 2 kg, 2.5 kg and 3 kg divided into three sections in relation to wedge size.
Foam supports:
Prone wedge and abduction block: high density compressed (reconstituted flexible PU) foam 960; Hardness factor 38 - 60; Density 96+10% kg/m³. (SANS 642 – 1976).
Cover:
• Vinyl with zip or hand-closed
RT233-7.3 | STANDING FRAMES | Indicate the item bidding for below bby stating "Yes bidding for this item". Item left blank will be assumed it is not bidded for | Complies Yes/No | Comments (Technical Specification Deviation)
RT233-7.3.1 | POSITIONER – UPRIGHT STANDING FRAMES
RT233-7.3.1.1 | UPRIGHT STANDING FRAME
Type:
• Adjustable upright standing support • Full trunk support • Upright posture support. Lower limb, pelvis, trunk, and head support (optional) • Allow age-appropriate abduction of hips • Allow age-appropriate hip and knee extension (this means a degree of flexion for certain ages)
Size range (variance + 5 cm):
• Extra small: Heel to Head 65cm-100cm ;Hip width maximum 30cm
• Small: Heel to head 85cm-120cm ;Hip width maximum 30cm
• Medium: Heel to head 100cm-135cm;Hip width maximum 35cm
• Large: Heel to head 135cm-160cm;Hip width maximum 40cm
Pelvic support:
• Pelvic support adjustable, height and tension
Knee support:
• To provide support under patella (patella should be free). Support height adjustable and adequately padded.
Foot support:
• Forward and backward adjustability and age-appropriate abduction and lateral rotation
Trunk side support:
• Design of standing frame should provide adequate trunk side support.
Lap tray/ worktable:
• A multi-purpose tray / worktable that provides a stable surface for upper limb support. The surface extensions should provide support for the elbows, the curve around the body should be cut in such a way to prevent a large opening between the body and the tray where the elbows, forearms or hands can slip through. • Height, depth, and angle adjustable independently from tilt in space
• Thickness: Between 0.6 cm and 1.5 cm
• Material: Wood (Supawood covered/protected with sealer and wood varnish)/ plastic with washable vinyl top
• Quality: Smooth, washable, splinter free surface, splinter proof, rounded edges, easily wipe-able and cleanable
RT233-7.3.1.2 | UPRIGHT STANDING FRAME – ADULT
Type:
• Adjustable upright standing support • Full trunk support • Upright posture support. Lower limb, pelvis, trunk • Allow age-appropriate abduction of hips • Allow age-appropriate hip and knee extension (this means a degree of flexion for certain ages)
Size range (variance + 5 cm):
| • 78cm width, 84cm depth, height adjustable from 102 -122 cm |  |  | Incorrect Sizes
Pelvic support:
• Pelvic support adjustable, height and tension
Knee support:
• To provide support under patella (patella should be free). Support height adjustable and adequately padded.
Trunk side support:
• Design of standing frame should provide adequate trunk side support.
Lap tray/ worktable:
• When attached, the curve around the body should be cut in such a way to prevent a large opening between the body and the tray where the elbows, forearms or hands can slip through.
Material:
• Wood (Supawood covered/protected with sealer and wood varnish)/ plastic with washable vinyl top
Quality:
• Smooth, washable, splinter free surface, splinter proof, rounded edges, easily wipe-able and cleanable
RT233-7.3.2 | POSITIONER – PRONE STANDING FRAME, ADJUSTABLE FORWARD TILT TO PRONE 30 DEGREES
Type:
• Adjustable prone standing support • Adjustable forward tilt to prone minimum 30 degrees • Adjustable supine standing support- backward tilt to supine minimum 30 degrees • Should be able to tilt to horizontal for ease of transfer of patients in/out of frame • Upright posture support. Lower limb, pelvis, trunk • Allow age-appropriate abduction of hips • Allow age-appropriate hip and knee extension (this means a degree of flexion for certain ages)
Size range (variance + 5 cm):
• Extra small: Heel to Head 65cm-100cm ;Hip width maximum 30cm
• Small: Heel to head 85cm-120cm ;Hip width maximum 30cm
• Medium: Heel to head 100cm-135cm;Hip width maximum 35cm
• Large: Heel to head 135cm-160cm;Hip width maximum 40cm
Tilt:
Variable (mechanical /Pneumatic/hydraulic) and can be set at minimum 30 degrees Needs appropriate back up safety locking system
Pelvic support:
Pelvic support, adjustable, height and tension
Knee support:
• To provide support under patella (patella should be free). Support height adjustable and adequately padded.
Foot support
• Forward and backward adjustability and age-appropriate abduction and lateral rotation
Trunk side support:
• Design of standing frame should provide adequate trunk side support.
Castors:
• Lockable castors should be robust to allow for pushing over uneven surfaces /outside. Minimum size 5cm diameter, 2cm width
Lap tray/ worktable:
• Easily removable (user should be able to fit and remove table by themselves), multi-purpose tray and provides stable surface. When attached, the surface extensions should provide support for the elbows but be sufficiently narrow not to inhibit self-propelling • Table to remain horizontal, independent of the tilt of the frame, dependent on tilt to 180 to assist with transfers in a horizontal position. • When attached, the curve around the body should be cut in such a way to prevent a large opening between the body and the tray where the elbows, forearms or hands can slip through • Height, depth, and angle adjustable independently from tilt in space
Thickness:
Between 0.6 cm and 1.5 cm
Material:
• Wood (Supawood covered/protected with sealer and wood varnish)/ plastic with washable vinyl top
Quality:
• Smooth, washable, splinter free surface, splinter proof, rounded edges, easily wipe-able and cleanable