## 1 Title
ANNEXURE 2 TECHNICAL SPECIFICATION AND PRICING SCHEDULE
BID NUMBER : RT24-1-2021
THE SUPPLY, DELIVERY, INSTALLATION, COMMISSIONING AND MAINTENANCE OF HOSPITAL FURNITURE AND WARD REQUIREMENTS TO THE STATE FOR THE PERIOD ENDING 30 APRIL 2026
CLOSING DATE : 20 OCTOBER 2021
NATIONAL TREASURY
TRANSVERSAL CONTRACTING

## 2 Main Pricing Schedule
All items being tendered for must comply with the following technical specification, failure to comply with any of the conditions set out below may result in immediate disqualification. Please note that where the technical specification calls for certification, this certification must accompany the bid, failure to provide such certification will result in immediate disqualification. Bidders must give details of items being tendered for in the COMMENTS section where there are variation for further deliberation. Complete relevant columns for each line item
1. MEDICAL BOWLS, KIDNEY DISHES, MEASURING JUGS | [18 empty] | Warranty |  | Comprehensive extended warranty
Item No. | Item Code | Description | Series award | Attached Reference | Quantity | Minimum order Quantities (where applicable) | Import Percentage (where applicable) | Local Percentage (where applicable) | Country of Manufacture (Local or Imported) | Foreign Currency Name | Brand Name | Model Number | Delivery Period (Days) | Unit Of Measure (UOM) | UOM Offered by Bidder | Equipment Expected Lifespan (where applicable) | Initial Manufacture Model Range Date | Unit Price | Maintenance Price Year 1 | Maintenance Price Year 2 | Maintenance Price Year 3 | Maintenance Price Year 4 | Maintenance Price Year 5 | Maintenance Price Year 6 | Maintenance Price Year 7 | Total Maintenance Cost | Comments
4 | RT24-01-004ME | Medical Bowl, type B, 200mm
8 | RT24-01-008ME | Kidney Dish, 260mm
13 | RT24-01-013ME | Measuring Jug, Stainless steel, 1000ml
14 | RT24-01-014ME | Measuring Jug, Stainless steel, 2000ml
2. HOSPITAL BEDS, MATTRESSES | [18 empty] | Warranty |  | Comprehensive extended warranty
Item No. | Item Code | Description | Series award | Attached Reference | Quantity | Minimum order Quantities (where applicable) | Import Percentage (where applicable) | Local Percentage (where applicable) | Country of Manufacture (Local or Imported) | Foreign Currency Name | Brand Name | Model Number | Delivery Period (Days) | Unit Of Measure (UOM) | UOM Offered by Bidder | Equipment Expected Lifespan (where applicable) | Initial Manufacture Model Range Date | Unit Price | Maintenance Price Year 1 | Maintenance Price Year 2 | Maintenance Price Year 3 | Maintenance Price Year 4 | Maintenance Price Year 5 | Maintenance Price Year 6 | Maintenance Price Year 7 | Total Maintenance Cost | Comments
1 | RT24-02-001ME | Gynaecology Examination Chair, Electric
2 | RT24-02-002ME | Gynaecology Examination Chair, Hydraulic
3 | RT24-02-003ME | Delivery Bed, Electric | Series award with RT24-02-005ME
4 | RT24-02-004ME | Delivery Bed, Hydraulic | Series award with RT24-02-005ME
5 | RT24-02-005ME | Delivery Bed Mattress | Series award with RT24-02-006ME
6 | RT24-02-006ME | Delivery Bed Mattress cover
7 | RT24-02-007ME | Chemo-Dialysis Chair
8 | RT24-02-008ME | ICU Bariatric Bed, 4-Section, Electric | Series award with RT24-02-010ME
9 | RT24-02-009ME | ICU Bariatric Bed, 4-Section + Scale, Electric | Series award with RT24-02-010ME
10 | RT24-02-010ME | Bariatric Bed Mattress | Series award with RT24-02-011ME
11 | RT24-02-011ME | Bariatric Bed Mattress cover
17 | RT24-02-017ME | Bed, High Low, Electric, 4-section | Series award with RT24-02-022ME
18 | RT24-02-018ME | Highcare Bed, Electric | Series award with RT24-02-024ME
20 | RT24-02-020ME | ICU Bed, 4-Section + Scale, Electric | Series award with RT24-02-024ME
22 | RT24-02-022ME | Mattress for hospital bed
24 | RT24-02-024ME | Pressure Mattress | Series award with RT24-02-025ME
25 | RT24-02-025ME | Pressure Mattress Cover
26 | RT24-02-026ME | Pressure Relief Mattress system with pump (200kg), Overlay
27 | RT24-02-027ME | Pressure Relief Mattress system with pump (200kg), Replacement
28 | RT24-02-028ME | Pressure Relief Mattress system with pump (300kg), Overlay
29 | RT24-02-029ME | Pressure Relief Mattress system with pump (300kg), Replacement
30 | RT24-02-030ME | Adolecent Bed, Electric | Series award with RT24-02-031ME
31 | RT24-02-031ME | Mattress, Adolecent Bed | Series award with RT24-02-032ME
32 | RT24-02-032ME | Mattress Cover, Adolecent Bed
36 | RT24-02-036ME | Infant COT | Series award with RT24-02-038ME
37 | RT24-02-037ME | Infant Cot, Rising backrest | Series award with RT24-02-038ME
38 | RT24-02-038ME | Infant Cot Mattress | Series award with RT24-02-039ME
39 | RT24-02-039ME | Infant Cot Mattress cover
43 | RT24-02-043ME | Trolley/Stretcher, Recovery room | Series award with RT24-02-046ME
44 | RT24-02-044ME | Trolley/Stretcher, emergency, Basic | Series award with RT24-02-046ME
45 | RT24-02-045ME | Trolley/Stretcher, emergency, patient, Electric | Series award with RT24-02-046ME
46 | RT24-02-046ME | Mattress, for trolley/Stretcher | Series award with RT24-02-047ME
47 | RT24-02-047ME | Cover, for mattress
3. BEDPANS AND URINALS | [18 empty] | Warranty |  | Comprehensive extended warranty
Item No. | Item Code | Description | Series award | Attached Reference | Quantity | Minimum order Quantities (where applicable) | Import Percentage (where applicable) | Local Percentage (where applicable) | Country of Manufacture (Local or Imported) | Foreign Currency Name | Brand Name | Model Number | Delivery Period (Days) | Unit Of Measure (UOM) | UOM Offered by Bidder | Equipment Expected Lifespan (where applicable) | Initial Manufacture Model Range Date | Unit Price | Maintenance Price Year 1 | Maintenance Price Year 2 | Maintenance Price Year 3 | Maintenance Price Year 4 | Maintenance Price Year 5 | Maintenance Price Year 6 | Maintenance Price Year 7 | Total Maintenance Cost | Comments
2 | RT24-03-002ME | Bedpan
8 | RT24-03-008ME | Bedpan, with Lid
4. CABINET UNITS | [18 empty] | Warranty |  | Comprehensive extended warranty
Item No. | Item Code | Description | Series award | Attached Reference | Quantity | Minimum order Quantities (where applicable) | Import Percentage (where applicable) | Local Percentage (where applicable) | Country of Manufacture (Local or Imported) | Foreign Currency Name | Brand Name | Model Number | Delivery Period (Days) | Unit Of Measure (UOM) | UOM Offered by Bidder | Equipment Expected Lifespan (where applicable) | Initial Manufacture Model Range Date | Unit Price | Maintenance Price Year 1 | Maintenance Price Year 2 | Maintenance Price Year 3 | Maintenance Price Year 4 | Maintenance Price Year 5 | Maintenance Price Year 6 | Maintenance Price Year 7 | Total Maintenance Cost | Comments
3 | RT24-04-003ME | Cabinet, medicine, mobile (ABS plastic)
8 | RT24-04-008ME | Cabinet, utility, mobile (ABS Plastic) |  | ME00.30.35 - Annexure H
7. LOCKER , TABLE UNITS AND OVERBED | [18 empty] | Warranty |  | Comprehensive extended warranty
Item No. | Item Code | Description | Series award | Attached Reference | Quantity | Minimum order Quantities (where applicable) | Import Percentage (where applicable) | Local Percentage (where applicable) | Country of Manufacture (Local or Imported) | Foreign Currency Name | Brand Name | Model Number | Delivery Period (Days) | Unit Of Measure (UOM) | UOM Offered by Bidder | Equipment Expected Lifespan (where applicable) | Initial Manufacture Model Range Date | Unit Price | Maintenance Price Year 1 | Maintenance Price Year 2 | Maintenance Price Year 3 | Maintenance Price Year 4 | Maintenance Price Year 5 | Maintenance Price Year 6 | Maintenance Price Year 7 | Total Maintenance Cost | Comments
2 | RT24-07-002ME | Locker, bedside, ABS plastic
8. SCALES | [18 empty] | Warranty |  | Comprehensive extended warranty
Item No. | Item Code | Description | Series award | Attached Reference | Quantity | Minimum order Quantities (where applicable) | Import Percentage (where applicable) | Local Percentage (where applicable) | Country of Manufacture (Local or Imported) | Foreign Currency Name | Brand Name | Model Number | Delivery Period (Days) | Unit Of Measure (UOM) | UOM Offered by Bidder | Equipment Expected Lifespan (where applicable) | Initial Manufacture Model Range Date | Unit Price | Maintenance Price Year 1 | Maintenance Price Year 2 | Maintenance Price Year 3 | Maintenance Price Year 4 | Maintenance Price Year 5 | Maintenance Price Year 6 | Maintenance Price Year 7 | Total Maintenance Cost | Comments
3 | RT24-08-003ME | Scale, Analog, Neonatal |  | Appendix H
9. PATIENT TRANSFER | [18 empty] | Warranty |  | Comprehensive extended warranty
Item No. | Item Code | Description | Series award | Attached Reference | Quantity | Minimum order Quantities (where applicable) | Import Percentage (where applicable) | Local Percentage (where applicable) | Country of Manufacture (Local or Imported) | Foreign Currency Name | Brand Name | Model Number | Delivery Period (Days) | Unit Of Measure (UOM) | UOM Offered by Bidder | Equipment Expected Lifespan (where applicable) | Initial Manufacture Model Range Date | Unit Price | Maintenance Price Year 1 | Maintenance Price Year 2 | Maintenance Price Year 3 | Maintenance Price Year 4 | Maintenance Price Year 5 | Maintenance Price Year 6 | Maintenance Price Year 7 | Total Maintenance Cost | Comments
7 | RT24-09-007ME | Transfer sheet/system, 180 cm x 70 cm | Series award with RT24-09-008ME & RT24-09-009ME
8 | RT24-09-008ME | Transfer sheet/system cover, Durable disposable plastic, 180 cm x 70 cm
9 | RT24-09-009ME | Transfer sheet/system cover, Durable reusable plastic, 180 cm x 70 cm
12 | RT24-09-012ME | Patient Hoist with Scale, Electric
13 | RT24-09-013ME | Bariatric Patient Hoist with scale, Electric
11. STOOLS | [18 empty] | Warranty |  | Comprehensive extended warranty
Item No. | Item Code | Description | Series award | Attached Reference | Quantity | Minimum order Quantities (where applicable) | Import Percentage (where applicable) | Local Percentage (where applicable) | Country of Manufacture (Local or Imported) | Foreign Currency Name | Brand Name | Model Number | Delivery Period (Days) | Unit Of Measure (UOM) | UOM Offered by Bidder | Equipment Expected Lifespan (where applicable) | Initial Manufacture Model Range Date | Unit Price | Maintenance Price Year 1 | Maintenance Price Year 2 | Maintenance Price Year 3 | Maintenance Price Year 4 | Maintenance Price Year 5 | Maintenance Price Year 6 | Maintenance Price Year 7 | Total Maintenance Cost | Comments
3 | RT24-11-003ME | Foot stool, Double, Rubber top

## 3 Optional Accessories Prices
All items being tendered for must comply with the following technical specification, failure to comply with any of the conditions set out below may result in immediate disqualification. Please note that where the technical specification calls for certification, this certification must accompany the bid, failure to provide such certification will result in immediate disqualification. Bidders must give details of items being tendered for in the COMMENTS section where there are variation for further deliberation. Complete relevant columns for each line item
1. MEDICAL BOWLS, KIDNEY DISHES, MEASURING JUGS | [18 empty] | Warranty |  | Comprehensive extended warranty
Item No. | Item Code | Description | Series award | Attached Reference | Quantity | Minimum order Quantities (where applicable) | Import Percentage (where applicable) | Local Percentage (where applicable) | Country of Manufacture (Local or Imported) | Foreign Currency Name | Brand Name | Model Number | Delivery Period (Days) | Unit Of Measure (UOM) | UOM Offered by Bidder | Equipment Expected Lifespan (where applicable) | Initial Manufacture Model Range Date | Unit Price | Maintenance Price Year 1 | Maintenance Price Year 2 | Maintenance Price Year 3 | Maintenance Price Year 4 | Maintenance Price Year 5 | Maintenance Price Year 6 | Maintenance Price Year 7 | Total Maintenance Cost | Comments
4 | RT24-01-004ME | Medical Bowl, type B, 200mm
8 | RT24-01-008ME | Kidney Dish, 260mm
13 | RT24-01-013ME | Measuring Jug, Stainless steel, 1000ml
14 | RT24-01-014ME | Measuring Jug, Stainless steel, 2000ml
2. HOSPITAL BEDS, MATTRESSES | [18 empty] | Warranty |  | Comprehensive extended warranty
Item No. | Item Code | Description | Series award | Attached Reference | Quantity | Minimum order Quantities (where applicable) | Import Percentage (where applicable) | Local Percentage (where applicable) | Country of Manufacture (Local or Imported) | Foreign Currency Name | Brand Name | Model Number | Delivery Period (Days) | Unit Of Measure (UOM) | UOM Offered by Bidder | Equipment Expected Lifespan (where applicable) | Initial Manufacture Model Range Date | Unit Price | Maintenance Price Year 1 | Maintenance Price Year 2 | Maintenance Price Year 3 | Maintenance Price Year 4 | Maintenance Price Year 5 | Maintenance Price Year 6 | Maintenance Price Year 7 | Total Maintenance Cost | Comments
1 | RT24-02-001ME | Gynaecology Examination Chair, Electric
2 | RT24-02-002ME | Gynaecology Examination Chair, Hydraulic
3 | RT24-02-003ME | Delivery Bed, Electric | Series award with RT24-02-005ME
4 | RT24-02-004ME | Delivery Bed, Hydraulic | Series award with RT24-02-005ME
5 | RT24-02-005ME | Delivery Bed Mattress | Series award with RT24-02-006ME
6 | RT24-02-006ME | Delivery Bed Mattress cover
7 | RT24-02-007ME | Chemo-Dialysis Chair
8 | RT24-02-008ME | ICU Bariatric Bed, 4-Section, Electric | Series award with RT24-02-010ME
9 | RT24-02-009ME | ICU Bariatric Bed, 4-Section + Scale, Electric | Series award with RT24-02-010ME
10 | RT24-02-010ME | Bariatric Bed Mattress | Series award with RT24-02-011ME
11 | RT24-02-011ME | Bariatric Bed Mattress cover
17 | RT24-02-017ME | Bed, High Low, Electric, 4-section | Series award with RT24-02-022ME
18 | RT24-02-018ME | Highcare Bed, Electric | Series award with RT24-02-024ME
20 | RT24-02-020ME | ICU Bed, 4-Section + Scale, Electric | Series award with RT24-02-024ME
22 | RT24-02-022ME | Mattress for hospital bed
24 | RT24-02-024ME | Pressure Mattress | Series award with RT24-02-025ME
25 | RT24-02-025ME | Pressure Mattress Cover
26 | RT24-02-026ME | Pressure Relief Mattress system with pump (200kg), Overlay
27 | RT24-02-027ME | Pressure Relief Mattress system with pump (200kg), Replacement
28 | RT24-02-028ME | Pressure Relief Mattress system with pump (300kg), Overlay
29 | RT24-02-029ME | Pressure Relief Mattress system with pump (300kg), Replacement
30 | RT24-02-030ME | Adolecent Bed, Electric | Series award with RT24-02-031ME
31 | RT24-02-031ME | Mattress, Adolecent Bed | Series award with RT24-02-032ME
32 | RT24-02-032ME | Mattress Cover, Adolecent Bed
36 | RT24-02-036ME | Infant COT | Series award with RT24-02-038ME
37 | RT24-02-037ME | Infant Cot, Rising backrest | Series award with RT24-02-038ME
38 | RT24-02-038ME | Infant Cot Mattress | Series award with RT24-02-039ME
39 | RT24-02-039ME | Infant Cot Mattress cover
43 | RT24-02-043ME | Trolley/Stretcher, Recovery room | Series award with RT24-02-046ME
44 | RT24-02-044ME | Trolley/Stretcher, emergency, Basic | Series award with RT24-02-046ME
45 | RT24-02-045ME | Trolley/Stretcher, emergency, patient, Electric | Series award with RT24-02-046ME
46 | RT24-02-046ME | Mattress, for trolley/Stretcher | Series award with RT24-02-047ME
47 | RT24-02-047ME | Cover, for mattress
3. BEDPANS AND URINALS | [18 empty] | Warranty |  | Comprehensive extended warranty
Item No. | Item Code | Description | Series award | Attached Reference | Quantity | Minimum order Quantities (where applicable) | Import Percentage (where applicable) | Local Percentage (where applicable) | Country of Manufacture (Local or Imported) | Foreign Currency Name | Brand Name | Model Number | Delivery Period (Days) | Unit Of Measure (UOM) | UOM Offered by Bidder | Equipment Expected Lifespan (where applicable) | Initial Manufacture Model Range Date | Unit Price | Maintenance Price Year 1 | Maintenance Price Year 2 | Maintenance Price Year 3 | Maintenance Price Year 4 | Maintenance Price Year 5 | Maintenance Price Year 6 | Maintenance Price Year 7 | Total Maintenance Cost | Comments
2 | RT24-03-002ME | Bedpan
8 | RT24-03-008ME | Bedpan, with Lid
4. CABINET UNITS | [18 empty] | Warranty |  | Comprehensive extended warranty
Item No. | Item Code | Description | Series award | Attached Reference | Quantity | Minimum order Quantities (where applicable) | Import Percentage (where applicable) | Local Percentage (where applicable) | Country of Manufacture (Local or Imported) | Foreign Currency Name | Brand Name | Model Number | Delivery Period (Days) | Unit Of Measure (UOM) | UOM Offered by Bidder | Equipment Expected Lifespan (where applicable) | Initial Manufacture Model Range Date | Unit Price | Maintenance Price Year 1 | Maintenance Price Year 2 | Maintenance Price Year 3 | Maintenance Price Year 4 | Maintenance Price Year 5 | Maintenance Price Year 6 | Maintenance Price Year 7 | Total Maintenance Cost | Comments
3 | RT24-04-003ME | Cabinet, medicine, mobile (ABS plastic)
8 | RT24-04-008ME | Cabinet, utility, mobile (ABS Plastic) |  | ME00.30.35 - Annexure H
7. LOCKER , TABLE UNITS AND OVERBED | [18 empty] | Warranty |  | Comprehensive extended warranty
Item No. | Item Code | Description | Series award | Attached Reference | Quantity | Minimum order Quantities (where applicable) | Import Percentage (where applicable) | Local Percentage (where applicable) | Country of Manufacture (Local or Imported) | Foreign Currency Name | Brand Name | Model Number | Delivery Period (Days) | Unit Of Measure (UOM) | UOM Offered by Bidder | Equipment Expected Lifespan (where applicable) | Initial Manufacture Model Range Date | Unit Price | Maintenance Price Year 1 | Maintenance Price Year 2 | Maintenance Price Year 3 | Maintenance Price Year 4 | Maintenance Price Year 5 | Maintenance Price Year 6 | Maintenance Price Year 7 | Total Maintenance Cost | Comments
2 | RT24-07-002ME | Locker, bedside, ABS plastic
8. SCALES | [18 empty] | Warranty |  | Comprehensive extended warranty
Item No. | Item Code | Description | Series award | Attached Reference | Quantity | Minimum order Quantities (where applicable) | Import Percentage (where applicable) | Local Percentage (where applicable) | Country of Manufacture (Local or Imported) | Foreign Currency Name | Brand Name | Model Number | Delivery Period (Days) | Unit Of Measure (UOM) | UOM Offered by Bidder | Equipment Expected Lifespan (where applicable) | Initial Manufacture Model Range Date | Unit Price | Maintenance Price Year 1 | Maintenance Price Year 2 | Maintenance Price Year 3 | Maintenance Price Year 4 | Maintenance Price Year 5 | Maintenance Price Year 6 | Maintenance Price Year 7 | Total Maintenance Cost | Comments
3 | RT24-08-003ME | Scale, Analog, Neonatal |  | Appendix H
9. PATIENT TRANSFER | [18 empty] | Warranty |  | Comprehensive extended warranty
Item No. | Item Code | Description | Series award | Attached Reference | Quantity | Minimum order Quantities (where applicable) | Import Percentage (where applicable) | Local Percentage (where applicable) | Country of Manufacture (Local or Imported) | Foreign Currency Name | Brand Name | Model Number | Delivery Period (Days) | Unit Of Measure (UOM) | UOM Offered by Bidder | Equipment Expected Lifespan (where applicable) | Initial Manufacture Model Range Date | Unit Price | Maintenance Price Year 1 | Maintenance Price Year 2 | Maintenance Price Year 3 | Maintenance Price Year 4 | Maintenance Price Year 5 | Maintenance Price Year 6 | Maintenance Price Year 7 | Total Maintenance Cost | Comments
7 | RT24-09-007ME | Transfer sheet/system, 180 cm x 70 cm | Series award with RT24-09-008ME & RT24-09-009ME
8 | RT24-09-008ME | Transfer sheet/system cover, Durable disposable plastic, 180 cm x 70 cm
9 | RT24-09-009ME | Transfer sheet/system cover, Durable reusable plastic, 180 cm x 70 cm
12 | RT24-09-012ME | Patient Hoist with Scale, Electric
13 | RT24-09-013ME | Bariatric Patient Hoist with scale, Electric
11. STOOLS | [18 empty] | Warranty |  | Comprehensive extended warranty
Item No. | Item Code | Description | Series award | Attached Reference | Quantity | Minimum order Quantities (where applicable) | Import Percentage (where applicable) | Local Percentage (where applicable) | Country of Manufacture (Local or Imported) | Foreign Currency Name | Brand Name | Model Number | Delivery Period (Days) | Unit Of Measure (UOM) | UOM Offered by Bidder | Equipment Expected Lifespan (where applicable) | Initial Manufacture Model Range Date | Unit Price | Maintenance Price Year 1 | Maintenance Price Year 2 | Maintenance Price Year 3 | Maintenance Price Year 4 | Maintenance Price Year 5 | Maintenance Price Year 6 | Maintenance Price Year 7 | Total Maintenance Cost | Comments
3 | RT24-11-003ME | Foot stool, Double, Rubber top

## 4 Category 1
CATEGORY 1
All items being tendered for must comply with the following technical specification, failure to comply with any of the conditions set out below may result in immediate disqualification. Please note that where the technical specification calls for certification, this certification must accompany the bid, failure to provide such certification will result in immediate disqualification. Bidders must give details of items being tendered for in the COMMENTS section where there are variation for further deliberation
MEDICAL BOWLS
RT24-01-004ME | Medical Bowl, type B, 200mm | Complies Yes/No | Comments (Technical Specification Deviation)
All items being tendered for must comply with the following technical specification, failure to comply with any of the conditions set out below may result in immediate disqualification. Please note that where the technical specification calls for certification, this certification must accompany the bid, failure to provide such certification will result in immediate disqualification. Bidders must give details of items being tendered for in the COMMENTS section where there are variation for further deliberation
1 | Technical Specification
1.1 | Bowl shall be made from stainless steel
1.2 | To comply to the latest issue of SANS 417, subsection 4 and 5.4, type B, size 4
1.3 | Rim to conform to fig.1(a) of SANS 417
1.4 | Chemical composition of stainless steel shall comply with the requirements given in table 1 of SANS 417, SABS test report to be submitted and only Grade A will be acceptable
1.5 | Dimensions:
1.5.1 | Internal depth 90 - 95mm
1.5.2 | Internal diameter at top 200 mm
1.5.3 | The sides must be straight and parallel
1.6 | The base must be flat
2 | Conformity Compliance (Please attach proof of compliance/Certificate)
2.1 | SANS 417
RT24-01-008ME | Kidney dish, 260mm | Complies Yes/No | Comments (Technical Specification Deviation)
All items being tendered for must comply with the following technical specification, failure to comply with any of the conditions set out below may result in immediate disqualification. Please note that where the technical specification calls for certification, this certification must accompany the bid, failure to provide such certification will result in immediate disqualification. Bidders must give details of items being tendered for in the COMMENTS section where there are variation for further deliberation
1 | Technical Specification
1.1 | Dish shall be made from stainless steel
1.2 | To comply to the latest issue of SANS 417, subsection 4 and 5.2, type A, size 3
1.3 | Rim to conform to fig.1(d) of SANS 417
1.4 | Chemical composition of stainless steel shall comply with the requirements given in table 1 of SANS 417, SABS test report to be submitted and only Grade A will be acceptable
1.5 | Dimensions:
1.5.1 | Overall length: 260 mm (±3 mm)
2 | Conformity Compliance (Please attach proof of compliance/Certificate)
2.1 | SANS 417
MEASURING JUGS
RT24-01-013ME | Measuring Jug, Stainless steel, 1000ml | Complies Yes/No | Comments (Technical Specification Deviation)
All items being tendered for must comply with the following technical specification, failure to comply with any of the conditions set out below may result in immediate disqualification. Please note that where the technical specification calls for certification, this certification must accompany the bid, failure to provide such certification will result in immediate disqualification. Bidders must give details of items being tendered for in the COMMENTS section where there are variation for further deliberation
1 | Technical Specification
1.1 | Stainless steel, to comply with the latest issue of SANS 417, subsection 4 and 5.17
1.2 | Graduation markings shall be on the inside surface and shall be in millilitres
1.3 | The markings shall be in 50 mL steps and every 100 mL shall be denominated
1.4 | Jug capacity: 1000ml
1.5 | Dimensions and graduation markings must be equivalent to the latest issue of SANS 417, subsection 5.17
2 | Conformity Compliance (Please attach proof of compliance/Certificate)
2.1 | SANS 417
RT24-01-014ME | Measuring Jug, Stainless steel, 2000ml | Complies Yes/No | Comments (Technical Specification Deviation)
All items being tendered for must comply with the following technical specification, failure to comply with any of the conditions set out below may result in immediate disqualification. Please note that where the technical specification calls for certification, this certification must accompany the bid, failure to provide such certification will result in immediate disqualification. Bidders must give details of items being tendered for in the COMMENTS section where there are variation for further deliberation
1 | Technical Specification
1.1 | Stainless steel, to comply with the latest issue of SANS 417, subsection 4 and 5.17
1.2 | Graduation markings shall be on the inside surface and shall be in millilitres
1.3 | The markings shall be in 50 mL steps and every 100 mL shall be denominated
1.4 | Jug capacity: 2000ml
1.5 | Dimensions and graduation markings must be equivalent to the latest issue of SANS 417, subsection 5.17
2 | Conformity Compliance (Please attach proof of compliance/Certificate)
1.7 | The markings must be etched or indillible for the life of the unit

## 5 Category 2
CATEGORY 2
HOSPITAL BEDS, MATTRESSES
All items being tendered for must comply with the following technical specification, failure to comply with any of the conditions set out below may result in immediate disqualification. Please note that where the technical specification calls for certification, this certification must accompany the bid, failure to provide such certification will result in immediate disqualification. Bidders must give details of items being tendered for in the COMMENTS section where there are variation for further deliberation
RT24-02-001ME | Gynae Examination Chair, Electric | Complies Yes/No | Comments (Technical Specification Deviation)
1 | Technical Specification
1 | The chair must be designed specifically for the gynaecology examination according to the most advanced production criteria in the field of medical equipments aiming at the maximum functional use and comfort for both operator and patient.
1.1 | The frame shall of mild steel with epoxy/nylon powder-coated finish to comply with SANS 778 paragraph 5.2
1.2 | Solid construction in Mild steel with epoxy powder coated refined with elegant covers moulded in ABS at high resistance
1.3 | Bed must support a patient mass up to a maximum of 250 kg
1.4 | The upholstered parts of the chair must be composed of high-density and undeformable foam pelted with bi-elastic, tear resistant, antimicrobial tissue, very easy to clean and disinfect.
1.5 | Trendelenburg inclination: 0 - 20°
1.6 | Reverse trendelenburg position: 0 - 10°
1.7 | It must be able to achieve backrest angle of 60°
1.8 | Backrest, height and trendelenburg positions must be possible with electric motor using hand control
1.9 | The height adjustment shall be by hand control and operated by means mains power
1.1 | Height adjustment shall be from atleast 700mm to 1100mm
1.11 | Couple of arm rests in metal (removable) designed to grant a safe grip to patient (to be used as handles)
1.12 | 4 Stainless steel rails set on both sides of seat and back rest
1.13 | Single-section leg rest electrically operated granting the chair to reach the bed position
1.14 | Lumbar cushion for a comfortable back position of the patient when lying down
1.15 | Couple of upholstered anatomical leg rests provided with stainless steel supports and manual adjustment in height and longitudinal direction
1.16 | The chair should have following features:
1.16.1 | Footrest with arch-designed supports in stainless steel and with refined foot plates in supple foam and pelted with same sky as the rest of chair’s upholstery – clamp fitting system in to the seat with manual adjustment in height and longitudinal direction
1.16.2 | Footrest must be detacheable
1.16.3 | Practical central cushion for filling the empy space between seat and single-leg rest to grant easy bed position
1.16.4 | Paper-roll holder fixed at the back rest of the chair
1.16.5 | Secretion basin in stainless steel provided with drainage system for waste liquids through a coiled tube
1.16.6 | Detachable head cushion
1.16.7 | Lithotomy poles that are attached to the framework underneath the seat section.
1.16.8 | It must be possible to position the chair in a flat examination couch setup
1.16.9 | Compatable matching operator's chair
2 | WARRANTY & MAINTENANCE
2.1 | A 2-year warranty inclusive maintenance and service (Year 1 to Year 2) should be quoted for, as per the SCC document
2.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
2.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
2.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
3 | Conformity Compliance (Please attach proof of compliance/Certificate)
3.1 | Manufacturer must comply to ISO 13485/ISO 9001 standards
3.2 | Units being quoted for must be CE or SABS certified.
3.3 | SANS 778 (Subsection 5.2)
3.4 | IEC 60601-1-2
RT24-02-002ME | Gynae Examination Chair, Hydraulic | Complies Yes/No | Comments (Technical Specification Deviation)
1 | Technical Specification
1 | The chair must be designed specifically for the gynaecology examination according to the most advanced production criteria in the field of medical equipments aiming at the maximum functional use and comfort for both operator and patient.
1.1 | The frame shall of mild steel with epoxy/nylon powder-coated finish to comply with SANS 778 paragraph 5.2
1.2 | Solid construction in Mild steel with epoxy powder coated refined with elegant covers moulded in ABS at high resistance
1.3 | Bed must support a patient mass up to a maximum of 250 kg
1.4 | The upholstered parts of the chair must be composed of high-density and undeformable foam pelted with bi-elastic, tear resistant, antimicrobial tissue, very easy to clean and disinfect.
1.5 | Trendelenburg inclination: 0 - 20°
1.6 | Reverse trendelenburg position: 0 - 10°
1.7 | It must be able to achieve backrest angle of 60°
1.8 | Backrest, seat, height and trendelenburg positions must be possible using foot switch
1.9 | All positions must be clearly marked and visible on the foot switch
1.1 | Height adjustment shall be from atleast 700mm to 1100mm
1.11 | Couple of arm rests in metal (removable) designed to grant a safe grip to patient (to be used as handles)
1.12 | 4 Stainless steel rails set on both sides of seat and back rest
1.13 | It must be possible to bring the chair to the bed position using a single push on the foot switch
1.14 | Lumbar cushion for a comfortable back position of the patient when lying down
1.15 | Couple of upholstered anatomical leg rests provided with stainless steel supports and manual adjustment in height and longitudinal direction
1.16 | The chair should have following features:
1.16.1 | Footrest with arch-designed supports in stainless steel and with refined foot plates in supple foam and pelted with same sky as the rest of chair’s upholstery – clamp fitting system in to the seat with manual adjustment in height and longitudinal direction
1.16.2 | Footrest must be detacheable
1.16.3 | Practical central cushion for filling the empy space between seat and single-leg rest to grant easy bed position
1.16.4 | Paper-roll holder fixed at the back rest of the chair
1.16.5 | Secretion basin in stainless steel provided with drainage system for waste liquids through a coiled tube
1.16.6 | Detachable head cushion
1.16.7 | Lithotomy poles that are attached to the framework underneath the seat section.
1.16.8 | It must be possible to position the chair in a flat examination couch setup
1.16.9 | Compatable matching operator's chair
2 | WARRANTY & MAINTENANCE
2.1 | A 2-year warranty inclusive maintenance and service (Year 1 to Year 2) should be quoted for, as per the SCC document
2.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
2.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
2.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
3 | Conformity Compliance (Please attach proof of compliance/Certificate)
3.1 | Manufacturer must comply to ISO 13485/ISO 9001 standards
3.2 | Units being quoted for must be CE or SABS certified.
3.3 | SANS 778 (Subsection 5.2)
3.4 | IEC 60601-1-2
RT24-02-003ME | Delivery Bed, Electric | Complies Yes/No | Comments (Technical Specification Deviation)
All items being tendered for must comply with the following technical specification, failure to comply with any of the conditions set out below may result in immediate disqualification. Please note that where the technical specification calls for certification, this certification must accompany the bid, failure to provide such certification will result in immediate disqualification. Bidders must give details of items being tendered for in the COMMENTS section where there are variation for further deliberation
1 | Technical Specification
1.1 | It must be possible to adjust trendelenburg, reverse trendelenburg, backrest and height of the bed
1.2 | Safe workload up to at least 280kg
1.3 | Structural metal parts shall be protected by a durable epoxy/nylon powder-coated finish to comply with SANS 778 which is resistant to chips and scratches
1.4 | Electrically operated height adjustment, backrest, trendelenburg and reverse trendelenburg
1.5 | The bed shall be able to deliver a highly versatile support surface which can be used for labour, delivery and recovery
1.6 | The unit shall allow optimum access to the mother and baby through all stages of the birthing process
1.7 | It shall be possible to convert quickly from a standard bed to a versatile delivery platform, and vice-versa.
1.8 | Manual and electrical emergency CPR must be available
1.9 | All controls shall be easily accessible. No controls shall be behind the headboard
1.1 | Compact size to fit into small delivery rooms
1.11 | Retractable platform, optimizing access to perineal area, mother and baby during delivery procedures
1.12 | Positions: The bed shall offer a choice of safe and comfortable positions that can be adopted during labour and delivery, minimum positions offered shall include:
1.12.1 | Lithotomy position with leg rest and fluid collection bowl
1.12.2 | Semi-chair with high foot support
1.12.3 | Chair with high foot support and hand grips
1.12.4 | Semi- chair with low foot support
1.12.5 | Supported kneeling
1.12.6 | Chair with handgrips and step
1.13 | Removable 2-part mattress, with perineal cut-out seat section
1.14 | Adjustable safety sides
1.15 | Central brake on all four castors
1.16 | Easy steer castor system for optimum mobility/stability
1.17 | Removable, moulded head and foot panels
1.18 | Easy to clean
1.19 | Industrial antistatic, non marking, non staining , ± 125mm castors, with a sealed casing
1.2 | Built-in sockets for IV rod and Patient lifting pole, drop in accessories, no clamps shall be used.
1.21 | Revolving rubber buffer wheels fitted to the four corners of the bed
1.22 | Removable fire retardant head and foot ends
1.23 | Counterbalanced base
1.24 | Storage of accessories: convenient storage for accessories shall be available and included in this offer
1.25 | Complete with a comprehensive range of accessories storable away from the bed, the flexibility and ease-of-use offered
1.26 | Sheet steel mattress platform for improved infection control. No protruding parts that can damage mattress or cover
1.27 | Fire retardant, ABS cover to the undercarriage
1.28 | Battery back-up in the event of mains power failure
1.29 | Dimensions:
1.29.1 | Length: ±2200mm , Width: ± 900mm
1.29.2 | Backrest Angle: 70o +5o
1.29.3 | Height adjustable: 600 – 900mm +40mm
1.29.4 | Trendelenburg angle: 15o +2o
1.29.5 | Reverse trendelenburg angle: 17o +2o
2 | Easy to fit accessories to be supplied as standard:
2.1 | Fixed, heavy duty safety sides
2.2 | Bed control unit / remote handset
2.3 | Electric motors 230V
2.4 | Accessories storage trolley / system
2.5 | Adjustable leg rests
2.6 | IV Rod –height adjustable
2.7 | Hand grips
2.8 | High foot supports
2.9 | Low foot supports
2.1 | Removable step
2.11 | Mattress (2 Sections complete), specification and compatibility as per item RT24-02-005ME
2.12 | Mattress Cover, specification and compatibility as per item RT24-02-006ME
3 | WARRANTY & MAINTENANCE
3.1 | A 2-year warranty inclusive maintenance and service (Year 1 to Year 2) for the main unit should be quoted for, as per the SCC document
3.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
3.2 | Mattress warranty: At least 4 years
3.3 | Cover warranty: At least 2 years
3.4 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
3.4.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
4 | Conformity Compliance (Please attach proof of compliance/Certificate)
4.1 | Manufacturer must comply to ISO 13485/ISO 9001 standards
4.2 | Units being quoted for must be CE or SABS certified.
4.3 | IEC 60601-1-2
RT24-02-004ME | Delivery Bed, Hydraulic | Complies Yes/No | Comments (Technical Specification Deviation)
All items being tendered for must comply with the following technical specification, failure to comply with any of the conditions set out below may result in immediate disqualification. Please note that where the technical specification calls for certification, this certification must accompany the bid, failure to provide such certification will result in immediate disqualification. Bidders must give details of items being tendered for in the COMMENTS section where there are variation for further deliberation
1 | Technical Specification
1.1 | It must be manually operated
1.2 | Safe workload up to at least 280kg
1.3 | Structural metal parts shall be protected by a durable epoxy/nylon powder-coated finish to comply with SANS 778 which is resistant to chips and scratches
1.4 | Height adjustment by hydraulic foot switch from approximately 360 - 760mm
1.5 | The bed shall be able to deliver a highly versatile support surface which can be used for labour, delivery and recovery (All phases of child birth)
1.6 | The unit shall allow optimum access to the mother and baby through all stages of the birthing process
1.7 | It shall be possible to convert quickly from a standard bed to a versatile delivery platform, and vice-versa.
1.8 | Dual side manual CPR levers
1.9 | All controls shall be easily accessible. No controls shall be behind the headboard
1.1 | Compact size to fit into small delivery rooms
1.11 | Retractable platform, optimizing access to perineal area, mother and baby during delivery procedures
1.12 | Positions: The bed shall offer a choice of safe and comfortable positions that can be adopted during labour and delivery, minimum positions offered shall include:
1.12.1 | Lithotomy position with leg rest and fluid collection bowl
1.12.2 | Semi-chair with high foot support
1.12.3 | Chair with high foot support and hand grips
1.12.4 | Semi- chair with low foot support
1.12.5 | Supported kneeling
1.12.6 | Chair with handgrips and step
1.13 | Removable 2-part mattress, with perineal cut-out seat section
1.14 | Adjustable safety sides
1.15 | Central brake on all four castors
1.16 | Easy steer castor system for optimum mobility/stability
1.17 | Removable, moulded head and foot panels
1.18 | Easy to clean
1.19 | Industrial antistatic, non marking, non staining , ± 125mm castors, with a sealed casing
1.2 | Built-in sockets for IV rod and Patient lifting pole, drop in accessories, no clamps shall be used.
1.21 | Revolving rubber buffer wheels fitted to the four corners of the bed
1.22 | Removable fire retardant head and foot ends
1.23 | Counterbalanced base
1.24 | Storage of accessories: convenient storage for accessories shall be available and included in this offer
1.25 | Complete with a comprehensive range of accessories storable away from the bed, the flexibility and ease-of-use offered
1.26 | Sheet steel mattress platform for improved infection control. No protruding parts that can damage mattress or cover
1.27 | Fire retardant, ABS cover to the undercarriage
1.28 | Battery back-up in the event of mains power failure
1.29 | Dimensions:
1.29.1 | Length: ±2200mm , Width: ± 900mm
1.29.2 | Backrest Angle: 70o +5o
1.29.3 | Height adjustable: 260 – 600mm +40mm
1.29.4 | Trendelenburg angle: 15o +2o
1.29.5 | Reverse trendelenburg angle: 17o +2o
2 | Easy to fit accessories to be supplied as standard:
2.1 | Fixed, heavy duty safety sides
2.2 | Bed control unit / remote handset
2.3 | Electric motors 230V
2.4 | Accessories storage trolley / system
2.5 | Adjustable leg rests
2.6 | IV Rod –height adjustable
2.7 | Hand grips
2.8 | High foot supports
2.9 | Low foot supports
2.1 | Removable step
2.11 | Mattress (2 Sections complete), specification and compatibility as per item RT24-02-005ME
2.12 | Mattress Cover, specification and compatibility as per item RT24-02-006ME
3 | WARRANTY & MAINTENANCE
3.1 | A 2-year warranty inclusive maintenance and service (Year 1 to Year 2) should be quoted for, as per the SCC document
3.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
3.2 | Mattress warranty: At least 4 years
3.3 | Cover warranty: At least 2 years
3.4 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
3.4.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
4 | Conformity Compliance (Please attach proof of compliance/Certificate)
4.1 | Manufacturer must comply to ISO 13485/ISO 9001 standards
4.2 | Units being quoted for must be CE or SABS certified.
4.3 | IEC 60601-1-2
RT24-02-005ME | Delivery bed Mattress | Complies Yes/No | Comments (Technical Specification Deviation)
All items being tendered for must comply with the following technical specification, failure to comply with any of the conditions set out below may result in immediate disqualification. Please note that where the technical specification calls for certification, this certification must accompany the bid, failure to provide such certification will result in immediate disqualification. Bidders must give details of items being tendered for in the COMMENTS section where there are variation for further deliberation
1 | Technical Specification
1.1 | Must be compatible with items RT24-02-003ME and RT24-02-004ME
1.2 | Including mattress which adheres to the following specs: Description:
1.2.1 | 2 section / parts which merge into one unit when the retractable platform is joined with the upper section of the bed
1.2.2 | It shall be able to retract the lower part of the mattress with the retractable platform to optimize access to perineal area, mother and baby during delivery procedures
1.2.3 | Flexible high density foam mattress with fully enclosing 2 way stretch cover
1.3 | Mattress core data:
1.3.1 | Mattress depth ±170mm profile cut and zoned melamine free
1.3.2 | Density:38 – 42 Kg/m3
1.3.3 | Hardness: 115 – 150 N
1.4 | Cover data:
1.4.1 | Fully enclosing 2 way stretch, Vapour permeable, Water resistant cover, with an Interstitial Anti-microbial agent and concealed zip, to prevent nosocomial and cross infections
1.4.2 | Fire retardant
2 | Warranty:
2.1 | Mattress: At least 4 years
2.2 | Cover: At least 2 years
3 | Conformity Compliance (Please attach proof of compliance/Certificate)
3.1 | Units being quoted for must be CE or SABS certified.
RT24-02-006ME | Delivery bed Mattress Cover | Complies Yes/No | Comments (Technical Specification Deviation)
All items being tendered for must comply with the following technical specification, failure to comply with any of the conditions set out below may result in immediate disqualification. Please note that where the technical specification calls for certification, this certification must accompany the bid, failure to provide such certification will result in immediate disqualification. Bidders must give details of items being tendered for in the COMMENTS section where there are variation for further deliberation
1 | Technical Specification
1.1 | To comply with the latest issue of SANS 1291-1, proof of compliance must be submitted
1.2 | Must be compatible with RT24-02-005ME
1.3 | Must be fully enclosing 2-way stretch;
1.3.1 | Vapour permeable
1.3.2 | Water resistant
1.3.3 | Interstitial anti-microbial agent
1.3.4 | Radiolucent
1.4 | Where a zip is incorporated it shall be: a heavy duty, large toothed synthetic zip to allow removal of the inner mattress totally covered by a flap extending over the zip to prevent ingress of fluids
2 | Warranty:
2.2 | Cover: At least 2 years
3 | Conformity Compliance (Please attach proof of compliance/Certificate)
3.1 | SANS 1291-1
RT24-02-007ME | Chemotherapy and Dialysis Chair | Complies Yes/No | Comments (Technical Specification Deviation)
All items being tendered for must comply with the following technical specification, failure to comply with any of the conditions set out below may result in immediate disqualification. Please note that where the technical specification calls for certification, this certification must accompany the bid, failure to provide such certification will result in immediate disqualification. Bidders must give details of items being tendered for in the COMMENTS section where there are variation for further deliberation
1 | Technical Specification
1.1 | The specification establishes the requirements for, supply, delivery, end user training, demonstration, commissioning and installation of a fully automated mobile Chemotherapy and Dialysis chair.
1.2 | The chair must offer :
1.2.1 | High stability
1.2.2 | Safety against overturning.
1.2.3 | Motorised position adjustment.
1.3 | High quality upholstery is a must for any good treatment couch being the interface between the couch and patient it must remain clean and resistant to damage.
1.4 | The exchangeable upholstery must designed such that it can be easily removed with few movement for cleaning or maintenance.
1.5 | Blood resistant artificial leather , which is resistant to disinfectants, blood and urine
1.6 | The chair must have a rechargeable battery back-up operation in case of power failure.
1.7 | Length: 2000mm approx
1.8 | Width: 1000mm including armrest approx
1.9 | Height: 520mm – 1020mm approx (20” – 40”)approx
1.1 | Chair must offer maximum patient weight of atleast 200kg
1.11 | The chair must be comfortable with soft armrests, which can be swivelled, folded upwards and adjusted in the height to offer perfect comfort for the patient.
1.12 | Motorized/electronic adjustment of:
1.12.1 | Backrest of 0 - 80°
1.12.2 | Footrest of 0 - 40°
1.12.3 | Height adjustment from 55cm (min) to 75cm (Max)
1.12.4 | Reverse trendelenburg function of 0 - 5°
1.12.5 | Trendelburg function of 0 - 15°
1.13 | Foam armrests tiltable on both sides of the chair to allow the patient to easily enter or leave the chair
1.14 | Antibacteria surface for upholstery with special soft back and seat part
1.15 | Cleanable protective cover for foot part and all armrest
1.16 | Electric height, tilt, foot & back operation
1.17 | Safe working load 180kgs or above
1.18 | Upholstery material
1.19 | The material for Upholstery material should be either UV-proof, impermeable to moisture,able to check for biocompatinilty and toxicity , flame retardant. Please state the material used on the bidded chair.
1.2 | Equipotential bonding cable(to be used when using intercardiac procedures)
1.21 | Hand switch with buttons for basic settings
1.22 | Relax pillow in upholstery colour
1.23 | Elegant underframe- and side-part covers
1.24 | Bed position
2 | The chair must feature the following of which accessories must be included in the bid price:
2.1 | Infusion bottle holder(IV pole) with height adjustment
2.2 | Reading lamp with flexible arm which must be connected directly to the chair.
2.3 | Paper roll holder
2.4 | Two packets of these paper rolls must be supplied with chair and the price be included in the bid price.
2.5 | Push bar made up of stainless steel.
2.6 | Lumbar support
2.7 | Detachable table
2.8 | 2 hook IV pole
2.9 | Hand remote control
3 | WARRANTY & MAINTENANCE
3.1 | A 2-year warranty inclusive maintenance and service (Year 1 to Year 2) should be quoted for, as per the SCC document
3.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
3.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
3.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
4 | Conformity Compliance (Please attach proof of compliance/Certificate)
4.1 | Manufacturer must comply to ISO 13485/ISO 9001 standards
4.2 | Units being quoted for must be CE or SABS certified.
4.3 | SANS 778 (Subsection 5.2)
4.4 | IEC 60601-1-2
RT24-02-008ME | Bariatric Bed, ICU, Electric | Complies Yes/No | Comments (Technical Specification Deviation)
All items being tendered for must comply with the following technical specification, failure to comply with any of the conditions set out below may result in immediate disqualification. Please note that where the technical specification calls for certification, this certification must accompany the bid, failure to provide such certification will result in immediate disqualification. Bidders must give details of items being tendered for in the COMMENTS section where there are variation for further deliberation
1 | Technical Specification
1.1 | Must comply with IEC 60601-2-52 (Particular requirements for basic safety and essential performance of medical beds) paragraphs: 201.1, 201.3, 201.7, 201.9, 201.13, 201.15, annex BB and annex CC
1.2 | Electrically operated variable height adjustment, backrest, knee-break, reverse trendelenburg and trendelenburg
1.3 | The height of the mattress support shall be adjustable in the range 260 – 600 mm (measured from the floor to the top of the mattress support), with tolerance of +40 mm
1.4 | Backrest angle: 65o (+5o)
1.5 | Knee-break angle: 35o (+5o)
1.6 | Trendelenburg angle: 15o (+3o)
1.7 | Reverse trendelenburg angle: 15o (+3o)
1.8 | The bed must be equiped with backrest angle indicator
1.9 | The bed shall be clearly marked with the corresponding Maximum Patient Weight and Safe Working Load in kg.
1.1 | Maximum Patient Weight shall not be less than 450kg
1.11 | Safe Working Load shall not be less than 500kg
1.12 | The bed must have controls for both patient with limited positions and for endusers accessible in both sides with all positions including lock key
1.13 | Mild steel frame with epoxy/nylon powder-coated finish to comply with SANS 778 paragraph 5.2, proof of compliance must be submitted.Epoxy/nylon powder coating colours: white, cream or grey
1.14 | Collapsible safety sides must be of full length to prevent patient from falling off the bed
1.15 | The bed must be fitted with Oxygen cylinder holder
1.16 | To be fitted with castors with a central castor locking system
1.17 | Castors must comply with the latest issue of SANS 621, proof of compliance must be submitted. Where castors are fitted into steel tubular legs, the tube shall be of wall thickness not less than 2,0mm
1.18 | Removable head and foot ends (ABS material may be offered)
1.19 | Mattress support: mattress support other than weldmesh is required, provide details
1.2 | Extension of bed of at least 25cm must be possible and extension mattress piece must be provided
1.21 | The following accessories MUST be offered to fit on the bed. The price for these accessories must not be included on this item bid price.
1.21.1 | Driprod
1.21.2 | Patient lifting pole with chain or strap and handle, must support a mass of 75 kg
1.21.3 | Traction pole with pulleys and weights
1.22 | The bed must have a battery back-up in case of power failure
1.23 | The capacity of battery when fully charged must last for no less than 4 hours
1.26 | The following accessories must meet the above specification, be compatible and supplied with the main unit
2 | Standard Accessories:
2.1 | Mattress, specification and compatibility as per item RT24-02-010ME
2.2 | Mattress Cover, specification and compatibility as per item RT24-02-011ME
3 | WARRANTY & MAINTENANCE
3.1 | A 2-year warranty inclusive maintenance and service (Year 1 to Year 2) should be quoted for, as per the SCC document
3.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
3.2 | Mattress warranty: At least 4 years
3.3 | Cover warranty: At least 2 years
3.4 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
3.4.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
4 | Conformity Compliance (Please attach proof of compliance/Certificate)
4.1 | Manufacturer must comply to ISO 13485/ISO 9001 standards
4.2 | Units being quoted for must be CE or SABS certified.
4.3 | SANS 778 (Subsection 5.2)
4.4 | SANS 621
4.5 | IEC 60601-2-52
4.6 | IEC 60601-1-2
RT24-02-009ME | Bariatric Bed, ICU, Electric, with Intergrated electronic weighting scale | Complies Yes/No | Comments (Technical Specification Deviation)
All items being tendered for must comply with the following technical specification, failure to comply with any of the conditions set out below may result in immediate disqualification. Please note that where the technical specification calls for certification, this certification must accompany the bid, failure to provide such certification will result in immediate disqualification. Bidders must give details of items being tendered for in the COMMENTS section where there are variation for further deliberation
1 | Technical Specification
1.1 | Must comply with IEC 60601-2-52 (Particular requirements for basic safety and essential performance of medical beds) paragraphs: 201.1, 201.3, 201.7, 201.9, 201.13, 201.15, annex BB and annex CC
1.2 | Electrically operated variable height adjustment, backrest, knee-break, reverse trendelenburg and trendelenburg
1.3 | The height of the mattress support shall be adjustable in the range 260 – 600 mm (measured from the floor to the top of the mattress support), with tolerance of +40 mm
1.4 | Backrest angle: 65o (+5o)
1.5 | Knee-break angle: 35o (+5o)
1.6 | Trendelenburg angle: 15o (+3o)
1.7 | Reverse trendelenburg angle: 15o (+3o)
1.8 | The bed must be equiped with backrest angle indicator
1.9 | The bed shall be clearly marked with the corresponding Maximum Patient Weight and Safe Working Load in kg.
1.1 | Maximum Patient Weight shall not be less than 450kg
1.11 | Safe Working Load shall not be less than 500kg
1.12 | The bed must have controls for both patient with limited positions and for endusers accessible in both sides with all positions including lock key
1.13 | Scale
1.13.1 | The bed must have an intergrated weight scale with auto compesation mode.
1.13.2 | The scale must have weighting capacity of atleast 450kg
1.13.3 | The scale must automatic switch-off feature after few seconds of weight display
1.14 | The bed must have controls for both patient with limited positions and for endusers with all positions including lock key
1.15 | Mild steel frame with epoxy/nylon powder-coated finish to comply with SANS 778 paragraph 5.2, proof of compliance must be submitted.Epoxy/nylon powder coating colours: white, cream or grey
1.16 | Collapsible safety sides must be of full length to prevent patient from falling off the bed
1.17 | The bed must be fitted with Oxygen cylinder holder : Cylinder holder must accommodate up to size E cylinders
1.18 | To be fitted with castors with a central castor locking system
1.19 | Castors must comply with the latest issue of SANS 621, proof of compliance must be submitted. Where castors are fitted into steel tubular legs, the tube shall be of wall thickness not less than 2,0mm
1.2 | Removable head and foot ends (ABS material may be offered)
1.21 | Mattress support: mattress support other than weldmesh is required, provide details
1.22 | Extension of bed of at least 25cm must be possible
1.23 | The following accessories MUST be offered to fit on the bed. The price for these accessories must not be included on this item bid price.
1.23.1 | Driprod
1.23.2 | Patient lifting pole with chain or strap and handle, must support a mass of 75 kg
1.23.3 | Traction pole with pulleys and weights
1.24 | The bed must have a battery back-up in case of power failure
1.25 | The capacity of battery when fully charged must last for no less than 4 hours
1.28 | The following accessories must meet the above specification, be compatible and supplied with the main unit
2 | Standard Accessories:
2.1 | Mattress, specification and compatibility as per item RT24-02-010ME
2.2 | Mattress Cover, specification and compatibility as per item RT24-02-011ME
3 | WARRANTY & MAINTENANCE
3.1 | A 2-year warranty inclusive maintenance and service (Year 1 to Year 2) should be quoted for, as per the SCC document
3.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
3.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
3.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
3.3 | Warranty of accessories:
3.3.1 | Mattress: At least 4 years
3.3.2 | Cover: At least 2 years
4 | Conformity Compliance (Please attach proof of compliance/Certificate)
4.1 | Manufacturer must comply to ISO 13485/ISO 9001 standards
4.2 | Units being quoted for must be CE or SABS certified.
4.3 | SANS 778 (Subsection 5.2)
4.4 | SANS 621
4.5 | IEC 60601-2-52
4.6 | IEC 60601-1-2
RT24-02-010ME | Mattress for Bariatric hospital bed | Complies Yes/No | Comments (Technical Specification Deviation)
All items being tendered for must comply with the following technical specification, failure to comply with any of the conditions set out below may result in immediate disqualification. Please note that where the technical specification calls for certification, this certification must accompany the bid, failure to provide such certification will result in immediate disqualification. Bidders must give details of items being tendered for in the COMMENTS section where there are variation for further deliberation
1 | Technical Specification
1.1 | To comply with the latest issue of SANS 883 AND 1291-1 (type 2), except for thickness indicated below. Proof of compliance must be submitted
1.2 | Must be compatible with items RT24-02-008ME, and RT24-02-009ME
1.3 | Density: No less than 42 Kg/m3, Please state the density of the item on offer
1.4 | The mattress with cover must be flexible to accommodate anatomical contouring of all the angles of the bed
1.5 | Mattress must support a patient of at least 450 kg and return to original shape when not in use
1.6 | Mattress must be radiolucent. To be constructed of flexible polyurethane foam complying with class 30, grade no.12 of SANS 883
1.7 | Overall dimensions:
1.7.1 | Length: 1 980 mm (+40mm)
1.7.2 | Width: 900 mm (+20mm)
1.7.3 | Thickness: 150 mm (+5mm)
1.8 | Must include removable cover according to RT24-02-011ME
1.9 | Manufacturer must supply a 4 year warranty on the mattress.
2 | Conformity Compliance (Please attach proof of compliance/Certificate)
2.1 | SANS 1291-1
2.2 | SANS 883
RT24-02-011ME | Cover for mattress, for Bariatric hospital bed | Complies Yes/No | Comments (Technical Specification Deviation)
All items being tendered for must comply with the following technical specification, failure to comply with any of the conditions set out below may result in immediate disqualification. Please note that where the technical specification calls for certification, this certification must accompany the bid, failure to provide such certification will result in immediate disqualification. Bidders must give details of items being tendered for in the COMMENTS section where there are variation for further deliberation
1 | Technical Specification
1.1 | To comply with the latest issue of SANS 1291-1, proof of compliance must be submitted
1.2 | Must be compatible with RT24-02-010ME
1.3 | Must be fully enclosing 2-way stretch;
1.3.1 | Vapour permeable
1.3.2 | Water resistant
1.3.3 | Interstitial anti-microbial agent
1.3.4 | radiolucent
1.4 | Where a zip is incorporated it shall be: a heavy duty, large toothed synthetic zip to allow removal of the inner mattress totally covered by a flap extending over the zip to prevent ingress of fluids
1.5 | Manufacturer must supply a 2 year warranty on the cover
2 | Conformity Compliance (Please attach proof of compliance/Certificate)
2.1 | SANS 1291-1
RT24-02-017ME | Hospital bed, High low, Electric (4- Section) | Complies Yes/No | Comments (Technical Specification Deviation)
All items being tendered for must comply with the following technical specification, failure to comply with any of the conditions set out below may result in immediate disqualification. Please note that where the technical specification calls for certification, this certification must accompany the bid, failure to provide such certification will result in immediate disqualification. Bidders must give details of items being tendered for in the COMMENTS section where there are variation for further deliberation
1 | Technical Specification
1.1 | Must comply with IEC 60601-2-52 (Particular requirements for basic safety and essential performance of medical beds) paragraphs: 201.1, 201.3, 201.7, 201.9, 201.13, 201.15, annex BB and annex CC
1.2 | Electrically operated variable height adjustment, backrest, knee break, reverse trendelenburg and trendelenburg
1.3 | The height of the mattress support shall be adjustable in the range (minimum 260 to 490) and (highest 500 to 900) mm - measured from the floor to the top of the mattress support
1.4 | Backrest angle: 65o (+5o)
1.5 | Trendelenburg angle: 15o (+3o)
1.6 | Reverse trendelenburg angle: 15o (+3o)
1.7 | Knee-break angle: 35o (+5o)
1.8 | One button Semi chair position of the bed must be possible
1.9 | Mattress support: mattress support other than weldmesh is required, provide details
1.1 | Collapsible safety sides must be of full length to prevent patient from falling off the bed. If the bed can be in a full chair position then the safety sides can be half size in length
1.11 | The bed must have controls for both patient with limited positions and for endusers accessible in both sides with all positions including lock key
1.12 | It must be possible to access the central brake mechanism from all side of the bed
1.13 | Mild steel frame with epoxy/nylon powder-coated finish to comply with SANS 778 paragraph 5.2, proof of compliance must be submitted.Epoxy/nylon powder coating colours: white, cream or grey
1.14 | The bed must have maximum patient weight of at least 250 kg
1.15 | The hand control must have graphical representation of all positions
1.16 | A buffer wheel shall be fitted to each of the 4 corners of the bed
1.17 | Electrically adjustable backrest.
1.18 | To be fitted with castors with a central castor locking system
1.19 | Castors must comply with the latest issue of SANS 621, proof of compliance must be submitted.
1.2 | Where castors are fitted into steel tubular legs, the tube shall be of wall thickness not less than 2,0mm
1.21 | Removable head and foot ends (ABS material may be offered)
1.22 | Extension of bed of at least 25cm must be possible and mattress piece of the extension must be provided with the unit
1.23 | The following accessories must be accommodated to fit on the bed. The price for these accessories must not be included on this item bid price.
1.23.1 | Driprod
1.23.2 | Patient lifting pole with chain or strap and handle, must support a mass of 75 kg
1.23.3 | Traction pole with pulleys and weights
1.24 | The bed must be equiped with CPR mechanism
1.25 | The bed dimensions must be compatible with the mattress dimensions
1.28 | The following accessories must meet the above specification, be compatible and supplied with the main unit
2 | Standard Accessories:
2.1 | Mattress, specification and compatibility as per item RT24-02-022ME
2.2 | Mattress Cover, specification and compatibility as per item RT24-02-023ME
3 | WARRANTY & MAINTENANCE
3.1 | A 2-year warranty inclusive maintenance and service (Year 1 to Year 2) for the main unit should be quoted for, as per the SCC document
3.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
3.2 | Mattress warranty of atleast 4 years
3.2 | Cover warranty of atleast 2 years
3.4 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
3.4.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
4 | Conformity Compliance (Please attach proof of compliance/Certificate)
4.1 | Manufacturer must comply to ISO 13485/ISO 9001 standards
4.2 | Units being quoted for must be CE or SABS certified.
4.3 | SANS 778 (Subsection 5.2)
4.4 | SANS 621
4.5 | IEC 60601-2-52
4.6 | IEC 60601-1-2
RT24-02-018ME | Bed, Highcare, Electric | Complies Yes/No | Comments (Technical Specification Deviation)
All items being tendered for must comply with the following technical specification, failure to comply with any of the conditions set out below may result in immediate disqualification. Please note that where the technical specification calls for certification, this certification must accompany the bid, failure to provide such certification will result in immediate disqualification. Bidders must give details of items being tendered for in the COMMENTS section where there are variation for further deliberation
1 | Technical Specification
1.1 | Must comply with IEC 60601-2-52 (Particular requirements for basic safety and essential performance of medical beds) paragraphs: 201.1, 201.3, 201.7, 201.9, 201.13, 201.15, annex BB and annex CC
1.2 | Electrically operated variable height adjustment, backrest, reverse trendelenburg and trendelenburg
1.3 | The height of the mattress support shall be adjustable in the range (minimum 260 to 490) and (highest 500 to 900) mm - measured from the floor to the top of the mattress support
1.4 | Backrest angle: 65o (+5o)
1.5 | Trendelenburg angle: 15o (+3o)
1.6 | Reverse trendelenburg angle: 15o (+3o)
1.7 | The bed must be equiped with backrest angle indicator
1.8 | The bed shall be clearly marked with the corresponding Maximum Patient Weight and Safe Working Load in kg.
1.9 | The bed must have maximum patient weight of at least 250 kg
1.1 | Safe Working Load shall not be less than 280kg
1.11 | The bed must have controls for both patient with limited positions and for endusers accessible in both sides with all positions including lock key
1.12 | The bed must have controls for all position and is a 4-section bed
1.13 | Mild steel frame with epoxy/nylon powder-coated finish to comply with SANS 778 paragraph 5.2, proof of compliance must be submitted.Epoxy/nylon powder coating colours: white, cream or grey
1.14 | Collapsible safety sides must be of full length to prevent patient from falling off the bed. If the bed can be in a full chair position then the safety sides can be half size in length
1.15 | The bed must be fitted with Oxygen cylinder holder
1.16 | To be fitted with castors with a central castor locking system
1.17 | Castors must comply with the latest issue of SANS 621, proof of compliance must be submitted. Where castors are fitted into steel tubular legs, the tube shall be of wall thickness not less than 2,0mm
1.18 | Removable head and foot ends (ABS material may be offered)
1.19 | Mattress support: mattress support other than weldmesh is required, provide details
1.2 | Extension of bed of at least 25cm must be possible and mattress piece of the extension must be provided with the unit
1.21 | The following accessories MUST be offered to fit on the bed. The price for these accessories must not be included on this item bid price.
1.21.1 | Driprod
1.21.2 | Patient lifting pole with chain or strap and handle, must support a mass of 75 kg
1.21.3 | Traction pole with pulleys and weights
1.24 | The following accessories must meet the above specification, be compatible and supplied with the main unit
2 | Standard Accessories:
2.1 | Mattress, specification and compatibility as per item RT24-02-024ME
2.2 | Mattress Cover, specification and compatibility as per item RT24-02-025ME
3 | WARRANTY & MAINTENANCE
3.1 | A 2-year warranty inclusive maintenance and service (Year 1 to Year 2) should be quoted for, as per the SCC document
3.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
3.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
3.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
3.3 | Warranty of accessories:
3.3.1 | Mattress: At least 4 years
3.3.2 | Cover: At least 2 years
4 | Conformity Compliance (Please attach proof of compliance/Certificate)
4.1 | Manufacturer must comply to ISO 13485/ISO 9001 standards
4.2 | Units being quoted for must be CE or SABS certified.
4.3 | SANS 778 (Subsection 5.2)
4.4 | SANS 621
4.5 | IEC 60601-2-52
4.6 | IEC 60601-1-2
RT24-02-020ME | ICU Bed, 4 sections, Electric, with Intergrated electronic weighting scale | Complies Yes/No | Comments (Technical Specification Deviation)
All items being tendered for must comply with the following technical specification, failure to comply with any of the conditions set out below may result in immediate disqualification. Please note that where the technical specification calls for certification, this certification must accompany the bid, failure to provide such certification will result in immediate disqualification. Bidders must give details of items being tendered for in the COMMENTS section where there are variation for further deliberation
1 | Technical Specification
1.1 | Must comply with IEC 60601-2-52 (Particular requirements for basic safety and essential performance of medical beds) paragraphs: 201.1, 201.3, 201.7, 201.9, 201.13, 201.15, annex BB and annex CC
1.2 | Electrically operated variable height adjustment, backrest, knee-break, reverse trendelenburg and trendelenburg
The bed should have the ability to be go on full chair position to enable easy patient egress
1.3 | The height of the mattress support shall be adjustable in the range (minimum 260 to 490) and (highest 500 to 900) mm - measured from the floor to the top of the mattress support
1.4 | Backrest angle: 65o (+5o)
1.5 | Knee-break angle: 35o (+5o)
1.6 | Trendelenburg angle: 15o (+3o)
1.7 | Reverse trendelenburg angle: 15o (+3o)
1.8 | The bed must be equiped with backrest angle indicator
1.9 | The bed shall be clearly marked with the corresponding Maximum Patient Weight and Safe Working Load in kg.
1.1 | Maximum Patient Weight shall not be less than 250kg
1.11 | Safe Working Load shall not be less than 280kg
1.12 | The bed must have an intergrated weight scale with auto compesation mode.
1.13 | The scale must have weighting capacity of atleast 250kg
1.14 | The scale must automatic switch-off feature after few seconds of weight display
1.15 | The bed must have controls for both patient with limited positions and for endusers accessible in both sides with all positions including lock key
1.16 | Mild steel frame with epoxy/nylon powder-coated finish to comply with SANS 778 paragraph 5.2, proof of compliance must be submitted.Epoxy/nylon powder coating colours: white, cream or grey
1.17 | Collapsible safety sides must be of full length to prevent patient from falling off the bed. If the bed can be in a full chair position then the safety sides can be half size in length
1.18 | The bed must be fitted with Oxygen cylinder holder
1.19 | To be fitted with castors with a central castor locking system
1.2 | Castors must comply with the latest issue of SANS 621, proof of compliance must be submitted. Where castors are fitted into steel tubular legs, the tube shall be of wall thickness not less than 2,0mm
1.21 | Removable head and foot ends (ABS material may be offered)
1.22 | Mattress support: mattress support other than weldmesh is required, provide details
1.23 | Extension of bed of at least 25cm must be possible
1.24 | The following accessories MUST be offered to fit on the bed. The price for these accessories must not be included on this item bid price.
1.24.1 | Driprod
1.24.2 | Patient lifting pole with chain or strap and handle, must support a mass of 75 kg
1.24.3 | Traction pole with pulleys and weights
1.25 | The bed must have a battery back-up in case of power failure
1.26 | The capacity of battery when fully charged must last for no less than 4 hours
1.29 | The following accessories must meet the above specification, be compatible and supplied with the main unit
2 | Standard Accessories:
2.1 | Mattress, specification and compatibility as per item RT24-02-024ME
2.2 | Mattress Cover, specification and compatibility as per item RT24-02-025ME
3 | WARRANTY & MAINTENANCE
3.1 | A 2-year warranty inclusive maintenance and service (Year 1 to Year 2) should be quoted for, as per the SCC document
3.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
3.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
3.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
3.3 | Warranty of accessories:
3.3.1 | Mattress: At least 4 years
3.3.2 | Cover: At least 2 years
4 | Conformity Compliance (Please attach proof of compliance/Certificate)
4.1 | Manufacturer must comply to ISO 13485/ISO 9001 standards
4.2 | Units being quoted for must be CE or SABS certified.
4.3 | SANS 778 (Subsection 5.2)
4.4 | SANS 621
4.5 | IEC 60601-2-52
4.6 | IEC 60601-1-2
RT24-02-022ME | Mattress for hospital bed | Complies Yes/No | Comments (Technical Specification Deviation)
All items being tendered for must comply with the following technical specification, failure to comply with any of the conditions set out below may result in immediate disqualification. Please note that where the technical specification calls for certification, this certification must accompany the bid, failure to provide such certification will result in immediate disqualification. Bidders must give details of items being tendered for in the COMMENTS section where there are variation for further deliberation
1 | Technical Specification
1.1 | To comply with the latest issue of SANS 883 AND 1291-1 (type 2), except for thickness indicated below. Proof of compliance must be submitted
1.2 | Must be compatible with items as per profile
1.3 | Density: 32 - 38 Kg/m3, Please state the density of the item on offer
1.4 | The mattress with cover must be flexible to accommodate anatomical contouring of all the angles of the bed
1.5 | Mattress must support a patient of at least 250 kg and return to original shape when not in use
1.6 | Mattress must be radiolucent. To be constructed of flexible polyurethane foam complying with class 30, grade no.12 of SANS 883
1.7 | Overall dimensions:
1.7.1 | Length: 1 980 mm (+40mm)
1.7.2 | Width: 860 mm (+20mm)
1.7.3 | Thickness: 150 mm (+5mm)
1.8 | Must include removable cover according to RT24-02-023ME
1.9 | Manufacturer must supply a 4 year warranty on the mattress.
2 | Conformity Compliance (Please attach proof of compliance/Certificate)
2.1 | SANS 1291-1
2.2 | SANS 883
RT24-02-024ME | Pressure Mattress | Complies Yes/No | Comments (Technical Specification Deviation)
All items being tendered for must comply with the following technical specification, failure to comply with any of the conditions set out below may result in immediate disqualification. Please note that where the technical specification calls for certification, this certification must accompany the bid, failure to provide such certification will result in immediate disqualification. Bidders must give details of items being tendered for in the COMMENTS section where there are variation for further deliberation
1 | Technical Specification
1.1 | It must be a two sided multi-zoned flexible high density foam mattress with fully enclosing 2 way stretch cover
1.2 | The mattress shall be able to turn in 4 different ways (4-way turn), bidder to describe it is achieved
1.3 | Must be a Pressure reducing mattress
1.4 | The mattresses’ cover shall be clearly marked with a turntable guide and washing instruction
1.5 | The foam core must be of very high density and quality
1.6 | Mattress core data required: Bidder to specify the core data of the mattress offered:
1.6.1 | Density: No less than 42 Kg/m3
1.6.2 | Hardness: 115 – 150 N
1.6.3 | Mattress depth: ±170mm profile cut and zoned melamine free
1.6.4 | Classification: V
1.6.5 | Size: 1980 x 860 x 170mm (78” x 33¾” x 6¾”).
1.6.6 | Cover data: Fully enclosing 2 way stretch, Vapour permeable, Water resistant cover, with an Interstitial Anti-microbial agent to prevent nosocomial and cross infections.
1.6.7 | To prevent patient injuries and cross contamination the cover shall have a concealed zip
1.6.8 | The mattress shall be fire retardant
1.6.9 | Patient SWL 250 kg
1.6.10 | Mattress shall be X-Ray translucency and shall be suitable for CPR
2 | Standard supply:
2.1 | Mattress Cover, specification and compatibility as per item RT24-02-025ME
3 | Warranty:
3.1 | Mattress: No less than 4 years
3.2 | Cover: Atleast 2 years
4 | Conformity Compliance (Please attach proof of compliance/Certificate)
4.1 | System must comply to ISO 9000 and ISO 9001 standards
4.2 | Units being quoted for must be CE or SABS certified.
RT24-02-025ME | Pressure Mattress Cover | Complies Yes/No | Comments (Technical Specification Deviation)
All items being tendered for must comply with the following technical specification, failure to comply with any of the conditions set out below may result in immediate disqualification. Please note that where the technical specification calls for certification, this certification must accompany the bid, failure to provide such certification will result in immediate disqualification. Bidders must give details of items being tendered for in the COMMENTS section where there are variation for further deliberation
1 | Technical Specification
1.1 | Cover data: Fully enclosing 2 way stretch,
1.2 | Must be high elongation in warp and welt cover
1.3 | Must have an Interstitial Anti-microbial agent to prevent nosocomial and cross infections
1.4 | Water resistant and made of waterproof fabric
1.5 | Vapour permeable and water vapour breathable of no less than 500 g/m2/24hrs
1.6 | To prevent patient injuries and cross contamination the cover shall have a concealed zip
1.7 | The mattresses’ cover shall be clearly marked with a turntable guide and washing instruction
2 | Warranty:
2.1 | Cover: Atleast 2 years
3 | Conformity Compliance (Please attach proof of compliance/Certificate)
3.1 | Units being quoted for must comply with the SABS fireproof regulation.
RT24-02-026ME | Pressure Relief Mattress system with pump (200kg), Overlay | Complies Yes/No | Comments (Technical Specification Deviation)
All items being tendered for must comply with the following technical specification, failure to comply with any of the conditions set out below may result in immediate disqualification. Please note that where the technical specification calls for certification, this certification must accompany the bid, failure to provide such certification will result in immediate disqualification. Bidders must give details of items being tendered for in the COMMENTS section where there are variation for further deliberation
Technical Specifications
1 | Mattress
1.1 | Mattress shall support patients up to 200 kg.
1.2 | Mattress cell configuration shall consist of:
1.3 | One or more static head cells that must be replaceable
1.4 | Alternating cells along the length of the mattress, that must be replaceable
1.5 | The mattress shall provide a lift of approximately 120 mm
1.6 | The mattress shall be of modular design allowing for easy and cost effective repairs
1.7 | The mattress shall fit on any hospital bed
1.8 | The alternating operating cycle shall be variable up to 10 minutes (approximately)
1.9 | The mattress shall automatically adjust pressures to optimal level according to patient weight, build and position as set by dial on pump
1.1 | The mattress shall have a single hand action, CPR release and deflate within 20 sec
1.11 | Dimensions:
1.11.1 | Length: 2000mm
1.11.2 | Max height at edges must be no less than 115mm
1.11.3 | Width: 900mm
1.11.4 | Cell height inflated: 115mm
2 | Cover
2.1 | The mattress cover shall be two way stretch and loose fitting
2.2 | The cover shall be water resistant
2.3 | The cover shall be vapour permeable
2.4 | The cover shall be bacteriostatic, fungistatic and antimicrobial
2.5 | The cover shall resist development of mould and mildew
2.6 | The cover shall be completely removable
2.7 | The cover shall be fire retardant
3 | Pump
3.1 | The pump shall be ergonomically designed, slim line and unobtrusive with approximate size of 150 x 270 x 160 (mm)
3.2 | The pump shall be whisper quiet and vibration free
3.3 | The pump shall have a carry handle
3.4 | The pump shall have a pressure control dial for setting of patient weight and position
3.5 | The pump shall have a hanging bracket ensuring pump can be fixed to most beds
3.6 | The pump shall have a recessed control panel to prevent accidental operations
3.7 | Interface pressure measured shall be below 20 mmHg for 50% of the cycle time
3.8 | Highest interface pressure shall be 45 mmHg reducing to 5 mmHg during cycle time
3.9 | Pump shall be able to distinguish between genuine alarms and transitory patient movement
3.1 | Pump shall comply with internationally recognized safety standards
3.11 | Operating cycle: variable, 0-10 minutes
4 | Tube Connectors
4.1 | It must be a snap fit tube connectors
4.2 | Connecting and disconnecting of tubing shall be a single handed operation
4.3 | Tubing shall be designed to prevent accidental disconnection
4.4 | Tube set shall be designed to eliminate incorrect connection
4.5 | Tubing shall be anti-kinking preventing blockage
5 | Alarms
5.1 | Alarm conditions to be recorded in memory for diagnostic purposes
5.2 | Mute button to silence audible alarms
5.3 | The following alarms must be available in both audible and visual:
5.4 | Low pressure alarm
5.5 | High pressure alarm
5.6 | Power failure alarm
5.7 | Pump fault alarm
6 | Power Supply
6.1 | The unit shall operate from 220 - 240 V/50Hz mains supply without a transformer and tolerate a 10% variation in voltage
6.2 | The unit shall have a built in rechargeable battery for alarm indication only
7 | Replacement Accessories (To be priced individually)
7.1 | Mattress Cell
7.2 | Mattress
7.3 | Mattress cover
7.4 | Pump
8 | WARRANTY & MAINTENANCE
8.1 | A 2-year warranty inclusive maintenance and service (Year 1 to Year 2) should be quoted for, as per the SCC document
8.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
8.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
8.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
9 | Conformity Compliance (Please attach proof of compliance/Certificate)
9.1 | Manufacturer must comply to ISO 13485/ISO 9001 standards
9.2 | Units being quoted for must be CE or SABS certified.
9.3 | IEC 60601-1-2
RT24-02-027ME | Pressure Relief Mattress system with pump (200kg), Replacement | Complies Yes/No | Comments (Technical Specification Deviation)
All items being tendered for must comply with the following technical specification, failure to comply with any of the conditions set out below may result in immediate disqualification. Please note that where the technical specification calls for certification, this certification must accompany the bid, failure to provide such certification will result in immediate disqualification. Bidders must give details of items being tendered for in the COMMENTS section where there are variation for further deliberation
Technical Specifications
1 | Mattress
1.1 | Mattress shall support patients up to 200 kg.
1.2 | Mattress cell configuration shall consist of:
1.3 | One or more static head cells that must be replaceable
1.4 | Alternating cells along the length of the mattress, that must be replaceable
1.5 | The mattress shall provide a lift of approximately 120 mm
1.6 | The mattress shall be of modular design allowing for easy and cost effective repairs
1.7 | The mattress shall fit on any hospital bed
1.8 | The alternating operating cycle shall be variable up to 10 minutes (approximately)
1.9 | The mattress shall automatically adjust pressures to optimal level according to patient weight, build and position as set by dial on pump
1.1 | The mattress shall have a single hand action, CPR release and deflate within 20 sec
1.11 | Dimensions:
1.11.1 | Length: 2000mm
1.11.2 | Max height at edges must be no less than 200 mm
1.11.3 | Width: 900mm
1.11.4 | Cell height inflated: 200 mm
2 | Cover
2.1 | The mattress cover shall be two way stretch and loose fitting
2.2 | The cover shall be water resistant
2.3 | The cover shall be vapour permeable
2.4 | The cover shall be bacteriostatic, fungistatic and antimicrobial
2.5 | The cover shall resist development of mould and mildew
2.6 | The cover shall be completely removable
2.7 | The cover shall be fire retardant
3 | Pump
3.1 | The pump shall be ergonomically designed, slim line and unobtrusive with approximate size of 150 x 270 x 160 (mm)
3.2 | The pump shall be whisper quiet and vibration free
3.3 | The pump shall have a carry handle
3.4 | The pump shall have a pressure control dial for setting of patient weight and position
3.5 | The pump shall have a hanging bracket ensuring pump can be fixed to most beds
3.6 | The pump shall have a recessed control panel to prevent accidental operations
3.7 | Interface pressure measured shall be below 20 mmHg for 50% of the cycle time
3.8 | Highest interface pressure shall be 45 mmHg reducing to 5 mmHg during cycle time
3.9 | Pump shall be able to distinguish between genuine alarms and transitory patient movement
3.1 | Pump shall comply with internationally recognized safety standards
3.11 | Operating cycle: variable, 0-10 minutes
4 | Tube Connectors
4.1 | It must be a snap fit tube connectors
4.2 | Connecting and disconnecting of tubing shall be a single handed operation
4.3 | Tubing shall be designed to prevent accidental disconnection
4.4 | Tube set shall be designed to eliminate incorrect connection
4.5 | Tubing shall be anti-kinking preventing blockage
5 | Alarms
5.1 | Alarm conditions to be recorded in memory for diagnostic purposes
5.2 | Mute button to silence audible alarms
5.3 | The following alarms must be available in both audible and visual:
5.4 | Low pressure alarm
5.5 | High pressure alarm
5.6 | Power failure alarm
5.7 | Pump fault alarm
6 | Power Supply
6.1 | The unit shall operate from 220 - 240 V/50Hz mains supply without a transformer and tolerate a 10% variation in voltage
6.2 | The unit shall have a built in rechargeable battery for alarm indication only
7 | Replacement Accessories (To be priced individually)
7.1 | Mattress Cell
7.2 | Mattress
7.3 | Mattress cover
7.4 | Pump
8 | WARRANTY & MAINTENANCE
8.1 | A 2-year warranty inclusive maintenance and service (Year 1 to Year 2) should be quoted for, as per the SCC document
8.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
8.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
8.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
9 | Conformity Compliance (Please attach proof of compliance/Certificate)
9.1 | Manufacturer must comply to ISO 13485/ISO 9001 standards
9.2 | Units being quoted for must be CE or SABS certified.
9.3 | IEC 60601-1-2
RT24-02-028ME | Pressure relief mattress system with pump (300kg), Overlay | Complies Yes/No | Comments (Technical Specification Deviation)
All items being tendered for must comply with the following technical specification, failure to comply with any of the conditions set out below may result in immediate disqualification. Please note that where the technical specification calls for certification, this certification must accompany the bid, failure to provide such certification will result in immediate disqualification. Bidders must give details of items being tendered for in the COMMENTS section where there are variation for further deliberation
Technical Specifications
1 | Mattress:
1.1 | Mattress shall support patients up to 300 kg.
1.2 | Mattress cell configuration shall consist of:
1.3 | One or more static head cells that must be replaceable
1.4 | Alternating cells along the length of the mattress, that must be replaceable
1.5 | The mattress shall provide a lift of approximately 135 mm
1.6 | The mattress shall be of modular design allowing for easy and cost effective repairs
1.7 | The mattress shall fit on any hospital bed
1.8 | The alternating operating cycle shall be variable up to 10 minutes (approximately)
1.9 | The mattress shall automatically adjust pressures to optimal level according to patient weight, build and position as set by dial on pump
1.1 | The mattress shall have a single hand action, CPR release and deflate within 15 sec
1.11 | Dimensions:
1.11.1 | Length: 2000mm
1.11.2 | Max height at edges: 135mm
1.11.3 | Width: 850mm
1.11.4 | Cell height inflated: 135mm
2 | Cover:
2.1 | The mattress cover shall be two way stretch and loose fitting
2.2 | The cover shall be water resistant
2.3 | The cover shall be vapour permeable
2.4 | The cover shall be bacteriostatic, fungistatic and antimicrobial
2.5 | The cover shall resist development of mould and mildew
2.6 | The cover shall be completely removable
2.7 | The cover shall be fire retardant
3 | Pump:
3.1 | The pump shall be ergonomically designed, slim line and unobtrusive with approximate size of 350 x 200 x 100 (mm) {LxHxD}
3.2 | The pump shall be whisper quiet and vibration free
3.3 | The pump shall have a carry handle
3.4 | The pump shall have a pressure control dial for setting of patient weight and position
3.5 | The pump shall have a hanging bracket ensuring pump can be fixed to most beds
3.6 | The pump shall have a recessed control panel to prevent accidental operations
3.7 | Interface pressure measured shall be below 20 mmHg for 50% of the cycle time
3.8 | Highest interface pressure shall be 45 mmHg reducing to 5 mmHg during cycle time
3.9 | Pump shall be able to distinguish between genuine alarms and transitory patient movement
3.1 | Pump shall comply with internationally recognized safety standards
3.11 | Operating cycle: variable, 0-10 minutes
4 | Multiway Tube-set:
4.1 | Connecting and disconnecting of tubing shall be a single handed operation
4.2 | Tubing shall be anti-kinking preventing blockage
4.3 | Tube-set shall be interchangeable thus eliminating incorrect connection
5 | Alarms:
5.1 | Alarm conditions to be recorded in memory for diagnostic purposes
5.2 | Mute button to silence audible alarms
5.3 | The following alarms must be available in both audible and visual:
5.4 | Low pressure alarm
5.5 | High pressure alarm
5.6 | Power failure alarm
5.7 | Pump fault alarm
6 | Power Supply:
6.1 | The unit shall operate from 220 - 240 V/50Hz mains supply without a transformer and tolerate a 10% variation in voltage
6.2 | The unit shall have a built in rechargeable battery for alarm indication only
7 | Replacement Accessories (To be priced individually)
7.1 | Mattress Cell
7.2 | Mattress
7.3 | Mattress cover
7.4 | Pump
8 | WARRANTY & MAINTENANCE
8.1 | A 2-year warranty inclusive maintenance and service (Year 1 to Year 2) should be quoted for, as per the SCC document
8.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
8.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
8.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
9 | Conformity Compliance (Please attach proof of compliance/Certificate)
9.1 | Manufacturer must comply to ISO 13485/ISO 9001 standards
9.2 | Units being quoted for must be CE or SABS certified.
9.3 | IEC 60601-1-2
RT24-02-029ME | Pressure relief mattress system with pump (300kg), Replacement | Complies Yes/No | Comments (Technical Specification Deviation)
All items being tendered for must comply with the following technical specification, failure to comply with any of the conditions set out below may result in immediate disqualification. Please note that where the technical specification calls for certification, this certification must accompany the bid, failure to provide such certification will result in immediate disqualification. Bidders must give details of items being tendered for in the COMMENTS section where there are variation for further deliberation
Technical Specifications
1 | Mattress:
1.1 | Mattress shall support patients up to 300 kg.
1.2 | Mattress cell configuration shall consist of:
1.3 | One or more static head cells that must be replaceable
1.4 | Alternating cells along the length of the mattress, that must be replaceable
1.5 | The mattress shall provide a lift of approximately 135 mm
1.6 | The mattress shall be of modular design allowing for easy and cost effective repairs
1.7 | The mattress shall fit on any hospital bed
1.8 | The alternating operating cycle shall be variable up to 10 minutes (approximately)
1.9 | The mattress shall automatically adjust pressures to optimal level according to patient weight, build and position as set by dial on pump
1.1 | The mattress shall have a single hand action, CPR release and deflate within 15 sec
1.11 | Dimensions:
1.11.1 | Length: 2000mm
1.11.2 | Max height at edges: 200mm
1.11.3 | Width: 850mm
1.11.4 | Cell height inflated: 200mm
2 | Cover:
2.1 | The mattress cover shall be two way stretch and loose fitting
2.2 | The cover shall be water resistant
2.3 | The cover shall be vapour permeable
2.4 | The cover shall be bacteriostatic, fungistatic and antimicrobial
2.5 | The cover shall resist development of mould and mildew
2.6 | The cover shall be completely removable
2.7 | The cover shall be fire retardant
3 | Pump:
3.1 | The pump shall be ergonomically designed, slim line and unobtrusive with approximate size of 350 x 200 x 100 (mm) {LxHxD}
3.2 | The pump shall be whisper quiet and vibration free
3.3 | The pump shall have a carry handle
3.4 | The pump shall have a pressure control dial for setting of patient weight and position
3.5 | The pump shall have a hanging bracket ensuring pump can be fixed to most beds
3.6 | The pump shall have a recessed control panel to prevent accidental operations
3.7 | Interface pressure measured shall be below 20 mmHg for 50% of the cycle time
3.8 | Highest interface pressure shall be 45 mmHg reducing to 5 mmHg during cycle time
3.9 | Pump shall be able to distinguish between genuine alarms and transitory patient movement
3.1 | Pump shall comply with internationally recognized safety standards
3.11 | Operating cycle: variable, 0-10 minutes
4 | Multiway Tube-set:
4.1 | Connecting and disconnecting of tubing shall be a single handed operation
4.2 | Tubing shall be anti-kinking preventing blockage
4.3 | Tube-set shall be interchangeable thus eliminating incorrect connection
5 | Alarms:
5.1 | Alarm conditions to be recorded in memory for diagnostic purposes
5.2 | Mute button to silence audible alarms
5.3 | The following alarms must be available in both audible and visual:
5.4 | Low pressure alarm
5.5 | High pressure alarm
5.6 | Power failure alarm
5.7 | Pump fault alarm
6 | Power Supply:
6.1 | The unit shall operate from 220 - 240 V/50Hz mains supply without a transformer and tolerate a 10% variation in voltage
6.2 | The unit shall have a built in rechargeable battery for alarm indication only
7 | Replacement Accessories (To be priced individually)
7.1 | Mattress Cell
7.2 | Mattress
7.3 | Mattress cover
7.4 | Pump
8 | WARRANTY & MAINTENANCE
8.1 | A 2-year warranty inclusive maintenance and service (Year 1 to Year 2) should be quoted for, as per the SCC document
8.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
8.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
8.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
9 | Conformity Compliance (Please attach proof of compliance/Certificate)
9.1 | Manufacturer must comply to ISO 13485/ISO 9001 standards
9.2 | Units being quoted for must be CE or SABS certified.
9.3 | IEC 60601-1-2
RT24-02-030ME | Bed, Adolecent, Electric | Complies Yes/No | Comments (Technical Specification Deviation)
All items being tendered for must comply with the following technical specification, failure to comply with any of the conditions set out below may result in immediate disqualification. Please note that where the technical specification calls for certification, this certification must accompany the bid, failure to provide such certification will result in immediate disqualification. Bidders must give details of items being tendered for in the COMMENTS section where there are variation for further deliberation
1 | Technical Specification
1.1 | Must comply with IEC 60601-2-52 (Particular requirements for basic safety and essential performance of medical beds) paragraphs: 201.1, 201.3, 201.7, 201.9, 201.13, 201.15, annex BB and annex CC
1.2 | Electrical adjustment of height adjustment, backrest, reverse trendelenburg and trendelenburg must be possible
1.3 | The height of the mattress support shall be adjustable in the range 260 – 600 mm (measured from the floor to the top of the mattress support), with tolerance of +40 mm
1.4 | Backrest angle: 65o (+5o)
1.5 | Trendelenburg angle: 15o (+3o)
1.6 | Reverse trendelenburg angle: 15o (+3o)
1.7 | The bed must be equiped with backrest angle indicator
1.8 | The bed shall be clearly marked with the corresponding Maximum Patient Weight and Safe Working Load in kg.
1.9 | It must have a Maximum Patient Weight of at least 150kg
1.1 | Safe Working Load shall not be less than 180kg
1.11 | The bed must have controls for all position
1.12 | Mild steel frame with epoxy/nylon powder-coated finish to comply with SANS 778 paragraph 5.2, proof of compliance must be submitted.Epoxy/nylon powder coating colours: white, cream or grey
1.13 | Collapsible safety sides must be of full length to prevent patient from falling off the bed
1.14 | The bed must have controls for both patient with limited positions and for endusers accessible in both sides with all positions including lock key
1.15 | The bed must be fitted with Oxygen cylinder holder
1.16 | To be fitted with castors with a central castor locking system
1.17 | Castors must comply with the latest issue of SANS 621, proof of compliance must be submitted. Where castors are fitted into steel tubular legs, the tube shall be of wall thickness not less than 2,0mm
1.18 | Removable head and foot ends (ABS material may be offered)
1.19 | Mattress support: mattress support other than weldmesh is required, provide details
1.2 | Extension of bed of at least 20cm must be possible and the additional mattress for the extension must be provided with the unit
1.21 | The following accessories MUST be offered to fit on the bed. The price for these accessories must not be included on this item bid price.
1.21.1 | Driprod
1.21.2 | Patient lifting pole with chain or strap and handle, must support a mass of 75 kg
1.21.3 | Traction pole with pulleys and weights
1.22 | The bed dimensions must be compatible with the mattress dimensions
1.23 | The bed must be equiped with CPR mechanism
1.24 | The following accessories must meet the above specification, be compatible and supplied with the main unit
2 | Standard Accessories:
2.1 | Mattress, specification and compatibility as per item RT24-02-031ME
2.2 | Mattress Cover, specification and compatibility as per item RT24-02-032ME
3 | WARRANTY & MAINTENANCE
3.1 | A 2-year warranty inclusive maintenance and service (Year 1 to Year 2) for the main unit should be quoted for, as per the SCC document
3.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
3.2 | Mattress warranty of atleast 4 years
3.2 | Cover warranty of atleast 2 years
3.4 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
3.4.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
4 | Conformity Compliance (Please attach proof of compliance/Certificate)
4.1 | Manufacturer must comply to ISO 13485/ISO 9001 standards
4.2 | Units being quoted for must be CE or SABS certified.
4.3 | SANS 778 (Subsection 5.2)
4.4 | SANS 621
4.5 | IEC 60601-2-52
4.6 | IEC 60601-1-2
RT24-02-031ME | Bed Mattress, Adolecent | Complies Yes/No | Comments (Technical Specification Deviation)
All items being tendered for must comply with the following technical specification, failure to comply with any of the conditions set out below may result in immediate disqualification. Please note that where the technical specification calls for certification, this certification must accompany the bid, failure to provide such certification will result in immediate disqualification. Bidders must give details of items being tendered for in the COMMENTS section where there are variation for further deliberation
1 | Technical Specification
1.24.1 | To comply with the latest issue of SANS 883 AND 1291-1 (type 2), except for thickness indicated below. Proof of compliance must be submitted
1.24.2 | Must be flexible high-density foam with fully enclosing 2-way stretch cover
1.24.3 | Density: 32 - 38 Kg/m3, Please state the density of the on offer
1.24.4 | The mattress with cover must be flexible to accommodate anatomical contouring of all the angles of the bed
1.24.5 | Mattress must support a patient of at least 150 kg and return to original shape when not in use
1.24.6 | Mattress must be radiolucent. To be constructed of flexible polyurethane foam complying with class 30, grade no.12 of SANS 883
1.24.7 | Overall dimensions:
1.24.7.1 | Length: 1 700 mm (+40mm)
1.24.7.2 | Width: 750 mm (+20mm)
1.24.7.3 | Thickness: 120 mm (+5mm)
1.26 | The following accessories must meet the above specification, be compatible and supplied with the main unit
2 | Standard Accessories:
2.1 | Mattress Cover, specification and compatibility as per item RT24-02-032ME
3 | WARRANTY
3.1 | Mattress warranty of atleast 4 years
3.2 | Cover warranty of atleast 2 years
RT24-02-032ME | Bed Mattress Cover, Adolecent | Complies Yes/No | Comments (Technical Specification Deviation)
All items being tendered for must comply with the following technical specification, failure to comply with any of the conditions set out below may result in immediate disqualification. Please note that where the technical specification calls for certification, this certification must accompany the bid, failure to provide such certification will result in immediate disqualification. Bidders must give details of items being tendered for in the COMMENTS section where there are variation for further deliberation
1 | Technical Specification
1.25.1 | To comply with the latest issue of SANS 1291-1, proof of compliance must be submitted
1.25.2 | Must be fully enclosing 2-way stretch;
1.25.2.1 | Vapour permeable
1.25.2.2 | Water resistant
1.25.2.3 | Interstitial anti-microbial agent
1.25.2.4 | radiolucent
1.25.3 | Where a zip is incorporated it shall be: Concealed heavy duty, large toothed synthetic zip to allow removal of the inner mattress.
2 | WARRANTY
2.1 | Cover warranty of atleast 2 years
RT24-02-036ME | Cot, Infant, complete with mattress | Complies Yes/No | Comments (Technical Specification Deviation)
All items being tendered for must comply with the following technical specification, failure to comply with any of the conditions set out below may result in immediate disqualification. Please note that where the technical specification calls for certification, this certification must accompany the bid, failure to provide such certification will result in immediate disqualification. Bidders must give details of items being tendered for in the COMMENTS section where there are variation for further deliberation
Technical Specification
1 | To comply with the latest issue SANS 521, subsection 5.5, fig. 6
2 | Where the item offered differs from the specification in SANS 521, except for the items specified below, the supplier must indicate the deviation and supply relevant details
3 | Must comply with IEC 60601-2-52 (Particular requirements for basic safety and essential performance of medical beds) paragraphs: 201.1, 201.3, 201.7, 201.9, 201.13, 201.15, annex BB and annex CC
4 | Mild steel frame with epoxy/nylon powder-coated finish to comply with SANS 778 paragraph 5.2, proof of compliance must be submitted.Epoxy/nylon powder coating colours: white, cream or grey
5 | Dimension:
5.1 | Length: 980 mm (± 5 mm)
5.2 | Width: 560 mm (± 5 mm)
5.3 | Mattress support: mattress support other than weldmesh is required, provide details
6 | The following accessories must meet the above specification, be compatible and supplied with the main unit
7 | Standard Accessories:
7.1 | Mattress, specification and compatibility as per item RT24-02-038ME
7.2 | Mattress Cover, specification and compatibility as per item RT24-02-039ME
8 | Warranty
8.1 | Mattress warranty of atleast 4 years
8.2 | Cover warranty of atleast 2 years
9 | Conformity Compliance (Please attach proof of compliance/Certificate)
9.1 | IEC 60601-2-52
9.2 | SANS 778 (Subsection 5.2)
9.3 | SANS 521
RT24-02-037ME | Cot, Infant, with rising backrest and Trendelenberg, with mattress | Complies Yes/No | Comments (Technical Specification Deviation)
All items being tendered for must comply with the following technical specification, failure to comply with any of the conditions set out below may result in immediate disqualification. Please note that where the technical specification calls for certification, this certification must accompany the bid, failure to provide such certification will result in immediate disqualification. Bidders must give details of items being tendered for in the COMMENTS section where there are variation for further deliberation
1 | Technical Specification
1.1 | To comply with the latest issue SANS 521, subsection 5.4, fig. 5
1.2 | Where the item offered differs from the specification in SANS 521, except for the items specified below, the supplier must indicate the deviation and supply relevant details
1.3 | Must comply with IEC 60601-2-52 (Particular requirements for basic safety and essential performance of medical beds) paragraphs: 201.1, 201.3, 201.7, 201.9, 201.13, 201.15, annex BB and annex CC
1.4 | Mild steel frame with epoxy/nylon powder-coated finish to comply with SANS 778 paragraph 5.2, proof of compliance must be submitted.Epoxy/nylon powder coating colours: white, cream or grey
1.5 | Rising backrest support.Trendelenberg and anti-Trendelenberg positions
1.6 | Dimension:
1.6.1 | Length: 980 mm (+ 30 mm)
1.6.2 | Width: 560 mm (+ 20 mm)
1.7 | Thickness: 100 mm (+ 5 mm)
1.8 | The following accessories must meet the above specification, be compatible and supplied with the main unit
2 | Standard Accessories:
2.1 | Mattress, specification and compatibility as per item RT24-02-038ME
2.2 | Mattress Cover, specification and compatibility as per item RT24-02-039ME
3 | Warranty
3.1 | Mattress warranty of atleast 4 years
3.2 | Cover warranty of atleast 2 years
4 | Conformity Compliance (Please attach proof of compliance/Certificate)
4.1 | SANS 778 (subsection 5.2)
4.2 | SANS 521
4.3 | IEC 60601-2-52
RT24-02-038ME | Mattress for Infant Cot | Complies Yes/No | Comments (Technical Specification Deviation)
All items being tendered for must comply with the following technical specification, failure to comply with any of the conditions set out below may result in immediate disqualification. Please note that where the technical specification calls for certification, this certification must accompany the bid, failure to provide such certification will result in immediate disqualification. Bidders must give details of items being tendered for in the COMMENTS section where there are variation for further deliberation
1 | Technical Specification
1.1 | To comply with the latest issue of SANS 883 AND 1291-1 (type 2), except for thickness indicated below. Proof of compliance must be submitted
1.2 | Must be flexible high-density foam with fully enclosing 2-way stretch cover
1.3 | Density: 32 - 38 Kg/m3, Please state the density of the on offer
1.4 | Must be compatible with items RT24-02-036ME, and RT24-02-037ME
1.5 | The mattress with cover must be flexible to accommodate anatomical contouring of all the angles of the bed
1.6 | Mattress must support a patient of at least 180 kg and return to original shape when not in use
1.7 | Mattress must be radiolucent. To be constructed of flexible polyurethane foam complying with class 30, grade no.12 of SANS 883
1.8 | Overall dimensions:
1.8.1 | Length: 940 mm (+ 30 mm)
1.8.2 | Width: 540 mm (+ 20 mm)
1.8.3 | Thickness: 100 mm (+ 5 mm)
1.9 | Must include removable cover according to RT24-02-039ME
1.11 | Manufacturer must supply a 4 year warranty on the mattress.
2 | Conformity Compliance (Please attach proof of compliance/Certificate)
2.1 | SANS 1291-1
2.2 | SANS 883
RT24-02-039ME | Cover for Infant cot mattress | Complies Yes/No | Comments (Technical Specification Deviation)
All items being tendered for must comply with the following technical specification, failure to comply with any of the conditions set out below may result in immediate disqualification. Please note that where the technical specification calls for certification, this certification must accompany the bid, failure to provide such certification will result in immediate disqualification. Bidders must give details of items being tendered for in the COMMENTS section where there are variation for further deliberation
1 | Technical Specification
1.1 | To comply with the latest issue of SANS 1291-1, proof of compliance must be submitted
1.2 | Must be compatible with RT24-02-038ME
1.3 | Must be fully enclosing 2-way stretch;
1.3.1 | Vapour permeable
1.3.2 | Water resistant
1.3.3 | Interstitial anti-microbial agent
1.3.4 | Radiolucent
1.4 | Where a zip is incorporated it shall be: Concealed heavy duty, large toothed synthetic zip to allow removal of the inner mattress.
1.5 | Manufacturer must supply a 2 year warranty on the cover
2 | Conformity Compliance (Please attach proof of compliance/Certificate)
2.1 | SANS 1291-1
RT24-02-043ME | Recovery Room Trolley/Stretcher | Complies Yes/No | Comments (Technical Specification Deviation)
All items being tendered for must comply with the following technical specification, failure to comply with any of the conditions set out below may result in immediate disqualification. Please note that where the technical specification calls for certification, this certification must accompany the bid, failure to provide such certification will result in immediate disqualification. Bidders must give details of items being tendered for in the COMMENTS section where there are variation for further deliberation
Technical Specifications
1 | Mattress
1.1 | Adjustable height , hydraulic, with foot pedal in both sides
1.2 | Trolley structure must have shocks/springs to provide comfort ride for patient
1.3 | Mild steel with epoxy/nylon powder-coated finish to comply with SANS 778 paragraph 5.2
1.4 | Trolley must be able to support bariatric patient weight of minimum 300 kg
1.5 | Must have Rising back rest, gas spring assisted
1.6 | Trendelenberg and anti-Trendelenberg tilting of +18o must be available using a foot pedal
1.7 | Safety sides must be light weight and collapsible
1.8 | Must have Oxygen cylinder holder
1.9 | Must have Drip rod included
1.1 | Must have Castors of 200 mm diameter
1.11 | Castors must comply with the latest issue of SANS 621
1.12 | Castors must be easy to maneuver in all direction, please state how that maneuvering is achieved
1.13 | The trolley must have padded push handles in both sides with option to store away and retract them
1.14 | The trolley must be equiped with roller bumpers in all four corners
1.15 | It must be easy to activate brake-and-go pedal
1.16 | X-ray cassette system (for full body length)
1.16 | The following accessories must be compatible and supplied with the main unit
2 | Standard Accessories:
2.1 | Mattress as per specification of item RT24-12-030ME
2.2 | Mattress Cover as per specification of item RT24-12-031ME
3 | Conformity Compliance (Please attach proof of compliance/Certificate)
3.1 | SANS 621
3.2 | SANS 778 (subsection 5.2)
3 | WARRANTY & MAINTENANCE
3.1 | A 2-year warranty inclusive maintenance and service (Year 1 to Year 2) should be quoted for, as per the SCC document
3.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Hydraulic cylinder), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
3.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
3.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Hydraulic cylinder), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
RT24-02-044ME | Emergency Patient Trolley/Stretcher, Basic | Complies Yes/No | Comments (Technical Specification Deviation)
All items being tendered for must comply with the following technical specification, failure to comply with any of the conditions set out below may result in immediate disqualification. Please note that where the technical specification calls for certification, this certification must accompany the bid, failure to provide such certification will result in immediate disqualification. Bidders must give details of items being tendered for in the COMMENTS section where there are variation for further deliberation
Technical Specifications
1.1 | Adjustable height , hydraulic, with foot pedal
1.2 | Rising back rest, gas spring assisted
1.3 | X-ray cassette system (for full body length)
1.4 | Must have Trendellenburg and Anti-Trendellenburg mechanism with no less than 12 degrees movement
1.5 | Must be able support a patient of mass 180 kg
1.6 | Must have Oxygen cylinder holder
1.7 | Must be supplied with Utility tray
1.8 | Telescopic/adjustable drip rod at head end of trolley
1.9 | Side rails must be collapsable
1.1 | Dimensions (approximate):
1.10.1 | Length: 2 100 mm
1.10.2 | Width: 750 mm
1.10.3 | Height: Adjustable from 600 to 900 mm
1.11 | Castors at least 200 mm diameter
1.12 | Centralized braking system with directional lock
1.13 | Castors must comply with the latest issue of SANS 621
1.14 | Where castors are fitted into steel tubular legs, the tube shall be of wall thickness not less than 2,0mm and castors shall be fixed to the tube by one of the following methods:
1.14.1 | a. Solid plug (long) complying with SANS 621 subsection 3.5.6 or
1.14.2 | b. Screwed into 35mm long sleeves welded into the tubular members and locked in an acceptable manner
1.14.3 | c. Rubber or plastic expanding sleeves for fitting castors are not acceptable
1.15 | Trolley must be easy to maneuver and steer in all direction using retractable 5th wheel
1.16 | The trolley must have padded push handles in both sides with option to store away and retract them
1.17 | The trolley must be equiped with roller bumpers in all four corners
1.18 | It must be easy to activate brake-and-go pedal
1.19 | Manual CPR lever must be accessible from both side of the trolley
1.2 | The following accessories must be compatible and supplied with the main unit
2 | Standard Accessories:
2.1 | Mattress as per specification of item RT24-12-030ME
2.2 | Mattress Cover as per specification of item RT24-12-031ME
3 | Conformity Compliance (Please attach proof of compliance/Certificate)
3.1 | SANS 621
3.2 | SANS 883
3.3 | SANS 1291-1
3 | WARRANTY & MAINTENANCE
3.1 | A 2-year warranty inclusive maintenance and service (Year 1 to Year 2) should be quoted for, as per the SCC document
3.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Hydraulic cylinder), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
3.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
3.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Hydraulic cylinder), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
RT24-02-045ME | Emergency Patient Trolley/Stretcher, Electric | Complies Yes/No | Comments (Technical Specification Deviation)
All items being tendered for must comply with the following technical specification, failure to comply with any of the conditions set out below may result in immediate disqualification. Please note that where the technical specification calls for certification, this certification must accompany the bid, failure to provide such certification will result in immediate disqualification. Bidders must give details of items being tendered for in the COMMENTS section where there are variation for further deliberation
1 | Technical Specification
1.1 | Adjustable height , electric, with controls accessible in both sides
1.2 | Trolley structure must have shocks/springs to provide comfort ride for patient
1.3 | Electric operated rising back rest from 0o - 90o
1.4 | X-ray cassette system (for full body length)
1.5 | Must have Trendelenburg and Anti-Trendelenburg mechanism also electric operated with no less than 18o movement
1.6 | Must be able support a bariatric patient of mass atleast 300 kg
1.7 | Must have Oxygen cylinder holder
1.8 | Must be supplied with Utility tray
1.9 | Telescopic/adjustable drip rod at head end of trolley
1.1 | Side rails must be collapsable
1.11 | Dimensions (approximate):
1.11.1 | Length: 2 100 mm
1.11.2 | Width: 750 mm
1.11.3 | Height: Adjustable from 550 to 900 mm
1.12 | Castors at least 200 mm diameter
1.13 | Centralized braking system with directional lock
1.14 | Castors must comply with the latest issue of SANS 621
1.15 | Trolley must be easy to maneuver and steer in all direction using retractable 5th wheel
1.16 | The trolley must have padded push handles in both sides with option to store away and retract them
1.17 | The trolley must be equiped with roller bumpers in all four corners
1.18 | It must be easy to activate brake-and-go pedal
1.19 | Electrically powered by 220V using retractable power cord
1.2 | The unit must have a back-up battery
1.21 | The following accessories must be compatible and supplied with the main unit
2 | Standard Accessories:
2.1 | Mattress as per specification of item RT24-12-030ME
2.2 | Mattress Cover as per specification of item RT24-12-031ME
3 | WARRANTY & MAINTENANCE
3.1 | A 2-year warranty inclusive maintenance and service (Year 1 to Year 2) should be quoted for, as per the SCC document
3.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Hydraulic cylinder), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
3.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
3.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Hydraulic cylinder), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
3.3 | Warranty of accessories:
3.3.1 | Mattress: At least 4 years
3.3.2 | Cover: At least 2 years
4 | Conformity Compliance (Please attach proof of compliance/Certificate)
4.1 | Manufacturer must comply to ISO 13485/ISO 9001 standards
4.2 | Units being quoted for must be CE or SABS certified.
4.3 | SANS 778 (Subsection 5.2)
4.4 | SANS 621
4.5 | IEC 60601-1-2
RT24-02-046ME | Trolley/Stretcher Mattress | Complies Yes/No | Comments (Technical Specification Deviation)
All items being tendered for must comply with the following technical specification, failure to comply with any of the conditions set out below may result in immediate disqualification. Please note that where the technical specification calls for certification, this certification must accompany the bid, failure to provide such certification will result in immediate disqualification. Bidders must give details of items being tendered for in the COMMENTS section where there are variation for further deliberation
1 | Technical Specification
1 | Mattress:
1.1 | Mattress to be a good quality high density foam of no less than 32 Kg/m3
1.2 | Must be constructed of high resilience polyurethane foam, to comply with SANS 883, Grade no. 12, Class 30
1.3 | Mattress must be radio translucent, waterproof, non-slip or with retaining straps/hooks
1.3.1 | Dimension:
1.3.2 | Length: 1 950 mm
1.3.3 | Width: 600 mm
1.4.3 | Thickness: minimum 80 mm
1.4 | The mattress must be supplied complete with the cover which meet the RT24-12-032ME specification
2 | Warranty:
2.1 | Mattress: At least 4 years
2.2 | Cover: At least 2 years
RT24-02-047ME | Trolley Mattress Cover | Complies Yes/No | Comments (Technical Specification Deviation)
All items being tendered for must comply with the following technical specification, failure to comply with any of the conditions set out below may result in immediate disqualification. Please note that where the technical specification calls for certification, this certification must accompany the bid, failure to provide such certification will result in immediate disqualification. Bidders must give details of items being tendered for in the COMMENTS section where there are variation for further deliberation
1 | Technical Specification
1.1 | Cover to comply with the latest issue of SANS 1291-1
1.2 | Must be removable and complying with the following characteristcs;
1.2.1 | Vapour permeable
1.2.2 | Water resistant
1.2.3 | Interstitial anti-microbial agent
1.2.4 | Radiolucent
1.3 | Where a zip is incorporated it shall be: Concealed heavy duty, large toothed synthetic zip to allow removal of the inner mattress.
1.4 | Cover must be non-slip
1.5 | The cover must fit the mattress specified in item RT24-12-031ME
2 | Warranty:
2.1 | Cover: At least 2 years

## 6 Category 3
CATEGORY 3
BEDPANS AND URINALS
RT24-03-002ME | Bedpan | Complies Yes/No | Comments (Technical Specification Deviation)
All items being tendered for must comply with the following technical specification, failure to comply with any of the conditions set out below may result in immediate disqualification. Please note that where the technical specification calls for certification, this certification must accompany the bid, failure to provide such certification will result in immediate disqualification. Bidders must give details of items being tendered for in the COMMENTS section where there are variation for further deliberation
1 | Technical Specification
1.1 | To comply with the latest issue of SANS 417, subsection 4 and 5.8
1.2 | Stainless steel, to comply with the latest issue of SANS 417, subsection 4.1.2
1.3 | The thickness of the material must be atleast 0.7mm
1.4 | Size: Adult
2 | Conformity Compliance (Please attach proof of compliance/Certificate)
2.1 | SANS 417
RT24-03-008ME | Urinal, Male, Stainless Steel | Complies Yes/No | Comments (Technical Specification Deviation)
All items being tendered for must comply with the following technical specification, failure to comply with any of the conditions set out below may result in immediate disqualification. Please note that where the technical specification calls for certification, this certification must accompany the bid, failure to provide such certification will result in immediate disqualification. Bidders must give details of items being tendered for in the COMMENTS section where there are variation for further deliberation
1 | Technical Specification
1.1 | The design and dimensions to conform to type A or B of the latest issue of SANS 417, subsection 5.9
1.2 | Material: Stainless steel, type A or B, to comply with subsection 4.1.2 of SANS 417
1.3 | Capacity: 1000ml
2 | Conformity Compliance (Please attach proof of compliance/Certificate)
2.1 | SANS 417

## 7 Category 4
CATEGORY 4
CABINET UNITS
RT24-04-003ME | Medicine Cabinet, Mobile (ABS) | Complies Yes/No | Comments (Technical Specification Deviation)
All items being tendered for must comply with the following technical specification, failure to comply with any of the conditions set out below may result in immediate disqualification. Please note that where the technical specification calls for certification, this certification must accompany the bid, failure to provide such certification will result in immediate disqualification. Bidders must give details of items being tendered for in the COMMENTS section where there are variation for further deliberation
1 | Technical Specification
1.1 | Top surface to be flat, suitable for writting and dispensing
1.2 | The unit must have at least five drawers as a minimum with top 4 having dividers
1.3 | Two large storage bottom drawers, no dividers
1.4 | Drawers must have smooth, durable sliding mechanism with stop to prevent total extraction of drawer
1.5 | Dimensions:
1.5.1 | Height: 950 mm (+50mm)
1.5.2 | Width: 750 mm
1.5.3 | Depth: 500 mm
1.6 | Mounted on 4 heavy duty swivel castors 125 mm diameter, two swivel and two braking (Tolerance ± 20mm):
1.7 | Castors must comply wih SANS 621
1.8 | Castors must be antistatic
1.9 | Where castors are fitted into steel tubular legs, the tube shall be of wall thickness not less than 2,0mm and castors shall be fixed to the tube by one of the following methods:
1.9.1 | Solid plug (long) complying with SANS 621 subsection 3.5.6 or
1.9.2 | Screwed into 35mm long sleeves welded into the tubular members and locked in an acceptable manner
1.9.3 | Rubber or plastic expanding sleeves for fitting castors are not acceptable
1.1 | Materials which are easily cleaned and of lasting finish, and low noise generating nature, shall be used
2 | Conformity Compliance (Please attach proof of compliance/Certificate)
2.1 | SANS 621
RT24-04-008ME | Utility Cabinet, Mobile (ABS plastic) | Complies Yes/No | Comments (Technical Specification Deviation)
All items being tendered for must comply with the following technical specification, failure to comply with any of the conditions set out below may result in immediate disqualification. Please note that where the technical specification calls for certification, this certification must accompany the bid, failure to provide such certification will result in immediate disqualification. Bidders must give details of items being tendered for in the COMMENTS section where there are variation for further deliberation
1 | Technical Specification
1.1 | To serve the same function as drawing on Appendix H, drawing ME 00.30.35, see attached
1.2 | ABS Material where appropriate
1.3 | The cabinet must be fitted with two pull out or swivel drawers, one on either sides which must comply with the latest issue of SANS 343, par. 3.3.4
1.4 | All drawers must have a durable sliding mechanism with stop to prevent total extraction of drawer
1.5 | The doors of the top compartment must be fitted with a door stopper mechanism (hard rubber and countersink screws)
1.6 | The doors must be fitted with durable hinges
1.7 | The retractable worktop must be of rigid construction and capable of being extended (not less than 300 mm) beyond the face of the cabinet without canting
1.8 | Mounted on 4 heavy duty swivel castors 125 mm diameter, two swivel and two braking
1.9 | Castors must comply with the latest issue SANS 621
1.1 | Castors must be antistic:
1.10.1 | Solid plug (long) complying with SANS 621 subsection 3.5.6 or
1.10.2 | Screwed into 35mm long sleeves welded into the tubular members and locked in an acceptable manner
1.10.3 | Rubber or plastic expanding sleeves for fitting castors are not acceptable
1.11 | Materials which are easily cleaned and of lasting finish, and low noise generating nature, shall be used.
2 | Conformity Compliance (Please attach proof of compliance/Certificate)
2.1 | SANS 621
2.2 | SANS 343

## 8 Category 7
CATEGORY 7
LOCKER , TABLE UNITS AND OVERBED
RT24-07-002ME | Bedside Locker, ABS plastic | Complies Yes/No | Comments (Technical Specification Deviation)
All items being tendered for must comply with the following technical specification, failure to comply with any of the conditions set out below may result in immediate disqualification. Please note that where the technical specification calls for certification, this certification must accompany the bid, failure to provide such certification will result in immediate disqualification. Bidders must give details of items being tendered for in the COMMENTS section where there are variation for further deliberation
1 | Technical Specification
1.1 | To comply with the latest issue of SANS 343, Type 1, with the following exception:
1.2 | Material: Durable ABS
1.3 | The ABS must be resistant to physical impact, chemical corrosion, and high thermal stress.
1.4 | The unit must withstand heavy-use and adverse environmental condition
1.5 | Must have a rail on both sides and Pull-out drawer
1.6 | Locker mounted on 4 castors.Castors must comply with the latest issue of SANS 621
1.7 | Where castors are fitted into steel tubular legs, the tube shall be of wall thickness not less than 2,0mm and castors shall be fixed to the tube by one of the following methods:
1.7.1 | a. Solid plug (long) complying with SANS 621 subsection 3.5.6 or
1.7.2 | b. Screwed into 35mm long sleeves welded into the tubular members and locked in an acceptable manner
1.7.3 | c. Rubber or plastic expanding sleeves for fitting castors are not acceptable
1.8 | Drawer must have a durable sliding mechanism and stop to prevent total extraction of drawer
1.9 | Materials which are easily cleaned and of lasting finish, and low noise generating nature, must be used
2 | Conformity Compliance (Please attach proof of compliance/Certificate)
2.1 | SANS 621
2.2 | SANS 343

## 9 Category 8
CATEGORY 8
SCALES
RT24-08-003ME | Scale, Analog, Neonatal | Complies Yes/No | Comments (Technical Specification Deviation)
All items being tendered for must comply with the following technical specification, failure to comply with any of the conditions set out below may result in immediate disqualification. Please note that where the technical specification calls for certification, this certification must accompany the bid, failure to provide such certification will result in immediate disqualification. Bidders must give details of items being tendered for in the COMMENTS section where there are variation for further deliberation
Product Offered
Brand:
Model:
Initial Manufacturing year of Model:
1 | Technical Specification
1.1 | To comply with the specifications on Appendix H, Type 2, see attached
1.2 | Weighing capacity: 20 kg
1.3 | The unit must have visible dial display
1.4 | Spring lever mechanism with vanadium steel spring
1.5 | Heavy, hardened steel pivots and bearing
1.6 | The unit to have zeroing or calibration dial
1.7 | Scale must be of robust construnction
2 | WARRANTY & MAINTENANCE
2.1 | A 2-year warranty inclusive maintenance and service (Year 1 to Year 2) should be quoted for, as per the SCC document
2.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
3 | Conformity Compliance (Please attach proof of compliance/Certificate)
3.1 | Manufacturer must comply to ISO 13485/ISO 9001 standards
3.2 | Units being quoted for must be CE or SABS certified.

## 10 Category 9
CATEGORY 9
PATIENT TRANSFER
RT24-09-007ME | Transfer Sheet, 180 x 70 cm | Complies Yes/No | Comments (Technical Specification Deviation)
All items being tendered for must comply with the following technical specification, failure to comply with any of the conditions set out below may result in immediate disqualification. Please note that where the technical specification calls for certification, this certification must accompany the bid, failure to provide such certification will result in immediate disqualification. Bidders must give details of items being tendered for in the COMMENTS section where there are variation for further deliberation
1 | Technical Specification
1.1 | It must be lightweight sheet for easy transportability
1.2 | The unit must be comfortable with soft core
1.3 | The unit must not have hard metal rollers but must have similar transfer technique
1.4 | Size: 180 cm x 70 cm (± 20 cm)
1.5 | Radiotransluscent With handles
1.6 | The starter pack on initial delivery must have both reusable and disposable cover
1.7 | With roll mat and folding
1.8 | All accessories and consumables listed must be compatible with main unit on offer
2 | Standard Accessories and Consumables:
2.1 | Durable disposable plastic cover (Pk of 50) as per item RT24-09-008ME
2.2 | Reusable Plastic Cover as per item RT24-09-009ME
RT24-09-008ME | Transfer board cover, Disposable | Complies Yes/No | Comments (Technical Specification Deviation)
All items being tendered for must comply with the following technical specification, failure to comply with any of the conditions set out below may result in immediate disqualification. Please note that where the technical specification calls for certification, this certification must accompany the bid, failure to provide such certification will result in immediate disqualification. Bidders must give details of items being tendered for in the COMMENTS section where there are variation for further deliberation
1 | Technical Specification
1.1 | To fit the transfer board Size: 180 cm x 70 cm (± 20 mm)
1.2 | Must be made of plastic material
1.3 | The starter pack of 50 disposable covers must be supplied as a pack/set
RT24-09-009ME | Transfer sheet cover, Re-usable | Complies Yes/No | Comments (Technical Specification Deviation)
All items being tendered for must comply with the following technical specification, failure to comply with any of the conditions set out below may result in immediate disqualification. Please note that where the technical specification calls for certification, this certification must accompany the bid, failure to provide such certification will result in immediate disqualification. Bidders must give details of items being tendered for in the COMMENTS section where there are variation for further deliberation
1 | Technical Specification
1.1 | To fit the transfer board Size: 80 cm x 70 cm (± 20 mm)
1.2 | Must be made of durable plastic material
1.3 | Must be supplied in a quantity of 1 (each)
RT24-09-012ME | Bariatric Patient Hoist, Electric, Mobile | Complies Yes/No | Comments (Technical Specification Deviation)
All items being tendered for must comply with the following technical specification, failure to comply with any of the conditions set out below may result in immediate disqualification. Please note that where the technical specification calls for certification, this certification must accompany the bid, failure to provide such certification will result in immediate disqualification. Bidders must give details of items being tendered for in the COMMENTS section where there are variation for further deliberation
1 | Technical Specification
1.1 | The hoist frame must be made from durable steel material, please state the steel type
1.2 | The hanger bar assembly must be constructed in stable and solid manner that will prevent patient from swinging
1.3 | If the frame is made of Mild steel with epoxy/nylon powder-coated finish, it must comply with SANS 778 paragraph 5.2
1.4 | The unit must provide stable lifting of the stated Safe Working Load
1.5 | The unit must be equiped with Emergency Quick Release system
1.6 | Safe Working Load: 450kg
1.7 | Minimum Lifting Height: 600mm
1.8 | Maximum Lifting Height: 2000mm
1.9 | Handset operated with spiral cable
1.1 | The unit must have non-slip handles to assist with easy manuevering
1.11 | The unit must have battery back-up
1.12 | A fully charged battery back-up must last no less than 5 hours
1.13 | Both rear castors to be fitted with a locking device
1.14 | Castors must comply to the latest issue of SANS 621
1.15 | Where castors are fitted into steel tubular legs, the tube shall be of wall thickness not less than 2,0mm and castors shall be fixed to the tube by one of the following methods:
1.15.1 | Solid plug (long) complying with SANS 621 subsection 3.5.6 or
1.15.2 | Screwed into 35mm long sleeves welded into the tubular members and locked in an acceptable manner
1.15.3 | Rubber or plastic expanding sleeves for fitting castors are not acceptable
1.16 | Provision shall be made on the underside of the seat to hold the slop bucket and to allow its easy insertion and withdrawal
1.17 | The following accessories must meet the above specification, be compatible and supplied with the main unit
2 | Standard Accessories:
2.1 | Universal padded sling
2.2 | Sling with commode opening
3 | WARRANTY & MAINTENANCE
3.1 | A 2-year warranty inclusive maintenance and service (Year 1 to Year 2) should be quoted for, as per the SCC document
3.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries, Flow sensors and Oxygen sensor), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
3.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
3.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries, Flow sensors and Oxygen sensor), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
4 | Conformity Compliance (Please attach proof of compliance/Certificate)
4.1 | Manufacturer must comply to ISO 13485/ISO 9001 standards
4.2 | Units being quoted for must be CE or SABS certified.
4.3 | SANS 778 (Subsection 5.2)
4.4 | SANS 621
4.5 | IEC 60601-1-2
4.3 | SANS 778 (Subsection 5.2)
4.4 | SANS 621
4.5 | IEC 60601-1-2
RT24-09-013ME | Bariatric Patient Hoist, Electric, Mobile with integrated scale | Complies Yes/No | Comments (Technical Specification Deviation)
All items being tendered for must comply with the following technical specification, failure to comply with any of the conditions set out below may result in immediate disqualification. Please note that where the technical specification calls for certification, this certification must accompany the bid, failure to provide such certification will result in immediate disqualification. Bidders must give details of items being tendered for in the COMMENTS section where there are variation for further deliberation
1 | Technical Specification
1.1 | The hoist frame must be made from durable steel material, please state the steel type
1.2 | The hanger bar assembly must be constructed in stable and solid manner that will prevent patient from swinging
1.3 | If the frame is made of Mild steel with epoxy/nylon powder-coated finish, it must comply with SANS 778 paragraph 5.2
1.4 | The unit must provide stable lifting of the stated Safe Working Load
1.5 | The unit must be equiped with Emergency Quick Release system
1.6 | Safe Working Load: 450kg
1.7 | Minimum Lifting Height: 600mm
1.8 | Maximum Lifting Height: 2000mm
1.9 | Handset operated with spiral cable
1.1 | The unit must have non-slip handles to assist with easy manuevering
1.11 | The unit must have battery back-up
1.12 | A fully charged battery back-up must last no less than 5 hours
1.13 | The unit must have integrated electronic scale
1.13.2 | Weighting capacity: at least 450kg
1.13.2 | must have a Zeroing capability
1.14 | Both rear castors to be fitted with a locking device
1.15 | Castors must comply to the latest issue of SANS 621
1.16 | Where castors are fitted into steel tubular legs, the tube shall be of wall thickness not less than 2,0mm and castors shall be fixed to the tube by one of the following methods:
1.16.1 | Solid plug (long) complying with SANS 621 subsection 3.5.6 or
1.16.2 | Screwed into 35mm long sleeves welded into the tubular members and locked in an acceptable manner
1.16.3 | Rubber or plastic expanding sleeves for fitting castors are not acceptable
1.17 | Provision shall be made on the underside of the seat to hold the slop bucket and to allow its easy insertion and withdrawal
1.18 | The following accessories must meet the above specification, be compatible and supplied with the main unit
2 | Standard Accessories:
2.1 | Universal padded sling
2.2 | Sling with commode opening
3 | WARRANTY & MAINTENANCE
3.1 | A 2-year warranty inclusive maintenance and service (Year 1 to Year 2) should be quoted for, as per the SCC document
3.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries, Flow sensors and Oxygen sensor), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
3.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
3.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries, Flow sensors and Oxygen sensor), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
4 | Conformity Compliance (Please attach proof of compliance/Certificate)
4.1 | Manufacturer must comply to ISO 13485/ISO 9001 standards
4.2 | Units being quoted for must be CE or SABS certified.
4.3 | SANS 778 (Subsection 5.2)
4.4 | SANS 621
4.5 | IEC 60601-1-2

## 11 Category 11
CATEGORY 11
STOOLS
RT24-11-003ME | Foot Stool, Double, Rubber top | Complies Yes/No | Comments (Technical Specification Deviation)
All items being tendered for must comply with the following technical specification, failure to comply with any of the conditions set out below may result in immediate disqualification. Please note that where the technical specification calls for certification, this certification must accompany the bid, failure to provide such certification will result in immediate disqualification. Bidders must give details of items being tendered for in the COMMENTS section where there are variation for further deliberation
1 | Technical Specification
1.1 | The unit offered to be of high quality in terms of workmanship with ability to withstand specified weight capacity over time
1.2 | Steel plate top with non-slip rubber sheeting of 4mm thickness
1.3 | Stainless steel edging
1.4 | Dimensions (Tolerance +10mm):
1.4.1 | Overall height: 340mm
1.4.2 | First step height: 170mm
1.4.3 | Width: 500mm
1.4.4 | Overall length: 530mm
1.4.5 | First step length: 225mm
1.5 | Weight capacity: 200kg
1.6 | Mild steel frame and legs with epoxy/nylon powder-coated finish, to comply with SANS 778 paragraph 5.2
1.7 | Legs of the stool must be fitted with rubber tips or equivalent to prevent slipping
1.8 | The stool must be of heavy duty construction with steel band between steps for added strength
2 | Conformity Compliance (Please attach proof of compliance/Certificate)
2.1 | SANS 778 (subsection 5.2)