## RT274-2023 Pricing Schedule
LINE # | ITEM NUMBER | ITEM DESCRIPTION | OFFER # | UNIT OF MEASURE | DELIVERY PRICE IN ZAR INCL. 15% VAT | MAINTANANCE YEAR 1 (ONLY APPLICABLE TO ITEM RT274-01-036 TO 064) INCL. 15% VAT | MAINTANANCE YEAR 2 (ONLY APPLICABLE TO ITEM RT274-01-036 TO 064) INCL. 15% VAT | MAINTANANCE YEAR 3 (ONLY APPLICABLE TO ITEM RT274-01-036 TO 064) INCL. 15% VAT | MAINTANANCE YEAR 4 (ONLY APPLICABLE TO ITEM RT274-01-036 TO 064) INCL. 15% VAT | MAINTANANCE YEAR 5 (ONLY APPLICABLE TO ITEM RT274-01-036 TO 064) INCL. 15% VAT | BRAND NAME | NAME OF MANUFACTURE | COUNTRY OF ORIGIN | LOCAL COMPONENT % | IMPORT COMPONENT | FOREIGN CURRENCY | LEAD TIME
1 | RT274-01-001 | Pre-cut cleaning floss kit for OAE machines & tympanometer probes pack of (Large) 30 pieces of thick floss | Offer 1 | Pack of 30 pieces
2 | RT274-01-002 | Pre-cut cleaning floss kit for OAE machines & tympanometers probes pack of (Small and Large) 100 pieces of thin floss & 30 pieces of thick floss | Offer 1 | Pack of 100 pieces
3 | RT274-01-003 | Pre-cut cleaning floss kit for oae machines & tympanometers probes (small) pack of 100 pieces of thin floss | Offer 1 | Pack of 100 pieces
4 | RT274-01-004 | Probe tip kit for OAE screening and diagnostic equipment, containing a set of probe tips of different sizes for neonates. Supplier must specify with which makes/models/ the probe tips are compatible and probe material. Supplier must specify number and sizes of probe tips. | Offer 1 | Each
5 | RT274-01-005 | Probe tip kit for OAE screening and diagnostic equipment, containing a set of probe tips of different sizes for paediatrics. Supplier must specify with which makes/models/ the probe tips are compatible and probe material. Supplier must specify number and sizes of probe tips. | Offer 1 | Each
6 | RT274-01-006 | Probe tip kit for OAE screening and diagnostic equipment, containing a set of probe tips of different sizes for adults. Supplier must specify with which makes/models the probe tips are compatible and probe material. Supplier must specify number and sizes of probe tips. | Offer 1 | Each
7 | RT274-01-007 | Probe tip kit for middle ear analyzer equipment, containing a set of probe tips of different sizes for neonates. Supplier must specify with which makes/models the probe tips are compatible and probe material. Supplier must specify number and sizes of probe tips. | Offer 1 | Each
8 | RT274-01-008 | Probe tip kit for middle ear analyzer equipment, containing a set of probe tips of different sizes for paediatrics. Supplier must specify with which makes/models the probe tips are compatible and probe material. Supplier must specify number and sizes of probe tips. | Offer 1 | Each
9 | RT274-01-009 | Probe tip kit for middle ear analyzer equipment, containing a set of probe tips of different sizes for adults. Supplier must specify with which makes/models the probe tips are compatible and probe material. Supplier must specify number and sizes of probe tips. | Offer 1 | Each
10 | RT274-01-010 | Electrode cover for AABR | Offer 1 | Each
11 | RT274-01-011 | Eartip, disposable, polyurethane foam, for use with tubal insert earphones & all evoked potential systems. Bidder to specify dimension and quantity per pack. Supplier to stipulate which equipment items are compatible with. | Offer 1 | Each
12 | RT274-01-012 | EEG paste, conductive, 400g water-miscible adhesive paste intended to reduce skin impedance and used with non-disposable neurodiagnostic electrodes, used to stick electrodes directly to the skin for a secure connection | Offer 1 | Each
13 | RT274-01-013 | EEG abrasive skin prepping gel, intended for use when a reduction in skin impedance would enhance a test result carries a ce mark. Supplier must specify packaging size and cost per ml/gram | Offer 1 | Each
14 | RT274-01-014 | Hearing aid programmer (USB), not limited to specific brand of hearing aid | Offer 1 | Each
15 | RT274-01-015 | Wireless hearing aid programmer, not limited to specific brand of hearing aid | Offer 1 | Each
16 | RT274-01-016 | Modular software for programming hearing aids, compatible with USB and wireless programmer, complete with installation guide, and assistance | Offer 1 | Each
17 | RT274-01-017 | Disposable electrodes for electrophysiologic testing diagnostic.Bider to specify number in pack and price per unit. Supplier to specify equipment compatibility | Offer 1 | Each
18 | RT274-01-018 | Disposable electrodes for for electrophysiologic testing screening ABR. Bidder to specify number in pack and price per unit. Supplier to specify equipment compatibility | Offer 1 | Each
19 | RT274-01-019 | Electrodes, disposable EP/ENG snap, wet gel, Supplier to specify number in pack and price per unit. | Offer 1 | Each
20 | RT274-01-020 | Electrodes, EEG cup, reusable, adult, electrodes & lead wires are durable electrode cup has cosistent thickness & surface area easy to hold hub cup & hub interface design provides stability during recordings lead wires are tangle free, flexible, kink-resistant & strong. Dimensions 1m lead wire, 10m diameter cup silver of 87 durometer. Packaged as single unit | Offer 1 | Each
21 | RT274-01-021 | Electrodes, EEG cup, reusable paediatric, electrodes & lead wires are durable electrode cup has cosistent thickness & surface area easy to hold hub cup & hub interface design provides stability during recordings lead wires are tangle free, flexible, kink-resistant & strong. Dimesnsions: 1m lead wire, 6mm diametere cup silver of 87 durometer. Packaged as a single unit. | Offer 1 | Each
22 | RT274-01-022 | One-channel portable audiometer capable of performing unmasked puretone air and bone conduction testing. Must come with earphones, bone conductor, response button and carrycase. Must test 125-8 000Hz from 0-100dB in increments of 5db. Must come with instruction manual. Bidder must state power source and additional features and options. | Offer 1 | Each
23 | RT274-01-023 | One-channel portable audiometer capable of performing unmasked puretone air and bone conduction testing. Must come with earphones, bone conductor, response button and carrycase. Must test 125-16 000Hz from 0-100dB in increments of 5db. Must come with instruction manual. Bidder must state power source and additional features and options. | Offer 1 | Each
24 | RT274-01-024 | Hand-held tympanometry and acoustic reflex screener. Must have standard 226hz probe tone and must test reflexes at 250, 1000, 2000 and 4000Hz. Must include an internal printer, probe assembly, two rolls of paper, memory for at least 12 tests, carry case, instruction manual and box of eartips of at least six different sizes, with three tips of each size. Multiple award, supplier must state power source and additional features and options. (MIOS/SITA certified ) | Offer 1 | Each
25 | RT274-01-025 | A hand-held portable unit capable of conducting screening DPOAEs and automated ABRs. Must consist of the hand-held screening unit, probe, system cables (USB to connect system to computer; ABR cable; power cable, charger unit and power cord, carry case, instruction manual, data management software CD and eartip accessory box, including probe replacement filters and 50 eartips of various sizes, including for babies. Must store at least 50 tests and have a built-in normative databases to reference | Offer 1 | Each
26 | RT274-01-026 | Pair of 90dB wall-mountable speakers with cable to connect to audiometer. Multiple award, must be compatible with awarded audiometers | Offer 1 | Each
27 | RT274-01-027 | Adjustable earphone headset that complies with the requirements of ANSI S3.6 – 2010 and ISO389-1 section 4.3. must have less than 1% harmonic distortion at 120dB at 1000Hz. Must be compatible with awarded audiometers. Multiple award | Offer 1 | Each
28 | RT274-01-028 | Paired insert earphones, compatible with awarded audiometer, ASSR, AEP or REM. Must consist of earphones, two replacement cables, 50 disposable eartips of various sizes and 213cm (7’) cord with dual (left and right) 1/4” mono colour-coded phono plugs. Velcro clip, lanyard and neck loop must be provided for earphones for use with audiometer and REM. Multiple award. | Offer 1 | Each
29 | RT274-01-029 | Flexible metal headband with bone conductor and cable compatible with awarded audiometer or AEP on one end. Bone conductor must meet ANSI/ASA S3.6-2010 standard. | Offer 1 | Each
30 | RT274-01-030 | Compact CD player for playing recorded speech materials. Must include CD player, instruction manual, power cable and cable to link to audiometer. Must be compatible with awarded audiometer. Multiple award. | Offer 1 | Each
31 | RT274-01-031 | Roll of thermal paper for printing results of tympanometry testing. Multiple award, must be compatible with awarded tympanometers, tympanometer-combination machines, OAE machines and real ear measurements. Supplier to specify size of roll and equipment compatibility. | Offer 1 | Each
32 | RT274-01-032 | Printer compatible with awarded audiometer, to print patient test record. Must include spare printer cartridge, power cable and connection cable to audiometer. MIOS/SITA certified. | Offer 1 | Each
33 | RT274-01-033 | Printer compatible with middle-ear analyser, to print results. Must include three rolls of thermal paper (if the printer uses this), spare printer cartridge (if it doesn’t use thermal paper), connection cable and power cable. (MIOS/SITA certified ) | Offer 1 | Each
34 | RT274-01-034 | Colour printer compatible with OAE, AEP and ASSR, to print test results. Must include spare printer cartridge, connection cable and power cable. (MIOS/SITA certified ) | Offer 1 | Each
35 | RT274-01-035 | Colour printer compatible with awarded REM and hearing aid analyser, to print test results. Must include spare printer cartridge, connection cable and power cable. (MIOS/SITA certified ) | Offer 1 | Each
36 | RT274-01-036 | Screening base unit: A base unit to which screening modules and their necessary hardware can be added or connected; must have capability for multiple modules added, must allow for subsequent addition of additional modules. Must be portable and must come with protective carrycase. Item must be compatible with items RT274-01-36 to 49. Bidder must provide pricing for the maintenance plan as provided for on the pricing schedule. | Offer 1 | Each
37 | RT274-01-037 | Screening audiometer base unit: One-channel audiometer base unit, capable of performing unmasked puretone air and bone conduction testing, must come with earphones, bone conductor and response button, must test 125-8 000Hz from 0-100dB in increments of 5db. Must be portable and must come with protective carrycase. Item must be compatible with items RT274-01-36 to 49. Bidder must provide pricing for the maintenance plan as provided for on the pricing schedule. | Offer 1 | Each
38 | RT274-01-038 | Screening high-freequency audiometer base unit: One-channel audiometer audiometer base unit, capable of performing unmasked puretone air and bone conduction testing, must come with earphones, bone conductor and response buitton, must test 125-16 000Hz from 0-100dB in increments of 5db. Must be portable and must come with protective carrycase. Item must be compatible with items RT274-01-36 to 49. Bidder must provide pricing for the maintenance plan as provided for on the pricing schedule. | Offer 1 | Each
39 | RT274-01-039 | Screening tympanometer with reflexes base unit: Hand-held tympanometry and acoustic reflex screener, must have standard 226hz probe tone and must test reflexes at 250, 1 000, 2 000 and 4 000Hz, must include an internal printer, probe assembly, two rolls of paper, memory for at least 12 tests, must include a box of eartips of at least six different sizes, with three tips of each size. (MIOS/SITA certified). Must be portable and must come with protective carrycase. Item must be compatible with items RT274-01-36 to 49. Bidder must provide pricing for the maintenance plan as provided for on the pricing schedule. | Offer 1 | Each
40 | RT274-01-040 | Screening high-frequency tympanometer with reflexes base unit: Hand-held tympanometry and acoustic reflex screene, must have standard 226Hz, 678Hz and 1 000Hz probe tones and must test reflexes at 250, 1 000, 2 000 and 4 000Hz, must include an internal printer, probe assembly, two rolls of paper, memory for at least 12 tests, must include a box of eartips of at least six different sizes, with three tips of each size. (MIOS/SITA certified). Must be portable and must come with protective carrycase. Item must be compatible with items RT274-01-36 to 49. Bidder must provide pricing for the maintenance plan as provided for on the pricing schedule. | Offer 1 | Each
41 | RT274-01-041 | Screening DPOAE and TDOAE base unit: A handheld unit capable of measuring DPOAEs and TDOAEs, must store at least 50 tests and have a built-in normative database to reference pass/refer results, must have software for download to PC based database, must include the hand-held unit, OAE probe, probe calibration test cavity, database software and micro USB charging / transfer cable, must include eartip kit and replacement probe tubes, must test the range of 2000-6000Hz at 65/55dB (DPOAE) and 1500-4000Hz at 80-85Hz per SPL. Must be portable and must come with protective carrycase. Item must be compatible with items RT274-01-36 to 49. Bidder must provide pricing for the maintenance plan as provided for on the pricing schedule. | Offer 1 | Each
42 | RT274-01-042 | Screening high-frequency DPOAE base unit: A handheld unit capable of measuring DPOAEs, must store at least 50 tests and have a built-in normative database to reference pass/refer results, must have software for download to PC based database, must include the hand-held unit, OAE probe, probe calibration test cavity, database software and micro USB charging / transfer cable, must include eartip kit and replacement probe tubes, must test the range of 2000-16 000Hz at 65/55dB. Must be portable and must come with protective carrycase. Item must be compatible with items RT274-01-36 to 49. Bidder must provide pricing for the maintenance plan as provided for on the pricing schedule. | Offer 1 | Each
43 | RT274-01-043 | Screening audiometer module: One-channel audiometer module, capable of performing unmasked puretone air and bone conduction testing, must come with earphones, bone conductor and response button, must test 125-8 000Hz from 0-100dB in increments of 5db. Must be portable and must come with protective carrycase. Item must be compatible with items RT274-01-36 to 49. Bidder must provide pricing for the maintenance plan as provided for on the pricing schedule. | Offer 1 | Each
44 | RT274-01-044 | Screening high-frequency audiometer module: One-channel audiometer module, capable of performing unmasked puretone air and bone conduction testing, Must come with earphones, bone conductor and respons button, must test 125-16 000Hz from 0-100dB in increments of 5db. Must be portable and must come with protective carrycase. Item must be compatible with items RT274-01-36 to 49. Bidder must provide pricing for the maintenance plan as provided for on the pricing schedule. | Offer 1 | Each
45 | RT274-01-045 | Screening tympanometer with reflexes module: Hand-held tympanometry and acoustic reflex screening module, must have standard 226hz probe tone and must test reflexes at 250, 1 000, 2 000 and 4 000Hz, must include an internal printer, probe assembly, two rolls of paper, memory for at least 12 tests, must include a box of eartips of at least six different sizes, with three tips of each size. (MIOS/SITA certified ). Must be portable and must come with protective carrycase. Item must be compatible with items RT274-01-36 to 49. Bidder must provide pricing for the maintenance plan as provided for on the pricing schedule. | Offer 1 | Each
46 | RT274-01-046 | Screening high-frequency tympanometer with reflexes module: Hand-held tympanometry and acoustic reflex screening module, must have standard 226Hz, 678Hz and 1 000Hz probe tones and must test reflexes at 250, 1 000, 2 000 and 4 000Hz, must include an internal printer, probe assembly, two rolls of paper, memory for at least 12 tests, must include a box of eartips of at least six different sizes, with three tips of each size. (MIOS/SITA certified ). Must be portable and must come with protective carrycase. Item must be compatible with items RT274-01-36 to 49. Bidder must provide pricing for the maintenance plan as provided for on the pricing schedule. | Offer 1 | Each
47 | RT274-01-047 | Screening DPOAE and TDOAE module: A module capable of measuring DPOAEs and TDOAEs, must store at least 50 tests and have a built-in normative database to reference pass/refer results, must have software for download to PC based databas, must include the hand-held unit, OAE probe, probe calibration test cavity, database software and micro USB charging / transfer cable, must include eartip kit and replacement probe tubes, must test the range of 2000-6000Hz at 65/55dB (DPOAE) and 1500-4000Hz at 80-85Hz per SPL. Must be portable and must come with protective carrycase. Item must be compatible with items RT274-01-36 to 49. Bidder must provide pricing for the maintenance plan as provided for on the pricing schedule. | Offer 1 | Each
48 | RT274-01-048 | Screening high-frequency DPOAE module: A module capable of measuring DPOAEs, must store at least 50 tests and have a built-in normative database to reference pass/refer results, must have software for download to PC based database, must include the hand-held unit, OAE probe, probe calibration test cavity, database software and micro USB charging / transfer cable, must include eartip kit and replacement probe tubes, must test the range of 2000-16 000Hz at 65/55dB. Must be portable and must come with protective carrycase. Item must be compatible with items RT274-01-36 to 49. Bidder must provide pricing for the maintenance plan as provided for on the pricing schedule. | Offer 1 | Each
49 | RT274-01-049 | Screening AABR module: A module capable of conducting screening automated ABRs, must consist of the hand-held screening unit, probe, system cables (USB to connect system to computer; ABR cable; power cable, data management software CD and eartip accessory box, including probe replacement filters and 50 eartips of various sizes, including for babies, must store at least 50 tests and have a built-in normative database to reference pass/refer results. Must be portable and must come with protective carrycase. Item must be compatible with items RT274-01-36 to 49. Bidder must provide pricing for the maintenance plan as provided for on the pricing schedule. | Offer 1 | Each
50 | RT274-01-050 | Diagnostic audiometer base unit: Two-channel audiometer capable of performing masked puretone and warbletone air and bone conduction and speech testing; masking sounds must include narrowband, white and speech noise, must come with earphones, paired insert phones, bone conductor, response button, talk-forward, built-in or gooseneck microphone and built-in pre-recorded speech material in South African accents and languages, and USB ports to connect to an external CD player for recorded tests, must have the following special tests: SISI, ABLB, Stenger, Bekesy, Weber, tone-in-noise, speech-in-noise, must test 125-8 000Hz from -10-120dB in increments of 5db. Base unit: A base unit to which screening modules and their necessary hardware can be added or connected, must have capability for multiple modules to added, must allow for subsequent addition of additional modules. Bidder must provide pricing for the maintenance plan as provided for on the pricing schedule. | Offer 1 | Each
51 | RT274-01-051 | Diagnostic high-frequency audiometer base unit: Two-channel audiometer capable of performing masked puretone and warbletone air and bone conduction and speech testing; masking sounds must include narrowband, white and speech noise, must come with earphones, paired insert phones, bone conductor, response button, talk-forward, built-in or gooseneck microphone and built-in pre-recorded speech material in South African accents and languages, and USB ports to connect to an external CD player for recorded tests, must have the following special tests: SISI, ABLB, Stenger, Bekesy, Weber, tone-in-noise, speech-in-noise, must test 125-20 000Hz from -10-120dB (100dB for frequencies above 12 000Hz) in increments of 5db. | Offer 1 | Each
52 | RT274-01-052 | Diagnostic middle-ear analyser base unit: Middle-ear analyser capable of performing the following tests: Tympanometry, ipsi- and contra-lateral reflexes, reflex decay and Eustachian tube function (intact and perforated tympanic membrane), must have 226Hz probe tone and test reflexes at 250, 500, 1000, 2000 and 4000Hz (puretone), 400-4000Hz BBN and clicks in the range of 35-120dB, must test tympanometry across range of +200 - -400daPa with accuracy of ±10daPa, must perform ipsi- and contra-lateral reflex testing, must come with the following: probe assembly (including contra-lateral insert earphone), probe mount kit/accessory (or shoulder strap), eartips (minimum six different sizes, with three of each tip size), calibration test cavity, cleaning kit and USB PC interface. | Offer 1 | Each
53 | RT274-01-053 | Diagnostic high frequency middle-ear analyser base unit: Middle-ear analyser capable of performing the following tests: Tympanometry, ipsi- and contra-lateral reflexes, reflex decay and Eustachian tube function (intact and perforated tympanic membrane; must have 226Hz, 678Hz and 1000Hz probe tone and test reflexes at 250, 500, 1000, 2000 and 4000Hz (puretone), 400-4000Hz BBN and clicks in the range of 35-120dB, must test tympanometry across range of +200 - -400daPa with accuracy of ±10daPa, must perform ipsi- and contra-lateral reflex testing, must come with the following: probe assembly (including contra-lateral insert earphone), probe mount kit/accessory (or shoulder strap), eartips (minimum six different sizes, with three of each tip size), calibration test cavity, cleaning kit and USB PC interface. Base unit: A base unit to which screening modules and their necessary hardware can be added or connected, must have capability for multiple modules to added, must allow for subsequent addition of additional modules. Bidder must provide pricing for the maintenance plan as provided for on the pricing schedule. | Offer 1 | Each
54 | RT274-01-054 | Diagnostic DPOAE and TDOAE base unit: A diagnostic unit capable of measuring DPOAEs and TDOAE’s, with at least six configurable protocols testing from 1.5-12000Hz, must store at least 50 tests and display normative data to assist with interpretation of results, must have software for download to PC based database, must include the unit, cradle, if relevant, OAE probe, probe calibration test cavity, database software, micro USB charging / transfer cable, carry case, eartip kit and replacement probe tubes and calibration certificate. Base unit: A base unit to which screening modules and their necessary hardware can be added or connected, must have capability for multiple modules to added, must allow for subsequent addition of additional modules. Bidder must provide pricing for the maintenance plan as provided for on the pricing schedule. | Offer 1 | Each
55 | RT274-01-055 | Diagnostic high frequency DPOAE and TDOAE base unit: A diagnostic unit capable of measuring DPOAEs and TDOAE’s, with at least six configurable protocols testing from 1.5-16 000Hz, must store at least 50 tests and display normative data to assist with interpretation of results, must have software for download to PC based databas, must include the unit, cradle, if relevant, OAE probe, probe calibration test cavity, database software, micro USB charging / transfer cable, carry case, eartip kit and replacement probe tubesl and calibration certificate. Base unit: A base unit to which screening modules and their necessary hardware can be added or connected, must have capability for multiple modules to added, must allow for subsequent addition of additional modules. Bidder must provide pricing for the maintenance plan as provided for on the pricing schedule. | Offer 1 | Each
56 | RT274-01-056 | Diagnostic audiometer module:Two-channel audiometer module capable of performing masked puretone and warbletone air and bone conduction and speech testing; masking sounds must include narrowband, white and speech noise, must come with earphones, paired insert phones, bone conductor, response button, talk-forward, built-in or gooseneck microphone and built-in pre-recorded speech material in South African accents and languages, and USB ports to connect to an external CD player for recorded tests, must have the following special tests: SISI, ABLB, Stenger, Bekesy, Weber, tone-in-noise, speech-in-noise, must test 125-8 000Hz from -10-120dB in increments of 5db. Base unit: A base unit to which screening modules and their necessary hardware can be added or connected, must have capability for multiple modules to added, must allow for subsequent addition of additional modules. Bidder must provide pricing for the maintenance plan as provided for on the pricing schedule. | Offer 1 | Each
57 | RT274-01-057 | Diagnostic high-frequency audiometer module: Two-channel audiometer module capable of performing masked puretone and warbletone air and bone conduction and speech testing;, masking sounds must include narrowband, white and speech noise, must come with earphones, paired insert phones, bone conductor, response button, talk-forward, built-in or gooseneck microphone and built-in pre-recorded speech material in South African accents and languages, and USB ports to connect to an external CD player for recorded tests, must have the following special tests: SISI, ABLB, Stenger, Bekesy, Weber, tone-in-noise, speech-in-noise, must test 125-20 000Hz from -10-120dB (100dB for frequencies above 12 000Hz) in increments of 5db. Base unit: A base unit to which screening modules and their necessary hardware can be added or connected, must have capability for multiple modules to added, must allow for subsequent addition of additional modules. Bidder must provide pricing for the maintenance plan as provided for on the pricing schedule. | Offer 1 | Each
58 | RT274-01-058 | Diagnostic middle-ear analyser module: Middle-ear analyser capable of performing the following tests: Tympanometry, ipsi- and contra-lateral reflexes, reflex decay and Eustachian tube function (intact and perforated tympanic membrane), must have 226Hz probe tone and test reflexes at 250, 500, 1000, 2000 and 4000Hz (puretone), 400-4000Hz BBN and clicks in the range of 35-120dB, must test tympanometry across range of +200 - -400daPa with accuracy of ±10daPa, must perform ipsi- and contra-lateral reflex testing, must come with the following: probe assembly (including contra-lateral insert earphone), probe mount kit/accessory (or shoulder strap), eartips (minimum six different sizes, with three of each tip size), calibration test cavity, cleaning kit and USB PC interface. Base unit: A base unit to which screening modules and their necessary hardware can be added or connected, must have capability for multiple modules to added, must allow for subsequent addition of additional modules. Bidder must provide pricing for the maintenance plan as provided for on the pricing schedule. | Offer 1 | Each
59 | RT274-01-059 | Diagnostic high frequency middle-ear analyser module: Middle-ear analyser module capable of performing the following tests: Tympanometry, ipsi- and contra-lateral reflexes, reflex decay and Eustachian tube function (intact and perforated tympanic membrane), must have 226Hz, 678Hz and 1000Hz probe tone and test reflexes at 250, 500, 1000, 2000 and 4000Hz (puretone), 400-4000Hz BBN and clicks in the range of 35-120dB, must test tympanometry across range of +200 - -400daPa with accuracy of ±10daPa, must perform ipsi- and contra-lateral reflex testing, must come with the following: probe assembly (including contra-lateral insert earphone), probe mount kit/accessory (or shoulder strap), eartips (minimum six different sizes, with three of each tip size), calibration test cavity, cleaning kit and USB PC interface. Base unit: A base unit to which screening modules and their necessary hardware can be added or connected, must have capability for multiple modules to added, must allow for subsequent addition of additional modules. Bidder must provide pricing for the maintenance plan as provided for on the pricing schedule. | Offer 1 | Each
60 | RT274-01-060 | Diagnostic DPOAE and TDOAE module: A diagnostic module capable of measuring DPOAEs and TDOAE’s, with at least six configurable protocols testing from 1.5-12000Hz, must store at least 50 tests and display normative data to assist with interpretation of results, must have software for download to PC based database, must include the unit, cradle, if relevant, OAE probe, probe calibration test cavity, database software, micro USB charging / transfer cable, carry case, eartip kit and replacement probe tubes and calibration certificate. Base unit: A base unit to which screening modules and their necessary hardware can be added or connected, must have capability for multiple modules to added, must allow for subsequent addition of additional modules. Bidder must provide pricing for the maintenance plan as provided for on the pricing schedule. | Offer 1 | Each
61 | RT274-01-061 | Diagnostic high frequency DPOAE and TDOAE module: A diagnostic unit capable of measuring DPOAEs and TDOAE’s, with at least six configurable protocols testing from 1.5-16 000Hz, must store at least 50 tests and display normative data to assist with interpretation of results, must have software for download to PC based database, must include the unit, cradle, if relevant, OAE probe, probe calibration test cavity, database software, micro USB charging / transfer cable, carry case, eartip kit and replacement probe tubes and calibration certificate. Base unit: A base unit to which screening modules and their necessary hardware can be added or connected, must have capability for multiple modules to added, must allow for subsequent addition of additional modules. Bidder must provide pricing for the maintenance plan as provided for on the pricing schedule. | Offer 1 | Each
62 | RT274-01-062 | AEP module: Must have diagnostic Auditory Evoked Potential (ABR, ECochG, AMLR and LLR), must present click, tone pip and tone burst stimuli with possibility of contralateral masking, must have continuous live display of stimuli, must present all test results of the patient on one chart, must consist of AEP unit, left, right and bone transducers, 50 soft tips of various sizes and 100 foam tips of various sizes to fit newborn to adult patients, three electrode cables, 100 disposable self-adhesive electrodes, laptop with separate output jacks for each transducer and one for the freefield speaker, freefield speaker, two-channel capability for ipsi/contralateral ABR recordings, USB cable to connect unit to laptop, data management software, power cable. Base unit: A base unit to which screening modules and their necessary hardware can be added or connected, must have capability for multiple modules to added, must allow for subsequent addition of additional modules. Bidder must provide pricing for the maintenance plan as provided for on the pricing schedule. | Offer 1 | Each
63 | RT274-01-063 | ASSR module: A diagnostic module capable of measuring ASSRs, using AM/FM modulated tonal stimuli for the 250 - 8,000 Hz frequency range, must have on-screen impedance measurement, must have override facility for audiologist to adapt the test protocol. Base unit: A base unit to which screening modules and their necessary hardware can be added or connected, must have capability for multiple modules to added, must allow for subsequent addition of additional modules. Bidder must provide pricing for the maintenance plan as provided for on the pricing schedule. | Offer 1 | Each
64 | RT274-01-064 | REM module: Module to assess the functioning of the ear without hearing amplification, specifically for the purposes of hearing aid fitting, must comply with IEC 61669:2015/ ISO 12124 / ANSI S3.46-2013, must use three options of pure tone, warble tone, iCRA, Chirp, pink noise and white noise, and Real Speech as stimuli, must be able to perform the following tests: REUR / REUG, REAR / REAG, REIG, REOR / REOG, MPO, Advanced, Live speech, RECD, must include paired insert phones, in-situ headset, probe kit (tubes, O-rings, probe tube guides), monitor headset with boom microphone, Headphones for counselling, software CD, medical grade power supply, USB connection cable. Base unit: A base unit to which screening modules and their necessary hardware can be added or connected, must have capability for multiple modules to added, must allow for subsequent addition of additional modules. Bidder must provide pricing for the maintenance plan as provided for on the pricing schedule. | Offer 1 | Each
65 | RT274-02-001 | Otolight, complete with compatible batteries, hard casing and minimum of 2 tapered tips (hard transparent plastic tips), for earmould impressions | Offer 1 | Each
66 | RT274-02-002 | Replacement otolight lightbulbs. Must be compatible with otolights supplied on RT274 contracts since 2012 | Offer 1 | Each
67 | RT274-02-003 | LED replacement otolight lightbulbs. Must be compatible with otolights supplied on RT274 contracts since 2012 | Offer 1 | Each
68 | RT274-02-004 | Replacement otolight tips, pack of two | Offer 1 | Each
69 | RT274-02-005 | Retubing kit in a zip durable pouch complete with stainless steel tube expander, scalpel, tweezer, hearing aid screw driver, clear strong instant glue, air blower for earmould, needle nose plier, a tube threader | Offer 1 | Each
70 | RT274-02-006 | Small electric hand drill complete with drill bits suitable for drilling earmoulds, pack of 5 different sizes 0.5, 1, 1.5, 2 and 2.5mm drill bits, buffer ball, grindstone and sand paper | Offer 1 | Each
71 | RT274-02-007 | Set of drillbits and accessories for a small electric earmould drill consisting of 5 different sizes 0.5, 1, 1.5, 2 and 2.5mm drill bits, buffer ball, grindstone and sand paper, for ear mould. To be compatible with awarded drill | Offer 1 | Each
72 | RT274-02-008 | Stetoclip, with 2 earpieces and silicone listening tip for attachment to hearing aid | Offer 1 | Each
73 | RT274-02-009 | Speculae, cone shaped, reusable fittings for otoscope that is made of hard durable plastic in various lengths and diameters. Bidder to specify quantity of sizes per pack and price per unit. Company to specify compatibility with brand of device. | Offer 1 | Each
74 | RT274-02-010 | Otoscope, electric, wall-mounted, complete with 12 specula in various sizes. Item RT274-02-014 and RT274-02-018 must b awarded as a series | Offer 1 | Each
75 | RT274-02-011 | Speculum reusable tips, made of hard, durable, plastic in various sizes for wall-mounted otoscope. Bidder to specify compatibility with brand and quantity of sizes per pack and price per unit. | Offer 1 | Each
76 | RT274-02-012 | Covers, audiometer earphone, each, aucostically transparent. Bidder to specify quantity per pack and price per unit. | Offer 1 | Each
77 | RT274-02-013 | Replacement lightbulb for wall-mounted electric otoscope | Offer 1 | Each
78 | RT274-02-014 | Blower air, egg shaped, for earmould | Offer 1 | Each
79 | RT274-03-001 | Kidney bowl, stainless steel specialized ear irrigation bowl, 25cm in length | Offer 1 | Each
80 | RT274-03-002 | Kidney bowl, plastic specialized ear irrigation bowl, 25cm in length | Offer 1 | Each
81 | RT274-03-003 | Cut-out cup, ear syringing, each, stainless steel, receiver 500ml | Offer 1 | Each
82 | RT274-03-004 | Ear irrigation sprayer: 1 spray wash bottle (470ml) with 1 irrigation basin and Bidder to specify quantity of irrigation tips per pack and unit price | Offer 1 | Each
83 | RT274-03-005 | Ear irrigation tips, screw onto special ear irrigation syringes, flared tips and produce divergent streams of water, supplier to specify quantity per pack and unit price | Offer 1 | Each
84 | RT274-03-006 | Syringe, ear irrigation, metal | Offer 1 | Each
85 | RT274-03-007 | Syringe, ear irrigation, plastic | Offer 1 | Each
86 | RT274-03-008 | Curette, cerumen, paediatric, stainless steel | Offer 1 | Each
87 | RT274-03-009 | Curette, cerumen, adult, stainless steel | Offer 1 | Each
88 | RT274-03-010 | Hook, cerumen, for extraction of hard wax and other obstacles from the ear canal | Offer 1 | Each
89 | RT274-03-011 | Forceps, cerumen, lighted, supplier to specify quantity per pack and unit price | Offer 1 | Each
90 | RT274-03-012 | Probe, each, for hearing aid, wax removal and cotton application | Offer 1 | Each
91 | RT274-03-013 | Forceps, crocodile, stainless steel | Offer 1 | Each
92 | RT274-03-014 | Ear loop, billeau, small, stainless steel | Offer 1 | Each
93 | RT274-03-015 | Lighted cerumen curettes, with magnification, 1 light source and 1 magnifier. Bidder to specify quantity of included curette tips per pack and unit price. | Offer 1 | Each
94 | RT274-03-016 | Lightsource, replacement for lighted cerumen management tips and forceps | Offer 1 | Each
95 | RT274-03-017 | Day hook, stainless steel for extraction of curemen and other obstructions from the ear canal | Offer 1 | Each
96 | RT274-03-018 | Curettes, plastic cerumen, adult, supplier to specify quantity per pack and unit price | Offer 1 | Each
97 | RT274-03-019 | Lightsource, replacement for lighted cerumen management tips and forceps | Offer 1 | Each
98 | RT274-03-020 | Day hook, stainless steel for extraction of curemen and other obstructions from the ear canal | Offer 1 | Each
99 | RT274-03-021 | Headlight, focused, battery operated, cordless, portable, each, for conducting various audiological procedures | Offer 1 | Each
100 | RT274-04-001 | A digital bone conduction hearing aid spectacles suitable for patients with conductive and mixed hearing losses or single sided deafness complete with plano lenses (may be removable to accomodate prescription lenses if applicable). A bone conduction hearing temple/sound processor attached to the arm of the spectacle that sits against the bone behind the ear may be detacheable and must have the following minimum specifications: a battery casing, 8 channels, automatic feedback cancellation, automatic noise reduction, omindirectional/adaptive directionality, MPO programmable, high and low cut filters, volume control, everyday/activity case, a pack of 6 appropriately sized batteries. Must have various programming options (standalone or interfaced software). May have telephone compatibility, wireless directional microphone and direct audio input . Frames/spectacle to be available in various sizes and colours where applicable. a 2 year warranty on the sound processor must be guaranteed by the supplier. | Offer 1 | Each
101 | RT274-04-002 | Detacheable, passive, digital, bone conduction sound processor with variety of wearing options/mounting options,such as soft headband, cap or adhesive surgical tape (Kit). Sound processor to be housed in a single casing that is water and sweat resistant with indicator lights and tamperproof locks. Should be suitable for single sided deafness, conductive and mixed hearing losses with bone conduction thresholds of up to 45 dB (mild to moderate option) and the power option for bone conduction thresholds of up to 65dB. may have telephone compatibility, direct audio input, ability to do insitu audiometry/threshold, datalogging and FM sytem compatibility where applicable. The sound processor to be available in a range of colours. Various mounting/wearing options (cap, softband, adhesive tape) to be available in various sizes and colours where applicable. Should Allow for programming via standalone or interfaced software. Must have a minimum of 8 channels. Must have adaptive feedback cancellation and adaptive noise reduction functions, fixed and adaptive directionality. should come with a pack of appropriately sized disposable batteries/ appropriate rechargeable battery where indicated. A 2 year warranty on the sound must be guaranteed by the supplier. Processor must have reliability reporting and manufacturer certification (as per submitted requirements). | Offer 1 | Each
102 | RT274-04-003 | Hearing aid, bone conductor, complete with adjustable head band, bone vibrator, cord, pack of 6 batteries and hearing aid, behind the ear, digital, hearing aid, for mild - moderate loss (30+db). Items RT274-04-003-RT274-04-004 must be compatible and items RT274-04-006-RT274-04-008 must be awarded as a series | Offer 1 | Each
103 | RT274-04-004 | Hearing aid, bone conductor, complete with wide head band, bone vibrator, cord, pack of 6 batteries and hearing aid, behind the ear, digital, hearing aid, for mild to moderate loss (30+db) | Offer 1 | Each
104 | RT274-04-005 | Detacheable, passive digital bone conduction sound processor (only) suitable for single sided deafness, conductive and mixed hearing loss, with bone conduction thresholds of 45 dB and power option for thresholds of up to 65dB. should be compatible with various wearing/mounting options (cap, softband, adhesive surgical tape) awarded. Ability to do insitu audiometry/ threshold, datalogging, FM system compatibility where applicable. A pack of 6 approriately sized disposabled batteries/appropriate rechargeable battery where indicated, sound processor cover, test rod, activity/everyday case. A 2 year warranty on the sound must be guaranteed by the supplier. Processor must have reliability reporting and manufacturer certification (as per submitted requirements). | Offer 1 | Each
105 | RT274-04-006 | Cord, bone conductor, each, behind the ears, compatible with all bone conductor hearing aids awarded | Offer 1 | Each
106 | RT274-04-007 | Headband, adjustable, for bone conductor hearing aid. Item RT274-04-007 must be compatible with items RT274-04-002-RT274-04-006 | Offer 1 | Each
107 | RT274-04-008 | Headband, wide band, for bone conductor hearing aid. Item RT274-04-008 must be compatible with items RT274-04-002-RT274-04-006 | Offer 1 | Each
108 | RT274-04-009 | Replacement bone conduction spectacle hearing aid temple in various colours and sizes. Suitable for conductive and mixed hearing losses as well as single sided deafness. Must be compatible with all bone conduction hearing aids spectacles awarded (RT274-04-001). | Offer 1 | Each
109 | RT274-04-010 | Replacement extension eartips for bone conduction hearing aid spectacle in various colours. Must be compatible with all bone conduction spectacles hearing aids awarded (RT 274-04-001) | Offer 1 | Each
110 | RT274-04-011 | Replacement frames indifferent colours and sizes for bone conduction hearing aid spectacles. Must be compatible with all bone conduction hearing aid spectacles awarded (RT274-04-001) | Offer 1 | Each
111 | RT274-04-012 | Replacement plano lenses for digital bone conduction hearing aid spectacles, must be compatible with RT274-04-001 | Offer 1 | Each
112 | RT274-04-013 | mountable Soft headband various colours and sizes where applicable, for passive digital bone conduction sound processor, Must be compatible with RT274-04-005 | Offer 1 | Each
113 | RT274-04-014 | mountable/adhesive connector/pad/adaptor, hypo allergic, medical grade, water resistant, for passive bone conduction sound processor. Must be compatible with RT 274-04-005 | Offer 1 | Each
114 | RT274-04-015 | mountable cap, various colours and sizes where applicable,for digital, passive bone conduction sound processor, must be compatible with RT 274-04-005 | Offer 1
115 | RT274-04-016 | Direct streaming device to facilitate ease of listening in difficult listening situations compatible with all passive bone conduction sound processors awarded on the tender. Must be compatible with RT 274-04-005 | Offer 1
116 | RT274-04-017 | Repair and service of all detacheable digital passive bone conduction sound processors awarded on the tender. Must be compatible with RT 274-04-005 | Offer 1
117 | RT274-05-001 | Digital rechargeable BTE for mild-moderate hearing loss (25-55dB), MPO: 75-125dB gain 40-60dB, with Bluetooth compatibility, with tinnitus management, with frequency-lowering. | Offer 1 | Each
118 | RT274-05-002 | Digital non-rechargeable BTE for mild-moderate hearing loss (25-55dB), MPO: 75-125dB gain 40-60dB, with Bluetooth compatibility, with tinnitus management, with frequency-lowering. | Offer 1 | Each
119 | RT274-05-003 | Digital rechargeable BTE for mild-moderate hearing loss (25-55dB), MPO: 75-125dB gain 40-60dB, without Bluetooth compatibility, with tinnitus management, with frequency-lowering. | Offer 1 | Each
120 | RT274-05-004 | Digital non-rechargeable BTE for mild-moderate hearing loss (25-55dB), MPO: 75-125dB gain 40-60dB, without Bluetooth compatibility, with tinnitus management, with frequency-lowering. | Offer 1 | Each
121 | RT274-05-005 | Digital rechargeable BTE for mild-moderate hearing loss (25-55dB), MPO: 75-125dB gain 40-60dB, with Bluetooth compatibility, without tinnitus management, with frequency-lowering. | Offer 1 | Each
122 | RT274-05-006 | Digital non-rechargeable BTE for mild-moderate hearing loss (25-55dB), MPO: 75-125dB gain 40-60dB, with Bluetooth compatibility, without tinnitus management, with frequency-lowering. | Offer 1 | Each
123 | RT274-05-007 | Digital rechargeable BTE for mild-moderate hearing loss (25-55dB), MPO: 75-125dB gain 40-60dB, without Bluetooth compatibility, without tinnitus management, with frequency-lowering. | Offer 1 | Each
124 | RT274-05-008 | Digital non-rechargeable BTE for mild-moderate hearing loss (25-55dB), MPO: 75-125dB gain 40-60dB, without Bluetooth compatibility, without tinnitus management, with frequency-lowering. | Offer 1 | Each
125 | RT274-05-009 | Digital rechargeable BTE for mild-moderate hearing loss (25-55dB), MPO: 75-125dB gain 40-60dB, with Bluetooth compatibility, with tinnitus management, without frequency-lowering. | Offer 1 | Each
126 | RT274-05-010 | Digital non-rechargeable BTE for mild-moderate hearing loss (25-55dB), MPO: 75-125dB gain 40-60dB, with Bluetooth compatibility, with tinnitus management, without frequency-lowering. | Offer 1 | Each
127 | RT274-05-011 | Digital rechargeable BTE for mild-moderate hearing loss (25-55dB), MPO: 75-125dB gain 40-60dB, without Bluetooth compatibility, with tinnitus management, without frequency-lowering. | Offer 1 | Each
128 | RT274-05-012 | Digital non-rechargeable BTE for mild-moderate hearing loss (25-55dB), MPO: 75-125dB gain 40-60dB, without Bluetooth compatibility, with tinnitus management, without frequency-lowering. | Offer 1 | Each
129 | RT274-05-013 | Digital rechargeable BTE for mild-moderate hearing loss (25-55dB), MPO: 75-125dB gain 40-60dB, with Bluetooth compatibility, without tinnitus management, without frequency-lowering. | Offer 1 | Each
130 | RT274-05-014 | Digital non-rechargeable BTE for mild-moderate hearing loss (25-55dB), MPO: 75-125dB gain 40-60dB, with Bluetooth compatibility, without tinnitus management, without frequency-lowering. | Offer 1 | Each
131 | RT274-05-015 | Digital rechargeable BTE for mild-moderate hearing loss (25-55dB), MPO: 75-125dB gain 40-60dB, without Bluetooth compatibility, without tinnitus management, without frequency-lowering. | Offer 1 | Each
132 | RT274-05-016 | Digital non-rechargeable BTE for mild-moderate hearing loss (25-55dB), MPO: 75-125dB gain 40-60dB, without Bluetooth compatibility, without tinnitus management, without frequency-lowering. | Offer 1 | Each
133 | RT274-05-017 | Digital rechargeable BTE for moderate-severe hearing loss (56-70dB), MPO: 116-135dB gain 46-70dB, with Bluetooth compatibility, with tinnitus management, with frequency-lowering. | Offer 1 | Each
134 | RT274-05-018 | Digital non-rechargeable BTE for moderate-severe hearing loss (56-70dB), MPO: 116-135dB gain 46-70dB, with Bluetooth compatibility, with tinnitus management, with frequency-lowering. | Offer 1 | Each
135 | RT274-05-019 | Digital rechargeable BTE for moderate-severe hearing loss (56-70dB), MPO: 116-135dB gain 46-70dB, without Bluetooth compatibility, with tinnitus management, with frequency-lowering. | Offer 1 | Each
136 | RT274-05-020 | Digital non-rechargeable BTE for moderate-severe hearing loss (56-70dB), MPO: 116-135dB gain 46-70dB, without Bluetooth compatibility, with tinnitus management, with frequency-lowering. | Offer 1 | Each
137 | RT274-05-021 | Digital rechargeable BTE for moderate-severe hearing loss (56-70dB), MPO: 116-135dB gain 46-70dB, with Bluetooth compatibility, without tinnitus management, with frequency-lowering. | Offer 1 | Each
138 | RT274-05-022 | Digital non-rechargeable BTE for moderate-severe hearing loss (56-70dB), MPO: 116-135dB gain 46-70dB, with Bluetooth compatibility, without tinnitus management, with frequency-lowering. | Offer 1 | Each
139 | RT274-05-023 | Digital rechargeable BTE for moderate-severe hearing loss (56-70dB), MPO: 116-135dB gain 46-70dB, without Bluetooth compatibility, without tinnitus management, with frequency-lowering. | Offer 1 | Each
140 | RT274-05-024 | Digital non-rechargeable BTE for moderate-severe hearing loss (56-70dB), MPO: 116-135dB gain 46-70dB, without Bluetooth compatibility, without tinnitus management, with frequency-lowering. | Offer 1 | Each
141 | RT274-05-025 | Digital rechargeable BTE for moderate-severe hearing loss (56-70dB), MPO: 116-135dB gain 46-70dB, with Bluetooth compatibility, with tinnitus management, without frequency-lowering. | Offer 1 | Each
142 | RT274-05-026 | Digital non-rechargeable BTE for moderate-severe hearing loss (56-70dB), MPO: 116-135dB gain 46-70dB, with Bluetooth compatibility, with tinnitus management, without frequency-lowering. | Offer 1 | Each
143 | RT274-05-027 | Digital rechargeable BTE for moderate-severe hearing loss (56-70dB), MPO: 116-135dB gain 46-70dB, without Bluetooth compatibility, with tinnitus management, without frequency-lowering. | Offer 1 | Each
144 | RT274-05-028 | Digital non-rechargeable BTE for moderate-severe hearing loss (56-70dB), MPO: 116-135dB gain 46-70dB, without Bluetooth compatibility, with tinnitus management, without frequency-lowering. | Offer 1 | Each
145 | RT274-05-029 | Digital rechargeable BTE for moderate-severe hearing loss (56-70dB), MPO: 116-135dB gain 46-70dB, with Bluetooth compatibility, without tinnitus management, without frequency-lowering. | Offer 1 | Each
146 | RT274-05-030 | Digital non-rechargeable BTE for moderate-severe hearing loss (56-70dB), MPO: 116-135dB gain 46-70dB, with Bluetooth compatibility, without tinnitus management, without frequency-lowering. | Offer 1 | Each
147 | RT274-05-031 | Digital rechargeable BTE for moderate-severe hearing loss (56-70dB), MPO: 116-135dB gain 46-70dB, without Bluetooth compatibility, without tinnitus management, without frequency-lowering. | Offer 1 | Each
148 | RT274-05-032 | Digital non-rechargeable BTE for moderate-severe hearing loss (56-70dB), MPO: 116-135dB gain 46-70dB, without Bluetooth compatibility, without tinnitus management, without frequency-lowering. | Offer 1 | Each
149 | RT274-05-033 | Digital rechargeable BTE for severe-profound hearing loss (71-90dB), MPO: MPO: 126-145dB gain 61-85dB, with Bluetooth compatibility, with tinnitus management, with frequency-lowering. | Offer 1 | Each
150 | RT274-05-034 | Digital non-rechargeable BTE for severe-profound hearing loss (71-90dB), MPO: MPO: 126-145dB gain 61-85dB, with Bluetooth compatibility, with tinnitus management, with frequency-lowering. | Offer 1 | Each
151 | RT274-05-035 | Digital rechargeable BTE for severe-profound hearing loss (71-90dB), MPO: MPO: 126-145dB gain 61-85dB, without Bluetooth compatibility, with tinnitus management, with frequency-lowering. | Offer 1 | Each
152 | RT274-05-036 | Digital non-rechargeable BTE for severe-profound hearing loss (71-90dB), MPO: MPO: 126-145dB gain 61-85dB, without Bluetooth compatibility, with tinnitus management, with frequency-lowering. | Offer 1 | Each
153 | RT274-05-037 | Digital rechargeable BTE for severe-profound hearing loss (71-90dB), MPO: MPO: 126-145dB gain 61-85dB, with Bluetooth compatibility, without tinnitus management, with frequency-lowering. | Offer 1 | Each
154 | RT274-05-038 | Digital non-rechargeable BTE for severe-profound hearing loss (71-90dB), MPO: MPO: 126-145dB gain 61-85dB, with Bluetooth compatibility, without tinnitus management, with frequency-lowering. | Offer 1 | Each
155 | RT274-05-039 | Digital rechargeable BTE for severe-profound hearing loss (71-90dB), MPO: MPO: 126-145dB gain 61-85dB, without Bluetooth compatibility, without tinnitus management, with frequency-lowering. | Offer 1 | Each
156 | RT274-05-040 | Digital non-rechargeable BTE for severe-profound hearing loss (71-90dB), MPO: MPO: 126-145dB gain 61-85dB, without Bluetooth compatibility, without tinnitus management, with frequency-lowering. | Offer 1 | Each
157 | RT274-05-041 | Digital rechargeable BTE for severe-profound hearing loss (71-90dB), MPO: MPO: 126-145dB gain 61-85dB, with Bluetooth compatibility, with tinnitus management, without frequency-lowering. | Offer 1 | Each
158 | RT274-05-042 | Digital non-rechargeable BTE for severe-profound hearing loss (71-90dB), MPO: MPO: 126-145dB gain 61-85dB, with Bluetooth compatibility, with tinnitus management, without frequency-lowering. | Offer 1 | Each
159 | RT274-05-043 | Digital rechargeable BTE for severe-profound hearing loss (71-90dB), MPO: MPO: 126-145dB gain 61-85dB, without Bluetooth compatibility, with tinnitus management, without frequency-lowering. | Offer 1 | Each
160 | RT274-05-044 | Digital non-rechargeable BTE for severe-profound hearing loss (71-90dB), MPO: MPO: 126-145dB gain 61-85dB, without Bluetooth compatibility, with tinnitus management, without frequency-lowering. | Offer 1 | Each
161 | RT274-05-045 | Digital rechargeable BTE for severe-profound hearing loss (71-90dB), MPO: MPO: 126-145dB gain 61-85dB, with Bluetooth compatibility, without tinnitus management, without frequency-lowering. | Offer 1 | Each
162 | RT274-05-046 | Digital non-rechargeable BTE for severe-profound hearing loss (71-90dB), MPO: MPO: 126-145dB gain 61-85dB, with Bluetooth compatibility, without tinnitus management, without frequency-lowering. | Offer 1 | Each
163 | RT274-05-047 | Digital rechargeable BTE for severe-profound hearing loss (71-90dB), MPO: MPO: 126-145dB gain 61-85dB, without Bluetooth compatibility, without tinnitus management, without frequency-lowering. | Offer 1 | Each
164 | RT274-05-048 | Digital non-rechargeable BTE for severe-profound hearing loss (71-90dB), MPO: MPO: 126-145dB gain 61-85dB, without Bluetooth compatibility, without tinnitus management, without frequency-lowering. | Offer 1 | Each
165 | RT274-05-049 | Super power digital rechargeable BTE for profound hearing loss (90+dB), MPO: 140+dB, gain 80+dB, with Bluetooth compatibility, with tinnitus management, with frequency-lowering. | Offer 1 | Each
166 | RT274-05-050 | Super power digital non-rechargeable BTE for profound hearing loss (90+dB), MPO: 140+dB, gain 80+dB, with Bluetooth compatibility, with tinnitus management, with frequency-lowering. | Offer 1 | Each
167 | RT274-05-051 | Super power digital rechargeable BTE for profound hearing loss (90+dB), MPO: 140+dB, gain 80+dB, without Bluetooth compatibility, with tinnitus management, with frequency-lowering. | Offer 1 | Each
168 | RT274-05-052 | Super power digital non-rechargeable BTE for profound hearing loss (90+dB), MPO: 140+dB, gain 80+dB, without Bluetooth compatibility, with tinnitus management, with frequency-lowering. | Offer 1 | Each
169 | RT274-05-053 | Super power digital rechargeable BTE for profound hearing loss (90+dB), MPO: 140+dB, gain 80+dB, with Bluetooth compatibility, without tinnitus management, with frequency-lowering. | Offer 1 | Each
170 | RT274-05-054 | Super power digital non-rechargeable BTE for profound hearing loss (90+dB), MPO: 140+dB, gain 80+dB, with Bluetooth compatibility, without tinnitus management, with frequency-lowering. | Offer 1 | Each
171 | RT274-05-055 | Super power digital rechargeable BTE for profound hearing loss (90+dB), MPO: 140+dB, gain 80+dB, without Bluetooth compatibility, without tinnitus management, with frequency-lowering. | Offer 1 | Each
172 | RT274-05-056 | Super power digital non-rechargeable BTE for profound hearing loss (90+dB), MPO: 140+dB, gain 80+dB, without Bluetooth compatibility, without tinnitus management, with frequency-lowering. | Offer 1 | Each
173 | RT274-05-057 | Super power digital rechargeable BTE for profound hearing loss (90+dB), MPO: 140+dB, gain 80+dB, with Bluetooth compatibility, with tinnitus management, without frequency-lowering. | Offer 1 | Each
174 | RT274-05-058 | Super power digital non-rechargeable BTE for profound hearing loss (90+dB), MPO: 140+dB, gain 80+dB, with Bluetooth compatibility, with tinnitus management, without frequency-lowering. | Offer 1 | Each
175 | RT274-05-059 | Super power digital rechargeable BTE for profound hearing loss (90+dB), MPO: 140+dB, gain 80+dB, without Bluetooth compatibility, with tinnitus management, without frequency-lowering. | Offer 1 | Each
176 | RT274-05-060 | Super power digital non-rechargeable BTE for profound hearing loss (90+dB), MPO: 140+dB, gain 80+dB, without Bluetooth compatibility, with tinnitus management, without frequency-lowering. | Offer 1 | Each
177 | RT274-05-061 | Super power digital rechargeable BTE for profound hearing loss (90+dB), MPO: 140+dB, gain 80+dB, with Bluetooth compatibility, without tinnitus management, without frequency-lowering. | Offer 1 | Each
178 | RT274-05-062 | Super power digital non-rechargeable BTE for profound hearing loss (90+dB), MPO: 140+dB, gain 80+dB, with Bluetooth compatibility, without tinnitus management, without frequency-lowering. | Offer 1 | Each
179 | RT274-05-063 | Super power digital rechargeable BTE for profound hearing loss (90+dB), MPO: 140+dB, gain 80+dB, without Bluetooth compatibility, without tinnitus management, without frequency-lowering. | Offer 1 | Each
180 | RT274-05-064 | Super power digital non-rechargeable BTE for profound hearing loss (90+dB), MPO: 140+dB, gain 80+dB, without Bluetooth compatibility, without tinnitus management, without frequency-lowering. | Offer 1 | Each
181 | RT274-06-001 | Digital rechargeable ITE for mild-moderate hearing loss (25-55dB), MPO: 75-125dB gain 40-60dB, with Bluetooth compatibility, with tinnitus management, with frequency-lowering. | Offer 1 | Each
182 | RT274-06-002 | Digital non-rechargeable ITE for mild-moderate hearing loss (25-55dB), MPO: 75-125dB gain 40-60dB, with Bluetooth compatibility, with tinnitus management, with frequency-lowering. | Offer 1 | Each
183 | RT274-06-003 | Digital rechargeable ITE for mild-moderate hearing loss (25-55dB), MPO: 75-125dB gain 40-60dB, without Bluetooth compatibility, with tinnitus management, with frequency-lowering. | Offer 1 | Each
184 | RT274-06-004 | Digital non-rechargeable ITE for mild-moderate hearing loss (25-55dB), MPO: 75-125dB gain 40-60dB, without Bluetooth compatibility, with tinnitus management, with frequency-lowering. | Offer 1 | Each
185 | RT274-06-005 | Digital rechargeable ITE for mild-moderate hearing loss (25-55dB), MPO: 75-125dB gain 40-60dB, with Bluetooth compatibility, without tinnitus management, with frequency-lowering. | Offer 1 | Each
186 | RT274-06-006 | Digital non-rechargeable ITE for mild-moderate hearing loss (25-55dB), MPO: 75-125dB gain 40-60dB, with Bluetooth compatibility, without tinnitus management, with frequency-lowering. | Offer 1 | Each
187 | RT274-06-007 | Digital rechargeable ITE for mild-moderate hearing loss (25-55dB), MPO: 75-125dB gain 40-60dB, without Bluetooth compatibility, without tinnitus management, with frequency-lowering. | Offer 1 | Each
188 | RT274-06-008 | Digital non-rechargeable ITE for mild-moderate hearing loss (25-55dB), MPO: 75-125dB gain 40-60dB, without Bluetooth compatibility, without tinnitus management, with frequency-lowering. | Offer 1 | Each
189 | RT274-06-009 | Digital rechargeable ITE for mild-moderate hearing loss (25-55dB), MPO: 75-125dB gain 40-60dB, with Bluetooth compatibility, with tinnitus management, without frequency-lowering. | Offer 1 | Each
190 | RT274-06-010 | Digital non-rechargeable ITE for mild-moderate hearing loss (25-55dB), MPO: 75-125dB gain 40-60dB, with Bluetooth compatibility, with tinnitus management, without frequency-lowering. | Offer 1 | Each
191 | RT274-06-011 | Digital rechargeable ITE for mild-moderate hearing loss (25-55dB), MPO: 75-125dB gain 40-60dB, without Bluetooth compatibility, with tinnitus management, without frequency-lowering. | Offer 1 | Each
192 | RT274-06-012 | Digital non-rechargeable ITE for mild-moderate hearing loss (25-55dB), MPO: 75-125dB gain 40-60dB, without Bluetooth compatibility, with tinnitus management, without frequency-lowering. | Offer 1 | Each
193 | RT274-06-013 | Digital rechargeable ITE for mild-moderate hearing loss (25-55dB), MPO: 75-125dB gain 40-60dB, with Bluetooth compatibility, without tinnitus management, without frequency-lowering. | Offer 1 | Each
194 | RT274-06-014 | Digital non-rechargeable ITE for mild-moderate hearing loss (25-55dB), MPO: 75-125dB gain 40-60dB, with Bluetooth compatibility, without tinnitus management, without frequency-lowering. | Offer 1 | Each
195 | RT274-06-015 | Digital rechargeable ITE for mild-moderate hearing loss (25-55dB), MPO: 75-125dB gain 40-60dB, without Bluetooth compatibility, without tinnitus management, without frequency-lowering. | Offer 1 | Each
196 | RT274-06-016 | Digital non-rechargeable ITE for mild-moderate hearing loss (25-55dB), MPO: 75-125dB gain 40-60dB, without Bluetooth compatibility, without tinnitus management, without frequency-lowering. | Offer 1 | Each
197 | RT274-06-017 | Digital rechargeable ITE for moderate-severe hearing loss (56-70dB), MPO: 116-135dB gain 46-70dB, with Bluetooth compatibility, with tinnitus management, with frequency-lowering. | Offer 1 | Each
198 | RT274-06-018 | Digital non-rechargeable ITE for moderate-severe hearing loss (56-70dB), MPO: 116-135dB gain 46-70dB, with Bluetooth compatibility, with tinnitus management, with frequency-lowering. | Offer 1 | Each
199 | RT274-06-019 | Digital rechargeable ITE for moderate-severe hearing loss (56-70dB), MPO: 116-135dB gain 46-70dB, without Bluetooth compatibility, with tinnitus management, with frequency-lowering. | Offer 1 | Each
200 | RT274-06-020 | Digital non-rechargeable ITE for moderate-severe hearing loss (56-70dB), MPO: 116-135dB gain 46-70dB, without Bluetooth compatibility, with tinnitus management, with frequency-lowering. | Offer 1 | Each
201 | RT274-06-021 | Digital rechargeable ITE for moderate-severe hearing loss (56-70dB), MPO: 116-135dB gain 46-70dB, with Bluetooth compatibility, without tinnitus management, with frequency-lowering. | Offer 1 | Each
202 | RT274-06-022 | Digital non-rechargeable ITE for moderate-severe hearing loss (56-70dB), MPO: 116-135dB gain 46-70dB, with Bluetooth compatibility, without tinnitus management, with frequency-lowering. | Offer 1 | Each
203 | RT274-06-023 | Digital rechargeable ITE for moderate-severe hearing loss (56-70dB), MPO: 116-135dB gain 46-70dB, without Bluetooth compatibility, without tinnitus management, with frequency-lowering. | Offer 1 | Each
204 | RT274-06-024 | Digital non-rechargeable ITE for moderate-severe hearing loss (56-70dB), MPO: 116-135dB gain 46-70dB, without Bluetooth compatibility, without tinnitus management, with frequency-lowering | Offer 1 | Each
205 | RT274-06-025 | Digital rechargeable ITE for moderate-severe hearing loss (56-70dB), MPO: 116-135dB gain 46-70dB, with Bluetooth compatibility, with tinnitus management, without frequency-lowering. | Offer 1 | Each
206 | RT274-06-026 | Digital non-rechargeable ITE for moderate-severe hearing loss (56-70dB), MPO: 116-135dB gain 46-70dB, with Bluetooth compatibility, with tinnitus management, without frequency-lowering. | Offer 1 | Each
207 | RT274-06-027 | Digital rechargeable ITE for moderate-severe hearing loss (56-70dB), MPO: 116-135dB gain 46-70dB, without Bluetooth compatibility, with tinnitus management, without frequency-lowering. | Offer 1 | Each
208 | RT274-06-028 | Digital non-rechargeable ITE for moderate-severe hearing loss (56-70dB), MPO: 116-135dB gain 46-70dB, without Bluetooth compatibility, with tinnitus management, without frequency-lowering. | Offer 1 | Each
209 | RT274-06-029 | Digital rechargeable ITE for moderate-severe hearing loss (56-70dB), MPO: 116-135dB gain 46-70dB, with Bluetooth compatibility, without tinnitus management, without frequency-lowering | Offer 1 | Each
210 | RT274-06-030 | Digital non-rechargeable ITE for moderate-severe hearing loss (56-70dB), MPO: 116-135dB gain 46-70dB, with Bluetooth compatibility, without tinnitus management, without frequency-lowering. | Offer 1 | Each
211 | RT274-06-031 | Digital rechargeable ITE for moderate-severe hearing loss (56-70dB), MPO: 116-135dB gain 46-70dB, without Bluetooth compatibility, without tinnitus management, without frequency-lowering | Offer 1 | Each
212 | RT274-06-032 | Digital non-rechargeable ITE for moderate-severe hearing loss (56-70dB), MPO: 116-135dB gain 46-70dB, without Bluetooth compatibility, without tinnitus management, without frequency-lowering. | Offer 1 | Each
213 | RT274-06-033 | Digital rechargeable ITE for severe-profound hearing loss (71-90dB), MPO: MPO: 126-145dB gain 61-85dB, with Bluetooth compatibility, with tinnitus management, with frequency-lowering. | Offer 1 | Each
214 | RT274-06-034 | Digital non-rechargeable ITE for severe-profound hearing loss (71-90dB), MPO: MPO: 126-145dB gain 61-85dB, with Bluetooth compatibility, with tinnitus management, with frequency-lowering. | Offer 1 | Each
215 | RT274-06-035 | Digital rechargeable ITE for severe-profound hearing loss (71-90dB), MPO: MPO: 126-145dB gain 61-85dB, without Bluetooth compatibility, with tinnitus management, with frequency-lowering. | Offer 1 | Each
216 | RT274-06-036 | Digital non-rechargeable ITE for severe-profound hearing loss (71-90dB), MPO: MPO: 126-145dB gain 61-85dB, without Bluetooth compatibility, with tinnitus management, with frequency-lowering. | Offer 1 | Each
217 | RT274-06-037 | Digital rechargeable ITE for severe-profound hearing loss (71-90dB), MPO: MPO: 126-145dB gain 61-85dB, with Bluetooth compatibility, without tinnitus management, with frequency-lowering. | Offer 1 | Each
218 | RT274-06-038 | Digital non-rechargeable ITE for severe-profound hearing loss (71-90dB), MPO: MPO: 126-145dB gain 61-85dB, with Bluetooth compatibility, without tinnitus management, with frequency-lowering. | Offer 1 | Each
219 | RT274-06-039 | Digital rechargeable ITE for severe-profound hearing loss (71-90dB), MPO: MPO: 126-145dB gain 61-85dB, without Bluetooth compatibility, without tinnitus management, with frequency-lowering. | Offer 1 | Each
220 | RT274-06-040 | Digital non-rechargeable ITE for severe-profound hearing loss (71-90dB), MPO: MPO: 126-145dB gain 61-85dB, without Bluetooth compatibility, without tinnitus management, with frequency-lowering. | Offer 1 | Each
221 | RT274-06-041 | Digital rechargeable ITE for severe-profound hearing loss (71-90dB), MPO: MPO: 126-145dB gain 61-85dB, with Bluetooth compatibility, with tinnitus management, without frequency-lowering. | Offer 1 | Each
222 | RT274-06-042 | Digital non-rechargeable ITE for severe-profound hearing loss (71-90dB), MPO: MPO: 126-145dB gain 61-85dB, with Bluetooth compatibility, with tinnitus management, without frequency-lowering. | Offer 1 | Each
223 | RT274-06-043 | Digital rechargeable ITE for severe-profound hearing loss (71-90dB), MPO: MPO: 126-145dB gain 61-85dB, without Bluetooth compatibility, with tinnitus management, without frequency-lowering. | Offer 1 | Each
224 | RT274-06-044 | Digital non-rechargeable ITE for severe-profound hearing loss (71-90dB), MPO: MPO: 126-145dB gain 61-85dB, without Bluetooth compatibility, with tinnitus management, without frequency-lowering. | Offer 1 | Each
225 | RT274-06-045 | Digital rechargeable ITE for severe-profound hearing loss (71-90dB), MPO: MPO: 126-145dB gain 61-85dB, with Bluetooth compatibility, without tinnitus management, without frequency-lowering. | Offer 1 | Each
226 | RT274-06-046 | Digital non-rechargeable ITE for severe-profound hearing loss (71-90dB), MPO: MPO: 126-145dB gain 61-85dB, with Bluetooth compatibility, without tinnitus management, without frequency-lowering. | Offer 1 | Each
227 | RT274-06-047 | Digital rechargeable ITE for severe-profound hearing loss (71-90dB), MPO: MPO: 126-145dB gain 61-85dB, without Bluetooth compatibility, without tinnitus management, without frequency-lowering. | Offer 1 | Each
228 | RT274-06-048 | Digital non-rechargeable ITE for severe-profound hearing loss (71-90dB), MPO: MPO: 126-145dB gain 61-85dB, without Bluetooth compatibility, without tinnitus management, without frequency-lowering. | Offer 1 | Each
229 | RT274-06-049 | Digital rechargeable ITC for mild-moderate hearing loss (25-55dB), MPO: 75-125dB gain 40-60dB, with Bluetooth compatibility, with tinnitus management, with frequency-lowering. | Offer 1 | Each
230 | RT274-06-050 | Digital non-rechargeable ITC for mild-moderate hearing loss (25-55dB), MPO: 75-125dB gain 40-60dB, with Bluetooth compatibility, with tinnitus management, with frequency-lowering. | Offer 1 | Each
231 | RT274-06-051 | Digital rechargeable ITC for mild-moderate hearing loss (25-55dB), MPO: 75-125dB gain 40-60dB, without Bluetooth compatibility, with tinnitus management, with frequency-lowering | Offer 1 | Each
232 | RT274-06-052 | Digital non-rechargeable ITC for mild-moderate hearing loss (25-55dB), MPO: 75-125dB gain 40-60dB, without Bluetooth compatibility, with tinnitus management, with frequency-lowering. | Offer 1 | Each
233 | RT274-06-053 | Digital rechargeable ITC for mild-moderate hearing loss (25-55dB), MPO: 75-125dB gain 40-60dB, with Bluetooth compatibility, without tinnitus management, with frequency-lowering. | Offer 1 | Each
234 | RT274-06-054 | Digital non-rechargeable ITC for mild-moderate hearing loss (25-55dB), MPO: 75-125dB gain 40-60dB, with Bluetooth compatibility, without tinnitus management, with frequency-lowering. | Offer 1 | Each
235 | RT274-06-055 | Digital rechargeable ITC for mild-moderate hearing loss (25-55dB), MPO: 75-125dB gain 40-60dB, without Bluetooth compatibility, without tinnitus management, with frequency-lowering. | Offer 1 | Each
236 | RT274-06-056 | Digital non-rechargeable ITC for mild-moderate hearing loss (25-55dB), MPO: 75-125dB gain 40-60dB, without Bluetooth compatibility, without tinnitus management, with frequency-lowering. | Offer 1 | Each
237 | RT274-06-057 | Digital rechargeable ITC for mild-moderate hearing loss (25-55dB), MPO: 75-125dB gain 40-60dB, with Bluetooth compatibility, with tinnitus management, without frequency-lowering. | Offer 1 | Each
238 | RT274-06-058 | Digital non-rechargeable ITE for mild-moderate hearing loss (25-55dB), MPO: 75-125dB gain 40-60dB, with Bluetooth compatibility, with tinnitus management, without frequency-lowering. | Offer 1 | Each
239 | RT274-06-059 | Digital rechargeable ITC for mild-moderate hearing loss (25-55dB), MPO: 75-125dB gain 40-60dB, without Bluetooth compatibility, with tinnitus management, without frequency-lowering. | Offer 1 | Each
240 | RT274-06-060 | Digital non-rechargeable ITC for mild-moderate hearing loss (25-55dB), MPO: 75-125dB gain 40-60dB, without Bluetooth compatibility, with tinnitus management, without frequency-lowering. | Offer 1 | Each
241 | RT274-06-061 | Digital rechargeable ITC for mild-moderate hearing loss (25-55dB), MPO: 75-125dB gain 40-60dB, with Bluetooth compatibility, without tinnitus management, without frequency-lowering. | Offer 1 | Each
242 | RT274-06-062 | Digital non-rechargeable ITC for mild-moderate hearing loss (25-55dB), MPO: 75-125dB gain 40-60dB, with Bluetooth compatibility, without tinnitus management, without frequency-lowering. | Offer 1 | Each
243 | RT274-06-063 | Digital rechargeable ITC for mild-moderate hearing loss (25-55dB), MPO: 75-125dB gain 40-60dB, without Bluetooth compatibility, without tinnitus management, without frequency-lowering. | Offer 1 | Each
244 | RT274-06-064 | Digital non-rechargeable ITC for mild-moderate hearing loss (25-55dB), MPO: 75-125dB gain 40-60dB, without Bluetooth compatibility, without tinnitus management, without frequency-lowering. | Offer 1 | Each
245 | RT274-06-065 | Digital rechargeable ITC for moderate-severe hearing loss (56-70dB), MPO: 116-135dB gain 46-70dB, with Bluetooth compatibility, with tinnitus management, with frequency-lowering. | Offer 1 | Each
246 | RT274-06-066 | Digital non-rechargeable ITC for moderate-severe hearing loss (56-70dB), MPO: 116-135dB gain 46-70dB, with Bluetooth compatibility, with tinnitus management, with frequency-lowering. | Offer 1 | Each
247 | RT274-06-067 | Digital rechargeable ITC for moderate-severe hearing loss (56-70dB), MPO: 116-135dB gain 46-70dB, without Bluetooth compatibility, with tinnitus management, with frequency-lowering. | Offer 1 | Each
248 | RT274-06-068 | Digital non-rechargeable ITC for moderate-severe hearing loss (56-70dB), MPO: 116-135dB gain 46-70dB, without Bluetooth compatibility, with tinnitus management, with frequency-lowering. | Offer 1 | Each
249 | RT274-06-069 | Digital rechargeable ITC for moderate-severe hearing loss (56-70dB), MPO: 116-135dB gain 46-70dB, with Bluetooth compatibility, without tinnitus management, with frequency-lowering. | Offer 1 | Each
250 | RT274-06-070 | Digital non-rechargeable ITC for moderate-severe hearing loss (56-70dB), MPO: 116-135dB gain 46-70dB, with Bluetooth compatibility, without tinnitus management, with frequency-lowering. | Offer 1 | Each
251 | RT274-06-071 | Digital rechargeable ITC for moderate-severe hearing loss (56-70dB), MPO: 116-135dB gain 46-70dB, without Bluetooth compatibility, without tinnitus management, with frequency-lowering, | Offer 1 | Each
252 | RT274-06-072 | Digital non-rechargeable ITC for moderate-severe hearing loss (56-70dB), MPO: 116-135dB gain 46-70dB, without Bluetooth compatibility, without tinnitus management, with frequency-lowering. | Offer 1 | Each
253 | RT274-06-073 | Digital rechargeable ITC for moderate-severe hearing loss (56-70dB), MPO: 116-135dB gain 46-70dB, with Bluetooth compatibility, with tinnitus management, without frequency-lowering. | Offer 1 | Each
254 | RT274-06-074 | Digital non-rechargeable ITC for moderate-severe hearing loss (56-70dB), MPO: 116-135dB gain 46-70dB, with Bluetooth compatibility, with tinnitus management, without frequency-lowering. | Offer 1 | Each
255 | RT274-06-075 | Digital rechargeable ITC for moderate-severe hearing loss (56-70dB), MPO: 116-135dB gain 46-70dB, without Bluetooth compatibility, with tinnitus management, without frequency-lowering. | Offer 1 | Each
256 | RT274-06-076 | Digital non-rechargeable ITC for moderate-severe hearing loss (56-70dB), MPO: 116-135dB gain 46-70dB, without Bluetooth compatibility, with tinnitus management, without frequency-lowering. | Offer 1 | Each
257 | RT274-06-077 | Digital rechargeable ITC for moderate-severe hearing loss (56-70dB), MPO: 116-135dB gain 46-70dB, with Bluetooth compatibility, without tinnitus management, without frequency-lowering. | Offer 1 | Each
258 | RT274-06-078 | Digital non-rechargeable ITC for moderate-severe hearing loss (56-70dB), MPO: 116-135dB gain 46-70dB, with Bluetooth compatibility, without tinnitus management, without frequency-lowering. | Offer 1 | Each
259 | RT274-06-079 | Digital rechargeable ITC for moderate-severe hearing loss (56-70dB), MPO: 116-135dB gain 46-70dB, without Bluetooth compatibility, without tinnitus management, without frequency-lowering. | Offer 1 | Each
260 | RT274-06-080 | Digital non-rechargeable ITC for moderate-severe hearing loss (56-70dB), MPO: 116-135dB gain 46-70dB, without Bluetooth compatibility, without tinnitus management, without frequency-lowering. | Offer 1 | Each
261 | RT274-06-081 | Digital rechargeable ITC for severe-profound hearing loss (71-90dB), MPO: MPO: 126-145dB gain 61-85dB, with Bluetooth compatibility, with tinnitus management, with frequency-lowering. | Offer 1 | Each
262 | RT274-06-082 | Digital non-rechargeable ITC for severe-profound hearing loss (71-90dB), MPO: MPO: 126-145dB gain 61-85dB, with Bluetooth compatibility, with tinnitus management, with frequency-lowering. | Offer 1 | Each
263 | RT274-06-083 | Digital rechargeable ITC for severe-profound hearing loss (71-90dB), MPO: MPO: 126-145dB gain 61-85dB, without Bluetooth compatibility, with tinnitus management, with frequency-lowering. | Offer 1 | Each
264 | RT274-06-084 | Digital non-rechargeable ITC for severe-profound hearing loss (71-90dB), MPO: MPO: 126-145dB gain 61-85dB, without Bluetooth compatibility, with tinnitus management, with frequency-lowering. | Offer 1 | Each
265 | RT274-06-085 | Digital rechargeable ITC for severe-profound hearing loss (71-90dB), MPO: MPO: 126-145dB gain 61-85dB, with Bluetooth compatibility, without tinnitus management, with frequency-lowering. | Offer 1 | Each
266 | RT274-06-086 | Digital rechargeable ITC for severe-profound hearing loss (71-90dB), MPO: MPO: 126-145dB gain 61-85dB, with Bluetooth compatibility, without tinnitus management, with frequency-lowering. | Offer 1 | Each
267 | RT274-06-087 | Digital non-rechargeable ITC for severe-profound hearing loss (71-90dB), MPO: MPO: 126-145dB gain 61-85dB, with Bluetooth compatibility, without tinnitus management, with frequency-lowering. | Offer 1 | Each
268 | RT274-06-088 | Digital rechargeable ITC for severe-profound hearing loss (71-90dB), MPO: MPO: 126-145dB gain 61-85dB, without Bluetooth compatibility, without tinnitus management, with frequency-lowering. | Offer 1 | Each
269 | RT274-06-089 | Digital non-rechargeable ITC for severe-profound hearing loss (71-90dB), MPO: MPO: 126-145dB gain 61-85dB, without Bluetooth compatibility, without tinnitus management, with frequency-lowering. | Offer 1 | Each
270 | RT274-06-090 | Digital rechargeable ITC for severe-profound hearing loss (71-90dB), MPO: MPO: 126-145dB gain 61-85dB, with Bluetooth compatibility, with tinnitus management, without frequency-lowering. | Offer 1 | Each
271 | RT274-06-091 | Digital non-rechargeable ITC for severe-profound hearing loss (71-90dB), MPO: MPO: 126-145dB gain 61-85dB, with Bluetooth compatibility, with tinnitus management, without frequency-lowering. | Offer 1 | Each
272 | RT274-06-092 | Digital rechargeable ITC for severe-profound hearing loss (71-90dB), MPO: MPO: 126-145dB gain 61-85dB, without Bluetooth compatibility, with tinnitus management, without frequency-lowering. | Offer 1 | Each
273 | RT274-06-093 | Digital non-rechargeable ITC for severe-profound hearing loss (71-90dB), MPO: MPO: 126-145dB gain 61-85dB, without Bluetooth compatibility, with tinnitus management, without frequency-lowering. | Offer 1 | Each
274 | RT274-06-094 | Digital rechargeable ITC for severe-profound hearing loss (71-90dB), MPO: MPO: 126-145dB gain 61-85dB, with Bluetooth compatibility, without tinnitus management, without frequency-lowering. | Offer 1 | Each
275 | RT274-06-095 | Digital non-rechargeable ITC for severe-profound hearing loss (71-90dB), MPO: MPO: 126-145dB gain 61-85dB, with Bluetooth compatibility, without tinnitus management, without frequency-lowering. | Offer 1 | Each
276 | RT274-06-096 | Digital rechargeable ITC for severe-profound hearing loss (71-90dB), MPO: MPO: 126-145dB gain 61-85dB, without Bluetooth compatibility, without tinnitus management, without frequency-lowering | Offer 1 | Each
277 | RT274-06-097 | Digital non-rechargeable ITC for severe-profound hearing loss (71-90dB), MPO: MPO: 126-145dB gain 61-85dB, without Bluetooth compatibility, without tinnitus management, without frequency-lowering. | Offer 1 | Each
278 | RT274-06-098 | Digital rechargeable CIC for mild-moderate hearing loss (25-55dB), MPO: 75-125dB gain 40-60dB, with Bluetooth compatibility, with tinnitus management, with frequency-lowering. | Offer 1 | Each
279 | RT274-06-099 | Digital non-rechargeable CIC for mild-moderate hearing loss (25-55dB), MPO: 75-125dB gain 40-60dB, with Bluetooth compatibility, with tinnitus management, with frequency-lowering. | Offer 1 | Each
280 | RT274-06-100 | Digital rechargeable CIC for mild-moderate hearing loss (25-55dB), MPO: 75-125dB gain 40-60dB, without Bluetooth compatibility, with tinnitus management, with frequency-lowering. | Offer 1 | Each
281 | RT274-06-101 | Digital non-rechargeable CIC for mild-moderate hearing loss (25-55dB), MPO: 75-125dB gain 40-60dB, without Bluetooth compatibility, with tinnitus management, with frequency-lowering. | Offer 1 | Each
282 | RT274-06-102 | Digital rechargeable CIC for mild-moderate hearing loss (25-55dB), MPO: 75-125dB gain 40-60dB, with Bluetooth compatibility, without tinnitus management, with frequency-lowering. | Offer 1 | Each
283 | RT274-06-103 | Digital non-rechargeable CIC for mild-moderate hearing loss (25-55dB), MPO: 75-125dB gain 40-60dB, with Bluetooth compatibility, without tinnitus management, with frequency-lowering. | Offer 1 | Each
284 | RT274-06-104 | Digital rechargeable CIC for mild-moderate hearing loss (25-55dB), MPO: 75-125dB gain 40-60dB, without Bluetooth compatibility, without tinnitus management, with frequency-lowering. | Offer 1 | Each
285 | RT274-06-105 | Digital non-rechargeable CIC for mild-moderate hearing loss (25-55dB), MPO: 75-125dB gain 40-60dB, without Bluetooth compatibility, without tinnitus management, with frequency-lowering. | Offer 1 | Each
286 | RT274-06-106 | Digital rechargeable CIC for mild-moderate hearing loss (25-55dB), MPO: 75-125dB gain 40-60dB, with Bluetooth compatibility, with tinnitus management, without frequency-lowering. | Offer 1 | Each
287 | RT274-06-107 | Digital non-rechargeable ITE for mild-moderate hearing loss (25-55dB), MPO: 75-125dB gain 40-60dB, with Bluetooth compatibility, with tinnitus management, without frequency-lowering. | Offer 1 | Each
288 | RT274-06-108 | Digital rechargeable CIC for mild-moderate hearing loss (25-55dB), MPO: 75-125dB gain 40-60dB, without Bluetooth compatibility, with tinnitus management, without frequency-lowering. | Offer 1 | Each
289 | RT274-06-109 | Digital non-rechargeable CIC for mild-moderate hearing loss (25-55dB), MPO: 75-125dB gain 40-60dB, without Bluetooth compatibility, with tinnitus management, without frequency-lowering. | Offer 1 | Each
290 | RT274-06-110 | Digital rechargeable CIC for mild-moderate hearing loss (25-55dB), MPO: 75-125dB gain 40-60dB, with Bluetooth compatibility, without tinnitus management, without frequency-lowering. | Offer 1 | Each
291 | RT274-06-111 | Digital non-rechargeable CIC for mild-moderate hearing loss (25-55dB), MPO: 75-125dB gain 40-60dB, with Bluetooth compatibility, without tinnitus management, without frequency-lowering. | Offer 1 | Each
292 | RT274-06-112 | Digital rechargeable CIC for mild-moderate hearing loss (25-55dB), MPO: 75-125dB gain 40-60dB, without Bluetooth compatibility, without tinnitus management, without frequency-lowering. | Offer 1 | Each
293 | RT274-06-113 | Digital non-rechargeable CIC for mild-moderate hearing loss (25-55dB), MPO: 75-125dB gain 40-60dB, without Bluetooth compatibility, without tinnitus management, without frequency-lowering. | Offer 1 | Each
294 | RT274-06-114 | Digital rechargeable CIC for moderate-severe hearing loss (56-70dB), MPO: 116-135dB gain 46-70dB, with Bluetooth compatibility, with tinnitus management, with frequency-lowering. | Offer 1 | Each
295 | RT274-06-115 | Digital non-rechargeable CIC for moderate-severe hearing loss (56-70dB), MPO: 116-135dB gain 46-70dB, with Bluetooth compatibility, with tinnitus management, with frequency-lowering. | Offer 1 | Each
296 | RT274-06-116 | Digital rechargeable CIC for moderate-severe hearing loss (56-70dB), MPO: 116-135dB gain 46-70dB, without Bluetooth compatibility, with tinnitus management, with frequency-lowering. | Offer 1 | Each
297 | RT274-06-117 | Digital non-rechargeable CIC for moderate-severe hearing loss (56-70dB), MPO: 116-135dB gain 46-70dB, without Bluetooth compatibility, with tinnitus management, with frequency-lowering. | Offer 1 | Each
298 | RT274-06-118 | Digital rechargeable CIC for moderate-severe hearing loss (56-70dB), MPO: 116-135dB gain 46-70dB, with Bluetooth compatibility, without tinnitus management, with frequency-lowering. | Offer 1 | Each
299 | RT274-06-119 | Digital non-rechargeable CIC for moderate-severe hearing loss (56-70dB), MPO: 116-135dB gain 46-70dB, with Bluetooth compatibility, without tinnitus management, with frequency-lowering. | Offer 1 | Each
300 | RT274-06-120 | Digital rechargeable CIC for moderate-severe hearing loss (56-70dB), MPO: 116-135dB gain 46-70dB, without Bluetooth compatibility, without tinnitus management, with frequency-lowering. | Offer 1 | Each
301 | RT274-06-121 | Digital non-rechargeable CIC for moderate-severe hearing loss (56-70dB), MPO: 116-135dB gain 46-70dB, without Bluetooth compatibility, without tinnitus management, with frequency-lowering. | Offer 1 | Each
302 | RT274-06-122 | Digital rechargeable CIC for moderate-severe hearing loss (56-70dB), MPO: 116-135dB gain 46-70dB, with Bluetooth compatibility, with tinnitus management, without frequency-lowering. | Offer 1 | Each
303 | RT274-06-123 | Digital non-rechargeable CIC for moderate-severe hearing loss (56-70dB), MPO: 116-135dB gain 46-70dB, with Bluetooth compatibility, with tinnitus management, without frequency-lowering | Offer 1 | Each
304 | RT274-06-124 | Digital rechargeable CIC for moderate-severe hearing loss (56-70dB), MPO: 116-135dB gain 46-70dB, without Bluetooth compatibility, with tinnitus management, without frequency-lowering. | Offer 1 | Each
305 | RT274-06-125 | Digital non-rechargeable CIC for moderate-severe hearing loss (56-70dB), MPO: 116-135dB gain 46-70dB, without Bluetooth compatibility, with tinnitus management, without frequency-lowering. | Offer 1 | Each
306 | RT274-06-126 | Digital rechargeable CIC for moderate-severe hearing loss (56-70dB), MPO: 116-135dB gain 46-70dB, with Bluetooth compatibility, without tinnitus management, without frequency-lowering. | Offer 1 | Each
307 | RT274-06-127 | Digital non-rechargeable CIC for moderate-severe hearing loss (56-70dB), MPO: 116-135dB gain 46-70dB, with Bluetooth compatibility, without tinnitus management, without frequency-lowering. | Offer 1 | Each
308 | RT274-06-128 | Digital rechargeable CIC for moderate-severe hearing loss (56-70dB), MPO: 116-135dB gain 46-70dB, without Bluetooth compatibility, without tinnitus management, without frequency-lowering. | Offer 1 | Each
309 | RT274-06-129 | Digital non-rechargeable CIC for moderate-severe hearing loss (56-70dB), MPO: 116-135dB gain 46-70dB, without Bluetooth compatibility, without tinnitus management, without frequency-lowering. | Offer 1 | Each
310 | RT274-06-130 | Digital rechargeable CIC for severe-profound hearing loss (71-90dB), MPO: MPO: 126-145dB gain 61-85dB, with Bluetooth compatibility, with tinnitus management, with frequency-lowering. | Offer 1 | Each
311 | RT274-06-131 | Digital non-rechargeable CIC for severe-profound hearing loss (71-90dB), MPO: MPO: 126-145dB gain 61-85dB, with Bluetooth compatibility, with tinnitus management, with frequency-lowering. | Offer 1 | Each
312 | RT274-06-132 | Digital rechargeable CIC for severe-profound hearing loss (71-90dB), MPO: MPO: 126-145dB gain 61-85dB, without Bluetooth compatibility, with tinnitus management, with frequency-lowering. | Offer 1 | Each
313 | RT274-06-133 | Digital non-rechargeable CIC for severe-profound hearing loss (71-90dB), MPO: MPO: 126-145dB gain 61-85dB, without Bluetooth compatibility, with tinnitus management, with frequency-lowering. | Offer 1 | Each
314 | RT274-06-134 | Digital rechargeable CIC for severe-profound hearing loss (71-90dB), MPO: MPO: 126-145dB gain 61-85dB, with Bluetooth compatibility, without tinnitus management, with frequency-lowering. | Offer 1 | Each
315 | RT274-06-135 | Digital non-rechargeable CIC for severe-profound hearing loss (71-90dB), MPO: MPO: 126-145dB gain 61-85dB, with Bluetooth compatibility, without tinnitus management, with frequency-lowering. | Offer 1 | Each
316 | RT274-06-136 | Digital rechargeable CIC for severe-profound hearing loss (71-90dB), MPO: MPO: 126-145dB gain 61-85dB, without Bluetooth compatibility, without tinnitus management, with frequency-lowering. | Offer 1 | Each
317 | RT274-06-137 | Digital non-rechargeable CIC for severe-profound hearing loss (71-90dB), MPO: MPO: 126-145dB gain 61-85dB, without Bluetooth compatibility, without tinnitus management, with frequency-lowering. | Offer 1 | Each
318 | RT274-06-138 | Digital rechargeable CIC for severe-profound hearing loss (71-90dB), MPO: MPO: 126-145dB gain 61-85dB, with Bluetooth compatibility, with tinnitus management, without frequency-lowering. | Offer 1 | Each
319 | RT274-06-139 | Digital non-rechargeable CIC for severe-profound hearing loss (71-90dB), MPO: MPO: 126-145dB gain 61-85dB, with Bluetooth compatibility, with tinnitus management, without frequency-lowering. | Offer 1 | Each
320 | RT274-06-140 | Digital rechargeable CIC for severe-profound hearing loss (71-90dB), MPO: MPO: 126-145dB gain 61-85dB, without Bluetooth compatibility, with tinnitus management, without frequency-lowering. | Offer 1 | Each
321 | RT274-06-141 | Digital non-rechargeable CIC for severe-profound hearing loss (71-90dB), MPO: MPO: 126-145dB gain 61-85dB, without Bluetooth compatibility, with tinnitus management, without frequency-lowering. | Offer 1 | Each
322 | RT274-06-142 | Digital rechargeable CIC for severe-profound hearing loss (71-90dB), MPO: MPO: 126-145dB gain 61-85dB, with Bluetooth compatibility, without tinnitus management, without frequency-lowering. | Offer 1 | Each
323 | RT274-06-143 | Digital non-rechargeable CIC for severe-profound hearing loss (71-90dB), MPO: MPO: 126-145dB gain 61-85dB, with Bluetooth compatibility, without tinnitus management, without frequency-lowering. | Offer 1 | Each
324 | RT274-06-144 | Digital rechargeable CIC for severe-profound hearing loss (71-90dB), MPO: MPO: 126-145dB gain 61-85dB, without Bluetooth compatibility, without tinnitus management, without frequency-lowering. | Offer 1 | Each
325 | RT274-06-145 | Digital non-rechargeable CIC for severe-profound hearing loss (71-90dB), MPO: MPO: 126-145dB gain 61-85dB, without Bluetooth compatibility, without tinnitus management, without frequency-lowering | Offer 1 | Each
326 | RT274-06-146 | Wax guards, product specific, small plastic attachment for in-the-ear, in-the-canal and complete-in-canal hearing instruments for the prevention of wax blockages | Offer 1 | Pack of 6
327 | RT274-07-001 | Batteries, dry cell, size AA, card of 4, pack of 5 | Offer 1 | Pack of 5
328 | RT274-07-002 | Batteries, dry cell, size AAA, card of 4, pack of 5 | Offer 1 | Pack of 5
329 | RT274-07-003 | Alkaline 9V battery, IEC type 6lr61, flat type, 9V-500mal | Offer 1 | Each
330 | RT274-07-004 | Batteries, size, 675, each, for hearing aid | Offer 1 | Each
331 | RT274-07-005 | Batteries, size, 13, each, for hearing aid | Offer 1 | Each
332 | RT274-07-006 | Batteries, size, 312, each, for hearing aid | Offer 1 | Each
333 | RT274-07-007 | Batteries, size, 10, each, for hearing aid | Offer 1 | Each
334 | RT274-07-008 | Rechargeable batteries size 13, for hearing aid | Offer 1 | Each
335 | RT274-07-009 | Rechargeable batteries size 675, for hearing aid | Offer 1 | Each
336 | RT274-07-010 | Rechargeable batteries size 10, for hearing aid | Offer 1 | Each
337 | RT274-07-011 | Rechargeable batteries size 312, for hearing aid | Offer 1 | Each
338 | RT274-07-012 | Charger for rechargeable hearing aid batteries | Offer 1 | Each
339 | RT274-07-013 | Tester, premium zinc air batteries, battery sizes 10, 13, 312 , 675, with flat terminals, 1.4V, for hearing aid | Offer 1 | Each
340 | RT274-07-014 | Microphone protection filter for BTEs, pack of 10 | Offer 1 | Pack of 10
341 | RT274-07-015 | Microphone protection filter for ITEs, pack of 10 | Offer 1 | Pack of 10
342 | RT274-07-016 | Microphone protection filter for ITCs, pack of 10 | Offer 1 | Pack of 10
343 | RT274-07-017 | Replacement adult tone hook that connects the earmould to the hearing aid. Must be compatible with awarded BTEs and BTEs in circulation since 2018 | Offer 1 | Each
344 | RT274-07-018 | Replacement adult tone hook with damper that connects the earmould to the hearing aid. Must be compatible with awarded BTEs and BTEs in circulation since 2018 | Offer 1 | Each
345 | RT274-07-019 | Replacement paediatric tone hook that connects the earmould to the hearing aid. Must be compatible with awarded BTEs and BTEs in circulation since 2018 | Offer 1 | Each
346 | RT274-07-020 | Replacement paediatric tone hook with damper that connects the earmould to the hearing aid. Must be compatible with awarded BTEs and BTEs in circulation since 2018 | Offer 1 | Each
347 | RT274-07-021 | Tubing, thin-walled, pre-bent, made of soft pliable transparent plastic, for ear mould 3mm in diameter and 10-12cm in length, pack of 10 | Offer 1 | Pack of 10
348 | RT274-07-022 | Tubing, medium-walled pre-bent, made of soft pliable transparent plastic for ear mould 3.3mm in diameter and 10-12cm in length, pack of 10 | Offer 1 | Pack of 10
349 | RT274-07-023 | Tubing, thick-walled, pre-bent, made of soft pliable transparent plastic, for ear mould 3.6mm in diameter and 10-12cm in length, pack of 10 | Offer 1 | Pack of 10
350 | RT274-07-024 | Glue, clear, strong, instant drying, 5g for re-tubing purposes pack of 10 | Offer 1
351 | RT274-07-025 | Slim tube for BTE hearing aid, pack of 10 | Offer 1 | Pack of 10
352 | RT274-07-026 | Dome to fit slim tube for BTE hearing aid, pack of 10 | Offer 1 | Pack of 10
353 | RT274-07-027 | Earmold, acrylic micro, each, for hearing aid | Offer 1 | Each
354 | RT274-07-028 | Earmold, acrylic full, each, for hearing aid | Offer 1 | Each
355 | RT274-07-029 | Earmold, acrylic skeleton, each, for hearing aid | Offer 1 | Each
356 | RT274-07-030 | Earmold, acrylic open canal, each, for hearing aid | Offer 1 | Each
357 | RT274-07-031 | Earmold, soft silicone, each, for hearing aid | Offer 1 | Each
358 | RT274-07-032 | Retainer, pliable plastic, to secure hearing aid to an adult client`s ear, consisting of: a small circular piece of plastic, approximately 2mm in diameter that attaches to the client`s ear, 2 wider plastic attachments that attach to the hearing aid | Offer 1 | Each
359 | RT274-07-033 | Retainer, pliable plastic, used to secure the hearing aid to a paediatric client`s ear. Should consist of a small circular piece of plastic of approximately 2mm in diameter that attaches to the client`s ear and 2 wider plastic attachments that attach to the hearing aid | Offer 1 | Each
360 | RT274-07-034 | Adult hearing aid care kit, in a durable zip pouch complete with cleaning cup, cleaning capsule, air blower, battery tester, metal clothing clip, retainer and stetoclip | Offer 1 | Each
361 | RT274-07-035 | Paediatric hearing aid care kit, in a durable zip pouch complete with cleaning cup, cleaning capsule, air blower, battery tester and stetoclip | Offer 1 | Each
362 | RT274-07-036 | Paediatric metal clip that attaches to the hearing aid to ensure that the hearing aid does not get lost | Offer 1 | Each
363 | RT274-07-037 | Sachet of silica gel, with saturation indicator, pack of 10 | Offer 1 | Pack of 10
364 | RT274-07-038 | Drying cup, 6 capsules, set, for hearing aids | Offer 1
365 | RT274-07-039 | Capsules, drying, pack of 6, for hearing aids | Offer 1 | Pack of 6
366 | RT274-07-040 | Dry aid, electronic, each, for hearing aids, universal, kit | Offer 1
367 | RT274-07-041 | Capsules, cleaning, dissolvable tablets used for cleaning ear mould pack of 6 | Offer 1 | Pack of 6
368 | RT274-07-042 | Disinfectant cleaner for hearing aid and earmould, cleans, disinfects, deodorises hearing aid and ear mould | Offer 1
369 | RT274-07-043 | Otostops, small, made of soft sponge/foam with cotton thread attachment, 8mm in diameter and 5mm in length (pack of 50) | Offer 1 | Pack of 50
370 | RT274-07-044 | Otostops, medium, made of soft sponge/foam with cotton thread attachment, 10mm in diameter and 1cm in length (pack of 50) | Offer 1 | Pack of 50
371 | RT274-07-045 | Otostops, large, made of soft sponge/foam with cotton thread attachment, 13mm in diameter and 1cm in length (pack of 50) | Offer 1 | Pack of 50
372 | RT274-07-046 | Syringe, 10-15cm in length, double plunge, suitable for making earmould impressions | Offer 1
373 | RT274-07-047 | Syringe, impression, each | Offer 1 | Each
374 | RT274-07-048 | Impression material, silicon based , 800g tub, with hardener, 800g tub, 40 shore | Offer 1 | Each
375 | RT274-07-049 | Impression material, silicon based, 800g tub, with tube of catalyst, 40 shore. Items RT274-07-017 and RT274-07-018 must be awarded as a series. | Offer 1 | Each
376 | RT274-07-050 | Tube, catalyst, 20g, for ear mould impression material. Must be compatible with item RT274-07-017 | Offer 1 | Each
377 | RT274-07-051 | Remote control, for Bluetooth behind the ear hearing aid, each, volume and programme adjustment | Offer 1 | Each
378 | RT274-07-052 | Wireless assistive device that does the following: must connect hearing aid via bluetooth to cellphone, must make adjustments to volume and programmes on wireless hearing aid, must have microphone function for assisting user in noisy environment | Offer 1 | Each
379 | RT274-07-053 | Wax guards, product specific for ITE, ITC and CIC hearing instruments for the prevention of wax blockages , pack of 6 | Offer 1 | Each
380 | RT274-07-054 | Replacement ITE/ITC/CIC casing | Offer 1 | Each
381 | RT274-08-001 | Digital rechargeable RIC for mild-moderate hearing loss (25-55dB), MPO: 75-120dB gain 40-60dB, with Bluetooth compatibility, with tinnitus management, with frequency-lowering. | Offer 1 | Each
382 | RT274-08-002 | Digital non-rechargeable RIC for mild-moderate hearing loss (25-55dB), MPO: 75-120dB gain 40-60dB, with Bluetooth compatibility, with tinnitus management, with frequency-lowering | Offer 1 | Each
383 | RT274-08-003 | Digital rechargeable RIC for mild-moderate hearing loss (25-55dB), MPO: 75-120dB gain 40-60dB, without Bluetooth compatibility, with tinnitus management, with frequency-lowering | Offer 1 | Each
384 | RT274-08-004 | Digital non-rechargeable RIC for mild-moderate hearing loss (25-55dB), MPO: 75-120dB gain 40-60dB, without Bluetooth compatibility, with tinnitus management, with frequency-lowering. | Offer 1 | Each
385 | RT274-08-005 | Digital rechargeable RIC for mild-moderate hearing loss (25-55dB), MPO: 75-120dB gain 40-60dB, with Bluetooth compatibility, without tinnitus management, with frequency-lowering. | Offer 1 | Each
386 | RT274-08-006 | Digital non-rechargeable RIC for mild-moderate hearing loss (25-55dB), MPO: 75-120dB gain 40-60dB, with Bluetooth compatibility, without tinnitus management, with frequency-lowering. | Offer 1 | Each
387 | RT274-08-007 | Digital rechargeable RIC for mild-moderate hearing loss (25-55dB), MPO: 75-120dB gain 40-60dB, without Bluetooth compatibility, without tinnitus management, with frequency-lowering. | Offer 1 | Each
388 | RT274-08-008 | Digital non-rechargeable RIC for mild-moderate hearing loss (25-55dB), MPO: 75-120dB gain 40-60dB, without Bluetooth compatibility, without tinnitus management, with frequency-lowering. | Offer 1 | Each
389 | RT274-08-009 | Digital rechargeable RIC for mild-moderate hearing loss (25-55dB), MPO: 75-120dB gain 40-60dB, with Bluetooth compatibility, with tinnitus management, without frequency-lowering. | Offer 1 | Each
390 | RT274-08-010 | Digital non-rechargeable RIC for mild-moderate hearing loss (25-55dB), MPO: 75-125dB gain 40-60dB, with Bluetooth compatibility, with tinnitus management, without frequency-lowering. | Offer 1 | Each
391 | RT274-08-011 | Digital rechargeable RIC for mild-moderate hearing loss (25-55dB), MPO: 75-120dB gain 40-60dB, without Bluetooth compatibility, with tinnitus management, without frequency-lowering. | Offer 1 | Each
392 | RT274-08-012 | Digital non-rechargeable RIC for mild-moderate hearing loss (25-55dB), MPO: 75-120dB gain 40-60dB, without Bluetooth compatibility, with tinnitus management, without frequency-lowering. | Offer 1 | Each
393 | RT274-08-013 | Digital rechargeable RIC for mild-moderate hearing loss (25-55dB), MPO: 75-120dB gain 40-60dB, with Bluetooth compatibility, without tinnitus management, without frequency-lowering. | Offer 1 | Each
394 | RT274-08-014 | Digital non-rechargeable RIC for mild-moderate hearing loss (25-55dB), MPO: 75-120dB gain 40-60dB, with Bluetooth compatibility, without tinnitus management, without frequency-lowering. | Offer 1 | Each
395 | RT274-08-015 | Digital rechargeable RIC for mild-moderate hearing loss (25-55dB), MPO: 75-120dB gain 40-60dB, without Bluetooth compatibility, without tinnitus management, without frequency-lowering | Offer 1 | Each
396 | RT274-08-016 | Digital non-rechargeable RIC for mild-moderate hearing loss (25-55dB), MPO: 75-120dB gain 40-60dB, without Bluetooth compatibility, without tinnitus management, without frequency-lowering. | Offer 1 | Each
397 | RT274-08-017 | Digital rechargeable RIC for moderate-severe hearing loss (56-70dB), MPO: 116-130dB gain 46-65dB, with Bluetooth compatibility, with tinnitus management, with frequency-lowering. | Offer 1 | Each
398 | RT274-08-018 | Digital non-rechargeable RIC for moderate-severe hearing loss (56-70dB), MPO: 116-130dB gain 46-65dB, with Bluetooth compatibility, with tinnitus management, with frequency-lowering. | Offer 1 | Each
399 | RT274-08-019 | Digital rechargeable RIC for moderate-severe hearing loss (56-70dB), MPO: 116-130dB gain 46-65dB, without Bluetooth compatibility, with tinnitus management, with frequency-lowering. | Offer 1 | Each
400 | RT274-08-020 | Digital non-rechargeable RIC for moderate-severe hearing loss (56-70dB), MPO: 116-130dB gain 46-65dB, without Bluetooth compatibility, with tinnitus management, with frequency-lowering. | Offer 1 | Each
401 | RT274-08-021 | Digital rechargeable RIC for moderate-severe hearing loss (56-70dB), MPO: 116-130dB gain 46-65dB, with Bluetooth compatibility, without tinnitus management, with frequency-lowering. | Offer 1 | Each
402 | RT274-08-022 | Digital non-rechargeable RIC for moderate-severe hearing loss (56-70dB), MPO: 116-130dB gain 46-65dB, with Bluetooth compatibility, without tinnitus management, with frequency-lowering. | Offer 1 | Each
403 | RT274-08-023 | Digital rechargeable RIC for moderate-severe hearing loss (56-70dB), MPO: 116-130dB gain 46-65dB, without Bluetooth compatibility, without tinnitus management, with frequency-lowering. | Offer 1 | Each
404 | RT274-08-024 | Digital non-rechargeable RIC for moderate-severe hearing loss (56-70dB), MPO: 116-130dB gain 46-65dB, without Bluetooth compatibility, without tinnitus management, with frequency-lowering. | Offer 1 | Each
405 | RT274-08-025 | Digital rechargeable RIC for moderate-severe hearing loss (56-70dB), MPO: 116-130dB gain 46-65dB, with Bluetooth compatibility, with tinnitus management, without frequency-lowering. | Offer 1 | Each
406 | RT274-08-026 | Digital non-rechargeable RIC for moderate-severe hearing loss (56-70dB), MPO: 116-130dB gain 46-65dB, with Bluetooth compatibility, with tinnitus management, without frequency-lowering. | Offer 1 | Each
407 | RT274-08-027 | Digital rechargeable RIC for moderate-severe hearing loss (56-70dB), MPO: 116-130dB gain 46-65dB, without Bluetooth compatibility, with tinnitus management, without frequency-lowering. | Offer 1 | Each
408 | RT274-08-028 | Digital non-rechargeable RIC for moderate-severe hearing loss (56-70dB), MPO: 116-130dB gain 46-65dB, without Bluetooth compatibility, with tinnitus management, without frequency-lowering. | Offer 1 | Each
409 | RT274-08-029 | Digital rechargeable RIC for moderate-severe hearing loss (56-70dB), MPO: 116-130dB gain 46-65dB, with Bluetooth compatibility, without tinnitus management, without frequency-lowering. | Offer 1 | Each
410 | RT274-08-030 | Digital non-rechargeable RIC for moderate-severe hearing loss (56-70dB), MPO: 116-130dB gain 46-65dB, with Bluetooth compatibility, without tinnitus management, without frequency-lowering. | Offer 1 | Each
411 | RT274-08-031 | Digital rechargeable RIC for moderate-severe hearing loss (56-70dB), MPO: 116-130dB gain 46-65dB, without Bluetooth compatibility, without tinnitus management, without frequency-lowering. | Offer 1 | Each
412 | RT274-08-032 | Digital non-rechargeable RIC for moderate-severe hearing loss (56-70dB), MPO: 116-130dB gain 46-65dB, without Bluetooth compatibility, without tinnitus management, without frequency-lowering. | Offer 1 | Each
413 | RT274-08-033 | Digital rechargeable RIC for severe-profound hearing loss (71-100dB), MPO: MPO: 120-145dB gain 60-85dB, with Bluetooth compatibility, with tinnitus management, with frequency-lowering. | Offer 1 | Each
414 | RT274-08-034 | Digital non-rechargeable RIC for severe-profound hearing loss (71-100dB), MPO: MPO: 120-145dB gain 60-85dB, with Bluetooth compatibility, with tinnitus management, with frequency-lowering. | Offer 1 | Each
415 | RT274-08-035 | Digital rechargeable RIC for severe-profound hearing loss (71-100dB), MPO: MPO: 120-145dB gain 60-85dB, without Bluetooth compatibility, with tinnitus management, with frequency-lowering. | Offer 1 | Each
416 | RT274-08-036 | Digital non-rechargeable RIC for severe-profound hearing loss (71-100dB), MPO: MPO: 120-145dB gain 60-85dB, without Bluetooth compatibility, with tinnitus management, with frequency-lowering. | Offer 1 | Each
417 | RT274-08-037 | Digital rechargeable RIC for severe-profound hearing loss (71-100dB), MPO: MPO: 120-145dB gain 60-85dB, with Bluetooth compatibility, without tinnitus management, with frequency-lowering. | Offer 1 | Each
418 | RT274-08-038 | Digital non-rechargeable RIC for severe-profound hearing loss (71-100dB), MPO: MPO: 120-145dB gain 60-85dB, with Bluetooth compatibility, without tinnitus management, with frequency-lowering. | Offer 1 | Each
419 | RT274-08-039 | Digital rechargeable RIC for severe-profound hearing loss (71-100dB), MPO: MPO: 120-145dB gain 60-85dB, without Bluetooth compatibility, without tinnitus management, with frequency-lowering. | Offer 1 | Each
420 | RT274-08-040 | Digital non-rechargeable RIC for severe-profound hearing loss (71-100dB), MPO: MPO: 120-145dB gain 60-85dB, without Bluetooth compatibility, without tinnitus management, with frequency-lowering. | Offer 1 | Each
421 | RT274-08-041 | Digital rechargeable RIC for severe-profound hearing loss (71-100dB), MPO: MPO: 120-145dB gain 60-85dB, with Bluetooth compatibility, with tinnitus management, without frequency-lowering. | Offer 1 | Each
422 | RT274-08-042 | Digital non-rechargeable RIC for severe-profound hearing loss (71-100dB), MPO: MPO: 120-145dB gain 60-85dB, with Bluetooth compatibility, with tinnitus management, without frequency-lowering. | Offer 1 | Each
423 | RT274-08-043 | Digital rechargeable RIC for severe-profound hearing loss (71-100dB), MPO: MPO: 120-145dB gain 60-85dB, without Bluetooth compatibility, with tinnitus management, without frequency-lowering. | Offer 1 | Each
424 | RT274-08-044 | Digital non-rechargeable RIC for severe-profound hearing loss (71-100dB), MPO: MPO: 120-145dB gain 60-85dB, without Bluetooth compatibility, with tinnitus management, without frequency-lowering. | Offer 1 | Each
425 | RT274-08-045 | Digital rechargeable RIC for severe-profound hearing loss (71-100dB), MPO: MPO: 120-145dB gain 60-85dB, with Bluetooth compatibility, without tinnitus management, without frequency-lowering. | Offer 1 | Each
426 | RT274-08-046 | Digital non-rechargeable RIC for severe-profound hearing loss (71-100dB), MPO: MPO: 120-145dB gain 60-85dB, with Bluetooth compatibility, without tinnitus management, without frequency-lowering. | Offer 1 | Each
427 | RT274-08-047 | Digital rechargeable RIC for severe-profound hearing loss (71-100dB), MPO: MPO: 120-145dB gain 60-85dB, without Bluetooth compatibility, without tinnitus management, without frequency-lowering. | Offer 1 | Each
428 | RT274-08-048 | Digital non-rechargeable RIC for severe-profound hearing loss (71-100dB), MPO: MPO: 120-145dB gain 60-85dB, without Bluetooth compatibility, without tinnitus management, without frequency-lowering. | Offer 1 | Each
429 | RT274-08-049 | Receiver, 35-45db, for receiver in canal | Offer 1 | Each
430 | RT274-08-050 | Wax guards, product specific, small plastic attachment for RIC hearing instruments for the prevention of wax blockages , | Offer 1 | Each
431 | RT274-08-051 | Replacement RIC earmould | Offer 1 | Each
432 | RT274-08-052 | Replacement RIC receiver and dome | Offer 1 | Each
433 | RT274-09-001 | Hearing aid, cross fitting, for contralateral routing of signal, used to fit asymmetrical hearing losses. Receiver in canal, complete with suitable batteries, earmould, dry aid kit, cord and receiver | Offer 1 | Each
434 | RT274-09-002 | Hearing aid, bi-cross fitting, for contralateral routing of signal, used to fit asymmetrical hearing losses. Receiver in canal, complete with suitable batteries, earmould, dry aid kit, cord and receiver | Offer 1 | Each
435 | RT274-09-003 | RIC hearing aid, cross fitting set, to fit asymmetric hearing loss, complete with suitable batteries, earmould, dry aid kit | Offer 1 | Each
436 | RT274-09-004 | Digital soundfield amplification system, consisting of a wireless transmitter with a microphone worn by the teacher, and an array of speakers which distribute and amplify speech/sound evenly across classroom and improve SNR by at least 10dB. | Offer 1 | Each
437 | RT274-09-005 | Digital FM system complete with microphone, transmitter, ear level receiver. Items RT274-09-005, RT274-09-008 and RT274-09-011 must be awarded as a series | Offer 1 | Each
438 | RT274-09-006 | Digital FM system complete with microphone, transmitter, bodyworn receiver. Items RT274-09-006, RT274-09-010 and RT274-09-011 must be awarded as a series | Offer 1 | Each
439 | RT274-09-007 | Digital FM system complete with microphone, transmitter, neckloop receiver. Items RT274-09-007, RT274-09-009 and RT274-09-011 must be awarded as a series | Offer 1 | Each
440 | RT274-09-008 | Digital FM system ear-level receiver only. Must be compatible with item RT274-09-005. | Offer 1 | Each
441 | RT274-09-009 | Digital FM system neckloop only. Must be compatible with item RT274-09-007. | Offer 1 | Each
442 | RT274-09-010 | Digital FM system bodyworn receiver only. Must be compatible with RT274-09-006 | Offer 1 | Each
443 | RT274-09-011 | Digital FM system transmitter only. Must be compatible with items RT274-09-005 and/or RT274-09-006 and/or RT274-09-007. | Offer 1 | Each
444 | RT274-09-012 | Direct audio input shoe, each, for BTE hearing aid, allowing direct audio input with assistive listening devices, and other audio devices | Offer 1 | Each
445 | RT274-09-013 | Table-top noise generator, with mains and battery power, with built-in timer and at least 6 different tracks, including white noise, rain and ocean | Offer 1 | Each
446 | RT274-10-01 | Cochlear implant system including: one sound processor; one cochlear implant and electrode array; and sound processor kit. Implant and electrode array with a minimum of 12 electrodes with a range of electrode options for different patient anatomy available at no additional cost; ability to conduct telemetry measurements and objective evoked potential measures to assess device functioning and integrity testing to determine the status of potentially non-functional implants; MRI compatibility for 3 Tesla with magnet in situ. Sound processor kit to comprise of a minimum of: one sound processing unit with the option of both omnidirectional and directional microphones, battery modules; battery charger; a range of magnet strength options; coil and coil cable to connect sound processor to coil if required for sound processor, microphone protection covers; an electronic device dry and store kit; an activity/everyday case; tamperproof locks and indicator lights for paediatric patients; datalogging; FM system compatibility. Sound processor to be available in a range of colours and wearing options, including paediatric configurations and to include: automatic auditory scene classifier; water resistance rating of IP57 for temporary emersion and an option of accessories to allow for IP68 for full emersion. Kit to include an FM system or equivalent for listening in challenging listening conditions. Processor must have reliability reporting and manufacturer certification (as per submitted requirements), a 3-year warranty on sound processor, minimum 1-year warranty on parts and maintenance plan to support patients with repairs and consumables. | Offer 1 | Each
447 | RT274-10-02 | Cochlear implant system including: one implant and electrode array; two sound processors and sound processor kits (to provide the patient with redundancy in the event of a faulty or obsolete processor). The implant and electrode array are to have a minimum of 12 electrodes with a range of electrode options for different patient anatomy available at no additional cost; ability to conduct telemetry measurements and objective evoked potential measures to assess device functioning and integrity testing to determine the status of potentially non-functional implants; MRI compatibility for 3 Tesla with magnet in situ. Sound processor kit to comprise of a minimum of: one sound processing unit with the option of both omnidirectional and directional microphones, battery modules; battery charger; a range of magnet strength options; coil and coil cable to connect sound processor to coil if required for sound processor, microphone protection covers; an electronic device dry and store kit; an activity/everyday case; tamperproof locks and indicator lights for paediatric patients; datalogging; FM system compatibility. Sound processor to be available in a range of colours and wearing options, including paediatric configurations and to include: automatic audnd sitory scene classifier; water resistance rating of IP57 for temporary emersion and an option of accessories to allow for IP68 for full emersion. Kit to include an FM system or equivalent for listening in challenging listening conditions. Processor must have reliability reporting and manufacturer certification (as per submitted requirements), a 3-year warranty on sound processor, minimum 1-year warranty on parts and maintenance plan to support patients with repairs and consumables. | Offer 1
448 | RT274-10-03 | Cochlear implant sound processor (one processor only; no kit) compatible with all cochlear implants and electrode arrays awarded on the tender as well as all previous devices issued in South Africa over the past 20 years. Sound processor to include one sound processing unit with the option of both omnidirectional and directional microphones; datalogging; FM system compatibility. Sound processor to be available in a range of colours and wearing options, including paediatric configurations and to include: automatic auditory scene classifier; water resistance rating of IP57 for temporary emersion and an option of accessories to allow for IP68 for full emersion. Processor must have reliability reporting and manufacturer certification (as per submitted requirements), a 3-year warranty on sound processor, minimum 1-year warranty on parts and maintenance plan to support patients with repairs and consumables. | Offer 1 | Each
449 | RT274-10-04 | Cochlear implant sound processors (two processors only; no kits) compatible with all cochlear implants and electrode arrays awarded on the tender as well as all previous devices issued in South Africa over the past 20 years. Sound processors to include two sound processing units (for redundancy in the event of a faulty or obsolete processor) with the option of both omnidirectional and directional microphones; datalogging; FM system compatibility. Sound processor to be available in a range of colours and wearing options, including paediatric configurations and to include: automatic auditory scene classifier; water resistance rating of IP57 for temporary emersion and an option of accessories to allow for IP68 for full emersion. Processors must have reliability reporting and manufacturer certification (as per submitted requirements), a 3-year warranty on sound processor, minimum 1-year warranty on parts and maintenance plan to support patients with repairs and consumables. | Offer 1
450 | RT274-10-05 | Cochlear implant and electrode array with a minimum of 12 electrodes with a range of options for different patient anatomy available at no additional cost; ability to conduct telemetry measurements to assess device functioning, including integrity testing of to determine the status of potentially non-functional implants; objective evoked potential measures; MRI compatibility for 3 Tesla with magnet in situ. Compatible with all sound processors awarded as well as all sound processors issued in South Africa over the past 5 years. | Offer 1 | Each
451 | RT274-10-06 | USB cable connection for programming a cochlear implant, compatible with all sound processors awarded on the tender as well as all previous devices issued within South Africa over the past 5 years | Offer 1 | Each
452 | RT274-10-07 | Programming cable for cochlear implant, compatible with all sound processors awarded on the tender as well as all previous devices issued within South Africa over the past 5 years | Offer 1 | Each
453 | RT274-10-08 | Programming pod compatible with all cochlear implant sound processors awarded on the tender as well as all previous devices issued within South Africa over the past 5 years | Offer 1 | Each
454 | RT274-10-09 | Cochlear implant sound processor rechargeable battery module compatible with all sound processors awarded on the tender as well as all previous devices issued within South Africa over the past 5 years, available in a range of colours including black, brown and beige. | Offer 1 | Each
455 | RT274-10-10 | Cochlear implant sound processor disposable battery holder compatible with all sound processors awarded on the tender as well as all previous devices issued within South Africa over the past 5 years, available in a range of colours including black, brown and beige. | Offer 1 | Each
456 | RT274-10-11 | Cochlear implant sound processor disposable battery rack compatible with all sound processors awarded on the tender as well as all previous devices issued within South Africa over the past 5 years, available in a range of colours including black, brown and beige. | Offer 1 | Each
457 | RT274-10-12 | Cochlear implant coils and cables compatible with all sound processors awarded on the tender as well as all previous devices issued within South Africa over the past 5 years available in a range of colours including black, brown and beige. | Offer 1 | Each
458 | RT274-10-13 | Magnets compatible with all cochlear implant sound processors awarded on the tender as well as all previous devices issued within South Africa over the past 5 years available in a range of strength options for retention and a range of colours including black, brown and beige. | Offer 1 | Each
459 | RT274-10-14 | Magnet insertion tool compatible with all cochlear implant sound processors awarded on the tender as well as all previous devices issued within South Africa over the past 5 years | Offer 1 | Each
460 | RT274-10-15 | Earhooks for retention of cochlear implant sound processor on the ear, compatible with all sound processors awarded on the tender as well as all previous devices issued within South Africa over the past 5 years and in a range of sizes from infant to adult | Offer 1 | Each
461 | RT274-10-16 | Charger for sound processor rechargeable battery compatible with all cochlear implant sound processors awarded on the tender as well as all previous devices issued within South Africa over the past 5 years | Offer 1 | Each
462 | RT274-10-17 | Remote control device to allow for control of cochlear implant sound processor, include volume and programme controls, compatible with all sound processors awarded on the tender as well as all previous devices issued within South Africa over the past 5 years | Offer 1 | Each
463 | RT274-10-18 | Waterproof cover to allow cochlear implant sound processor to be submerged in water with a minimum IP68 rating compatible with all sound processors awarded on the tender as well as all previous devices issued within South Africa over the past 5 years | Offer 1 | Each
464 | RT274-10-19 | Retention device for cochlear implant sound processor - to secure sound processor on the ear to prevent damage to or loss of processor, compatible with all awarded sound processors and those in circulation over the past 5 years | Offer 1 | Each
465 | RT274-10-20 | Retention device for cochlear implant sound processor - to secure sound processor to the head or hair to prevent damage to or loss of processor, compatible with all awarded sound processors and those in circulation over the past 5 years | Offer 1 | Each
466 | RT274-10-21 | Battery charger set with charger and cable compatible with all cochlear implant sound processors awarded on the tender as well as all previous devices issued within South Africa over the past 5 years | Offer 1 | Each
467 | RT274-10-22 | Battery charger compatible with all cochlear implant sound processors awarded on the tender as well as all previous devices issued within South Africa over the past 5 years | Offer 1 | Each
468 | RT274-10-23 | Battery charger cable compatible with all cochlear implant sound processors awarded on the tender as well as all previous devices issued within South Africa over the past 5 years | Offer 1 | Each
469 | RT274-10-24 | Portable battery charging cable in the event of inability to charge via electricity compatible with all cochlear implant sound processors awarded on the tender as well as all previous devices issued within South Africa over the past 5 year | Offer 1 | Each
470 | RT274-10-25 | Power bank for charging cochlear implant sound processor in the event of inability to charge via electricity | Offer 1 | Each
471 | RT274-10-26 | Electronic drying kit for cochlear implant sound processor | Offer 1 | Each
472 | RT274-10-27 | Additional parts required for functioning of all cochlear implant sound processors awarded on the tender as well as all previous devices issued within South Africa over the past 5 years | Offer 1 | Each
473 | RT274-10-28 | Dessicant drying bricks for electronic drying kit for maintenance of cochlear implant sound processors | Offer 1 | Each
474 | RT274-10-29 | Batteries for cochlear implant, 1.4V premium, disposable, pack of 60 batteries - must explicitly state compatibility with cochlear implants | Offer 1 | Each
475 | RT274-10-30 | Microphone cover protectors compatible with all cochlear implant sound processors awarded on the tender as well as all previous devices issued within South Africa over the past 5 years | Offer 1 | Each
476 | RT274-10-31 | Monitor earphones to allow for daily monitoring of microphone and processor functioning for paediatric patients, compatible with all cochlear implant sound processors awarded on the tender as well as all previous devices issued within South Africa over the past 5 years. | Offer 1 | Each
477 | RT274-10-32 | Additional consumables required for functioning of all cochlear implant sound processors awarded on the tender as well as all previous devices issued within South Africa over the past 5 years | Offer 1 | Each
478 | RT274-10-33 | Direct streaming device for listening to the television/media compatible with all cochlear implant sound processors awarded on the tender as well as all previous devices issued within South Africa over the past 5 years | Offer 1 | Each
479 | RT274-10-34 | Direct streaming device for listening to access to a cellphone compatible with all cochlear implant sound processors awarded on the tender as well as all previous devices issued within South Africa over the past 5 years | Offer 1 | Each
480 | RT274-10-35 | Direct streaming device to facilitate ease of listening in difficult listening situations compatible with all cochlear implant sound processors awarded on the tender as well as all previous devices issued within South Africa over the past 5 years | Offer 1 | Each
481 | RT274-10-36 | Device to enable Bluetooth wireless streaming to cochlear implant sound processor where direct streaming is not available in a cochlear implant sound processor | Offer 1 | Each
482 | RT274-10-37 | Repair and service of all cochlear implant sound processors awarded on the tender as well as all previous devices issued within South Africa over the past 5 years | Offer 1 | Each
483 | RT274-10-38 | A Percutaneous bone conduction hearing system: one detachable external digital sound processor, one implantable electromagnetic vibrator/abutment coupled to a titanium screw, one surgical kit; one sound processor kit. The sound processor should be suitable for single sided deafness, conductive and mixed hearing losses with bone conduction thresholds of up to 45 dB (mild to moderate option) and the power option for bone conduction thresholds of up to 65dB. Sound processor kit to comprise of a minimum of: one sound processing unit, one test rod/skull simulator where indicated, pack of 6 appropriately sized disposable batteries/appropriate rechargeable battery where indicated, drying capsule/kit, sound processor cover (optional), an activity/everyday case, tamperproof locks and indicator lights. The sound processor should with have the option of both omnidirectional and directional microphones, have capability to be programmed as a standalone or via an interfaced software, have telephone compatibility, direct audio input, ability to do insitu audiometry/threshold, datalogging and FM system compatibility and streaming compatibility. Sound processor to be available in a range of colours, water resistance rating of IP57 for temporary emersion. Processor must have reliability reporting and manufacturer certification (as per submitted requirements), a 2-year warranty on sound processor and aftercare service must be guaranteed by the manufacturer. | Offer 1
484 | RT274-10-39 | Passive transcutaneous bone conduction hearing system: one detachable external magnetic digital sound processor, one subcutaneous magnet coupled to one or two titanium implants imbedded in the skull, one surgical kit and one sound processor kit. The sound processor should be suitable for single sided deafness, conductive and mixed hearing losses with bone conduction thresholds of up to 45 dB (mild to moderate option) and the power option for bone conduction thresholds of up to 65dB where indicated. Sound processor kit to comprise of a minimum: one sound processing unit, one test rod where indicated, magnetic spacers with variable strength to ensure patient comfort and secure fit, pack of 6 appropriately sized disposable batteries/appropriate rechargeable battery where indicated, drying capsule/kit, sound processor cover (optional), an activity/everyday case, tamperproof locks and indicator lights. The sound processor should with have the option of both omnidirectional and directional microphones, have capability to be programmed as a standalone or via an interfaced software, have telephone compatibility, ability to do insitu audiometry/threshold, datalogging capabilities, FM system compatibility and streaming capabilities. Sound processor to be available in a range of colours, water resistance rating of IP57 for temporary emersion. Processor must have reliability reporting and manufacturer certification (as per submitted requirements), a 2-year warranty on sound processor and aftercare service must be guaranteed by the manufacturer. | Offer 1
485 | RT274-10-40 | Active transcutaneous bone conduction hearing system: one external magnetic digital sound processor, one subcutaneous receiver stimulator, an active component capable of providing direct vibrational energy (via a floating mass transducer or piezoelectric motor), one surgical kit and one sound processor kit. The sound processor should be suitable for single sided deafness, conductive and mixed hearing losses with bone conduction thresholds of up to 45 dB (mild to moderate option) and the power option for bone conduction thresholds of up to 65dB where indicated. Sound processor kit to comprise of a minimum: one sound processing unit, one test rod where indicated, magnetic spacers with variable strength to ensure patient comfort and secure fit, pack of 6 appropriately sized disposable batteries/appropriate rechargeable battery where indicated, drying capsule/kit, sound processor cover (optional), an activity/everyday case, tamperproof locks and indicator lights. The sound processor should with have the option of both omnidirectional and directional microphones, have capability to be programmed as a standalone or via an interfaced software, have telephone compatibility, ability to do insitu audiometry/threshold, datalogging capabilities, FM system compatibility and/streaming capabilities. Sound processor to be available in a range of colours, water resistance rating of IP57 for temporary emersion. Processor must have reliability reporting and manufacturer certification (as per submitted requirements), a 2-year warranty on sound processor and aftercare service must be guaranteed by the manufacturer. | Offer 1
486 | RT274-10-41 | Digital bone conduction sound processor (one processor only; no surgical kit) compatible with all percutaneous bone conduction implants awarded on the tender as well as all previous devices issued in South Africa over the past 20 years. Sound processor unit to comprise of a minimum: one sound processing unit, pack of 6 appropriately sized disposable batteries/appropriate rechargeable battery where indicated, drying capsule/kit, sound processor cover (optional), an activity/everyday case, tamperproof locks and indicator lights. Processor must have reliability reporting and manufacturer certification (as per submitted requirements), a 2-year warranty on sound processor and aftercare service must be guaranteed by the manufacturer. | Offer 1
487 | RT274-10-42 | Magnetic detachable digital bone conduction sound processor (one processor only; no surgical kit) compatible with all passive transcutaneous bone conduction hearing implants awarded on the tender as well as all previous devices issued in South Africa over the past 10 years. Sound processor unit to comprise of a minimum: one sound processing unit, pack of 6 appropriately sized disposable batteries/appropriate rechargeable battery where indicated, drying capsule/kit, sound processor cover (optional), an activity/everyday case, tamperproof locks and indicator lights. Processor must have reliability reporting and manufacturer certification (as per submitted requirements), a 2-year warranty on sound processor and aftercare service must be guaranteed by the manufacturer. | Offer 1
488 | RT274-10-43 | Magnetic digital bone conduction sound processor (one processor only; no surgical kit) compatible with all active transcutaneous bone conduction hearing implants awarded on the tender as well as all previous devices issued in South Africa over the past 10 years. Sound processor unit to comprise of a minimum: one sound processing unit, pack of 6 appropriately sized disposable batteries/appropriate rechargeable battery where indicated, drying capsule/kit, sound processor cover (optional), an activity/everyday case, tamperproof locks and indicator lights. Processor must have reliability reporting and manufacturer certification (as per submitted requirements), a 2-year warranty on sound processor and aftercare service must be guaranteed by the manufacturer. | Offer 1
489 | RT274-10-44 | An percutaneous bone conduction implant and surgical kit only (excluding the sound processor): consisting of one titanium abutment that penetrates the skin and connected to an osseo-integrated implant fixture in the temporal bone, one surgical kit. Suitable for mild to moderate conductive and mixed hearing losses with bone conduction thresholds of up to 45 dBHL – 65dBHL as well as single sided deafness. Should be compatible with all sound processors awarded as well as all sound processors issues in South Africa over the past 5 years. | Offer 1
490 | RT274-10-45 | Passive transcutaneous bone conduction implant and surgical kit only (excluding the sound processor): one subcutaneous magnet, with coupling abilities to one or two titanium implants imbedded in the skull, one surgical kit. Suitable for mild to moderate conductive and mixed hearing losses with bone conduction thresholds of up to 45 dBHL – 65dBHL as well as single sided deafness. should be compatible with all sound processors awarded as well as all sound processors issues in South Africa over the past 5 years. | Offer 1
491 | RT274-10-46 | Active transcutaneous bone conduction implant and surgical kit only (excluding the sound processor): one subcutaneous magnetic receiver stimulator, an active component capable of providing direct vibrational energy (via a floating mass transducer or piezoelectric motor), one surgical kit. Suitable for mild to moderate conductive and mixed hearing losses with bone conduction thresholds of up to 45 dBHL – 65dBHL as well as single sided deafness. should be compatible with all sound processors awarded as well as all sound processors issues in South Africa over the past 5 years. | Offer 1
492 | RT274-10-47 | Surgical Vibrating ossicular prosthesis/ active middle ear implant system: one external sound processor, one vibrating ossicular or round window prosthesis surgically implantable in the middle ear with a wire that connects to, one partially implantable receptor and transducer capable of providing piezoelectric, electromagnetic or electromechanical vibrational energy. Suitable for the treatment of certain sensorineural hearing losses, conductive hearing losses and mixed hearing losses. Sound processor kit to comprise of a minimum: one sound processing unit, one test rod where indicated, magnetic spacers with variable strength to ensure patient comfort and secure fit, pack of 6 appropriately sized disposable batteries/appropriate rechargeable battery where indicated, drying capsule/kit, sound processor cover (optional), an activity/everyday case, tamperproof locks and/or indicator lights. The sound processor should with have the option of both omnidirectional and directional microphones, have capability to be programmed as a standalone or via an interfaced software, have telephone compatibility, ability to do insitu audiometry/threshold, datalogging capabilities, FM system compatibility and/streaming capabilities. Sound processor to be available in a range of colours, water resistance rating of IP57 for temporary emersion. Processor must have reliability reporting and manufacturer certification (as per submitted requirements), a 2-year warranty on sound processor and aftercare service must be guaranteed by the manufacturer. | Offer 1
493 | RT274-10-48 | Surgical Vibrating ossicular prosthesis/ active middle ear implant only with surgical kit (excluding the sound processor): one vibrating ossicular or round window prosthesis surgically implantable in the middle ear with a wire that connects to, one partially implantable receptor and transducer capable of providing piezoelectric, electromagnetic or electromechanical vibrational energy, one surgical kit. Suitable for the treatment of certain sensorineural hearing losses, conductive hearing losses and mixed hearing losses. should be compatible with all sound processors awarded as well as all sound processors issues in South Africa over the past 5 years. | Offer 1
494 | RT274-10-49 | Magnetic digital sound processor (one processor only; no surgical kit) compatible with all surgical vibrating ossicular/round window prosthesis/active middle ear implants awarded on the tender as well as all previous devices issued in South Africa over the past 5 years. Sound processor unit to comprise of a minimum: one sound processing unit, pack of 6 appropriately sized disposable batteries/appropriate rechargeable battery where indicated, drying capsule/kit, sound processor cover (optional), an activity/everyday case, tamperproof locks and/or indicator lights. Processor must have a reliability reporting and manufacturer certification (as per submitted requirements), a 2 year warranty on sound processor and aftercare service must be guaranteed by the manufacturer. | Offer 1
495 | RT274-10-50 | Direct streaming device for listening to the television/media compatible with all implantable bone conduction implants sound processors awarded on the tender as well as all previous devices issued within South Africa over the past 5 years | Offer 1
496 | RT274-10-51 | Direct streaming device for listening to access to a cellphone compatible with all implantable bone conduction implants sound processors awarded on the tender as well as all previous devices issued within South Africa over the past 5 years | Offer 1
497 | RT274-10-52 | Direct streaming device to facilitate ease of listening in difficult listening situations compatible with all implantable bone conduction implants sound processors awarded on the tender as well as all previous devices issued within South Africa over the past 5 years | Offer 1
498 | RT274-10-53 | Device to enable Bluetooth wireless streaming to implantable bone conduction implants sound processor where direct streaming is not available in a implantable bone conduction implants sound processor | Offer 1
499 | RT274-10-54 | Repair and service of all implantable bone conduction implants sound processors awarded on the tender as well as all previous devices issued within South Africa over the past 5 years | Offer 1
500 | RT274-11-001 | Books on clinical management of children with cochlear implants: Latest version of Clinical Resource book including the following content: - Candidacy - Multidisciplinary team intervention - Audiological management - Speech, language and literacy development | Offer 1 | Each
501 | RT274-11-002 | Books on auditory-verbal therapy for hearing impaired children: Latest version of Clinical Resource book Including: Strategies for the development of listening skills in hearing impaired children Strategies for the development of speech in hearing impaired children | Offer 1 | Each
502 | RT274-11-003 | Books on Language development in hearing impaired children: Latest version of Clinical Resource book Including the following content: - Strategies for developing language, reading and writing skills in Deaf and hard of hearing children | Offer 1 | Each
503 | RT274-11-004 | Books on Aural rehabilitation: Overview of aural rehabilitation - Considerations for the hearing impaired - Decision making on assistive devices for hearing loss - Techniques/strategies in speech and language development | Offer 1 | Each
504 | RT274-11-005 | Books on Hearing losses and associated conditions: Latest version of Clinical Resource book Including: - Types of hearing losses - Aetiology - Syndromes - Genetic factors - Clinical presentation - Management (treatment) | Offer 1 | Each
505 | RT274-11-006 | Audiology assessments: Latest version of Clinical Resource book Including: - Basic audiometry testing - Paediatric testing - Speech audiometry - Site of lesion tests - Tympanometry testing - Electrophysiology | Offer 1 | Each
506 | RT274-11-007 | South African Sign language dictionary: Dictionary with the following - Sign and words - Recorded verson of signs on electronic medium | Offer 1 | Each
507 | RT274-12-001 | Calibration Service of existing equipment in clinical facilities (limited to audiology equipments). Calibration of each model of equipment supplied to the state since 2013, the calibration cost must be inclusive of tranvelling cost, and any other cost to enable the service provider to provide the service. Bidders must provide a list of equipment that they offer calibration which have been supplied to the state already. The list must be provided on the pricing schedule (Relevant sheet) provided | Offer 1 | As per list to be provided by bidder
508 | RT274-13-002 | Spares parts for repairs on equipments - The bidders must provide a list of spare parts on the pricing schedule (relevant sheet) provided. This are the spare parts which will be procured by institutions to repair existing and new equipments. | Offer 1 | As per list to be provided by bidder

## Pricing Schedule for Calibratio
LINE # | ITEM NUMBER | EQUIPMENT DESCRIPTION | BRAND NAME | MODEL NUMBER | MODEL YEAR | NAME OF MANUFACTURE | COUNTRY OF ORIGIN | LEAD TIME
1 | Product Code
2
3
4
5
6
7
8
9
10
11
12
13
14
15
16
17
18
19
20
21
22
23
24
25
26
27
28
29
30
31

## Pricing Schedule for Repairs
LINE # | ITEM NUMBER | ITEM DESCRIPTION | MODEL NUMBER | MODEL YEAR | NAME OF MANUFACTURE | COUNTRY OF ORIGIN | LEAD TIME
1 | Product Code
2
3
4
5
6
7
8
9
10
11
12
13
14
15
16
17
18
19
20
21
22
23
24
25
26
27
28
29
30
31