## Detail Specifications
LINE # | ITEM NUMBER | SUB-CATEGORY DESCRIPTION | ITEM DESCRIPTION
1 | RT275-01-001 | Low-Tech Communication and Literacy | Fabric literacy strip Single strip for top or bottom use. Single 10cm x 24cm fabric literacy strip. Single 4" Fabric Literacy strip.
2 | RT275-01-002 | Low-Tech Communication and Literacy | Fabric literacy flip strip Allows simple addition of pictures/words and ability to flip to create a new sentence each time. Pages made of loop fabric Total of 16 pages. Dimension: Individual pages must be 10cm x 10cm and the entire size of the strip being 4cm x 10cm Pages grouped in sets of 4, aligned horizontally. Pages affixed to plastic stabilisation board with metal rings
3 | RT275-01-003 | Low-Tech Communication and Literacy | Clear literacy flip strip Allows simple addition of pictures/words and ability to flip to create a new sentence each time. Total of 16 pages. Pages made of clear pocket pages into which to slip pictures Dimension: Individual pages are 10cm x 10cm and the entire size of the strip is 4cm x 10cm. Pages grouped in sets of 4, aligned horizontally. Pages affixed to plastic stabilisation board with metal rings.
4 | RT275-01-004 | Low-Tech Communication and Literacy | Portable waist communicator Equipped with 6 fabric pages measuring 10cm x 22cm Back of cover must have loops to fit any belt up to 5cm wide (excluding belt). Zipper on outside cover
5 | RT275-01-005 | Low-Tech Communication and Literacy | Waist communication apron Contains pockets to store unused pictures until needed. Minimum of 2 pockets. Fabric must allow male (hook-side) velcro to stick onto it. Dimensions: 45cm x 35cm Packaging: single unit
6 | RT275-01-006 | Low-Tech Communication and Literacy | Wrap-around activity communication mat Female loop fabric measures approx. 37cm on sides and 58cm along the top. Width of strips: 10cm. Both sides must be usable. Must be bundled with removable clear cover that covers both sides
7 | RT275-01-007 | Low-Tech Communication and Literacy | Portable display board Approx. 122cm x 41cm (when open) board with webbing carrying handle and closure strap. One side of the board must allow for male hook velcro to stick on it, the other side of the board must be durable fabric
8 | RT275-01-008 | Low-Tech Communication and Literacy | Portable object communication unit Portable unit with adjustable webbing handle that accommodates objects up to 15cm in height. Measures 76cm x 40cm open and 30cm x 40cm x 15cm closed. Fabric surface inside and nylon material on the outside, filled with plastic board for stiffened surface.
9 | RT275-01-009 | Low-Tech Communication and Literacy | Horizontal/vertical picture schedule Symbol display accessory that allows placement of pictures or items for the activities to be completed in the picture pocket on the right/bottom. Dimensions: 50 x 55 x 10cm
10 | RT275-01-010 | Low-Tech Communication and Literacy | Eye-gaze frame with stand Durable transperant light weight (e.g perspex or plexiglass frame and stand with detachable base to use on a desk or be hand-held.
11 | RT275-01-011 | Low-Tech Communication and Literacy | Velcro fastener Dot: Male 10-15mm Self-adhesive male hook fastener dots, 10-15mm diameter, 25m roll
12 | RT275-01-012 | Low-Tech Communication and Literacy | Velcro fastener Dot: Female 10-15mm Self-adhesive female loop fastener dots, 10-15mm diameter, 25m roll
13 | RT275-01-013 | Low-Tech Communication and Literacy | Velcro fastener Strip Female 20mm Self-adhesive male hook fastener strips, 20mm width, 25m roll
14 | RT275-01-014 | Low-Tech Communication and Literacy | Velcro fastener Strip Female 20mm Self-adhesive male hook fastener strips, 20mm width, 25m roll
15 | RT275-01-015 | Low-Tech Communication and Literacy | Robotic tool for learners with Autism (Avatar) Robot designed to model human expressions and emotions, with digitised symbol display, and built in comprehensive curriculum targeted at social-emotional skills for learners with autism
16 | RT275-02-001 | Communication Devices | Voice output device, single message Voice output communication aid activated by touch. Supports recording and re-recording of single messages. Removable clear cover for inserting symbols or photographs. Rechargeable. Available in sets.
17 | RT275-02-002 | Communication Devices | Voice output device, single message, audio input Voice output communication aid activated by touch. Supports recording and re-recording of single messages. Removable clear cover for inserting symbols or photographs. Includes audio-in port to record audio directly from a CD player or computer. Rechargeable. Available in sets.
18 | RT275-02-003 | Communication Devices | Voice output device, 100 messages Voice output communication aid activated by touch. 25 large activation surfaces. Pre-recording of up to 25 messages across 5 levels. Up to 100 messages can be stored. Built-in overlay storage. Record and level lock, volume control. Battery-operated. 3cm x 3cm recording keys.
19 | RT275-02-004 | Communication Devices | Voice output device, 8 minutes, large Large Voice output communication aid activated by touch. Recording time of up to 8 minutes. Battery-operated.
20 | RT275-02-005 | Communication Devices | Voice output device, 8 minutes, small Small voice output communication aid activated by touch. Recording time of up to 8 minutes. Battery-operated.
21 | RT275-02-006 | Communication Devices | Voice output device, 3 messages, large Large voice output communication aid with two modes - single message and tap to talk. 1 48-second message or 3 24-second messages. Battery-operated.
22 | RT275-02-007 | Communication Devices | Voice output device, 40 messages, large Large voice output communication aid with two modes - single message and tap to talk. 10 levels, 40 message capacity, LCD readout. Battery-operated.
23 | RT275-02-008 | Communication Devices | Voice output device, 45 messages, large Voice output communication aid, 12 large touch activation surfaces. Recording of up to 45 messages across 5 levels.
24 | RT275-02-009 | Communication Devices | Voice output device, 45 messages, large Voice output communication aid activated by touch. 12 large activation surfaces. Recording of up to 12 messages (3 of which remain constant) across 5 levels. Up to 45 messages can be stored. built-in overlay storage. Record and level lock, volume control. Battery-operated. 5.5cm x 4.25cm recording keys.
25 | RT275-02-010 | Communication Devices | Voice output device, 4 messages, small Small voice output communication aid activated by touch. Lightweight with powerful, high-quality speaker; Supports recording of up to 4 messages; buttons measure 2.5x2.5cm.
26 | RT275-02-011 | Communication Devices | Voice output device, 9 messages, wearable Wearable voice output communication aid activated by touch. 3 programmable message buttons; Up to 9 messages can be stored in at least 2 levels. Rechargeable
27 | RT275-02-012 | Communication Devices | Voice output device, 4 minutes, large Large voice output communication aid activated by touch. Recording of multiple messages in sequence for up to 4 minutes. Battery-operated.
28 | RT275-02-013 | Communication Devices | Voice output device, 4 minutes, small Small voice output communication aid activated by touch. Recording of multiple messages in sequence for up to 4 minutes. Battery-operated.
29 | RT275-02-014 | Communication Devices | Voice output device, 2 messages Voice output communication aid activated by touch. 2 or 4 activation surfaces that allow for dual message recording. Recording a number of choices. Battery-operated.
30 | RT275-02-015 | Communication Devices | Voice output device, 64 messages, switch-activated Voice output communication aid activated by a switch. Choice to record 1, 2, 4 or 8 messages for up to 16 minutes, or a total of 64 messages. Overlays (keyguard) for 1, 2, 4 and 8 messages. Step-ahead jack for single switch sequencial message access. 8 input jacks for single switch access to any specific location. Battery-operated.
31 | RT275-02-016 | Communication Devices | Voice output device, 163 messages Voice output communication aid activated by touch. 35 activation surfaces. Pre-recording of up to 35 messages (3 of which remain constant) across at least 5 levels. Up to 163 messages can be stored. Built-in overlay storage. Record and level lock, volume control. Battery-operated.
32 | RT275-02-017 | Communication Devices | Voice output device, 160 messages, scanning Voice output communication aid activated both via touch and switch if required. 32 activation surfaces and recording across 5 levels. Choice to scan 2, 4, 8, 16 or 32 messages with scan level feature. Option for visual or auditory cueing or direct access. Scanning pattern options include step, linear or group, row/column and speed. Can play multiple messages in sequence. Headphone access available. Built-in overlays storage. Step-ahead jack for single switch sequential message access. Record and level lock, volume control. Battery-operated.
33 | RT275-02-018 | Communication Devices | Voice output device, 6 minutes Voice output communication aid activated by touch. 7 message locations and 5 levels, recording time of 6 minutes. Includes carry handle. Battery-operated.
34 | RT275-02-019 | Communication Devices | Voice output device, 20 minutes Voice output communication aid activated by touch. 23 message locations and 5 levels, recording time of 20 minutes, Includes carry handle. Battery-operated.
35 | RT275-02-020 | Communication Devices | Induction Loop System, Fixed Integrates with built-in receiver (telecoil) in hearing aids and cochlear implants to transmit audio to user via induction loop installed in room. Integrates with PA system. Supports a range of room sizes, from small rooms up to large halls. Must conform to IEC 60118-4:2006.
36 | RT275-02-021 | Communication Devices | Induction Loop System, Portable Portable unit integrates with built-in receiver (telecoil) in hearing aids and cochlear implants to transmit audio to user via induction loop (integrated or external). Supports counters and service desks. Must conform to IEC 60118-4:2006. Includes carry case.
37 | RT275-02-022 | Communication Devices | Induction Loop Receiver Receiver unit with built-in telecoil, for use with installed induction loops. For users whose hearing aids do not have a telecoil or who don't use hearing aids. Must be used with headphones with safe limit.
38 | RT275-02-023 | Communication Devices | Portable Voice Amplifier Compact voice amplification system that amplifies user's faint/quiet/whispered voice. Various microphone options depending on user. 8+ hours of talk time with built-in rechargeable battery.
39 | RT275-02-024 | Communication Devices | Portable Voice Amplifier - Waist worn Worn by the user to amplify their voice. Suitable for one-to-one speech or group use. Worn around the waist with included secure waist band. Includes headset with boom microphone. Rechargeable battery.
40 | RT275-02-025 | Communication Devices | Portable Voice Amplifier - Wireless Wireless microphone system consisting of a lightweight lapel microphone transmitter and a mini receiver. The receiver picks up speech from the transmitter and can be used as an audio source for a wide variety of systems such as speaker amplifiers, communication equipment and personal listeners. Speakers and amplifier not included.
41 | RT275-02-026 | Communication Devices | Dual Headset Specialised dual headset for switchboard use. Separate inputs for PC and phone system; Various cabling configurations
42 | RT275-02-027 | Communication Devices | Dynamic Communication Device Specialised portable dedicated communication device with software. Generates synthesised or recorded speech by selecting words or symbols on the screen Customisable interface, compatible with various access methods. Mounting available.
43 | RT275-02-028 | Communication Devices | Wearable Warning System Customisable system of transmitters and receivers to alert user about various events in their environment, such as a bell or smoke alarm; Choice of transmitters and alert types based on required function (e.g. shaker, flashing light, extra loud noise)
44 | RT275-02-029 | Communication Devices | Timer / Stopwatch Measure amount of time available or elapsed; Used for time management or scheduling; Hours, minutes and seconds with alarm function
45 | RT275-02-030 | Communication Devices | Inclusive Interactive Music System Music system that can be accessed in a variety of ways to promote inclusion; Make music and sounds via methods such as touch, movement, switch or eye gaze. Can be mounted.
46 | RT275-03-001 | Assistive Accessories | Small assistive switch Small (<75mm), durable, quality single switch compatible with assistive communication device via wired connection. Activation by pressure anywhere on top surface. Must provide audible and tactile feedback. Must have mechanisms to prevent slipping and allow for secure mounting. Available in several bright colours.
47 | RT275-03-002 | Assistive Accessories | Small assistive switch, wireless Small (<75mm), durable, quality single switch compatible with assistive communication device via wireless connection. Activation by pressure anywhere on top surface. Must provide audible and tactile feedback. Must have mechanisms to prevent slipping and allow for secure mounting. Available in several bright colours.
48 | RT275-03-003 | Assistive Accessories | Large assistive switch Large (75mm+), durable, quality single switch compatible with assistive communication device via wired connection. Activation by pressure anywhere on top surface. Assists people with moderate to severe upper extremity and motor disabilities. Must provide audible and tactile feedback. Must have mechanisms to prevent slipping and allow for secure mounting. Available in several bright colours.
49 | RT275-03-004 | Assistive Accessories | Large assistive switch, wireless Large (75mm+), durable, quality single switch compatible with assistive communication device via wired connection. Activation by pressure anywhere on top surface. Assists people with moderate to severe upper extremity and motor disabilities. Must provide audible and tactile feedback. Must have mechanisms to prevent slipping and allow for secure mounting. Available in several bright colours
50 | RT275-03-005 | Assistive Accessories | Head/cheek assistive switch Durable single switch compatible with assistive communication device, enables head and/or cheek activation Must be smooth and soft Must have a reliable mechanism for fastening to a pillow or wheelchair cushion Must provide audible and tactile feedback.
51 | RT275-03-006 | Assistive Accessories | Light touch assistive switch Durable single switch, activated through light touch; Must provide audible and tactile feedback; Activating force 10-11g; Requires compatible mounting plate
52 | RT275-03-007 | Assistive Accessories | Proximity assistive switch, 10mm Compact, sensitive, wired proximity sensor switch that is activated when a body part is within 10mm of the switch top. Must have visual and auditory feedback. Magnetic mounting. Packaging: Each
53 | RT275-03-008 | Assistive Accessories | Proximity assistive switch, 25mm Large, sensitive, wired proximity sensor switch that is activated when a body part is within 25mm of the switch top, aproximately 90mm wide activation surface. Must have visual and auditory feedback. Magnetic mounting.
54 | RT275-03-009 | Assistive Accessories | Joystick assistive switch, 4-way Durable joystick compatible with assistive access device to control Android and iOS devices. Supports 4 switches to be activated by moving the joystick in 4 directions. Must have at least two tips: flexible rubber and hard plastic.
55 | RT275-03-010 | Assistive Accessories | Joystick assistive switch, 5-way Durable joystick compatible with optional mounting plate, allows 4 switches to be activated via 4 joystick directions. 5th switch is activated by pressing down the joystick. Must have at least two tips: flexible rubber and hard plastic. Must provide audible and tactile feedback; Mounting via compatible mounting plate; requires optional 9-pin serial interface
56 | RT275-03-011 | Assistive Accessories | Squeeze/pinch assistive switch Durable switch compatible with optional mounting plate. Tube shaped switch activated by squeeze and/or pinch. Allows user approximately 2 seconds to release grip in order to prevent switch latching. Soft to touch. Tactile feedback. Mounting via compatible mounting plate
57 | RT275-03-012 | Assistive Accessories | Bend assistive switch Durable switch compatible with optional mounting plate. A switch activated by bending in either direction. Soft to touch. Moisture resistant. Tactile feedback. Mounting via compatible mounting plate
58 | RT275-03-013 | Assistive Accessories | Infra-Red assistive switch Durable infra-red switch compatible with assistive communication device. Momentary contact optical switch that can be controlled with eye blink, eyebrow movement, finger movement, head movement, or facial muscle movement. Sound feedback that works by detecting a beam of pulsed infrared light. Good for users who have difficulty activating push button switches
59 | RT275-03-014 | Assistive Accessories | Low-profile assistive switch Durable single switch compatible with assistive communication device. Activation of switch done by pressure application anywhere on switch’s top surface. Switch Low profile surface angled towards the user with limited movement to allow for easier pressing.
60 | RT275-03-015 | Assistive Accessories | Pneumatic assistive switch Durable switch compatible with assistive communication device. Activated through puff of the lips or sip. The supplier to indicate if there is an internal saliva trap for cleaning. Mounting required.
61 | RT275-03-016 | Assistive Accessories | Bite assistive switch Switch with silicone cover activated via bite action
62 | RT275-03-017 | Assistive Accessories | USB adapter for assistive switch Adapter cable to plug switches into a computer through the USB port Supports connection of one or two switches Standard 3.5mm switch inputs
63 | RT275-03-018 | Assistive Accessories | Advanced USB adapter for assistive switch Adapter cable to plug switches into a computer through the USB port Supports connection of one or two switches; includes configuration software Standard 3.5mm switch inputs
64 | RT275-03-019 | Assistive Accessories | Dual wireless assistive switch interface Bluetooth switch interface supporting at least 2 switches. Can synchronise with smartphone or tablet to activate switch scanning mode
65 | RT275-03-020 | Assistive Accessories | Switch adapter for battery-operated device (AA) Adapter cable that can convert a battery-operated toy or appliance into a switch activated device. Compatible with AA battery size
66 | RT275-03-021 | Assistive Accessories | Switch adapter for battery-operated device (C/D) Adapter cable that can convert a Battery-operated toy or appliance into a switch activated device. Compatible with C/D battery size
67 | RT275-03-022 | Assistive Accessories | USB adapter for assistive switch - 6-way An adapter cable to connect switches into a computer through the USB port. Supports connection of up to 6 switches. With configurable software.
68 | RT275-03-023 | Assistive Accessories | USB Interface Kit for assistive switch - 6-way wireless An adapter cable to connect switches to a computer through the USB port type receiver and wireless transmitter. Supports connection of up to 6 switches and computer accessories via wireless receiver and transmitter kit
69 | RT275-03-024 | Assistive Accessories | 2-way adapter for Assistive switch Switch accessory that allows two separate switch-adapted devices to be activated simultaneously
70 | RT275-03-025 | Assistive Accessories | 2-way adapter for Assistive switch, wireless Switch accessory that allows two separate switch-adapted devices to be activated simultaneously. Wireless capabilities with wireless switches
71 | RT275-03-026 | Assistive Accessories | Switch Interface Box Device that allows the connection of a standard 3.5 mm switch cable to a computer or device requiring a USB connection. Available in various models, for connection of 2 to 12 switches; Some replicate basic keyboard commands such as Spacebar and Enter; Some are programmable, offering more functionality
72 | RT275-03-027 | Assistive Accessories | Wireless Transmitter for Switch Wireless transmitter to make a wired switch wireless. Connection from wired switch into the transmitter; Requires wireless receiver for complete solution
73 | RT275-03-028 | Assistive Accessories | Wireless Receiver USB device to receive wireless signal from a compatible assistive device such as switch, joystick, rollerball, keyboard. Designed for compatible devices; Operating range up to 10m; Compatible with Windows and Mac
74 | RT275-03-029 | Assistive Accessories | Assistive switch-adapted toy, plastic On/off type toy adapted for use with AAC switch. Battery-operated. Durable, strong plastic. Toy to facilitate development of prelinguistic skills as a precursor to further use of devices
75 | RT275-03-030 | Assistive Accessories | Assistive switch-adapted toy, plush On/off type toy adapted for use with AAC switch. Battery-operated. Durable plush toy. Toy to facilitate development of prelinguistic skills as a precursor to further use of devices
76 | RT275-03-031 | Assistive Accessories | 2-way adapter for Assistive switch - 2 toys Switch accessory that allows two separate switch-adapted devices to be activated simultaneously. Wireless capabilities with wireless switches; switches available separately.
77 | RT275-03-032 | Assistive Accessories | 2-way wireless adapter for Assistive switch toys - 1 toy Allows remote control of single switch toy. Attaches to toy with cable provided. Requires a compatible wireless switch, or any standard switch connected to a compatible receiver. Toys, switches and receiver available separately.
78 | RT275-03-033 | Assistive Accessories | 2-way wireless adapter for Assistive switch toys - 1 toy, with timer and latch Allows remote control of single switch toy. Attaches to toy with cable provided. Requires a compatible wireless switch, or any standard switch connected to a compatible receiver. Toys, switches and receiver available separately.
79 | RT275-03-034 | Assistive Accessories | 2-way wireless adapter with proximity switch Proximity switch transmitter with external additional switch jack; and receiver with switch latch and timer functions, and 9m range.
80 | RT275-03-035 | Assistive Accessories | Switch Toy Control via computer Connects up to 4 switch adapted toys, with customisble software, for access with any input device such as eye control, switch, touch. Computer and switch toys available separately.
81 | RT275-03-036 | Assistive Accessories | Switch adapted Mains Appliance Control Controller to allow two mains appliances to be remotely operated using wired or wireless switches. Must provide galvanic electrical isolation between user and appliance to maximise user safety. Each appliance may be controlled separately in a number of different ways, for example latching outputs and timed outputs. Preferably also allows users to co-operate to turn on appliances. Unit must have either South African sockets or be accompanied by SA travel adapters
82 | RT275-03-037 | Assistive Accessories | Mounting plate for bend switch Mounting plate compatible with a bend switch. Must be compatible with relevant switch
83 | RT275-03-038 | Assistive Accessories | Mounting plate for squeeze switch or joystick Mounting plate compatible with a switch that supports activation with a squeeze or moving a joystick. Must be compatible with relevant switch or joystick
84 | RT275-03-039 | Assistive Accessories | Mounting plate for light-touch switch Mounting plate compatible with a switch that supports activation with very minimal light touch. Must be compatible with relevant switch
85 | RT275-03-040 | Assistive Accessories | Universal switch mounting plate kit Mounting plate kit compatible with a variety of switches
86 | RT275-03-041 | Assistive Accessories | Device mounting plate Mounting plate compatible with AAC device
87 | RT275-03-042 | Assistive Accessories | Mounting arm and clamp for communication device Attachment to mount communication device to prevent any unwanted movement as well as facilitate access to communication device. Must include a clamp for connection to tables as well as tubes
88 | RT275-03-043 | Assistive Accessories | Universal mounting system Mounting system for secure placement of any switch or device. Arm must be extendable and adjustable. Must include a super clamp to clamp onto flat or tubular surfaces, and mounting plate. Must have Friction knob. Type of mounting plate to be supplied must be specified on order
89 | RT275-03-044 | Assistive Accessories | Universal mounting system Mounting system for secure placement of any switch or device. Arm must be extendable and adjustable. Must include a super clamp to clamp onto flat or tubular surfaces, and mounting plate. Must have 3 pivot points, quick release levers to extend 51cm, can hold up to 3.2kg. Type of mounting plate to be supplied must be specified on order
90 | RT275-03-045 | Assistive Accessories | Floor stand and VESA mounting plate Rollable upright mount for the positioning of communication devices to prevent unwanted movement and facilitate access to communication devices, e.g. eye gaze devices over beds or wheelchairs. Includes VESA mounting plate (GA 100 X 100), a device adapter.
91 | RT275-03-046 | Assistive Accessories | Floor stand and VESA mounting plate, telescopic Rollable upright mount with telescopic height adjustment, for the positioning of communication devices to prevent unwanted movement and facilitate access to communication devices, e.g. eye gaze devices over beds or wheelchairs. Includes VESA mounting plate (GA 100 X 100), a device adapter.
92 | RT275-03-047 | Assistive Accessories | Mounting arm with clamp for heavy devices One armed aluminuim mounting sytems with table clamp, keeps devices weighing 6-10kg "floating" for easy re-positioning. Device can be moved 90 degrees vertically and 360 degrees horizontally
93 | RT275-03-048 | Assistive Accessories | Device mounting plate and arm Mounting plate and arm compatible with AAC device
94 | RT275-03-049 | Assistive Accessories | Mounting arm with clamp for large devices Mounting arm for tablets, larger communication devices, monitors, laptops; Vertical reach Up to 75cm, Horizontal up to 67cm; Weight capacity 5kg; Aluminium; 3 Pivot Points, Quick Shift levers Tilt, Pan, Rotate, Diagonal; Super Clamp for poles or surfaces 1-5.5cm thick. Requires mounting plate.
95 | RT275-03-050 | Assistive Accessories | Mounting Arm with table stand for large devices For iPads, Large Communication Devices, Tablets, Monitor, Laptops; Height 5-50cm Weight capacity up to 7.5kg; Aluminium; 3 Pivot Points, QuickShift levers Tilt, Pan, Rotate, Diagonal; Stands on Table or Floor; Requires mounting plate.
96 | RT275-03-051 | Assistive Accessories | Mounting arm with clamp, turn knobs For Switches, iPads, Light-Medium Communication Devices, Tablets; reach Up to 52cm; Weight up to 2.3kg; Aluminium; 3 Pivot Points, levers at each Turn knob to secure all; Super Clamp for poles or surfaces 1-5.5cm thick; Requires adapter for device.
97 | RT275-03-052 | Assistive Accessories | Lightweight mounting arm for switches Reach Up to 50-54cm; weight up to 0.5kg; Aluminium; 3 Pivot Points, locks at each; Super Clamp for poles or surfaces 1-5.5cm thick; with universal mounting plate.
98 | RT275-03-053 | Assistive Accessories | Lightweight mounting arm kit for switches Reach Up to 40cm; weight up to 0.5kg; Aluminium; 3 Pivot Points, locks at each; Super Clamp for poles or surfaces 1.3-3.5cm thick; with 3 plates for popular switches.
99 | RT275-03-054 | Assistive Accessories | Gooseneck mounting arm For switches, Light Communication Devices; reach up to 50cm; weight capacity up to 0.5kg; Brass; Flexible Bending Arm; Super Clamp for poles or surfaces 1-5.5cm thick; with Universal Mounting Plate
100 | RT275-03-055 | Assistive Accessories | Mounting arm for laptop or monitors 75 x 75 or 100 x 100 VESA holes or 11-15.6" Laptops; Height 5-50cm; Weight from 1.5-4kg; Aluminium; 85° Tilt, Swivel, Pan, Rotate, upward and downward tilt, and 360°rotation. Gas spring loaded. Desk Clamp 2-10cm deep, or Grommet for drilled hole up to 7cm wide, 9cm deep. VESA for Monitors, Tray for Laptops. Flexible easy swing motion to position.
101 | RT275-03-056 | Assistive Accessories | Mounting plate for tablet Mounting Plate with tablet holder that can be secured to a table top Ball joints allow a wide range of easy adjustment of angle and choice of landscape or portrait mode. Base with screw fixing holes to allow firm mounting
102 | RT275-03-057 | Assistive Accessories | Mounting plate for 10" tablet (holder with clamp mount) Mounting Plate with clamp to mount a tablet to a round bar or table edge.
103 | RT275-03-058 | Assistive Accessories | Mounting plate for 11" tablet (holder with clamp mount) Mounting Plate with clamp to mount a tablet to a round bar or table edge.
104 | RT275-03-059 | Assistive Accessories | Mounting plate for 12" tablet (holder with clamp mount) Mounting Plate with clamp to mount a tablet to a round bar or table edge.
105 | RT275-03-060 | Assistive Accessories | Mounting plate for tablet (adjustable holder with clamp mount) Mounting Plate with clamp to mount a tablet to a round bar or table edge.
106 | RT275-03-061 | Assistive Accessories | Mounting plate for tablet (holder with mounting arm and clamp) Mounting Plate with clamp and flexible gooseneck arm to mount a tablet to a round bar or table edge. Includes universal switch mounting plate.
107 | RT275-03-062 | Assistive Accessories | Clamp for mounting arm Clamp for use with mounting arms. Connects the arms to table and tubes. Maximum bite 5cm.
108 | RT275-03-063 | Assistive Accessories | Universal Battery Charger Charger for rechargeable batteries: NiMH, compatible with 9V, AAA, AA
109 | RT275-03-064 | Assistive Accessories | 9V Battery Charger Charger for 9V NiMH rechargeable batteries
110 | RT275-03-065 | Assistive Accessories | Rechargeable AAA battery NiMH, 700mAh, 500 charge cycles, single unit
111 | RT275-03-066 | Assistive Accessories | Rechargeable AAA battery NiMH, 1000mAh, 500 charge cycles, single unit
112 | RT275-03-067 | Assistive Accessories | Rechargeable AA battery NiMH, 1500mAh, 500 charge cycles, single unit
113 | RT275-03-068 | Assistive Accessories | Rechargeable AA battery NiMH, 2000mAh, 500 charge cycles, single unit
114 | RT275-03-069 | Assistive Accessories | Rechargeable 9V battery NiMH, 200mAh, 500 charge cycles, single unit
115 | RT275-03-070 | Assistive Accessories | 9V Li-Ion battery Disposable, 9V, Li-Ion, single unit
116 | RT275-04-001 | ICT Base Devices and Solutions | Basic 7" tablet for AAC solutions Form factor: Slate tablet Windows 10/Android 5, 7" display, 720 lines, multi-touch Quad-core ARM/Atom CPU (or equivalent), 1GB RAM, 32GB storage Weight & dimensions: 900g, <1.5cm I/O: USB 2 + 3 ports, HDMI, Audio on-board, WiFi4, Bluetooth 4 2MP+1MP camera, 5 hours battery life Proof of compliance issued by SITA to be submitted with bid Proof of compliance issued by ICASA to be submitted with bid
117 | RT275-04-002 | ICT Base Devices and Solutions | Entry-level 11" tablet for AAC solutions Form factor: Slate tablet/Convertible/Detachable Windows 10, 11" display, 768 lines, multi-touch Quad-core Celeron CPU (or equivalent), 2GB RAM, 64GB storage Weight & dimensions: 900kg, <1.5cm I/O: USB 2 + 3 ports, HDMI, Audio on-board, WiFi4, Bluetooth 4 2MP+1MP camera, 6 hours battery life Proof of compliance issued by SITA to be submitted with bid Proof of compliance issued by ICASA to be submitted with bid
118 | RT275-04-003 | ICT Base Devices and Solutions | Midrange 13" tablet for AAC solutions Form factor: Slate tablet/Convertible/Detachable Windows 10, 13" display, 768 lines, multi-touch Quad-core Core i3 CPU (or equivalent), 4GB RAM, 128GB storage Weight & dimensions: 1.2kg, <1.5cm I/O: USB 2 + 3 ports, HDMI, Audio on-board, WiFi4, Bluetooth 4 4MP+1MP camera, 8 hours battery life Proof of compliance issued by SITA to be submitted with bid Proof of compliance issued by ICASA to be submitted with bid
119 | RT275-04-004 | ICT Base Devices and Solutions | High-end 14" tablet for AAC solutions Form factor: Slate tablet/Convertible/Detachable Windows 10, 14" display, 1080 lines, multi-touch Quad-core Core i5 CPU (or equivalent), 8GB RAM, 256GB storage Weight & dimensions: 1.2kg, <1.5cm I/O: USB 2 + 3 ports, HDMI, Audio on-board, WiFi4, Bluetooth 4 5MP+2MP camera, 8 hours battery life Proof of compliance issued by SITA to be submitted with bid Proof of compliance issued by ICASA to be submitted with bid
120 | RT275-04-005 | ICT Base Devices and Solutions | Dedicated 10" Windows tablet for AAC solution Windows 10, 10" display, 1080 lines, multi-touch 4-core 1.8GHz low-power (e.g. Intel Atom), 4GB RAM, 64GB storage Connectivity: WiFi5, Bluetooth 4.0 LE Ports: 2xUSB 2.0; 1xUSB 3.0, Headphone/mic; Battery life: 8 hours Sensors: 3-axis gyro, accelerometer, compass, ambient light sensor Cameras: 4MP rear, 1MP front Audio: Powerful forward-facing stereo speakers + Microphone Weight 800g max, thickness 4cm max Ruggedised: Toughened glass, IP22, 80cm drop spec, magnesium or equivalent strength body. Durable case with strap, kickstand and mount plate included. Proof of compliance issued by SITA to be submitted with bid Proof of compliance issued by ICASA to be submitted with bid
121 | RT275-04-006 | ICT Base Devices and Solutions | Text-to-speech solution with indigenous SA languages Communication device with text-to-voice software in an indigenous South African language, loaded onto a midrange tablet. Language and gender to be specified at order time: Afrikaans, English, isiXhosa, isiZulu, Sepedi, isiNdebele, Tshivenda, Sesotho, Setswana, SiSwati, Xitsonga. The bidder to indicate if the device is compatible with windows/android/IOS. Proof of compliance issued by SITA to be submitted with bid Proof of compliance issued by ICASA to be submitted with bid
122 | RT275-04-007 | ICT Base Devices and Solutions | Text-to-speech solution with indigenous SA language (Afrikaans & English) Communication device with text-to-voice software in Afrikaans and English with translated text and symbol interfaces. The bidder to indicate if the device is compatible with windows/android/IOS. Proof of compliance issued by SITA to be submitted with bid Proof of compliance issued by ICASA to be submitted with bid
123 | RT275-04-008 | ICT Base Devices and Solutions | Tablet communication Device with AAC software Midrange AAC tablet, preloaded with software with a PCS symbol library and communication device templates and AAC software with a symbol library and free downloadable Core First Lessons and Books. Proof of compliance issued by SITA to be submitted with bid Proof of compliance issued by ICASA to be submitted with bid
124 | RT275-04-009 | ICT Base Devices and Solutions | Tablet communication Device with AAC software Midrange AAC tablet, preloaded with software with a symbol library and communication device templates. Includes Widgit, Symbol Stix and IT Pictures symbols. Proof of compliance issued by SITA to be submitted with bid Proof of compliance issued by ICASA to be submitted with bid
125 | RT275-04-010 | ICT Base Devices and Solutions | Tablet communication Device with AAC software Midrange AAC tablet, preloaded with software to support literacy and symbol-based communication, and is a primary writing tool for emerging readers and writers. Proof of compliance issued by SITA to be submitted with bid Proof of compliance issued by ICASA to be submitted with bid
126 | RT275-04-011 | ICT Base Devices and Solutions | Tablet communication Device with AAC software Midrange AAC tablet, preloaded with software with a symbol library, translated interfaces in English and Afrikaans, adult and child voices and text or symbol communication. Proof of compliance issued by SITA to be submitted with bid Proof of compliance issued by ICASA to be submitted with bid
127 | RT275-04-012 | ICT Base Devices and Solutions | Dedicated 10" tablet with AAC applications Dedicated AAC tablet preloaded with software: Snap + Core First or Communicator 5 Proof of compliance issued by SITA to be submitted with bid Proof of compliance issued by ICASA to be submitted with bid
128 | RT275-04-013 | ICT Base Devices and Solutions | Speech-Generating Tablet Windows-based speech-generating tablet with built-in dual 3W speakers, Symbol-based AAC software built around evidence-based core-word framework, accessible via touch, switch, and mouse. Support for Windows access via optional dongle. Durable case with strap, kickstand and mount plate included. Proof of compliance issued by SITA to be submitted with bid Proof of compliance issued by ICASA to be submitted with bid
129 | RT275-04-014 | ICT Base Devices and Solutions | Tablet keyguard Keyguard to improve touch accuracy on tablet. Must be compitable with the offered tablet item RT275-04-013
130 | RT275-04-015 | ICT Base Devices and Solutions | Eyegaze-based speech generating system Built-in switch access hardware, IR and prolonged battery and powerful speakers Purpose-built bracket solution for mobile communication and computer access, built-in eye tracking device, speech recognition, 2 switch ports and IR for hands-free access to windows-based tablets, 2 applications for windows access and control. Tablet: 11-13", Intel Core i5, 4GB RAM, 128GB storage, Windows 10 Pro with Windows Control software. Mounting plate included. Mounting arm not included. Communication software must be added separately. Proof of compliance issued by SITA to be submitted with bid Proof of compliance issued by ICASA to be submitted with bid
131 | RT275-04-016 | ICT Base Devices and Solutions | Ruggedised AAC Tablet, 8" Portable communication aid with amplified sound. Used with touch, switches and other USB devices. 8” toughened touch display, HD. Wheelchair mounting plate, built-in stand, shoulder strap and handle included. Rugged casing. Includes AAC communication software in English and Afrikaans text & symbols. Mounting arm not included. Proof of compliance issued by SITA to be submitted with bid Proof of compliance issued by ICASA to be submitted with bid
132 | RT275-04-017 | ICT Base Devices and Solutions | Ruggedised AAC Tablet, 10" Portable communication aid with amplified sound. Used with touch, switches and other USB devices. 10” toughened touch display, HD. Wheelchair mounting plate, built-in stand, shoulder strap and handle included. Rugged casing. Includes AAC communication software in English and Afrikaans text & symbols. Mounting arm not included. Proof of compliance issued by SITA to be submitted with bid Proof of compliance issued by ICASA to be submitted with bid
133 | RT275-04-018 | ICT Base Devices and Solutions | Ruggedised AAC Tablet, 12" Portable communication aid with amplified sound. Used with touch, switches and other USB devices. Eye gaze option. 1m drop spec. 12.5” full HD toughened touch display. Remote power button. Tactile, water-resistant buttons. Multi-angle stand and wheelchair mounting plate included. Includes AAC communication software in English and Afrikaans text & symbols. Mounting arm not included. Proof of compliance issued by SITA to be submitted with bid Proof of compliance issued by ICASA to be submitted with bid
134 | RT275-04-019 | ICT Base Devices and Solutions | Ruggedised Speech Generating Tablet with Eye Tracking Dedicated speech generating tablet with bult-in advanced eyetracker, Windows 10, 12"-15" Gorilla Glass display, IR enabled, powerful world-facing speakers for clear and loud augmentative and alternative communication. Mounting not included. Proof of compliance issued by SITA to be submitted with bid Proof of compliance issued by ICASA to be submitted with bid
135 | RT275-04-020 | ICT Base Devices and Solutions | High-visibility keyboard Large keyboard, high-contrast keys to improve readability and reduce eyestrain. 20+mm square keys. The bidder to indicate if the device is compatible with windows/android/IOS device
136 | RT275-04-021 | ICT Base Devices and Solutions | Metal cover for high-visibility keyboard Metal frame that fits securely over alternative keyboard to provide physical separation (Compatible with, award with High-visibility keyboard)
137 | RT275-04-022 | ICT Base Devices and Solutions | Clear cover for high-visibility keyboard Clear perspex frame that fits securely over alternative keyboard to provide physical separation (Compatible with, award with High-visibility keyboard)
138 | RT275-04-023 | ICT Base Devices and Solutions | High-visibility keyboard, wireless Large wireless keyboard, high-contrast/multi-colour keys to improve readability and reduce eyestrain. 20+mm square keys
139 | RT275-04-024 | ICT Base Devices and Solutions | Clear cover for wireless high-visibility keyboard Clear perspex frame that fits securely over alternative keyboard to provide physical separation (Compatible with, award with High-visibility keyboard, wireless)
140 | RT275-04-025 | ICT Base Devices and Solutions | Large-Print Keyboard Full-size keyboard with high-contrast, large-print bold keycaps; Various colour schemes; USB connection
141 | RT275-04-026 | ICT Base Devices and Solutions | Keyguard for Large-Print Keyboard Frame fitted over the top of an alternative keyboard to provide physical separation between each key.
142 | RT275-04-027 | ICT Base Devices and Solutions | Large-Print Keyboard with magnification support Full-size keyboard with high-contrast, large-print bold keycaps; Various colour schemes; USB connection; Dedicated keys for screen magnifier software
143 | RT275-04-028 | ICT Base Devices and Solutions | Large-Key Keyboard Keyboard with large keys for easier identification and selection of the keys.
144 | RT275-04-029 | ICT Base Devices and Solutions | Keyguard for Large-Key Keyboard Frame fitted over the top of an alternative keyboard to provide physical separation between each key.
145 | RT275-04-030 | ICT Base Devices and Solutions | Single-handed Keyboard Ergonomic keyboard to enable one-hand typing. Special shape and letter layout; Match natural hand movement; Minimise finger movement; Increase typing speed; Available for right-handed or left-handed typist; USB connection
146 | RT275-04-031 | ICT Base Devices and Solutions | Pointing device - Joystick Joystick with speed control key, latching drag switch and button that allow only up/down or left/right movement of the cursor; 5 different cursors Proof of compliance issued by SITA to be submitted with bid Proof of compliance issued by ICASA to be submitted with bid
147 | RT275-04-032 | ICT Base Devices and Solutions | Pointing device - Advanced joystick Joystick with speed control key, latching drag switch and button that allows only up/down or left/right movement of the cursor; 5 different cursors, anti-tremor function, drag and drop, pan feature, dwell clicker feature, light touch.The bidder to indicate if the device is compatible with windows/android/IOS, and how does it connect to device -USB/Bluetooth/wireless/3.5mm jack, battery oprated? Proof of compliance issued by SITA to be submitted with bid Proof of compliance issued by ICASA to be submitted with bid
148 | RT275-04-033 | ICT Base Devices and Solutions | Pointing device - Trackball Trackball with speed control key, latching drag switch and button, allows only up/down or left/right movement of the cursor. 62-64mm ball Proof of compliance issued by SITA to be submitted with bid Proof of compliance issued by ICASA to be submitted with bid
149 | RT275-04-034 | ICT Base Devices and Solutions | Pointing device - Advanced trackball Trackball with speed control key, latching drag switch and button allows only up/down or left/right movement of the cursor; Dwell Clicker feature, Pan feature, switchable left/right click, double click feature, drag/lock feature, anti-tremor feature, audible and visual warnings. 62-64mm ball Proof of compliance issued by SITA to be submitted with bid Proof of compliance issued by ICASA to be submitted with bid
150 | RT275-04-035 | ICT Base Devices and Solutions | Pointing device - Head mouse Wearable head-controlled pointing device with bite switch. Multi-platform (Windows and Android) mouse control. Bluetooth, Mini USB port, Switch input (3.5mm jack)
151 | RT275-04-036 | ICT Base Devices and Solutions | Pointing Device - Head / Limb Mouse Mouse alternative for users with limited mobility, enables hands-free computer use. Provide mouse movement by reading user head/limb movements; Clicking achieved through dwelling or by activating a separate switch; Wireless receiver with USB connection Proof of compliance issued by SITA to be submitted with bid Proof of compliance issued by ICASA to be submitted with bid
152 | RT275-04-037 | ICT Base Devices and Solutions | Pointing Device - Colour-coded Mouse Alternative mouse with colour-coded buttons. Small, fits easily into a child's hand; Coloured left and right buttons to distinguish between left click and right click; Scroll wheel; USB connection Proof of compliance issued by SITA to be submitted with bid Proof of compliance issued by ICASA to be submitted with bid
153 | RT275-04-038 | ICT Base Devices and Solutions | Pointing device - Entry-level eyetracker Basic eye-tracking device with PC-based processing, 30Hz data rate, supports 12-19" display, USB port; learning software bundled. The bidder to indicate if the device is compatible with windows/android/IOS, and how does it connect to device -USB/Bluetooth/wireless/3.5mm jack, battery oprated? Proof of compliance issued by SITA to be submitted with bid Proof of compliance issued by ICASA to be submitted with bid
154 | RT275-04-039 | ICT Base Devices and Solutions | Pointing device - Entry-level eyetracker with OS control Basic eye-tracking device with PC-based processing, 30Hz data rate, supports 12-19" display, full Windows control and mouse functions, USB port.The bidder to indicate if the device is compatible with windows/android/IOS. Proof of compliance issued by SITA to be submitted with bid Proof of compliance issued by ICASA to be submitted with bid
155 | RT275-04-040 | ICT Base Devices and Solutions | Pointing device - Midrange eyetracker Midrange eye-tracking device with embedded processing, 40Hz data rate, supports 12-19" display, USB port; learning software bundled. The bidder to indicate if the device is compatible with windows/android/IOS. The bidder to specify the learning software - what areas does it target- attention/scanning/memory? Proof of compliance issued by SITA to be submitted with bid Proof of compliance issued by ICASA to be submitted with bid
156 | RT275-04-041 | ICT Base Devices and Solutions | Pointing device - Midrange eyetracker with OS control Midrange eye-tracking device with embedded processing, 40Hz data rate, supports 12-19" display, full Windows control and mouse functions, USB port. The bidder to indicate if the device is compatible with windows/android/IOS. Proof of compliance issued by SITA to be submitted with bid Proof of compliance issued by ICASA to be submitted with bid
157 | RT275-04-042 | ICT Base Devices and Solutions | Pointing device - Advanced eyetracker Advanced eye-tracking device with embedded processing, 50Hz data rate, supports 12-27" display, USB port, switch port, microphone array; learning software bundled.The bidder to indicate if the device is compatible with windows/android/IOS. Proof of compliance issued by SITA to be submitted with bid Proof of compliance issued by ICASA to be submitted with bid
158 | RT275-04-043 | ICT Base Devices and Solutions | Pointing device - Advanced eyetracker with OS control Advanced eye-tracking device with embedded processing, 50Hz data rate, supports 12-27" display, full Windows control and mouse functions, USB port, switch port, microphone array. The bidder to indicate if the device is compatible with windows/android/IOS. Proof of compliance issued by SITA to be submitted with bid Proof of compliance issued by ICASA to be submitted with bid
159 | RT275-04-044 | ICT Base Devices and Solutions | Advanced Eyetracker Kit for Rugged AAC tablet With full windows control software for eye tracker. Eye tracker with 40Hz data rate, on board processing. The bidder to specify the specifications of the table. Proof of compliance issued by SITA to be submitted with bid Proof of compliance issued by ICASA to be submitted with bid
160 | RT275-04-045 | ICT Base Devices and Solutions | Optical Character Recognition (OCR) + Camera Bundle Software/hardware bundle that captures hard-copy or electronic documents and converts them into electronic text; Up to 20 pages per minute; Automatic mode with motion sensor; Text-to-speech output; Low-vision features allow customising on-screen display: font, size, spacing and colours. Proof of compliance issued by SITA to be submitted with bid Proof of compliance issued by ICASA to be submitted with bid
161 | RT275-04-046 | ICT Base Devices and Solutions | Entry-level eyetracker with Track and Learn Basic eyetracker with software bundle: analysis and point and click only, for use with learning and communication software.The bidder to indicate if the device is compatible with windows/android/IOS. Proof of compliance issued by SITA to be submitted with bid Proof of compliance issued by ICASA to be submitted with bid
162 | RT275-04-047 | ICT Base Devices and Solutions | Entry-level eyetracker with 3 Graded Inclusive Eye Gaze Learning Software titles Basic eyetracker with software bundle: analysis and point and click software only, for use with learning and communication software. Includes Inclusive Eye Gaze - All 3 titles. The bidder to indicate whether single licence/site licence/multi-device licence/single user licence - multiple devices vs single user - single device licence Proof of compliance issued by SITA to be submitted with bid Proof of compliance issued by ICASA to be submitted with bid
163 | RT275-04-048 | ICT Base Devices and Solutions | Entry-level eyetracker with software for Looking to Learn Peripheral eye-tracking device with eyechip with fully embedded processing with USB, 40cm per second head movement compensation; provides point and click only, analysis software, for use with learning software and communication software. includes inclusive eyegaze learning software with 40 activities across 5 ocular motor skill areas: sensory, explore, target, choose, control. Proof of compliance issued by SITA to be submitted with bid Proof of compliance issued by ICASA to be submitted with bid
164 | RT275-04-049 | ICT Base Devices and Solutions | Advanced eye tracker with Gaze Point and Viewing Software Peripheral eye-tracking device with eyechip with fully embedded processing with USB, 4 microphone array, 3.5mm jack integrated switch port, 40cm per second head movement compensation, connects to windows-based tablets, supports 12-27" screen size. Includes infrared dongle for environmental control, provides point and click only, for use with learning software and communication software, and analysis software Proof of compliance issued by SITA to be submitted with bid Proof of compliance issued by ICASA to be submitted with bid
165 | RT275-04-050 | ICT Base Devices and Solutions | Advanced eye tracker for Inclusive Eyegaze Peripheral eye-tracking device with eyechip with fully embedded processing with USB, 4 microphone array, 3.5mm jack integrated switch port, 40cm per second head movement compensation, connects to windows-based tablets, supports 12-27" screen size. Includes infrared dongle for environemntal control, provides point and click, includes inclusive eyegaze learning software, all three titles Proof of compliance issued by SITA to be submitted with bid Proof of compliance issued by ICASA to be submitted with bid
166 | RT275-04-051 | ICT Base Devices and Solutions | Mini Eye Tracking Solution - High end tablet Mini eye tracker with Advanced AAC tablet. Durable mounting case, with mounting plate included. Communication software or learning content must be added separately. Mounting arm must be added separately Proof of compliance issued by SITA to be submitted with bid Proof of compliance issued by ICASA to be submitted with bid
167 | RT275-04-052 | ICT Base Devices and Solutions | Entry-level Eye Tracking solution - Midrange tablet Mini eyetracker with midrange AAC tablet. Durable mounting case, with mounting plate included. Communication software, learning content and mounting arm not included Proof of compliance issued by SITA to be submitted with bid Proof of compliance issued by ICASA to be submitted with bid
168 | RT275-04-053 | ICT Base Devices and Solutions | Advanced Eye Tracking Solution - Advanced tablet Advanced Eye Tracker with advanced AAC tablet. Includes OS control software, durable case for mounting and mounting plate; Mounting arm not included. Communication software, learning content and mounting arm not included Proof of compliance issued by SITA to be submitted with bid Proof of compliance issued by ICASA to be submitted with bid
169 | RT275-04-054 | ICT Base Devices and Solutions | Digital Braille Writer Mechanical braille typewriter with digital features; Six-key braille keyboard; Paper size up to A4; text to speech output Proof of compliance issued by SITA to be submitted with bid Proof of compliance issued by ICASA to be submitted with bid
170 | RT275-04-055 | ICT Base Devices and Solutions | Digital Braille Writer, Advanced Mechanical braille typewriter with digital features; Six-key braille keyboard; Paper size up to A4; Built-in display for braille and large-print output; text to speech output Proof of compliance issued by SITA to be submitted with bid Proof of compliance issued by ICASA to be submitted with bid
171 | RT275-04-056 | ICT Base Devices and Solutions | Portable Refreshable Braille Display Enables blind computer user to read text output; 12 refreshable braille cells; Cursor router buttons; Braille keyboard; Navigation keys; USB and Bluetooth; Compatible with screen reading software; Multi-platform: Windows, MacOS, iOS, Android; Rechargeable battery Proof of compliance issued by SITA to be submitted with bid Proof of compliance issued by ICASA to be submitted with bid
172 | RT275-04-057 | ICT Base Devices and Solutions | Portable Refreshable Braille Display with Note-Taking Enables blind computer user to read text output; 14 refreshable braille cells; Cursor router buttons; Braille keyboard; Navigation keys; Standalone note-taking and book reading capability; USB and Bluetooth; Compatible with screen reading software; Multi-platform: Windows, MacOS, iOS, Android; Rechargeable battery Proof of compliance issued by SITA to be submitted with bid Proof of compliance issued by ICASA to be submitted with bid
173 | RT275-04-058 | ICT Base Devices and Solutions | Desktop Refreshable Braille Display Enables blind computer user to read text output; 40 refreshable braille cells; Cursor router buttons; Braille keyboard; Navigation keys; USB and Bluetooth; Compatible with screen reading software; Multi-platform: Windows, MacOS, iOS, Android; Rechargeable battery Proof of compliance issued by SITA to be submitted with bid Proof of compliance issued by ICASA to be submitted with bid
174 | RT275-04-059 | ICT Base Devices and Solutions | Advanced Refreshable Braille Display Enables blind computer user to read text output; 80 refreshable braille cells; Cursor router buttons; Braille keyboard; Navigation keys; Standalone note-taking and book reading capability; USB and Bluetooth; Compatible with screen reading software; Multi-platform: Windows, MacOS, iOS, Android; Rechargeable battery Proof of compliance issued by SITA to be submitted with bid Proof of compliance issued by ICASA to be submitted with bid
175 | RT275-04-060 | ICT Base Devices and Solutions | Braille Reader and Notetaker Portable refreshable braille device with display and braille keyboard; 20 eight-dot braille cells; Book reader, note-taker and braille display; Language independent; SD card support; USB and Bluetooth connectivity; Rechargeable and user-replaceable battery Proof of compliance issued by SITA to be submitted with bid Proof of compliance issued by ICASA to be submitted with bid
176 | RT275-04-061 | ICT Base Devices and Solutions | Braille Personal Digital Assistant Portable device with braille keyboard for integrated computing experience; Speech synthesiser and 18-cell braille display; Cursor routing keys; Read any content in speech, braille or both; Support for multiple languages; Stereo speakers and microphone; Built-in productivity applications; Internal storage and USB/SD card support for expansion; USB, Wi-Fi, Bluetooth and GPS; Rechargeable and user-replaceable battery Proof of compliance issued by SITA to be submitted with bid Proof of compliance issued by ICASA to be submitted with bid
177 | RT275-04-062 | ICT Base Devices and Solutions | Braille Personal Digital Assistant, 32-cell Portable device with braille keyboard for integrated computing experience; Speech synthesiser and 32-cell braille display; Cursor routing keys; Read any content in speech, braille or both; Support for multiple languages; Stereo speakers and microphone; Built-in productivity applications; Internal storage and USB/SD card support for expansion; USB, Wi-Fi, Bluetooth and GPS; Rechargeable and user-replaceable battery Proof of compliance issued by SITA to be submitted with bid Proof of compliance issued by ICASA to be submitted with bid
178 | RT275-04-063 | ICT Base Devices and Solutions | Braille Personal Digital Assistant - Android Portable device with braille keyboard for integrated computing experience; Speech synthesiser and 20-cell braille display; Android platform; Cursor routing keys; Read any content in speech, braille or both; Support for multiple languages; Stereo speakers and microphone; Built-in productivity applications; Access to mainstream apps; Internal storage and USB/SD card support for expansion; HDMI, USB, Wi-Fi, Bluetooth and GPS; Camera for OCR capture; Rechargeable and user-replaceable battery Proof of compliance issued by SITA to be submitted with bid Proof of compliance issued by ICASA to be submitted with bid
179 | RT275-04-064 | ICT Base Devices and Solutions | Braille Personal Digital Assistant - Android, 32-cell Portable device with braille keyboard for integrated computing experience; Speech synthesiser and 20-cell braille display; Android platform; Cursor routing keys; Read any content in speech, braille or both; Support for multiple languages; Stereo speakers and microphone; Built-in productivity applications; Access to mainstream apps; Internal storage and USB/SD card support for expansion; HDMI, USB, Wi-Fi, Bluetooth and GPS; Camera for OCR capture; Rechargeable and user-replaceable battery Proof of compliance issued by SITA to be submitted with bid Proof of compliance issued by ICASA to be submitted with bid
180 | RT275-04-065 | ICT Base Devices and Solutions | Braille Personal Digital Assistant - Windows Portable device with braille keyboard for integrated computing experience; Screen reader (JAWS) and 40-cell braille display; Windows 10 operating system; Cursor routing keys; Read any content in speech, braille or both; Support for multiple languages; Stereo speakers and microphone; Internal storage and USB/SD card support for expansion; HDMI, USB, Wi-Fi, Bluetooth and GPS; Rechargeable battery Proof of compliance issued by SITA to be submitted with bid Proof of compliance issued by ICASA to be submitted with bid
181 | RT275-04-066 | ICT Base Devices and Solutions | Low-Volume Braille Printer Embosses text as braille using braille translation software; 60 cps print speed; Tractor-fed fan-folded paper 11.5" x 11"; Single-sided printing; Includes tactile graphics software; USB connection Proof of compliance issued by SITA to be submitted with bid Proof of compliance issued by ICASA to be submitted with bid
182 | RT275-04-067 | ICT Base Devices and Solutions | Low-Volume Braille Printer with Ink Printing Embosses text as braille using braille translation software and prints text with 300 dpi colour and black ink; 60 cps print speed; Cut-sheet paper up to A4; Single-sided printing; Includes tactile graphics software; USB connection Proof of compliance issued by SITA to be submitted with bid Proof of compliance issued by ICASA to be submitted with bid
183 | RT275-04-068 | ICT Base Devices and Solutions | Medium-Volume Braille Printer Embosses text as braille using braille translation software; 120 cps print speed; Tractor-fed fan-folded paper 11.5" x 11"; Single and double-sided printing; Able to produce tactile graphics with optional software; USB, Ethernet or Wi-Fi connection; Print from portable devices; Optional acoustic hood Proof of compliance issued by SITA to be submitted with bid Proof of compliance issued by ICASA to be submitted with bid
184 | RT275-04-069 | ICT Base Devices and Solutions | Medium-Volume Braille Printer, Cut Sheet Embosses text as braille using braille translation software; 120 cps print speed; Cut-sheet paper up to A3; Single and double-sided printing; Optional tactile graphics software; USB or Ethernet connection
185 | RT275-04-070 | ICT Base Devices and Solutions | Medium-Volume Braille Printer with Tactile Graphics Embosses text as braille using braille translation software; 120 cps print speed; Tractor-fed fan-folded paper 11.5" x 11"; Single and double-sided printing; Includes tactile graphics software; USB or Ethernet connection Proof of compliance issued by SITA to be submitted with bid Proof of compliance issued by ICASA to be submitted with bid
186 | RT275-04-071 | ICT Base Devices and Solutions | Medium-Volume Braille Printer with Tactile Graphics, Cut Sheet Embosses text as braille using braille translation software; 120 cps print speed; Cut-sheet paper up to A3; Single and double-sided printing; Includes tactile graphics software; USB or Ethernet connection Proof of compliance issued by SITA to be submitted with bid Proof of compliance issued by ICASA to be submitted with bid
187 | RT275-04-072 | ICT Base Devices and Solutions | High-Volume Braille Printer Embosses text as braille using braille translation software; 100 cps print speed; Tractor-fed fan-folded paper 11.5" x 11"; Single and double-sided printing; Includes tactile graphics software; USB or Ethernet connection; Optional ink attachment Proof of compliance issued by SITA to be submitted with bid Proof of compliance issued by ICASA to be submitted with bid
188 | RT275-04-073 | ICT Base Devices and Solutions | High-Volume Braille Printer, 200 cps Embosses text as braille using braille translation software; 200 cps print speed; Tractor-fed fan-folded paper 11.5" x 11"; Single and double-sided printing; Includes tactile graphics software; USB or Ethernet connection; Optional ink attachment Proof of compliance issued by SITA to be submitted with bid Proof of compliance issued by ICASA to be submitted with bid
189 | RT275-04-074 | ICT Base Devices and Solutions | High-Volume Braille Printer, Cut Sheet, Noise Cancelling Embosses text as braille using braille translation software; 300 cps print speed; Noise-cancelling design; Tractor-fed fan-folded paper 11.5" x 11"; Single and double-sided printing; Able to produce tactile graphics with optional software; USB, Ethernet or Wi-Fi connection; Print from portable devices; Includes braille stapler Proof of compliance issued by SITA to be submitted with bid Proof of compliance issued by ICASA to be submitted with bid
190 | RT275-04-075 | ICT Base Devices and Solutions | High-Volume Braille Printer, Cut Sheet, Noise Cancelling Embosses text as braille using braille translation software; 300 cps print speed; Noise-cancelling design; Cut-sheet paper up to A3; Single and double-sided printing; Automatic booklet format printing; Able to produce tactile graphics with optional software; USB, Ethernet or Wi-Fi connection; Print from portable devices; Includes braille stapler Proof of compliance issued by SITA to be submitted with bid Proof of compliance issued by ICASA to be submitted with bid
191 | RT275-04-076 | ICT Base Devices and Solutions | High-Volume Braille Printer Heavy Duty Produces braille books at high speed; Heavy-duty construction; Speed of 300 cps; Noise-cancelling design; Tractor-fed fan-folded paper 11.5" x 11"; Single and double-sided printing; USB or Ethernet connection; Includes braille translation software Proof of compliance issued by SITA to be submitted with bid Proof of compliance issued by ICASA to be submitted with bid
192 | RT275-04-077 | ICT Base Devices and Solutions | High-Volume Braille Printer Heavy Duty, 600 cpis Produces braille books at high speed; Heavy-duty construction; Speed of 600 cps; Noise-cancelling design; Uses paper rolls; Single and double-sided printing; Built-in paper cutter and stacker; USB or Ethernet connection; Includes braille translation software Proof of compliance issued by SITA to be submitted with bid Proof of compliance issued by ICASA to be submitted with bid
193 | RT275-04-078 | ICT Base Devices and Solutions | Braille Paper, tractor Continuous tractor feed 11.5" x 11" 135g; 1000 sheets per box
194 | RT275-04-079 | ICT Base Devices and Solutions | Braille Paper, A4 Single sheet 135g A4 (210 x 297 mm); 500 sheets per pack
195 | RT275-04-080 | ICT Base Devices and Solutions | Braille Paper, A3 Single sheet 135g A3 (297 x 420 mm); 500 sheets per pack
196 | RT275-04-081 | ICT Base Devices and Solutions | Braille Paper, reel Continuous reel 135g +/- 190 kg; 280 mm wide 1.2 m OD
197 | RT275-04-082 | ICT Base Devices and Solutions | Tactile Duplicator - Swell Produces graphics using raised surfaces to convey non-textual information e.g. tactile diagrams, maps and graphs. Reproduce tactile graphics using swell paper Proof of compliance issued by SITA to be submitted with bid Proof of compliance issued by ICASA to be submitted with bid
198 | RT275-04-083 | ICT Base Devices and Solutions | Tactile Duplicator Consumables - Swell Swell paper: A4 (200 sheets per pack) or A3 (100 sheets per pack)
199 | RT275-04-084 | ICT Base Devices and Solutions | Tactile Duplicator - Thermoform Produces graphics using raised surfaces to convey non-textual information e.g. tactile diagrams, maps and graphs; Reproduce tactile graphics using thermoform paper; Make multiple copies from an original Proof of compliance issued by SITA to be submitted with bid Proof of compliance issued by ICASA to be submitted with bid
200 | RT275-04-085 | ICT Base Devices and Solutions | Tactile Duplicator Consumables - Thermoform Thermoform paper: 11" x 11.5" or 8.5" x 11" (500 per pack)
201 | RT275-04-086 | ICT Base Devices and Solutions | Tactile Touchpad and Audio Software Incorporates audio feedback with tactile graphics, providing interactive way to understand complex, visually-oriented information. Creates audio-tactile graphics from standard image formats; Offers various types of textures for ease of interpretation; Includes OCR software; Touchpad A4 size minimum; USB connection Proof of compliance issued by SITA to be submitted with bid Proof of compliance issued by ICASA to be submitted with bid
202 | RT275-04-087 | ICT Base Devices and Solutions | Specialised e-Reader Accessible mobile device for reading digital publications such as e-books and periodicals. Adjustable font size; Text-to-speech; Adjustable background colour scheme; Accessible menus (voice-over); Optional adjustable font colour scheme Proof of compliance issued by SITA to be submitted with bid Proof of compliance issued by ICASA to be submitted with bid
203 | RT275-04-088 | ICT Base Devices and Solutions | Multimedia Player Portable multimedia player and e-book reader (including audio and PDF formats); Male/female text-to-speech voices; Automatic bookmarking; Voice recorder; Rechargeable battery. The bidder to specify the life-span of battery. Proof of compliance issued by SITA to be submitted with bid Proof of compliance issued by ICASA to be submitted with bid
204 | RT275-04-089 | ICT Base Devices and Solutions | Multimedia Player - Wi-Fi Portable multimedia player and e-book reader (including audio and PDF formats); Male/female text-to-speech voices; Automatic bookmarking; Voice recorder; Rechargeable battery; Wi-Fi connectivity for internet radio and podcast feeds Proof of compliance issued by SITA to be submitted with bid Proof of compliance issued by ICASA to be submitted with bid
205 | RT275-04-090 | ICT Base Devices and Solutions | Multimedia Player - Wi-Fi, OCR Proof of compliance issued by SITA to be submitted with bid Proof of compliance issued by ICASA to be submitted with bid Portable multimedia player and e-book reader (including audio and PDF formats); Male/female text-to-speech voices; Automatic bookmarking; Voice recorder; Rechargeable battery; Wi-Fi connectivity for internet radio and podcast feeds; OCR feature to capture and read aloud printed documents
206 | RT275-04-091 | ICT Base Devices and Solutions | Digital Voice Recorder Recording device to record speech and other audio for later playback. High-quality recording and playback; Built-in microphone and speaker; At least 8 hours of recording time; USB port; Battery-operated Proof of compliance issued by SITA to be submitted with bid Proof of compliance issued by ICASA to be submitted with bid
207 | RT275-04-092 | ICT Base Devices and Solutions | Text-to-Speech Handheld Scanner Reads English text aloud when scanned over print to promote independent reading. Hear words and lines of text read aloud; Dictionary feature; Scan directly to a document on a computer; Access previously scanned documents on the device Proof of compliance issued by SITA to be submitted with bid Proof of compliance issued by ICASA to be submitted with bid
208 | RT275-04-093 | ICT Base Devices and Solutions | Advanced Visualiser for e-Learning and Signing High resolution (up to 4K) and refresh rate of 30 frames per second, making it suitable for capturing live signing. For e-learning, lecture capturing, distance learning. Built-in microphone with noise reduction, HDMI output. Proof of compliance issued by SITA to be submitted with bid Proof of compliance issued by ICASA to be submitted with bid
209 | RT275-04-094 | ICT Base Devices and Solutions | Carry Case for Tablet. Carry case that can comfortably fit and secure tablet with case compatible with tablets offered. The bidder to specify device size, compatibility of case. Size/fit. Ruggedised. Proof of compliance issued by SITA to be submitted with bid Proof of compliance issued by ICASA to be submitted with bid
210 | RT275-04-095 | ICT Base Devices and Solutions | Protective Frame for Tablet Protective frame for tablet that is lightweight, easy to grab and hold; Material is non-toxic, BPA, lead and latex free. Sturdy stand included; Compatible with tablets offered Proof of compliance issued by SITA to be submitted with bid Proof of compliance issued by ICASA to be submitted with bid
211 | RT275-04-096 | ICT Base Devices and Solutions | Protective carry case for tablet Durable case for specialised AAC tablet. Includes carrying strap, mounting plate, handle and kick stand Proof of compliance issued by SITA to be submitted with bid Proof of compliance issued by ICASA to be submitted with bid
212 | RT275-04-097 | ICT Base Devices and Solutions | Tablet case Rubber tablet holding case with optional strap for shoulder use. Compatible with tablets offered Proof of compliance issued by SITA to be submitted with bid Proof of compliance issued by ICASA to be submitted with bid
213 | RT275-04-098 | ICT Base Devices and Solutions | Mounting plate for tablet Tablet holder to mount tablet onto wheelchair or table. Includes mounting plate. Mounting Arm not included. Compatible with tablets offered. The bidder to specify if the mounting plate for tablet is compatible with carry cases. Sizes of the mouting plates for various laptop/tablet sizes. Proof of compliance issued by SITA to be submitted with bid Proof of compliance issued by ICASA to be submitted with bid
214 | RT275-05-001 | Assistive Software | Symbol-based communication software Enables people with limited or unclear speech to use a computer as a voice output communication aid. Supports communication for emerging to independent communicators and includes both picture symbols and text-based communication. Accessible via touch, switch, eyegaze and mouse. Includes core vocabulary, topics, instant conversational phrases and keyboard. (e.g. Snap + Core First) Proof of compliance issued by SITA to be submitted with bid
215 | RT275-05-002 | Assistive Software | Software for communication and computer access Enables people with limited or unclear speech to use a computer as a voice output communication aid, using symbols or text to build sentences. Accepts input from eye gaze, switches, head pointer, touchscreen, mouse and other options. Available in multiple languages, text or symbol interfaces, and includes learning games, core vocabulary communication, environment control and computer control software. (e.g. Grid) Proof of compliance issued by SITA to be submitted with bid
216 | RT275-05-003 | Assistive Software | Software for AAC, Communication and Environmental Control Enables people with limited or unclear speech to use a computer as a voice output communication aid, using either symbols or text to build sentences. Supports long-distance communication and accepts input from eye gaze, switches, head pointer, touchscreen, mouse and other options. For emerging communication, symbol communication and text communication supported by intelligent prediction. Includes environmental control software. (e.g. Communicator 5). The bidder to indicate if the device is compatible with windows/android Proof of compliance issued by SITA to be submitted with bid
217 | RT275-05-004 | Assistive Software | Software for communication and access Enables people with limited or unclear speech to use a computer as a voice output communication aid, using symbols or text to build sentences. Accepts input from eye gaze, switches, head pointer, touchscreen, mouse and other options. Choice of level most appropriate for the user, including choices between emerging communication, symbol communication and text communication. Ready-made text or symbol interfaces in English, or make your own in other languages. Includes learning games, core vocabulary communication, environment control and computer control software. (e.g. Tobii Dynavox Communicator; key2Enable) Proof of compliance issued by SITA to be submitted with bid
218 | RT275-05-005 | Assistive Software | Literacy development and curriculum delivery software Cross-curricular literacy toolkit for learners from ECD upwards. Educators and therapist can use to develop and support literacy and learning via symbol support, sentence-building sets, matching activities, verbal storytelling templates, mindmapping, art and paint features, and books for learners experiencing barriers to accessing the curriculum. Caters for barriers to learning including but not limited to autism, cerebral palsy, dyslexia, low vision, intellectual impairment and hearing impairment. Features text-to-speech word processor, voice recording, accessibility features for eye-gaze, switches and touch; Customisable for South African languages and South African curriculum. Multi-user version available. (e.g. Clicker 7, speechlink, naturalreader, giner) Proof of compliance issued by SITA to be submitted with bid
219 | RT275-05-006 | Assistive Software | Literacy Development Software with Reading and Spelling Support Literacy development and curriculum delivery software that caters for a variety of barriers to learning. With spelling and reading support, mind mapping and screen appearance support. Proof of compliance issued by SITA to be submitted with bid
220 | RT275-05-007 | Assistive Software | Software to create activity boards Includes drawing program, symbol database and templates for AAC devices, templates for learning activities and interactive samples (e.g. Boardmaker Studio) Proof of compliance issued by SITA to be submitted with bid
221 | RT275-05-008 | Assistive Software | Symbol-Resource Software with Speech Output Software that turns a computer into a speech output device.Includes drawing program, symbol database, templates and interactive samples.(e.g. Boardmaker with Speaking Dynamically Pro V6) Proof of compliance issued by SITA to be submitted with bid
222 | RT275-05-009 | Assistive Software | Software to make communication overlays and educational resources Includes two symbol libraries. Print boards for communication devices or make therapy /educational resources. User-friendly interface (e.g. MatrixMaker) Proof of compliance issued by SITA to be submitted with bid
223 | RT275-05-010 | Assistive Software | Simple Literacy and Typing Software with Symbol Support Fun introduction to the keyboard by using speech, phonics and symbols in colourful and engaging letter exercises to teach literacy skills. (e.g. First Keys 3) Proof of compliance issued by SITA to be submitted with bid
224 | RT275-05-011 | Assistive Software | SA text to speech voice (single language) Realistic text-to-speech voices available in eleven South African languages. Specify language, Male/Female, Adult/Child at order time. Single- and multi-user options
225 | RT275-05-012 | Assistive Software | Symbol communication book system Ready-made communication books for different levels of communication. Requires software that uses PCS symbols to edit the books before printing. For users who can point at symbols on the page with a hand or fist. (e.g. PODD - Direct Access)
226 | RT275-05-013 | Assistive Software | Symbol communication book system, alternative access Ready-made communication books for different levels of communication. Requires software that uses PCS symbols to edit the books before printing. Alternative layout options for eye pointing or partner assisted scanning. (e.g. PODD - Alternative Access)
227 | RT275-05-014 | Assistive Software | Software to create personal communication boards Contains overlay cells with multiple images and text boxes. Supports creation and editing of personal communication boards. Allows single installation. (e.g. GoTalk overlay software). Single- and multi-user options Proof of compliance issued by SITA to be submitted with bid
228 | RT275-05-015 | Assistive Software | Switch skills software Themes to help early switch users understand cause and effect, to develop switch use. Activities encourage speaking and listening, helps to develop concentration skills. Bright clear graphics are suitable for those with visual impairment, includes sound and music. Collection of easy switch timing activities for all ages. (e.g. Switch Skills 1, 2 or 3) Proof of compliance issued by SITA to be submitted with bid
229 | RT275-05-016 | Assistive Software | Switch skills software with scanning Switch activities designed for patients who are learning how to use switches to make choices either within software or scanning communication devices. (e.g. Switch Skills - Scanning) Proof of compliance issued by SITA to be submitted with bid
230 | RT275-05-017 | Assistive Software | Language and switch skills development software Enables patients to increase their ability to use switches and to demonstrate understanding of language concepts. Stimulates and encourages movement and hand function, develop visual skills, and provide users with access to a range of learning situations, particularly those limited in physical movement who cannot access traditional methods. Bidders may specify up to 10 different themes. Single- and multi-user options. (e.g. Switch It) Proof of compliance issued by SITA to be submitted with bid
231 | RT275-05-018 | Assistive Software | Language and skills development software, multi-user Enables patients to increase their ability to use switches and to demonstrate understanding of language concepts. Stimulates and encourages movement and hand function, develop visual skills, and provide users with access to a range of learning situations, particularly those limited in physical movement who cannot access traditional methods. Bidders may specify up to 10 different themes. (e.g. Switch It) Proof of compliance issued by SITA to be submitted with bid
232 | RT275-05-019 | Assistive Software | Screen Magnification Software Magnifies and enhances computer display content. Various magnification levels, colour and contrast settings. Mouse customisation available. Reading of documents, web pages and email Proof of compliance issued by SITA to be submitted with bid
233 | RT275-05-020 | Assistive Software | Screen Magnification Software with Speech Magnifies and enhances computer display content. Various magnification levels, colour and contrast settings. Mouse customisation available. Reading of documents, web pages and email. Screen reading feature with human-sounding text-to-speech synthesiser; Recorder to convert text documents into audio recordings Proof of compliance issued by SITA to be submitted with bid
234 | RT275-05-021 | Assistive Software | Screen Reading Software Converts computer screen content into synthesised speech and/or braille output (with optional braille display); Human-sounding text-to-speech synthesiser; Access to Windows, MS Office applications and web browsers. Built-in OCR Proof of compliance issued by SITA to be submitted with bid
235 | RT275-05-022 | Assistive Software | Screen Magnification and Screen Reading Software Bundle Magnifies computer screen content; Enhances colour contrast and focus; Magnification up to 60x; Various magnification levels, colour schemes and reading views; Pointer and cursor enhancements; Converts computer screen content into synthesised speech and/or braille output (with optional braille display); Human-sounding text-to-speech synthesiser; Access to Windows, MS Office applications and web browsers; Quick navigation keys to navigate documents; Built-in OCR Proof of compliance issued by SITA to be submitted with bid
236 | RT275-05-023 | Assistive Software | Optical Character Recognition (OCR) Software Captures hard-copy or electronic documents and converts them into electronic text. Proof of compliance issued by SITA to be submitted with bid Proof of compliance issued by ICASA to be submitted with bid
237 | RT275-05-024 | Assistive Software | Optical Character Recognition (OCR) Software for Visually Impaired Captures hard-copy or electronic documents and converts them into electronic text; Speech features with human-sounding text-to-speech synthesiser; Low vision features tailor the appearance of text on the screen - font style, size, spacing and colours Proof of compliance issued by SITA to be submitted with bid
238 | RT275-05-025 | Assistive Software | Simple Computing Environment Easy-to-use software program comprising user interface, email, word processing, book reading and internet access. Simplifies computing tasks. Menus with magnification, high-contrast colours and speech output; Requires minimal computer knowledge or touch typing skills. Can be disabled to access the standard desktop environment Proof of compliance issued by SITA to be submitted with bid
239 | RT275-05-026 | Assistive Software | Braille Translation Software Translates text to braille for embossing on a braille printer/embosser and formats documents according to braille layout rules. Supports all major South African languages as well as Unified English Braille Code (UEB); Offers contracted and uncontracted braille; Accessible for blind users. Proof of compliance issued by SITA to be submitted with bid
240 | RT275-05-027 | Assistive Software | Braille Translation Software for Mathematics Used for creation of mathematical notation for inclusion in braille translation software. Includes equation editor. Integrates with leading braille translation software that support UEB maths Proof of compliance issued by SITA to be submitted with bid
241 | RT275-05-028 | Assistive Software | Tactile Graphics Software Produces images using raised surfaces to convey non-textual information e.g. tactile pictures, diagrams, maps and graphs. Convert images to tactile graphics; Includes basic drawing and editing tools. Proof of compliance issued by SITA to be submitted with bid
242 | RT275-05-029 | Assistive Software | DAISY Reading Software Reads documents in Digital Accessible Information System (DAISY) format. Full navigational features; Output in synchronised speech and text. Proof of compliance issued by SITA to be submitted with bid
243 | RT275-05-030 | Assistive Software | Speech Recognition Software Recognises speech and converts it to text. Create, modify and format documents using speech. Insert frequently used phrases with voice commands. Control the computer and search the internet with voice commands.Bundled with headset. Proof of compliance issued by SITA to be submitted with bid
244 | RT275-05-031 | Assistive Software | Sign Language Development Software Software to teach South African Sign Language to learners and their communication partners such as teachers, parents and friends. Proof of compliance issued by SITA to be submitted with bid
245 | RT275-05-032 | Assistive Software | Sign Language Literacy Software Software to develop reading and writing skills in learners who use South African Sign Language. To accommodate for all grades and include adapted storybooks.
246 | RT275-05-033 | Assistive Software | Keyboard Literacy Software Teach keyboard touch typing skills; Accessible to blind and low-vision users; Built-in curriculum, allows creation of own lessons; Progress reports Proof of compliance issued by SITA to be submitted with bid
247 | RT275-05-034 | Assistive Software | Eye Gaze Learning Software Engaging activities that teach and encourage the development of eye tracking skills. Includes tools to assess and track user performance showing where the user looked and whether their eye gaze was appropriate for the activity (e.g. EyeGaze Learning Curve).
248 | RT275-05-035 | Assistive Software | Eye Gaze Learning Software - For Severe Cortical-Visual Impairment Facilitates development of sensory eye gaze skills following a hierarchical structure, from fundamental skills to more refined and controlled eye gaze movements. (e.g. Sensory EyeFX) Proof of compliance issued by SITA to be submitted with bid
249 | RT275-05-036 | Assistive Software | Eye Gaze User Assessment Software Identifies a user's ability to use an eye gaze device and their potential cognitive functions. Records eye tracking data from various applications and uses map superimposed over application to report on gaze data (E.g. Gaze Viewer) Proof of compliance issued by SITA to be submitted with bid
250 | RT275-05-037 | Assistive Software | Literacy Software for Reading and Writing Support Assists learners with reading and writing difficulties. Visual enhancements, audio and visual feedback when reading text. With support for reading, writing, study and research and can be used with any documents and web pages. Read accessible and inaccessible PDF files. Includes word prediction and dictionary and create audio file from text. Single- and multi-user options. (e.g. TextHelp Read & Write) Proof of compliance issued by SITA to be submitted with bid
251 | RT275-05-038 | Assistive Software | Maths Development Software Supports needs of users with difficulties with maths, including severe and persistent difficulties. Multi-modal activities include lesson plans, computer-based activities, number games, printable worksheets and motivating reward certificates. Optional assessment tool to identify areas of difficulty and personalise remedial action. Can be run online or offline. Proof of compliance issued by SITA to be submitted with bid
252 | RT275-05-039 | Assistive Software | Keyboard Skills Software Develops keyboard skills. Typing tutor programs to teach the keyboard layout, develop touch typing skills, improve spelling and other keyboard skills; Can be customised to suit the user's needs (e.g. Talking Typing Tutor) Proof of compliance issued by SITA to be submitted with bid
253 | RT275-05-040 | Assistive Software | Keyboard Skills Software for Learners with learning and literacy difficulties. Develops keyboard skills for learners who battle with reading and who require high repetition and reduced visual clutter. Program teaches keyboard layout, develops touch typing and can be customised to user's needs. Proof of compliance issued by SITA to be submitted with bid
254 | RT275-05-041 | Assistive Software | Mouse Skills Software Develops mouse skills. Activities to practice mouse clicking, double clicking, drag and drop and other mouse skills; Can be customised to suit the user's needs (e.g. IT Mouse Skills) Proof of compliance issued by SITA to be submitted with bid
255 | RT275-05-042 | Assistive Software | Curriculum Activities Software: Literacy Develops literacy skills. For a variety of Grades. Proof of compliance issued by SITA to be submitted with bid
256 | RT275-05-043 | Assistive Software | Curriculum Activities Software: Numeracy Develops numeracy skills. For a variety of Grades. Proof of compliance issued by SITA to be submitted with bid
257 | RT275-05-044 | Assistive Software | Curriculum Activities Software: Science Develops science skills. For a variety of Grades. Proof of compliance issued by SITA to be submitted with bid
258 | RT275-05-045 | Assistive Software | Curriculum Activities Software: Creativity Develops creativity skills. For a viariety of Grades. Proof of compliance issued by SITA to be submitted with bid
259 | RT275-05-046 | Assistive Software | Mind Mapping Literacy Support Software Mind mapping and idea capture software. Supports reading and writing tasks within a word processor, with text-to-speech and advanced mind-mapping to enable users of all abilities to plan and execute writing tasks. Accessible for eye gaze, switch, touch and mouse. Single- and multi-user options. (e.g. DocsPlus) Proof of compliance issued by SITA to be submitted with bid
260 | RT275-05-047 | Assistive Software | Simulation Software for Maths, Science, Technology and Computing Develop various skills. Variety of themes including Maths, Science, Technology and Computing. Allows users with disabilities to create simulations and conduct a range of experiments in a virtual science lab Proof of compliance issued by SITA to be submitted with bid
261 | RT275-05-048 | Assistive Software | Tablet with AAC Software (Therapist) Midrange AAC tablet, preloaded with AAC software, symbol library and communication device template software; for therapist use Proof of compliance issued by SITA to be submitted with bid
262 | RT275-05-049 | Assistive Software | Software to Treat Articulation Speech Sounds on Cue for Windows: Used for treatment of articulation difficulty caused by apraxia of speech or other speech difficulties. Increases the intensity of treatment of people with apraxia and uses the Rosenbek's integral stimulation approach with Hill's carrier phrase facilitation drills Proof of compliance issued by SITA to be submitted with bid
263 | RT275-06-001 | Low Vison Items | White Cane Mobility tool used by people who are visually impaired to detect objects in their path. Collapsible, with roller tip; Various lengths available
264 | RT275-06-002 | Low Vison Items | Head Pointer Head-worn pointer protruding from the chin or forehead to enable efficient pointing and typing. Special tip required for touch screen
265 | RT275-06-003 | Low Vison Items | Money stick Small aluminium measuring tool with step-like markers to measure the different sizes of banknotes.
266 | RT275-06-004 | Low Vison Items | Liquid level indicator Device that can be hooked onto the edge of a cup or glass and will indicate by using sound or vibration when the cup/glass is full.
267 | RT275-06-005 | Low Vison Items | Signature guide Flat guide with a window cut into it to guide the signatory to sign in the correct place.
268 | RT275-06-006 | Low Vison Items | Goal ball Football or basketball-sized ball with a sound-generator inside it to allow it to be tracked via sound
269 | RT275-06-007 | Low Vison Items | Typing Stick For users with limited hand mobility or dexterity. Examples: hand, wrist or mouth stick; For use with keyboard, calculator, telephone and other devices
270 | RT275-06-008 | Low Vison Items | Handheld Optical Magnifier Lens with handle that magnifies objects and text. Integrated lighting; Various fixed magnification levels available
271 | RT275-06-009 | Low Vison Items | Optical Magnifier with Stand Lens with handle and stand that magnifies objects and text. Integrated lighting; Various fixed magnification levels available
272 | RT275-06-010 | Low Vison Items | Optical Reading Magnifier, Bar Bar-shaped lens that magnifies text. Various fixed magnification levels available; integrated reading line
273 | RT275-06-011 | Low Vison Items | Optical Reading Magnifier, Dome Dome-shaped lens that magnifies text. Various fixed magnification levels available
274 | RT275-06-012 | Low Vison Items | Manual Braille Slate Braille slate and stylus for writing braille on paper. Full page A4 size; 27 lines, 30 cells
275 | RT275-06-013 | Low Vison Items | Manual Braille Writer Mechanical or electronic typewriter for typing braille on special braille paper. Six-key braille keyboard; Paper size up to 11” x 11.5”"
276 | RT275-06-014 | Low Vison Items | Braille Paper, single sheet Single sheet 135g, 255 x 340 mm; 500 sheets per pack
277 | RT275-06-015 | Low Vison Items | Braille Paper, A4 Single sheet 135g A4 (210 x 297 mm); 500 sheets per pack
278 | RT275-06-016 | Low Vison Items | Desktop Digital Magnifier Device with integrated high-definition autofocus camera and 16"+ display. Digitally magnifies and enhances contrast of printed materials. Displays real-time view of document; adjustable magnification up to 50X; colour contrast options; allows user to write under the camera. Proof of compliance issued by SITA to be submitted with bid Proof of compliance issued by ICASA to be submitted with bid
279 | RT275-06-017 | Low Vison Items | Advanced Desktop Digital Magnifier Device with integrated high-definition autofocus camera and 20"+ display. Digitally magnifies and enhances contrast of printed materials. Displays real-time view of document; adjustable magnification up to 60X; colour contrast options; allows user to write under the camera. Proof of compliance issued by SITA to be submitted with bid
280 | RT275-06-018 | Low Vison Items | Desktop Digital Magnifier with Text to Speech Device with integrated camera and 16"+ display that digitally magnifies and enhances contrast of printed materials. High-definition autofocus camera; Displays real-time view of document; Adjustable magnification up to 50X; Colour contrast options; Built-in XY table; Allows user to write under the camera; text to speech output Proof of compliance issued by SITA to be submitted with bid
281 | RT275-06-019 | Low Vison Items | Desktop Digital Magnifier with Text to Speech - 20" Device with integrated camera and 20"+ display that digitally magnifies and enhances contrast of printed materials. High-definition autofocus camera; Displays real-time view of document; Adjustable magnification up to 60X; Colour contrast options; Built-in XY table; Allows user to write under the camera; text to speech output Proof of compliance issued by SITA to be submitted with bid
282 | RT275-06-020 | Low Vison Items | Desktop Document Magnifier/Camera Rotatable high-definition autofocus camera on mount, allowing magnification, distance viewing or webcam functionality; Displays real-time view of document; Adjustable magnification level, up to 40X; Colour contrast options; Allows user to write under the camera; output to external monitor Proof of compliance issued by SITA to be submitted with bid
283 | RT275-06-021 | Low Vison Items | Desktop/Portable Digital Magnifier, battery-powered Device with integrated camera and 15"+ display that digitally magnifies and enhances contrast of printed materials. High-definition autofocus camera; Displays real-time view of document; Adjustable magnification level up to 30X; Colour contrast options; Allows user to write under the camera; Battery-powered Proof of compliance issued by SITA to be submitted with bid
284 | RT275-06-022 | Low Vison Items | Portable Digital Magnifier Performs digital magnification and contrast of printed materials with output to computer or monitor; Suitable for long periods of reading; High definition autofocus camera; At least 30x magnification; Distance viewing, document reading and self-view; Colour contrast options; Supports writing under the camera; Optional XY table available for certain models; Compatible with most popular screen magnification software; Optional speech feedback; USB and HDMI outputs. Proof of compliance issued by SITA to be submitted with bid
285 | RT275-06-023 | Low Vison Items | Portable Magnifier with Speech Performs digital magnification and contrast of printed materials with output to computer, monitor or tablet; High-definition autofocus camera for self-viewing, reading or distance viewing (3-in-1 camera); Magnification up to 30x; Displays real-time view of document; Colour contrast options; Supports writing under the camera; Full Page Text-to-Speech (OCR); Multiple language recognition; Male & female premium voices; Compatible with most popular screen magnification software; Lightweight and foldable; Optional suction base; USB, HDMI and Wi-Fi outputs; Battery operated - rechargeable and user-replaceable batteries included Proof of compliance issued by SITA to be submitted with bid
286 | RT275-06-024 | Low Vison Items | Handheld Magnifier, 4" Performs digital magnification and contrast of printed materials; Built-in camera and 4" display; High definition autofocus camera; Up to 10x magnification; LED lighting; Colour contrast options; Freeze image function; Handle and optional reading stand; Lightweight; Rechargeable battery Proof of compliance issued by SITA to be submitted with bid
287 | RT275-06-025 | Low Vison Items | Handheld Magnifier with Stand, 5" Performs digital magnification and contrast of printed materials; Built-in camera and 5" display; High definition autofocus camera; Up to 12x magnification; LED lighting; Colour contrast options; Freeze image function; Handle and built-in reading stand; Lightweight; Rechargeable battery Proof of compliance issued by SITA to be submitted with bid
288 | RT275-06-026 | Low Vison Items | Handheld Magnifier with Stand, 7" Performs digital magnification and contrast of printed materials; Built-in camera and 7" display; High definition autofocus camera; Up to 18x magnification; LED lighting; Colour contrast options; Freeze image function; Built-in reading stand; Distance viewing; Lightweight; HDMI output; Rechargeable battery. Proof of compliance issued by SITA to be submitted with bid
289 | RT275-06-027 | Low Vison Items | Handheld Magnifier with Speech, 10" touch Performs digital magnification and contrast of printed materials; Built-in camera and 7" display; High definition autofocus camera; Up to 18x magnification; LED lighting; Colour contrast options; Freeze image function; Built-in reading stand; Distance viewing; Text to speech; Lightweight; HDMI output; Rechargeable battery Proof of compliance issued by SITA to be submitted with bid
290 | RT275-06-028 | Low Vison Items | Mouse-type Digital Magnifier Performs digital magnification and contrast of printed materials; Support for external display or computer preferred; Adjustable magnification and colour contrast options
291 | RT275-06-029 | Low Vison Items | Handheld Monocular View and magnify objects in the near distance as well as close up; Adjustable magnification and colour contrast options; Large field of view; Rechargeable battery
292 | RT275-06-030 | Low Vison Items | Wearable Digital Magnifier Lightweight head-worn glasses that performs digital magnification and contrast; High-speed autofocus camera, adjustable magnification; Brightness, colour contrast, outline and freeze functions; Rechargeable battery
293 | RT275-06-031 | Low Vison Items | Large Display Calculator Large display and buttons; Supports basic arithmetic operations, percentage, square root, nth power; Memory keys; Battery-operated
294 | RT275-06-032 | Low Vison Items | Advanced Large Display Calculator Large display and buttons; Supports basic arithmetic operations as well as Scientific or Financial functions (separate models); Battery-operated
295 | RT275-06-033 | Low Vison Items | Talking Calculator Calculator with audio and speech output. Supports basic arithmetic operations, percentage, square root, nth power; Memory keys; Adjustable volume; Earphone output; Battery-operated
296 | RT275-06-034 | Low Vison Items | Advanced Talking Calculator Calculator with audio and speech output. Supports basic arithmetic operations as well as Scientific, Financial or Graphing functions (separate models), Adjustable volume; Earphone output; Battery-operated
297 | RT275-06-035 | Low Vison Items | Talking Watch Talking watch for low-vision users. Bold high-contrast hands and numbers; Speak time and date; Alarm setting; Medium and large sizes available
298 | RT275-06-036 | Low Vison Items | Optical Character Recognition (OCR) Device - battery-powered Standalone device that captures hard-copy documents and converts them into electronic text; Reads printed text aloud using text-to-speech; Several male/female voices available; Displays magnified text in high-contrast colours on built-in screen; Portable and lightweight; Rechargeable battery Proof of compliance issued by SITA to be submitted with bid
299 | RT275-06-037 | Low Vison Items | Optical Character Recognition (OCR) Device Standalone device that captures hard-copy documents and converts them into electronic text; Reads printed text aloud using text-to-speech; Several male/female voices available; Connect to external monitor to display magnified text in high-contrast colours; Portable and lightweight Proof of compliance issued by SITA to be submitted with bid
300 | RT275-06-038 | Low Vison Items | Audio Labelling Device Label items with voice memos for later identification. Record a description (voice memo) of the item; The device identifies the item and speaks the recorded memo Proof of compliance issued by SITA to be submitted with bid
301 | RT275-06-039 | Low Vison Items | Navigational Aid for Visually Impaired Wrist-worn navigational aid for increased awareness of obstacles.
302 | RT275-06-040 | Low Vison Items | Cane-Mountable Travel Aid Sensor-based device that detects the presence of obstacles from knee to head height and pre-warns from a distance of up to a few meters. Proof of compliance issued by SITA to be submitted with bid
303 | RT275-06-041 | Low Vison Items | Handheld Talking GPS Device that voices street names, intersections and landmarks when approaching them. Audio feedback as the user walks; Announce location and surrounding areas; Customised route can be set; Route can be reversed Proof of compliance issued by SITA to be submitted with bid
304 | RT275-07-001 | Feeding and Oral Motor | Feeding Bottle: Wide-necked plastic baby bottle, BPA-free - Size: 125-150ml. Bottles must have millilitre markings in addition to total capacity marking. All items must be sterilizable. Where applicable, products must be made from non-toxic material. All products must be PFAS chemical free.
305 | RT275-07-002 | Feeding and Oral Motor | Feeding Bottle: Wide-necked plastic baby bottle, BPA-free - Size: 250ml. Bottles must have millilitre markings in addition to total capacity marking. All items must be sterilizable. Where applicable, products must be made from non-toxic material. All products must be PFAS chemical free.
306 | RT275-07-003 | Feeding and Oral Motor | Feeding Bottle: Wide-necked anti-colic plastic baby bottle, BPA-free -Size: minimum 125ml. Bottles must have millilitre markings in addition to total capacity marking. All items must be sterilizable. Where applicable, products must be made from non-toxic material. All products must be PFAS chemical free.
| RT275-07-004 | Feeding and Oral Motor | Teat for premature infants: 0-6 months For infants weighing less than 2000 g BPA free. Latex or silicone teat shaped to fit oral cavity and to simulate form and function of the mother’s breast when feeding Designed to stimulate and train the sucking reflex. Size is smaller than standard teat for 0-6 months. Item to be compatible with item RT275-07-001, RT275-07-002 and RTRT275-07-003
| RT275-07-005 | Feeding and Oral Motor | Teat for Premature Babies: Silicone orthodontic teat for use with wide-necked bottle. Item to be compatible with item RT275-07-001, RT275-07-002 and RTRT275-07-003
| RT275-07-006 | Feeding and Oral Motor | Teat for Premature Babies: Silicone non-orthodontic teat for use with wide-necked bottle. I Item to be compatible with item RT275-07-001, RT275-07-002 and RTRT275-07-003
| RT275-07-007 | Feeding and Oral Motor | Teat for Premature Babies: Silicone non-orthodontic teat with variable flow rate, for use with wide-necked bottle. Item to be compatible with item RT275-07-001, RT275-07-002 and RTRT275-07-003
| RT275-07-008 | Feeding and Oral Motor | Teat: Silicone orthodontic teat for use with wide-necked bottle. Size: Bidder to quote for all sizes up to 12 months. (e.g. 0-3 month; 3-6 months; 6-12 months/ sizing range offered). Item to be compatible with item RT275-07-001, RT275-07-002 and RTRT275-07-003
| RT275-07-009 | Feeding and Oral Motor | Teat: Silicone non-orthodontic teat for use with wide-necked bottle. Size: Bidder to quote for all sizes up to 12 months. (e.g. 0-3 month; 3-6 months; 6-12 months/ sizing range offered). Item to be compatible with item RT275-07-001, RT275-07-002 and RTRT275-07-003
| RT275-07-031 | Feeding and Oral Motor | Teat: Silicone cross-cut non-orthodontic teat for use with wide-necked bottle. Size: Bidder to quote for all sizes up to 12 months. (e.g. 0-3 month; 3-6 months; 6-12 months/ sizing range offered). Item to be compatible with item RT275-07-001, RT275-07-002 and RTRT275-07-003
| RT275-07-010 | Feeding and Oral Motor | Feeding Bottle: Regular-necked plastic baby bottle, BPA-free -Size: 125-150ml. Bottles must have millilitre markings in addition to total capacity marking. All items must be sterilizable. Where applicable, products must be made from non-toxic material. All products must be PFAS chemical free.
| RT275-07-011 | Feeding and Oral Motor | Feeding Bottle: Regular-necked plastic baby bottle, BPA-free - Size: 250ml. Bottles must have millilitre markings in addition to total capacity marking. All items must be sterilizable. Where applicable, products must be made from non-toxic material. All products must be PFAS chemical free.
| RT275-07-012 | Feeding and Oral Motor | Feeding Bottle: Regular-necked plastic baby bottle, soft squeezable, BPA-free - Size: minimum 125ml. All items must be sterilizable. Where applicable, products must be made from non-toxic material. All products must be PFAS chemical free.
| RT275-07-013 | Feeding and Oral Motor | Feeding Bottle: Regular-necked anti-colic plastic baby bottle, BPA-free - Size: minimum 125ml. Bottles must have millilitre markings in addition to total capacity marking. All items must be sterilizable. Where applicable, products must be made from non-toxic material. All products must be PFAS chemical free.
| RT275-07-014 | Feeding and Oral Motor | Teat for Premature Babies: Silicone orthodontic teat for use with regular-necked bottle. Item to be compatible with item RT275-07-010, RT275-07-011 , RT275-07-012 and RT275-07-013
| RT275-07-015 | Feeding and Oral Motor | Teat for Premature Babies: Silicone non-orthodontic teat for use with regular-necked bottle. Item to be compatible with item RT275-07-010, RT275-07-011 , RT275-07-012 and RTRT275-07-013
| RT275-07-016 | Feeding and Oral Motor | Teat for Premature Babies:Silicone non-orthodontic teat with variable flow rate, for use with regular-necked bottle. Item to be compatible with item RT275-07-010, RT275-07-011 , RT275-07-012 and RTRT275-07-013
| RT275-07-017 | Feeding and Oral Motor | Teat for Premature Babies: Latex orthodontic teat for use with regular-necked bottle. Item to be compatible with item RT275-07-010, RT275-07-011 , RT275-07-012 and RTRT275-07-013
| RT275-07-018 | Feeding and Oral Motor | Teat for Premature Babies: Latex non-orthodontic teat for use with regular-necked bottle. Item to be compatible with item RT275-07-010, RT275-07-011 , RT275-07-012 and RTRT275-07-013
| RT275-07-019 | Feeding and Oral Motor | Teat for Premature Babies: Latex non-orthodontic teat with variable flow rate, for use with regular-necked bottle. Item to be compatible with item RT275-07-010, RT275-07-011 , RT275-07-012 and RTRT275-07-013
| RT275-07-020 | Feeding and Oral Motor | Teat: Silicone orthodontic teat for use with regular-necked bottle. Size: Bidder to quote for all sizes up to 12 months. (e.g. 0-3 month; 3-6 months; 6-12 months/ sizing range offered) . Item to be compatible with item RT275-07-010, RT275-07-011 , RT275-07-012 and RTRT275-07-013
| RT275-07-021 | Feeding and Oral Motor | Teat: Silicone cross-cut orthodontic teat for use with wide-necked bottle. Size: Bidder to quote for all sizes up to 12 months. (e.g. 0-3 month; 3-6 months; 6-12 months/ sizing range offered). Item to be compatible with item RT275-07-010, RT275-07-011 , RT275-07-012 and RTRT275-07-013
| RT275-07-022 | Feeding and Oral Motor | Teat:Silicone cross-cut orthodontic teat for use with regular-necked bottle. Size: Bidder to quote for all sizes up to 12 months. (e.g. 0-3 month; 3-6 months; 6-12 months/ sizing range offered) . Item to be compatible with item RT275-07-010, RT275-07-011 , RT275-07-012 and RTRT275-07-013
| RT275-07-023 | Feeding and Oral Motor | Teat: Silicone non-orthodontic teat for use with regular-necked bottle. Size: Bidder to quote for all sizes up to 12 months. (e.g. 0-3 month; 3-6 months; 6-12 months/ sizing range offered). Item to be compatible with item RT275-07-003, RT275-07-010 , RT275-07-011 ,RT275-07-012 and RT275-07-013
| RT275-07-024 | Feeding and Oral Motor | Teat: Silicone cross-cut non-orthodontic teat for use with regular-necked bottle. Size: Bidder to quote for all sizes up to 12 months. (e.g. 0-3 month; 3-6 months; 6-12 months/ sizing range offered). RT275-07-010 , RT275-07-011 ,RT275-07-12 and RT275-07-013
| RT275-07-025 | Feeding and Oral Motor | Teat: Silicone non-orthodontic teat with variable flow rate, for use with wide-necked bottle. Size: Bidder to quote for all sizes up to 12 months. (e.g. 0-3 month; 3-6 months; 6-12 months/ sizing range offered). RT275-07-010 , RT275-07-011 ,RT275-07-012 and RT275-07-013
| RT275-07-026 | Feeding and Oral Motor | Teat :Silicone non-orthodontic teat with variable flow rate, for use with regular-necked bottle. Size: Bidder to quote for all sizes up to 12 months. (e.g. 0-3 month; 3-6 months; 6-12 months/ sizing range offered). RT275-07-010 , RT275-07-011 ,RT275-07-012 and RT275-07-013
| RT275-07-027 | Feeding and Oral Motor | Teat:Latex orthodontic teat for use with regular-necked bottle. Size: Bidder to quote for all sizes up to 12 months. (e.g. 0-3 month; 3-6 months; 6-12 months/ sizing range offered). RT275-07-010 , RT275-07-011 ,RT275-07-012 and RT275-07-013
| RT275-07-028 | Feeding and Oral Motor | Teat:Latex cross-cut orthodontic teat for use with regular-necked bottle. Size: Bidder to quote for all sizes up to 12 months. (e.g. 0-3 month; 3-6 months; 6-12 months/ sizing range offered). RT275-07-010 , RT275-07-011 ,RT275-07-012 and RT275-07-013
| RT275-07-029 | Feeding and Oral Motor | Teat:Latex non-orthodontic teat for use with regular-necked bottle. Size: Bidder to quote for all sizes up to 12 months. (e.g. 0-3 month; 3-6 months; 6-12 months/ sizing range offered). RT275-07-010 , RT275-07-011 ,RT275-07-012 and RT275-07-013
| RT275-07-030 | Feeding and Oral Motor | Teat:Latex cross-cut non-orthodontic teat for use with regular-necked bottle. Size: Bidder to quote for all sizes up to 12 months. (e.g. 0-3 month; 3-6 months; 6-12 months/ sizing range offered). RT275-07-010 , RT275-07-011 ,RT275-07-012 and RT275-07-013
| RT275-07-031 | Feeding and Oral Motor | Teat: Latex non-orthodontic teat with variable flow rate, for use with regular-necked bottle. Size: Bidder to quote for all sizes up to 12 months. (e.g. 0-3 month; 3-6 months; 6-12 months/ sizing range offered). RT275-07-010 , RT275-07-011 ,RT275-07-012 and RT275-07-013
| RT275-07-032 | Feeding and Oral Motor | Feeding Bottle: teat with 4.cm milk reservoir and 3cm teat Special-needs feeding bottle, with one-way valve and squeezable variable-flow silicone teat with 4.5cm milk reservoir and 3cm teat for premature baby Size: 80ml capacity. Bottles must have millilitre markings in addition to total capacity marking. All items must be sterilizable. Where applicable, products must be made from non-toxic material. All products must be PFAS chemical free.
| RT275-07-033 | Feeding and Oral Motor | Teat: teat with 4.5cm milk reservior and 3cm teat Replacement squeezable variable-flow teat with 4.5cm milk reservoir and 3cm teat for premature baby for special-needs bottle. It must be compatible with item RT275-07-032.
| RT275-07-034 | Feeding and Oral Motor | Feeding Bottle: teat with 4.5cm milk reservoir and 3.5cm teat Special-needs feeding bottle, with one-way valve and squeezable variable-flow silicone teat with 4.5cm milk reservoir and 3.5cm teat for full-term baby. size: 80ml capacity. All items must be sterilizable. Where applicable, products must be made from non-toxic material. All products must be PFAS chemical free.
| RT275-07-035 | Feeding and Oral Motor | Teat: teat with 4.5cm milk reservoir and 3.5cm teat Replacement squeezable variable-flow teat with 4.5cm milk reservoir and 3.5cm teat for full-term baby for special-needs bottle. It must be compatible with item RT275-07-034
| RT275-07-036 | Feeding and Oral Motor | Orthodontic teat for cleft palate. Specify bottle item is compatible with. Teat shaped to fit oral cavity and to simulate form and function of the mother’s breast when feeding The teat is modified to cover the cleft and facilitate feeding in a baby with a cleft. Teat must come with the specialized ring to fit on standard bottle if appropriate. Bidder must submit compatible bottle, ring and teat.
| RT275-07-037 | Feeding and Oral Motor | Teat for cleft lip. The base of the teat is modified to compensate for the cleft of the lip during feeding. Latex teat shaped to fit oral cavity and to simulate form and function of the mother’s breast when feeding. Teat must come with the specialized ring to fit on standard bottle if appropriate.
| RT275-07-038 | Feeding and Oral Motor | Cleft lip/palate nurser Made of low density polyethylene Non-Disposable and sterilizable Flattened oval shaped bottle Pliable bottle that aids milk flow Cross cut feeding hole on teat Size Capacity of minimum 125ml with 15ml graduations
| RT275-07-039 | Feeding and Oral Motor | Special feeder for pre-term babies with feeding difficulties Silicone teat Polypropylene bottles and valves Teat separated from bottle by 1 way valve Washable, re-usable, can be sterilised Consists of bottle; mini teat; collar; valve membrane and disc Dimensions: minimum 80ml bottle
| RT275-07-040 | Feeding and Oral Motor | System for Supplemental Nursing System used for feeding babies with supplement or expressed breast milk at the breast. Simulates experience of breastfeeding Used for premature babies, weak and slow weight gaining infants, building milk supply, relactation and adopted children. Consists of the following parts: Neck Cord; Bottle; Three appropriately sized tubing; Valve holder; Attachment ring; Cover for the bottle to prevent spillage and for storage Two rolls of hypoallergenic paper tape Dimensions: minimum 125 ml bottle
| RT275-07-041 | Feeding and Oral Motor | Pacifiers - premature soother For infants weighing less than 2000g Teat shaped to fit oral cavity and to simulate form and function of the mother’s breast when feeding Made of latex Designed to stimulate and train the sucking reflex under therapist supervision Size is smaller than standard teat for 0-6 months
| RT275-07-042 | Feeding and Oral Motor | Orthodontic Pacifiers for infants: 0-6 months: Made of Latex Designed to stimulate and train the sucking reflex Made of latex Size: 0-6 months
| RT275-07-043 | Feeding and Oral Motor | Orthodontic Pacifiers for infants: 0-6 months - Made of Silicone Designed to stimulate and train the sucking reflex Made of silicone Size: 0-6 months
| RT275-07-044 | Feeding and Oral Motor | Orthodontic Pacifiers for infants: 6-12 months - Made of Latex Designed to stimulate and train the sucking reflex Made of latex Size: 6-12 months
| RT275-07-045 | Feeding and Oral Motor | Orthodontic Pacifiers for infants: 6-12 months - Made of Silicone Designed to stimulate and train the sucking reflex Made of silicone Size: 6-12 months
| RT275-07-046 | Feeding and Oral Motor | Pacifiers for infants cherry shaped teat: Made of latex. Size: 0-6 months
| RT275-07-047 | Feeding and Oral Motor | Pacifiers for infants cherry shaped teat: Made of silicone. Size: 0-6 months
| RT275-07-048 | Feeding and Oral Motor | Pacifiers for infants cherry shaped teat: Made of latex. Size: 6-12 months
| RT275-07-049 | Feeding and Oral Motor | Pacifiers for infants cherry shaped teat : Made of silicone. Size: 6-12 months
| RT275-07-050 | Feeding and Oral Motor | Peristaltic Teat - vented, soft silicone with matt surface for use with wide neck or narrow neck bottle. Super small hole, small hole, medium hole (Size: Quote for all sizes)
| RT275-07-052 | Feeding and Oral Motor | Feeding Cup: Medium Cup for cup-feeding babies: minimum 60ml capacity graduated translucent plastic cup with flange
| RT275-07-053 | Feeding and Oral Motor | Nipple shield: Thin Nipple shield: thin silicone sheet with perforated dome to fit over nipple, to protect the nipple and assist with latching
| RT275-07-054 | Feeding and Oral Motor | Nipple shield: Medium dome: 19-22mm dome length Nipple shield with medium dome: thin silicone sheet with perforated dome to fit over nipple, to protect the nipple and assist with latching: 19-22mm dome length
| RT275-07-055 | Feeding and Oral Motor | Nipple shield: Large dome: 23-26mm dome length Nipple shield with large dome: thin silicone sheet with perforated dome to fit over nipple, to protect the nipple and assist with latching: 23-26mm dome length
| RT275-07-056 | Feeding and Oral Motor | Nipple shield:Extra-large dome: 27mm dome length or longer Nipple shield with extra-large dome: thin silicone sheet with perforated dome to fit over nipple, to protect the nipple and assist with latching: 27mm dome length or longer
| RT275-07-057 | Feeding and Oral Motor | Nipple sizer for determining nipple shield size. Measures size (in mm) of the nipple to determine size of the dome and size of the flange.
| RT275-07-058 | Feeding and Oral Motor | Nipple Correctors - Electric
| RT275-07-059 | Feeding and Oral Motor | Nipple Correctors - Manual
| RT275-07-060 | Feeding and Oral Motor | Vibrator and Oral Foam Swabs Hand-held vibrator with attached replaceable toothettes, battery-operated Provides a technique for promoting awareness of the oral cavity, to relax "fixing" and to aid in the diagnosis of oral sensitivity Packaging: Set including 20 toothettes. Item must be compatible with item RT275-07-061
| RT275-07-061 | Feeding and Oral Motor | Oral Foam Swabs For increasing or decreasing oral sensitivity prior to feeding, oral exercise, or speech sound production work Foam swabsticks for attachment to hand-held vibrator, soft, secure foam heads with distinct ridges. Oral swabs must be compatible with the vibrator. Packaging: 250 per pack Item must be compatible with item RT275-07-060
| RT275-07-062 | Feeding and Oral Motor | Intraoral Massager Safe tool for initial oral-motor therapy and stimulation that can be used inside and outside the mouth on babies, children and adults to improve oral-motor awareness Soft, nubby brush tips can stimulate, massage, or desensitise gums, lips, tongues, and cheeks that would in turn assist sucking and swallowing for feeding, reduce hypersensitivity and ease oral defensiveness in individuals who do not like the feel of food, toothbrushes, straws, and other objects in their mouths Thick easy to grip handle, textured tip with soft rubbery flanges
| RT275-07-063 | Feeding and Oral Motor | Finger intraoral massager Finger cuff that fits on adult finger to teach motor planning for chewing utilising the phasic bite Used to stimulate lateral tongue reflex and oral sensory awareness Made of soft rubber, latex free
| RT275-07-064 | Feeding and Oral Motor | Sensory Bean Bags To be used as a means to benefit patients with tactile defensive behaviour To be used with patients who are hypersensitive to touch on or near the face as well as to stimulate patients who are hyposensitive in the facial area To normalise sensitivity on body and face by introducing varying textures in a controlled way Bean bags of varying textures including silky, furry, bumpy and scratchy.
| RT275-07-065 | Feeding and Oral Motor | Tubes for jaw closure and lip rounding To be used in a graduated program to teach volitional jaw closure and lip rounding which is beneficial to then progress to straw drinking and speech development, particularly for use with patients with cerebral palsy or traumatic brain injury Plastic tubes Packaging: Set of varied sizes
| RT275-07-066 | Feeding and Oral Motor | Spirometer Use to target abdominal grading and controlled airflow Does not produce sound therefore is a good tool for patients who exhibit auditory defensiveness For patients with velo-pharyngeal incompetence and for those who need visual cues To help build intraoral pressure and controlled airflow for adults as well as children which then results in longer utterances, increased breath support and coordination of tongue retraction with open mouth posture required for vowel sounds Dial to adjust level of difficulty, includes two mouthpieces: one flat and one round, includes nose clip set.
| RT275-07-067 | Feeding and Oral Motor | Nasometer Adevice that measures the ratio of nasal and oral airflow
| RT275-07-068 | Feeding and Oral Motor | Sensory Oral Vibration Tool Soothing vibrating tool that teaches individuals how to lateralise their tingue, elevate their tongue tip, make a tongue bowl, close their lips, learn a rotatory chew, position the articulators for sound production, and various other oral motor skills Handheld vibrator with 30+ interchangeable tips also having the added bonus of a chewable loop attached to the plastic handle. For those patients who need vibration and sensory stimulation to transition them into working on biting and chewing skills with the textured tip and chewable loop Used for patients who need tactile sensory input needed to develop speech and feeding skills Chewable loop handle, provided with one textured chewable tip and additional tip. Standard AAA battery included.
| RT275-07-069 | Feeding and Oral Motor | Tip attachment for oral vibratory tool To provide tactile input to the gums, palate, lips, cheeks, and tongue rectangular shape with three different surfaces: bumpy, striated, and smooth. Packaging: 3 per set. This item must be compatible with item RT275-07-68
| RT275-07-070 | Feeding and Oral Motor | Tip attachment for oral vibratory tool To provide tactile input to the gums, palate, lips, cheeks, and tongue rectangular shape with three different surfaces: bumpy, striated, and smooth, softer sensory friendly material. Packaging: 3 per set. This item must be compatible with item RT275-07-68
| RT275-07-071 | Feeding and Oral Motor | Tip attachment for oral vibratory tool Specifically for babies' mouths To provide tactile input to the gums, palate, lips, cheeks, and tongue rectangular shape with three different surfaces: bumpy, striated, and smooth Packaging: 3 per set. This item must be compatible with item RT275-07-68
| RT275-07-072 | Feeding and Oral Motor | Tip attachment for oral vibratory tool Specifically for babies' mouths To provide tactile input to the gums, palate, lips, cheeks, and tongue rectangular shape with three different surfaces: bumpy, striated, and smooth, softer sensory friendly material. Packaging: 3 per set. This item must be compatible with item RT275-07-68
| RT275-07-073 | Feeding and Oral Motor | Tip attachment for oral vibratory tool Allows for a "rolling" action to increase awareness to the cheeks, lips etc cylindrical with long ridges along its circumference. Packaging: each. This item must be compatible with item RT275-07-68
| RT275-07-074 | Feeding and Oral Motor | Tip attachment for oral vibratory tool Allows for a "rolling" action to increase awareness to the cheeks, lips etc cylindrical with long ridges along its circumference, softer sensory material. Packaging: each This item must be compatible with item RT275-07-68
| RT275-07-075 | Feeding and Oral Motor | Spoon tip attachment for oral vibratory tool Sensory feeding tool and oral stimulation Allows for oral awareness, stimulation to the lips, tongue, cheeks and jaw Shallow bowl of spoon, harder material . Packaging: each. This item must be compatible with item RT275-07-68
| RT275-07-076 | Feeding and Oral Motor | Spoon tip attachment for oral vibratory tool Sensory feeding tool and oral stimulation Allows for oral awareness, stimulation to the lips, tongue, cheeks and jaw For use in patients with a bite reflex and /or sensory issues Shallow bowl of spoon, softer material. Packaging: each. This item must be compatible with item RT275-07-68
| RT275-07-077 | Feeding and Oral Motor | Spoon tip attachment for oral vibratory tool Sensory feeding tool and oral stimulation Allows for oral awareness, stimulation to the lips, tongue, cheeks and jaw for patients with food aversions and provide additional tongue stability Shallow bowl of spoon, harder textured material. Packaging: each. This item must be compatible with item RT275-07-68
| RT275-07-078 | Feeding and Oral Motor | Spoon tip attachment for oral vibratory tool Sensory feeding tool and oral stimulation Allows for oral awareness, stimulation to the lips, tongue, cheeks and jaw for patients with food aversions and provide additional tongue stability Shallow bowl of spoon, softer textured material Packaging: each. This item must be compatible with item RT275-07-68
| RT275-07-079 | Feeding and Oral Motor | Tip attachment for oral vibratory tool Allows for targeting of biting and chewing skills including a sustained bite, rhythmic chewing, and jaw strength and stability Tip ends are smooth, round and chewable Packaging: each. This item must be compatible with item RT275-07-68
| RT275-07-080 | Feeding and Oral Motor | Tip attachment for oral vibratory tool Allows for targeting of biting and chewing skills including a sustained bite, rhythmic chewing, and jaw strength and stability Able to reach to back molar area Tip ends are smooth, round and chewable, extra length Packaging: each. This item must be compatible with item RT275-07-68
| RT275-07-081 | Feeding and Oral Motor | Tip attachment for oral vibratory tool Allows for targeting of biting and chewing skills including a sustained bite, rhythmic chewing, and jaw strength and stability Able to reach to back molar area Tip ends are smooth, round and chewable, extra length, softer sensory material Packaging: Each. This item must be compatible with item RT275-07-68
| RT275-07-082 | Feeding and Oral Motor | Tip attachment for oral vibratory tool Allows for targeting of biting and chewing skills including a sustained bite, rhythmic chewing, and jaw strength and stability Able to reach to back molar area Allows for more textured input Tip ends are textured, round and chewable, extra length Packaging: each. This item must be compatible with item RT275-07-68
| RT275-07-083 | Feeding and Oral Motor | Tip attachment for oral vibratory tool Allows for targeting of biting and chewing skills including a sustained bite, rhythmic chewing, and jaw strength and stability Able to fill tube with soft, pureed foods for targeted functional chewing practice, to establish the concept of biting/chewing to get food, and to work on motor planning Tip ends are smooth, round and chewable, extra length with a hollow tube, soft material Packaging: each. This item must be compatible with item RT275-07-68
| RT275-07-084 | Feeding and Oral Motor | Tip attachment for oral vibratory tool Allows for targeting of biting and chewing skills including a sustained bite, rhythmic chewing, and jaw strength and stability Able to fill tube with soft, pureed foods for targeted functional chewing practice, to establish the concept of biting/chewing to get food, and to work on motor planning Tip ends are smooth, round and chewable, extra length with a hollow tube, soft textured material Packaging: each. This item must be compatible with item RT275-07-68
| RT275-07-085 | Feeding and Oral Motor | Tip attachment for oral vibratory tool Allows for targeting of biting and chewing skills including a sustained bite, rhythmic chewing, and jaw strength and stability Able to fill tube with soft, pureed foods for targeted functional chewing practice, to establish the concept of biting/chewing to get food, and to work on motor planning Tip ends are smooth, round and chewable, extra length with a hollow tube, hard textured material Packaging: each. This item must be compatible with item RT275-07-68
| RT275-07-086 | Feeding and Oral Motor | Tip attachment for oral vibratory tool Allows for sensory toothbrushing Allows for tactile and proprioceptive input to the tongue, cheeks, gums, palate and lips Firmer bristles Packaging: each. This item must be compatible with item RT275-07-68
| RT275-07-087 | Feeding and Oral Motor | Tip attachment for oral vibratory tool Allows for sensory toothbrushing Allows for tactile and proprioceptive input to the tongue, cheeks, gums, palate and lips Pliable bristles Packaging: each. This item must be compatible with item RT275-07-68
| RT275-07-088 | Feeding and Oral Motor | Tip attachment for oral vibratory tool Allows for targeting of biting and chewing skills, oral awareness, jaw grading and stability Hard resilient material, including 'dog', 'mouse' and 'cat' shapes Packaging: 3 per set including dog, mouse and cat shapes. This item must be compatible with item RT275-07-68
| RT275-07-089 | Feeding and Oral Motor | Tip attachment for oral vibratory tool Allows for targeting of biting and chewing skills, oral awareness, jaw grading and stability Softer material, including 'dog', 'mouse' and 'cat' shapes Packaging: 3 per set including dog, mouse and cat shapes. This item must be compatible with item RT275-07-68
| RT275-07-090 | Feeding and Oral Motor | Tip attachment for oral vibratory tool Allows for isolation of specific spots within the oral cavity Fine narrow tip Packaging: each. This item must be compatible with item RT275-07-68
| RT275-07-091 | Feeding and Oral Motor | Tip attachment for oral vibratory tool Allows for isolation of specific spots within the oral cavity Fine narrow tip, softer sensory material Packaging: each. This item must be compatible with item RT275-07-68
| RT275-07-092 | Feeding and Oral Motor | Adapter tip attachment for oral vibratory tool Allows for the insertion of lollipops, foam tips to be used with vibration To target oral hygiene, lip strength, oral motor strength Packaging: each. This item must be compatible with item RT275-07-68
| RT275-07-093 | Feeding and Oral Motor | Tip attachment for oral vibratory tool Allows for soothing and stress/pain-relieving input to muscles, tendons and nerves, joints and trigger points around the oral cavity Allows for proprioceptive input and sensory stimulation outside of the mouth (not for intra-oral exercises) Cylindrical with a built-in stainless steel ball Packaging: each. This item must be compatible with item RT275-07-68
| RT275-07-094 | Feeding and Oral Motor | Oral Vibration Tool and attachment tips - kit Kit consisting of a slow vibrating hand-held tool used to generate different sensations and proprioceptive feedback within the oral cavity. Included 3 tips, including 1) rectangular shape with different surfaces, 2) mini rectangular shape with different surfaces and 3) cylindrical with long ridges along its circumference. To come in a sturdy travel case. Packaging: Each
| RT275-07-095 | Feeding and Oral Motor | Oral Vibration Tool and attachment tips - kit Kit consisting of a slow vibrating hand-held tool used to generate different sensations and proprioceptive feedback within the oral cavity. Included 3 tips made of softer sensory friendly material, including 1) rectangular shape with different surfaces, 2) mini rectangular shape with different surfaces and 3) cylindrical with long ridges along its circumference. To come in a sturdy travel case. Packaging: Each
| RT275-07-096 | Feeding and Oral Motor | Oral Vibration Tool and attachment tips - kit Kit consisting of a slow vibrating hand-held tool used to generate different sensations and proprioceptive feedback within the oral cavity. Included 5 tips made of softer sensory friendly material, including 1) rectangular shape with different surfaces, 2) mini rectangular shape with different surfaces, 3) cylindrical with long ridges along its circumference 4) shallow bowl of spoon, harder material and 5) shallow bowl of spoon, softer material. To come in a sturdy travel case. Packaging: Each
| RT275-07-097 | Feeding and Oral Motor | Oral Vibration Tool and attachment tips - kit Kit consisting of a slow vibrating hand-held tool used to generate different sensations and proprioceptive feedback within the oral cavity. Included 5 tips made of softer sensory friendly material, including 1) rectangular shape with different surfaces, 2) mini rectangular shape with different surfaces, 3) cylindrical with long ridges along its circumference 4) shallow bowl of spoon, harder textured material and 5) shallow bowl of spoon, softer textured material. To come in a sturdy travel case. Packaging: Each
| RT275-07-098 | Feeding and Oral Motor | Oral Vibration Tool and attachment tips - kit Kit consisting of a slow vibrating hand-held tool used to generate different sensations and proprioceptive feedback within the oral cavity. Included 5 tips made of softer sensory friendly material, including 1) rectangular shape with different surfaces, 2) mini rectangular shape with different surfaces, 3) cylindrical with long ridges along its circumference, 4) fine narrow tip, 5) shallow bowl of spoon, softer textured material, 6) smooth, round and chewable, 7) smooth, round and chewable extra length, and 8) soft material mouse shape. To come in a sturdy travel case. Packaging: Each
| RT275-07-099 | Feeding and Oral Motor | Oral Vibration Tool and attachment tips - kit Kit consisting of a slow vibrating hand-held tool used to generate different sensations and proprioceptive feedback within the oral cavity. Included 5 tips made of softer sensory friendly material, including 1) shallow bowl of spoon, harder material, 2) shallow bowl of spoon, softer material, 3) shallow bowl of spoon, harder textured material and 4) shallow bowl of spoon, softer textured material. To come in a sturdy travel case. Packaging: Each
| RT275-07-100 | Feeding and Oral Motor | Oral Motor kit Kit consisting of 1) a horn that sounds off with both inhalation and exhalation, 2) set of 8 bean shaped bags of varying textures including silky, furry, bumpy and scratchy, 3) a drinking cup with two handles and two twist-off recessed lids, made of hard reusable plastic, 4) flavoured tongue depressors, 5) a hand-held vibrator with attached replaceabe toothettes, including batteries and 20 toothettes, 6) buttons with string to target increased lip strength, 7) jaw grading bite tool constructed from latex-free, durable, hard resistent material. 8) jaw grading bite tool constructed from latex-free, durable sensory friendly material, 9) tongue tip lateralisation and elevation tools, 10) set of 4 different kinds of biting tubes of different shapes, hardness and softness, 11) Book - "Oral Placement Therapy for Speech Clarity and Feeding, 12) Book - "The Homework Book", 13) set of 6 platic tubes for jaw closure and lip rounding, 14) jaw exerciser kit, 15) tubes to blow on that can be extended or collapsed, 16) oral motor control and range kit, consisting of a combination of 11 specialised straws to be used in a hierarchial order, made of flexible and durable plastic, 17) Oral airflow device kit, consisting of 12 horns made of durable plastic with some metal parts to be used in a hierarchial order, 18) mini mirror with case, 19) hand puppet, 20) Bottle to cup feed transitioner made from pliable bottle that allows fluid to flow up the straw when squeezed gently, including spill proof lid, made of soft latex free flexible plastics, in a set supplied with one replaceable flexi plastic straw pre-cut to a length suitable for use in the specific bottle, 21) bubble kit. To come in durable toolbox. Packaging: Each
| RT275-07-101 | Feeding and Oral Motor | Oral Motor kit Kit consisting of 1) a horn that sounds off with both inhalation and exhalation, 2) set of 8 bean shaped bags of varying textures including silky, furry, bumpy and scratchy, 3) a drinking cup with two handles and two twist-off recessed lids, made of hard reusable plastic, 4) a hand-held vibrator with attached replaceabe toothettes, including batteries and 20 toothettes, 5) jaw grading bite tool constructed from latex-free, durable, hard resistant material. 6) jaw grading bite tool constructed from latex-free, durable sensory friendly material, 7) set of 4 different kinds of biting tubes of different shapes, hardness and softness, 8) Book - "Oral Placement Therapy for Speech Clarity and Feeding, 9) Book - "The Homework Book", 10) tubes to blow on that can be extended or collapsed, 11) oral motor control and range kit, consisting of a combination of 11 specialised straws to be used in a hierarchial order, made of flexible and durable plastic, 12) Oral airflow device kit, consisting of 12 horns made of durable plastic with some metal parts to be used in a hierarchial order, 13) Bottle to cup feed transitioner made from pliable bottle that allows fluid to flow up the straw when squeezed gently, including spill proof lid, made of soft latex free flexible plastics, in a set supplied with one replaceable flexi plastic straw pre-cut to a length suitable for use in the specific bottle, 14) bubble kit. To come in durable case. Packaging: Each
| RT275-07-102 | Feeding and Oral Motor | Jaw Grading Bite Tools - Set (Hard) To promote symmetrical jaw strength, jaw stability, and grading for patients with weak jaw musculature evidenced during speech sound production or feeding (i. e., reduced mobility, sliding, or jutting) For use with dysarthric apraxic/dyspraxic patients in order to teach jaw position and muscle memory. Items should be constructed from latex-free, durable material.
| RT275-07-103 | Feeding and Oral Motor | Jaw Grading Bite Tools - Set (Soft) To promote symmetrical jaw strength, jaw stability, and grading for patients with weak jaw musculature evidenced during speech sound production or feeding (i. e., reduced mobility, sliding, or jutting) For use with dysarthric apraxic/dyspraxic patients in order to teach jaw position and muscle memory. For patients who need extra sensory input and therefore tolerate a softer more pliable object. Constructed from latex-free, durable, sensory friendly material.
| RT275-07-104 | Feeding and Oral Motor | Bite Tools Set of 4 different biting instruments (e.g. tubes of different shapes, hardness and softness) to improve chewing skills. To teach jaw grading and teach children and adults how to chew foods for improved feeding safety and nutritional intake. To be used in a hierarchy of increasing difficulty for development of jaw stability Latex free and lead free.
| RT275-07-105 | Feeding and Oral Motor | Chewing Tool Provide a resilient, non-food, chewable surface for practicing biting and chewing skills and to develop jaw motion for biting and chewing skills. For use on children and adults as a safe and effective tool for practising biting and chewing skills. For use in treatment of individuals who are developing initial oral motor skills, or rehabilitating the jaw after surgery, trauma or stroke. Made from flexible durable plastic that is latex free and lead free. Available in assorted colours of 4 different shapes and of different hardness and softness.
| RT275-07-106 | Feeding and Oral Motor | Sensory Chewing Tool An oral motor chew tool with three different extensions - one smooth, one ribbed and one bumpy Used to assist in oral exploration to increase movement in cheeks, tongue, lips and jaw The extensions reach the upper and lower back teeth for molar munching.
| RT275-07-107 | Feeding and Oral Motor | Sensory Chewing Tool An oral motor chew tool with three different ends - one with ridges, one with bumps and one with swirls Provide sensory input for oral structures. Made out of soft, chewy, resillient material.
| RT275-07-108 | Feeding and Oral Motor | Sensory Chewing Tool An oral motor chew tool The textures simulate the feel of actual food for realistic non-food chewing practice and also add extra tactile sensation to the lips, cheeks, gums, and tongue to help desensitise the mouth, and decrease oral defensiveness/aversions, and transition individuals with sensory issues from pureed to textured foods; Made out of soft, chewy, resillient material with three different surfaces to chew on: one smooth, one with circular bumps, and one with rectangular bumps.
| RT275-07-109 | Feeding and Oral Motor | Sensory Chewing Tool An oral motor chew tube for added sensory input that is scented: e. g chocolate, vanilla and fruity scents For patients who exert more jaw pressure To exercise the lips, tongue, cheeks and especially the jaw while building oral strength, mobility and control. Made out of soft, chewy, resillient material infused with four different scents.
| RT275-07-110 | Feeding and Oral Motor | Sensory Chewing Tool For babies in the mouthing stage of development Soothing chewy tool that promotes oral exploration, biting and chewing skills while providing calming oral input. Made of a soft, latex-free material with a short handle designed especially for a baby's mouth and hands.
| RT275-07-111 | Feeding and Oral Motor | Sensory Chewing Tool Kit Variety of mouthing tools, that provides opportunity for the mouth to learn through jaw movement, tongue movement and oral exploration. Used to help with oral defensiveness, to decrease food aversions, and/or as a safe and durable alternative to toothgrinding and chewing on hazardous non-food items Used to work on a wide variety of oral motor exercises targeting the improvement of feeding skills. Must have in several different colours, textures, scents and sizes.
| RT275-07-112 | Feeding and Oral Motor | Biomechanical tape for lip-taping: 5m x 5cm roll 5m x 5cm roll of smooth four-way stretch biomechanical tape with long-lasting, hypoallergenic, latex-free skin adhesive. Must be available in beige or black. Adhesive must not leave a residue on the skin. To be used for pre-surgical taping of cleft lips.
| RT275-07-113 | Feeding and Oral Motor | Biomechanical tape for lip-taping: 31m x 5cm 31m x 5cm roll of smooth four-way stretch biomechanical tape with long-lasting, hypoallergenic, latex-free skin adhesive. Must be available in beige or black. Adhesive must not leave a residue on the skin. To be used for pre-surgical taping of cleft lips.
| RT275-07-114 | Feeding and Oral Motor | Dynamic Tape: Tape, rehabilitative, strong elastic/nylon-lycra recoil, 4-way stretch, to absorb load and allow full range of motion, heat activated acrylic adhesive, latex free, +/- 200% elongation from original length, with no rigid end point, breathable and water resistant, supplied in 50mm width x 5-meter roll, available in various colours.
| RT275-07-115 | Feeding and Oral Motor | Specialized feeding spoon for Paediatrics Shallow bowl of spoon Made of soft, flexible silicone /plastic that is durable. Reusable and break-resistant. Size: about 2.5cm in width (bowl of spoon)
| RT275-07-116 | Feeding and Oral Motor | Specialized feeding textured spoon for Paediatric Clients: Made of hard plastic Shallow bowl of spoon Made of hard plastic that is durable.Reusable and break-resistant. Size: 2cm bowl width of spoon and is textured underneath
| RT275-07-117 | Feeding and Oral Motor | Specialized feeding textured spoon for Paediatric Clients: Made of silicone/soft plastic: Size: small bowl Shallow bowl of spoon Made of silicone/soft plastic that is durable. Reusable and break-resistant. Dimensions: Size: small bowl width of spoon and is textured underneath
| RT275-07-118 | Feeding and Oral Motor | Specialized feeding spoon for Paediatric Clients: Made of hard plastic: Size 1.5cm bowl width of spoon Shallow bowl of spoon Made of hard plastic that is durable. Reusable and break-resistant. Size: 1.5cm bowl width of spoon
| RT275-07-119 | Feeding and Oral Motor | Specialized feeding spoon for Paediatric Clients (Infants aged 4 months and older) Shallow bowl of spoon that is coated with silicone. Made of metal. Size: about 2.5cm in width (bowl of spoon)
| RT275-07-120 | Feeding and Oral Motor | Specialized feeding spoon for Paediatric Clients: Angles easy grip spoon: Shallow bowl For children with limited fine motor control Angled, easy grip spoon Shallow bowl. Reusable and break-resistant. Made of durable hard plastic.
| RT275-07-121 | Feeding and Oral Motor | Specialized feeding fork for Paediatric Clients For children with limited fine motor control Angled, easy grip fork Made of durable hard plastic. Reusable and break-resistant.
| RT275-07-122 | Feeding and Oral Motor | Specialized feeding equipment for Paediatric Clients Made of netting material into which solid food is placed and that can be attached to a hard plastic handle. To encourage safe oral feeding and chewing. Made of durable hard plastic & reusable netting.
| RT275-07-123 | Feeding and Oral Motor | Specialized feeding Cup: hard plastic: Paediatric size Transition feeding cup with hard durable spouted lid without handles Spill proof. Made of durable hard plastic that is transparent. Size: Paediatric
| RT275-07-124 | Feeding and Oral Motor | Specialized feeding Cup: soft plastic without handles Paediatric size Transition feeding cup with soft plastic spouted lid without handles Spill proof. Made of durable hard plastic that is transparent Size: Paediatric
| RT275-07-125 | Feeding and Oral Motor | Specialized feeding Cup: hard durable with handles: Paediatric size Transition feeding cup with hard durable spouted lid with handles Spill proof. Made of durable hard plastic that is transparent. Size: Paediatric
| RT275-07-126 | Feeding and Oral Motor | Specialized feeding Cup: soft plastic with handles: paediatric size Transition feeding cup with soft plastic spouted lid with handles Spill proof. Made of durable hard plastic that is transparent. Size: Paediatric
| RT275-07-127 | Feeding and Oral Motor | Specialized feeding Cup: soft durable spouted lid with handles smooth: Paediatrics Transition feeding cup with soft durable spouted lid with handles smooth Spill proof. Made of durable hard plastic that is transparent Size: For Paediatrics
| RT275-07-128 | Feeding and Oral Motor | Specialized feeding Cup: hard durable spouted lid with handles smooth: size Paediatrics Transition feeding cup with hard durable spouted lid with handles smooth Spill proof Made of durable hard plastic that is transparent Size: For Paediatrics
| RT275-07-129 | Feeding and Oral Motor | Feeding Cup: Pliable: Transparent. Size small paediatric Pliable Plastic cup that can be cut out if needed for paediatric patients that require assisted feeding. Transparent. Size: Small Paediatric
| RT275-07-130 | Feeding and Oral Motor | Feeding Cup: Pliable: Size medium Paediatric Pliable Plastic cup that can be cut out if needed for paediatric patients that require assisted feeding Size: Medium Paediatric
| RT275-07-131 | Feeding and Oral Motor | Feeding Cup: Pliable: Size Large Paediatric Pliable Plastic cup that can be cut out if needed for paediatric patients that require assisted feeding Size: Large Paediatric .
| RT275-07-132 | Feeding and Oral Motor | Advanced Cup Feeder Cup feeder with special one-way valve system Special mouthpiece prevents spillage of breast milk Soft feeding pad. Dimensions: 80ml-250ml bottle with 1ml and 5ml graduations
| RT275-07-133 | Feeding and Oral Motor | Specialised Feeding Spoons: Bowl of spoon measuring approximately 2.5cm width Narrow spoon with shallow bowl that makes it ideal for feeding therapy. Made of ABS plastic. Must be durable, reusable and break-resistant. Appropriate width for working with children with problems such as poor lip closure, oral hypersensitivity, or tongue thrust. For use with babies or children up till 4 years of age Bowl of spoon measuring approximately 2.5cm width
| RT275-07-134 | Feeding and Oral Motor | Specialised Feeding Spoons : Bowl of spoon measuring approximately 2.5cm Spoon with shallow bowl that makes it ideal for feeding therapy. Made of ABS plastic. Must be durable, reusable and break-resistant. For working with older children and adults with problems such as poor lip closure, oral hypersensitivity, or tongue thrust. For use on children or adults who are oral-tactile defensive Allows food to slide off easily. Bowl of spoon measuring approximately 3cm
| RT275-07-135 | Feeding and Oral Motor | Specialised Feeding Spoon Textured bottom that provides increased awareness and stability during feeding Promotes transitioning across food textures Size: Larger than 2cm
| RT275-07-136 | Feeding and Oral Motor | Specialised feeding cup with straw Drinking cup with two handles to keep hands at midline and a recessed lid that encourages tongue retraction and improved lip closure. For teaching transition from bottle feeding to cup drinking. Cup that controls the flow of the fluid to the patient while the therapist can also work on coordination of the cup-in-hand-movement Made of hard reusable durable plastic, a cup that comes with twist-off lid for straw drinking with replaceable straws. Item to run in series with straw cup. It must be compatible with item item RT275-07-137
| RT275-07-137 | Feeding and Oral Motor | Flexible Tubing Flexible straw tubing for use in drinking bottle with spill proof lid Made from latex free, flexible, durable plastic. Must be compatible with bottle to cup feed transitioner. Item to run in series with straw cup. It must be compatible with item item RT275-07-136
| RT275-07-138 | Feeding and Oral Motor | Specialised feeding cup - 70ml soft oval silicone cup with reservoir in the front. Designed for feeding babies with cleft and other feeding difficulties.
| RT275-07-139 | Feeding and Oral Motor | Cut Out Cups Allows a child to drink without head or neck extension Flexible cups with cutouts can be gently squeezed to change the shape of the cup lip Allows liquids to be visible through the plastic Latex free Size: 30ml volume
| RT275-07-140 | Feeding and Oral Motor | Cut Out Cups Allows a child to drink without head or neck extension Flexible cups with cutouts can be gently squeezed to change the shape of the cup lip Allows liquids to be visible through the plastic Latex free Size: 60ml volume
| RT275-07-141 | Feeding and Oral Motor | Cut Out Cups Allows older child or adult to drink without head or neck extension Flexible cups with cutouts can be gently squeezed to change the shape of the cup lip Allows liquids to be visible through the plastic Latex free Size: 200ml volume
| RT275-07-142 | Feeding and Oral Motor | Bottle to cup feed transitioner Allows for the control of the flow of liquid into a child's mouth, from pliable bottle that allows fluid to flow up the straw when squeezed gently Encourages children to learn how to straw drink Used to transition from bottle-feeding to cup drinking Includes spill-proof lid, made of soft latex free flexible plastics Packaging: A set, supplied with one replaceable flexi plastic straw pre-cut to a length suitable for use in the specific bottle
| RT275-07-143 | Feeding and Oral Motor | Lip Block Mouthpiece that is inserted into the top of a standard drinking straw which prevents individuals from biting the straw and/or inserting the straw too far into the mouth. Can be cleaned for use with multiple patients, made of flexible plastic
| RT275-07-144 | Feeding and Oral Motor | Laryngeal Mirror - Kit Mirror for dysphagia intervention. Metal laryngeal mirror with metal handle. Autoclavable.
| RT275-07-145 | Feeding and Oral Motor | supportive positioning chair/device for babies under 10kg to be for used for children in videoflouroscopy. To position and support babies, toddlers, children when conducting swallow studies in radiology. Adjustable height and head support, with safety straps. Must not be radio-opaque, must be washable. Same description for children 10-20kg.
| RT275-08-001 | Laryngectomy | Cleaning brush for voice prosthesis (compatible with prosthesis): 4.5-10mm length For cleaning inside of the voice prosthesis without removing it from the standard te-fistula. Washable and re-usable for several weeks size: brush head standard size, suitable for use in voice prostheses from 4.5-10mm length
| RT275-08-002 | Laryngectomy | Cleaning brush for voice prosthesis (compatible with prosthesis): 12-18mm length For cleaning inside of voice prosthesis without removing it from the te-fistula, washable and re-usable for several weeks size: brush head extra large, from 12-18mm length.
| RT275-08-003 | Laryngectomy | Indwelling voice prosthesis - puncture kit: 8mm, 22.5fr Sterile, 8mm, 22.5fr. Silicone prosthesis with one-way, low airflow resistance flap valve and puncture set, designed for retrograde primary placement within tracheoesophageal fistula, for voice restoration post laryngectomy
| RT275-08-004 | Laryngectomy | Indwelling voice prosthesis - puncture kit: 10mm, 22.5fr Sterile, 10mm, 22.5fr. Silicone prosthesis with one-way, low airflow resistance flap valve and puncture set, designed for retrograde primary placement within tracheoesophageal fistula, for voice restoration post laryngectomy
| RT275-08-005 | Laryngectomy | Indwelling voice prosthesis - puncture kit: 12.5mm, 22.5fr Sterile, 12.5mm, 22.5fr. Silicone prosthesis with one-way, low airflow resistance flap valve and puncture set, designed for retrograde primary placement within tracheoesophageal fistula, for voice restoration post laryngectomy
| RT275-08-006 | Laryngectomy | Indwelling voice prosthesis - puncture kit: 8mm, 20fr Sterile, 8mm, 20fr. Silicone prosthesis with one-way, low airflow resistance flap valve and puncture set, designed for retrograde primary placement within tracheoesophageal fistula, for voice restoration post laryngectomy
| RT275-08-007 | Laryngectomy | Indwelling voice prosthesis - puncture kit: 10mm, 20fr Sterile, 10mm, 20fr. Silicone prosthesis with one-way, low airflow resistance flap valve and puncture set, designed for retrograde primary placement within tracheoesophageal fistula, for voice restoration post laryngectomy
| RT275-08-008 | Laryngectomy | Indwelling voice prosthesis - puncture kit: 12,5mm, 20fr Sterile, 12.5mm, 20fr. Silicone prosthesis with one-way, low airflow resistance flap valve and puncture set, designed for retrograde primary placement within tracheoesophageal fistula, for voice restoration post laryngectomy
| RT275-08-009 | Laryngectomy | Indwelling voice prosthesis - puncture kit: 8mm, 17fr Sterile, 8mm, 17fr. Silicone prosthesis with one-way, low airflow resistance flap valve and puncture set, designed for retrograde primary placement within tracheoesophageal fistula, for voice restoration post laryngectomy
| RT275-08-010 | Laryngectomy | Indwelling voice prosthesis - puncture kit: 10mm, 17fr Sterile, 10mm, 17fr. Silicone prosthesis with one-way, low airflow resistance flap valve and puncture set, designed for retrograde primary placement within tracheoesophageal fistula, for voice restoration post laryngectomy
| RT275-08-011 | Laryngectomy | Indwelling voice prosthesis - puncture kit: 12.5mm, 17fr Sterile, 12.5mm, 17fr. Silicone prosthesis with one-way, low airflow resistance flap valve and puncture set, designed for retrograde primary placement within tracheoesophageal fistula, for voice restoration post laryngectomy
| RT275-08-012 | Laryngectomy | Indwelling voice prosthesis - puncture kit: 15mm, 17fr Sterile, 15mm, 17fr. Silicone prosthesis with one-way, low airflow resistance flap valve and puncture set, designed for retrograde primary placement within tracheoesophageal fistula, for voice restoration post laryngectomy
| RT275-08-013 | Laryngectomy | Indwelling voice prosthesis: 4mm, 22.5fr 4mm, 22.5fr. Silicone prosthesis with one-way, low airflow resistance flap valve, with inserter, designed for anterograde secondary placement within a tracheo-esophageal fistula, for voice restoration post laryngectomy
| RT275-08-014 | Laryngectomy | Indwelling voice prosthesis: 6mm, 22.5fr 6mm, 22.5fr. Silicone prosthesis with one-way, low airflow resistance flap valve, designed for anterograde secondary placement within a tracheo-esophageal fistula, for voice restoration post laryngectomy
| RT275-08-015 | Laryngectomy | Indwelling voice prosthesis: 8mm, 22.5fr 8mm, 22.5fr. Silicone prosthesis with one-way, low airflow resistance flap valve, designed for anterograde secondary placement within a tracheo-esophageal fistula, for voice restoration post laryngectomy
| RT275-08-016 | Laryngectomy | Indwelling voice prosthesis: 10mm, 22.5fr 10mm, 22.5fr. Silicone prosthesis with one-way, low airflow resistance flap valve, designed for anterograde secondary placement within a tracheo-esophageal fistula, for voice restoration post laryngectomy
| RT275-08-017 | Laryngectomy | Indwelling voice prosthesis: 12,5mm,22.5fr 12.5mm, 22.5fr. Silicone prosthesis with one-way, low airflow resistance flap valve, designed for anterograde secondary placement within a tracheo-esophageal fistula, for voice restoration post laryngectomy
| RT275-08-018 | Laryngectomy | Indwelling voice prosthesis: 15mm, 22.5fr. Silicone prosthesis with one-way, low airflow resistance flap valve, designed for anterograde secondary placement within a tracheo-esophageal fistula, for voice restoration post laryngectomy
| RT275-08-019 | Laryngectomy | Indwelling voice prosthesis: 4mm, 20fr. Silicone prosthesis with one-way, low airflow resistance flap valve, designed for anterograde secondary placement within a tracheo-esophageal fistula, for voice restoration post laryngectomy
| RT275-08-020 | Laryngectomy | Indwelling voice prosthesis: 6mm, 20fr. Silicone prosthesis with one-way, low airflow resistance flap valve, designed for anterograde secondary placement within a tracheo-esophageal fistula, for voice restoration post laryngectomy
| RT275-08-021 | Laryngectomy | Indwelling voice prosthesis: 8mm, 20fr. Silicone prosthesis with one-way, low airflow resistance flap valve, designed for anterograde secondary placement within a tracheo-esophageal fistula, for voice restoration post laryngectomy
| RT275-08-022 | Laryngectomy | Indwelling voice prosthesis: 10mm, 20fr. Silicone prosthesis with one-way, low airflow resistance flap valve, designed for anterograde secondary placement within a tracheo-esophageal fistula, for voice restoration post laryngectomy
| RT275-08-023 | Laryngectomy | Indwelling voice prosthesis: 12.5mm, 20fr. Silicone prosthesis with one-way, low airflow resistance flap valve, designed for anterograde secondary placement within a tracheo-esophageal fistula, for voice restoration post laryngectomy
| RT275-08-024 | Laryngectomy | Indwelling voice prosthesis: 15mm, 20fr. Silicone prosthesis with one-way, low airflow resistance flap valve, designed for anterograde secondary placement within a tracheo-esophageal fistula, for voice restoration post laryngectomy
| RT275-08-025 | Laryngectomy | Indwelling voice prosthesis: 4mm, 17fr. Silicone prosthesis with one-way, low airflow resistance flap valve, preloaded inside the inserter, designed for anterograde secondary placement within a tracheo-esophageal fistula, for voice restoration post laryngectomy
| RT275-08-026 | Laryngectomy | Indwelling voice prosthesis: 6mm, 17fr. Silicone prosthesis with one-way, low airflow resistance flap valve, designed for anterograde secondary placement within a tracheo-esophageal fistula, for voice restoration post laryngectomy
| RT275-08-027 | Laryngectomy | Indwelling voice prosthesis: 8mm, 17fr. Silicone prosthesis with one-way, low airflow resistance flap valve, designed for anterograde secondary placement within a tracheo-esophageal fistula, for voice restoration post laryngectomy
| RT275-08-028 | Laryngectomy | Indwelling voice prosthesis: 10mm, 17fr. Silicone prosthesis with one-way, low airflow resistance flap valve, designed for anterograde secondary placement within a tracheo-esophageal fistula, for voice restoration post laryngectomy
| RT275-08-029 | Laryngectomy | Indwelling voice prosthesis: 12.5mm, 17fr. Silicone prosthesis with one-way, low airflow resistance flap valve, designed for anterograde secondary placement within a tracheo-esophageal fistula, for voice restoration post laryngectomy
| RT275-08-030 | Laryngectomy | Indwelling voice prosthesis: 15mm, 17fr. Silicone prosthesis with one-way, low airflow resistance flap valve, designed for anterograde secondary placement within a tracheo-esophageal fistula, for voice restoration post laryngectomy
| RT275-08-031 | Laryngectomy | Non-indwelling voice prosthesis: 4mm, 16fr. Silicone, patient-changeable prosthesis with one-way, low airflow resistance flap valve, designed for anterograde secondary placement within tracheoesophageal fistula, for voice restoration post laryngectomy.
| RT275-08-032 | Laryngectomy | Non-indwelling voice prosthesis: 6mm, 16fr. Silicone, patient-changeable prosthesis with one-way, low airflow resistance flap valve, designed for anterograde secondary placement within tracheoesophageal fistula, for voice restoration post laryngectomy
| RT275-08-033 | Laryngectomy | Non-indwelling voice prosthesis: 8mm, 16fr.Silicone, patient-changeable prosthesis with one-way, low airflow resistance flap valve, designed for anterograde secondary placement within tracheoesophageal fistula, for voice restoration post laryngectomy
| RT275-08-034 | Laryngectomy | Non-indwelling voice prosthesis: 10mm, 16fr. Silicone, patient-changeable prosthesis with one-way, low airflow resistance flap valve, designed for anterograde secondary placement within tracheoesophageal fistula, for voice restoration post laryngectomy
| RT275-08-035 | Laryngectomy | Non-indwelling voice prosthesis: 12mm, 16fr. Silicone, patient-changeable prosthesis with one-way, low airflow resistance flap valve, designed for anterograde secondary placement within tracheoesophageal fistula, for voice restoration post laryngectomy
| RT275-08-036 | Laryngectomy | Non-indwelling voice prosthesis: 14mm, 16fr. Silicone, patient-changeable prosthesis with one-way, low airflow resistance flap valve, designed for anterograde secondary placement within tracheoesophageal fistula, for voice restoration post laryngectomy
| RT275-08-037 | Laryngectomy | Non-indwelling voice prosthesis: 18mm, 16fr. Silicone, patient-changeable prosthesis with one-way, low airflow resistance flap valve, designed for anterograde secondary placement within tracheoesophageal fistula, for voice restoration post laryngectomy
| RT275-08-038 | Laryngectomy | Non-indwelling voice prosthesis: 22mm, 16fr. Silicone, patient-changeable prosthesis with one-way, low airflow resistance flap valve, designed for anterograde secondary placement within tracheoesophageal fistula, for voice restoration post laryngectomy
| RT275-08-039 | Laryngectomy | Non-indwelling voice prosthesis: 25mm, 16fr. Silicone, patient-changeable prosthesis with one-way, low airflow resistance flap valve, designed for anterograde secondary placement within tracheoesophageal fistula, for voice restoration post laryngectomy
| RT275-08-040 | Laryngectomy | Non-indwelling voice prosthesis: 4mm, 17fr. Silicone, patient-changeable prosthesis with one-way, low airflow resistance flap valve, designed for anterograde secondary placement within tracheoesophageal fistula, for voice restoration post laryngectomy
| RT275-08-041 | Laryngectomy | Non-indwelling voice prosthesis: 6mm, 17fr. Silicone, patient-changeable prosthesis with one-way, low airflow resistance flap valve, designed for anterograde secondary placement within tracheoesophageal fistula, for voice restoration post laryngectomy
| RT275-08-042 | Laryngectomy | Non-indwelling voice: 8mm, 17fr. Silicone, patient-changeable prosthesis with one-way, low airflow resistance flap valve, designed for anterograde secondary placement within tracheoesophageal fistula, for voice restoration post laryngectomy
| RT275-08-043 | Laryngectomy | Non-indwelling voice prosthesis: 10mm, 17fr. Silicone, patient-changeable prosthesis with one-way, low airflow resistance flap valve, designed for anterograde secondary placement within tracheoesophageal fistula, for voice restoration post laryngectomy
| RT275-08-044 | Laryngectomy | Non-indwelling voice prosthesis: 12mm, 17fr. Silicone, patient-changeable prosthesis with one-way, low airflow resistance flap valve, designed for anterograde secondary placement within tracheoesophageal fistula, for voice restoration post laryngectomy
| RT275-08-045 | Laryngectomy | Non-indwelling voice prosthesis: 14mm, 17fr. Silicone, patient-changeable prosthesis with one-way, low airflow resistance flap valve, designed for anterograde secondary placement within tracheoesophageal fistula, for voice restoration post laryngectomy
| RT275-08-046 | Laryngectomy | Non-indwelling voice prosthesis: 18mm, 17fr. Silicone, patient-changeable prosthesis with one-way, low airflow resistance flap valve, designed for anterograde secondary placement within tracheoesophageal fistula, for voice restoration post laryngectomy
| RT275-08-047 | Laryngectomy | Non-indwelling voice prosthesis: 22mm, 17fr. Silicone, patient-changeable prosthesis with one-way, low airflow resistance flap valve, designed for anterograde secondary placement within tracheoesophageal fistula, for voice restoration post laryngectomy
| RT275-08-048 | Laryngectomy | Non-indwelling voice prosthesis: 25mm, 17fr. Silicone, patient-changeable prosthesis with one-way, low airflow resistance flap valve, designed for anterograde secondary placement within tracheoesophageal fistula, for voice restoration post laryngectomy
| RT275-08-049 | Laryngectomy | Non-indwelling voice prosthesis: 4mm, 20fr. Silicone, patient-changeable prosthesis with one-way, low airflow resistance flap valve, designed for anterograde secondary placement within tracheoesophageal fistula, for voice restoration post laryngectomy
| RT275-08-050 | Laryngectomy | Non-indwelling voice prosthesis: 6mm, 20fr. Silicone, patient-changeable prosthesis with one-way, low airflow resistance flap valve, designed for anterograde secondary placement within tracheoesophageal fistula, for voice restoration post laryngectomy
| RT275-08-051 | Laryngectomy | Non-indwelling voice prosthesis: 8mm, 20fr. Silicone, patient-changeable prosthesis with one-way, low airflow resistance flap valve, designed for anterograde secondary placement within tracheoesophageal fistula, for voice restoration post laryngectomy
| RT275-08-052 | Laryngectomy | Non-indwelling voice prosthesis: 10mm, 20fr. Silicone, patient-changeable prosthesis with one-way, low airflow resistance flap valve, designed for anterograde secondary placement within tracheoesophageal fistula, for voice restoration post laryngectomy
| RT275-08-053 | Laryngectomy | Non-indwelling voice prosthesis: 12mm, 20fr. Silicone, patient-changeable prosthesis with one-way, low airflow resistance flap valve, designed for anterograde secondary placement within tracheoesophageal fistula, for voice restoration post laryngectomy
| RT275-08-054 | Laryngectomy | Non-indwelling voice prosthesis: 14mm, 20fr. Silicone, patient-changeable prosthesis with one-way, low airflow resistance flap valve, designed for anterograde secondary placement within tracheoesophageal fistula
| RT275-08-055 | Laryngectomy | Non-indwelling voice prosthesis: 18mm, 20fr. Silicone, patient-changeable prosthesis with one-way, low airflow resistance flap valve, designed for anterograde secondary placement within tracheoesophageal fistula, for voice restoration post laryngectomy
| RT275-08-056 | Laryngectomy | Non-indwelling voice prosthesis: 22mm, 20fr. Silicone, patient-changeable prosthesis with one-way, low airflow resistance flap valve, designed for anterograde secondary placement within tracheoesophageal fistula, for voice restoration post laryngectomy
| RT275-08-057 | Laryngectomy | Non-indwelling voice prosthesis: 25mm, 20fr. Silicone, patient-changeable prosthesis with one-way, low airflow resistance flap valve, designed for anterograde secondary placement within tracheoesophageal fistula, for voice restoration post laryngectomy
| RT275-08-058 | Laryngectomy | Non-indwelling voice prosthesis: 6mm, 16fr. Silicone, patient-changeable prosthesis with one-way, slit valve, designed for anterograde secondary placement within tracheoesophageal fistula, for voice restoration post laryngectomy
| RT275-08-059 | Laryngectomy | Non-indwelling voice prosthesis: 8mm, 16fr. Silicone, patient-changeable prosthesis with one-way, slit valve, designed for anterograde secondary placement within tracheoesophageal fistula, for voice restoration post laryngectomy
| RT275-08-060 | Laryngectomy | Non-indwelling voice prosthesis: 10mm, 16fr. Silicone, patient-changeable prosthesis with one-way, slit valve, designed for anterograde secondary placement within tracheoesophageal fistula, for voice restoration post laryngectomy
| RT275-08-061 | Laryngectomy | Non-indwelling voice prosthesis: 12mm, 16fr. Silicone, patient-changeable prosthesis with one-way, slit valve, designed for anterograde secondary placement within tracheoesophageal fistula, for voice restoration post laryngectomy
| RT275-08-062 | Laryngectomy | Non-indwelling voice prosthesis: 14mm, 16fr. Silicone, patient-changeable prosthesis with one-way, slit valve, designed for anterograde secondary placement within tracheoesophageal fistula, for voice restoration post laryngectomy
| RT275-08-063 | Laryngectomy | Non-indwelling voice prosthesis: 18mm, 16fr. Silicone, patient-changeable prosthesis with one-way, slit valve, designed for anterograde secondary placement within tracheoesophageal fistula, for voice restoration post laryngectomy
| RT275-08-064 | Laryngectomy | Indwelling voice prosthesis with enlarged esophageal flange: 4mm, 17fr. Silicone prosthesis with enlarged esophageal flange to manage leakage around voice prosthesis, with one-way, low airflow resistance flap valve, with insertion device, designed for anterograde secondary placement within a tracheo-esophageal fistula, for voice restoration post laryngectomy
| RT275-08-065 | Laryngectomy | Indwelling voice prosthesis with enlarged esophageal flange: 6mm, 17fr. Silicone prosthesis with enlarged esophageal flange to manage leakage around voice prosthesis, with one-way, low airflow resistance flap valve, with insertion device, designed for anterograde secondary placement within a tracheo-esophageal fistula, for voice restoration post laryngectomy
| RT275-08-066 | Laryngectomy | Indwelling voice prosthesis with enlarged esophageal flange: 8mm, 17fr. Silicone prosthesis with enlarged esophageal flange to manage leakage around voice prosthesis, with one-way, low airflow resistance flap valve, with insertion device, designed for anterograde secondary placement within a tracheo-esophageal fistula, for voice restoration post laryngectomy
| RT275-08-067 | Laryngectomy | Indwelling voice prosthesis with enlarged esophageal flange: 10mm, 17fr. Silicone prosthesis with enlarged esophageal flange to manage leakage around voice prosthesis, with one-way, low airflow resistance flap valve, with insertion device, designed for anterograde secondary placement within a tracheo-esophageal fistula, for voice restoration post laryngectomy
| RT275-08-068 | Laryngectomy | Indwelling voice prosthesis with enlarged esophageal flange: 12.5mm, 17fr. Silicone prosthesis with enlarged esophageal flange to manage leakage around voice prosthesis, with one-way, low airflow resistance flap valve, with insertion device, designed for anterograde secondary placement within a tracheo-esophageal fistula, for voice restoration post laryngectomy
| RT275-08-069 | Laryngectomy | Indwelling voice prosthesis with enlarged esophageal flange: 15mm, 17fr. Silicone prosthesis with enlarged esophageal flange to manage leakage around voice prosthesis, with one-way, low airflow resistance flap valve, with insertion device, designed for anterograde secondary placement within a tracheo-esophageal fistula, for voice restoration post laryngectomy
| RT275-08-070 | Laryngectomy | Indwelling voice prosthesis with enlarged esophageal flange: 4mm, 20fr. Silicone prosthesis with enlarged esophageal flange to manage leakage around voice prosthesis, with one-way, low airflow resistance flap valve, with insertion device, designed for anterograde secondary placement within a tracheo-esophageal fistula, for voice restoration post laryngectomy
| RT275-08-071 | Laryngectomy | Indwelling voice prosthesis with enlarged esophageal flange: 6mm, 20fr. Silicone prosthesis with enlarged esophageal flange to manage leakage around voice prosthesis, with one-way, low airflow resistance flap valve, with insertion device, designed for anterograde secondary placement within a tracheo-esophageal fistula, for voice restoration post laryngectomy
| RT275-08-072 | Laryngectomy | Indwelling voice prosthesis with enlarged esophageal flange: 8mm, 20fr. Silicone prosthesis with enlarged esophageal flange to manage leakage around voice prosthesis, with one-way, low airflow resistance flap valve, with insertion device, designed for anterograde secondary placement within a tracheo-esophageal fistula, for voice restoration post laryngectomy
| RT275-08-073 | Laryngectomy | Indwelling voice prosthesis with enlarged esophageal flange 10mm, 20fr. Silicone prosthesis with enlarged esophageal flange to manage leakage around voice prosthesis, with one-way, low airflow resistance flap valve, with insertion device, designed for anterograde secondary placement within a tracheo-esophageal fistula, for voice restoration post laryngectomy
| RT275-08-074 | Laryngectomy | Indwelling voice prosthesis with enlarged esophageal flange 12.5mm, 20fr. Silicone prosthesis with enlarged esophageal flange to manage leakage around voice prosthesis, with one-way, low airflow resistance flap valve, with insertion device, designed for anterograde secondary placement within a tracheo-esophageal fistula, for voice restoration post laryngectomy
| RT275-08-075 | Laryngectomy | Indwelling voice prosthesis with enlarged esophageal flange 15mm, 20fr. Silicone prosthesis with enlarged esophageal flange to manage leakage around voice prosthesis, with one-way, low airflow resistance flap valve, with insertion device, designed for anterograde secondary placement within a tracheo-esophageal fistula, for voice restoration post laryngectomy
| RT275-08-076 | Laryngectomy | Indwelling voice prosthesis with enlarged esophageal flange 4mm, 22.5fr. Silicone prosthesis with enlarged esophageal flange to manage leakage around voice prosthesis, with one-way, low airflow resistance flap valve, with insertion device, designed for anterograde secondary placement within a tracheo-esophageal fistula, for voice restoration post laryngectomy
| RT275-08-077 | Laryngectomy | Indwelling voice prosthesis with enlarged esophageal flange: 6mm, 22.5fr. Silicone prosthesis with enlarged esophageal flange to manage leakage around voice prosthesis, with one-way, low airflow resistance flap valve, with insertion device, designed for anterograde secondary placement within a tracheo-esophageal fistula, for voice restoration post laryngectomy
| RT275-08-078 | Laryngectomy | Indwelling voice prosthesis with enlarged esophageal flange: 8mm, 22.5fr. Silicone prosthesis with enlarged esophageal flange to manage leakage around voice prosthesis, with one-way, low airflow resistance flap valve, with insertion device, designed for anterograde secondary placement within a tracheo-esophageal fistula, for voice restoration post laryngectomy
| RT275-08-079 | Laryngectomy | Indwelling voice prosthesis with enlarged esophageal flange: 10mm, 22.5fr. Silicone prosthesis with enlarged esophageal flange to manage leakage around voice prosthesis, with one-way, low airflow resistance flap valve, with insertion device, designed for anterograde secondary placement within a tracheo-esophageal fistula, for voice restoration post laryngectomy
| RT275-08-080 | Laryngectomy | Indwelling voice prosthesis with enlarged esophageal flange: 12.5mm, 22.5fr 12.5mm, 22.5fr. Silicone prosthesis with enlarged esophageal flange to manage leakage around voice prosthesis, with one-way, low airflow resistance flap valve, with insertion device, designed for anterograde secondary placement within a tracheo-esophageal fistula, for voice restoration post laryngectomy
| RT275-08-081 | Laryngectomy | Indwelling voice prosthesis with enlarged esophageal flange: 15mm, 22.5fr 15mm, 22.5fr. Silicone prosthesis with enlarged esophageal flange to manage leakage around voice prosthesis, with one-way, low airflow resistance flap valve, with insertion device, designed for anterograde secondary placement within a tracheo-esophageal fistula, for voice restoration post laryngectomy
| RT275-08-082 | Laryngectomy | Indwelling voice prosthesis: 4mm, 16Fr. 4mm, 16FrSilicone prosthesis with one-way, low airflow resistance flap valve, designed for anterograde secondary placement within a tracheo-esophageal fistula, for voice restoration post laryngectomy
| RT275-08-083 | Laryngectomy | Indwelling voice prosthesis: 6mm, 16Fr. 6mm, 16Fr Silicone prosthesis with one-way, low airflow resistance flap valve, designed for anterograde secondary placement within a tracheo-esophageal fistula, for voice restoration post laryngectomy
| RT275-08-084 | Laryngectomy | Indwelling voice prosthesis: 8mm, 16Fr. Silicone prosthesis with one-way, low airflow resistance flap valve, designed for anterograde secondary placement within a tracheo-esophageal fistula, for voice restoration post laryngectomy
| RT275-08-085 | Laryngectomy | Indwelling voice prosthesis: 10mm, 16Fr. 10mm, 16Fr Silicone prosthesis with one-way, low airflow resistance flap valve, designed for anterograde secondary placement within a tracheo-esophageal fistula, for voice restoration post laryngectomy
| RT275-08-086 | Laryngectomy | Indwelling voice prosthesis: 12mm, 16Fr. 12mm, 16FrSilicone prosthesis with one-way, low airflow resistance flap valve, designed for anterograde secondary placement within a tracheo-esophageal fistula, for voice restoration post laryngectomy
| RT275-08-087 | Laryngectomy | Indwelling voice prosthesis: 14mm, 16Fr. 14mm, 16Fr Silicone prosthesis with one-way, low airflow resistance flap valve, designed for anterograde secondary placement within a tracheo-esophageal fistula, for voice restoration post laryngectomy
| RT275-08-088 | Laryngectomy | Indwelling voice prosthesis: 16mm, 16Fr. 16mm, 16Fr Silicone prosthesis with one-way, low airflow resistance flap valve, designed for anterograde secondary placement within a tracheo-esophageal fistula, for voice restoration post laryngectomy
| RT275-08-089 | Laryngectomy | Indwelling voice prosthesis: 18mm, 16Fr. 18mm, 16Fr Silicone prosthesis with one-way, low airflow resistance flap valve, designed for anterograde secondary placement within a tracheo-esophageal fistula, for voice restoration post laryngectomy
| RT275-08-090 | Laryngectomy | Indwelling voice prosthesis: 20mm, 16Fr. 20mm, 16Fr Silicone prosthesis with one-way, low airflow resistance flap valve, designed for anterograde secondary placement within a tracheo-esophageal fistula, for voice restoration post laryngectomy
| RT275-08-091 | Laryngectomy | Indwelling voice prosthesis: 12mm, 20Fr. 12mm, 20Fr Silicone prosthesis with one-way, low airflow resistance flap valve, designed for anterograde secondary placement within a tracheo-esophageal fistula, for voice restoration post laryngectomy
| RT275-08-092 | Laryngectomy | Indwelling voice prosthesis: 14mm, 20Fr. 14mm, 20Fr Silicone prosthesis with one-way, low airflow resistance flap valve, designed for anterograde secondary placement within a tracheo-esophageal fistula, for voice restoration post laryngectomy
| RT275-08-093 | Laryngectomy | Indwelling voice prosthesis: 16mm, 20Fr. 16mm, 20Fr Silicone prosthesis with one-way, low airflow resistance flap valve, designed for anterograde secondary placement within a tracheo-esophageal fistula, for voice restoration post laryngectomy
| RT275-08-094 | Laryngectomy | Indwelling voice prosthesis: 18mm, 20Fr. 18mm, 20Fr Silicone prosthesis with one-way, low airflow resistance flap valve, designed for anterograde secondary placement within a tracheo-esophageal fistula, for voice restoration post laryngectomy
| RT275-08-095 | Laryngectomy | Indwelling voice prosthesis: 20mm, 20Fr. 20mm, 20Fr Silicone prosthesis with one-way, low airflow resistance flap valve, designed for anterograde secondary placement within a tracheo-esophageal fistula, for voice restoration post laryngectomy
| RT275-08-096 | Laryngectomy | Indwelling voice prosthesis: 12mm, 20Fr. Silicone prosthesis with one-way, low airflow resistance flap valve, designed for anterograde secondary placement within a tracheo-esophageal fistula, for voice restoration post laryngectomy
| RT275-08-097 | Laryngectomy | Indwelling voice prosthesis : 5mm, 16Fr. Silicone prosthesis with one-way, low airflow resistance flap valve, designed for anterograde secondary placement within a tracheo-esophageal fistula, for voice restoration post laryngectomy
| RT275-08-098 | Laryngectomy | Indwelling voice prosthesis: 7mm, 16Fr. Silicone prosthesis with one-way, low airflow resistance flap valve, designed for anterograde secondary placement within a tracheo-esophageal fistula, for voice restoration post laryngectomy
| RT275-08-099 | Laryngectomy | Indwelling voice prosthesis : 9mm, 16Fr. Silicone prosthesis with one-way, low airflow resistance flap valve, designed for anterograde secondary placement within a tracheo-esophageal fistula, for voice restoration post laryngectomy
| RT275-08-100 | Laryngectomy | Indwelling voice prosthesis: 5mm, 20Fr. 5mm, 20Fr Silicone prosthesis with one-way, low airflow resistance flap valve, designed for anterograde secondary placement within a tracheo-esophageal fistula, for voice restoration post laryngectomy
| RT275-08-101 | Laryngectomy | Indwelling voice prosthesis: 7mm, 20Fr. 7mm, 20Fr Silicone prosthesis with one-way, low airflow resistance flap valve, designed for anterograde secondary placement within a tracheo-esophageal fistula, for voice restoration post laryngectomy
| RT275-08-102 | Laryngectomy | Indwelling voice prosthesis: 9mm, 20Fr 9mm, 20Fr. Silicone prosthesis with one-way, low airflow resistance flap valve, designed for anterograde secondary placement within a tracheo-esophageal fistula, for voice restoration post laryngectomy
| RT275-08-103 | Laryngectomy | Indwelling voice prosthesis with antimicrobial resistance (Silver Oxide): 4mm,20Fr 4mm, 20Fr. Silicone prosthesis with one-way, low airflow resistance flap valve, designed for anterograde secondary placement within a tracheo-esophageal fistula, for voice restoration post laryngectomy. The prosthesis silicone flap valve is embedded with silver oxide as a material preservative
| RT275-08-104 | Laryngectomy | Indwelling voice prosthesis with antimicrobial resistance (Silver Oxide): 6mm, 20Fr 6mm, 20Fr. Silicone prosthesis with one-way, low airflow resistance flap valve, designed for anterograde secondary placement within a tracheo-esophageal fistula, for voice restoration post laryngectomy. The prosthesis silicone flap valve is embedded with silver oxide as a material preservative
| RT275-08-105 | Laryngectomy | Indwelling voice prosthesis with antimicrobial resistance (Silver Oxide): 8mm, 20Fr 8mm, 20Fr. Silicone prosthesis with one-way, low airflow resistance flap valve, designed for anterograde secondary placement within a tracheo-esophageal fistula, for voice restoration post laryngectomy. The prosthesis silicone flap valve is embedded with silver oxide as a material preservative
| RT275-08-106 | Laryngectomy | Indwelling voice prosthesis with antimicrobial resistance (Silver Oxide): 10mm, 20Fr 10mm, 20Fr. Silicone prosthesis with one-way, low airflow resistance flap valve, designed for anterograde secondary placement within a tracheo-esophageal fistula, for voice restoration post laryngectomy. The prosthesis silicone flap valve is embedded with silver oxide as a material preservative
| RT275-08-107 | Laryngectomy | Indwelling voice prosthesis with antimicrobial resistance (Silver Oxide): 12mm, 20Fr 12mm, 20Fr. Silicone prosthesis with one-way, low airflow resistance flap valve, designed for anterograde secondary placement within a tracheo-esophageal fistula, for voice restoration post laryngectomy. The prosthesis silicone flap valve is embedded with silver oxide as a material preservative
| RT275-08-108 | Laryngectomy | Indwelling voice prosthesis with antimicrobial resistance (Silver Oxide): 14mm, 20Fr 14mm, 20Fr. Silicone prosthesis with one-way, low airflow resistance flap valve, designed for anterograde secondary placement within a tracheo-esophageal fistula, for voice restoration post laryngectomy. The prosthesis silicone flap valve is embedded with silver oxide as a material preservative
| RT275-08-109 | Laryngectomy | Indwelling voice prosthesis with antimicrobial resistance (Silver Oxide): 4mm, 16Fr 4mm, 16Fr. Silicone prosthesis with one-way, low airflow resistance flap valve, designed for anterograde secondary placement within a tracheo-esophageal fistula, for voice restoration post laryngectomy. The prosthesis silicone flap valve is embedded with silver oxide as a material preservative
| RT275-08-110 | Laryngectomy | Indwelling voice prosthesis with antimicrobial resistance (Silver Oxide): 6mm, 16Fr 6mm, 16Fr. Silicone prosthesis with one-way, low airflow resistance flap valve, designed for anterograde secondary placement within a tracheo-esophageal fistula, for voice restoration post laryngectomy. The prosthesis silicone flap valve is embedded with silver oxide as a material preservative
| RT275-08-111 | Laryngectomy | Indwelling voice prosthesis with antimicrobial resistance (Silver Oxide): 8mm, 16Fr 8mm, 16Fr. Silicone prosthesis with one-way, low airflow resistance flap valve, designed for anterograde secondary placement within a tracheo-esophageal fistula, for voice restoration post laryngectomy. The prosthesis silicone flap valve is embedded with silver oxide as a material preservative
| RT275-08-112 | Laryngectomy | Indwelling voice prosthesis with antimicrobial resistance (Silver Oxide): 10mm, 16Fr 10mm, 16Fr. Silicone prosthesis with one-way, low airflow resistance flap valve, designed for anterograde secondary placement within a tracheo-esophageal fistula, for voice restoration post laryngectomy. The prosthesis silicone flap valve is embedded with silver oxide as a material preservative
| RT275-08-113 | Laryngectomy | Indwelling voice prosthesis with antimicrobial resistance (Silver Oxide): 12mm,16Fr 12mm, 16Fr. Silicone prosthesis with one-way, low airflow resistance flap valve, designed for anterograde secondary placement within a tracheo-esophageal fistula, for voice restoration post laryngectomy. The prosthesis silicone flap valve is embedded with silver oxide as a material preservative
| RT275-08-114 | Laryngectomy | Indwelling voice prosthesis with antimicrobial resistance (Silver Oxide): 14mm, 16Fr 14mm, 16Fr. Silicone prosthesis with one-way, low airflow resistance flap valve, designed for anterograde secondary placement within a tracheo-esophageal fistula, for voice restoration post laryngectomy. The prosthesis silicone flap valve is embedded with silver oxide as a material preservative
| RT275-08-115 | Laryngectomy | Indwelling voice prosthesis with dual valve and antimicrobial resistance (Silver Oxide): 6mm, 20Fr 6mm, 20Fr. Silicone prosthesis with two, one-way, low airflow resistance flap valves, designed for anterograde secondary placement within a tracheo-esophageal fistula, for voice restoration post laryngectomy. The prosthesis silicone flap valve is embedded with silver oxide as a material preservative. The second valve is intended to prevent leakage after the first valve fails.
| RT275-08-116 | Laryngectomy | Indwelling voice prosthesis with dual valve and antimicrobial resistance (Silver Oxide): 8mm, 20Fr 8mm, 20Fr. Silicone prosthesis with two, one-way, low airflow resistance flap valves, designed for anterograde secondary placement within a tracheo-esophageal fistula, for voice restoration post laryngectomy. The prosthesis silicone flap valve is embedded with silver oxide as a material preservative. The second valve is intended to prevent leakage after the first valve fails.
| RT275-08-117 | Laryngectomy | Indwelling voice prosthesis with dual valve and antimicrobial resistance (Silver Oxide): 10mm, 20Fr 10mm, 20Fr. Silicone prosthesis with two, one-way, low airflow resistance flap valves, designed for anterograde secondary placement within a tracheo-esophageal fistula, for voice restoration post laryngectomy. The prosthesis silicone flap valve is embedded with silver oxide as a material preservative. The second valve is intended to prevent leakage after the first valve fails.
| RT275-08-118 | Laryngectomy | Indwelling voice prosthesis with dual valve and antimicrobial resistance (Silver Oxide): 12mm, 20Fr 12mm, 20Fr. Silicone prosthesis with two, one-way, low airflow resistance flap valves, designed for anterograde secondary placement within a tracheo-esophageal fistula, for voice restoration post laryngectomy. The prosthesis silicone flap valve is embedded with silver oxide as a material preservative. The second valve is intended to prevent leakage after the first valve fails.
| RT275-08-119 | Laryngectomy | Indwelling voice prosthesis with dual valve and antimicrobial resistance (Silver Oxide): 14mm, 20Fr 14mm, 20Fr. Silicone prosthesis with two, one-way, low airflow resistance flap valves, designed for anterograde secondary placement within a tracheo-esophageal fistula, for voice restoration post laryngectomy. The prosthesis silicone flap valve is embedded with silver oxide as a material preservative. The second valve is intended to prevent leakage after the first valve fails.
| RT275-08-120 | Laryngectomy | Indwelling voice prosthesis with dual valve, large flange and antimicrobial resistance (Silver Oxide): 6mm, 20Fr Large Flange 6mm, 20Fr Large Flange. Silicone prosthesis with two, one-way, low airflow resistance flap valves, designed for anterograde secondary placement within a tracheo-esophageal fistula, for voice restoration post laryngectomy. The prosthesis silicone flap valve is embedded with silver oxide as a material preservative. The second valve is intended to prevent leakage after the first valve fails.
| RT275-08-121 | Laryngectomy | Indwelling voice prosthesis with dual valve, large flange and antimicrobial resistance (Silver Oxide): 8mm, 20Fr Large Flange 8mm, 20Fr Large Flange. Silicone prosthesis with two, one-way, low airflow resistance flap valves, designed for anterograde secondary placement within a tracheo-esophageal fistula, for voice restoration post laryngectomy. The prosthesis silicone flap valve is embedded with silver oxide as a material preservative. The second valve is intended to prevent leakage after the first valve fails.
| RT275-08-122 | Laryngectomy | Indwelling voice prosthesis with dual valve, large flange and antimicrobial resistance (Silver Oxide): 10mm, 20Fr Large Flange 10mm, 20Fr Large Flange. Silicone prosthesis with two, one-way, low airflow resistance flap valves, designed for anterograde secondary placement within a tracheo-esophageal fistula, for voice restoration post laryngectomy. The prosthesis silicone flap valve is embedded with silver oxide as a material preservative. The second valve is intended to prevent leakage after the first valve fails.
| RT275-08-123 | Laryngectomy | Indwelling voice prosthesis with dual valve, large flange and antimicrobial resistance (Silver Oxide): 12mm, 20Fr Large Flange 12mm, 20Fr Large Flange. Silicone prosthesis with two, one-way, low airflow resistance flap valves, designed for anterograde secondary placement within a tracheo-esophageal fistula, for voice restoration post laryngectomy. The prosthesis silicone flap valve is embedded with silver oxide as a material preservative. The second valve is intended to prevent leakage after the first valve fails.
| RT275-08-124 | Laryngectomy | Indwelling voice prosthesis with dual valve, large flange and antimicrobial resistance (Silver Oxide) : 14mm, 20Fr Lange Flange 14mm, 20Fr Large Flange. Silicone prosthesis with two, one-way, low airflow resistance flap valves, designed for anterograde secondary placement within a tracheo-esophageal fistula, for voice restoration post laryngectomy. The prosthesis silicone flap valve is embedded with silver oxide as a material preservative. The second valve is intended to prevent leakage after the first valve fails.
| RT275-08-125 | Laryngectomy | Sterile Tracheoesophageal puncture set Set used for secondary placement of a tracheoesophageal prosthesis. The set includes a disposable 16 gauge curved introducer needle with a 14 gauge Teflon sheath, a 66cm silicone catheter, with a tapered reinforced tip attached to a 61cm Teflon coated stainless steel leader cable, and a fenestrated surgical draper with adhesive around the aperture.
| RT275-08-126 | Laryngectomy | Sterile Voice Prosthesis placement surgical kit Sterile, 16Fr voice prosthesis placement kit for primary or secondary puncture. puncture set, designed for retrograde primary placement within tracheoesophageal fistula, for voice restoration post laryngectomy. It must be compatible with item: RT275-08-127, RT275-08-128, RT275-08-129, RT275-08-130.
| RT275-08-127 | Laryngectomy | Sterile Indwelling Voice Prosthesis: 8mm, 16Fr Clinician placed sterile, 8mm, 16Fr. Silicone prosthesis with one-way, airflow resistance flap valve, designed for primary tracheoesophageal puncture and placement within the tracheoesophageal fistula, for voice restoration post laryngectomy. It must be compatible with item: RT275-08-126, RT275-08-128, RT275-08-129, RT275-08-130.
| RT275-08-128 | Laryngectomy | Sterile Indwelling Voice Prosthesis: 10mm, 16Fr Clinician placed sterile, 10mm, 16Fr. Silicone prosthesis with one-way, airflow resistance flap valve, designed for primary tracheoesophageal puncture and placement within the tracheoesopha fistula, for voice restoration post laryngectomy. It must be compatible with item: RT275-08-126, RT275-08-127, RT275-08-129, RT275-08-130
| RT275-08-129 | Laryngectomy | Sterile Indwelling Voice Prosthesis: 12mm, 16Fr Clinician placed sterile, 12mm, 16Fr. Silicone prosthesis with one-way, airflow resistance flap valve, designed for primary tracheoesophageal puncture and placement within the tracheoesophageal fistula, for voice restoration post laryngectomy. It must be compatible with item:RT275-08-126, RT275-08-127, RT275-08-128, RT275-08-130
| RT275-08-130 | Laryngectomy | Sterile Indwelling Voice Prosthesis Clinician placed sterile, 14mm, 16Fr. Silicone prosthesis with one-way, airflow resistance flap valve, designed for primary tracheoesophageal puncture and placement within the tracheoesophageal fistula, for voice restoration post laryngectomy. It must be compatible with item:RT275-08-126, RT275-08-127, RT275-08-128, RT275-08-129
| RT275-08-131 | Laryngectomy | Sterile Voice Prosthesis placement surgical kit Sterile, 20Fr voice prosthesis placement kit for primary or secondary puncture. puncture set, designed for retrograde primary placement within tracheoesophageal fistula, for voice restoration post laryngectomy. It must be compatible with item: RT275-08-131, RT275-08-132, RT275-08-133, RT275-08-134, RT275-08-135
| RT275-08-132 | Laryngectomy | Sterile Indwelling Voice Prosthesis Clinician placed sterile, 8mm, 20Fr. Silicone prosthesis with one-way, airflow resistance flap valve, designed for primary tracheoesophageal puncture and placement within the tracheoesophageal fistula, for voice restoration post laryngectomy. It must be compatible with item:RT275-08-131, RT275-08-132, RT275-08-134, RT275-08-135
| RT275-08-133 | Laryngectomy | Sterile Indwelling Voice Prosthesis Clinician placed sterile, 10mm, 20Fr. Silicone prosthesis with one-way, airflow resistance flap valve, designed for primary tracheoesophageal puncture and placement within the tracheoesophageal fistula, for voice restoration post laryngectomy. It must be compatible with item: RT275-08-88, RT275-08-89, RT275-08-91, RT275-08-92
| RT275-08-134 | Laryngectomy | Sterile Indwelling Voice Prosthesis Clinician placed sterile, 12mm, 20Fr. Silicone prosthesis with one-way, airflow resistance flap valve, designed for primary tracheoesophageal puncture and placement within the tracheoesophageal fistula, for voice restoration post laryngectomy. It must be compatible with item: RT275-07-131, RT275-07-132, RT275-07-133, RT275-07-135
| RT275-08-135 | Laryngectomy | Sterile Indwelling Voice Prosthesis Clinician placed sterile, 14mm, 20Fr. Silicone prosthesis with one-way, airflow resistance flap valve, designed for primary tracheoesophageal puncture and placement within the tracheoesophageal fistula, for voice restoration post laryngectomy. It must be compatible with item: RT275-08-131, RT275-08-132, RT275-08-133, RT275-08-134
| RT275-08-136 | Laryngectomy | Heat moisture exchanger (standard) for inactive patients For post-laryngectomy use; provides for sufficient moisture retention to humidify air breathed in by laryngectomy patients, compatible with adhesive laryngectomy button, and laryngectomy tube.
| RT275-08-137 | Laryngectomy | Heat moisture exchanger (high flow) for active patients For post-laryngectomy use; provides for sufficient moisture retention to humidify air breathed in by laryngectomy patients, compatible with adhesive laryngectomy button, and laryngectomy
| RT275-08-138 | Laryngectomy | Heat moisture exchanger (high moisture) for nightime use For post-laryngectomy use, provides for sufficient moisture retention to humidify air breathed in by laryngectomy patients, compatible with adhesive lary button, and laryngectomy tube.
| RT275-08-139 | Laryngectomy | Handsfree Voice Kit Used for laryngectomised patients to have hands free stoma occlusion and voicing Used in conjunction with voice prosthesis and compatible heat moisture exchanger (HME)
| RT275-08-140 | Laryngectomy | Hands free cassette (light membrane) Compatible with HME, used for hands free speech for laryngectomised patients.
| RT275-08-141 | Laryngectomy | Hands free cassette (medium membrane) Compatible with HME, used for hands free speech for laryngectomised patients.
| RT275-08-142 | Laryngectomy | Hands free cassette (strong membrane) Compatible with HME, used for hands free speech for laryngectomised patients.
| RT275-08-143 | Laryngectomy | Hands free cassette (adjustable membrane strength) Compatible with HME, used for hands free speech for laryngectomised patients.
| RT275-08-144 | Laryngectomy | Heat moisture exchanger Compatible with hands free cassette.
| RT275-08-145 | Laryngectomy | Tracheostomy cleaning brush Soft tipped, angled to accommodate bending angle of inner cannula, soft bristles to remove crusted secretions, size: 4.5-10mm diameter, 21cm length.
| RT275-08-146 | Laryngectomy | Electrolarynx device Battery-operated, neck type device with oral adapter which produces vibrations to allow voicing for laryngectomy patients. Hand-held, rechargeable unit, intra-oral attachment, charging unit and batteries included
| RT275-08-147 | Laryngectomy | Mechanical larynx device With oral tubing, placed over tracheostomy, lung driven (pneumatic voicing), for laryngectomised patients.
| RT275-08-148 | Laryngectomy | Ring laryngectomy tube Soft tube used to keep stoma of laryngectomy patient open. Compatible with HME system or hands-free system. 8/36: inner diameter: 9.5mm, outer diameter: 12mm.
| RT275-08-149 | Laryngectomy | Ring laryngectomy tube Soft tube used to keep stoma of laryngectomy patient open. Compatible with HME system or hands-free system. 8/55 inner diameter: 9.5mm, outer diameter: 12mm.
| RT275-08-150 | Laryngectomy | Ring laryngectomy tube Soft tube used to keep stoma of laryngectomy patient open. Compatible with HME system or hands-free system. 9/36: inner diameter: 10.5mm, outer diameter: 13.5mm.
| RT275-08-151 | Laryngectomy | Ring laryngectomy tube Soft tube used to keep stoma of laryngectomy patient open. Compatible with HME system or hands-free system. 9/55: inner diameter: 10.5mm, outer diameter: 13.5mm.
| RT275-08-152 | Laryngectomy | Ring laryngectomy tube Soft tube used to keep stoma of laryngectomy patient open. Compatible with HME system or hands-free system. 10/36: inner diameter: 12mm, outer diameter: 15mm.
| RT275-08-153 | Laryngectomy | Ring laryngectomy tube Soft tube used to keep stoma of laryngectomy patient open. Compatible with HME system or hands-free system. 10/55: inner diameter 12mm, outer diameter 15mm,
| RT275-08-154 | Laryngectomy | Silicone Sheeting. Thin silicone sheeting that can be used for voice prosthesis fitting to allow for minimal periprosthetic leak. Dimensions: 0.5mm thick.
| RT275-08-155 | Laryngectomy | Sizer measuring device Compatible with voice prosthesis, reusable and sterilisable, allows the clinician to measure the length of the te fistula.
| RT275-08-156 | Laryngectomy | Sizer flanges Compatible with sizer, non sterile measuring flanges that attach to sizer. For single use. Size: standard.
| RT275-08-157 | Laryngectomy | Dilator Compatible with voice prosthesis, increase the diameter of the tracheoesophageal (te) fistula, size: standard.
| RT275-08-158 | Laryngectomy | Flush/pipette Compatible with voice prosthesis, flush is a cleaning device that allows air or fresh water to be flushed into the lumen of the prosthesis, thereby helping to clean debris and crusts from the prosthesis. Size: standard.
| RT275-08-159 | Laryngectomy | Shower aid Compatible with adhesive, laryngectomy button and laryngectomy tube. Allows patient to shower safely without the risk of aspiration, size: standard.
| RT275-08-160 | Laryngectomy | Oval adhesive (strong seal) Base plate adhesive, for deeper stoma, supports speaking hands free.
| RT275-08-161 | Laryngectomy | Oval adhesive (standard) Compatible with HME, standard oval adhesive for flat area around stoma which allows for the use of HME, size: standard.
| RT275-08-162 | Laryngectomy | Oval adhesive (sensitive) Compatible with HME, oval adhesive for patients with sensitive skin which allows for the use of HME, size: standard (sensitive).
| RT275-08-163 | Laryngectomy | Oval adhesive (strong and sensitive) Compatible with HME, strong oval adhesive for irregular neck around stoma which allows for the use of HME, strong and sensitive, size: standard (strong and sensitive)
| RT275-08-164 | Laryngectomy | Round adhesive (strong seal) Base plate adhesive, for flatter stoma, supports speaking hands free.
| RT275-08-165 | Laryngectomy | Round adhesive (standard) Compatible with HME, standard round adhesive, for flat area around stoma which allows for the use of HME, size: standard.
| RT275-08-166 | Laryngectomy | Round adhesive (sensitive) Compatible with HME standard round adhesive, for patients with sensitive skin which allows for the use of HME, size: standard (sensitive).
| RT275-08-167 | Laryngectomy | Round adhesive (sensitive and strong) Compatible with HME, strong round adhesive, for patients with irregular neck around stoma which allows for the use of HME, size: standard (sensitive and strong).
| RT275-08-168 | Laryngectomy | Laryngectomy button, compatible with HME. Dimension: 12/8. Short self-retaining silicone tracheal cannula skin that maintains opening of tracheostoma and acts as holder for HME. Helps prevent shrinking of tracheostoma, dimension: 12/8.
| RT275-08-169 | Laryngectomy | Laryngectomy button, compatible with HME. Dimension: 14/8 Short self-retaining silicone tracheal cannula skin that maintains opening of tracheostoma and acts as holder for HME. Helps prevent shrinking of tracheostoma, dimension: 14/8.
| RT275-08-170 | Laryngectomy | Laryngectomy button, compatible with HME. Dimension: 16/8 Short self-retaining silicone tracheal cannula skin that maintains opening of tracheostoma and acts as holder for HME. Helps prevent shrinking of tracheostoma dimension: 16/8.
| RT275-08-171 | Laryngectomy | Laryngectomy button, compatible with HME. Dimension: 18/8 Short self-retaining silicone tracheal cannula skin that maintains opening of tracheostoma and acts as holder for HME. Helps prevent shrinking of tracheostoma dimension: 18/8.
| RT275-08-172 | Laryngectomy | Laryngectomy button, compatible with HME. Dimension: 12/18 Short self-retaining silicone tracheal cannula skin that maintains opening of tracheostoma and acts as holder for HME. Helps prevent shrinking of tracheostoma, dimension: 12/18.
| RT275-08-173 | Laryngectomy | Laryngectomy button, compatible with HME. Dimension: 14/18 Short self-retaining silicone tracheal cannula skin that maintains opening of tracheostoma and acts as holder for HME. Helps prevent shrinking of tracheostoma, dimension: 14/18.
| RT275-08-174 | Laryngectomy | Laryngectomy button, compatible with HME. Dimension: 16/18 Short self-retaining silicone tracheal cannula skin that maintains opening of tracheostoma and acts as holder for HME. Helps prevent shrinking of tracheostoma, dimension: 16/18
| RT275-08-175 | Laryngectomy | Laryngectomy button, compatible with HME. Dimension: 18/18 Short self-retaining silicone tracheal cannula skin that Maintains opening of tracheostoma and acts as holder for HME. Helps prevent shrinking of tracheostoma, dimension: 18/18
| RT275-08-176 | Laryngectomy | Laryngectomy tube. Dimension: 8/36 Soft, durable, lightweight, silicone tube used to keep the tracheostoma open after a laryngectomy. Does not hold an HME. Size: inner diameter 9.5mm, outer diameter 12mm, length 36mm dimension: 8/36
| RT275-08-177 | Laryngectomy | Laryngectomy tube. Dimension: 8/55 Soft, durable, lightweight, silicone tube used to keep the tracheostoma open after a laryngectomy. Does not hold an HME.Size: inner diameter 9.5mm, outer diameter 12mm, length 55mm dimension: 8/55
| RT275-08-178 | Laryngectomy | Laryngectomy tube. Dimension: 9/36 Soft, durable, lightweight, silicone tube used to keep the tracheostoma open after a laryngectomy. Does not hold an HME. Size: inner diameter 10.5mm, outer diameter 13.5mm, length 36mm dimension: 9/36
| RT275-08-179 | Laryngectomy | Laryngectomy tube. Dimension: 9/55 Soft, durable, lightweight, silicone tube used to keep the tracheostoma open after a laryngectomy. Does not hold an HME. Size: inner diameter 10.5mm, outer diameter 13.5mm, length 55mm dimension: 9/55
| RT275-08-180 | Laryngectomy | Laryngectomy tube.Dimension: 10/36 Soft, durable, lightweight, silicone tube used to keep the tracheostoma open after a laryngectomy. Does not hold an HME. Size: inner diameter 12mm, outer diameter 15mm, length 36mm dimension: 10/36
| RT275-08-181 | Laryngectomy | Laryngectomy tube.Dimension: 10/55 Soft, durable, lightweight, silicone tube used to keep the tracheostoma open after a laryngectomy. Does not hold an HME .Size: inner diameter 12mm, outer diameter 15mm, length 55mm dimension: 10/55
| RT275-08-182 | Laryngectomy | Laryngectomy tube.Dimension:12/36 Soft, durable, lightweight, silicone tube used to keep the tracheostoma open after a laryngectomy. Does not hold an HME. Size: inner diameter 12mm, outer diameter 15mm, length 55mm dimension: 12/36
| RT275-08-183 | Laryngectomy | Laryngectomy tube.Dimension:12/55 Soft, durable, lightweight, silicone tube used to keep the tracheostoma open after a laryngectomy. Does not hold an HME.0 Size: inner diameter 12mm, outer diameter 15mm, length 55mm dimension: 12/55
| RT275-08-184 | Laryngectomy | Unfenestrated laryngectomy tube, compatible with HME. Size: inner diameter 9.5mm, outer diameter 12mm, length 36mm dimension: 8/36 Soft silicone tube that maintains the opening of tracheostoma, compatible with HME or hands free system. Size: inner diameter 9.5mm, outer diameter 12mm, length 36mm dimension: 8/36
| RT275-08-185 | Laryngectomy | Unfenestrated laryngectomy tube, compatible with HME. Size: inner diameter 9.5mm, outer diameter 12mm, length 55mm dimension: 8/55 Soft silicone tube that maintains the opening of tracheostoma, compatible with HME or hands free system. Size: inner diameter 9.5mm, outer diameter 12mm, length 55mm dimension: 8/55
| RT275-08-186 | Laryngectomy | Unfenestrated laryngectomy tube, compatible with HME. Size: inner diameter 10.5mm, outer diameter 13.5mm, length 36mm dimension: 9/36 Soft silicone tube that maintains the Opening of tracheostoma, compatible with HME or hands free system. Size: inner diameter 10.5mm, outer diameter 13.5mm, length 36mm dimension: 9/36
| RT275-08-187 | Laryngectomy | Unfenestrated laryngectomy tube, compatible with HME. Size: inner diameter 10.5mm, outer diameter 13.5mm, length 55mm, dimension: 9/55 Soft silicone tube that maintains the opening of tracheostoma, compatible with HME or hands free system. Size: inner diameter 10.5mm, outer diameter 13.5mm, length 55mm, dimension: 9/55
| RT275-08-188 | Laryngectomy | Unfenestrated laryngectomy tube, compatible with HME. Size: inner diameter 10.5mm, outer diameter 13.5mm, length 55mm dimension: 10/36 Soft silicone tube that maintains the opening of tracheostoma, compatible with HME or hands free system. Size: inner diameter 10.5mm, outer diameter 13.5mm, length 55mm dimension: 10/36
| RT275-08-189 | Laryngectomy | Unfenestrated laryngectomy tube, compatible with HME. Size: inner diameter 10.5mm, outer diameter 13.5mm, length 55mm dimension: 10/55 Soft silicone tube that maintains the opening of tracheostoma, compatible with HME or hands free system. Size: inner diameter 10.5mm, outer diameter 13.5mm, length 55mm dimension: 10/55
| RT275-08-190 | Laryngectomy | Fenestrated laryngectomy tube, compatible with HME. Size: inner diameter 9.5mm, outer diameter 12mm, length 36mm dimension: 8/36 Soft silicone tube that maintains the opening of tracheostoma and allows for simultaneous voicing, compatible with HME or hands free system. Size: inner diameter 9.5mm, outer diameter 12mm, length 36mm dimension: 8/36
| RT275-08-191 | Laryngectomy | Fenestrated laryngectomy tube, compatible with HME. Size: inner diameter 10.5mm, outer diameter 13.5mm, length 36mm dimension: 9/36 Soft silicone tube that maintains the opening of tracheostoma and allows for simultaneous voicing, compatible with HME or hands free system. Size: inner diameter 9.5mm, outer diameter 12mm, length 55mm dimension: 8/55
| RT275-08-192 | Laryngectomy | Fenestrated laryngectomy tube, compatible with HME. Size: inner diameter 10.5mm, outer diameter 13.5mm, length 36mm dimension: 9/36 Soft silicone tube that maintains the opening of tracheostoma and allows for simultaneous voicing, compatible with HME or hands free system. Size: inner diameter 10.5mm, outer diameter 13.5mm, length 36mm dimension: 9/36
| RT275-08-193 | Laryngectomy | Fenestrated laryngectomy tube, compatible with HME. Size: inner diameter 10.5mm, outer diameter 13.5mm, length 55mm dimension: 9/55 Soft silicone tube that maintains the opening of tracheostoma and allows for simultaneous voicing, compatible with HME or hands free system. Size: inner diameter 10.5mm, outer diameter 13.5mm, length 55mm dimension: 9/55
| RT275-08-194 | Laryngectomy | Fenestrated laryngectomy tube, compatible with HME. Size: inner diameter 12mm, outer diameter 15mm, length 36mm dimension: 10/36 Soft silicone tube that maintains the opening of tracheostoma and allows for simultaneous voicing, compatible with HME or hands free system. Size: inner diameter 12mm, outer diameter 15mm, length 36mm dimension: 10/36
| RT275-08-195 | Laryngectomy | Fenestrated laryngectomy tube, compatible with HME. Size: inner diameter 12mm, outer diameter 15mm, length 55mm dimension: 10/55 Soft silicone tube that maintains the opening of tracheostoma and allows for simultaneous voicing, compatible with HME or hands free system. Size: inner diameter 12mm, outer diameter 15mm, length 55mm dimension: 10/55
| RT275-08-196 | Laryngectomy | Fenestrated laryngectomy tube. size: inner diameter 12mm, outer diameter 12mm, length 36mm Soft silicone tube used to keep stoma of laryngectomy patient open, compatible with HME system or hands-free system, size: inner diameter 12mm, outer diameter 12mm, length 36mm
| RT275-08-197 | Laryngectomy | Silicone washer, compatible with voice prosthesis. size 17 fr Thin silicone washer that is placed around voice prosthesis to reduce periprosthetic leakage around, in patients using indwelling voice prosthesis size 17fr
| RT275-08-198 | Laryngectomy | Silicone washer, compatible with voice prosthesis. size 20fr Thin silicone washer that is placed around voice prosthesis to reduce periprosthetic leakage around, in patients using indwelling voice prosthesis size 20fr
| RT275-08-199 | Laryngectomy | Silicone washer, compatible with voice prosthesis. size 22.5 fr Thin silicone washer that is placed around voice prosthesis to reduce per prosthetic leakage around, in patients using indwelling voice prosthesis size 22.5 fr
| RT275-08-200 | Laryngectomy | Silicone washer, compatible with voice prosthesis.size 24fr Thin silicone washer that is placed around voice prosthesis to reduce per prosthetic leakage around, in patients using indwelling voice prosthesis size 24 fr
| RT275-08-201 | Laryngectomy | Prosthesis plug, compatible with voice prosthesis.size 16fr Tool for temporarily stopping leakage through the voice prosthesis, size 16fr
| RT275-08-202 | Laryngectomy | Prosthesis plug, compatible with voice prosthesis. size 17fr Tool for temporarily stopping leakage through the voice prosthesis, size 17fr
| RT275-08-203 | Laryngectomy | Prosthesis plug, compatible with voice prosthesis.20fr Tool for temporarily stopping leakage through the voice prosthesis, size 20fr
| RT275-08-204 | Laryngectomy | Prosthesis plug, compatible with voice prosthesis. 22.5fr Tool for temporarily stopping leakage through the voice prosthesis, size 22.5fr
| RT275-08-205 | Laryngectomy | Prosthesis plug, compatible with prosthesis. .size: 24fr Tool that helps to temporarily stop leakage through voice prosthesis, size: 24fr
| RT275-09-001 | Tracheostomy | Tracheostoma spacer. Size: internal diameter: 8mm, 3mm thick Silicone gel spacer which is round, tacky and laminated. Tool for improving seal around tracheostoma device fittings, improve cushioning and reduce air leakage. Size: internal diameter: 8mm, 3mm thick
| RT275-09-002 | Tracheostomy | Tracheostoma spacer. Size: internal diameter: 12mm, 3mm thick Silicone gel spacer which is round, tacky and laminated. Tool for improving seal around tracheostoma device fittings, improve cushioning and reduce air leakage. Size: internal diameter: 12mm, 3mm thick
| RT275-09-003 | Tracheostomy | Tracheostoma spacer. Size: internal diameter: 8mm, 5mm thick Silicone gel spacer which is round, tacky and laminated. Tool for improving seal around tracheostoma device fittings, improve cushioning and reduce air leakage. Size: internal diameter: 8mm, 5mm thick
| RT275-09-004 | Tracheostomy | Tracheostoma spacer. Size: internal diameter: 12mm, 5mm thick Silicone gel spacer which is round, tacky and laminated. Tool for improving seal around tracheostoma device fittings, improve cushioning and reduce air leakage. Size: internal diameter: 12mm, 5mm thick
| RT275-09-005 | Tracheostomy | Protective tracheostomy / laryngectomy cloth protectors: Size: small with adjustable hook and loop or tie-up fastener Protective tracheostomy cover for filtration and warming up of the respiratory air, washable, made of polyester with filter. Thick, durable and includes a sponge layer which provides humidification and filtration of the incoming air. Size: small with adjustable hook and loop or tie-up fastener
| RT275-09-006 | Tracheostomy | Protective tracheostomy / laryngectomy cloth protectors. Size: medium, with adjustable hook and loop or tie-up fastener Protective tracheostomy cover for filtration and warming up of the respiratory air, washable, made of polyester with filter. Thick, durable and includes a sponge layer which provides humidification and filtration of the incoming air. Size: medium, with adjustable hook and loop or tie-up fastener
| RT275-09-007 | Tracheostomy | Protective tracheostomy / laryngectomy cloth protectors. Size: large with adjustable hook and loop or tie-up fastener Protective tracheostomy cover for filtration and warming up of the respiratory air, washable, made of polyester with filter. Thick, durable and includes a sponge layer which provides humidification and filtration of the incoming air. Size: large with adjustable hook and loop or tie-up fastener
| RT275-09-008 | Tracheostomy | Protective tracheostomy / laryngectomy cloth protectors. Size: small with adjustable velcro fastener Protective tracheostomy cover for filtration and warming up of the respiratory air, washable, made of polyester with filter. Thick, durable and includes a sponge layer which provides humidification and filtration of the incoming air. Size: small with adjustable velcro fastener
| RT275-09-009 | Tracheostomy | Protective tracheostomy / laryngectomy cloth protectors. Size: medium with adjustable velcro fastener Protective tracheostomy cover for filtration and warming up of the respiratory air, washable, made of polyester with filter. Thick, durable and includes a sponge layer which provides humidification and filtration of the incoming air. Size: medium with adjustable velcro fastener
| RT275-09-010 | Tracheostomy | Protective tracheostomy / laryngectomy cloth protectors. Size: large with adjustable velcro fastener Protective tracheostomy cover for filtration and warming up of the respiratory air, washable, made of polyester with filter. Thick, durable and includes a sponge layer which provides humidification and filtration of the incoming air. Size: large with adjustable velcro fastener
| RT275-09-011 | Tracheostomy | Tracheostomy tube neck strap: Size Adult Broad cotton cover with velcro fastener to hold tracheostomy tube in place size adult
| RT275-09-012 | Tracheostomy | Tracheostomy tube neck strap Broad cotton cover with velcro fastener to hold tracheostomy tube in place size peadiatric
| RT275-09-013 | Tracheostomy | Stoma Button for surgically created tracheal stomas For safe decannulation and weaning, designed to have the full lumen of trachea, the patient is immediately able to cough up the stimulus secretion, because of airtight sealing of the stoma the patient is immediately able to speak Made out of soft and flexible high quality silicone Length of 45mm Packaging: 1 unit supplied in sterile package
| RT275-09-014 | Tracheostomy | Stoma Button for percutaneously created tracheal stomas For safe decannulation and weaning, designed to have the full lumen of trachea, the patient is immediately able to cough up the stimulus secretion, because of airtight sealing of the stoma the patient is immediately able to speak Made out of soft and flexible high quality silicone Length of 70mm Packaging: 1 unit supplied in sterile package
| RT275-09-015 | Tracheostomy | Tracheostomy cap/plug Latex free disposable decannulation plug aimed to close the proximal end of fenestrated tracheostomy tube to enable breathing through the fenestration and upper airway. Compatible with single lumen tracheostomy tube
| RT275-09-016 | Tracheostomy | Tracheostomy cap/plug Latex free disposable decannulation plug aimed to close the proximal end of fenestrated tracheostomy tube to enable breathing through the fenestration and upper airway. Compatible with double lumen tracheostomy tube
| RT275-09-017 | Tracheostomy | Speech Valve for Tracheostomy Tube 2-way valve placed on the external part of the tracheostomy tube Allows the patient to voice
| RT275-09-018 | Tracheostomy | Speaking Valve and oxygen nipple to allow oxygen therapy 2-way valve placed on the external part of the tracheostomy tube Allows the patient to voice as well as receive the benefits of oxygen therapy at the same time.
| RT275-09-019 | Tracheostomy | Speaking Valve with humidifier (HME) and oxygen nipple to allow oxygen therapy 2-way valve placed on the external part of the tracheostomy tube Allows the patient to voice as well as receive the benefits of a humidifier and oxygen therapy at the same time
| RT275-09-020 | Tracheostomy | Modified Speech Valve for use with a ventilator Speech valve with additional fitting for placement in line during ventilation with a ventilator Allows the patient to voice and speak during ventilation therapy (with male fitting)
| RT275-09-021 | Tracheostomy | Tracheostomy Tube Protector Allows for protection of the tracheostomy tube from accidental occlusion by having a distal grid and side apertures
| RT275-09-022 | Tracheostomy | Humidification Moisture Exchange (for Tracheostomy) Allows for trapping and absorbing moisture and heat during expiration through a medical grade sponge Size: Infants Dimensions: 25mm external diameter, 15mm internal diameter, 30mm length
| RT275-09-023 | Tracheostomy | Humidification Moisture Exchange (for tracheostomy) Allows for trapping and absorbing moisture and heat during expiration through a special sponge Size: Children & adults Dimensions: for attachment to all tubes with 15mm connector
| RT275-09-024 | Tracheostomy | Replacement sponges for HMEs (for tracheostomy) Medical Grade Sponges used within the HMEs to trap and absorb moisture and heat during expiration
| RT275-09-025 | Tracheostomy | HME (for tracheostomy) with removable oxygen port (for adults) Allows for trapping and absorbing moisture and heat during expiration through a medical grade sponge Dual functions (can be used with or without oxygen) Latex free
| RT275-09-026 | Tracheostomy | Hand-held pulse oximeter Must be lightweight, made of durable material, easy to transport - must be in a small durable, lightweight carry bag, Battery-operated, easy to use, able to measure blood oxygen saturation levels and pulse rate accurately & reliably for all ages ( from neonates to adults), sensors for adults, paediatrics & neonates (all 3 are needed and not 1 or the other), with LCD display, numeric display with plethysmogram display, pulse tone modulation, adjustable audio & visual alarms.
| RT275-09-027 | Tracheostomy | Tracheostoma spacer Silicone gel spacer which is round, tacky and laminated. Tool for improving seal around tracheostoma device fittings, improve cushioning and reduce air leakage. Size: internal diameter: 8mm, 3mm thick
| RT275-09-028 | Tracheostomy | Tracheostoma spacer Silicone gel spacer which is round, tacky and laminated. Tool for improving seal around tracheostoma device fittings, improve cushioning and reduce air leakage. Size: internal diameter: 12mm, 3mm thick
| RT275-09-029 | Tracheostomy | Tracheostoma spacer Silicone gel spacer which is round, tacky and laminated. Tool for improving seal around tracheostoma device fittings, improve cushioning and reduce air leakage. Size: internal diameter: 8mm, 5mm thick
| RT275-09-030 | Tracheostomy | Tracheostoma spacer Silicone gel spacer which is round, tacky and laminated. Tool for improving seal around tracheostoma device fittings, improve cushioning and reduce air leakage. Size: internal diameter: 12mm, 5mm thick
| RT275-09-031 | Tracheostomy | Protective tracheostomy / laryngectomy cloth protectors Protective tracheostomy cover for filtration and warming up of the respiratory air, washable, made of polyester with filter. Thick, durable and includes a sponge layer which provides humidification and filtration of the incoming air. Size: small with adjustable hook and loop or tie-up fastener
| RT275-09-032 | Tracheostomy | Protective tracheostomy / laryngectomy cloth protectors. Protective tracheostomy cover for filtration and warming up of the respiratory air, washable, made of polyester with filter. Thick, durable and includes a sponge layer which provides humidification and filtration of the incoming air. Size: medium, with adjustable hook and loop or tie-up fastener
| RT275-09-033 | Tracheostomy | Protective tracheostomy / laryngectomy cloth protectors Protective tracheostomy cover for filtration and warming up of the respiratory air, washable, made of polyester with filter. Thick, durable and includes a sponge layer which provides humidification and filtration of the incoming air. Size: large with adjustable hook and loop or tie-up fastener
| RT275-09-034 | Tracheostomy | Protective tracheostomy / laryngectomy cloth protectors Protective tracheostomy cover for filtration and warming up of the respiratory air, washable, made of polyester with filter. Thick, durable and includes a sponge layer which provides humidification and filtration of the incoming air. Size: small with adjustable velcro fastener
| RT275-09-035 | Tracheostomy | Protective tracheostomy / laryngectomy cloth protectors Protective tracheostomy cover for filtration and warming up of the respiratory air, washable, made of polyester with filter. Thick, durable and includes a sponge layer which provides humidification and filtration of the incoming air. Size: medium with adjustable velcro fastener
| RT275-09-036 | Tracheostomy | Protective tracheostomy / laryngectomy cloth protectors Protective tracheostomy cover for filtration and warming up of the respiratory air, washable, made of polyester with filter. Thick, durable and includes a sponge layer which provides humidification and filtration of the incoming air. Size: large with adjustable velcro fastener
| RT275-09-037 | Tracheostomy | Tracheostomy tube neck strap Broad cotton cover with velcro fastener to hold tracheostomy tube in place size adult
| RT275-09-038 | Tracheostomy | Tracheostomy tube neck strap Broad cotton cover with velcro fastener to hold tracheostomy tube in place size peadiatric
| RT275-09-039 | Tracheostomy | Stoma Button for surgically created tracheal stomas For safe decannulation and weaning, designed to have the full lumen of trachea, the patient is immediately able to cough up the stimulus secretion, because of airtight sealing of the stoma the patient is immediately able to speak Made out of soft and flexible high quality silicone Length of 45mm Packaging: 1 unit supplied in sterile package
| RT275-09-040 | Tracheostomy | Stoma Button for percutaneously created tracheal stomas For safe decannulation and weaning, designed to have the full lumen of trachea, the patient is immediately able to cough up the stimulus secretion, because of airtight sealing of the stoma the patient is immediately able to speak Made out of soft and flexible high quality silicone Length of 70mm Packaging: 1 unit supplied in sterile package
| RT275-09-041 | Tracheostomy | Tracheostomy cap/plug Latex free disposable decannulation plug aimed to close the proximal end of fenestrated tracheostomy tube to enable breathing through the fenestration and upper airway. Compatible with single lumen tracheostomy tube
| RT275-09-042 | Tracheostomy | Tracheostomy cap/plug Latex free disposable decannulation plug aimed to close the proximal end of fenestrated tracheostomy tube to enable breathing through the fenestration and upper airway. Compatible with double lumen tracheostomy tube
| RT275-09-043 | Tracheostomy | Speech Valve for Tracheostomy Tube 2-way valve placed on the external part of the tracheostomy tube Allows the patient to voice
| RT275-09-044 | Tracheostomy | Speaking Valve and oxygen nipple to allow oxygen therapy 2-way valve placed on the external part of the tracheostomy tube Allows the patient to voice as well as receive the benefits of oxygen therapy at the same time.
| RT275-09-045 | Tracheostomy | Speaking Valve with humidifier (HME) and oxygen nipple to allow oxygen therapy 2-way valve placed on the external part of the tracheostomy tube Allows the patient to voice as well as receive the benefits of a humidifier and oxygen therapy at the same time
| RT275-09-046 | Tracheostomy | Modified Speech Valve for use with a ventilator Speech valve with additional fitting for placement in line during ventilation with a ventilator Allows the patient to voice and speak during ventilation therapy (with male fitting)
| RT275-09-047 | Tracheostomy | Tracheostomy Tube Protector Allows for protection of the tracheostomy tube from accidental occlusion by having a distal grid and side apertures
| RT275-09-048 | Tracheostomy | Humidification Moisture Exchange (for Tracheostomy) Allows for trapping and absorbing moisture and heat during expiration through a medical grade sponge Size: Infants Dimensions: 25mm external diameter, 15mm internal diameter, 30mm length
| RT275-09-049 | Tracheostomy | Humidification Moisture Exchange (for tracheostomy) Allows for trapping and absorbing moisture and heat during expiration through a special sponge Size: Children & adults Dimensions: for attachment to all tubes with 15mm connector
| RT275-09-050 | Tracheostomy | Replacement sponges for HMEs (for tracheostomy) Medical Grade Sponges used within the HMEs to trap and absorb moisture and heat during expiration
| RT275-09-051 | Tracheostomy | HME (for tracheostomy) with removable oxygen port (for adults) Allows for trapping and absorbing moisture and heat during expiration through a medical grade sponge Dual functions (can be used with or without oxygen) Latex free
| RT275-09-052 | Tracheostomy | Hand-held pulse oximeter Must be lightweight, made of durable material, easy to transport - must be in a small durable, lightweight carry bag, Battery-operated, easy to use, able to measure blood oxygen saturation levels and pulse rate accurately & reliably for all ages ( from neonates to adults), sensors for adults, paediatrics & neonates (all 3 are needed and not 1 or the other), with LCD display, numeric display with plethysmogram display, pulse tone modulation, adjustable audio & visual alarms.
| RT275-10-01 | Assessment Tools | Goldman Fristoe Test of Articulation Standardised test that provides information about a child's articulation by sampling both spontaneous and imitated sound. Measures articulation of constant sounds and determines types of misarticultation. Age range between 2 - 21 years.
| RT275-10-02 | Assessment Tools | Scales of Cognitive Ability for Traumatic Brain Injury (SCATBI) Standardised test that assesses cognitive abilties of head injured patients. Subtests include perception/discrimination, orientation, organisation, recall and reasoning. Results can be used to establish severity of injury and to chart progress during recovery. Items progress in difficulty level. Targeted for adolescents and adults. Must be bundled with carry case/bag.
| RT275-10-03 | Assessment Tools | Toddler Language Scale Rossetti Infant - Standardised test that identifies preverbal and verbal language development problems. Uses observation, elicitation, and direct reports to determine if a behaviour is mastered. Age range between birth - 3 years.
| RT275-10-04 | Assessment Tools | Stuttering Severity Instrument Standardised test that measures stuttering severity of both adults and children for clinical and research use. Measures frequency of repetition and prolongations, duration of blocks and physical concomitants.
| RT275-10-05 | Assessment Tools | Assessment of the Speakers Experience of Stuttering (OASES) Standardised test that Overal measures the impact of stuttering on a person's life. This brief, yet comprehensive self report is built on solid theoretical foundation to help assess the impact of stuttering in multiple life situations. Age range of 7 years to adult.
| RT275-10-06 | Assessment Tools | Peabody Picture Vocabulary Test Standardised test that measures receptive single word vocabulary. Age range of 2.5 - 90+ years.
| RT275-10-07 | Assessment Tools | Western Aphasia Battery Standardised test that is individually administered for adults with acquired neurological disorders. Assesses linguistic skills most often affected by aphasia. Also assesses non-linguistic skills. Provides differential diagnosis information. Age range of 18 - 89 years.
| RT275-10-08 | Assessment Tools | Functional Assessment of Communication Skills for Adults (ASHA FACS) Standardised test that assists in measuring and recording the functional communication of adults with speech, language and cognitive communcation disorders. Assesses social communication, communication of basic needs, reading, writing and number concepts, and daily planning. Age range of 16 years - adult.
| RT275-10-09 | Assessment Tools | Clinical Evaluation of Language Fundamentals (UK - version) Standardised test to determine language strengths and weaknesses. Provides information on receptive language, expressive language, language structure, content and memory. Age range of 5 - 21 years.
| RT275-10-10 | Assessment Tools | Preschool Language Scale - UK version 5 Standardised test of developmental language. The items range from preverbal, interaction-based skills to emerging language to early literacy. Interactive, play-based assessment addressing language skills for children. Age range of birth - 7 years.
| RT275-10-11 | Assessment Tools | Kindergarten Language Screening Test Standardised screening assessment that identifies children who need further diagnostic testing to determine if they have language deficits that will impede academic achievement. Assesses the child's ability to demonstrate common preschool knowledge, understand questions, follow commands, repeat sentences, compare and contrast objects and use spontaneous speech. Age range of 3 - 6 years.
| RT275-10-12 | Assessment Tools | Test of Auditory Processing Skills Standardised test for children that have diagnoses of auditory perceptual difficulties, imperceptions of auditory modality, language problems and/or learning problems. Age range of 4 - 18 years.
| RT275-10-13 | Assessment Tools | Test of Language Development - Primary Standardised test assessing picture vocabulary, oral vocabulary, grammatical understanding, sentence immitation, grammatical completion, word discrimination and word articulation. Age range of 4 - 8 years.
| RT275-10-14 | Assessment Tools | Oral Speech Mechanism Screening Examination Standardised test that evaluated the patient's lips, tongue, jaw, teeth, palate, pharynx, velopharyngeal mechanism, and diadochokinetic rates. Age range of 5 - 78 years.
| RT275-10-15 | Assessment Tools | New Reynell Development Language Scale (NRDLS) Standardised language test to assess childrens production and understanding of spoken language. Suitable for ages ranging 3-8 years. Consists of multilingual toolkit and additional language.
| RT275-10-16 | Assessment Tools | South African Language Assessment (SALA) Standardised language assessment tool for zulu speaking children. Age range of 3 years to 7 years.
| RT275-10-17 | Assessment Tools | Renfrew action Picture Test Standardised test to obtain samples of spoken language to be evaluated in terms of information given and grammatical structures
| RT275-10-18 | Assessment Tools | Test of aided communication symbol preferences Test that allows a starting point for designing or selecting appropriate AAC device page set up. Test forms must reproducible
| RT275-10-19 | Assessment Tools | Shipley and McAffee, A Resource Manual, Fourth Edition Resource manual with a wide range of procedures and materials for obtaining, interpreting and reporting assessment data for a range of various speech, language pathologies. Consists of a book and CD
| RT275-10-20 | Assessment Tools | Visual Stress Assessment Provides anyone involved in detailed assessment of the colour overlay needs of people with a full set of information and assessment tools, presented in a clear, easy to follow format that is accessible to professionals and non-professionals alike.
| RT275-10-21 | Assessment Tools | The Discalculia Assessment Designed to explore which aspects of numeracy the child is struggling to acquire. The evidence from the assessment can then be used to draw up a personalized teaching plan. With clear, step-by-step instructions alongside photocopiable or downloadable assessment sheets.
| RT275-10-22 | Assessment Tools | Social Language Development Test Assesses language skill for social interactions. Subtests consist of question answering tasks, interpretations of photographed scenes, and verbal explanations.
| RT275-10-23 | Assessment Tools | Test of Integrated Language and Literacy Skills™ (TILLS™) Examiner's Kit Examiner Test Kitis the reliable, valid assessment professionals need to test oral and written language skills in students ages 6–18 years. TILLS is a comprehensive, norm-referenced test that has been standardized for three purposes: To identify language/literacy disorders, to document patterns of relative strengths and weaknesses, and to track changes in language and literacy skills over time. This Examiner’s Kit contains everything you need to administer, score, and interpret the TILLS test.
| RT275-10-24 | Assessment Tools | LiPS® - Fourth Edition - Complete Kit The LiPS: Lindamood Phoneme Sequencing® Program for Reading, Spelling, and Speech–Fourth Edition is a comprehensive multisensory program that uses explicit, systematic instruction to develop phonological awareness, decoding, spelling, and reading skills. The goal of LiPS is to develop fluent readers and competent spellers.
| RT275-10-25 | Assessment Tools | Autism Interventions: Exploring the Spectrum of Autism Second Edition Autism Interventions provides the most current assessment and intervention techniques, based on best practice and up-to-date research, for children on the Autism Spectrum. This practical and valuable tool offers holistic, dynamic, and effective intervention strategies and treatment plans to improve the functioning of individuals with autism.
| RT275-10-26 | Assessment Tools | CAAP-2: Clinical Assessment of Articulation and Phonology-Second Edition The CAAP-2 is the most current assessment for articulation and phonology. It is time-efficient, accurate, and yields results that are easy to score and interpret. You may administer the Articulation Inventory in 15-20 minutes.
| RT275-10-27 | Assessment Tools | Arizona-4: Arizona Articulation and Phonology Scale-Fourth Edition The Arizona Articulation and Phonology Scale- Fourth Edition measures speech intelligibility, articulatory impairment, and phonological impairment in one quick, easy to use assessment. The fourth edition of this widely used assessment retains the strengths of its predecessors and adds features that enhance its effectiveness in identifying individuals who have speech sound disorders.
| RT275-10-28 | Assessment Tools | Rey Auditory Verbal Learning Test - RAVLT The RAVLT is widely used to evaluate verbal learning and memory, including proactive inhibition, retroactive inhibition, retention, encoding versus retrieval, and subjective organization.
| RT275-10-29 | Assessment Tools | HAPP-3: Hodson Assessment of Phonological Patterns-Third Edition 1. coding and categorizing phonological deviations, 2. determining severity intervals/ratings (mild, moderate, severe, profound) along a continuum, 3. identifying major phonological patterns that need to be targeted (for goal statement), and 4. obtaining scores for documenting treatment effects over time (for evidence-based practice).
| RT275-10-30 | Assessment Tools | PAT-3: Photo Articulation Test-Third Edition The PAT-3 enables the clinician to rapidly and accurately assess and interpret articulation errors. The test consists of 72 color photographs (9 photos on each of eight sheets). The first 69 photos test consonants and all but one vowel and one diphthong. The remaining 3 pictures test connected speech and the remaining vowel and diphthong. A deck of the same 72 color photographs, each on a separate card, is provided for further diagnosis and may be used in speech-language remediation.
| RT275-10-31 | Assessment Tools | TOCS: Test of Childhood Stuttering The Test of Childhood Stuttering (TOCS) provides clinicians and researchers with a sound method for assessing speech fluency skills and stuttering-related behaviors in children 4 through 12 years of age. Its main purposes are to (1) identify children who stutter, (2) determine the severity of a child's stuttering, and (3) document changes in a child's fluency functioning over time. It can also be used as a tool in research on childhood stuttering.
| RT275-10-32 | Assessment Tools | ASIEP-3: Autism Screening Instrument for Educational Planning - Third Edition ASIEP-3 is a valid and reliable measure of autism and related behaviors. It was designed to identify individuals with autism; assist in planning appropriate educational programs for those individuals; monitor performance and progress; and research autism.
| RT275-10-33 | Assessment Tools | CARS-2: Childhood Autism Rating Scale-Second Edition CARS-2: Childhood Autism Rating Scale-Second EditionCARS-2 is a 15-item rating scale used to identify children with autism and distinguishing them from those with developmental disabilities. It is empirically validated and provides concise, objective, and quantifiable ratings based on direct behavioral observation. It was normed on a sample of 1,034 individuals with autism spectrum disorders.
| RT275-10-34 | Assessment Tools | DTVP-3: Developmental Test of Visual Perception - Third Edition The Gilliam Autism Rating Scale, now in its third edition, is one of the most widely used instruments for the assessment of Autism Spectrum Disorder in the world. The GARS-3 assists teachers, parents, and clinicians in identifying autism in individuals and estimating its severity.
| RT275-10-35 | Assessment Tools | LAC-3: Lindamood Auditory Conceptualization Test-Third Edition TACL-4 is the latest revision of this popular, individually administered test. It is a reliable and valid measure of a child's receptive spoken vocabulary, grammar, and syntax. The test measures children's ability to understand three language forms:
| RT275-10-36 | Assessment Tools | TACL-4: Test for Auditory Comprehension of Language-Fourth Edition TACL-4 is the latest revision of this popular, individually administered test. It is a reliable and valid measure of a child's receptive spoken vocabulary, grammar, and syntax. The test measures children's ability to understand three language forms:
| RT275-10-37 | Assessment Tools | LAC-3: Lindamood Auditory Conceptualization Test-Third Edition LAC-3 is a norm-referenced assessment that measures an individual's ability to perceive and conceptualize speech sounds using a visual medium. Because of the importance of these auditory skills to reading, the results are helpful for speech-language pathologists, special educators, and reading specialists.
| RT275-10-38 | Assessment Tools | (OWLS™-II) Oral and Written Language Scales, Second Edition Provides a complete and integrated picture of oral and written language skills across a wide age range test
| RT275-11-001 | Therapy Resources | My Play Home Doll House app The characters perform relevant daily activities related to the environment. Aids in developing pronouns, verbs, adjectives, imaginative play, sequencing etc. Compatible with 3 items below
| RT275-11-002 | Therapy Resources | My Stores/shop Doll House app The characters perform relevant daily activities related to the environment. Aids in developing pronouns, verbs, adjectives, imaginative play, sequencing etc.
| RT275-11-003 | Therapy Resources | My School Doll House The characters perform relevant daily activities related to the environment. Aids in developing pronouns, verbs, adjectives, imaginative play, sequencing etc.
| RT275-11-004 | Therapy Resources | My Hospital Doll House The characters perform relevant daily activities related to the environment. Aids in developing pronouns, verbs, adjectives, imaginative play, sequencing etc.
| RT275-11-005 | Therapy Resources | Letter Sounds Reading Doctor app Assists in teaching association between graphemes and phonemes. App uses games with repetition to teach children to associate letters and letter patterns with the sound they make.The supplier to indicate the operating system.
| RT275-11-006 | Therapy Resources | Aphasia Language therapy (Tactus Therapy Solutions) 4 in 1 toolkit of apps to include Comprehension, Naming, Writing and Reading Must have core nouns, verbs, & adjectives run throughout all 4 apps Target receptive & expressive goals Built-in hierarchy of cues encourages independence Completely customizable with words, pictures, prompts, & sounds Assessment and therapy exercises for clinical and independent use.The supplier to indicate the operating system.
| RT275-11-007 | Therapy Resources | Dysphagia Therapy (Tactus Therapy Solutions) Description: application designed to help clinicians assess swallowing disorders and determine appropriate rehabilitation options. Consists of four sections: Assessment, Treatment, Therapy Finder, and Resources. App to contain information about swallowing anatomy and physiology, the cranial nerves, treatment options, and downloadable patient handouts. Thebidder to indicate the operating system.
| RT275-11-008 | Therapy Resources | Speech Therapy vocabulary picture cards - food Set of high-quality photograph cards depicting various food types, including fruit, vegetables, prepared foods, drinks and snacks. Helps to improve attention, language and observation skills e. g. listening, understanding, vocabulary, writing, memory and language for both adult and children patients. Minimum of 20 cards per set.
| RT275-11-009 | Therapy Resources | Speech Therapy vocabulary picture cards - everyday objects Set of high-quality photograph cards with large format with different colour backgrounds that differentiate different categories of the common everyday nouns. Helps to improve attention, contration, vocabulary building, developing comprehension, encouraging expressive language, improving communication skills, grouping and classification for both adult and children patients. Minimum of 20 cards per set.
| RT275-11-010 | Therapy Resources | Speech Therapy vocabulary picture cards - basic verbs Set of high-quality photograph cards to illustrate frequently used verbs. Helps to improve verbal comprehension, encouraging expressive language, improving communication and teaching grammatical structure for both adult and children patients. Minimum of 20 cards per set.
| RT275-11-011 | Therapy Resources | Speech Therapy vocabulary picture cards - familiar verbs Set of high-quality photograph cards to illustrate frequently used and easily recorgnised verbs. Associated objects have been added to the images to give an idea of place and to encourage expressive langugage for both adult and children patients. Minimum of 20 cards per set.
| RT275-11-012 | Therapy Resources | Speech Therapy vocabulary picture cards - object groups Set of high-quality photograph cards with to illustrate groups of objects on a single card where all but one of the objects is from a single category. Contains three levels of difficulty. Minimum of 20 cards per set.
| RT275-11-013 | Therapy Resources | Speech Therapy vocabulary picture cards - basic sequences Set of high-quality photograph cards with to illustrate different three step sequences of well known everyday activities. Helps to establish basic concepts, understanding of order and direction, logical thought and expressive language skills, sequencing, vocabulary, improving communication skills and doing everyday tasks. Suitable for both adults and paediatric patients. Minimum of 4 sets.
| RT275-11-014 | Therapy Resources | Speech Therapy vocabulary picture cards - cause and effect Set of high-quality photograph cards to illustrate the sequential nature of day to day events. The notion of one event leads to another is shown in a number of different ways using two card sequences and ranges in it's level of complexity. Minimum of 20 cards per set.
| RT275-11-015 | Therapy Resources | Speech Therapy vocabulary picture cards - emotions Set of high-quality photograph cards with illustrations of different emotions. Encourages the understanding of other peoples feelings, allows for individual intepretation and response, has the potential to illustrate important issues surrounding the subject of emotions and feelings. Minimum of 20 cards per set.
| RT275-11-016 | Therapy Resources | Speech Therapy vocabulary picture cards - early objects Set of high-quality photograph cards that are pocket sized, illustrating early objects that are familiar in a child's surroundings. Allows for naming vocabulary through play. Minimum of 20 cards per set.
| RT275-11-017 | Therapy Resources | Speech Therapy vocabulary picture cards - early actions Set of high-quality photograph cards that are pocket sized, illustrating early actions that are familiar in a child's surroundings. Allows for the development of vocabulary of verbs, and assists children putting two or three words together. Minimum of 20 cards per set.
| RT275-11-018 | Therapy Resources | Speech Therapy vocabulary picture cards - early opposites Set of high-quality photograph cards that are pocket sized, illustrating opposites essential to the teaching of basic language. Contains different adjectives and prepositions to help increase a child's vocabulary. Minimum of 20 cards per set.
| RT275-11-019 | Therapy Resources | Speech Therapy vocabulary picture cards - emotions (child) Set of high-quality photograph cards that are pocket sized, illustrating simple feelings and attitudes. Minimum of 20 cards per set.
| RT275-11-020 | Therapy Resources | Treatment Resource Manual for Speech-Language Pathology, Roth and Worthington Resource manual with intervention strategies for unique child and adult populations. Covers metalinguistic skills, intervention for phonological disorders and language disorders in children.
| RT275-11-021 | Therapy Resources | Resource for Pediatric Dysphagia Resource manual with tools and goals to organise treatment and includes content on anatomy and physiology, clinical evaluation, instrumentation, treatment, managing dysphagia in the school setting and treatment of infants. Includes case history form, evaluation forms for 0-4 months, 4 months - 5 years and 5 - 18 years. In-service guide on dysphagia in school age children, education handouts, observation forms, questionaire for children with severe drooling, tips for developing good mealtime behaviours.
| RT275-11-022 | Therapy Resources | Resource for Dysphagia Resource manual with information on preparing for patient assessment, screening and bedside clinical evaluation, education and advocacy, ethics, instrumental assessment of swallowing (e. g. VFFS and FEES), planning dysphagia treatment, documenting treatment, special considerations for intensive care, and using data to manage dysphagia. The resource manual comes with bedside evaluation form, FEES report form, MBS report form, patient interview form, case examples, patient/family education materials, suctioning competency evaluation tool, sample letters to physicians, efficacy references, referral forms, and handouts of swallowing guidelines. Consists of book and CD.
| RT275-11-023 | Therapy Resources | Speech Therapy Assessment Tool Phonological Assessment Battery (PhAB), children aged 6-14 years, for whom English is an additional language, included data on special studies
| RT275-11-024 | Therapy Resources | Speech Therapy Assessment Tool Comprehensive Test Of Phonological Processing (CTOPP), Children Aged 6-14 Years, Testing Time 40 Min
| RT275-11-025 | Therapy Resources | Speech Therapy Assessment Tool Autism Spectrum Rating Scales By S Goldstein & Naglieri, JA
| RT275-11-026 | Therapy Resources | Speech Therapy Assessment Tool New Reynell Developmental Language Scales (NRDLS), age range: 3-8 years, for whom English is an additional language
| RT275-11-027 | Therapy Resources | Therapy resource Picture Exchange Communication System (PECS), AAC training system
| RT275-11-028 | Therapy Resources | Resource book Speech Language Therapist's Handbook for Inclusive School Practices: Julie Causton & Chelsea Tracy-Bronson (2014)
| RT275-11-029 | Therapy Resources | Therapy resource Speaking, Listening and Understanding: Games for Young Children, Catherine Delamin & Jill Spring,
| RT275-11-030 | Therapy Resources | Therapy resource Developing Language Concepts: Programmes for School-Aged Children, Bridget Burrows
| RT275-11-031 | Therapy Resources | Therapy resource Receptive and Expressive Activities for Language Therapy (REALt), Frenette Southwood and Ondene van Dulm. For children aged 4 and older. Available in English and Afrikaans.
| RT275-11-032 | Therapy Resources | Oral Placement Therapy for Speech Clarity and Feeding - S. Rosenfeld-Johnson
| RT275-11-033 | Therapy Resources | Oral Placement Therapy Homework Book - S. Rosenfeld-Johnson
| RT275-11-034 | Therapy Resources | Sensory Development: Mixed Senses Pack - Large Selection of developmental resources to stimulate vision, touch, auditory skills, sense of smell. Vision: examples include glowing and light-up products; Touch: examples include objects with a variety of textures; Auditory: examples include bells, musical instruments, etc.; Olfactory (smell): examples include essential oils and objects which have an aroma
| RT275-11-035 | Therapy Resources | Sensory Development: Vision Sensory Pack - Small Selection of developmental resources: Vision
| RT275-11-036 | Therapy Resources | Sensory Development: Tactile Fidget Pack - Small Selection of developmental resources: Tactile Fidget
| RT275-11-037 | Therapy Resources | Sensory Development: Aromas Pack - Small Selection of developmental resources: Aromas
| RT275-11-038 | Therapy Resources | Sensory Development: Sounds Pack - Small Selection of developmental resources: Sounds
| RT275-11-039 | Therapy Resources | Sensory Development: UV Glow Pack - Small Selection of developmental resources: UV Glow
| RT275-11-040 | Therapy Resources | Sensory Development: Fine Motor Sensory Pack - Small Selection of developmental resources: Fine Motor Sensory
| RT275-11-041 | Therapy Resources | Sensory Development: Pencil Grips Pack - Small Pencil Grips Selection Pack - set of 12
| RT275-11-042 | Therapy Resources | Sensory Development: Visual Stress Support Pack - Small Visual Stress Support - Reading Window
| RT275-11-043 | Therapy Resources | Sensory Development: Visual Stress Support Pack - Small Visual Stress Support - Reading Ruler
| RT275-11-044 | Therapy Resources | Sensory Development: Talking Therapy Pack - Small Talking Therapy Mirror - Record and See
| RT275-11-045 | Therapy Resources | Sensory Development: Daily Living Therapy Pack - Small Assisted Daily Living Therapy Pack
| RT275-11-046 | Therapy Resources | Sensory Development: Sensory Input Pack - Small Silicon chew pendant for sensory input / anxiety
| RT275-11-047 | Therapy Resources | Sensory Development: Fine Motor Pack - Small Fine Motor Maze Game with Magnetic Pen
| RT275-11-048 | Therapy Resources | Sensory Development: Fine Motor Pack - Small Fine Motor Threading Fruit / Veg
| RT275-11-049 | Therapy Resources | The Early Intervention Kit For early intervention with children aged 0 - 3 years. Addresses assessment, intervention, and documentation of pre-linguistic skills, speech and language development and sound production. Includes therapy guide, activity book, printable pages, sign language cards and vinyl folder. Nancy B Swigert.
| RT275-12-01 | Assistive Services | Remote Interpreting Interpretation service for sign language or spoken language; Includes device with high-quality audio and video capture and remote site with skilled interpreter
| RT275-12-02 | Assistive Services | Live Captioning Service that creates accurate, real-time speech-to-text from spoken content; Includes device with audio transmission (high-quality microphone) and text reception and display, and remote site with skilled captioner transcribing via stenograph or speech-to-text software; Optionally, with a good quality video feed, the service can offer sign language interpretation
| RT275-12-03 | Assistive Services | Video Closed Captioning Caption creation service to enhance existing video materials with displayed text; Transcription of the audio portion of a video; Provide additional, interpretive or environmental information; Captions can be enabled/disabled by the viewer
| RT275-12-04 | Assistive Services | A4 Portable Mirror Used for providing immediate visual feedback A4 mirror with strong protective frame. Surface of the mirror must be made of durable material, easy to clean and sterilisable. Must have a stand to allow positioning in portrait or landscape orientation on a flat surface. Mirror must be break-resistant and resistant to most chemicals.The mirror must be non-magnifying.
| RT275-12-05 | Assistive Services | Hand Held Mirror Mirror must be break-resistant and resistant to most chemicals. Surface of the mirror must be made of durable material, easy to clean and steralisable. Must have a strong plastic frame and handle. Must have a strong plastic frame and handle. Must be non-magnifying on one side and 2x magnification in the other side diameter of mirror must be 15cm.
| RT275-12-06 | Assistive Services | Palm Mirror Used for providing immediate visual feedback. Palm mirro must be break-resistant and resistant to most chemicals. Surfae of the mirror must be made of durable material, easy to clean and steralisable. Must be approximately 5cm x 7cm. Must be non-magnifying. Must have rounded corners and smoth, non-sharp edges. Must come in a protective pouch.
| RT275-12-007 | Assistive Services | Elecyromyography Unit with EMG. An electrical stimulation device and non-invasive method for treating dysphagia and artificially promoting swallowing by simultaneous electrical stimulation of the pharyngeal and facial muscles to include surface EMG. Unit must be portable, lighweight, battery operated, a fully charged battery must last for a minimum of 4 hours, the battery must be rechargeable. Must have the following feature programs: channel control panel with two or four channels, frequency selection, intensity control buttors, work and rest time parameters, protocol selection. The hardware components must be supplied with the following: amplifier, electrodes leads, out come data screen program selection switches/buttons, steam cables/codes, regarchable batteries, batteries charger. The unit must include the following: a complete user operation/maintenance manual x 2 Book/File; CD; DVD copies in English Language which must include the following information: fault finding guide; circuirt diagrams/schematics;circuit descriptions; PCB layouts; calibrtion guide; part numbers and exploded diagram of mechanical parts/panels. List of Accessories: amplifier, electrodes, leads, outcome data screen, program selection switches/buttons, steam cables/codes, regarchable batteries, batteries charger. The instrument should meet the following minumb requirements: ISO Standards FDR standards (provide certificates). Item must be compatible with RT275-12-008, RT275-12-009, RT275-12-010, RT275-12-011, RT275-12-012, RT275-12-013, RT275-12-014
| RT275-12-008 | Assistive Services | Amplifier for Electromyography unit (with EMG) Item must be compatible with RT275-12-07, Item must be compatible with RT275-12-007, RT275-12-009, RT275-12-010, RT275-12-011, RT275-12-012, RT275-12-013, RT275-12-014
| RT275-12-009 | Assistive Services | Electrode leads for Electromyography unit (with EMG) Item must be compatible with RT275-12-007, RT275-12-008, RT275-12-010, RT275-12-011, RT275-12-012, RT275-12-013, RT275-12-014
| RT275-12-010 | Assistive Services | Out come data screen for Electromyography unit (with EMG) Item must be compatible with RT275-12-007, RT275-12-008, RT275-12-009, RT275-12-011, RT275-12-012, RT275-12-013, RT275-12-014
| RT275-12-011 | Assistive Services | Program selection switches/buttons for Electromyography unit (with EMG) Item must be compatible with RT275-12-007, RT275-12-008, RT275-12-009, RT275-12-010, RT275-12-012, RT275-12-013, RT275-12-014
| RT275-12-012 | Assistive Services | Steam cables/codes for Electromyography unit (with EMG) Item must be compatible with RT275-12-007, RT275-12-008, RT275-12-009, RT275-12-010, RT275-12-011, RT275-12-013, RT275-12-014
| RT275-12-013 | Assistive Services | Regarchable batteries for Electromyography unit (with EMG) Item must be compatible with RT275-12-007, RT275-12-008, RT275-12-009, RT275-12-010, RT275-12-011, RT275-12-012, RT275-12-014
| RT275-12-014 | Assistive Services | Batteries charger for Electromyography unit (with EMG) Item must be compatible with RT275-12-007, RT275-12-008, RT275-12-009, RT275-12-010, RT275-12-011, RT275-12-012, RT275-12-013
| RT275-12-015 | Assistive Services | Elecyromyography Unit without EMG. An electrical stimulation device and non-invasive method for treating dysphagia and artificially promoting swallowing by simultaneous electrical stimulation of the pharyngeal and facial muscles. Unit must be portable, lighweight, battery operated, a fully charged battery must last for a minimum of 4 hours, the battery must be rechargeable. Must have the following feature programs: channel control panel with two or four channels, frequency selection, intensity control buttors, work and rest time parameters, protocol selection. The hardware components must be supplied with the following: amplifier, electrodes leads, out come data screen, program selection switches/buttons, steam cables/codes, regarchable batteries, batteries charger. The unit must include the following: a complete user operation/maintenance manual x 2 Book/File; CD; DVD copies in English Language which must include the following information: fault finding guide; circuirt diagrams/schematics;circuit descriptions; PCB layouts; calibrtion guide; part numbers and exploded diagram of mechanical parts/panels. List of Accessories: amplifier, electrodes, leads, outcome data screen, program selection switches/buttons, steam cables/codes, regarchable batteries, batteries charger. The instrument should meet the following minumb requirements: ISO Standards FDR standards (provide certificates) Item must be compatible with RT275-12-016, RT275-12-017, RT275-12-018, RT275-12-019, RT275-12-020, RT275-12-021, RT275-12-022
| RT275-12-016 | Assistive Services | Amplifier for Electromyography unit (without EMG) Item must be compatible with RT275-12-016, RT275-12-017, RT275-12-018, RT275-12-019, RT275-12-020, RT275-12-021, RT275-12-022
| RT275-12-017 | Assistive Services | Electrode leads for Electromyography unit (without EMG) Item must be compatible with RT275-12-016, RT275-12-017, RT275-12-018, RT275-12-019, RT275-12-020, RT275-12-021, RT275-12-022
| RT275-12-018 | Assistive Services | Out come data screen for Electromyography unit (without EMG)
| RT275-12-019 | Assistive Services | Program selection switches/buttons for Electromyography unit (without EMG) Item must be compatible with RT275-12-016, RT275-12-017, RT275-12-018, RT275-12-019, RT275-12-020, RT275-12-021, RT275-12-022
| RT275-12-020 | Assistive Services | Steam cables/codes for Electromyography unit (without EMG) Item must be compatible with RT275-12-016, RT275-12-017, RT275-12-018, RT275-12-019, RT275-12-020, RT275-12-021, RT275-12-022
| RT275-12-021 | Assistive Services | Regarchable batteries for Electromyography unit (without EMG)
| RT275-12-022 | Assistive Services | Batteries charger for Electromyography unit (without EMG)