## Technical Specifications
ANNEXURE B: TECHNICAL SPECIFICATIONS RT28-1-2024
Item Code | Category | Item Description | Unit of Measure | End-User
RT28-1-001 | Tapes | Instrument colour code tape, unique vinyl tape for marking all types of surgical instruments. Autoclavable, withstands repeated sterillization cycles. Must be extremely durable and can be used in all types of autoclaves Size: ±6,35mm x ±10m Colour: purple Roll | roll | NATIONAL
RT28-1-002 | Tapes | Instrument colour code tape, unique vinyl tape for marking all types of surgical instruments. Autoclavable, withstands repeated sterillization cycles. Must be extremely durable and can be used in all types of autoclaves Size: ±6,35mm x ±10m Colour: orange Roll | roll | NATIONAL
RT28-1-003 | Tapes | Instrument colour code tape, unique vinyl tape for marking all types of surgical instruments. Autoclavable, withstands repeated sterillization cycles. Must be extremely durable and can be used in all types of autoclaves Size: ±6,35mm x ±10m Colour: brown Roll | roll | NATIONAL
RT28-1-004 | Tapes | Manual portable dispenser. Cut Length: As required. Tape Width: 3-25mm Usable tape type: Autoclavable. Masking tape, Ethylene oxide tape Must have a weighted nonslip base. | each | NATIONAL
RT28-1-005 | Accessories & Consumables | Silicone spray, anti-spatter silicone 400ml-500ml aeresol container | each | NATIONAL
RT28-1-006 | Accessories & Consumables | 470 x 274 x 187 Bottom for 3/4 container height:187mm 3/4-size lid w/rentention Plate Silver 3/4 size perf basket 406x253x76mm Rack for bone punches and rongeurs Silicone basket liner 3/4 size | each | NATIONAL
RT28-1-007 | Accessories & Consumables | 592 x 274 x 187 Bottom for 1/1 container height:187mm Full-size lid w/rentention plate silver Mis storage rack with basket Silicone basket liner larger Size | each | NATIONAL
RT28-1-008 | Tubes | Tube, sterilization. Flat sterilizing reels: with a gusset Heat sealable reels of paper and plastic film Seams to peel open without linting and/or tearing film Sterilization process indicators for steam and ethylene oxide shall be printed at equal distances throughout the roll Suitable for packaging medical instruments, which have to be sterilised in steam or ethylene oxide. Quality of paper shall comply with SANS 50868 and DIN 50953-6 to test for the bacterial barrier. Size: 50mm x 200m Roll | roll | NATIONAL
RT28-1-009 | Tubes | Low termperature sterilisation reels (roll) Peelable reels designed for hydrogen peroxide, ethylene oxide, gamma radiation sterilization processess. Sterilization rols must be constructed of two pieces of web material which is sealed on the sides. The barrier/porous film must be a spunbound polyolefin 4057B with a build-in chemical indicator that change colour after being exposed to the sterilization process. The reel need to comply with EN 868-5, ISO 11607, 93/42/EEC Medical Device Regulation with the amendment of 2007/47/EC standards or equivalent. Size: 100mmx70m Roll | roll | NATIONAL
RT28-1-010 | Tapes | Lead free indicator tape for ethylene oxide sterilization. Pressure sensitive adhesive backing with chemical indicator stripes. Conforms to the International Standards Organization (ISO 11140-1) and European Norm (EN 867-2, Class A) as well as the British Standards The complete change indicating exposure to an ethylene oxide sterilization process shall not occur until the indicator has been exposed to 600 ± 30 mg/l ethylene oxide, 60 ± 10% RH at 30°C ± 1°C for not less than 5 minutes. Must adhere to woven, treated woven, non-woven, paper and paper/plastic wraps and must be easy to remove. | roll | NATIONAL
RT28-1-011 | Indicators | Steam sensitive chemical indicator pack for routine identification of air leaks and vacuum deficiencies vacuum system steam autoclave. Conforms to ISO 11140-4,5 standard. Consists of a test sheet positioned inside porous materials and sealed inside a disposable outer wrap. The test sheet will turn a uniform dark brown/black color, except when air removal failures such as air leaks occur. Diagonal light yellow chemical indicator stripes turn a dark brown/black when air removal is effective. Test sheets can be filed for quality assurance records. Lead-free construction. | each | NATIONAL
RT28-1-012 | Wraps | Wrap, sterilization, medical, non-woven Must be compatible to be use with steam, ethylene oxide, formaldehyde and Hydrogen peroxide gas sterilization processes Spunbond-Meltblown-Meltblown-Meltblown-Spunbond (SMMMS) To comply with the latest issue of SANS 50868-2 DIN 58953-6, SANS 11607-1 and/or EN868-2 Size: ±600 x 600mm | pack of 100 | NATIONAL
RT28-1-013 | Wraps | Wrap, sterilization, medical, non-woven Must be compatible to be use with steam, ethylene oxide, formaldehyde and Hydrogen peroxide gas sterilization processes Spunbond-Meltblown-Meltblown-Meltblown-Spunbond (SMMMS) To comply with the latest issue of SANS 50868-2 DIN 58953-6, SANS 11607-1 and/or EN868-2 Size: 900 x 900mm | pack of 100 | NATIONAL
RT28-1-014 | Labelling Devices | Labelling device for use with steam and gas and low themperature sterilization processess Labelling device must be print on a double adhesive layer label Must be able to print at least 5 different data information on the label device must have a alphanumeric print head Supplier is responsible to maintain the device at no extra cost during contract period on all movable parts of the device Each Item 53, 54, 55 and 56 must be compatible | each | NATIONAL
RT28-1-015 | Wraps | Wrap, Double, sterilization, medical, non-woven Spunbond-Meltblown-Spunbond (SMMS) To comply with the latest issue of SANS 50868-2 DIN 58953-6, SANS 11607-1 and/or EN868-2 Regular set Size: ±1200 x ±1200mm | pack of 100 | NATIONAL
RT28-1-016 | Wraps | Wrap, Double, sterilization, medical, non-woven Spunbond-Meltblown-Spunbond (SMMS) To comply with the latest issue of SANS 50868-2 DIN 58953-6, SANS 11607-1 and/or EN868-2 Mid-Weight set Size: ±600mm x ±600mm | pack of 100 | NATIONAL
RT28-1-017 | Wraps | Wrap, Double, sterilization, medical, non-woven Spunbond-Meltblown-Spunbond (SMMS) To comply with the latest issue of SANS 50868-2 and/or EN868-2 Mid-Weight set Size: ±900mm x ±900mm | pack of 100 | NATIONAL
RT28-1-018 | Wraps | Wrap, Double, sterilization, medical, non-woven Spunbond-Meltblown-Spunbond (SMMS) To comply with the latest issue of SANS 50868-2 DIN 58953-6, SANS 11607-1 and/or EN868-2 Mid-Weight set Size: ±1200mm x ±1200mm | pack of 100 | NATIONAL
RT28-1-019 | Wraps | Wrap, Double, sterilization, medical, non-woven Spunbond-Meltblown-Spunbond (SMMS) To comply with the latest issue of SANS 50868-2 DIN 58953-6, SANS 11607-1 2 and/or EN868-2 Heavy set Size: ±600mm x ±600mm | pack of 100 | NATIONAL
RT28-1-020 | Wraps | Wrap, Double, sterilization, medical, non-woven Spunbond-Meltblown-Spunbond (SMMS) To comply with the latest issue of SANS 50868-2 DIN 58953-6, SANS 11607-1 and/or EN868-2 Heavy duty Size: ±900mm x ±900mm | pack of 100 | NATIONAL
RT28-1-021 | Wraps | Wrap, Double, sterilization, medical, non-woven Spunbond-Meltblown-Spunbond (SMMS) To comply with the latest issue of SANS 50868-2 DIN 58953-6, SANS 11607-1 2 and/or EN868-2 Heavy set Size: ±1200mm x ±1200mm | pack of 100 | NATIONAL
RT28-1-022 | Indicators | Indicator vacuum test pack, bowie Dick. Must test efficiency of scope sterilation process at temperatures of 134°C for 3,5minutes or 121°C for 15 minutes. The contens of the pack must includes the following: 1 Helix Process Challenge Device, 500 indicator strips type 2 steam, cotton bag and 1 manual/colour reference. Packaging 500 indicators per box | each | NATIONAL
RT28-1-023 | Wraps | Wrap, single, sterilization, medical, non-woven Spunbond-Meltblown-Spunbond (SMS) To comply with the latest issue of SANS 50868-2 Regular Size: 600 x 600mm Pack of 240 Sheets | pack of 240 | NATIONAL
RT28-1-024 | Wraps | Wrap, single, sterilization, medical, non-woven Spunbond-Meltblown-Spunbond (SMS) To comply with the latest issue of SANS 50868-2 Regular Size: 600 x 900mm Pack of 240 Sheets | pack of 240 | NATIONAL
RT28-1-025 | Wraps | Wrap, single, sterilization, medical, non-woven Spunbond-Meltblown-Spunbond (SMS) To comply with the latest issue of SANS 50868-2 Regular Size: 900 x 900mm Pack of 240 Sheets | pack of 240 | NATIONAL
RT28-1-026 | Wraps | Wrap, single, sterilization, medical, non-woven Spunbond-Meltblown-Spunbond (SMS) To comply with the latest issue of SANS 50868-2 Regular Size: 1200 x 1200mm Pack of 240 Sheets | pack of 240 | NATIONAL
RT28-1-027 | Wraps | Wrap, single, sterilization, medical, non-woven Spunbond-Meltblown-Spunbond (SMS) To comply with the latest issue of SANS 50868-2 Regular Size: 1400 x 1400mm Pack of 240 Sheets | pack of 240 | NATIONAL
RT28-1-028 | Wraps | Wrap, single, sterilization, medical, non-woven Spunbond-Meltblown-Spunbond (SMS) To comply with the latest issue of SANS 50868-2 Regular Size: 1400 x 1830mm Pack of 240 Sheets | pack of 240 | NATIONAL
RT28-1-029 | Wraps | Wrap, single, sterilization, medical, non-woven Spunbond-Meltblown-Spunbond (SMS) To comply with the latest issue of SANS 50868-2 Mid-Weight Size: 600mm x 600mm Pack of 240 Sheets | pack of 240 | NATIONAL
RT28-1-030 | Wraps | Wrap, single, sterilization, medical, non-woven Spunbond-Meltblown-Spunbond (SMS) To comply with the latest issue of SANS 50868-2 Mid-Weight Size: 600mm x 900mm Pack of 240 Sheets | pack of 240 | NATIONAL
RT28-1-031 | Wraps | Wrap, single, sterilization, medical, non-woven Spunbond-Meltblown-Spunbond (SMS) To comply with the latest issue of SANS 50868-2 Mid-Weight Size: 900mm x 900mm Pack of 240 Sheets | pack of 240 | NATIONAL
RT28-1-032 | Wraps | Wrap, single, sterilization, medical, non-woven Spunbond-Meltblown-Spunbond (SMS) To comply with the latest issue of SANS 50868-2 Mid-Weight Size: 1200mm x 1200mm Pack of 240 Sheets | pack of 240 | NATIONAL
RT28-1-033 | Wraps | Wrap, single, sterilization, medical, non-woven Spunbond-Meltblown-Spunbond (SMS) To comply with the latest issue of SANS 50868-2 Mid-Weight Size: 1400mm x 1400mm Pack of 240 Sheets | pack of 240 | NATIONAL
RT28-1-034 | Wraps | Wrap, single, sterilization, medical, non-woven Spunbond-Meltblown-Spunbond (SMS) To comply with the latest issue of SANS 50868-2 Mid-Weight Size: 1400mm x 1830mm Pack of 240 Sheets | pack of 240 | NATIONAL
RT28-1-035 | Wraps | Wrap, single, sterilization, medical, non-woven Spunbond-Meltblown-Spunbond (SMS) To comply with the latest issue of SANS 50868-2 Heavy duty Size: 600mm x 600mm Pack of 120 Sheets | pack of 120 | NATIONAL
RT28-1-036 | Wraps | Wrap, single, sterilization, medical, non-woven Spunbond-Meltblown-Spunbond (SMS) To comply with the latest issue of SANS 50868-2 Heavy duty Size: 600mm x 900mm Pack of 120 Sheets | pack of 120 | NATIONAL
RT28-1-037 | Wraps | Wrap, single, sterilization, medical, non-woven Spunbond-Meltblown-Spunbond (SMS) To comply with the latest issue of SANS 50868-2 Heavy duty Size: 900mm x 900mm Pack of 120 Sheets | pack of 120 | NATIONAL
RT28-1-038 | Wraps | Wrap, single, sterilization, medical, non-woven Spunbond-Meltblown-Spunbond (SMS) To comply with the latest issue of SANS 50868-2 Heavy duty Size: 1200mm x 1200mm Pack of 120 Sheets | pack of 120 | NATIONAL
RT28-1-039 | Wraps | Wrap, single, sterilization, medical, non-woven Spunbond-Meltblown-Spunbond (SMS) To comply with the latest issue of SANS 50868-2 Heavy duty Size: 1400mm x 1400mm Pack of 48 Sheets | pack of 120 | NATIONAL
RT28-1-040 | Wraps | Wrap, single, sterilization, medical, non-woven Spunbond-Meltblown-Spunbond (SMS) To comply with the latest issue of SANS 50868-2 Heavy duty Size: 1400mm x 1830mm Pack of 120 Sheets | pack of 120 | NATIONAL
RT28-1-041 | Tubes | Tube, sterilization. Flat sterilizing reels: without gusset. Heat sealable reels of paper and plastic film. Quality of paper shall comply with SANS 50868 parts 1, 3, 4 and 5 and DIN 50953-6 to test for the bacterial barrier. Seams to peel open without linting and/or tearing film. Sterilization process indicators for steam and ethylene oxide shall be printed at equal distances throughout the roll. Suitable for packaging medical instruments, which have to be sterilised in steam or ethylene oxide. Size: 300mm x 200m Roll | pack of 120 | NATIONAL
RT28-1-042 | Cleaning Brushes | Toothbrush style cleaning brushes, nylon ergonomic handle: Double ended nylon bristles: brush width 6 mm, end to end 180mm | each | NATIONAL
RT28-1-043 | Cleaning Brushes | Toothbrush style cleaning brushes, nylon ergonomic handle: Three row nylon bristles, brush width 6mm, bristle length 29mm, end to end 180mm | each | NATIONAL
RT28-1-044 | Cleaning Brushes | Toothbrush style cleaning brushes, nylon ergonomic handle: Three row stainless steel bristles: brush width 6mm, bristle length 29mm, end to end 180mm | each | NATIONAL
RT28-1-045 | Cleaning Brushes | Burr cleaning brushes, guide powerful cleaning of orthopaedic devices like burrs, rasps and files with the burr style brush, High-grade plastic handles resist breaking stainless steel bristles provide powerful cleaning. Round burr brush: brush width 16,mm, bristle length 13mm, end to end 64mm | each | NATIONAL
RT28-1-046 | Cleaning Brushes | Burr cleaning brushes, guide powerful cleaning of orthopaedic devices like burrs, rasps and files with the burr style brush, High-grade plastic handles resist breaking stainless steel bristles provide powerful cleaning. Single row burr brush: brush width 3mm, bristle length 35mm, end to end 165mm | each | NATIONAL
RT28-1-047 | Cleaning Brushes | Burr cleaning brushes, guide powerful cleaning of orthopaedic devices like burrs, rasps and files with the burr style brush, High-grade plastic handles resist breaking stainless steel bristles provide powerful cleaning. Three row brush: brush width 16mm, bristle length 35mm, end to end 450mm | each | NATIONAL
RT28-1-048 | Cleaning Brushes | Instrument cleaning brushes(bottle brushes, nylon ergonomic bristles). Nylon bristles : brush width 2,mm, bristle length 13mm, end to end 150mm | each | NATIONAL
RT28-1-049 | Cleaning Brushes | Instrument cleaning brushes(bottle brushes, nylon ergonomic bristles). Nylon bristles : brush width 2,mm, bristle length 13mm, end to end 450mm | each | NATIONAL
RT28-1-050 | Cleaning Brushes | Instrument cleaning brushes(bottle brushes, nylon ergonomic bristles). Nylon bristles : brush width 2,mm, bristle length 13mm, end to end 600mm | each | NATIONAL
RT28-1-051 | Cleaning Brushes | Instrument cleaning brushes(bottle brushes, nylon ergonomic bristles). Nylon bristles : brush width 2,mm, bristle length 25mm, end to end 600mm | each | NATIONAL
RT28-1-052 | Cleaning Brushes | Instrument cleaning brushes(bottle brushes, nylon ergonomic bristles). Nylon bristles : brush width 2,mm, bristle length 13mm, end to end 450mm | each | NATIONAL
RT28-1-053 | Cleaning Brushes | Instrument cleaning brushes(bottle brushes, nylon ergonomic bristles). Nylon bristles : brush width 3,mm, bristle length 25mm, end to end 450mm | each | NATIONAL
RT28-1-054 | Cleaning Brushes | Instrument cleaning brushes(bottle brushes, nylon ergonomic bristles). Nylon bristles : brush width 3,mm, bristle length 50mm, end to end 450mm | each | NATIONAL
RT28-1-055 | Cleaning Brushes | Instrument cleaning brushes(bottle brushes, nylon ergonomic bristles). Nylon bristles : brush width 6mm, bristle length 50mm, end to end 450mm | each | NATIONAL
RT28-1-056 | Cleaning Brushes | Instrument cleaning brushes(bottle brushes, nylon ergonomic bristles). Nylon bristles : brush width 8mm, bristle length 75mm, end to end 600mm | each | NATIONAL
RT28-1-057 | Cleaning Brushes | Instrument cleaning brushes(bottle brushes, nylon ergonomic bristles). Nylon bristles : brush width 10mm, bristle length 38mm, end to end 150mm | each | NATIONAL
RT28-1-058 | Cleaning Brushes | Instrument cleaning brushes(bottle brushes, nylon ergonomic bristles). Nylon bristles : brush width 11mm, bristle length 75mm, end to end 600mm | each | NATIONAL
RT28-1-059 | Cleaning Brushes | Instrument cleaning brushes(bottle brushes, nylon ergonomic bristles). Nylon bristles : brush width 22mm, bristle length 89mm, end to end 450mm | each | NATIONAL
RT28-1-060 | Cleaning Brushes | Instrument cleaning brushes(bottle brushes, nylon ergonomic bristles). Nylon bristles : brush width 13mm, bristle length 75m, end to end 450mm | each | NATIONAL
RT28-1-061 | Cleaning Brushes | Instrument cleaning brushes(bottle brushes, nylon ergonomic bristles). Nylon bristles : brush width 11mm, bristle length 50m, end to end 150mm | each | NATIONAL
RT28-1-062 | Cleaning Brushes | Extra-long speciality cleaning brushes extra-long brush with nylon bristle tip attached to rigid tubing: Bristle on one end only : brush width 3.5mm, bristle length 10mm, end to end 2300mm | each | NATIONAL
RT28-1-063 | Cleaning Brushes | Extra-long speciality cleaning brushes extra-long brush with nylon bristle tip attached to rigid tubing: Bristle on one end only : brush width 3.5mm, bristle length 20mm, end to end 2300mm | each | NATIONAL
RT28-1-064 | Cleaning Brushes | Extra-long speciality cleaning brushes extra-long brush with nylon bristle tip attached to rigid tubing: Bristle on both ends : brush width 3.5mm, bristle length 10mm, end to end 2300mm | each | NATIONAL
RT28-1-065 | Cleaning Brushes | Extra-long speciality cleaning brushes extra-long brush with nylon bristle tip attached to rigid tubing: Bristle on both ends : brush width 3.5mm, bristle length 20mm, end to end 2300mm | each | NATIONAL
RT28-1-066 | Cleaning Brushes | Extra-long speciality cleaning brushes delicate pull-through brush. nylon bristle tip attached to stainless steel wire: Bristle on one end only : brush width 1mm, bristle length 10mm, end to end 650mm | each | NATIONAL
RT28-1-067 | Cleaning Brushes | Extra-long speciality cleaning brushes delicate pull-through brush. nylon bristle tip attached to stainless steel wire: Bristle on one end only : brush width 2mm, bristle length 10mm, end to end 1200mm | each | NATIONAL
RT28-1-068 | Cleaning Brushes | Extra-long speciality cleaning brushes delicate pull-through brush. nylon bristle tip attached to stainless steel wire: Bristle on one end only : brush width 3mm, bristle length 10mm, end to end 1200mm | each | NATIONAL
RT28-1-069 | Cleaning Brushes | Extra-long speciality cleaning brushes delicate pull-through brush. nylon bristle tip attached to stainless steel wire: Bristle on one end only : brush width 3mm, bristle length 10mm, end to end 1600mm | each | NATIONAL
RT28-1-070 | Cleaning Brushes | Extra-long speciality cleaning brushes extra-long nylon bristles attached to stainless steel coil wire. loop handle at one end and beaded tip at bristle end: Bristle on one end only : brush width 4.6mm, bristle length 10mm, end to end 2300mm | each | NATIONAL
RT28-1-071 | Cleaning Brushes | Extra-long speciality cleaning brushes Delicate pull-through brush. Nylon bristle tip attached to stainless steel wire: Bristle on one end only : brush width 1mm, bristle length 10mm, End to end 650mm | each | NATIONAL
RT28-1-072 | Cleaning Brushes | Extra-long speciality cleaning brushes Delicate pull-through brush. Nylon bristle tip attached to stainless steel wire: Bristle on one end only : brush width 2mm, bristle length 10mm, End to end 1200mm | each | NATIONAL
RT28-1-073 | Cleaning Brushes | Extra-long speciality cleaning brushes Delicate pull-through brush. Nylon bristle tip attached to stainless steel wire: Bristle on one end only : brush width 3mm, bristle length 10mm, End to end 1200mm | each | NATIONAL
RT28-1-074 | Cleaning Brushes | Extra-long speciality cleaning brushes Delicate pull-through brush. Nylon bristle tip attached to stainless steel wire: Bristle on one end only : brush width 3mm, bristle length 10mm, End to end 1600mm | each | NATIONAL
RT28-1-075 | Cleaning Brushes | Reusable handle & brush, tips for sterile processing departments that prefer the consistency the same handle provides, these brushes offer reusable shafts with fully interchangeable, nylon disposable brush heads. diameter (mm) 5,8,10 | each | NATIONAL