## Specs
RT286-01-2021 : ANNEXURE C PRICING SCHEDULE
ITEM NUMBER | ITEM CATEGORY | ITEM DESCRIPTION | Name of the Bidder | Name of the Manufacturer | Brand Name | Unit of measure (UOM) | Test Report/ Techincal Specification | Minimum Order Quantity | COUNTRY OF ORIGIN | Product Code Number/Model No. | EQUIPMENT/ ACCESSORY PRICE INCL. OF VAT, 1ST & 2ND YEAR MANTAINANCE | 3RD YEAR MANTAINANCE COST | 4th YEAR MANTAINANCE COST | 5th YEAR MANTAINANCE COST | TOTAL 5 YEARS MANTAINANCE COST | TOTAL PRICE ,EQUIPMENTS & MANTAINANCE
RT286-01-001 | Laboratory Equipment | Plaster Dispenser - Electric SPECIFICATION An electrically operated silo for the dry storge and controlled dispensing of plasters and stone in the dental laboratory. The unit shall have the following features:Suspended on the wall or bench mounted in a frameCapacity of the dispenser no less that 25kgConstruction of the dispenser:Outer housing made of stainless steel or plasticInner hopper of plastic for corrosion-free plaster / stone supplyBuilt-in electrical vibratorDispensing leverVibration action ensuring even dispensing of materialActivating dispensing lever setting vibrator into motionSABS Plug : Must comply with both SANS 164-0 and SANS 164-1 |  |  |  | Each
RT286-01-002 | Laboratory Equipment | Model Trimmer - Single Disc / BladeSPECIFICATION A single disc/ blade model trimmer utilised in a dental laboratory for the trimming of plaster, stone and super hard stone models. The unit shall have the following features:Single motor encased in water-proof casingDrive motor speed minimum 1800rpmON/OFF power switch & speed selector on casingTrimmer discs features: Reversible disc, Course, dia 10" or 12"Tap for adjusting water flow to the discWater jets for cleaning disc during operationDiscs covered with water splash guard, leaving grinding surface app 100mm x 140mmUnit mounted on anti-vibration rubber stoppersSplash shield Water spray attachmentOrthodontic work table including Vertisquare, Angle guide and Template SABS Plug : Must comply with both SANS 164-0 and SANS 164-1Additional AccessoriesReplacement disc Reversible, Course, dia 10" or 12"Replacement Vertisquare |  |  |  | Each
RT286-01-003 | Laboratory Equipment | Model Trimmer - Dual Disc / Blade for the Model Trimmer - Dual Disc / BladeSPECIFICATION A dual disc/ blade model trimmer utilised in a dental laboratory for the trimming of plaster, stone and super hard stone models as well as ortho stone orthodontic study models. The unit shall have the following features:Motor is Split Phase, Thermally Protected, Moisture ResistantTrimmer discs features: Reversible, One Coarse for trimming, One Fine for Finishing, diameter 12"One tap for adjusting water flow to the discWater jets for cleaning disc during operationDisc covered with water splash guard Each grinding table with one guide groove in line with each grinding discSquaring tool and an angulation tool to be supplied with each unit to fit into groovesUnit mounted on anti-vibration rubber stoppersAccessory Splash guard dual SABS Plug : Must comply with both SANS 164-0 and SANS 164-1 |  |  |  | Each
RT286-01-003 A (Aditional Accessories | Laboratory Equipment | Replacement disc Coarse for trimming diameter 12" |  |  |  | Each
RT286-01-003 B | Laboratory Equipment | Replacement disc Fine for Finishing, diameter 12" for the Model Trimmer - Dual Disc / Blade |  |  |  | Each
RT286-01-003 C | Laboratory Equipment | Replacement Vertisquare for the Model Trimmer - Dual Disc / Blade |  |  |  | Each
RT286-01-004 | Laboratory Equipment | Model Trimmer - Dry System with Partially Coated Disc / BladeSPECIFICATION A single disc, dry system, plaster model trimmer for the trimming of models in a dental laboratory. The unit shall have the following features:Powerful motor no less than 1600WQuiet running electric motor, not more than 75 dB(A)A partially diamond coated discElectronic instant-stop for trimmer discTool-free quick release lock on removable casing doorSafety switch to interrupt power on opening casing doorFitted with connection for external dust / particle extractor unitSABS Plug : Must comply with both SANS 164-0 and SANS 164-1 |  |  |  | Each
RT286-01-004 A (Additional Accessories) | Laboratory Equipment | Replacement partially diamond coated disc for the Model Trimmer - Dry System with Partially Coated Disc / Blade |  |  |  | Each
RT286-01-005 | Laboratory Equipment | Pin Hole Drilling UnitSPECIFICATION A benchtop, manually controlled unit utilised in the crown and bridge technique for drilling holes in the base of crown and bridge models, prior to the placement of die pins. The unit shall have the following features:Sturdy, stable baseSlip-free model table Model table designed to prevent dust accumulationPin drill mounted below model table pointing upwardsOverhead light / laser aiming deviceAiming device exact match to pin drillDrilling activated bmo pressing model table down/upChanging of drill bit quick and preferably tool-freeAdjustable drill depth: 0-12mmEngine / Drive shaft protected from plaster dustSABS Plug : Must comply with both SANS 164-0 and SANS 164-1 |  |  |  | Each
RT286-01-005 A (Additional Accessories) | Laboratory Equipment | Replacement pin drill for the Pin Hole Drilling Unit |  |  |  | Each
RT286-01-005 B (Additional Accessories) | Laboratory Equipment | Replacement LED Light for the Pin Hole Drilling Unit |  |  |  | Each
RT286-01-006 | Laboratory Equipment | Surveyor for Models - Single ArmSPECIFICATION A benchtop, manually controlled unit utilised in the cobalt-chrome technique for model analysis, measuring undercuts and marking the equator. The unit shall have the following features:Base of unit: approx. 120mm x 90mmOne double jointed armsOne accessories holdersFlat based, moveable cast holderCast holder features:Free-movingFlat baseAdjustable ball joint between base & holderOne fixed model gripsOne adjustable model gripShank diameter of accessories: 3mmUnit supplied with complete accessory set:Lead marker holderLeadWax knifeBracing holderBracing holderAnalysing pinUndercut measuring pin: 0,25mmUndercut measuring pin: 0,50mmUndercut measuring pin: 0,75mm |  |  |  | Each
RT286-01-006 A (Additional Accessories) | Laboratory Equipment | Lead marker holder for the Surveyor for Models - Single Arm |  |  |  | Each
RT286-01-006 B (Additional Accessories) | Laboratory Equipment | Lead for the Surveyor for Models - Single Arm |  |  |  | Each
RT286-01-006 C (Additional Accessories) | Laboratory Equipment | Wax knife / scraper for the Surveyor for Models - Single Arm |  |  |  | Each
RT286-01-006 D (Additional Accessories) | Laboratory Equipment | Bracing holder for the Surveyor for Models - Single Arm |  |  |  | Each
RT286-01-006 E (Additional Accessories) | Laboratory Equipment | Analysing pin for the Surveyor for Models - Single Arm |  |  |  | Each
RT286-01-006 F (Additional Accessories) | Laboratory Equipment | Undercut measuring pin: 0,25mm for the Surveyor for Models - Single Arm |  |  |  | Each
RT286-01-006 G (Additional Accessories) | Laboratory Equipment | Undercut measuring pin: 0,50mm for the Surveyor for Models - Single Arm |  |  |  | Each
RT286-01-006 H (Additional Accessories) | Laboratory Equipment | Undercut measuring pin: 0,75mm for the Surveyor for Models - Single Arm |  |  |  | Each
RT286-01-007 | Laboratory Equipment | Surveyor for Models - Double Arm SPECIFICATION A benchtop, manually controlled unit utilised in the cobalt-chrome technique for model analysis, measuring undercuts and marking the equator. The unit shall have the following features:Base of unit: approx. 120mm x 90mmTwo double jointed armsTwo accessories holdersFlat based, moveable cast holderCast holder features:Free-movingFlat baseAdjustable ball joint between base & holderOne fixed model gripsOne adjustable model gripShank diameter of accessories: 3mmUnit supplied with complete accessory set:Lead marker holderLeadWax knifeBracing holderAnalysing pinUndercut measuring pin: 0,25mmUndercut measuring pin: 0,50mmUndercut measuring pin: 0,75mm |  |  |  | Each
RT286-01-007 A (Additional Accessories) | Laboratory Equipment | Lead marker holder for the Surveyor for Models - Double Arm |  |  |  | Each
RT286-01-007 B (Additional Accessories) | Laboratory Equipment | Lead for the Surveyor for Models - Double Arm |  |  |  | Each
RT286-01-007 C (Additional Accessories) | Laboratory Equipment | Wax knife / scraper for the Surveyor for Models - Double Arm |  |  |  | Each
RT286-01-007 D (Additional Accessories) | Laboratory Equipment | Bracing holder for the Surveyor for Models - Double Arm |  |  |  | Each
RT286-01-007 E (Additional Accessories) | Laboratory Equipment | Analysing pin for the urveyor for Models - Double Arm |  |  |  | Each
RT286-01-007 F (Additional Accessories) | Laboratory Equipment | Undercut measuring pin: 0,25mm for the surveyor for Models - Double Arm |  |  |  | Each
RT286-01-007 G (Additional Accessories) | Laboratory Equipment | Undercut measuring pin: 0,50mm for the surveyor for Models - Double Arm |  |  |  | Each
RT286-01-007 H (Additional Accessories) | Laboratory Equipment | Undercut measuring pin: 0,75mm for the surveyor for Models - Double Arm |  |  |  | Each
RT286-01-008 | Laboratory Equipment | Boiling-Out Unit 6 or 12 (individual pricing) flask capacitySPECIFICATION A castor mounted, free standing unit utilised in a dental laboratory for automatically boiling-out wax up, scalding and polymerisation during production in a dental laboratory. The unit shall have the following features:Primary application: Boiling out of denture flasksAdditional functions:Automatic spraying bmo fixed spray bar jetsManual spraying bmo hand shower on hoseScalding (Steam cleaning) of flasksPolymerisation of acrylic in denture flasksPolymerisation at 95°CPermanent connection to fixed water supply (optional)Automatic water level control & fillingHousing made of stainless steelSwitches & dials mounted on top-rear of housingMinimum features of control panel:ON / OFF switchTemperature control dialTimer for a “wax-out” programTimer for boiling-out programON / OFF switch for hand spray/sRegulating valve for sprayingA 24 hour timerWater level warning lightAt least one integrated, adjustable hand sprayHand spray adjustable between jet to gentle sprayTank made of stainless steelTank covered with perforated stainless steel lidsOutlet of tank in baseAccess to drain from the front of the unitPump features:Corrosion resistant pump: Seal-free with 5 year guaranteeElectric pump motor positioned above water levelChangeable filter before pump chamberFlask baskets:6 or 12 supplied per unitStainless steel constructionFitted with insulated plastic / rubber handlesSABS Plug : Must comply with both SANS 164-0 and SANS 164-1 |  |  |  | Each
RT286-01-008 A (Additional Accessories) | Laboratory Equipment | Descaling liquid for the Boiling-Out Unit 6 or 12 (individual pricing) flask capacity |  |  |  | Each
RT286-01-008 B (Additional Accessories) | Laboratory Equipment | Hand shower spray head for the Boiling-Out Unit 6 or 12 (individual pricing) flask capacity |  |  |  | Each
RT286-01-008 C (Additional Accessories) | Laboratory Equipment | Replacement flask baskets for the Boiling-Out Unit 6 or 12 (individual pricing) flask capacity |  |  |  | Each
RT286-01-009 | Laboratory Equipment | Blow-Off Nozzle - Rubber |  |  |  | Each
RT286-01-009 A (Additional Accessories) | Laboratory Equipment | Replacement ring clamp for the Blow-Off Nozzle - Rubber |  |  |  | Each
RT286-01-010 | Laboratory Equipment | Plaster Dispenser - Manual SPECIFICATION An electrically operated silo for the dry storge and controlled dispensing of plasters and stone in the dental laboratory. The unit shall have the following features:Suspended on the wall or bench mounted in a frameCapacity of the dispenser no less that 25kgConstruction of the dispenser:Outer housing made of stainless steel or plasticInner hopper of plastic for corrosion-free plaster / stone supplyBuilt-in electrical vibratorDispensing leverVibration action ensuring even dispensing of materialActivating dispensing lever setting vibrator into motionSABS Plug : Must comply with both SANS 164-0 and SANS 164-1 |  |  |  | Each
RT286-01-011 | Laboratory Equipment | Duplicating Unit - Addition Cured Two-Component SiliconeSPECIFICATION A benchtop, electrically controlled, silicone mixing and metering unit, utilised for duplicating plaster and plastic models in a dental laboratory. The unit shall have the following features:Unit to consist of:Precision mixing headTwo rapid action couplingsMixing nozzles x 100Socket spannerMetering test cards x 10Mechanical mixingDeliver a mixing ratio 1:1SABS Plug : Must comply with both SANS 164-0 and SANS 164-1 |  |  |  | Each
RT286-02-001 | Laboratory Equipment | Laboratory Equipment SPECIFICATION A benchtop, free standing hydraulic flask press utilised in the acrylic section of a dental laboratory during flasking procedures. The unit shall have the following features:Capacity: 1 to 3 flasksMinimum vertical clearance at least 220mmMinimum clearance between uprights 128mmPress plate diameter no less than120mmPress piston vertical stroke approximately 30mmPress piston to return automatically to base after ventingSeparate pump and press cylindersPressure gauge to measure applied pressureHydraulic pressure applied bmo horisontally operating leverMinimum pressure exerted is at 200 Atm (8000Kg)Venting and locking by means of separate leversMechanical clamping of flasks by means of a vertical adjustable spindle with horizontal handle on top |  |  |  | Each
RT286-02-002 A Additional accesories | Laboratory Equipment | Pack of mixing nozzles x 100 for Duplicating Unit - Addition Cured Two-Component Silicone |  |  |  | Each
RT286-02-003 | Laboratory Equipment | Model Drying CabinetSPECIFICATION A benchtop, electrically heated drying cabinet for heating up and drying duplicate models, utilised during production in a dental laboratory. The unit shall have the following features:Operating temperature must reach a minimum of 200°CPre-selection of required temperatureInterior lightingAccommodate at least 12 models at a timeAt least two removable shelvesFront lid width of cabinet interior(Optional)Full width window in lidWindow of heat resistant glassUnit housing of corrosion resistant steelSABS Plug : Must comply with both SANS 164-0 and SANS 164-1 |  |  |  | Each
RT286-02-003 A (Additional accessories ) | Laboratory Equipment | Removable shelf for the Model Drying Cabinet |  |  |  | Each
RT286-02-003 B (Additional accessories ) | Laboratory Equipment | Replacement globe for the Model Drying Cabinet |  |  |  | Each
RT286-02-003 C (Additional accessories ) | Laboratory Equipment | Replacement door seal for the Model Drying Cabinet |  |  |  | Each
RT286-02-004 | Laboratory Equipment | Flask Press - Manual: 1 - 3 FlasksSPECIFICATION A benchtop, free standing manually operated flask press utilised in the acrylic section of a dental laboratory during flasking procedures. The unit shall have the following features:Capacity: 1 to 3 flasksMinimum vertical clearance at least 150mmMinimum clearance between uprights 130mmPressing plate diameter no less than 120mmClamping of flasks and application of pressure by means of a vertical spindle with horizontal handle on topBase of Flask Press to be supplied with holes for mounting to working surface |  |  |  | Each
RT286-02-005 | Laboratory Equipment | Flask Press - Hydraulic: 1 - 3 FlasksSPECIFICATION A benchtop, free standing hydraulic flask press utilised in the acrylic section of a dental laboratory during flasking procedures. The unit shall have the following features:Capacity: 1 to 3 flasksMinimum vertical clearance at least 220mmMinimum clearance between uprights 128mmPress plate diameter no less than120mmPress piston vertical stroke approximately 30mmPress piston to return automatically to base after ventingSeparate pump and press cylindersPressure gauge to measure applied pressureHydraulic pressure applied bmo horisontally operating leverMinimum pressure exerted is at 200 Atm (8000Kg)Venting and locking by means of separate leversMechanical clamping of flasks by means of a vertical adjustable spindle with horizontal handle on top |  |  |  | Each
RT286-02-006 | Laboratory Equipment | Thermal Moulding Device: Positive Pressure TechniqueSPECIFICATION A benchtop, electronically controlled machine used in the thermal moulding production of appliances for all types of dental applications. The unit shall have the following features:Operating pressure: 2.5 BarPressure adjustable depending on type of material being utilisedUtilise compressed air for pressure moulding techniqueHandle all types of thermo-forming materials ranging from Shellac to mouth guard materialsAccommodate (mould) material thickness from 0.1 to 5mmMaterial diameter must be 125mmUnit suitable for manufacturing following items:Mouth guardsCopingsSpecial Impression TraysFluoride TraysRetainersPositionersTMJ AppliancesTemporariesModel DuplicationHave an electrical or infrared heating elementHeating to be uniform and consistentOverheating protection by automatic switch off of the heater by the end of the heating timeAll functions electronically controlled, checked and displayed on a digital display panelFollowing programmable functions:Heating timeHeating temperatureHolding temperatureCooling timeSABS Plug : Must comply with both SANS 164-0 and SANS 164-1 |  |  |  | Each
RT286-02-006 A Additional accessories | Laboratory Equipment | Replacement set of O-rings |  |  |  | Each
RT286-02-006 B Additional accessories | Laboratory Equipment | Replacement locking plate for material |  |  |  | Each
RT286-02-006 C Additional accessories | Laboratory Equipment | Replacement pellets for embedding models in moulding device |  |  |  | Each
RT286-03-001 | Laboratory Equipment | Electronic Wax Modelling InstrumentSPECIFICATION A benchtop, electronically controlled modelling instrument for carrying out waxing-up techniques for work ranging from crown and bridge to partial dentures. The unit shall have the following features:Control Unit:Temperature display by means of an LED panelTemperature setting by means of a sealed touch padCasing angled at 20°Connector for handpiece on front of casingHolder for handpiece & tips on rear of casingHolder to accommodate 6 modelling tipsTemperature setting in CelsiusProgrammable temperature from 50°C to 200°CHandpiece:Heat insulatedSealed against liquid waxNatural cork, non-irritant instrument handleVoltage supplied to handpiece: 12 V DCCable length: no shorter than 1,2 mUnit to be supplied with:Control unit with integrated handle & tip holderHeating handle with cableTransformerInstruction manual with spare parts list Set of at least 7 modelling tipsSABS Plug : Must comply with both SANS 164-0 and SANS 164-1 |  |  |  | Each
RT286-03-001 A Additional Accessories | Laboratory Equipment | Replacement modelling tips for the Electronic Wax Modelling Instrument |  |  |  | Each
RT286-03-002 B Additional Accessories | Laboratory Equipment | Replacement Heating handle with cable for the Electronic Wax Modelling Instrument |  |  |  | Each
RT286-03-002 | Laboratory Equipment | Electronic Wax HeaterSPECIFICATION A benchtop, 3-pot electronic wax heater utilised in dental laboratory. The unit shall have the following features:Unit with integrated:Minimum 3 Wax pots with lids,Heating elements,Control panel,Temperature display, andHinged coverSeparate temperature control on each potHeating element connect direct to each potTemperature range between 60°C and 110°CLED bar graph pot temperature displaySealed touch control padSABS Plug : Must comply with both SANS 164-0 and SANS 164-1 |  |  |  | Each
RT286-03-003 | Laboratory Equipment | Bunsen Burner Lab Dual Pipe – Propane Gas |  |  |  | Each
RT286-03-003 A Additional accessories | Laboratory Equipment | Replacement ring clamp for the Bunsen Burner Lab Dual Pipe – Propane Gas |  |  |  | Each
RT286-03-004 | Laboratory Equipment | Bunsen Burner Lab Dual Pipe – Natural GasSPECIFICATION A free standing, tabletop, dual pipe, Bunsen burner for NATURAL gas. The unit shall have the following features:Gas supply setting via rocking leverIndividual control of oxygen supply via setting wheelControl of flame intensity by means of 2 ascending pipes:Small pipe: Pilot flameLarge pipe: Main flameSturdy metal base for secure standBurner removable from baseBurner tilt of 30° to either side Pressure hose for natural gas for the Bunsen Burner Lab Dual Pipe – Natural GasRing Clamps for the Bunsen Burner Lab Dual Pipe – Natural Gas |  |  |  | Each
RT286-03-004 A Additonal Accessories | Laboratory Equipment | Replacement pressure hose for natural gas for the Bunsen Burner Lab Dual Pipe – Natural Gas |  |  |  | Each
RT286-03-004 B Additonal Accessories | Laboratory Equipment | Replacement Ring Clamps for the Bunsen Burner Lab Dual Pipe – Natural Gas |  |  |  | Each
RT286-03-005 | Laboratory Equipment | Burner Outfit, Gas: Fully Automatic Ignition & Cut-OffSPECIFICATION A benchtop, electronically controlled fully automatic dental laboratory gas burner for utilisation in techniques requiring heating of hand held instruments. The unit shall have the following features:Connection to normal gas supplyConnection to electrical networkBurner operating on an infrared optical electronic principleFlame activated by hand movement over infrared sensorInfrared sensor not influenced by external lightFlame extinguishes automatically when burner is not in useSelector switch for automatic control or continuous burningAdjustable air/gas mixtureAdjustable flame sizeReduction of gas consumption by approximately 80%SABS Plug : Must comply with both SANS 164-0 and SANS 164-1 Pressure hose for gas for the Burner Outfit, Gas: Fully Automatic Ignition & Cut-Off Ring Clamps for the Burner Outfit, Gas: Fully Automatic Ignition & Cut-Off |  |  |  | Each
RT286-03-005 A Additional accessories | Laboratory Equipment | Replacement pressure hose for gas for the Burner Outfit, Gas: Fully Automatic Ignition & Cut-Off |  |  |  | Each
RT286-03-005 B Additional accessories | Laboratory Equipment | Replacement Ring Clamps for the Burner Outfit, Gas: Fully Automatic Ignition & Cut-Off |  |  |  | Each
RT286-03-006 | Laboratory Equipment | Micro Torch – Butane GasSPECIFICATION A refillable hand held butane torch utilised during dental laboratory work. The unit shall have the following features:Electronic trigger for flame ignitionFlame size adjustable bmo of adjusting handle/leverAir adjustable by means of adjusting handle/leverSafety Lock StopperIntegrated Butane Gas ReservoirFlame temperature no less than 1300°Cvertical or horisontal operating positionOperating time per refill between 30 –35 minutes |  |  |  | Each
RT286-03-006 A Additional Accessories | Laboratory Equipment | Bottle Butane gas refill for the Micro Torch – Butane Gas ( 300ml canister) |  |  |  | Each |  | Not Applicable
RT286-03-007 | Laboratory Equipment | Induction Heater Portable for Hand InstrumentsSPECIFICATION An electrical heater for dental hand instruments, in the form of a compact induction heating system replacing gas flames. The unit shall have the following features:Unit must be electrically operated - not gas Main supply : 230VElectromagnetic induction heating - no gas flame required No heat emmision into the environmentExternal sensors should emit heating only when there is no instrument in the chamber The modeling instrument must be heated inless than 3 seconds.SABS Plug : Must comply with both SANS 164-0 and SANS 164-1 |  |  |  | Each
RT286-03-007 A Additonal Accessories | Laboratory Equipment | Replacement silicone cup to protect the chamber |  |  |  | Each
RT286-03-008 | Laboratory Equipment | Soldering and Melting Torch: Microflame – Propane GasSPECIFICATION A hand held soldering and melting torch utilised for soldering and melting of all types of alloys during dental laboratory work, utilising PROPANE (LP-Liquid Petroleum) gas / oxygen mix. The unit shall have the following features:Operate with propane gas (liquid petroleum) and oxygenAccessories for installation shall consist of:Shut-off valve for propaneAnti-explosion device for propane gas at pressures between 20millibar and 1.5barPressure reducing valve for oxygenAnti-explosion device for oxygenSoldering and melting unit shall consist of:Handle2 m pressure hose for oxygen2 m pressure hose for propane gasSoldering insert, complete, with needle jets dia: 0.9, 1.2, 1.6 mmMelting insert for propane /oxygenMelting insert nozzle |  |  |  | Each
RT286-03-008 A Additional Accessories | Laboratory Equipment | Replacement needle jets diameter: 0.9, 1.2, 1.6 mm for the Induction Heater Portable for Hand Instruments |  |  |  | Each
RT286-03-008 B Additional Accessories | Laboratory Equipment | Replacement pressure hose for oxygen for the Induction Heater Portable for Hand Instruments |  |  |  | Each
RT286-03-008 C Additional Accessories | Laboratory Equipment | Replacement pressure hose for propane gas for the Induction Heater Portable for Hand Instruments |  |  |  | Each
RT286-03-008 D Additional Accessories | Laboratory Equipment | Replacement Ring Clamps for the Induction Heater Portable for Hand Instruments |  |  |  | Each
RT286-03-009 | Laboratory Equipment | Soldering and Melting Torch: Microflame – Natural GasSPECIFICATION A hand held soldering and melting torch utilised for soldering and melting of all types of alloys during dental laboratory work, utilising NATURAL (Acetelene) gas / oxygen mix. The unit shall have the following features:Operate with natural gas (acetelene) and oxygenAccessories for installation shall consist of:Shut-off valve for natural gasAnti-explosion device for natural gas at pressures between 20 millibar and 1.5 barPressure reducing valve for oxygenAnti-explosion device for oxygenSoldering and melting unit shall consist of:Handle2 m pressure hose for oxygen2 m pressure hose for natural gasSoldering insert, complete, with needle jets dia: 0.9, 1.2, 1.6 mmMelting insert for natural gas /oxygenMelting insert nozzle |  |  |  | Each
|  | Additional Accessories: |  |  |  | Each
RT286-03-009 A Additional Accessories | Laboratory Equipment | Replacement needle jets diameter: 0.9, 1.2, 1.6 mm for the Soldering and Melting Torch: Microflame – Natural Gas |  |  |  | Each
RT286-03-009 B Additional Accessories | Laboratory Equipment | Replacement pressure hose for oxygen for the Soldering and Melting Torch: Microflame – Natural Gas |  |  |  | Each
RT286-03-009 C Additional Accessories | Laboratory Equipment | Replacement pressure hose for natural gas for the Soldering and Melting Torch: Microflame – Natural Gas |  |  |  | Each
RT286-03-009 D Additional Accessories | Laboratory Equipment | Replacement Ring Clamps for the Soldering and Melting Torch: Microflame – Natural Gas |  |  |  | Each
RT286-04-001 | Laboratory Equipment | Vacuum Mixer: Plaster & Investment MaterialSPECIFICATION A wall mounted or benchtop mechanical vacuum mixer for the mixing of all dental plasters and investment compounds utilised in a dental laboratory. The unit shall have the following features:Single speed vertical drive shaft, approx. 450rpm’sPositive drive chuck to activate rotation of mixerRust resistant motorIntegrated vacuum pump, oiler unit and vacuum hose with coupling to mixing bowl providedON/OFF switch on the unit within easy reach of operatorMotor drive and vacuum activated by insertion of mixing bowlTwo mixing bowls provided per unit:Approximately 550mlApproximately 800ml SABS Plug : Must comply with both SANS 164-0 and SANS 164-1 |  |  |  | Each
RT286-04-001 A Additional accessories | Laboratory Equipment | Mixing bowls approximately 200ml - 1200ml for the Vacuum Mixer: Plaster & Investment Material |  |  |  | Each
RT286-04-001 B Additional accessories | Laboratory Equipment | Replacement mixing bowl O-Ring for the Vacuum Mixer: Plaster & Investment Material |  |  |  | Each
RT286-04-001 C Additional accessories | Laboratory Equipment | Replacement O-Ring for suction/vacuum nozzle for the Vacuum Mixer: Plaster & Investment Material |  |  |  | Each
RT286-04-002 | Laboratory Equipment | Vacuum Mixer: Plaster, Investment & Silicone Duplicating MaterialSPECIFICATION A wall mounted or benchtop mechanical vacuum mixer for the mixing of all dental plasters, investment compounds and silicone duplicating material utilised in a dental laboratory. The unit shall have the following features:Usable as either wall mounted or bench unit.Single speed vertical drive shaft, approx. 100 - 450rpm’sVacuum can be regulated from 70 - 100%Vacuum created by motor pump and not compressed air.Vacuum capacity 16 litre per minuteElectronic ControlsFunctions should include prespatulation, prevacuum, reverse rotation, post vacuum.Maintenance-free diaphragm pumpAccommodate three sizes mixing bowls: approx 250, 550, 850mlTwo mixing bowls provided per unit:Approximately 550mlApproximately 800ml SABS Plug : Must comply with both SANS 164-0 and SANS 164-1 |  |  |  | Each
RT286-04-002 A Additonal Accessories | Laboratory Equipment | Mixing bowls approximately 200ml - 1200ml for the Vacuum Mixer: Plaster, Investment & Silicone Duplicating Material |  |  |  | Each
RT286-04-003 B Additonal Accessories | Laboratory Equipment | Replacement mixing bowl O-Ring for the Vacuum Mixer: Plaster, Investment & Silicone Duplicating Material |  |  |  | Each
RT286-04-004 C Additonal Accessories | Laboratory Equipment | Replacement O-Ring for suction/vacuum nozzle for the Vacuum Mixer: Plaster, Investment & Silicone Duplicating Material |  |  |  | Each
| Laboratory Equipment | Laboratory Vibrator SPECIFICATION A dental laboratory vibrator utilised during mixing of plaster and investment material, the casting of models and pouring of investment materials. The unit shall have the following features:Small: Size no smaller than 200 x 200mm ON / OFF switch on front of control panelElectronically regulated variable speed controlHeavy Duty rubber cover on vibrating surfaceRubber cover designed with raised edge on outer rimUniform vibration to all parts of the trayFitted with vibration free rubber mountingsSABS Plug : Must comply with both SANS 164-0 and SANS 164-1 |  |  |  | Each
RT286-04-003 A Additional Accessories | Laboratory Equipment | Replacement Rubber cover for the Vacuum Mixer |  |  |  | Each
RT286-04-004 | Laboratory Equipment | Dispenser, Dental Acrylic Powder and Liquid: Orthodontic SPECIFICATION A benchtop, dispensing mechanism for the exact proportioning of acrylic powder and liquid in the removeable orthodontic application. The unit shall have the following features:Operate without compressed air or electricityDispensing Unit consists of:Powder dispenser for approximately 500gLiquid dispenser for approximately 500mlDispensers positioned above working area for gravity feedTwo pedal switches for operating dispensersMetal nozzlesExtraction system can be fittedAll sides of unit enclosed except frontCarbon filter to eliminate monomer vapoursIntegration of at least 4 dispensers |  |  |  | Each
RT286-04-004 A Accessories | Laboratory Equipment | Replacement metal nozzles for the Dispenser, Dental Acrylic Powder and Liquid: Orthodontic |  |  |  | Each
RT286-04-004 B Accessories | Laboratory Equipment | Replacement flexible pipes for the Dispenser, Dental Acrylic Powder and Liquid: Orthodontic |  |  |  | Each
RT286-04-004 C Additonal Accessories | Laboratory Equipment | Replacement carbon filter for the Dispenser, Dental Acrylic Powder and Liquid: Orthodontic |  |  |  | Each
RT286-04-004 D Accessories | Laboratory Equipment | Replacement pedal switches for the Dispenser, Dental Acrylic Powder and Liquid: Orthodontic |  |  |  | Each
RT286-04-005 | Laboratory Equipment | Polymerisation Bath – Electric SPECIFICATION A benchtop, fully automatic, electrically heated water or steam bath utilised as a polymerisation apparatus fitted with a thermostat and designed for long duration polymerisation of any resin in a flask. The unit shall have the following features:Capacity minimum of 10 flasksProtected heating elementOverload protectionOptical display of functionsON / OFF power switch on control panelHousing of stainless steelIntegrated hinged lidHeat insulated lid handleEntire polymerisation process fully automatic and programmableProgammable two-stage polymerisationProgramming options:Preselection of pre-polymerisation temperaturePreselection of duration of pre-polymerisation up to 12 hoursPreselection of final polymerisation temperaturePreselection of duration of final polymerisation up to 12 hoursUnit to switch off automatically at end of programmeSABS Plug : Must comply with both SANS 164-0 and SANS 164-1 |  |  |  | Each
RT286-04-006 | Laboratory Equipment | Polymerisation Vessel – Pneumatic SPECIFICATION A benchtop, manually operated pressure polymerisation apparatus utilised for polymerisation of any cold curing resins. The unit shall have the following features:Unit consists of:VesselLidHose & fitting to pressurise vesselCapacity: minimum of 4 denturesOperating pressure approx 2 barInside of vessel coated with anti-corrosion teflon coatingThe lid shall have the following features:Safety lock to prevent accidental openingLocking & pressure release switchPressure gaugeInlet valve for airPressure over-load safety valve |  |  |  | Each
RT286-04-006 A Accessoires | Laboratory Equipment | Replacement lid seal for the Polymerisation Vessel – Pneumatic |  |  |  | Each
RT286-04-007 | Laboratory Equipment | Polymerisation Vessel - Pneumatic & Heat SPECIFICATION A benchtop, fully automatic, thermo-regulated water heated polymerisation apparatus fitted with thermostat and designed for polymerisation of any cold curing resins. The unit shall have the following features:Unit consists of:Casing with pressure vesselIntegrated hinged lid with locking mechanismIntegrated safety featuresPermanent connection to compressed air supplyCasing, chamber & lid of corrosion resistant materialCapacity minimum 2 modelsOperating pressure maximum 5 barAutomatically sealing system for lidThermostatically regulated temperatureTemperature manually adjustableMaximum temperature 95°CUnit fitted with temperature indicator lightUnit fitted with pressure gaugeUnit fitted with a water drain valveUnit fitted with separate ON / OFF switchSABS Plug : Must comply with both SANS 164-0 and SANS 164-1 |  |  |  | Each
RT286-04-007 A | Laboratory Equipment | Replacement lid seal/O-ring for the Polymerisation Vessel - Pneumatic & Heat |  |  |  | Each
RT286-04-007 B | Laboratory Equipment | Replacement ring clamp for the Polymerisation Vessel - Pneumatic & Heat |  |  |  | Each
RT286-04-008 | Laboratory Equipment | Laboratory Vibrator SPECIFICATION A dental laboratory vibrator utilised during mixing of plaster and investment material, the casting of models and pouring of investment materials. The unit shall have the following features: Large : No larger than 270 x 380mm ON / OFF switch on front of control panelElectronically regulated variable speed controlHeavy Duty rubber cover on vibrating surfaceRubber cover designed with raised edge on outer rimUniform vibration to all parts of the trayFitted with vibration free rubber mountingsSABS Plug : Must comply with both SANS 164-0 and SANS 164-1 |  |  |  | Each
RT286-04-008 A Accessories | Laboratory Equipment | Replacement Rubber cover for the Vacuum Mixer |  |  |  | Each
RT286-05-001 | Laboratory Equipment | Pre-Heating & Burnout Furnace – SmallSPECIFICATION A benchtop, pre-heating and burnout furnace with electrical heating elements and electronic temperature controls, fitted with an extraction and catalyser system, utilised for the heating and burning out of casting rings for crowns, bridges and metal denture frames. The unit shall have the following features:Maximum temperature: 1100°CFurnace chamber volume: 3 litreFurnace chamber surface area: approx. 150mm x 150mmStainless steel furnace casingSolvent-free high grade enamel paintTouch-safe outer wall temperatureInsulation by means of a fibre moduleFurnace door to hinge downwards to horizontal positionObservation hole in doorHeating elements encased in ceramic platesHeating from at least two sidesExhaust duct in rear wall of furnaceExternal mounting device and connection for vapour vent with fan or catalyser by exhaust ductFurnace fitted vapour vent and fanFan connection to program controllerIntegrated program controller underneath furnace chamberFeatures of integrated program controller:Microprocessor controlledStorage for at least 5 programmesAt least 3 heating ramps & holding times per programProgram steps: 1°C in 1 minuteAudible alarmReal-time clock for programmingLED temperature displayTouch control keypadKeypad protected from dustUnit supplied with:Steel collecting tray x 1Ceramic collecting tray x 1Ready for use with 2 m electrical cable & plugSABS Plug : Must comply with both SANS 164-0 and SANS 164-1 |  |  |  | Each
RT286-05-001 A Accessories | Laboratory Equipment | Replacement heating elements ofr the Pre-Heating & Burnout Furnace – Small |  |  |  | Each
RT286-05-001 B Accessories | Laboratory Equipment | Replacement thermocouple for the Pre-Heating & Burnout Furnace – Small |  |  |  | Each
RT286-05-001 C Accessories | Laboratory Equipment | Replacement steel and ceramic collecting trays for the Pre-Heating & Burnout Furnace – Small |  |  |  | Each
RT286-05-002 | Laboratory Equipment | Pre-Heating & Burnout Furnace – MediumSPECIFICATION A benchtop, pre-heating and burnout furnace with electrical heating elements and electronic temperature controls, fitted with an extraction and catalyser system, utilised for the heating and burning out of casting rings for crowns, bridges and metal denture frames. The unit shall have the following features:Maximum temperature: 1100°CFurnace chamber volume: 5 litreFurnace chamber surface area: approx. 220mm x 180mmStainless steel furnace casingSolvent-free high grade enamel paintTouch-safe outer wall temperatureInsulation by means of a fibre moduleFurnace door to hinge downwards to horizontal positionObservation hole in doorHeating elements encased in ceramic platesHeating from at least two sidesExhaust duct in rear wall of furnaceExternal mounting device and connection for vapour vent with fan or catalyser by exhaust ductFurnace fitted vapour vent and fanFan connection to program controllerIntegrated program controller underneath furnace chamberFeatures of integrated program controller:Microprocessor controlledStorage for at least 5 programmesAt least 3 heating ramps & holding times per programProgram steps: 1°C in 1 minuteAudible alarmReal-time clock for programmingLED temperature displayTouch control keypad Keypad protected from dustUnit supplied with:Steel collecting tray x 1Ceramic collecting tray x 1Ready for use with 2 m electrical cable & plugSABS Plug : Must comply with both SANS 164-0 and SANS 164-1 |  |  |  | Each
RT286-05-002 A Accessories | Laboratory Equipment | Replacement heating elements for the Pre-Heating & Burnout Furnace – Medium |  |  |  | Each
RT286-05-002 B Accessories | Laboratory Equipment | Replacement thermocouple for the Pre-Heating & Burnout Furnace – Medium |  |  |  | Each
RT286-05-002 C Accessories | Laboratory Equipment | Replacement steel and ceramic collecting trays for the Pre-Heating & Burnout Furnace – Medium |  |  |  | Each
RT286-05-003 | Laboratory Equipment | Pre-Heating & Burnout Furnace – Large SPECIFICATION A benchtop, pre-heating and burnout furnace with electrical heating elements and electronic temperature controls, fitted with an extraction and catalyser system, utilised for the heating and burning out of casting rings for crowns, bridges and metal denture frames. The unit shall have the following features:Maximum temperature: 1100°CFurnace chamber volume: 9 litreFurnace chamber surface area: approx. 250mm x 250mmStainless steel furnace casingSolvent-free high grade enamel paintTouch-safe outer wall temperatureInsulation by means of a fibre moduleFurnace door to hinge downwards to horisontal positionObservation hole in doorHeating elements encased in ceramic platesHeating from at least two sidesExhaust duct in rear wall of furnaceExternal mounting device and connection for vapour vent with fan or catalyser by exhaust ductFurnace fitted vapour vent and fanFan connection to program controllerIntegrated program controller underneath furnace chamberFeatures of integrated program controller:Microprocessor controlledStorage for at least 5 programmesAt least 3 heating ramps & holding times per programProgram steps: 1°C in 1 minuteAudible alarmReal-time clock for programmingLED temperature displayTouch control keypadKeypad protected from dustUnit supplied with:Steel collecting tray x 1Ceramic collecting tray x 1Ready for use with 2 m electrical cable & plugSABS Plug : Must comply with both SANS 164-0 and SANS 164-1 |  |  |  | Each
RT286-05-003 A | Laboratory Equipment | Replacement heating elements- for the Pre-Heating & Burnout Furnace – Large |  |  |  | Each
RT286-05-003 B | Laboratory Equipment | Replacement thermocouple for the Pre-Heating & Burnout Furnace – Large |  |  |  | Each
RT286-05-003 C | Laboratory Equipment | Replacement steel and ceramic collecting trays for the Pre-Heating & Burnout Furnace – Large |  |  |  | Each
RT286-05-004 | Laboratory Equipment | Casting Machine Centrifugal – Mechanical: Bench-Top SPECIFICATION A bench top mounted, mechanically activated centrifugal casting machine with articulated casting arm, utilised for flame casting of all precious and non-precious dental alloys in the dental laboratory. The unit shall have the following features:Central base housing with integrated spiral springDrive shaft with articulated casting armAdjustable counterweightAdjustable crucible and muffle holderSlideable and Lockable crucible holderMechanical activation mechanism for centrifuge |  |  |  | Each
RT286-05-004 A Accessories | Laboratory Equipment | Replacement spiral spring for the Casting Machine Centrifugal – Mechanical: Bench-Top |  |  |  | Each
RT286-05-004 B Accessories | Laboratory Equipment | Replacement crucible for the Casting Machine Centrifugal – Mechanical: Bench-Top |  |  |  | Each
RT286-05-004 C Accessories | Laboratory Equipment | Replacement crucible insert for the Casting Machine Centrifugal – Mechanical: Bench-Top |  |  |  | Each
RT286-05-005 | Laboratory Equipment | Casting Machine Centrifugal – Motorised SPECIFICATION A free standing unit, electrically activated motorised centrifugal casting machine with articulated casting arm, utilised for flame casting of all precious and non-precious dental alloys in the dental laboratory. The unit shall have the following features:Centrifuge housed in free standing unitDrive shaft with articulated casting armAdjustable counterweightAdjustable crucible and muffle holderCrucible holder lockableElectrical activation of motorised mechanism for centrifugeMotor housed below centrifugeRear hinged safety cover on top of unitSABS Plug : Must comply with both SANS 164-0 and SANS 164-1 |  |  |  | Each
RT286-05-005 A Accessories | Laboratory Equipment | Replacement crucible for the Casting Machine Centrifugal – Motorised |  |  |  | Each
RT286-05-005 B Accessories | Laboratory Equipment | Replacement crucible insert for the Casting Machine Centrifugal – Motorised |  |  |  | Each
RT286-05-005 C Accessories | Laboratory Equipment | Replacement ring/muffle cradle for the Casting Machine Centrifugal – Motorised |  |  |  | Each
RT286-05-006 | Laboratory Equipment | Casting Machine Centrifugal – Induction (Bench-Top Model) SPECIFICATION A bench top mounted, high frequency induction casting machine for casting of all precious and non-precious dental alloys in the dental laboratory. The unit shall have the following features:Central base housing with integrated electric motor driveDrive shaft with casting armAdjustable counterweightAdjustable crucible and muffle holderCrucible holder lockableTinted inspection port on lid of unitElectrical activation of motorised mechanism for centrifugeInduction coil surrounding crucibleWater cooled unitSABS Plug : Must comply with both SANS 164-0 and SANS 164-1 |  |  |  | Each
RT286-05-006 A Additional Accessories | Laboratory Equipment | Replacement spiral heating coil/element for the Casting Machine Centrifugal – Induction (Bench-Top Model) |  |  |  | Each
RT286-05-006 B Additional Accessories | Laboratory Equipment | Replacement crucible for the Casting Machine Centrifugal – Induction (Bench-Top Model) |  |  |  | Each
RT286-05-006 C Additional Accessories | Laboratory Equipment | Replacement crucible insert for the Casting Machine Centrifugal – Induction (Bench-Top Model) |  |  |  | Each
RT286-05-006 D Additional Accessories | Laboratory Equipment | Replacement muffle/ring cradle for the Casting Machine Centrifugal – Induction (Bench-Top Model) |  |  |  | Each
RT286-05-006 E Additional Accessories | Laboratory Equipment | Replacement coloured observation port glass for the Casting Machine Centrifugal – Induction (Bench-Top Model) |  |  |  | Each
RT286-05-007 | Laboratory Equipment | Induction Casting Machine - Vacuum Pressure and High Frequency Melting SPECIFICATION Complete with dual optic touch free sensor to read temperature accurately to prevent overheating of alloy. Automatic and manual operation.Must store up to 640 casting logsRated voltage 220V 50/60 HzPower consumption 14 ampsMelting power 3.6kva 65khzCompressed air minimum 75psiWater consumption approximately 1.5 gal.minSABS Plug : Must comply with both SANS 164-0 and SANS 164-1 |  |  |  | Each
RT286-05-007 A Accessories | Laboratory Equipment | Replacement spiral heating coil/element for the Induction Casting Machine - Vacuum Pressure and High Frequency Melting |  |  |  | Each
RT286-05-007 B Accessories | Laboratory Equipment | Replacement crucible for the Induction Casting Machine - Vacuum Pressure and High Frequency Melting |  |  |  | Each
RT286-05-007 C Accessories | Laboratory Equipment | Replacement crucible insert for the Induction Casting Machine - Vacuum Pressure and High Frequency Melting |  |  |  | Each
RT286-05-007 D Accessories | Laboratory Equipment | Replacement muffle/ring cradle for the Induction Casting Machine - Vacuum Pressure and High Frequency Melting |  |  |  | Each
RT286-05-007 E Accessories | Laboratory Equipment | Replacement coloured observation port glass for the Induction Casting Machine - Vacuum Pressure and High Frequency Melting |  |  |  | Each
RT286-06-001 | Laboratory Equipment | Blast Cleaning Machine Cobalt–Chrome (Automatic) SPECIFICATION A benchtop, cobalt-chrome sandblaster unit for automatic blasting of up to 6 castings or manual blasting of individual castings, with integrated timer, regulator, gauges and multi-stage filtering capability. The unit shall have the following features:Integrated multi-stage filtering capability (at least four stages)On board filter module for self contained / independent operationPre-filter collection capacity approx. 1 kgFilter module operates automatically when blaster is switched onWarning signal when filter is fullSafety cut-out activated until full pre-filter bag is replacedTank capacity no less than 4,5 kgRe-circulating system for blasting materialUtilise different grades (micron) blasting materialAir consumption no less than approx. 100 l/minCompressed air connection between 4-6 barLighting: At least one lighting unitLighting unit/s: No less than 15 WattHousing of blaster:Plastic, CFC freeCorrosion resistantAngled viewing windowApertures sealed with resistant stretch material sleeveApertures sealed with resistant stretch material sleeveRotating sandblasing drum for at least 6 cobalt chrome framesDrum mounted in-line with with sandblasting nozzleON / OFF switch for rotating drumTimer for rotating drumSABS Plug : Must comply with both SANS 164-0 and SANS 164-1 |  |  |  | Each
RT286-06-001 A Accessories | Laboratory Equipment | Replacement rotating drum for the Blast Cleaning Machine Cobalt–Chrome (Automatic) |  |  |  | Each
RT286-06-001 B Accessories | Laboratory Equipment | Replacement Filter bags for the Blast Cleaning Machine Cobalt–Chrome (Automatic) |  |  |  | Each
RT286-06-001 C Accessories | Laboratory Equipment | Replacement sleeves for the Blast Cleaning Machine Cobalt–Chrome (Automatic) |  |  |  | Each
RT286-06-001 D Accessories | Laboratory Equipment | Replacement gloves for the Blast Cleaning Machine Cobalt–Chrome (Automatic) |  |  |  | Each
RT286-06-001 E Accessories | Laboratory Equipment | Replacement Protective glass for the Blast Cleaning Machine Cobalt–Chrome (Automatic) |  |  |  | Each
RT286-06-001 F Accessories | Laboratory Equipment | Replacement light bulbs for the Blast Cleaning Machine Cobalt–Chrome (Automatic) |  |  |  | Each
RT286-06-002 | Laboratory Equipment | Blast Cleaning Machine Crown, Bridge & Cobalt - Chrome (Manual) SPECIFICATION A benchtop, sandblaster unit for manual blasting of investment material in crown, bridge and partial denture work. The unit shall have the following features:Must have an extraction system for unusable fine dust to be extracted.Must have a minimum of two blasting tanks.One Blasting tank for 20 – 70 micron aluminium oxide.One Blasting tank for 70 – 250 micron aluminium oxide.Must have a minimum of two Sandblasting nozzles:One Sandblasting nozzle 0.8mm for 20 – 70 micron aluminium oxide.One Sandblasting nozzle 1.2mm for 70 – 250 micron aluminium oxide.SABS Plug : Must comply with both SANS 164-0 and SANS 164-1 ( SANS 164-2 ??) |  |  |  | Each
RT286-06-002 A Accessories | Laboratory Equipment | Replacement fixed nozzles for the Blast Cleaning Machine Crown, Bridge & Cobalt - Chrome (Manual) |  |  |  | Each
RT286-06-002 B- Accessories | Laboratory Equipment | Replacement sleeves for the Blast Cleaning Machine Crown, Bridge & Cobalt - Chrome (Manual) |  |  |  | Each
RT286-06-002 C Accessories | Laboratory Equipment | Replacement gloves for the Blast Cleaning Machine Crown, Bridge & Cobalt - Chrome (Manual) |  |  |  | Each
RT286-06-002 D Accessories | Laboratory Equipment | Filter bags for extraction unit for the Blast Cleaning Machine Crown, Bridge & Cobalt - Chrome (Manual) |  |  |  | Each
RT286-06-003 | Laboratory Equipment | Blast Cleaning Machine Crown & Bridge – Micro Sandblaster (Manual) SPECIFICATION A benchtop, micro sandblaster with blasting nozzles / pencil jets for manual blasting and polishing of crown & bridge prosthetic units. The unit shall have the following features:Compact size: No more than 400 x 300 x 560 mmCasing exterior & interior corrosion proof & wear resistantMinimum two sizes blasting nozzles / pencil jets with sand reservoirsDeliver 25-50 micron and 90-250 micron sandCapability to expand unit with two more sets of sand reservoirs and blasting nozzlesFoot switch controlled pneumatic valveImmediate Stop (IS) functionBlasting chamber internal features:Abrasion coatedFull tank width sieve above contaminated sand accumulation areaTank selection switchCompressed air regulatorFull tank width observation windowProtection screen for observation windowBuilt-in illuminationIllumination minimum 2400 LUXBlasting nozzles & hosesColour coded blasting unit, tank selector switch & blasting handlesConstant air pressure in reservoirs and hosesReservoir maintain pressure after operationValve per reservoir for manual depressurising before openingPressure gauge on top of casingTwo operating aperturesSturdy gloves fitted to aperturesVentilation slitsNoise Level: Less than 70 dB(A)Two sizes spare nozzles supplied with each unitSABS Plug : Must comply with both SANS 164-0 and SANS 164-1 |  |  |  | Each
RT286-06-003 A Accessories | Laboratory Equipment | Replacement light bulbs for the Blast Cleaning Machine Crown & Bridge – Micro Sandblaster (Manual) |  |  |  | Each
RT286-06-003 B Accessories | Laboratory Equipment | Protective gloves for the Blast Cleaning Machine Crown & Bridge – Micro Sandblaster (Manual) |  |  |  | Each
RT286-06-003 C Accessories | Laboratory Equipment | Replacement fixed nozzles for the Blast Cleaning Machine Crown & Bridge – Micro Sandblaster (Manual) |  |  |  | Each
RT286-06-003 D Accessories | Laboratory Equipment | Replacement protective glass for the Blast Cleaning Machine Crown & Bridge – Micro Sandblaster (Manual) |  |  |  | Each
RT286-06-003 E Accessories | Laboratory Equipment | Replacement foot switch for pneumatic valve for the Blast Cleaning Machine Crown & Bridge – Micro Sandblaster (Manual) |  |  |  | Each
RT286-06-004 | Laboratory Equipment | Dental Laboratory Dust Extraction System (Dry) – PortableSPECIFICATION A portable dust extractor for one or two stations that can be used when working with handpieces, benchtop grinders or other on-bench appliances to prevent solid particles or dust particles from contaminating the working environment. The unit shall have the following features:Suction Unit with two speed settingsTri-layered disposable filter bag for suction unitAccumulator for solids and most dustConnecting hose (+-5cm) between accumulator and suction unitT-Piece for accumulator2 x Hoses (+-5cm) for work stations2 x Work station attachments2 x Remote control for suction unitSABS Plug : Must comply with both SANS 164-0 and SANS 164-1 |  |  |  | Each
RT286-06-004 A Accesories | Laboratory Equipment | Replacement filter bag |  |  |  | Each
RT286-06-004 B Accessories | Laboratory Equipment | Replacement ring clasp |  |  |  | Each
RT286-06-005 | Laboratory Equipment | Dental Laboratory Dust Extraction System (Dry) – Mobile SPECIFICATION A mobile (with castors), dust extractor for lab work stations, blasting units and dry model trimmers, or any other on-bench appliances to prevent solid or dust particles from contaminating the working environment. The unit shall have the following features:Mobile unit on castorsBrushless motor: EC no-contact energy transferLifespan of motor: Approx 800 hoursMotor noise level: No more than 62 dB (A)Two stage filter systemParticle capture rate: 99,9%Filter material to capture health endangering dust of MAK value smaller than 1 mg/ cubic mCapacity of filter bag: 12kgOn-board automatic activator switchEach unit shall be supplied with:Tubing: 3mCloth Filter: x1Replacement filter bags: 1 set of 5SABS Plug : Must comply with both SANS 164-0 and SANS 164-1 |  |  |  | Each
RT286-06-005 A Accessories | Laboratory Equipment | Replacement filter bag for the Dust Extraction System (Dry) – Mobile |  |  |  | Each
RT286-06-005 B Accessories | Laboratory Equipment | Replacement cloth filter for the Dust Extraction System (Dry) – Mobile |  |  |  | Each
RT286-06-006 | Laboratory Equipment | Dental Laboratory Dust Extraction System (Wet) – Fixed SPECIFICATION An in-line wet dust extraction system utilised when working with benchtop grinders and other benchtop appliances, for the elimination of gypsum, metal and silicon carbide particles from the air. The unit shall have the following features:Primary filtering capability by means of waterSilent running of drive motorT-Piece and hoses to connect to at least two waste particle producing appliancesSABS Plug : Must comply with both SANS 164-0 and SANS 164-1 |  |  |  | Each
RT286-06-006 A Accessories | Laboratory Equipment | Replacement water filter bag for the Dust Extraction System (Wet) – Fixed |  |  |  | Each
RT286-06-006 B Accessories | Laboratory Equipment | Replacement ring clasp for the Dust Extraction System (Wet) – Fixed |  |  |  | Each
RT286-06-007 | Laboratory Equipment | Dental Laboratory Dust Extraction Drawer System (Dry) – Fixed SPECIFICATION An under-benchtop dust extraction drawer mounted below a lab workbench to suction gypsum, metal, acrylic and silicon carbide particles generated during working procedures of a dental technician. The unit shall have the following features:Fitted under workbench, without hindering freedom of movementMaximum dimensions: H180mm, W490mmMinimum extension: 500mmSuction hood and handrestAutomatic start-up linked to micromotorAutomatic switch-off after short micromotor pauseWorking surface: WoodMaintenance free induction motorFilter bag:Quick change capabilityLarge dimensional (Full drawer width)Unit supplied with two spare filter bagsSABS Plug : Must comply with both SANS 164-0 and SANS 164-1 |  |  |  | Each
RT286-06-007 A Accessories | Laboratory Equipment | Replacement filter bag for the Dust Extraction Drawer System (Dry) – Fixed |  |  |  | Each
RT286-06-008 | Laboratory Equipment | Suction Hood - Electrical SPECIFICATION A benchtop or over-bench mounted hood for use in a furnace room to catch and extract fumes, smoke and heat generated by furnaces. Hood to be connected to existing ventilation. The unit shall have the following features:Dimensions: At least 1 500mm wide and 700mm deepStainless steel constructionDesign to expedite removal of all fumes, smoke and heatElectrical fan mounted in top of hoodConnection to existing ventilation with flexible pipesLighting inside hood for improved visibilityON / OFF switch with pilot light installed on front of hoodAdditional carbon filterSABS Plug : Must comply with both SANS 164-0 and SANS 164-1 |  |  |  | Each
RT286-06-008 A | Laboratory Equipment | Replacement carbon filter for the Suction Hood - Electrical |  |  |  | Each
RT286-06-008 B | Laboratory Equipment | Replacement light bulb for the Suction Hood - Electrical |  |  |  | Each
RT286-07-001 | Laboratory Equipment | Pneumatic Chisel / Handpiece for Trimming of Plaster & Investment SPECIFICATION A pneumatic deflasking chisel / handpiece for the removal of plaster and investment material. The unit shall have the following features:Operating pressure 4-6 BarPiston frequency approximately 130 HzAdjustable power on chisel itselfRemovable and interchangeable chiselsHeavy duty components made of High-Tech steel with hardcoat layerFull strike power transferred straight to chisel insertNon-slip handleLight weight construction, approximately 400gSupplied with at least 3 different chiselsRing clasp for the Pneumatic Chisel / Handpiece for Trimming of Plaster & Investment |  |  |  | Each
RT286-07-001 A Accessories | Laboratory Equipment | Replacement chisel bits for the Pneumatic Chisel / Handpiece for Trimming of Plaster & Investment |  |  |  | Each
RT286-07-001 B Accessoties | Laboratory Equipment | Replacement ring clasp for the Pneumatic Chisel / Handpiece for Trimming of Plaster & Investment |  |  |  | Each
RT286-07-002 | Laboratory Equipment | Steam Cleaner-Portable SPECIFICATION A benchtop, electrically operated portable wet and dry steam cleaner utilised during production in a dental laboratory. The unit shall have the following features:Water tank capacity no less than 4.5 litresPortable due to manual filling of tank, no fixed connection to water supplyNo less than 4.5Bar of wet or dry steamReaching operating pressure no longer than 15 minutesAn adjustable steam pressure knob to accurately select flow level of steam requiredConstruction of heavy duty stainless steelThe following indicators / switches to be provided:Low water level light indicatorWater level alarm switchAutomatic pressure gauge protection switchAutomatic thermal safety system heating switchSABS Plug : Must comply with both SANS 164-0 and SANS 164-1 |  |  |  | Each
RT286-07-002 A Accessories | Laboratory Equipment | One spare pressure cap supplied per unit for the Steam Cleaner-Portable |  |  |  | Each
RT286-07-002 B Accessories | Laboratory Equipment | Replacement light bulb for the Steam Cleaner-Portable |  |  |  | Each
RT286-07-003 | Laboratory Equipment | Steam Cleaner-Fixed: Wet & Dry (High Performance) SPECIFICATION A benchtop, electrically operated fixed wet and dry steam cleaner utilised during production in a dental laboratory. The unit shall have the following features:The steam cleaner shall be mainly utilised for:Removing wax & polish residuesCleaning models & traysDegreasing metal components for ceramic applicationA selection can be made between:Wet steamDry steam (saturated steam)Mixture dry steam & waterUnit is fixed - permanently connected to main water supplySteam pressure approx 3 barSpray gun:Push button pressure switch on gunPressure switch for thumb useInsulated gun housingFlexible insulated hose from gun to unitFlexible hose at least 1 m longUnit equipped with demineralisation cartridgeUnit housing of corrosion resistant steel & plasticPressure gauge visible on front of housingDry/wet/water selector switch on front of housingOperating indicator LED’s on front of housingSafety features:Water leak monitoring sensor built into unitLeak sensor shuts off water supply automaticallyReleasing gun pressure switch cuts off steamInternal connections shall be copper pipes and threaded joints for safetyEmpty demineralisation cartridge shuts off water supplyRing spannerSABS Plug : Must comply with both SANS 164-0 and SANS 164-1 |  |  |  | Each
RT286-07-003 A Accessories | Laboratory Equipment | Ring spanner for the Steam Cleaner-Fixed: Wet & Dry (High Performance) |  |  |  | Each
RT286-07-003 B Accessories | Laboratory Equipment | Replacement full demineralisation cartridge for the Steam Cleaner-Fixed: Wet & Dry (High Performance) |  |  |  | Each
RT286-07-003 C Accessories | Laboratory Equipment | 2 Inserts for the Steam Cleaner-Fixed: Wet & Dry (High Performance) |  |  |  | Each
RT286-07-003 D Accessories | Laboratory Equipment | Replacement light bulb for the Steam Cleaner-Fixed: Wet & Dry (High Performance) |  |  |  | Each
RT286-07-003 E Accessories | Laboratory Equipment | Replacement flexible insulated hose from gun to unit for the Steam Cleaner-Fixed: Wet & Dry (High Performance) |  |  |  | Each
RT286-07-003 F Accessories | Laboratory Equipment | Replacement water supply hose for the Steam Cleaner-Fixed: Wet & Dry (High Performance) |  |  |  | Each
RT286-07-003 G Accessories | Laboratory Equipment | Replacement ring clasps for the Steam Cleaner-Fixed: Wet & Dry (High Performance) |  |  |  | Each
RT286-07-004 | Laboratory Equipment | Laboratory Motor Electric (Hanging) SPECIFICATION An electric laboratory motor with flexible cable transmission installation comprising of an electric motor, electronic foot control, flexible cable transmission drive and motor stand The unit shall have the following features:Electric MotorPower: 220 ‑ 240 V, 50 ‑ 60 HzCommutor drive unit with built‑in ventilator for automatic coolingMotor functioning on variable speeds between approximately 5 000-20 000 rpmTorque not less than 18 NcmOutput not less than 180 WattBall bearings maintenance freeRapid braking when switched offRadio interference freeVibration reducing hanging device part of motor assemblyElectronic foot-controlGuaranteed variable speed regulation for ranges of motorSealed housing for protection against dust and electrical shockDetachable from the motor without tools3 Pin 15 Amp plug with at least 2 meter 3 core flex shall be suppliedFlexible cable droveOuter cable consist of a special protective flexible spiral metal tube of approximately 12mm in diameterCable at motor connection detachable without special toolsCable at handpiece connection ball bearing Slip-Joint typeBall bearings at both ends of drive cableInner drive shaft not less than 5 mm in diameterLength not less than 1000 mmStandOver the bench rust protected steel projectionA minimum adjustable vertical length of 1000 mmG-type screw clamp shall be able to fix to bench tops of 17-61 mm in thicknessSABS Plug : Must comply with both SANS 164-0 and SANS 164-1 |  |  |  | Each
RT286-07-004 A Accessories | Laboratory Equipment | Replacement cable for the Laboratory Motor Electric (Hanging) |  |  |  | Each
RT286-07-004 B Accessories | Laboratory Equipment | Replacement hanging motor handpiece catch (end of cable) for the Laboratory Motor Electric (Hanging) |  |  |  | Each
RT286-07-005 | Laboratory Equipment | Handpiece for Laboratory Motor Electric (Hanging) SPECIFICATION A dental laboratory handpiece that is driven by a dental laboratory motor by means of a flexible transmission cable. The unit shall have the following features:Continuous running with drive speed up to 25 000 rpm.Exchangeable chucks to accept tools with shafts 2,35 to 3,175 mm in diameterChuck to accept 2,35 mm diameter tools shall be suppliedTwist chuck tension device for rapid tool insertChuck tension device coveredHandpiece equipped with three sealed bearingsBall bearings self-lubricatingConnection to cable of the slip-joint typeAnatomically shaped for the average size handMetal handpiece housing covered with heat-insulated plasticInsulation colourfast and not react chemically with dental materialsWeight not exceeding 170 g |  |  |  | Each
RT286-07-005 A Accessories | Laboratory Equipment | Replacement Chuck for the Handpiece for Laboratory Motor Electric (Hanging) |  |  |  | Each
RT286-07-006 | Laboratory Equipment | Dental Laboratory Micro Motor (Multi-Purpose) SPECIFICATION A dental laboratory micro-motor consisting of a drive system for intermediate and precision work with a hand guided, brushless, handpiece to be used for long periods of time at speeds between 1 000 & 45 000 Rpm’s in the crown & bridge, cobalt-chrome and light acrylic technique. Tabletop and / or foot control unit, forward/reverse, On/Off and foot pedal for speed control function. The unit shall have the following features:Motor: Brushless - InductionSpeed range: Approximately 1 000 to 45 000 Rpm'sMaximum torque: At least 7.8 NcmConcentricity: Approximately 0.02 mmMotor Cooling: Fan integrated in motor housingHandpiece:Design: Ergonomical shape for fitting in the handSupplied with 2.35 mm chuckTool insert / exchange by means of twist chuck - not leverRapid braking on switching off motorFlexible connect cable with socket from motor to control unitControl Unit:Table Top Unit with ON/OFF, Forward/Reverse, speed control & Foot pedalFoot Unit with ON/OFF, Forward/Reverse & variable speed control foot pedalHandpiece stand supplied with each unitSABS Plug : Must comply with both SANS 164-0 and SANS 164-1 |  |  |  | Each
RT286-07-006 A Accessories | Laboratory Equipment | Replacement Chuck for the Dental Laboratory Micro Motor (Multi-Purpose) |  |  |  | Each
RT286-07-006 B Accessories | Laboratory Equipment | Replacement fuse for the Dental Laboratory Micro Motor (Multi-Purpose) |  |  |  | Each
RT286-07-006 C Accessories | Laboratory Equipment | Replacement foot pedal for the Dental Laboratory Micro Motor (Multi-Purpose) |  |  |  | Each
RT286-07-007 | Laboratory Equipment | Dental Laboratory Micro Motor (Heavy Duty)SPECIFICATION A dental laboratory micro-motor consisting of a drive system for intermediate and precision work with a hand guided, brushless, handpiece to be used for long periods of time at speeds between 3 000 & 65 000 Rpm’s in the heavy acrylic and cobalt-chrome technique. Tabletop and / or foot control unit, forward/reverse, On/Off and foot pedal for speed control function. The unit shall have the following features:Motor: Brushless - InductionSpeed range: Approximately 3 000 to 65 000 Rpm'sMaximum torque: At least 8.7 NcmConcentricity: Approximately 0.02 mmMotor Cooling: Fan integrated in motor housingHandpiece:Design: Ergonomical shape for fitting in the handSupplied with 2.35 mm chuckTool insert / exchange by means of twist chuck - not leverRapid braking on switching off motorFlexible connect cable with socket from motor to control unitControl Unit:Table Top Unit with ON/OFF, Forward/Reverse, speed control & Foot pedalFoot Unit with ON/OFF, Forward/Reverse & variable speed control foot pedalHandpiece stand supplied with each unitSABS Plug : Must comply with both SANS 164-0 and SANS 164-1 |  |  |  | Each
RT286-07-007 A Accessories | Laboratory Equipment | Replacement Chuck for the Dental Laboratory Micro Motor (Heavy Duty) |  |  |  | Each
RT286-07-007 B Accessories | Laboratory Equipment | Replacement foot pedal for the Dental Laboratory Micro Motor (Heavy Duty) |  |  |  | Each
RT286-07-008 | Laboratory Equipment | Air Turbine Non - LED SPECIFICATION A benchtop, high-torque laboratory turbine straight handpiece unit, for precise shaping of hard and very hard materials in the production section of a dental laboratory. The unit shall have the following features:Air consumption: Between 40 & 50 litres / minuteOperating Pressure: Between 2,4 & 3,4 barMaximum speed: 300,000 rpm’sChuck diameter: 1,6 mmWall or under-bench mounting possibleUnit shall consist of:Turbine handpieceON / OFF foot controlControl Unit with filter & pressure gaugeHandpiece rest Rotor cartridge (Plus spare rotor cartridge)Water reservoirLubrication: Manual or automaticallyHandpiece features:Quick chuck mechanismCartridge systemTwo-fold precision ball bearingsExtremely flexible connecting hoseAnatomically shapedSilencerAir escape after turbine is drivenWater cooling and dispensingSABS Plug : Must comply with both SANS 164-0 and SANS 164-1 |  |  |  | Each
RT286-07-008 A Accessories | Laboratory Equipment | Replacement rotor cartridge for the Air Turbine Non - LED |  |  |  | Each
RT286-07-008 B Accessories | Laboratory Equipment | Replacement foot pedal for the Air Turbine Non - LED |  |  |  | Each
RT286-07-008 C Accessories | Laboratory Equipment | Replacement ring clasp for the Air Turbine Non - LED |  |  |  | Each
RT286-07-009 | Laboratory Equipment | Air Turbine LEDSPECIFICATION A benchtop, high-torque laboratory turbine straight handpiece unit, for precise shaping of hard and very hard materials in the production section of a dental laboratory. The unit shall have the following features:Air consumption: Between 40 & 50 litres / minuteOperating Pressure: Between 2,4 & 3,4 barMaximum speed: 300,000 rpm’sChuck diameter: 1,6 mmWall or under-bench mounting possibleUnit shall consist of:Turbine handpieceON / OFF foot controlControl Unit with filter & pressure gaugeHandpiece restRotor cartridge (Plus spare rotor cartridge)Water reservoirLubrication: Manual or automaticallyHandpiece features:Quick chuck mechanismCartridge systemTwo-fold precision ball bearingsExtremely flexible connecting hoseAnatomically shapedSilencerAir escape after turbine is drivenWater cooling and dispensingLED lightSABS Plug : Must comply with both SANS 164-0 and SANS 164-1 |  |  |  | Each
RT286-07-009 A Accessories | Laboratory Equipment | Replacement rotor cartridge for the Air Turbine LED |  |  |  | Each
RT286-07-009 B Accessories | Laboratory Equipment | Replacement LED light for the Air Turbine LED |  |  |  | Each
RT286-07-009 C Accessories | Laboratory Equipment | Replacement foot pedal for the Air Turbine LED |  |  |  | Each
RT286-07-009 D Accessories | Laboratory Equipment | Replacement ring clasp for the Air Turbine LED |  |  |  | Each
RT286-07-010 | Laboratory Equipment | Polishing & Grinding Machine with built-in Suction, Lighting & Eye Protection SPECIFICATION A polishing and grinding machine used in a dental laboratory, utilised for the rough trim, pumice and polish of plastic and metal dentures. The unit shall have the following features:Central drive unit with two shafts for mounting/fitting spindlesAt least two speed settingsIntegrated suction system per working areaFilter bags easily replaceableWaste drawer for solids built into base of unitSeparate polishing and extraction motorSeparate power and lighting switchIntegrated lighting per working areaLED lights providing lightingIntegrated eye protection shield per working areaShield of safety glassConstruction of stainless steel, dust tight and splash proofWorking surface anti-corodal aluminiumRest of unit shall be of lacquered stainless steel and / or high impact plasticSABS Plug : Must comply with both SANS 164-0 and SANS 164-1 |  |  |  | Each
RT286-07-010 A Accessories | Laboratory Equipment | Replacement filter bag for the Polishing & Grinding Machine with built-in Suction, Lighting & Eye Protection |  |  |  | Each
RT286-07-010 B Accessories | Laboratory Equipment | Replacement LED light for the Polishing & Grinding Machine with built-in Suction, Lighting & Eye Protection |  |  |  | Each
RT286-07-010 C Accessories | Laboratory Equipment | Replacement shield for safety glass for the Polishing & Grinding Machine with built-in Suction, Lighting & Eye Protection |  |  |  | Each
RT286-07-011 | Laboratory Equipment | High Speed Grinder with Internal Suction, Lighting & Eye Protection SPECIFICATION A benchtop high speed grinding machine used in a dental laboratory, utilised for finishing prosthetic castings made from non-precious and precious metal alloys. The unit shall have the following features:Drive unit power consumption approximately 250WDrive system to run on vibration free bearingsHigh torque at 24 000 rpm’sNo speed reduction under normal work loadNo switching off of motor when changing burs/grinding stones/toolsQuick-action collect chuck release for fast tool changeExtraction capacity shall be no less than 3 cubic m/minInternal lighting shall be providedWell lit working areaHinged eye-protection cover shall be provided for the working areaShield of safety glassSABS Plug : Must comply with both SANS 164-0 and SANS 164-1 |  |  |  | Each
RT286-07-011 A Accessories | Laboratory Equipment | Replacement chuck for the High Speed Grinder with Internal Suction, Lighting & Eye Protection |  |  |  | Each
RT286-07-011 B Accessories | Laboratory Equipment | Replacement light bulb for the High Speed Grinder with Internal Suction, Lighting & Eye Protection |  |  |  | Each
RT286-07-011 C Accessories | Laboratory Equipment | Replacement shield for safety glass for the High Speed Grinder with Internal Suction, Lighting & Eye Protection |  |  |  | Each
RT286-07-011 D Accessories | Laboratory Equipment | Replacement drive belt for the High Speed Grinder with Internal Suction, Lighting & Eye Protection |  |  |  | Each
RT286-07-012 | Laboratory Equipment | Electrolytic polishing Unit – One Casting SPECIFICATION A benchtop, electronically controlled unit utilised for electrolytic polishing of cobalt-chrome castings in a dental laboratory. The unit shall have the following features:Polishing of single cobalt-chrome castingsSeparate immersion polishing bath to control unitImmersion polishing bath capacity 1 litreImmersion bath sealed with lidFeatures on control panel of housing:ON / OFF power switchCurrent indicated on ammeterLED operating indicatorsAdjustable polishing current by means of a dialTimerConstant selected current valuePolishing current electronically controlledProtected clasp tipsNo striations on castings after polishingControl unit of corrosion resistant steel & plasticSABS Plug : Must comply with both SANS 164-0 and SANS 164-1 |  |  |  | Each
RT286-07-012 A Accessories | Laboratory Equipment | Replacement clamps and hook for the Electrolytic polishing Unit – One Casting |  |  |  | Each
RT286-07-013 | Laboratory Equipment | Electrolytic Polishing Unit - Two or More Castings SPECIFICATION A benchtop, electronically controlled unit utilised for electrolytic polishing of cobalt-chrome castings in a dental laboratory. The unit shall have the following features:Simultaneous polishing of at least two cobalt-chrome castingsIntegrated immersion polishing tub and control unitTwo laterally moving anode supportsAnode supports fitted with knurled screw clampsImmersion polishing tub capacity 2 litreImmersion tub covered by hinged lidLid fitted with observation windowFeatures on control panel of housing:ON / OFF power indicatorLED current indicatorsAdjustable polishing current bmo dialTimerSpecial features:Thermostatically controlled bath heatingAutomatic control of current intensityThree power settingsRing cathode Automatic control of object movementPolish completion in 2 x 4-6 minutesQuick action short-circuit cut-outNo striations on castings after polishingSupplementary cathode for deep palatal sectionsUnit made of corrosion resistant steel & plasticSABS Plug : Must comply with both SANS 164-0 and SANS 164-1 |  |  |  | Each
RT286-07-013 A Accessories | Laboratory Equipment | Replacement clamps and hook for the Electrolytic Polishing Unit - Two or More Castings |  |  |  | Each
RT286-07-014 | Laboratory Equipment | Ultrasonic Cleaning Unit- Laboratory: Large CapacitySPECIFICATION A benchtop, electronically controlled, ultrasonic cleaner designed for all types of cleaning of prosthetic appliances and instruments in a dental laboratory. The unit shall have the following features:Should consist of a hinge cover for easy one hand openingStainless steel tankTank dimensions: approx 19.7 x 14.6 x 10.5 cmTank capacity: approx 2 litreInsert basket for instrumentsAdjustable time cyclesDegas cycle for degassing fresh solutionSelectable run timesCountdown timerSolution tracking - to see how long solution has been in useTime cycle: 0 to 30 minutesTransducer frequency: approx 50 kHzThermostatically adjusted heatingSABS Plug : Must comply with both SANS 164-0 and SANS 164-1 |  |  |  | Each
RT286-07-014 A Accessories | Laboratory Equipment | Replacement insert basket for the Ultrasonic Cleaning Unit- Laboratory: Large Capacity |  |  |  | Each
RT286-07-014 B Accessories | Laboratory Equipment | Replacement ultrasonic solution for the Ultrasonic Cleaning Unit- Laboratory: Large Capacity |  |  |  | Each
RT286-07-015 | Laboratory Equipment | Ultrasonic Cleaning Unit- Laboratory: Small CapacitySPECIFICATION A benchtop, electronically controlled, ultrasonic cleaner designed for cleaning of single dentures (acrylic / metal) as well as crowns and bridges during the production process in a dental laboratory. The unit shall have the following features:Stainless steel casing & lidStainless steel tankTank dimensions: approx L:150, W:85, H:65mmTank capacity: approx 0,5 litreManual or automatic temperature & time controlTime cycle: 0 to 10 minutesTransducer frequency: approx 60 kHzThermostatically adjusted heatingSABS Plug : Must comply with both SANS 164-0 and SANS 164-1 |  |  |  | Each
RT286-07-015 A- Accessories | Laboratory Equipment | Replacement ultrasonic solution for the Ultrasonic Cleaning Unit- Laboratory: Small Capacity |  |  |  | Each
RT286-07-016 | Laboratory Equipment | Porcelain Firing Furnace - Press/CeramicsSPECIFICATION A benchtop dental porcelain firing furnace and vacuum pump, with automated and computerised control for firing of porcelain work with a minimum of 400 definable firing programmes. The unit shall have the following features:Should be a combination furnaceBoth a press and ceramic furnaceFurnace should specially coordinate with IPS e-max press materialshave at least 20 individual press programsShould contain at least 500 individual firing programsShould consist of a bottom reflector that ensures homogeneous heat distribution in the firing chamberShould consist of an integrated Crack Detection SystemShould contain Infrared Technology to recognize size and control tempreture of investment ringsFiring status indicated by means of LED lights and audio warning signalsContain optical status and progress display featureElectronic press drive with force sensor to produce high-quality resultsNo compressed air connection required.Automatic double-range temperature calibrationVacuum pump maximum power consumption of 250 WDimensions of firing chamber at least 90mm diameter and 80 mm highMaximum firing temperature of at least 1 200°CSoftware update via USB flash driveVacuum pump integrated with control unit and firing unit (Double Check)SABS Plug : Must comply with both SANS 164-0 and SANS 164-1 |  |  |  | Each
RT286-07-016 A Accessories | Laboratory Equipment | Replacement heating element for the Porcelain Firing Furnace - Press/Ceramics |  |  |  | Each
RT286-07-016 B Accessories | Laboratory Equipment | Replacement thermocouple for the Porcelain Firing Furnace - Press/Ceramics |  |  |  | Each
RT286-07-016 C Accessories | Laboratory Equipment | Firing tray for the Porcelain Firing Furnace - Press/Ceramics |  |  |  | Each
RT286-07-017 | Laboratory Equipment | Porcelain Firing FurnaceSPECIFICATION A benchtop dental porcelain firing furnace and vacuum pump, with automated and computerised control for firing of porcelain work with a minimum of 400 definable firing programmes. The unit shall have the following features:Firing Furnace system consists of firing unit, control unit and vacuum pumpCompatible with all materials and productsFiring unit:Fully automatic and microprocessor controlledFiring status indicated by means of LED lights and audio warning signalsFiring muffle of silica quartsIntegrated cooling traysSafety and convenience features:Controlled night mode and fast coolingAutomatic cleaning functionAutomatic temperature adjustment before every program startControl Unit:No smaller than 16 cmConnection by means of USBTouch screen with photo viewerChoice of different firing techniquesMinimum of 400 defineable firing programmesAble to monitor temperature and vacuumVacuum pump integrated with control unit and firing unitSABS Plug : Must comply with both SANS 164-0 and SANS 164-1 |  |  |  | Each
RT286-07-017 A Accessories | Laboratory Equipment | Replacement heating element for the Porcelain Firing Furnace |  |  |  | Each
RT286-07-017 B Accessories | Laboratory Equipment | Replacement thermocouple for the Porcelain Firing Furnace |  |  |  | Each
RT286-07-017 C Accessories | Laboratory Equipment | Replacement USB cable for the Porcelain Firing Furnace |  |  |  | Each
RT286-07-017 D Accessories | Laboratory Equipment | Firing tray for the Porcelain Firing Furnace |  |  |  | Each
RT286-07-018 | Laboratory Equipment | InLab CAD-CAM SPECIFICATION A benchtop Multifunctional CAD/CAM system for Dental Laboratories consisting of a Multi Axis Scanner, Milling Unit, Sintering Furnace and Operating System. The unit shall have the following features:Multi Axis Scanner:Table Top model with coaxial scanner and milling tools.Functional and user-friendly design.Manual and automatic functionQuick image processing at a rate of at least 1mil pixels in 0.140 seconds and capture entire Jaw in less than 23 seconds.Be able to position the most popular articulators. Be able to scan models (all popular plasters) and impressions.The export of scanning data must be in open STL format for processing on third party CAD/CAM systems.USB interphaseMilling Components: Self-calibrating and auto compensation for the wear of the grinding tools ensures utmost precision.Milling steps at least 7um during grinding of restoration.Grain roughness of the finished milled restorations must be at least @ 50 um.The unit must have at least a self contained water and coolant chamber.The unit must be able to use at least the following materials:a. Aluminium Oxideb. Feldspar/glass ceramicsd. Lithium disilicatee. Models from polyurethane polymers.f. Non Precious metals (pre sintered Cobalt Chrome)g. Resin NANO Ceramics h. Silicate ceramicsi. Ti Base sets for implantsj. Translucent zirconium oxide for full contour restorationsk. zirconium oxideVarious burs must be available for different materialsThe bidder must include all software components needed for Framework, VinCrOn and Wax UpThe unit must at least have the following wide-ranging of clinical indications:a. Abutments: (Custom hybrid titanium base for most popular implant systems)b. Attachments: for attachment pinsc. Bars: choice of several geometries, Design tools for customizationd. Crown Coping and bridge frameworks (METAL). Partially or fully reduced frameworkse. Crown and Bridge frameworks – metalf. Full Crownsg. Inlaysh. Veneers. (min thickness 300µmi. Onlaysj. Scanbodies for transferring exact position of implant to design softwarek. Stress breaker attachmentsl. Surgical guidem.TelescopesMust have 2 rotational and 3 translation axes Accommodate complex geometries such as long span bridges and bars, models, surgical guidesBe able to hold up to 6 tools per magazineHighly precise detection of the exact position of disks and blocks for processing of blocks with dentine structures and mesostructuresChamber. At least have a dia of 95mm and height 30mm and able to be used for Zirconium, PMMA, composite and waxIndications of models, splints, surgical guides, prosthesis etcOpen for disk supplies of all kindsMulti – block management for blocks up to 40 X 19 X 19 mm and be able to at least used for grinding glass ceramics, hybrid ceramic & lithium disilicate ceramicsExternal water tank for coolant be supplied External suction be supplied with automatic communication with the machine Sintering FurnaceThe machine must be programmable with the most common brands of zirconia materialMust have at least four pre-set programme positions for sintering with pre-drying and ventilationMust have an integrated gas management moduleAt least be able to sinter up to 50 units at onceOperating SystemMust be supplied with all hard- and software to run the systemThe software must support all working steps namely scanning, milling of restorations and models, abutments, bars, surgical guides for implant placementOpen system for flexible integration. STL data of all dental CAD softwareIntergratable with other systems including Cone beam machinesSABS Plug : Must comply with both SANS 164-0 and SANS 164-1 |  |  |  | Each
RT286-07-018 A Accessories | Laboratory Equipment | Replacement bur chuck for th e - CAD-CAM |  |  |  | Each
RT286-07-018 B Accessories | Laboratory Equipment | Replacement burs for the CAD-CAM |  |  |  | Each
RT286-07-018 C Accessories | Laboratory Equipment | Replacement water filter for the - CAD-CAM |  |  |  | Each
RT286-07-018 D Accessories | Laboratory Equipment | Replacement air filter for the I CAD-CAM |  |  |  | Each
RT286-07-018 E Accessories | Laboratory Equipment | Milling fluid for the - CAD-CAM |  |  |  | Each
RT286-08-001 | Laboratory Equipment | Heavy Duty Dental Spot Welder SPECIFICATION A benchtop, heavy-duty spot welding and soldering unit with electronic control for precise dose performance within the dental industry. The unit shall have the following features:Finely adjustable electrode distance to prevent distortion during spot weldingFinely adjustable current to enable addition of fine wires to heavy castingsSpot welding control via optical & acoustic signalPossible to spot weld and solder without changing electrodesFoot control for soldering as well as opening the fixed electrodesSpot welding features:Possible for wires, clasps, solderChoice between hand-held or fixed electrodesWelding time control for overheating protectionSoldering features:Freehand soldering with carbon electrodesModel soldering with copper electrodesSABS Plug : Must comply with both SANS 164-0 and SANS 164-1 |  |  |  | Each
RT286-08-001 A Accessories | Laboratory Equipment | Replacement electrodes for the Heavy Duty Dental Spot Welder |  |  |  | Each
RT286-08-002 | Laboratory Equipment | Inert Gas Welder SPECIFICATION A benchtop, heavy-duty inert gas welding unit and welding microscope with electronic control for applying welds to all common dental alloys and titanium, for work associated with the repairing or manufacturing of dental appliances for patients. The unit shall have the following features:Consists of control unit, welding handpiece with flexible pipe & welding microscopeControl unit:Main voltage: 230 V ACPower input approximately 400 VASpeed control for weldingMemory slots for for personalised welding parametersAdjustable Impulse timeAdjustable welding powerConnect to Medical grade Argon gas supply for inert gas weldingWelding handpiece:Tungsten electrodeUtilise shielding (inert) gasOxidation free metalLow heat generationMinimum thickness of sheet welding: 0.10 mmMinimum thickness of wire welding: 0.25 mmReplaceable tungsten welding electrodeWelds per electrode: at least 1800Stereo microscope:Mounted on adjustable standStand with supporting arm for welding hadpieceAt least 10x magnificationEye protection systemAutomatic shutter system for eye protectionSABS Plug : Must comply with both SANS 164-0 and SANS 164-1 |  |  |  | Each
RT286-08-002 A Accessories | Laboratory Equipment | Replacement electrodes for the Inert Gas Welder |  |  |  | Each
RT286-08-002 B Accessories | Laboratory Equipment | Replacement oxidation free metal for the Inert Gas Welder |  |  |  | Each
RT286-08-002 C Accessories | Laboratory Equipment | Replacement Argon gas regulator for the Inert Gas Welder |  |  |  | Each
RT286-08-002 D Accessories | Laboratory Equipment | Replacement Argon gas flexible pipe for the Inert Gas Welder |  |  |  | Each
RT286-08-002 E Accessories | Laboratory Equipment | Replacement ring clasps for the Inert Gas Welder |  |  |  | Each
RT286-08-003 | Laboratory Equipment | Laser WelderSPECIFICATION A high performance, inert gas, free standing laser-welding unit with integrated stereo-microscope. The unit shall have the following features:Microscope magnification: No less than 15-foldSighting device consisting of reticle in microscope, plus 2 crossed sighting light diodesParameter display at frontAdjustable manipulation openingsLighting in working chamber: Minimum 2 adjustable cold light reflectorsIntegrated cooling systemIntegrated argon inert gas systemInert gas jet: Minimum 1Integrated fume-gas extraction systemPulse energy: 50 joulesPulse length: 0.5 – 20 msMean Power: 20 WPeak power minimum: 4 kWSABS Plug : Must comply with both SANS 164-0 and SANS 164-1 |  |  |  | Each
RT286-08-003 A Accessories | Laboratory Equipment | Replacement electrodes for the Laser Welder |  |  |  | Each
RT286-08-003 B Accessories | Laboratory Equipment | Replacement Argon gas regulator for the Laser Welder |  |  |  | Each
RT286-08-003 C Accessories | Laboratory Equipment | Replacement Argon gas flexible pipe for the Laser Welder |  |  |  | Each
RT286-08-003 D Accessories | Laboratory Equipment | Replacement ring clasps for the Laser Welder |  |  |  | Each
RT286-09-001 | Laboratory Equipment | Chair Dental Laboratory upholstered SPECIFICATION An anatomically correct, fully adjustable, mobile upholstered chair utilised by dental technicians in the production section of a dental laboratory. The unit shall have the following features:Frame: Robust metal constructionMetal finish: Chrome/ anodised / powder coatedBase consist of five horizontal legsOne 360° double wheel castor per legSeat:Anatomically contouredPrevent pressure on thighsPrevent restriction of blood circulationAdjustable seat height: 500 - 600mmUpholstery breathing & perforatedBackrest:Anatomically contouredAdjust horizontally & vertically with body positionPrevent tiring of the lower backMinimum height adjustment: 200mmMinimum rake adjustment: 100mmUpholstery breathing & perforatedFoam minimum two layers with:Approx 100mm heavy density (VP45/ Chip 80 Foam) andWith 10mm light density |  |  |  | Each
RT286-09-002 | Laboratory Equipment | Chair Dental Laboratory with silicone SPECIFICATION An anatomically correct, fully adjustable, mobile silicone upholstered chair utilised by dental technicians in the production section of a dental laboratory. The unit shall have the following features:Frame: Robust metal constructionMetal finish: Chrome/ anodised / powder coatedBase consist of five horizontal legsOne 360° double wheel castor per legSeat:Anatomically contouredPrevent pressure on thighsPrevent restriction of blood circulationAdjustable seat height: 500 - 600mmBackrest:Anatomically contouredAdjust horizontally & vertically with body positionPrevent tiring of the lower backMinimum height adjustment: 200mmMinimum rake adjustment: 100mmFoam minimum two layers with:Approx 100mm heavy density (VP45/ Chip 80 Foam) andWith 10mm light density |  |  |  | Each
RT286-09-003 | Laboratory Equipment | Magnifying Light – Circular SPECIFICATION A combination light and magnifier for utilisation in the production section of a dental laboratory. The unit shall have the following features:Circular in designNo sharp edges to cause accidental injuryHousing of lamp impact resistant plasticON / OFF switch on top of lamp housingMagnifying lens of cut glass - 2 dioptreCircular lamp positioned around magnifying lensLED lightingLamp fitted into base of lamp housingHousing to prevent light shining into operator eyesMounting / base of lamp by means of:Bench Clamp, orFixing bracket for wall and / or bench, orWeighted baseLight head three dimensionally adjustableAdjustment by means of a double articulating arm pivoting in baseFeatures of lamp:Create natural lightConsistent light intensity by means of electronic controlIndependent of ambient lighting, daylight, or other colour influencing factorsElectric supply to light via base and arm to headSABS Plug : Must comply with both SANS 164-0 and SANS 164-1 |  |  |  | Each
RT286-09-003 A Accessories | Laboratory Equipment | Replacement circular LED light for the Magnifying Light – Circular |  |  |  | Each
RT286-09-003 B Accessories | Laboratory Equipment | Replacement magnifying lens for the Magnifying Light – Circular |  |  |  | Each
RT286-09-004 | Laboratory Equipment | Plaster Trap - Under Basin SPECIFICATION A plaster trap with separating compartments to collect solids and suspended sediments from water from the overhead drain basin, before draining into the regular sewerage system. The unit shall have the following features:Manufactured of corrosion resistant stainless steel or ABS plastic materialMounted underneath all commercially available basins and in cabinetsAt least 3 seperating compartmentsTight fitting cover / lid over plaster trap to eliminate odoursPlumbing connection to sewerage system according to building regulationsDesign to facilitate ease of cleaning |  |  |  | Each
RT286-10-001 | Radiological equipment | Intra-Oral X-Ray Processor-Automatic SPECIFICATION An automatic intra-oral x-ray processeor for all sizes of intra-oral x-rays analogue films. The unit must have the following features:Process intra-oral x-ray films up to 65 mmProcessing time approximately 2 minutes for wet endodonticProcessing time approximately 4 minutes for dryEquipped with integrated drierEquipped with daylight loaderSABS Plug : Must comply with both SANS 164-0 and SANS 164-1 |  |  |  | Each
RT286-10-002 | Radiological equipment | Intra-Oral X-Ray Processor-Manual SPECIFICATION A manual intra-oral x-ray processeor for peri-apical and bitewing dental film, with internal dipping containers for developing, fixing and cleaning (water) liquid. The unit must have the following features:Process intra-oral x-ray peri-apical and bitewing dental filmCasing manufactured from corosion resistant materialSupplied with 3 internal dipping containers for developing, fixing & cleaning liquidEquipped with viewing screen comrising lead acrylic material minimum 0.4 mm thickness and 0.5mmPbEquipped with sleeved hand access ports |  |  |  | Each
RT286-10-002 A Accessories | Radiological equipment | Replacement sleeves for Intra-Oral X-Ray Processor-Manual |  |  |  | Each
RT286-10-003 | Infection control | Heat Sealer SPECIFICATION A desktop unit utilised for heat sealing instruments or tools in sterilisation bags (plastic / paper combination prior to sterilisation in dental autoclaves). The unit shall have the following features:Include a built-in cutterMust have an automatic temperature controlMust have an automatic Start / Stop control |  |  |  | Each
RT286-10-003 A Accessories | Infection control | Replacement element for the Heat Sealer |  |  |  | Each
RT286-10-003 B Accessories | Infection control | Cover for the Heat Sealer |  |  |  | Each
RT286-10-003 C Accessories | Infection control | Insulator for the Heat Sealer |  |  |  | Each
RT286-10-003 D Accessories | Infection control | Cutter Balade for the Heat Sealer |  |  |  | Each
RT286-10-004 | Laser | Diode Laser Dental - Portable SPECIFICATION A portable dental diode laser for soft tissue surgery and bleaching in the dental surgery. The unit must have the following features:Application to include soft tissue surgery and bleachingOperating Voltage to 240 voltsWavelength between 445 to 980nmDisposable tips available in a variety of shapes sizes and anglesTouch Screen MenuAutoclavable handpieceHandpiece must accept: a) Standard canulas b) Multi tips c) Stripable FibreFoot operationSABS Plug : Must comply with both SANS 164-0 and SANS 164-1 |  |  |  | Each
RT286-10-004 A Accessories | Laser | Compatible Disposable tips available in a variety of shapes sizes and angles |  |  |  | Each
RT286-10-004 B Accessories | Laser | Standard canulas |  |  |  | Each
RT286-10-004 C Accessories | Laser | Stripable Fibre |  |  |  | Each
RT286-10-005 | Endodontic Equipment | Endodontic Obturator Oven SPECIFICATION A specifically designed heat source for heating endodontic obturators. The unit must have the following features:Give fast and predictable resultsHeats obturators in at least 17 secondsIntegrated timer for heating ovenSignal on timer when oven is at optimal temperatureA specifically designed heat source for heating endodontic obturators. The unit must have the following features:Give fast and predictable resultsHeats obturators in at least 17 secondsIntegrated timer for heating ovenSignal on timer when oven is at optimal temperature |  |  |  | Each
RT286-10-006 | Infection control | Water Demineraliser Unit SPECIFICATION A desktop water demineraliser unit for dental autoclaves. The unit MUST have the following features:Full installation of the Demineraliser unit on delivery No solenoid valvesunpressurized system for tank fillingMonitored by conductivity motorSABS Plug : Must comply with both SANS 164-0 and SANS 164-1 |  |  |  | Each
RT286-10-006 A Accessories | Infection control | Cartriges for the Water Demineraliser Unit |  |  |  | Each
RT286-10-007 | Infection control | Ultrasonic Cleaning Unit SPECIFICATION An ultrasonic bath utilised for the cleaning of dental instruments prior to sterilisation. The unit shall have the following features:Minimum 4.5 Litre tank (working capacity)Supplied with instrument basket, bur tray ( stainless steel material) and beaker (glass material)Equipped with an integrated timerIntegrated hinged lidSABS Plug : Must comply with both SANS 164-0 and SANS 164-1 |  |  |  | Each
RT286-10-007 A Accessories | Infection control | Basket ( stainless steel material) for the Ultrasonic Cleaning Unit |  |  |  | Each
RT286-10-007B Accessories | Infection control | Beaker( (glass material) for the Ultrasonic Cleaning Unit |  |  |  | Each
RT286-10-007 C Accessories | Infection control | Bur tray ( stainless steel material) for the Ultrasonic Cleaning Unit |  |  |  | Each
RT286-10-008 | Dental surgery small equipment | Amalgam Separator Multiple Surgeries SPECIFICATION A fixed installation amalgam separator for up multiple surgeries. The unit MUST have the following features:Fixed installationEquipped for connection of multiple treatment unitsServe number of surgeries:2 - 4Output power: Single-phase, Approximately 210WNoise level: No more than 60dB(A)Volume of collector vessel: Approximately 600 cm3Separation rate: At least 97%Flow rate at least 15 l/minAutomatic filling level checkAutomatic switch on when waste water is available from treatment unitCentrifugal separator with hydrodynamic self-cleaning function |  |  |  | Each
RT286-10-008 A Accessories | Dental surgery small equipment | Collection Container with recycle kit for the Amalgam Separator Multiple Surgeries |  |  |  | Each
RT286-10-009 | Dental surgery small equipment | Amalgam Separator 2 Surgeries SPECIFICATION A fixed installation amalgam separator for up to 2 surgeries. The unit MUST have the following features:Fixed installationEquipped for connection of several treatment unitsServe number of surgeries: Maximum 2Output power: Single-phase, Approximately 115WNoise level: No more than 60dB(A)Volume of collector vessel: Approximately 90 cm3Separation rate: At least 95%Flow rate at least 4 l/minAutomatic filling level checkAutomatic switch on when waste water is available from treatment unitCentrifugal separator with hydrodynamic self-cleaning function |  |  |  | Each
RT286-10-009 Aaccessories | Dental surgery small equipment | Collection Container with recycle kit for the Amalgam Separator 2 Surgeries |  |  |  | Each
RT286-10-010 | Dental surgery equipment | Suction Dry 2-3 Surgeries SPECIFICATION A fixed installation centralised dry aspiration plant for serving 3 surgeries. The unit MUST have the following features:Fixed installationServe number of surgeries: 2 - 3Output power: Single-phase, Approximately 1.8kW - 5A, 50 HzSpeed of motor: No less than 2800 rpmVacuum setting at approximately 170 mBarNoise level without housing: No more than 65dB(A) |  |  |  | Each
RT286-10-011 | Radiological equipment | Pan X-Ray Viewer SPECIFICATION An X-Ray viewer utilised for viewing panoramic and smaller size X-Ray photos. The unit shall have the following features:Utilised as a Countertop or Wall-mounted itemViewing area no smaller than 150 x 140mmEven illumination to be suppliedViewing panel of shatterproof plexiglassRustproofEasy-to-clean surfaces for infection controlSABS Plug : Must comply with both SANS 164-0 and SANS 164-1 |  |  |  | Each
RT286-10-012 | Dental surgery equipment | Oil Free air compressor for 2-4 surgery (60DB) SPECIFICATION A fixed installation dental compressor with dryer (oil-free) for serving 2-4 surgeries. The unit shall have the following features:Voltage 220/240 V 1,5 kW with built in thermo overload protection as per user requirementDelivery capacity not less than 150 1/min maximum pressure 1000 kpaStorage Capacity at least 35 litres - 75 litresPower 0.75Kw 1HPFitted with ¼” safety release valveTank pressure gaugePressure range control switch, 6 – 10 barNoise level maximum 60 DB (A)Either piston type or turbo typeReverse pressure exhaust valve for motor protectionAir filter for filtration of 10 micron particles, with automatic drain valveOil free with automatic drainage system |  |  |  | Each
RT286-10-021 | Dental surgery small equipment | Ultrasonic Scaler Multi-Function SPECIFICATION A multi-function ultrasonic (piezon) scaler that can also be used specifically for endodontic and post removal. The unit shall have the following features:Must be supplied with:Control unitHandpiece (Fibre Optic)Handpiece CordWater supply tube with water filterFoot PedalTip Replacement wrench if applicableWater filterUnit shall include the following kits or instruments:Kit for removing supra gingival calculus that include tips Work on large areas on all tooth surfacesTips for interproximal scaling Flattened round tip for large areas of calculusTip for molar interproximal surfacesEndodontic removal instrument kit with tips for the removal of fractured endodontic instruments as well as the opening of sclerotic canals:Angled small spherical diamond coated tipsAngled instrument tip with working length up to 10mm Angled instrument tip with working length up to 20mm Angled instrument with working length up to 24mmInstrument tip for crown, bridges and metal post removalsSABS Plug : Must comply with both SANS 164-0 and SANS 164-1 |  |  |  | Each
RT286-10-021 A Accessories | Dental surgery small equipment | Angled small spherical diamond coated tips |  |  |  | Each
RT286-10-021 B Accessories | Dental surgery small equipment | Angled instrument tip with working length up to 10mm |  |  |  | Each
RT286-10-021 C Accessories | Dental surgery small equipment | Angled instrument tip with working length up to 20mm |  |  |  | Each
RT286-10-021 D Accessories | Dental surgery small equipment | Kit for removing supra gingival calculus that include tips |  |  |  | Each
RT286-10-021 E Accessories | Dental surgery small equipment | Work on large areas on all tooth surfaces |  |  |  | Each
RT286-10-021 F Accessories | Dental surgery small equipment | Tips for interproximal scaling |  |  |  | Each
RT286-10-021 G Accessories | Dental surgery small equipment | Flattened round tip for large areas of calculus |  |  |  | Each
RT286-10-022 | Dental surgery small equipment | Ultrasonic Piezo Scaler SPECIFICATION A multi-function ultrasonic (piezon) scaler. The unit shall have the following features:Hand piece Control unitHandpiece CordWater supply tube with water filterFoot PedalTip Replacement wrench Water filterUnit shall include the following kits or instruments:Kit for removing supra gingival calculus that include tips forWork on large areas on all tooth surfacesTips for interproximal scalingFlattened round tip for large areas of calculusTip for molar interproximal surfacesSABS Plug : Must comply with both SANS 164-0 and SANS 164-1 |  |  |  | Each
RT286-10-023 | Dental surgery small equipment | Ultrasonic Scaler Multi-Function SPECIFICATION A multi-function ultrasonic (piezon) scaler. The unit shall have the following features:Must be supplied with:Control unitHandpiece CordWater supply tube with water filterFoot PedalTip Replacement wrench Water filterUnit shall include the following kits or instruments:Kit for removing supra gingival calculus that include tips forWork on large areas on all tooth surfacesTips for interproximal scalingFlattened round tip for large areas of calculusTip for molar interproximal surfaces |  |  |  | Each
RT286-10-024 | Dental surgery small equipment | Dental Ultrasonic Scaler with a polishing function (Insert 30K) SPECIFICATION A multi-function ultrasonic scaler wit a polishing capability. The unit shall have the following features:Portable Ultra compact bodySimple and intuitive operation control panelWater AdjustmentInserts washabble and autoclavable up to 135ºC Control unit with hand piece cord and holderWater filter setWater Supply connectorFoot controlAt least 3 scaling inserts 30KMust be supplied with a quick connector (Male & Female)Must be able to perform air polishingSABS Plug : Must comply with both SANS 164-0 and SANS 164-1 |  |  |  | Each
RT286-10-024 A Accessories | Dental surgery small equipment | Insert 30K |  |  |  | Each
RT286-10-024 B Accessories | Dental surgery small equipment | Polishing Insert |  |  |  | Each
RT286-10-024 C Accessories | Dental surgery small equipment | Prophy polishing powder |  |  |  | Each
RT286-10-025 | Dental surgery small equipment | Dental Ultrasonic Scaler (Inserts 25K) SPECIFICATION A multi-function ultrasonic scaler. The unit shall have the following features:Portable Ultra compact bodySimple and intuitive operation control panelWater adjustment switchInserts washable and autoclavable up to 135ºCControl unit wih handpiece cord and HolderWater filter setWater supply connectorFoot controlAt least 3 scaling insert (25K)Must be supplied with a quick connectorSABS Plug : Must comply with both SANS 164-0 and SANS 164-1 |  |  |  | Each
RT286-10-025 A Accessories | Dental surgery small equipment | Scaling insert (25K) |  |  |  | Each
RT286-10-026 | Dental surgery small equipment | Dental Air Scaler SPECIFICATION A handpiece air scaler to fit on a dental unit:Oscillation frequency of 5.500 to 6.500 HzHandpiece washable and autoclavable up to 135ºCScaling tip kitTip wrench with Torque limiterCompitable wih 4/6 hole tubbing |  |  |  | Each
RT286-10-026 A Accessories | Dental surgery small equipment | Scaling tip kit
RT286-10-027 | Infection control | Desktop Autoclave S-Type for Dental Instruments SPECIFICATION A S-type, casette loading desktop autoclave for sterilising wrapped dental instruments. The unit shall have the following features:Must comply with ISO Standard (PREN 13060) for S-Type autoclavesInclude baskets and accessories for sterilising wrapped instrumentsMust be a casette type loading autoclaveCasette mesurements no less than 28cm x 18cm x 3.5cmFull sterilisation cycle from 6-45 minutesSterilization chamber volume of minimum 1.8 litres Temperature Range :121 - 135⁰ CSABS Plug : Must comply with both SANS 164-0 and SANS 164-1The unit MUST measure chamber pressure with accuracy of ±3% of full scale.AlarmsWater supply failureCycle failureSafety FeaturesShut down when opened.Minimum capacity of 1.6 litre |  |  |  | Each
RT286-10-028 | Infection control | Desktop Autoclave S-Type for Dental Instruments SPECIFICATION A S-type, casette loading desktop autoclave for sterilising wrapped dental instruments. The unit shall have the following features:Must comply with ISO Standard (PREN 13060) for S-Type autoclavesInclude baskets and accessories for sterilising wrapped instrumentsMust be a casette type loading autoclaveCasette mesurements Large 38cmx18cmx7.5cmFull sterilisation cycle from 6-45 minutesSterilization chamber volume of minimum 1.8 litres Temperature Range :121 - 135⁰ CSABS Plug : Must comply with both SANS 164-0 and SANS 164-1The unit MUST measure chamber pressure with accuracy of ±3% of full scale.AlarmsWater supply failureCycle failureSafety FeaturesShut down when opened.Minimum capacity of 1.6 litre |  |  |  | Each
RT286-10-029 | Infection control | Desktop Autoclave S-Type for Dental Instruments SPECIFICATION A S-type, casette loading desktop autoclave for sterilising wrapped dental instruments. The unit shall have the following features:Must comply with ISO Standard (PREN 13060) for S-Type autoclavesInclude baskets and accessories for sterilising wrapped instrumentsMust be a casette type loading autoclaveCasette mesurements no less than 28cm x 18cm x 3.5cmFull sterilisation cycle from 6-45 minutesSterilization chamber volume of minimum 6 litres Temperature Range :121 - 135⁰ CSABS Plug : Must comply with both SANS 164-0 and SANS 164-1The unit MUST measure chamber pressure with accuracy of ±3% of full scale.AlarmsWater supply failureCycle failureSafety FeaturesShut down when opened.Minimum capacity of 1.6 litre |  |  |  | Each
RT286-10-030 | Infection control | Desktop Autoclave S-Type for Dental Instruments WITH water distiller SPECIFICATION A S-type, casette loading desktop autoclave for sterilising wrapped dental instruments. The unit shall have the following features:Must comply with ISO Standard (PREN 13060) for S-Type autoclavesInclude baskets and accessories for sterilising wrapped instrumentsMust be a casette type loading autoclaveCasette mesurements no less than 28cm x 18cm x 3.5cmFull sterilisation cycle from 6-45 minutesSterilization chamber volume of minimum 1.8 litres Temperature Range :121 - 135⁰ CSABS Plug : Must comply with both SANS 164-0 and SANS 164-1The unit MUST measure chamber pressure with accuracy of ±3% of full scale.AlarmsWater supply failureCycle failureSafety FeaturesShut down when opened.Minimum capacity of 1.6 litre |  |  |  | Each
RT286-10-031 | Infection control | Desktop Autoclave S-Type for Dental Instruments WITH water distiller SPECIFICATION A S-type, casette loading desktop autoclave for sterilising wrapped dental instruments. The unit shall have the following features:Must comply with ISO Standard (PREN 13060) for S-Type autoclavesInclude baskets and accessories for sterilising wrapped instrumentsMust be a casette type loading autoclaveCasette mesurements Large 38cmx18cmx7.5cmFull sterilisation cycle from 6-45 minutesSterilization chamber volume of minimum 1.8 litres Temperature Range :121 - 135⁰ CSABS Plug : Must comply with both SANS 164-0 and SANS 164-1The unit MUST measure chamber pressure with accuracy of ±3% of full scale.AlarmsWater supply failureCycle failureSafety FeaturesShut down when opened.Minimum capacity of 1.6 litre |  |  |  | Each
RT286-10-032 | dental equipment | Mobile Dental Unit non fibre optic without hand pieces SPECIFICATION A Complete mobile dental delivery system, non fiberoptic function without hand piecesDental cabinet houlsing the following:3 Handpiece outlets ( 4 - 6 Holes)3 in1 syringeSelector switchFoot controlMaster shut off assemblyCompressor (25L - 50L; ≥100L/Min)Cleanwater systemThe unit must operate on 240VAC 5060Hz fitted with SABS Plug : Must comply with both SANS 164-0 and SANS 164-1Unit to be fitted with smooth running four swivel lockable wheelsHeight = 700MM - 900MM Width = 400MM - 500MM Depth = 480MM - 600MM |  |  |  | Each
RT286-10-033 | dental equipment | Mobile Dental Unit non fibre optic with hand pieces SPECIFICATION A Complete mobile dental delivery sytem, non fiberoptic function witht hand piecesDental cabinet housing the following : 3 Handpiece outlets ( 4 - 6 Holes)3 in 1 SyringeSelector switchFoot controlMaster shut off assemblyCompressor (25L - 50L; ≥100L/Min)Cleanwater systemSelf illumination LED Handpiece with coupling TurbineInternal spray straight handpieceContra angle blue handpieceProphylaxis handpieceAir motor 4-holeThe unit must operate on 240VAC 50/60Hz fitted with 3-prongs plugUnit to be fitted with smooth running four swivel lockable wheelsHeight = 700MM - 900MM Width = 400MM - 500MM Depth = 480MM - 600MM |  |  |  | Each
RT286-10-034 | dental equipment | Mobile Dental Unit fibre optic capability function and built in suction with low and high volume tubing and hand pieces SPECIFICATION A Complete mobile dental delivery sytem with fiberoptic capability function complete with hand piecesDental cabinet housing the following:3 Handpiece outlets (4-6 holes)3 in 1 SyringeSelector switchFoot controlMaster shut off assemblyCompressor (25L - 50L; ≥100L/Min)Clean water reservoirSelf illumination handpiece with couplingInternal spray straight handpiece with fibre opticsContra angle blue handpiece with fibre opticsProphylaxis HandpieceAir motor 4-holeSuction with low and high volume tubing The unit must operate on 240VAC 50/60Hz fitted with 3-prongs plugThe unit to be fitted with smooth running four swivel lockable wheelsHeight = 700MM - 900MM Width = 400MM - 500MM Depth = 480MM - 600MM |  |  |  | Each
RT286-10-035 | Infection control | Dental Autoclave class S SPECIFICATION A Table top s class autoclave with a short cycle availble not exceeding 20 min for quick sterIlizingTable top steam cassette autoclave.It must include a water distillerIt MUST have programmable cycles : Wrapped & Unwrapped automatic.Temperature range : 100- 134⁰ with a heat shutoff function.The controls MUST be sealed against fluid penetration. Must Include Baskets and accessories for sterilising wrapped instrumentsThe unit MUST have chamber temperature monitoring, accurate to ± 1⁰C over the operating range.Must be able to complete a full cycle in less than 18 minutesThe unit MUST measure chamber pressure with accuracy of ±3% of full scale.AlarmsWater supply failureCycle failureSafety FeaturesShut down when opened.Minimum capacity of 1.6 litreSABS Plug : Must comply with both SANS 164-0 and SANS 164-1 |  |  |  | Each
RT286-10-036 | dental equipment | Portable dental unit SPECIFICATION The unit is a fully contained mobile dental apparatus for point of use applications for oral health treatment in community clinics and hospitals. With its compressor, gas tank pressure release filter valve, clean water bottle, saliva gathering bottle, on/off foot switch, it provides a reliable steady dry and uncontaminated air source available for high- speed turbine hand piece. The unit is portable, simply operated, easy to maintain and use widely by dentist, Oral Hygienists, Dental Therapists and doctors.Easily transported with pull-out handle and wheels.Weighs 15 to 30kg depending on configuration.220V compatible, oil less compressor (580W)7 L tank capacityStandard Features 3- way Syringe with hose Saliva ejector hose x1Oil free compressor x1Portable unit must include the following extras:1 Piezo ultrasonic scaler.2 LED curing light.Fast Hand Piece Hi speed hand piece 32 PB Borden /Midwest type connection.Triple spray distribution symmetrical in the direction of the bur tip.Rotation up to 420000 rpmRotors accurately balanced.Working air pressure between 32 and 35 means 3-bars to 3.5 Slow Hand Piece/ Contra-Angle Single External Spray: A silicone hose external to the handpiece is used for passage of the water from the body of the micro motor to the end of the handpiece used.Rotation:Maximum of 20,000 RPMTransmission 1:1.Ideal Weight:Nett weight under 71.0 gramsLow work noiseAir MotorStraight hand piece |  |  |  | Each
RT286-10-037 | Infection control | Autoclave for handpieces SPECIFICATION An Autoclave for different dental handpieces with an automatic oiling featureThe apparatus must be able to clean in all internal channels of the handpieces, contra angles and turbines internally as well as externally.Must be able to lubricate (oil) the handpieces after internal cleaning but before external hot cleaning and before the sterilization process started.Must have a drying process.Must have choice of sterilization programmes that can cater for quick processing (3 minutes) to lower heat at 121 degrees Celsius up to 20 minutes.Must be a low noise autoclaveMust at least conform to the European or American standard.A build-in water conductivity sensor that measures the water quality that activates if the quality deteriorates.The unit must include water distiller/demineraliser.Alarms: Water supply failureCycle failureSafety FeaturesShut down when opened.SABS Plug : Must comply with both SANS 164-0 and SANS 164-1 |  |  |  | Each
RT286-10-038 | Radiological Equipment | Digital Intraoral video camera system An inra-oral video camara for diagnostic and education purposes SPECIFICATION A direct intraoral Video camera shall consist of the following components: and/or the following features:The software must be compatible with existing software of the user/requestor.Must be equipped with a high resolution and at least a 3 to 4 point LED lightLED light must be integrated in the handpieceAt least a 80-900 single optic system (Intra and extra Oral)The camera must have: (a) Intraoral and extraoral views(b) Image freezing and saving(c) Direct and mirror vision(d) Be used on a computer or video monitor. |  |  |  | Each
RT286-10-039 | Radiological Equipment | Digital Intraoral imaging Photostimulable phosphor plate system with computer SPECIFICATION A X-ray sytem utilising plates for quick digilising of intra oral images. The unit MUST have the following features:Must be compatible with AC and DC intra oral X-ray Units.Software including implant module and implant archive. .Fast Ethernet connection. Must include all cabling and installation from the unit to the surgeries.4 X imaging plate sets per unit (1 for every surgery) MUST be supplied and includes (per surgery): a. 2 x 22x33mm plates. b. 2 X 24x40mm plates. c. 5 X 31x41 mm Plates d. 1 X 27x45mm plate e. 1 X storage box for plates. f. 100 pieces of hygiene bags per sizeThe unit must be provided with an up to date desktop/laptop computer with the following requirements:≥4GB Ram≥1TB Hard drive17 inch LED screen, minimum 1.3mp screen for reviewing images Software for processing all captured images, must be DICOM3 compatible, including DICOM worklist,DICOM PACS, DICOM Store, DICOM transfer, DISOM query/retrieve, DICOM print, DICOM MPPS and DICOM RDSRA keyboard and mouse (cordless/cord)The line frequency MUST be 50HzThe unit MUST use a line voltage of 220VAC to 240VACAt least windows 10The unit must be free standing with height adjustmentSABS Plug : Must comply with both SANS 164-0 and SANS 164-1 (2 fase) |  |  |  | Each
RT286-10-040 | Radiological Equipment | Digital Intraoral imaging Photostimulable phosphor plate system without computerSPECIFICATION An Intra-oral X-ray system utilising plates for quick digilising of x-ray images. The unit MUST have the following features:Must be compatible with AC and DC intra oral X-ray Units.Software including implant module and implant archive. .Fast Ethernet connection. Must include all cabling and installation from the unit to the surgeries.4 imaging plate sets per unit (1 for every surgery) MUST be supplies and includes (per surgery): a. 2 x 22x33mm plates. c. 5 X 31x41 mm Plates d. 1 X 27x45mm plate e. 1 X storage box for plates. f. 100 pieces of hygiene bags per sizeSABS Plug : Must comply with both SANS 164-0 and SANS 164-1 |  |  |  | Each
RT286-10-040 A Accessories | Radiological Equipment | 2 x 22x33mm plates. |  |  |  | Each
RT286-10-040 B Accessories | Radiological Equipment | 2 X 24x40mm plates. |  |  |  | Each
RT286-10-040 C Accessories | Radiological Equipment | 5 X 31x41 mm Plates |  |  |  | Each
RT286-10-040 D Accessories | Radiological Equipment | 1 X 27x45mm plate |  |  |  | Each
RT286-10-040 E Accessories | Radiological Equipment | 1 X storage box for plates. |  |  |  | Each
RT286-10-040F Accessories | Radiological Equipment | 100 pieces of hygiene bags per size |  |  |  | Each
RT286-10-041 | Radiological Equipment | Cone Beam Computed Tomography (CBCT) 3D Dental imaging system small field of view (FOV) SPECIFICATION A Cone Beam CT 3D Dental Imaging System for 3 dimention image at all planes in vieving field At least a 3cm X 3cm reconstruction volume.Standing or seated (wheelchair accessible) Patient position and able to accommodate children.Focal spot of 0.5mm, Anode voltage 60-90kV, X-ray Source – High Frequency, constant potential (DC). Image detector silicon flat panel to be appropriate with the 8cm X 8cm. Must be able to stitch two scans360 degree rotation for Image Acquisition.Automatically adjust radiation dosage to patient’s age and size. Preset functions to adjust for Field of View (FOV), patient size, scan location and radiation dosage.Must include scanner, cabinet with electronic circuitry. A Computer Workstation with all the software. Workstation with 64bit , min 16gig ram, 1tb mem, 2mp diagnostic monitor with gray scale.Software must be able to be programmed for Implants. Software capabilities must include addons for Implant planning, stent design, Endodontic evaluationMust include a computer with LED flat screen to run the system. (same as workstation)Must be able to handle multiple reconstructions from the same scan.Must include DICOM3 compatibility, including DICOM worklist, DICOM PACS, DICOM Store, DICOM transfer, DICOM query/retrieve, DICOM print, DICOM MPPS and DICOM RDSRMust be able to produce TMJ images: sagittal, panoramic sagittal and coronal reconstruction. Must cover the entire range of applications from single implant to complete ENT and Orthogenetic planning and diagnose. Software must make provision for accurate implant placement and implant guide planning Must be able to accommodate standing and sitting patients SABS Plug : Must comply with both SANS 164-0 and SANS 164-1 Pug (2 Phase) |  |  |  | Each
RT286-10-041 A Accessories | Radiological Equipment | Implant planning software for the manufacture of the surgical guide |  |  |  | Each
RT286-10-041 B Accessories | Radiological Equipment | Replacement battery pack compatible with the system supplied |  |  |  | Each
RT286-10-041 C Accessories | Radiological Equipment | Replacemnt bite blocks compatible with the system supplied |  |  |  | Each
RT286-10-042 Accessories | Radiological Equipment | Cone Beam Computed Tomography (CBCT) 3D Dental imaging system large field of view (FOV) SPECIFICATION A Cone Beam CT 3D Dental Imaging System for 3 dimension image at all planes in vieving field At least a 15.5cm X 15.5cm reconstruction volume. Must be capable of performing at least 18x18 FOVStanding or seated (wheelchair accessible) Patient position and able to accommodate children.Focal spot of 0.5mm X-ray Source – High Frequency, constant potential (DC). Image detector silicon flat/cone panel, at least 20cm X 25 cm (field of view). ????360 degree rotation for Image Acquisition.Automatically adjust radiation dosage to patient’s age and size. Preset functions to adjust for FOV, patient size, scan location and radiation dosage.Must include scanner, cabinet with electronic circuitry. A Computer Workstation with all the software. Workstation with 64bit , min 16gig ram, 1tb mem, 2mp diagnostic monitor with gray scale.Software must be able to be programmed for Implants. Software must be able to be programmed for Implants. Software capabilities must include addons for Implant planning, stent design, Endodontic evaluationMust include a computer with LED flat screen to run the system.Must be able to handle multiple recon-structions from the same scan.Software must be able to be converted to DICOM format. Must include DICOM3 compatibility, including DICOM worklist, DICOM PACS, DICOM Store, DICOM transfer, DISOM query/retrieve, DICOM print, DICOM MPPS and DICOM RDSRMust be able to produce TMJ images: sagittal, panoramic sagittal and coronal reconstruction. Must cover the entire range of applications from single implant to complete ENT and Orthogenetic planning and diagnose. Software must make provision for accurate implant placement and implant guide planning Must be able to accommodate standing and sitting patientsSABS Plug : Must comply with both SANS 164-0 and SANS 164-1 (2 Phase) |  |  |  | Each
RT286-10-042 A Accessories | Radiological Equipment | Implant planning software for the manufacture of the surgical guide |  |  |  | Each
RT286-10-042 B Accessories | Radiological Equipment | Replacement battery pack compatible with the system supplied ??? |  |  |  | Each
RT286-10-042 C Accessories | Radiological Equipment | Replacemnt bite blocks compatible with the system supplied, |  |  |  | Each
RT286-10-043 | Radiological Equipment | Dental digital panoramic X-Ray SPECIFICATION An X-ray unit capable of taking digital panoramic images of the mouth to include sinus, TMJ and the whole jawThe unit must be able to cater for adult and paediatric patientsThe tube output voltage must be between 60kV to 90kVFocal dimensions must be 0.5mmThe unit must be capable for the following applications:a. Panoramicb. Temporomandibular Joint (TMJ) ImagingThe unit must have a maximum current rating of 15mAThe maximum exposure time must be 15 secondsa. ≥4GB Ramb. ≥1TB Hard drivec. 17 inch LED screen with 1920x 1080 pixels for diagnosing and review (as per SAHPRA recommendation) d. Software for processing all captured imagese. A keyboardf. A mouse with a cord/ cordlessThe unit MUST use a line voltage of 220VAC to 240VACThe line frequency MUST be 50HzThe unit must be free standing with height adjustmentUPSmust include DICOM3 compatibility, including DICOM worklist, DICOM PACS, DICOM Store, DICOM transfer, DISOM query/retrieve, DICOM print, DICOM MPPS and DICOM RDSRSABS Plug : Must comply with both SANS 164-0 and SANS 164-1 (2 -phase) |  |  |  | Each
RT286-10-044 | Radiological Equipment | Intra Oral Dental X-Ray Dual Function (Analogue & Digital) wall Mounted SPECIFICATION A X-ray unit for capturing intra oral analgue and digital images The unit must be a wall mount type (with an option of a long arm) The tube head MUST be between 60kV to 70kV The unit must have a computerized timer with the following functions:Automatic setting of technique factors in mAsObject program selectionTooth typePatient sizeManual overrideHand switch with 3m cord or cordless The focal point of the unit MUST be from 0.4mm to 0.8mmThe power supply of the unit must be mounted on the unitThe unit MUST use a line voltage of 220VAC to 240VACThe line frequency MUST be 50HzAnalogue and digitalSABS Plug : Must comply with both SANS 164-0 and SANS 164-1 (2 PHase) |  |  |  | Each
RT286-10-045 | Radiological Equipment | Dental Digital panoramic X-Ray with cephalometric arm without a printer SPECIFICATION A Panoramic X-ray with the capability to cephalometric photos comlete without a printerThe unit must be able to cater for adult and paediatric patientsThe tube output voltage must be between 60kV to 90kVFocal dimensions must be 0.5mmThe unit must be capable for the following applications:.Panoramic.Temporomandibular Joint (TMJ) Imaging.Cephalometric The unit must have a maximum current rating of 15mAThe maximum exposure time must be 15 secondsThe unit must be provided with an up to date desktop computer with the following requirements:.≥4GB Ram.≥1TB Hard drive.17 inch LED screen 2 mega pixel.Software for processing all captured images and dicom campatible. Must include DICOM3 compatibility, including DICOM worklist, DICOM PACS, DICOM Store, DICOM transfer, DISOM query/retrieve, DICOM print, DICOM MPPS and DICOM RDSR.A keyboardThe unit must use a line voltage of 220VAC to 240VACThe line frequency must be 50HzThe unit must be free standing with height adjustmentUPSSABS Plug : Must comply with both SANS 164-0 and SANS 164-1 (2 Fase) |  |  |  | Each
RT286-10-046 | Radiological Equipment | Dental Digital panoramic X-Ray with cephalometric with a printer SPECIFICATION A Panoramic X-ray with the capability to cephalometric photos comlete with a printerThe unit must be able to cater for adult and paediatric patientsThe tube output voltage must be between 60kV to 90kVFocal dimensions must be 0.5mmThe unit must be capable for the following applications:PanoramicTemporomandibular Joint (TMJ) ImagingCephalometric The unit must have a maximum current rating of 15mAThe maximum exposure time must be 15 secondsThe unit must be provided with an up to date desktop computer with the following requirements:≥4GB Ram≥1TB Hard drive17 inch LED screen 1920x 1080 pixels for diagnosing and review (as per SAHPRA regulations) Software for processing all captured images and dicom campatible. Must include DICOM3 compatibility, including DICOM worklist, DICOM PACS, DICOM Store, DICOM transfer, DISOM query/retrieve, DICOM print, DICOM MPPS and DICOM RDSRA keyboardThe unit must use a line voltage of 220VAC to 240VACThe line frequency must be 50HzThe unit must be free standing with height adjustmentPrinter (capable to print panoramic as well as cephalometric images)UPSSABS Plug : Must comply with both SANS 164-0 and SANS 164-1 (2- Phase) |  |  |  | Each
RT286-10-047 | Dental surgery small equipment | Dental HandpiecesTURBINE FULL RANGE SPECIFICATION To fit 4-6 hole standard supply tubing.Speed 350 000 rpm minimum at 2.2 kg per cm.Ballbearing, output approximately 10 wattsField replaceable rotor, one spare rotor to be suppliedMetal bur change chuck, pushtype systemQuick action coupling for easy lubricationGuarantee time of turbine must be indicated – minimum 12 monthsAir consumption MUST exceed 36 1/mNoise level not to exceed 75 dbAutoclavable up to 135°CWith fibre optic function To come with durable metal casing (stainless steel) |  |  |  | Each
RT286-10-048 | Dental surgery small equipment | Turbine full range SPECIFICATION To fit 4-6 hole standard supply tubing.Speed 350 000 rpm minimum at 2.2 kg per cm.Ballbearing, output approximately 10 wattsField replaceable rotor, one spare rotor to be suppliedMetal bur change chuck, pushtype systemQuick action coupling for easy lubricationGuarantee time of turbine must be indicated – minimum 12 monthsAir consumption MUST exceed 36 1/mNoise level not to exceed 75 dbAutoclavable up to 135°CWith non fibre optic function To come with durable metal casing (stainless steel) |  |  |  | Each
RT286-10-052 | Dental surgery small equipment | Turbine full range SPECIFICATION To fit 4-6 hole standard supply tubing.Speed 350 000 rpm minimum at 2.2 kg per cm.Ballbearing, output approximately 10 wattsField replaceable rotor, one spare rotor to be suppliedMetal bur change chuck, pushtype systemQuick action coupling for easy lubricationGuarantee time of turbine must be indicated – minimum 12 monthsAir consumption MUST exceed 36 1/mNoise level not to exceed 75 dbAutoclavable up to 135°CWith fibre optic function To come with durable metal casing (titanium) |  |  |  | Each
RT286-10-053 | Dental surgery small equipment | Turbine full range SPECIFICATION To fit 4-6 hole standard supply tubing.Speed 350 000 rpm minimum at 2.2 kg per cm.Ballbearing, output approximately 10 wattsField replaceable rotor, one spare rotor to be suppliedMetal bur change chuck, pushtype systemQuick action coupling for easy lubricationGuarantee time of turbine must be indicated – minimum 12 monthsAir consumption MUST exceed 36 1/mNoise level not to exceed 75 dbAutoclavable up to 135°CWith non fibre optic function To come with durable metal casing (titanium) |  |  |  | Each
RT286-10-054 | Dental surgery small equipment | Led straight handpiece for micromotor - Air full range SPECIFICATION Standard 1 : 1 transmissionBallbearingTo take burs Iso type 104 44, 50 mm with diameter of 2,35 mmMaximum speed 40 000 rpmE-type quick disconnectionWater spray incorporated internallyAutoclavable up to 135°CQuick action coupling for easy lubricationTo come with durable metal casing (titanium) |  |  |  | Each
RT286-10-055 | Dental surgery small equipment | Led straight handpiece for micromotor - Air full range SPECIFICATION Standard 1 : 1 transmissionBallbearingTo take burs Iso type 104 44, 50 mm with diameter of 2,35 mmMaximum speed 40 000 rpmE-type quick disconnectionWater spray incorporated externallyAutoclavable up to 135°CQuick action coupling for easy lubricationTo come with durable metal casing (titanium) |  |  |  | Each
RT286-10-056 | Dental surgery small equipment | Led straight handpiece for micromotor - Air full range SPECIFICATION Standard 1 : 1 transmissionBallbearingTo take burs Iso type 104 44, 50 mm with diameter of 2,35 mmMaximum speed 40 000 rpmE-type quick disconnectionWater spray incorporated internallyAutoclavable up to 135°CQuick action coupling for easy lubricationTo come with durable metal casing (stainless steel) |  |  |  | Each
RT286-10-057 | Dental surgery small equipment | Led straight handpiece for micromotor - Air full range SPECIFICATION Standard 1 : 1 transmissionBallbearingTo take burs Iso type 104 44, 50 mm with diameter of 2,35 mmMaximum speed 40 000 rpmE-type quick disconnectionWater spray incorporated extenallyAutoclavable up to 135°CQuick action coupling for easy lubricationTo come with durable metal casing (stainless steel) |  |  |  | Each
RT286-10-058 | Dental surgery small equipment | LED contra-angle handpiece for micromotor - Air full range SPECIFICATION Standard transmission 1 : 1 bluePush button type systemBallbearing typeMaximum speed 40 000 rpmTo accommodate ISO type 204 2.35 mm diameter latch type bursE-type quick disconnectionWater spray incorporated internallyAutoclavable up to 135°CQuick action coupling for easy lubricationTo come with durable metal casing (stainless steel)0 |  |  |  | Each
RT286-10-059 | Dental surgery small equipment | LED contra-angle handpiece for micromotor - Air full range SPECIFICATION Standard transmission 1 : 1 bluePush button type systemBallbearing typeMaximum speed 40 000 rpmTo accommodate ISO type 204 2.35 mm diameter latch type bursE-type quick disconnectionWater spray incorporated internallyAutoclavable up to 135°CQuick action coupling for easy lubricationTo come with durable metal casing (titanium) |  |  |  | Each
RT286-10-060 | Dental surgery small equipment | LED contra-angle handpiece for micromotor - Air full rangeSPECIFICATION Reduction hand piece GreenPush button type systemBallbearing typeTo accommodate ISO type 204 2.35 mm diameter latch type bursE-type quick disconnectionWater spray incorporated internallyAutoclavable up to 135°CQuick action coupling for easy lubricationTo come with durable metal casing (stainless steel) |  |  |  | Each
RT286-10-061 | Dental surgery small equipment | LED contra-angle handpiece for micromotor - Air full range SPECIFICATION Reduction hand piece GreenPush button type systemBallbearing typeTo accommodate ISO type 204 2.35 mm diameter latch type bursE-type quick disconnectionWater spray incorporated internallyAutoclavable up to 135°CQuick action coupling for easy lubricationTo come with durable metal casing (titanium) |  |  |  | Each
RT286-10-062 | Dental surgery small equipment | LED contra-angle handpiece for micromotor - Air full range SPECIFICATION Speed increasing redPush button type systemBallbearing typeTo accommodate ISO type 204 2.35 mm diameter latch type bursE-type quick disconnectionWater spray incorporated internallyAutoclavable up to 135°CQuick action coupling for easy lubricationTo come with durable metal casing (stainless steel) |  |  |  | Each
RT286-10-063 | Dental surgery small equipment | LED contra-angle handpiece for micromotor - Air full range SPECIFICATION Speed increasing redPush button type systemBallbearing typeTo accommodate ISO type 204 2.35 mm diameter latch type bursE-type quick disconnectionWater spray incorporated internallyAutoclavable up to 135°CQuick action coupling for easy lubricationTo come with durable metal casing (titanium) |  |  |  | Each
RT286-10-064 | Dental surgery small equipment | LED prophylaxis handpiece full range SPECIFICATION Reduction GreenTo fit E-Type quick disconnectionSealed ballbearingTo accept snap-on and or screw in cups and brushesAutoclavable up to 135°CQuick action coupling for easy lubricationTo come with durable metal casing (titanium)0 |  |  |  | Each
RT286-10-065 | Dental surgery small equipment | LED prophylaxis handpiece full range SPECIFICATION Reduction GreenTo fit E-Type quick disconnectionSealed ballbearingTo accept snap-on and or screw in cups and brushesAutoclavable up to 135°CQuick action coupling for easy lubricationTo come with durable metal casing (stainless steel)0 |  |  |  | Each
RT286-10-066 | Dental surgery small equipment | Micromotor - Air full range SPECIFICATION Speed range 5 000 – 40 000 rpmOutput 20W Power – min Ncm Torque minE-Type quick couplingForward and reverse rotationWeight not to exceed 90 gm. Speed control selector by means of control ringWater spray/air facility to be incorporated internallyMaximum air consumption 45 1/mAutoclavable up to135°CQuick action coupling for easy lubricationTo come with durable metal casing (stainless steel)SPECIFICATION Speed range 5 000 – 40 000 rpmOutput 20W Power – min Ncm Torque minE-Type quick couplingForward and reverse rotationWeight not to exceed 90 gm. Speed control selector by means of control ringWater spray/air facility to be incorporated internallyMaximum air consumption 45 1/mAutoclavable up to135°CQuick action coupling for easy lubricationTo come with durable metal casing (stainless steel) |  |  |  | Each
RT286-10-067 | Dental surgery small equipment | Micromotor - Air full range SPECIFICATION Speed range 5 000 – 40 000 rpmOutput 20W Power – min Ncm Torque minE-Type quick couplingForward and reverse rotationWeight not to exceed 90 gm. Speed control selector by means of control ringWater spray/air facility to be incorporated internallyMaximum air consumption 45 1/mAutoclavable up to135°CQuick action coupling for easy lubricationTo come with durable metal casing (stainless steel)With LED capability |  |  |  | Each
RT286-10-068 | Dental surgery small equipment | Electric Micro Motors ( stand alone)SPECIFICATION Selection 16: 1 , 1: 5 and 1: 1LED OpticBrushless motor Internal water spraySmooth rotation & minimul vibrationSpeed range : 2000 to 40 000 rpmMaximum Torque 3.4 N/cm |  |  |  | Each
RT286-10-070 | Dental equipment | Dental operator stool SPECIFICATION An ergonomic stool used by the operator2-way adjustable backrest for lumbar support. Anatomical shaped.Contoured seat cushions with a “rolled” edge to reduce pressure on the legs.Easy to reach adjustment paddle for increase lumbar support.Tear resistant upholsterySeamless upholstery for easier cleaning.Upholstary, design and colour to be in line with the patient operator chair5 leg cast aluminium base Twin wheel castors.The unit must have height adjustableThe unit must have active and passive backrest support |  |  |  | Each
RT286-10-071 | Dental equipment | Dental delivery system SPECIFICATION A Dental delivery system, complete with tubing,spitoon, scaler, clean water system and curing lightUnit must have four/ six (4/6) Handpiece tube holes or more (all modes/functions) with self activationsAutomatic selection of cutting H/P3 in 1 Syringe with pinch valve technique and sterilizable sleeve and nozzleThe unit to have at least two brake handles (one on each side of delivery head)To have touchpad for controlsConsole with water flow control valve for each handpiece, if not incorporated in handpiece: must have a minimum of 4 tubes Console with 4 control switches for the following:a) on/off Handpiece water/air spray b) on/off handpiece air spray coolantIndividual handpiece pressure adjustment reliable water retraction system requiredThe heads must be durable, providing many years of reliabilityFree hanging straight/ whip arm air outlet hoses are preferred.The clean water bottle must be standard (1 litre)It must come with a piezo scaler system with foot switchThe curing light (LED with at least 3 settings levels) must also be supplied with the systemMaster on/off switch provided at connection boxAdditional Option : Over the patient deliveryRear mounted deliverySide deliveryRear mounted deliveryAs per surgery requirements the curing light supplied as a stand alone or intergrated into unitMaster on/off switch provided at connection boxSABS Plug : Must comply with both SANS 164-0 and SANS 164-1 (2 Phase)Additional Option : Over the patient deliverySide deliveryRear mounted deliveryAs per surgery requirements the curing light supplied as a stand alone or intergrated into unit |  |  |  | Each
RT286-10-072 | Dental equipment | Dental LED LightSPECIFICATION An operator light for ilumination of work fieldIt MUST be LED lightTouch less sensor ON/OFF and manual controlVoltage 220 – 240, 50 Hz Mounted to Chair / Ceiling Rotational axes: Horizontal, vertical, diagonalConstant colour temperature 5,000KTemperature increase not more than 5 degrees C above room temperature at maximum intensity.Light pattern: 85mm - 95mm x 145mm - 160mm at 700 mmAdjustable light intensity: 15 000Lux - 30 000LuxShadow less illuminationOn/Off switch in lamp yoke, and provision for altering light intensity, without changing colour temperaturePower Consumption: 12W - 17W or lessWith Mode SelectionThe light must have a shield that is scratch and discolouration resistant for protection against bulb shattering |  |  |  | Each
RT286-10-073 | Dental equipment | Dental assistant stool SPECIFICATION An ergonomic stool used by the Dental assistant2-way adjustable backrest for lumbar support. Anatomical shaped.Contoured seat cushions with a “rolled” edge to reduce pressure on the legs.Easy to reach adjustment paddle for increase lumbar support.Tear resistant upholsterySeamless upholstery for easier cleaning.Upholstary, design and colour to be in line with the patient operator chair5 leg cast aluminium base Twin wheel castors. Adjustable torso support that can be converted for left or right handed operators. Height adjustable foot ring). The unit must be height adjustable |  |  |  | Each
RT286-10-074 | Dental equipment | Water Unit / Cuspidor / Spitton SPECIFICATION A chairside water and spitoon systemChair-mounted either left or rightSpittoon must be removable without any tools, for easy cleaningCeramic bowls spittoonWarm water tumbler fillerSpittoon flush facilitiesSolids trap for spittoon bowlSelective holder consisting of:Solids collectorSolids collectorOne HVE suction tubeOne saliva ejector tubeCup fill shut off valve incorporatedLow voltage remote control for suction.Autoclavable HVE cannulas to be supplied |  |  |  | Each
RT286-10-074 A Accessories |  | 3 in 1 syringe |  |  |  | Each
RT286-10-074 B Accessories |  | The bowl can also be stainless steel |  |  |  | Each
RT286-10-075 | Dental equipment | Basic Dental chair SPECIFICATION A dental chair facility to do oprative procedures in the oral envioronment Dental ChairAnatomically contoured seat and backrest for comfortable full support in all positions suitable for treatment of both adults & children The chair offered must have double articulating headrest (Optional) which is reversable (cushion included) Adjustable backrest - thickness 25mm - 50mmUpholstery must be seamless durable waterproof material that can be easily cleaned and withstand disinfenctants and cleaning chemicals. Upholstery must be replaceableReplaceable plastic cover must be provided to protect upholstery in shoe areaArmrest: Positions swing out/ down ArmChair must be ProgrammableThe chair must be floor mountAccessory foot controls MUST be fitted to chair base to facilitate both height and trendelenburg adjustmentRange of lift height 342.90 mm to 800.10 mmRecline range +62° to -12°Total length at maximum extension 1820mm - 2000 mmBase (H x W x D) : 600mm -650mm x 500mm - 560mm x 490mm - 510mmElectro mechanical/hydraulic lift system fitted in chair baseThe seat surface must be 410mm - 480mm x 400mm - 480mm (LxB)Carry capacity of 200kgChair swivel 60° - 70° (≤35° each side of center)To operate on 240 VAC 50/60Hz, fitted with standard hospital 3-prongs plug. |  |  |  | Each
RT286-10-075 A Accessories |  | Control block with programable positions as well as to facilitate both height and trendelenburg adjustment |  |  |  | Each
RT286-10-076 | Dental equipment | Dental treatment centre (Surgery unit) "SPECIFICATION A dental surgery chair for either a L-hand or R-hand configuration. The unit must have the following features:" Chair mounted delivery system: The unit must operate by means of a reliable pneumatic/electronic control block and single diaphragm The unit must be centrally mounted to facilitate 270 degree movement front or back of the chair for left and right hand dentists. The facility must be immediate without the use of any technical assistance, plumbing and electrical plugs The design of the unit must be seamless with rounded edges and the carcass in the aerosol zone should be seamless and smooth The height adjustment must be manual and held firmly in position by a mechanism The unit must have automatic shut–off controls for all the handpieces Adjustment of water coolant must be provided for by precise needle valve flow controls knobs. These knobs must be easily be removed for chemical sterilisation. The unit must be supplied with 3 sets of removable knobs The controls of the unit must be positioned out of the aerosol zone Water retraction must be prohibited on all handpieces by means of a reliable non retractable valve system The dental unit must have an effective tubing and handpiece flushing system that allows increased water volume without running the handpieces Individual air pressure adjustment must be provided for each handpiece an optional analogue gauge must be fitted to the dental unit to indicate working air pressure of all handpieces A disposable cup holder for waste must be fitted to the dental unit The dental unit must accommodate a standard instrument tray (Approximately size of 250 X 350 mm) and the tray must be easily removable but firmly kept in place Two removable trays (Approximately size of 250 X 350 mm) must be provided The dental unit must accommodate the following: 1 x 3 way syringe complete 2 x Fibre optic tubing’s 1 x Electric Motor The 3 way syringe head and tip must be quick disconnectable and autoclavable at 135 degrees Celsius 3 x Non-disposable tips must be supplied per syringe The 3 way syringe head must be able to accommodate disposable tips The syringe head must be non water retracting Free hanging straight hoses (Tubing) must withstand at least 100PSI and proof of durability of tubing may be required The hoses must be smooth for quick and effective wipe down The power pack for the fibre optics must be incorporated into the dental unit A programmable touch pad option must be available on the unit Pneumatic/ electronic pedal type foot control for variable speed regulation of all handpieces to include water spray control and water On/Off switch An arm control must be provided and all umbilicals must be concealed and protected in the arm Load carrying capacity of the arm control must not be less than 5 Kg above the weight of the dental unit and must remain fixed A connection box must be provided to contain air and water filters. The connection box must supply services to the dental unit, the spittoon and chair without visible tubing The power supply (220-240 V, 50Hz fused mains input) to be provided from the connection box to the dental unit, spittoon and chair High performance operating LED lighting mounted and be able to accommodate Left and right handed dentists without any technical assistance, plumbing or electrical modifications The power supply must be 220 -240 V, 50HZ Temperature increase must not be more than 5 degrees Celsius above room temperature at a distance of 80 cm when the light is on maximum intensity Light must have a varying intensity device The light must be provided with an easy access On/Off switch on lamp head Operating field distance of 70 centimetres must be approximately 5 X 20 CM with no glare for the patient The light must have a shield that is scratch and discolouration resistant for protection against bulb shattering. Light must be rounded for easy cleaning and increases asepsis Lamp head to be adjustable in 125 degrees vertical and 180 degrees horizontal swivel The light must move smoothly and all pivots must be permanently lubricated The light arm must have a single rounded arm without transformer to support adjustable tension spring counter balance Light handles must be able to accept barrier protection bags An On/Off master shut-off must be located where it is visible in the centre. Cuspidor and assistant arm must have the following The units must have the ability to be swivel for left en right hand use as must be possible without outside technical support, plumbing or electricity change Spittoon bowl must be of an one removable piece, easy to clean design that allows easy access to inside of the housing Bowl flush system and cup filler must be manually controlled from 2 sides The bowl housing must be mounted to accommodate water cup and to prevent accidental slip of the cup The bowl housing must have two separate facilities for cup filling and spittoon rinse The spouts for rinsing and cup filling must be easy removable for autoclaving at 135 degrees Celsius and be constructed of a rust free material The spittoon bows must provide with a solid trap The dental assistant side central vacuum and water instrumentation be mounted on the cuspidor to be left and right interchangeable and must include: Easy clean high volume evacuation Saliva ejector x 1 Three way syringe x 1 Easy clean high volume evacuating valve must have positive On/Off controls and be autoclavable at 135 degrees Celsius The high volume autoclavable valve must be removable and quick disconnectable without the use of tools High volume autoclavable valves be supplied per dental assistant’s station. The saliva ejector must have a positive On/Off control and must have stop in full open and closed position and must be autoclavable at 135 degrees Celsius Saliva ejector must be easy removable without the use of tools Two saliva ejectors and one high volume suction with no rubber tips must be supplied per dental assistant station 3 x Non-disposable tips must be supplied per dental assistant’s station The 3 way syringe head must be able to accommodate disposable tips The syringe head must be non water retracting Free hanging straight hoses (Tubing) must withstand at least 100PSI and proof of durability of tubing may be required All suction hoses must be smooth and removable without any tools for easy disinfection and maintenance Cuspidors must have a factory fitted clean water system with a bottle design that is easy to remove and to clean Clean water system must have a minimum capacity of l (one) litre One additional bottle (in total two bottles of min 1 (one) litre capacity per bottle system ) must be supplied Connection (umbilical) from cuspidor/spittoon to floor Box must be supplied to conceal electronics components and should not interfere with accessibility to the chair The dental chair must have the following features: The seat and backrest must be anatomically contoured for full supporting all positions suitable for treatment of adults and children The backrest must provide adequate support to elbows and lower lumber part of the body, when the backrest is fully reclined Backrest with build in Trendelburg mechanism must have an extendable double articulating anatomically shaped headrest with suitable adjustment for children The headrest must be easily removable without the use of any tools The headrest ( optional) must be easily be reversed to allow for operating on patients outside the chair (for example special needs patient in a wheelchair) Backrest of the upper section not be thicker than 65 mm and must be rounded in shape for easy access with no obstructions to the operator Armrest must be fixed and must be movable with bushing pivot points to swivel out of the way The armrest must not interfere with the operator The minimum height of the seat must not be more than 450mm in the completely down position The seat must support a pressure or any load per square inch of a minimum of 100Kg in any area of the seat The stress points of the seat and backrest must be supported by cast aluminium or metal The seat and backrest must be upholstered in a durable, washable and cool breathing material with softly rounded upholster edges with no exposed seems at patient contact The upholster for the seat and backrest must be replaceable in the dental surgery without the use of outside technical assistance with minimal downtime Replaceable plastic cover must be provided to protect the upholstery in the shoe area and must be easily be changed and fastened under the chair with optional Velcro The base plate must not interfere with the operator The back of the chair must have the facility to allow for left and right rotating dental unit in order to ensure comfortable operation for either left or right handed operator. This facility must provide a path for eclectic wiring and must form an integrated part of the complete system Chair movement must be by means of low voltage switching and foot control must facilitate movement The foot control must be fitted to facilitate manual settings Adjustment of the foot control memory settings must be without the use of tools and must be user friendly The foot control must be made of a durable material The foot control must not be fitted to the base of the chair The foot control must be stable on the floor areas and must not slip away when operated A lift system must be completely sealed and be able to lift a capacity of minimum 180Kg The chair movement rate facilitated by the lift system must be adjustable for a minimum of four chair positions, i.e. base up, base down, back up, and back down The stationary portion of the base must house an integrated electric motor The reservoir must be transparent so that the fluid can be manually inspected and not be pressurised Fittings and tubings used in the low pressure portion of the system must be attached at all points with barbed fitting and retained upon the barb by a retaining clamp The stationary part of the base must provide for easy access to mechanical components by removal of a formed secured housing The system excludes hand-pieces, suction units and compressors The unit must have safety machanism to stop the movement of the chair when it hits resistance The unit must meet the relevant international safety standards SABS Plug : Must comply with both SANS 164-0 and SANS 164-1 (2 Phase) |  |  |  | Each
RT286-10-077 | Radiological Equipment | Medical Review Monitor Colour : White or Black Must comply with SAHPRA SABS Plug : Must comply with both SANS 164-0 and SANS 164-1 (2 Phase) |  |  |  | Each
RT286-10-078 | Dental surgery small equipment | Dental LED light - portable An operator light for illumination of work field Sensor ON/OFF and manual control Voltage 220- 240 , 50Hz Detachable base Easily detachable and/or bendable stand Constant colour temperature of up to 5000k Mode selection Temperature increase not more than 5 degree celsius above room temperature at maximum intensity Adjustable light field with maximum 7000 LUX Colour rendering index (CRI) must be above 95 Additional accessories Protective carrier bag for ease of transportation Optional Light shield that is scratch and discoloration resistant Attachable to the mobile chair |  |  |  | Each
RT286-11-002 | Dental surgery small equipment | Corded Laser Caries Detector (Tabletop) SPECIFICATION A laser caries detection aid that measures photological changes in tooth structure caused by caries." Must be able to be used for pit-and -fissure areas, even when outer surface is intact. LED light integrated in the handpiece Uses light energy with no danger to patient to operator. Standard cord connection between handpiece and control unit Must be portable. Initial delivery must include the following: 1 X Fiber -optic tubing. 1 X Unit 2 X tips 1 X sterilization cassette. 1 X autoclavable handpiece sheath 1 X calibration disk 1 X operating booklet Be able to accommodate a perio type tip for detecting subgingival caries SABS Plug : Must comply with both SANS 164-0 and SANS 164-1 |  |  |  | Each
RT286-11-002 A Accessories | Dental surgery small equipment | Replacement tip for the Corded Laser Caries Detector (Tabletop) |  |  |  | Each
RT286-11-002 B Accessories | Dental surgery small equipment | Replacement safety glasses for the Corded Laser Caries Detector (Tabletop) |  |  |  | Each
RT286-11-003 | Dental surgery small equipment | Single-Tooth Anesthesia System "SPECIFICATION A single tooth anesthesia system that controls the flow and pressure of dental anesthesia. Must be able to be used for all dental application" Initial delivery must include the following: 1 X Drive Unit 1 X foot control pedal 1 X instruction manual (special conditions) Box of 50 X handpieces with needles SABS Plug : Must comply with both SANS 164-0 and SANS 164-1 |  |  |  | Each
RT286-11-003 A Accessories | Dental surgery small equipment | Box of 50 X handpieces with needles for the Single-Tooth Anesthesia System |  |  |  | Each
RT286-11-004 | Dental surgery small equipment | Pulp Tester "SPECIFICATION A pulp tester for testing the lity of teeth." Must be able to be used on anterior and posterior teeth. Must be portable Digital display Offers different rates of stimulus Rechargeable batteries Initial delivery must include the following: 1 X Pulp tester Different sizes of autoclavable tip |  |  |  | Each
RT286-11-004 A Accessories | Dental surgery small equipment | Batteries for the Pulp Tester |  |  |  | Each
RT286-11-004 B Accessories | Dental surgery small equipment | Tips for the Pulp Tester |  |  |  | Each
RT286-11-005 | Endodontic equipment | Apex Locator "SPECIFICATION An apex locator for endodontic procedures" The product must have multi frequency technology being able to work in wet and dry canals Must be portable Readable display indicating different lengths of the root canal treatment Initial delivery must include the following: Rechargeable battery Battery charger 1 X Drive Unit 2 X Lip hooks Measuring wire Fork (Endo file holder) |  |  |  | Each
RT286-11-005 A Accessories | Endodontic equipment | Rechargeable battery for the Apex Locator |  |  |  | Each
RT286-11-005 B Accessories | Endodontic equipment | Battery charger for the Apex Locator |  |  |  | Each
RT286-11-005 C Accessories | Endodontic equipment | 1 X Drive Unit for the Apex Locator |  |  |  | Each
RT286-11-005 D Accessories | Endodontic equipment | 2 X Lip hooks for the Apex Locator |  |  |  | Each
RT286-11-005 E Accessories | Endodontic equipment | Measuring wire for the Apex Locator |  |  |  | Each
RT286-11-005 F Accessories | Endodontic equipment | Fork (Endo file holder) for the Apex Locator |  |  |  | Each
RT286-11-006 | Dental surgery small equipment | Fluorescence visualization unit "SPECIFICATION A flourescence visualization unit for early detecting mucosal abnormalities , including oral cancer before it can be seen under ordinary white light. " Must be non invasive. Must be portable. No rinses or or dyes to be used. Screening per patient not longer than 3 min Rechargeable for handheld use Initial delivery must include the following: 1 X flourecence handpiece 1 X rechargeable battery 1 X charging cradle 1 X external power supply-brick 1 X pair patient safty glasses 1 X examination guide Compatable Camera system for photo documentation for the Fluorescence visualization unit |  |  |  | Each
RT286-11-006 A Accessories | Dental surgery small equipment | Caps per box to fit unit for individual patients for the Fluorescence visualization unit |  |  |  | Each
RT286-11-006 B Accessories | Dental surgery small equipment | Sheets for the Fluorescence visualization unit |  |  |  | Each
RT286-11-006 C Accessories | Dental surgery small equipment | Compatable Camera system for photo documentation for the Fluorescence visualization unit |  |  |  | Each
RT286-11-006 D Accessories | Dental surgery small equipment | Rechargable batteries |  |  |  | Each
RT286-11-006 E Accessories | Dental surgery small equipment | Patient glassess |  |  |  | Each
RT286-11-007 | Dental equipment | Piezoelectric surgery unit "SPECIFICATION A piezoelectric unit for bone surgery." Must be able to deliver about 60 watts. Must be portable. Must include autoclavable LED handpiece. Must have a Large control screen Must have a pump controlled irrigation system for constant precise flow Initial delivery must include the following: 1 X Piezotome LED Handpiece 1 X foot control 1 X power cord 1 X Irrigation bracket 1 X Handpiece holder SABS Plug : Must comply with both SANS 164-0 and SANS 164-1 |  |  |  | Each
RT286-11-007 A Accessories | Dental equipment | 1 X Piezotome LED Handpiece for the Piezoelectric surgery unit |  |  |  | Each
RT286-11-007 B Accessories | Dental equipment | 1 X foot control for the Piezoelectric surgery unit |  |  |  | Each
RT286-11-007 C Accessories | Dental equipment | 1 X power cord for the Piezoelectric surgery unit |  |  |  | Each
RT286-11-007 D Accessories | Dental equipment | 1 X Irrigation bracket for the Piezoelectric surgery unit |  |  |  | Each
RT286-11-007 E Accessories | Dental equipment | 1 X Handpiece holder for the Piezoelectric surgery unit |  |  |  | Each
RT286-11-008 | Dental surgery small equipment | Surgical Contra Angle Handpieces with 45 degree angle ( 1:2.7 with light) "SPECIFICATION Surgical handpiece with light " Must 1:2.7ratio. Must be contra angled Must be push button type Must be able to hold 1.6mm di FG -type burs Maximum drive speed 40 000 rpm Must be able to work with any surgical unit Must have an external spray Must be able to be fully dismantled without tools for ease of cleaning prior to sterilization |  |  |  | Each
RT286-11-009 | Dental surgery small equipment | Prosthodontic Electronic Screwdriver "SPECIFICATION A cordless and electronic prostho screwdriver to work with all latch type burs and torque wrenches and drivers." Must have 1:2.7ratio. Torque control can be set from 8 to38 Ncm to replace all torque wrenches and drivers Initial delivery must include the following: 1 X Rechargeable battery 1 X Handpiece drive -contrangle 1 X Hexagon chucking system 1 X Charging station LCD Digital Display |  |  |  | Each
RT286-11-009 A Accessories | Dental surgery small equipment | 1 X Rechargeable battery for the Prosthodontic Electronic Screwdriver |  |  |  | Each
RT286-11-009 B Accessories | Dental surgery small equipment | 1 X Handpiece drive -contrangle for the Prosthodontic Electronic Screwdriver |  |  |  | Each
RT286-11-009 C Accessories | Dental surgery small equipment | 1 X Hexagon chucking system for the Prosthodontic Electronic Screwdriver |  |  |  | Each
RT286-11-009 D Accessories | Dental surgery small equipment | 1 X Charging station for the Prosthodontic Electronic Screwdriver |  |  |  | Each
RT286-11-010 | Maintenance | Cleaning, lubrication and purging system for dental handpieces "SPECIFICATION A cleaning, lubrication and purging system to maintain dental handpieces. " Must have a motor type that turns gears on attachment for maintenance of moving parts. Must have chuck maintenance section for bur retention. Must have an internal oil suction that removes oil and debris to minimize cleanup. Must have provision for an adapter e.g. multiflex for maintaining turbine handpieces. Must be able to accept other adapters to maintain various brands an handpiece types SABS Plug : Must comply with both SANS 164-0 and SANS 164-1 Initial delivery must include the following: 1 X Service adapter 1 X Charging cradle 1 X Air hose 1 X Window cover 1 X Compatible Oil |  |  |  | Each
RT286-11-010 A Accessories | Maintenance | 1 X Service adapter for the Cleaning, lubrication and purging system for dental handpieces |  |  |  | Each
RT286-11-010 B Accessories | Maintenance | 1 X Charging cradle for the Cleaning, lubrication and purging system for dental handpieces |  |  |  | Each
RT286-11-010 C Accessories | Maintenance | 1 X Air hose for the Cleaning, lubrication and purging system for dental handpieces |  |  |  | Each
RT286-11-010 D Accessories | Maintenance | 1 X Window cover for the Cleaning, lubrication and purging system for dental handpieces |  |  |  | Each
RT286-11-010 E Accessories | Maintenance | 1 X Compatible Oil for the Cleaning, lubrication and purging system for dental handpieces |  |  |  | Each
RT286-11-011 | Dental surgery small equipment | Hydro Abrasion Cavity Preparation and Polishing System "SPECIFICATION Hydro abrasion cavity preparation and polishing system for treatment without anesthetic." Must have a motor type that turns gears on attachment for maintenance of moving parts. Must have multiple tips for different procedures. Should have color coded cartridge system . Must be able to utilize 26 micron to 50 micron Aluminum Oxide Powder an and sodium bicarbonate . Must be able to utilize single patient nozzle and tip Foot Control SABS Plug : Must comply with both SANS 164-0 and SANS 164-1 Initial delivery must include the following: 1 X 29 micron Aluminum Oxide powder (Fine cutting Powder) container. 1 X 53 micron Aluminum Oxide powder(Course Cutting Powder) container. 1 X Stain Removal container (Sodium Bicarbonite) 1 X Polishing and profhylaxis Container1 50 X Disposable tips |  |  |  | Each
RT286-11-011 A Accessories | Dental surgery small equipment | 29 micron Aluminum Oxide powder (Fine cutting Powder) container for the Hydro Abrasion Cavity Preparation and Polishing System |  |  |  | Each
RT286-11-011 B Accessories | Dental surgery small equipment | 53 micron Aluminum Oxide powder(Course Cutting Powder) container for the Hydro Abrasion Cavity Preparation and Polishing System |  |  |  | Each
RT286-11-011 C Accessories | Dental surgery small equipment | Stain Removal container (Sodium Bicarbonite) for theHydro Abrasion Cavity Preparation and Polishing System |  |  |  | Each
RT286-11-011 D Accessories | Dental surgery small equipment | Polishing and profhylaxis Container1 for the Hydro Abrasion Cavity Preparation and Polishing System |  |  |  | Each
RT286-11-011 E Accessories | Dental surgery small equipment | Disposal tips sizes 0.4/0.6/08mm for the Hydro Abrasion Cavity Preparation and Polishing System |  |  |  | Each
RT286-11-012 | Endodontic equipment | Endodontic Irrigation and Debridement Apparatus "SPECIFICATION Endodontic irrigation and debridement apparatus for disruption of smear layer of biofilm" Must have a contra-angled design for access to posterior teeth. Must have multiple tips for different procedures. Should have at least 3 speed motor (sonic) for options ranging from 1800 up to 10200 cycles per minute Must be cordless. Tips must be snap-on/snap-off design. Tip sizes must be color coded Initial delivery must include the following: 25 X of each type of disposable tips, small, medium and large tips. 100 X protective sleeves for the unit 1 X sequence card Batteries |  |  |  | Each
RT286-11-012 A Accessories | Endodontic equipment | Disposable tips, small, medium and large tips for the Endodontic Irrigation and Debridement Apparatus |  |  |  | Each
RT286-11-012 B Accessories | Endodontic equipment | Protective sleeves for the unit for the Endodontic Irrigation and Debridement Apparatus |  |  |  | Each
RT286-11-012 C Accessories | Endodontic equipment | Sequence card for the Endodontic Irrigation and Debridement Apparatus |  |  |  | Each
RT286-11-013 | Endodontic equipment | Rotating Corded Endodontic Motor, Handpiece and Apex Locator Combination "SPECIFICATION Endodontic motor and handpiece combination for rotating NiTi files" Must have a contra-angled design for access to posterior teeth. Must have auto reverse function. Should have adjustable speed motor ranging from 250 up to 1200 rpm Adjustable tourque ranging from 0.6 - 5.2 N.cm Integrated apex locator Self callibration function Rechargeable batteries Stand (If separate from the charging unit) Charging Unit with SABS Plug : Must comply with both SANS 164-0 and SANS 164-1 (If separate from the stand) SABS Plug : Must comply with both SANS 164-0 and SANS 164-1 |  |  |  | Each
RT286-11-013 Aaccessories | Endodontic equipment | Rechargeable batteries for the Rotating Corded Endodontic Motor, Handpiece and Apex Locator Combination |  |  |  | Each
RT286-11-013 B Accessories | Endodontic equipment | Stand (If separate from the charging unit) for the Rotating Corded Endodontic Motor, Handpiece and Apex Locator Combination |  |  |  | Each
RT286-11-013 C Accessories | Endodontic equipment | Charging Unit with SABS Plug : Must comply with both SANS 164-0 and SANS 164-1 (If separate from the stand) for the Rotating Corded Endodontic Motor, Handpiece and Apex Locator Combination |  |  |  | Each
RT286-11-013 D Accessories | Endodontic equipment | Contra angle head 2.1 incl conection fork for the Rotating Corded Endodontic Motor, Handpiece and Apex Locator Combination |  |  |  | Each
RT286-11-013 E Accessories | Endodontic equipment | Handpiece and lead without the head for the Rotating Corded Endodontic Motor, Handpiece and Apex Locator Combination |  |  |  | Each
RT286-11-013 F Accessories | Endodontic equipment | 1 X mesurement cable and 2 hooks for the Rotating Corded Endodontic Motor, Handpiece and Apex Locator Combination |  |  |  | Each
RT286-11-013 G Accessories | Endodontic equipment | Duel hook for the Rotating Corded Endodontic Motor, Handpiece and Apex Locator Combination |  |  |  | Each
RT286-11-013 H Accessories | Endodontic equipment | Lip clips for the Rotating Corded Endodontic Motor, Handpiece and Apex Locator Combination |  |  |  | Each
RT286-11-013 I Accessories | Endodontic equipment | Fork for apex measurements for the Rotating Corded Endodontic Motor, Handpiece and Apex Locator Combination |  |  |  | Each
RT286-11-014 | Endodontic equipment | Rotating Cordless Endodontic Motor and Handpiece Combination "SPECIFICATION Endodontic motor and handpiece combination for rotating NiTi files- cordles" Must have a contra-angled design for access to posterior teeth. Must have auto reverse function. Should have adjustable speed motor ranging from 250 up to 1200 rpm Adjustable tourque ranging from 0.6 - 5.2 N.cm Self callibration function Rechargeable Batteries Stand (If separate from the charging unit) Charging Unit with SABS Plug : Must comply with both SANS 164-0 and SANS 164-1 (If separate from the stand) |  |  |  | Each
RT286-11-014 A Accessories | Endodontic equipment | Rechargeable Batteries for the Rotating Cordless Endodontic Motor and Handpiece Combination |  |  |  | Each
RT286-11-014 B Accessories | Endodontic equipment | Stand for the Rotating Cordless Endodontic Motor and Handpiece Combination |  |  |  | Each
RT286-11-014 C Accessories | Endodontic equipment | Charging Unit with SABS Plug : Must comply with both SANS 164-0 and SANS 164-1 for the Rotating Cordless Endodontic Motor and Handpiece Combination |  |  |  | Each
RT286-11-014 D Accessories | Endodontic equipment | Contra angle head for the Rotating Cordless Endodontic Motor and Handpiece Combination |  |  |  | Each
RT286-11-014 E Accessories | Endodontic equipment | Handpiece without the head for the Rotating Cordless Endodontic Motor and Handpiece Combination |  |  |  | Each
RT286-11-015 | Endodontic equipment | Corded Rotating and reciprocating Endodontic Motor and Handpiece Combination "SPECIFICATION Endodontic motor and handpiece combination for rotating and reciprocating NiTi files" Must have a contra-angled design for access to posterior teeth. Must have auto reverse function. Should have adjustable speed motor ranging from 250 up to 1200 rpm Adjustable tourque ranging from 0.6 - 5.2 N.cm Self callibration function Rechargeable batteries Stand (If separate from the charging unit) Charging Unit with SABS Plug : Must comply with both SANS 164-0 and SANS 164-1 (If separate from the stand) SABS Plug : Must comply with both SANS 164-0 and SANS 164-1 |  |  |  | Each
RT286-11-015 A Accessories | Endodontic equipment | Rechargeable batteries for the Corded Rotating and reciprocating Endodontic Motor and Handpiece Combination |  |  |  | Each
RT286-11-015 B Accessories | Endodontic equipment | Stand (If separate from the charging unit) for the Corded Rotating and reciprocating Endodontic Motor and Handpiece Combination |  |  |  | Each
RT286-11-015 C Accessories | Endodontic equipment | Charging Unit with SABS Plug : Must comply with both SANS 164-0 and SANS 164-1 (If separate from the stand) for the Corded Rotating and reciprocating Endodontic Motor and Handpiece Combination |  |  |  | Each
RT286-11-015 D Accessories | Endodontic equipment | Contra angle head for the Corded Rotating and reciprocating Endodontic Motor and Handpiece Combination |  |  |  | Each
RT286-11-015 E Accessories | Endodontic equipment | Handpiece without the head for the Rotating Cordless Endodontic Motor and Handpiece Combination |  |  |  | Each
RT286-11-015 F Accessories | Accessories | Handpiece Stand for the Corded Rotating and reciprocating Endodontic Motor and Handpiece Combination |  |  |  | Each
RT286-11-015 G Accessories | Endodontic equipment | Spray Nozzle (F-type) for the Corded Rotating and reciprocating Endodontic Motor and Handpiece Combination |  |  |  | Each
RT286-11-015 H Accessories | Endodontic equipment | Wave Contra angle 6:1 ratio for the Corded Rotating and reciprocating Endodontic Motor and Handpiece Combination |  |  |  | Each
RT286-11-015 I Accessories | Endodontic equipment | Foot Pedal for the Corded Rotating and reciprocating Endodontic Motor and Handpiece Combination |  |  |  | Each
RT286-11-016 | Endodontic equipment | Endodontic down Pack and Backfill Root Canal Machine "SPECIFICATION An Endodontic down Pack and backfill root canal machine" Must have a Down pack capability and have the following characteristic:. Use of pluggers of various sizes Pluggers must be able for use of down packing and searing of GP Must be able to utilise pluggers to be click-on/click of type Must have a thermal response tip for pulp lity testing Must have a Backfill capability and have the following : Temperature and flow rate to be set individually Precise release of GP by micro motor Quick change of cartridges through drop-in/drop out design Have flexible and extra heat conductive silver alloy cannulas that is suitable for curved canals One cartridge must be able to fill up to 5 canals Initial delivery must include the following: 1 X Plugger small ISO 40 with .03 taper 1 X pack of 10 cartridges 20 G 1 X pack of 10 cartridges 23G SABS Plug : Must comply with both SANS 164-0 and SANS 164-1 |  |  |  | Each
RT286-11-016 A Accessories | Endodontic equipment | 1 X Plugger small ISO 40 with .03 taper for the Endodontic down Pack and Backfill Root Canal Machine |  |  |  | Each
RT286-11-016 B Accessories | Endodontic equipment | 1 X Plugger medium ISO 50 with .05 taper for the Endodontic down Pack and Backfill Root Canal Machine |  |  |  | Each
RT286-11-016 C Accessories | Endodontic equipment | 1 X Plugger large ISO 60 with .06 taper for the Endodontic down Pack and Backfill Root Canal Machine |  |  |  | Each
RT286-11-016 D Accessories | Endodontic equipment | 1 X pack of 10 cartridges 20 G for the Endodontic down Pack and Backfill Root Canal Machine |  |  |  | Each
RT286-11-016 E Accessories | Endodontic equipment | 1 X pack of 10 cartridges 23G for the Endodontic down Pack and Backfill Root Canal Machine |  |  |  | Each
RT286-11-016 F Accessories | Endodontic equipment | Replacement cartridge nuts for the Endodontic down Pack and Backfill Root Canal Machine |  |  |  | Each
RT286-11-016 G Accessories | Endodontic equipment | Thermanal respose tip for the Endodontic down Pack and Backfill Root Canal Machine |  |  |  | Each
RT286-11-016 H Accessories | Endodontic equipment | Cartridge bending tool for the Endodontic down Pack and Backfill Root Canal Machine |  |  |  | Each
RT286-11-016 I Accessories | Endodontic equipment | Hand Cleaning Brush for the Endodontic down Pack and Backfill Root Canal Machine |  |  |  | Each
RT286-11-016 J Accessories | Endodontic equipment | Heat shield for the Endodontic down Pack and Backfill Root Canal Machine |  |  |  | Each
RT286-11-016 K Accessories | Endodontic equipment | GP Cartridges 23G for the Endodontic down Pack and Backfill Root Canal Machine |  |  |  | Each
RT286-11-017 | Radiological Equipment | Portable Handheld X Ray Unit "SPECIFICATION A Portable handheld X-ray for use were conventional mounted X-rays can not be used." Must be able to work with film, digital sensors and photostimulable phosphor plated It would be preferable if a attachable sensor and build in computer with a digital screen to see the radiographs "'lives" on the machine can be supplied The X-ray unit must be wireless, handheld and completely mobile. It must include a backscatter radiation shield. Must be for all dental use e.g. diagnostic and endodontics. Must be able to allow at least 250 images on a single charge. Must be supplied with a rechargeable battery. Backup rechargeable battery Charging unit The unit has the ability to transfer images wireless to a PC. SABS Plug : Must comply with both SANS 164-0 and SANS 164-1 |  |  |  | Each
RT286-11-017 A Accessories | Radiological Equipment | Rechargeable battery for the Portable Handheld X Ray Unit |  |  |  | Each
RT286-11-017 B Accessories | Radiological Equipment | Charging unit for the Portable Handheld X Ray Unit |  |  |  | Each
RT286-11-018 | Dental surgery small equipment | Hypodermic Needle incinerator portable "SPECIFICATION A device that incinerate needles to prevent needle injuries that is portable" A portable hypodemic needle incinirator for dental needles. Must comply with all health and safety regulations. Temperature range from 1500 - 2000 degrees celcius Indicator light Needle gauge 25mm and 38mm Dimensions Height 95mm x Width 170mm x Length 300mm Charging unit Batteries SABS Plug : Must comply with both SANS 164-0 and SANS 164-1 |  |  |  | Each
RT286-11-019 | Dental surgery small equipment | Hypodermic Needle Incinerator Desktop "SPECIFICATION A device that incinerate needles to prevent needle injuries that is portable" A desktopn hypodermic needle incinirator for dental needles. Must comply with all health and safety regulations. Temperature range from 1500 - 2000 degrees celcius Indicator light Needle gauge 25mm and 38mm Dimensions Height 95mm x Width 170mm x Length 300mm SABS Plug : Must comply with both SANS 164-0 and SANS 164-1 |  |  |  | Each
RT286-11-020 | Dental surgery small equipment | Platelet Rich Plasma Centrifuge (PRP) "SPECIFICATION Unstainless frame" Microprocessor control Brushless motor driving Monitored by conductivity motor LCD Display Protection of unbalanced an door cover with function of alarming SABS Plug : Must comply with both SANS 164-0 and SANS 164- |  |  |  | Each
RT286-11-020 A Accessories | Dental surgery small equipment | PRP Tubes |  |  |  | Each
RT286-11-020 B Accessories | Dental surgery small equipment | Activator PRP tubes |  |  |  | Each
RT286-11-021 | Dental surgery small equipment | Intra Oral Wireless Scanner with built in Colour Determination "SPECIFICATION Intra oral wireless scanner for use in digital impressions with the following features:" Realistic colours, shade measurement, and an intraoral camera Pen grip Wireless Portable, lightweight pod and scanner that connects to your laptop via USB Battery Charger with at least 2 X Batteries Must be able to include the widest range of treatment options with verified CAD/CAM lab workflows nl. Crowns Bridges Veneers, Inlays & Onlays Removable Partial Dentures Posts and Cores Abutments, Implant Bridges and Bars Envisaging Must be able to include the at least the following add on software: Orthodontics Implant Studio Design Studio Open system (Optional) SABS Plug : Must comply with both SANS 164-0 and SANS 164-1 |  |  |  | Each
RT286-11-022 | Infection control | Air purifiers SPECIFICATION Four or five stage filtration system HEPA filtration system, UVC , activated carbon Multi humidity settings With adjustable Timer Filter replacement indicator Reduces viruses and bacteria Multi speed setting Impregnable seal around filters Energy efficient With sleep mode Filter lifespan more than six months Low noise levels Adjustable fan speed Air quality indicators |  |  |  | Each
RT286-11-023 | Infection control | Hydroclorous acid generator SPECIFICATION Disinfecting manufacturing unit producing hypochlorous acid and sodium hypochloride. Capacity – 5lt to 20lt Working Time – 5 minutes to 1 Hour PPM – 300 to 1500mg/L Producing a Chlorine concentration of more than 1000mg/L that is effective against a broad range of microorganisms. Effective sterilization rate of 99%. Does not produce carcinogenic, teratogenic, mutagenic and other chlorinated compounds. Product with shelf life reaching and exceeding 30 days. Should include minimum of a two-year warranty. |  |  |  | Each
RT286-11-024 | Infection control | Extra oral vacuum/ extractor SPECIFICATION To remove high volume of droplets and aerosols produced during treatment Filter efficiency - > 99% elimination of 0.3 micron size particles For use in closed spaces Sound - < 100db high speed < or = 60 db low speed Triple H14 HEPA filter system or quadruple activated carbon and H14 HEPA filter system Contains UVC lamp working with the HEPA filter to inactivate viruses and bacteria Adjustable power levels With Auto HEPA filter replacement reminder Lightweight – portable Movable, flexible suction arm Long working time of approximately 4 hours |  |  |  | Each
RT286-11-025 | Infection control | Germicidal Ultraviolet Lamp SPECIFICATION · Wall mounted or Ceiling mounted · Lamp emitting UV-C light that is germicidal · Covering surface area up to 50 sqm · Safety Timer · Designed for consultation rooms, passages, waiting rooms, etc |  |  |  | Each
RT286-11-026 | Dental surgery small equipment | Hydrocyclone Amalgam IS18 SEPARATOR SPECIFICATION PAL 14, 28, 38, 50 IS 18 VE with Electronic Amalgam Level Control Components i) Centrifuge ii) Hydrocyclone iii) Pressure relief valve iv) Control panel v) Output Power 0.25KW - 0.28 A vi) Tested by TUV, according to ISO 11143 Standard vii) W=470mm, D=400mm, H=590mm |  |  |  | Each
RT286-11-027 | Dental surgery small equipment | Hydrocyclone ISO 6 Amalgam separator SPECIFICATION I) Fixing plates ii) Re-circulation hoses iii) TUV Tested according to ISO11143 standard |  |  |  | Each
RT286-11-028 | Dental surgery small equipment | Hydrocyclone Amalgam separator SPECIFICATION Components i)Hydrocyclone ISO 18 ii) Fixing plates iii) Re-circulation hoses iv)Tested by TUV according to ISO 11143 Stardard |  |  |  | Each
RT286-11-029 | Dental surgery small equipment | Hydrocyclone ISO 4 Amalgam separator SPECIFICATION Version for Turbo-Jet 1 Modular without a box Hydrocyclone Outlet connector with pressure relief valve Tested by TUV according to ISO 1143 Standard SPECIFICATION Multi-surgery version to be installed inside dental unit Components i) Mini-separator ii) Mignon Electro-pneumatic valve iii)Draining unit iv) AC 45 PB v)Ø25 Straight tube holder vi)Ø30 Straight tube holder vi)Ø25 bent tube holder vii)Ø30 bent tube holder viii) Canister W=195mm, D=135mm, H=340mm ix) Electrovalve W=120mm, D=105, H=160mm SPECIFICATION Single-surgery version to be installed inside dental unit Components i) Mini-separator ii) Drainig unit iii)AC 45 PB v)Ø25 Straight tube holder vi)Ø30 Straight tube holder vi)Ø25 bent tube holder vii)Ø30 bent tube holder viii) Canister W=195mm, D=135mm, H=340mm |  |  |  | Each
RT286-11-030 |  | A Conscious Sedation Machine with Automatic Flow Control.The machine must pass the tests laid down by CE directives. Conscious sedation machine with flow control box efficacious safety devices that do not permit errors, hence safeguarding the safety of the patient. An O2 device, to cut off the flow of N2O, delivering 100% oxygen both to the control balloon and to the patient being treated. An O2 knob, to set the flow required by the patient in litres per minute, and a special regulation knob that enables the percentage of nitrous oxide delivered to be automatically varied from 0 to 70% Columns indicating fast and easy reading. Couplings fitted on the back of the flow control box and their attachments are physically different in order to safeguard against the risk of inverting the hoses for the two gases. Circuit with mask Adaptible circuit: special breather hose and adapter pipes that fit tge nitrous oxide machine Masks. Adaptable masks with foreheat rest and backflow valve neck elastic and backfow valve, neck elastic band and soft silicone face mask. Rubber balloon/reservoir bag. Bladder type bag, made of latex-fee material or silicone. 3 Litres capacity Nitrous oxide scavenger pipe to enable adaption to high volume suction. |  |  |  | Each
RT286-11-030 A Accessories | Dental equipment | Mask Small |  |  |  | Each
RT286-11-030 B Accessories | Dental equipment | Mask Large |  |  |  | Each
RT286-11-030 C Accessories | Dental equipment | Nitrous oxide scavenger pipe |  |  |  | Each
RT286-11-030 D Accessories | Dental equipment | Adaptible circuit: |  |  |  | Each
RT286-11-030 E Accessories | Dental equipment | Rubber balloon/reservoir bag |  |  |  | Each
RT286-11-031 | Dental equipment | 0-190 Inclinicable Binocular Tube with PD Adjusting Knob Zoom Magnification, Manual 4K Camera with recording system Lens 200-450mm 9# LED Illumination (Second LED Module included) 1000mm Second Arm Super Balancing arm Ceiling Mount LAN (Wifi) Connectivity: Mobile phone, Tablets, Windows Computers One Hand Operation High Quality German Optics Integrated Yellow and Green Filters Spot illumination Mount optional: Wall, Ceiling, Mobile 4k-32 Inc TV Monitor included Bracket for TV included |  |  |  | Each
RT286-11-032 | Dental equipment | Computer integrated probe Complete system for periodontal probing and charting Provides interactive exam Constant pressure- automated Talking –probes records all parameters of periodontal exam including pocket depths to the nearest 0,2mm Software and hardware and training included Voiceworks bleeding command shortcuts New AAP Periodontal diagnosis classification Live chat support from within the softwrare Add sticky note to any tooth Ability to mark deciduous teeth Can be linked to multiple office locations Wireless network compatibility Options for pocket depth comparison arrows
RT286-11-033 | Dental equipment | Dental operator chair-portable Easily transportable Robust construction Quick assembly and dismantling at least within 30sec High seating comfort Height adjustable Bearing weight of 5.65kg Detachable armrest,base and sit Additional accessories Protective carrier bag for ease of transportation Additional Accessories Backrest Armrest |  |  |  | Each

## Removed Items
RT286-10-011 | Dental surgery equipment | Suction Wet 2-4 Surgeries |  | Awarded already, remove
SPECIFICATION A fixed installation centraised wet aspiration plant for serving 2-4 surgeries. The unit MUST have the following features:
Fixed installation
Serve number of surgeries: 2 - 4
Output power: Single-phase, Approximately 1.7kW - 9A
Maximum flow: No less than 820 l/min
Operating head for continuous service: Maximum 3300mm H2O
Average water consumption: Approximately 6 l/min
Piping pressure: Approximately 4 bar
Dimensions: Approximately 370mm W, 300mm D, 420mm H
RT286-10-012 | Dental surgery equipment | Suction Dry 4-6 Surgeries |  | Awarded already, remove
SPECIFICATION A fixed installation centraised dry aspiration plant for serving 6 surgeries. The unit MUST have the following features:
Fixed installation
Serve number of surgeries: 4 - 6
Output power: Three-phase, Approximately 2kW - 3.5A, 50 Hz
Speed of motor: No less than 2800 rpm
Vacuum setting at approximately 160 mBar
Noise level without housing: No more than 65dB(A)
RT286-10-014 | Dental surgery equipment | Suction Wet 1-2 Surgeries |  | Awarded already, remove
SPECIFICATION A fixed installation centraised wet aspiration plant for serving 1-2 surgeries. The unit MUST have the following features:
Fixed installation
Serve number of surgeries: 1 - 2
Output power: Single-phase, Approximately 0.9kW - 5A
Maximum flow: No less than 500 l/min
Operating head for continuous service: Maximum 2700mm H2O
Average water consumption: Approximately 2 l/min
Piping pressure: Approximately 4 bar
Dimensions: Approximately 550mm W, 350mm D, 360mm H
RT286-10-015 | Dental surgery equipment | Suction Dry 1 Surgery |  | Awarded already, remove
SPECIFICATION A fixed installation centraised dry aspiration plant for serving 1 surgery. The unit MUST have the following features:
Fixed installation
Serve number of surgeries: 1
Output power: Single-phase, Approximately 0.4kW - 2A, 50 Hz
Speed of motor: No less than 2800 rpm
Vacuum setting at approximately 170 mBar
Noise level without housing: No more than 65dB(A)
RT286-10-016 | Dental surgery equipment | Suction Dry Mobile 1 Surgery |  | Awarded already, remove
SPECIFICATION A mobile suction unit for continuous utilisation in a dental surgery. The unit MUST have the following features:
Mobile unit to be steered around chair and between surgeries
Equipped with 4 double wheel castor, one per corner
Designed for continuous service
Capacity of canister: minimum 4 Litres
Canister to be emptied manually
Noise level: No more than 65dB(A)
Dimensions: Approximately 550mm W, 350mm D, 750mm H
Net weight: Approximately 35kg
Gross weight: Approximately 40kg
RT286-10-036 | Radiological Equipment | Automatic Intra Oral X-Ray developer
Remove the item |  | SPECIFICATION An Automatic intra oral x-ray developer | spelling | to be removed , Duplicate on RT286-10-001
Processing time approximately 2 min(Wet endodontic) and 4 min(Dry).
Must have a drier
Daylight loader
SABS Plug : Must comply with both SANS 164-0 and SANS 164-1
RT286-11-001 | Dental surgery small equipment | Digital Amalgamator |  | Awarded already, remove
Remove awarded already |  | SPECIFICATION Noise level less than 65BD
Automatic swich off-electronic time control
Operation cycle up to 99S
Body Shock resistant plastoc body
The unit should operate on 240VAC 50/60Hz fitted with 3-prongs plug
Unit to have a sealed and isolated mixing chamber which deactivates on opeing
Replaceable practical and durable capsule holder (both amalgam and glass ionomer capsule)
SABS Plug : Must comply with both SANS 164-0 and SANS 164-1
RT286-10-073 | Dental equipment | Portabe Dental chair
|  | SPECIFICATION A Portable chair to be abled to be carried by an operator |  | Remove awarded already
The unit must be able to carry up to 200kg
It must have medium alloy construction
The seat must adjustable from 35 to 60cm
The unit must be able to be used as a sit up chair
The maximum weight must not be more than 12.5kg
The dimensions must not be more than 510 x 270 x1000mm
It must come complete with a back pack
The unit must make provision for a portable dental light, spit funnel and tray
RT286-10-018 | Dental surgery equipment | Compressor 1 Surgery - Oil free |  | Awarded already, remove
SPECIFICATION A fixed installation dental compressor with dryer (oil-free) for serving 1 surgery. The unit shall have the following features:
Serve number of surgeries: At least 1
Number of cylinders: Minimum 1
Number of dryers: At least 1
Output power: Approximately 1-50Hz, ± 0.5kW, ±3.5A
Air delivery: At least 66N l/min (pressure no less than 5 real bar)
Volume of air tank: No less than 30 Litres
Noise level: No more than 65dB(A)
Dimensions: Approximately 650mm W, 450mm D, 730mm H
Net weight: Approximately 45kg
Gross weight: Approximately 55kg
RT286-10-019 | Dental surgery equipment | Oil Free air compressor for 4-6 surgery (60DB) |  | Awarded already, remove
SPECIFICATION A fixed installation dental compressor with dryer (oil-free) for serving4 -6 surgeries. The unit shall have the following features:
Voltage 220/240 V 1,5 kW with built in thermo overload protection as per user requirement
Delivery capacity not less than 150 1/min maximum pressure 1000 kpa
Storage Capacity at least 35 litres - 75 litres
Fitted with ¼” safety release valve
Manual tank drain valve
Tank pressure gauge
Pressure range control switch, 6 – 10 bar
Noise level not more than 60 DB (A)
Either piston type or turbo type
Reverse pressure exhaust valve for motor protection
Air filter for filtration of 10 micron particles, with automatic drain valve
Oil free with automatic drianage system
RT286-10-043 | Radiological Equipment | Digital Radiography intra oral system-sensor without a computer | name amendment | Awarded already
|  | SPECIFICATION A Sensor for immediate capturing of digital X-ray intraoral images | spelling
Sensor with USB 2 connection
Mount for sensor storage
|  | Cable length 3 meters / cordless (standard is app 2m) | market available is 2mm
Hygienic sheaths X50
Twain Compatibility - Sensor can be integrated with any imaging and dental practice management software
Sensor positioning kit
The unit MUST use a line voltage of 220VAC to 240VAC
The line frequency MUST be 50Hz
Software including implant module and implant archive. Must include DICOM3 compatibility, including DICOM worklist, DICOM PACS, DICOM Store, DICOM transfer, DISOM query/retrieve, DICOM print, DICOM MPPS and DICOM RDSR
RT286-10-044 | Radiological Equipment | Digital Radiography intra oral system-sensor with a computer |  | Awarded already
|  | SPECIFICATION A Sensor for immediate capturing of digital X-ray intraoral images | spelling
Sensor with USB 2 connection
Mount for sensor storage
Cable length 3 meters / cordless (standard app 2m)
Hygienic sheaths X50
Twain Compatibility - Sensor can be integrated with any imaging and dental practice management software
Sensor positioning kit
The unit must be provided with an up to date desktop/laptop computer with the following requirements:
≥4GB Ram
≥1TB Hard drive
|  | 17 inch LCD screen with 1920x 1080 pixels for diagnosing and review (as per SAHPRA regulations) | as per SAHPRA guidelines
Software for processing all captured images. Must include DICOM3 compatibility, including DICOM worklist, DICOM PACS, DICOM Store, DICOM transfer, DISOM query/retrieve, DICOM print, DICOM MPPS and DICOM RDSR
A keyboard
The unit MUST use a line voltage of 220VAC to 240VAC
The line frequency MUST be 50Hz
The unit must be free standing with height adjustment
|  | Software including implant module and implant archive. - not applicable must include DICOM3 compatibility, including DICOM worklist, DICOM PACS, DICOM Store, DICOM transfer, DISOM query/retrieve, DICOM print, DICOM MPPS and DICOM RDSR | Panoramic machines do not normally have implant designing software (2d)
SABS Plug : Must comply with both SANS 164-0 and SANS 164-1
RT286-10-045 | Radiological Equipment | Intra Oral Dental X-ray (Analogue & Digital) - Mobile ????? |  | Awarded already
|  | SPECIFICATION A X-ray unit for capturing intra oral analgue and digital images (WALL Mouted) | you can not have a mobile unit that is wall mounted
The unit must be a mobile type (with an option of a long arm)
The tube head MUST be between 60kV to 70kV
The unit must have a computerized timer with the following functions:
Automatic setting of technique factors in mAs
Object program selection
Tooth type
Patient size
Manual override
Hand switch with 3m cord or cordless
The focal point of the unit MUST be from 0.4mm to 0.8mm
The power supply of the unit must be mounted on the unit
The unit MUST use a line voltage of 220VAC to 240VAC
The line frequency MUST be 50Hz
Analogue and digital
SABS Plug : Must comply with both SANS 164-0 and SANS 164-1 (2 fase)
RT286-10-046 | Radiological Equipment | Intra Oral Dental X-Ray (Analogue & Digital) Floor Mounted |  | Awarded already
SPECIFICATION A X-ray unit for capturing intra oral analgue and digital images (Floor Mouted)
The unit must be a floor mount type (with an option of a long arm)
The tube head MUST be between 60kV to 70kV
The unit must have a computerized timer with the following functions:
Automatic setting of technique factors in mAs
Object program selection
Tooth type
Patient size
Manual override
Hand switch with 3m cord or cordless
The focal point of the unit MUST be from 0.4mm to 0.8mm
The power supply of the unit must be mounted on the unit
The unit MUST use a line voltage of 220VAC to 240VAC
The line frequency MUST be 50Hz
Analogue and digital
SABS Plug : Must comply with both SANS 164-0 and SANS 164-1 (2 Fase)
RT286-10-032 | Infection control | Dental Autoclave B-Class
Awarded |  | SPECIFICATION A Table top B class autoclave without a PRION Cycle
Mains voltage MUST operate at 220 – 240V.
Current MUST not exceed 15 Amps
MUST have no re-circulation of used water.
Chamber MUST be constructed of stainless steel or copper brass alloy. MUST be ASME or SABS Plug : Must comply with both SANS 164-0 and SANS 164-1 approved.
The unit must have ability to heat the entire chamber efficiently.
MUST have heat insulated casing with built-in condensation coils and tank.
MUST have automatic timer to complete sterilizing cycle.
MUST be fitted with safety pressure valve.
MUST be fitted with temperature and pressure gauges.
Complete cycle after warm up MUST not exceed 30 min.
Chamber dimensions MUST be ±510 mm deep by ±360mm diameter.
Operating pressure MUST be 240 Kpa.
Operating temperature MUST not exceed – 135°C.
Minimum reservoir capacity 4 litres.
16 Amp plug MUST be supplied.
MUST have separate drying cycles which shuts off automatically – 7 – 12 minutes.
The unit MUST be of a horizontal, chamber-type autoclave, and the entire unit MUST be microprocessor controlled.
The unit MUST monitor the water level, and if low, switch off to protect the elements from overheating and damage. The unit MUST switch off when water level low, and not when overheating.
The unit MUST include at least 3 perforated trays with a detacheable handle.
Each tray MUST hold at least 10 dental instruments. The chamber MUST be long enough to sterilize instruments up to 300mm long
All cycles MUST be automatic – no user intervention MUST be required, except to select the correct cycle and start the cycle
End of cycle, and all error messages, MUST be indicated by a clear audible signal
It MUST be possible to terminate any cycle at any stage.
The unit MUST display the operating temperature in an electronic format, and also display pressure.
The unit MUST be automatically controlled by temperature. Control for pressure as an additional measure will serve as a recommendation.
The unit MUST automatically protect itself and the operator against over temperature and overpressure
The unit MUST be able to perform the following cycle:
134º with vacuum with drying
134º with vacuum without drying
134º non-vacuum with drying
134º non-vacuum without drying
121º with vacuum with drying
121º with vacuum without drying
121º non-vacuum with drying
121º non-vacuum without drying
Please provide cycle times for each of the programs for hot and cold cycle
The units MUST be able to sterilize unwrapped solids and hollow instruments, pouched solid and hollow instruments, single wrapped solid and hollow instruments, double wrapped solid and hollow instruments and porous items
The unit must also store data of at least the last 40 cycles. This data must be available in a printable and downloadable format. The unit MUST be able to connect to a computer, via a USB connection, to download the required information. If any software program is needed to perform this, it must be included in the bid price as an accessory.
Capacity of the chamber MUST be 16+ litres.
Water MUST be drained from the chamber after each and every cycle
External elements must be used, please state the wattage of the elements
The unit MUST have an insulated shell to protect operator from accidental heat injuries
Overpressure valve and all other safety valves must vent into a safe area, and not into working area
Unit must monitor the water quality in chamber, and stop cycle if the quality is not satisfactory
Reservoir and chamber to have a manual drain facility that is controlled with a positive on/off valve
Operator controls MUST be by touch-type, and not switches for infection control purposes
All electrical boards/electronic components MUST be well shielded/insulated from heat and potential water spillage
The door mechanism MUST have an auto-lock safety facility to prevent the door being opened during a cycle
It MUST not be possible to start a cycle if the door is not properly closed and sealed
Water distiller/demineraliser to be included for each autoclave. This must meet the required water quality for this autoclave. The price of this unit must be included in the bid price
The autoclave is to display time lapsed and the time remaining on cycle
All error codes to be supplied in digital format
The entire unit must form a self-contained working unit. It must be easy to move the unit to where it’s required
Unit must be supplied with a spare door seal
SABS Plug : Must comply with both SANS 164-0 and SANS 164-1
RT286-10-007 | Infection control | Water Distillation Unit |  | To remove this item as it is covered under the main contract
SPECIFICATION A desktop water distillation unit for dental autoclaves. The unit MUST have the following features:
Must be able to supply at least 4 Liters of distilled water to different types of autoclaves
SABS Plug : Must comply with both SANS 164-0 and SANS 164-1