## Sheet1
ANNEXURE C : PRICING SCHEDULE
BID RT286-2021
BID DESCRIPTON : SUPPLY, DELIVERY, INSTALLATION, COMMISSIONING AND MAINTENANCE OF DENTAL EQUIPMENT
TO THE STATE FOR THE PERIOD 1 OCTOBER 2021 TO 30 SEPTEMBER 2024
ITEM CODE | ITEM NUMBER | ITEM DESCRIPTION | CATEGORY NAME | BIDDER NAME | UOM | RESPONSE PRICE IN ZAR CURRENCY INCLUSIVE OF 15% VAT | BRAND NAME | COUNTRANTRY OF ORIGIN | FOREIGN EXCHANGE RATE | IMPORT COMPONENT OF PRICE AS %
42152214-00000 | RT286-01-001 | Plaster Dispenser - Electric | Work Preparation & Model Production |  | Each
SPECIFICATION An electrically operated silo for the dry storge and controlled dispensing of plasters and stone in the dental laboratory. The unit shall have the following features:
Suspended on the wall or bench mounted in a frame
Capacity of the dispenser no less that 25kg
Construction of the dispenser:
Outer housing made of stainless steel or plastic
Inner hopper of plastic for corrosion-free plaster / stone supply
Built-in electrical vibrator
Dispensing lever
Vibration action ensuring even dispensing of material
Activating dispensing lever setting vibrator into motion
SABS plug
42152212-00000 | RT286-01-002 | Model Trimmer - Single Disc / Blade |  |  | Each
SPECIFICATION A single disc/ blade model trimmer utilised in a dental laboratory for the trimming of plaster, stone and super hard stone models. The unit shall have the following features:
Single motor encased in water-proof casing
Drive motor speed minimum 1800rpm
ON/OFF power switch & speed selector on casing
Trimmer discs features: Reversible disc, Course, dia 10" or 12"
Tap for adjusting water flow to the disc
Water jets for cleaning disc during operation
Discs covered with water splash guard, leaving grinding surface app 100mm x 140mm
Unit mounted on anti-vibration rubber stoppers
Splash shield
Water spray attachment
Orthodontic work table including Vertisquare, Angle guide and Template
SABS plug
Optional items
Replacement disc Reversible, Course, dia 10" or 12"
Replacement Vertisquare
42152212-00001 | RT286-01-003 | Model Trimmer - Dual Disc / Blade |  |  | Each
SPECIFICATION A dual disc/ blade model trimmer utilised in a dental laboratory for the trimming of plaster, stone and super hard stone models as well as ortho stone orthodontic study models. The unit shall have the following features:
Motor is Split Phase, Thermally Protected, Moisture Resistant
Trimmer discs features: Reversible, One Coarse for trimming, One Fine for Finishing, diameter 12"
One tap for adjusting water flow to the disc
Water jets for cleaning disc during operation
Disc covered with water splash guard
Each grinding table with one guide groove in line with each grinding disc
Squaring tool and an angulation tool to be supplied with each unit to fit into grooves
Unit mounted on anti-vibration rubber stoppers
Accessory Splash guard dual
SABS plug
Optional items
A | RT286-01-003 A | Replacement disc Coarse for trimming diameter 12" |  |  | Each
B | RT286-01-003 B | Replacement disc Fine for Finishing, diameter 12" |  |  | Each
C | RT286-01-003 C | Replacement Vertisquare |  |  | Each
42152212-00002 | RT286-01-004 | Model Trimmer - Dry System with Partially Coated Disc / Blade |  |  | Each
SPECIFICATION A single disc, dry system, plaster model trimmer for the trimming of models in a dental laboratory. The unit shall have the following features:
Powerful motor no less than 1600W
Quiet running electric motor, not more than 75 dB(A)
A partially diamond coated disc
Electronic instant-stop for trimmer disc
Tool-free quick release lock on removable casing door
Safety switch to interrupt power on opening casing door
Fitted with connection for external dust / particle extractor unit
SABS plug
Optional items
| RT286-01-004 A | Replacement partially diamond coated disc |  |  | Each
42151617-00011 | RT286-01-005 | Pin Hole Drilling Unit |  |  | Each
SPECIFICATION A benchtop, manually controlled unit utilised in the crown and bridge technique for drilling holes in the base of crown and bridge models, prior to the placement of die pins. The unit shall have the following features:
Sturdy, stable base
Slip-free model table
Model table designed to prevent dust accumulation
Pin drill mounted below model table pointing upwards
Overhead light / laser aiming device
Aiming device exact match to pin drill
Drilling activated bmo pressing model table down/up
Changing of drill bit quick and preferably tool-free
Adjustable drill depth: 0-12mm
Engine / Drive shaft protected from plaster dust
SABS plug
Optional items
| RT286-01-005 A | Replacement pin drill |  |  | Each
| RT286-01-005 B | Replacement LED Light |  |  | Each
42151616-00004 | RT286-01-006 | Surveyor for Models - Single Arm |  |  | Each
SPECIFICATION A benchtop, manually controlled unit utilised in the cobalt-chrome technique for model analysis, measuring undercuts and marking the equator. The unit shall have the following features:
Base of unit: approx. 120mm x 90mm
One double jointed arms
One accessories holders
Flat based, moveable cast holder
Cast holder features:
Free-moving
Flat base
Adjustable ball joint between base & holder
One fixed model grips
One adjustable model grip
Shank diameter of accessories: 3mm
Unit supplied with complete accessory set:
Lead marker holder
Lead
Wax knife
Bracing holder
Analysing pin
Undercut measuring pin: 0,25mm
Undercut measuring pin: 0,50mm
Undercut measuring pin: 0,75mm
Optional items
| RT286-01-006 A | Lead marker holder |  |  | Each
| RT286-01-006 B | Lead |  |  | Each
| RT286-01-006 C | Wax knife / scraper |  |  | Each
| RT286-01-006 D | Bracing holder |  |  | Each
| RT286-01-006 E | Analysing pin |  |  | Each
| RT286-01-006 F | Undercut measuring pin: 0,25mm |  |  | Each
| RT286-01-006 G | Undercut measuring pin: 0,50mm |  |  | Each
| RT286-01-006 H | Undercut measuring pin: 0,75mm |  |  | Each
42151616-00005 | RT286-01-007 | Surveyor for Models - Double Arm |  |  | Each
SPECIFICATION A benchtop, manually controlled unit utilised in the cobalt-chrome technique for model analysis, measuring undercuts and marking the equator. The unit shall have the following features:
Base of unit: approx. 120mm x 90mm
Two double jointed arms
Two accessories holders
Flat based, moveable cast holder
Cast holder features:
Free-moving
Flat base
Adjustable ball joint between base & holder
One fixed model grips
One adjustable model grip
Shank diameter of accessories: 3mm
Unit supplied with complete accessory set:
Lead marker holder
Lead
Wax knife
Bracing holder
Analysing pin
Undercut measuring pin: 0,25mm
Undercut measuring pin: 0,50mm
Undercut measuring pin: 0,75mm
Optional items
| RT286-01-007 A | Lead marker holder |  |  | Each
| RT286-01-007 B | Lead |  |  | Each
| RT286-01-007 C | Wax knife / scraper |  |  | Each
| RT286-01-007 D | Bracing holder |  |  | Each
| RT286-01-007 E | Analysing pin |  |  | Each
| RT286-01-007 F | Undercut measuring pin: 0,25mm |  |  | Each
| RT286-01-007 G | Undercut measuring pin: 0,50mm |  |  | Each
| RT286-01-007 H | Undercut measuring pin: 0,75mm |  |  | Each
42152205-00000 | RT286-01-008 | Boiling-Out Unit 6 or 12 (individual pricing) flask capacity |  |  | Each
SPECIFICATION A castor mounted, free standing unit utilised in a dental laboratory for automatically boiling-out wax up, scalding and polymerisation during production in a dental laboratory. The unit shall have the following features:
Primary application: Boiling out of denture flasks
Additional functions:
Automatic spraying bmo fixed spray bar jets
Manual spraying bmo hand shower on hose
Scalding (Steam cleaning) of flasks
Polymerisation of acrylic in denture flasks
Polymerisation at 95°C
Permanent connection to fixed water supply (optional)
Automatic water level control & filling
Housing made of stainless steel
Switches & dials mounted on top-rear of housing
Minimum features of control panel:
ON / OFF switch
Temperature control dial
Timer for a “wax-out” program
Timer for boiling-out program
ON / OFF switch for hand spray/s
Regulating valve for spraying
A 24 hour timer
Water level warning light
At least one integrated, adjustable hand spray
Hand spray adjustable between jet to gentle spray
Tank made of stainless steel
Tank covered with perforated stainless steel lids
Outlet of tank in base
Access to drain from the front of the unit
Pump features:
Corrosion resistant pump: Seal-free with 5 year guarantee
Electric pump motor positioned above water level
Changeable filter before pump chamber
Flask baskets:
6 or 12 supplied per unit
Stainless steel construction
Fitted with insulated plastic / rubber handles
SABS plug
Optional items
| RT286-01-008 A | Descaling liquid |  |  | Each
| RT286-01-008 B | Hand shower spray head |  |  | Each
| RT286-01-008 C | Replacement flask baskets |  |  | Each
42151696-00000 | RT286-01-009 | Blow-Off Nozzle - Rubber |  |  | Each
SPECIFICATION A rubber blow-off nozzle for compressed air utilised during production in a dental laboratory. The unit shall have the following features:
Rubber exterior
Serrated metal attachment nipple at rear
Air activation bmo sideways pressure on front section of nozzle
Length: Approximately 120mm
Diameter: No more than 30mm
Supplied with ring clamp for attachment to air hose
Optional items
| RT286-01-009 A | Replacement ring clamp |  |  | Each
42151616-00006 | RT286-02-001 | Duplicating Unit - Reversible Duplicating Material | Duplicating & Hardening |  | Each
SPECIFICATION A benchtop, free standing hydraulic flask press utilised in the acrylic section of a dental laboratory during flasking procedures. The unit shall have the following features:
Capacity: 1 to 3 flasks
Minimum vertical clearance at least 220mm
Minimum clearance between uprights 128mm
Press plate diameter no less than120mm
Press piston vertical stroke approximately 30mm
Press piston to return automatically to base after venting
Separate pump and press cylinders
Pressure gauge to measure applied pressure
Hydraulic pressure applied bmo horisontally operating lever
Minimum pressure exerted is at 200 Atm (8000Kg)
Venting and locking by means of separate levers
Mechanical clamping of flasks by means of a vertical adjustable spindle with horizontal handle on top
42151616-00007 | RT286-02-002 | Duplicating Unit - Addition Cured Two-Component Silicone |  |  | Each
SPECIFICATION A benchtop, electrically controlled, silicone mixing and metering unit, utilised for duplicating plaster and plastic models in a dental laboratory. The unit shall have the following features:
Unit to consist of:
Precision mixing head
Two rapid action couplings
Mixing nozzles x 100
Socket spanner
Metering test cards x 10
Mechanical mixing
Deliver a mixing ratio 1:1
SABS plug
Optional items
| RT286-02-002 A | Pack of mixing nozzles x 100 |  |  | Each
41104510-00000 | RT286-02-003 | Model Drying Cabinet |  |  | Each
SPECIFICATION A benchtop, electrically heated drying cabinet for heating up and drying duplicate models, utilised during production in a dental laboratory. The unit shall have the following features:
Operating temperature must reach a minimum of 200°C
Pre-selection of required temperature
Interior lighting(Optional)
Accommodate at least 12 models at a time
At least two removable shelves(Optional)
Hinged front lid width of cabinet interior(Optional)
Full width window in hinged lid (Optional)
Window of heat resistant glass
Unit housing of corrosion resistant steel
SABS plug
Optional items
| RT286-02-003 A | Removable shelf |  |  | Each
| RT286-02-003 B | Replacement globe |  |  | Each
| RT286-02-003 C | Replacement door seal |  |  | Each
42152224-00002 | RT286-02-004 | Flask Press - Manual: 1 - 3 Flasks |  |  | Each
SPECIFICATION A benchtop, free standing manually operated flask press utilised in the acrylic section of a dental laboratory during flasking procedures. The unit shall have the following features:
Capacity: 1 to 3 flasks
Minimum vertical clearance at least 150mm
Minimum clearance between uprights 130mm
Pressing plate diameter no less than 120mm
Clamping of flasks and application of pressure by means of a vertical spindle with horizontal handle on top
Base of Flask Press to be supplied with holes for mounting to working surface
42152224-00003 | RT286-02-005 | Flask Press - Hydraulic: 1 - 3 Flasks |  |  | Each
SPECIFICATION A benchtop, free standing hydraulic flask press utilised in the acrylic section of a dental laboratory during flasking procedures. The unit shall have the following features:
Capacity: 1 to 3 flasks
Minimum vertical clearance at least 220mm
Minimum clearance between uprights 128mm
Press plate diameter no less than120mm
Press piston vertical stroke approximately 30mm
Press piston to return automatically to base after venting
Separate pump and press cylinders
Pressure gauge to measure applied pressure
Hydraulic pressure applied bmo horisontally operating lever
Minimum pressure exerted is at 200 Atm (8000Kg)
Venting and locking by means of separate levers
Mechanical clamping of flasks by means of a vertical adjustable spindle with horizontal handle on top
42152102-00022 | RT286-02-006 | Thermal Moulding Device: Positive Pressure Technique |  |  | Each
SPECIFICATION A benchtop, electronically controlled machine used in the thermal moulding production of appliances for all types of dental applications. The unit shall have the following features:
Operating pressure: 2.5 Bar
Pressure adjustable depending on type of material being utilised
Utilise compressed air for pressure moulding technique
Handle all types of thermo-forming materials ranging from Shellac to mouth guard materials
Accommodate (mould) material thickness from 0.1 to 5mm
Material diameter must be 125mm
Unit suitable for manufacturing following items:
Mouth guards
Copings
Special Impression Trays
Fluoride Trays
Retainers
Positioners
TMJ Appliances
Temporaries
Model Duplication
Have an electrical or infrared heating element
Heating to be uniform and consistent
Overheating protection by automatic switch off of the heater by the end of the heating time
All functions electronically controlled, checked and displayed on a digital display panel
Following programmable functions:
Heating time
Heating temperature
Holding temperature
Cooling time
SABS plug
Optional items
| RT286-02-006 A | Replacement set of O-rings |  |  | Each
| RT286-02-006 B | Replacement locking plate for material |  |  | Each
| RT286-02-006 C | Replacement pellets for embedding models in moulding device |  |  | Each
42152107-00000 | RT286-03-001 | Electronic Wax Modelling Instrument | Modelling |  | Each
SPECIFICATION A benchtop, electronically controlled modelling instrument for carrying out waxing-up techniques for work ranging from crown and bridge to partial dentures. The unit shall have the following features:
Control Unit:
Temperature display by means of an LED panel
Temperature setting by means of a sealed touch pad
Casing angled at 20°
Connector for handpiece on front of casing
Holder for handpiece & tips on rear of casing
Holder to accommodate 6 modelling tips
Temperature setting in Celsius
Programmable temperature from 50°C to 200°C
Handpiece:
Heat insulated
Sealed against liquid wax
Natural cork, non-irritant instrument handle
Voltage supplied to handpiece: 12 V DC
Cable length: no shorter than 1,2 m
Unit to be supplied with:
Control unit with integrated handle & tip holder
Heating handle with cable
Transformer
Instruction manual with spare parts list
Set of at least 7 modelling tips
SABS plug
Optional items
| RT286-03-001 A | Replacement modelling tips |  |  | Each
| RT286-03-002 B | Replacement Heating handle with cable |  |  | Each
42152107-00001 | RT286-03-002 | Electronic Wax Heater |  |  | Each
SPECIFICATION A benchtop, 3-pot electronic wax heater utilised in dental laboratory. The unit shall have the following features:
Unit with integrated:
Minimum 3 Wax pots with lids,
Heating elements,
Control panel,
Temperature display, and
Hinged cover
Separate temperature control on each pot
Heating element connect direct to each pot
Temperature range between 60°C and 110°C
LED bar graph pot temperature display
Sealed touch control pad
SABS plug
42152203-00014 | RT286-03-003 | Bunsen Burner Lab Dual Pipe – Propane Gas |  |  | Each
SPECIFICATION A free standing, tabletop, dual pipe, Bunsen burner for PROPANE gas. The unit shall have the following features:
Gas supply setting via rocking lever
Individual control of oxygen supply via setting wheel
Control of flame intensity by means of 2 ascending pipes:
Small pipe: Pilot flame
Large pipe: Main flame
Sturdy metal base for secure stand
Burner removable from base
Burner tilt of 30° to either side
Optional items
| RT286-03-003 A | Replacement ring clamp |  |  | Each
42152203-00015 | RT286-03-004 | Bunsen Burner Lab Dual Pipe – Natural Gas |  |  | Each
SPECIFICATION A free standing, tabletop, dual pipe, Bunsen burner for NATURAL gas. The unit shall have the following features:
Gas supply setting via rocking lever
Individual control of oxygen supply via setting wheel
Control of flame intensity by means of 2 ascending pipes:
Small pipe: Pilot flame
Large pipe: Main flame
Sturdy metal base for secure stand
Burner removable from base
Burner tilt of 30° to either side
Optional items
| RT286-03-004 A | Replacement pressure hose for natural gas |  |  | Each
| RT286-03-004 B | Replacement Ring Clamps |  |  | Each
42152203-00016 | RT286-03-005 | Burner Outfit, Gas: Fully Automatic Ignition & Cut-Off |  |  | Each
SPECIFICATION A benchtop, electronically controlled fully automatic dental laboratory gas burner for utilisation in techniques requiring heating of hand held instruments. The unit shall have the following features:
Connection to normal gas supply
Connection to electrical network
Burner operating on an infrared optical electronic principle
Flame activated by hand movement over infrared sensor
Infrared sensor not influenced by external light
Flame extinguishes automatically when burner is not in use
Selector switch for automatic control or continuous burning
Adjustable air/gas mixture
Adjustable flame size
Reduction of gas consumption by approximately 80%
SABS plug
Optional items
| RT286-03-005 A | Replacement pressure hose for gas |  |  | Each
| RT286-03-005 B | Replacement Ring Clamps |  |  | Each
42152203-00017 | RT286-03-006 | Micro Torch – Butane Gas |  |  | Each
SPECIFICATION A refillable hand held butane torch utilised during dental laboratory work. The unit shall have the following features:
Electronic trigger for flame ignition
Flame size adjustable eans of adjusting handle/lever
Air adjustable by means of adjusting handle/lever
Safety Lock Stopper
Integrated Butane Gas Reservoir
Flame temperature no less than 1300°C
vertical or horisontal operating position
Operating time per refill between 30 –35 minutes
Optional items
| RT286-03-006 A | Bottle Butane gas refill |  |  | Each
42152203-00018 | RT286-03-007 | Induction Heater Portable for Hand Instruments |  |  | Each
SPECIFICATION An electrical heater for dental hand instruments, in the form of a compact induction heating system replacing gas flames. The unit shall have the following features:
Unit must be electrically operated - not gas
Main supply : 230V
Electromagnetic induction heating - no gas flame required
No heat emmision into the environment
External sensors should emit heating only when there is no instrument in the chamber
The modeling instrument must be heated inless than 3 seconds.
SABS plug
Optional items
| RT286-03-007 A | Replacement silicone cup to protect the chamber |  |  | Each
42152203-00019 | RT286-03-008 | Soldering and Melting Torch: Microflame – Propane Gas |  |  | Each
SPECIFICATION A hand held soldering and melting torch utilised for soldering and melting of all types of alloys during dental laboratory work, utilising PROPANE (LP-Liquid Petroleum) gas / oxygen mix. The unit shall have the following features:
Operate with propane gas (liquid petroleum) and oxygen
Accessories for installation shall consist of:
Shut-off valve for propane
Anti-explosion device for propane gas at pressures between 20millibar and 1.5bar
Pressure reducing valve for oxygen
Anti-explosion device for oxygen
Soldering and melting unit shall consist of:
Handle
2 m pressure hose for oxygen
2 m pressure hose for propane gas
Soldering insert, complete, with needle jets dia: 0.9, 1.2, 1.6 mm
Melting insert for propane /oxygen
Melting insert nozzle
Optional items
| RT286-03-008 A | Replacement needle jets diameter: 0.9, 1.2, 1.6 mm |  |  | Each
| RT286-03-008 B | Replacement pressure hose for oxygen |  |  | Each
| RT286-03-008 C | Replacement pressure hose for propane gas |  |  | Each
| RT286-03-008 D | Replacement Ring Clamps |  |  | Each
42152203-00020 | RT286-03-009 | Soldering and Melting Torch: Microflame – Natural Gas |  |  | Each
SPECIFICATION A hand held soldering and melting torch utilised for soldering and melting of all types of alloys during dental laboratory work, utilising NATURAL (Acetelene) gas / oxygen mix. The unit shall have the following features:
Operate with natural gas (acetelene) and oxygen
Accessories for installation shall consist of:
Shut-off valve for natural gas
Anti-explosion device for natural gas at pressures between 20millibar and 1.5 bar
Pressure reducing valve for oxygen
Anti-explosion device for oxygen
Soldering and melting unit shall consist of:
Handle
2 m pressure hose for oxygen
2 m pressure hose for propane gas
Soldering insert, complete, with needle jets dia: 0.9, 1.2, 1.6 mm
Melting insert for natural gas /oxygen
Melting insert nozzle
Optional items
| RT286-03-009 A | Replacement needle jets diameter: 0.9, 1.2, 1.6 mm |  |  | Each
| RT286-03-009 B | Replacement pressure hose for oxygen |  |  | Each
| RT286-03-009 C | Replacement pressure hose for propane gas |  |  | Each
| RT286-03-009 D | Replacement Ring Clamps |  |  | Each
42152217-00000 | RT286-04-001 | Vacuum Mixer: Plaster & Investment Material | Investing & Polymerisation |  | Each
SPECIFICATION A wall mounted or benchtop mechanical vacuum mixer for the mixing of all dental plasters and investment compounds utilised in a dental laboratory. The unit shall have the following features:
Single speed vertical drive shaft, approx. 450rpm’s
Positive drive chuck to activate rotation of mixer
Rust resistant motor
Integrated vacuum pump, oiler unit and vacuum hose with coupling to mixing bowl provided
ON/OFF switch on the unit within easy reach of operator
Motor drive and vacuum activated by insertion of mixing bowl
Two mixing bowls provided per unit:
approximately 550ml
approximately 800ml
SABS plug
Optional items
| RT286-04-001 A | Mixing bowls approximately 200ml - 1200ml |  |  | Each
| RT286-04-001 B | Replacement mixing bowl O-Ring |  |  | Each
| RT286-04-001 C | Replacement O-Ring for suction/vacuum nozzle |  |  | Each
42152217-00001 | RT286-04-002 | Vacuum Mixer: Plaster, Investment & Silicone Duplicating Material |  |  | Each
SPECIFICATION A wall mounted or benchtop mechanical vacuum mixer for the mixing of all dental plasters, investment compounds and silicone duplicating material utilised in a dental laboratory. The unit shall have the following features:
Usable as either wall mounted or bench unit.
Single speed vertical drive shaft, approx. 100 - 450rpm’s
Vacuum can be regulated from 70 - 100%
Vacuum created by motor pump and not compressed air.
Vacuum capacity 16 litre per minute
Electronic Controls
Functions should include prespatulation, prevacuum, reverse rotation, post vacuum.
Maintenance-free diaphragm pump
Accommodate three sizes mixing bowls: approx 250, 550, 850ml
Two mixing bowls provided per unit:
approximately 550ml
approximately 800ml
SABS plug
Optional items
| RT286-04-002 A | Mixing bowls approximately 200ml - 1200ml |  |  | Each
| RT286-04-003 B | Replacement mixing bowl O-Ring |  |  | Each
| RT286-04-004 C | Replacement O-Ring for suction/vacuum nozzle |  |  | Each
42152218-00000 | RT286-04-003 | Laboratory Vibrator |  |  | Each
SPECIFICATION A dental laboratory vibrator utilised during mixing of plaster and investment material, the casting of models and pouring of investment materials. The unit shall have the following features:
Size no smaller than 200 x 200mm and no larger than 270 x 380mm
ON / OFF switch on front of control panel
Electronically regulated variable speed control
Heavy Duty rubber cover on vibrating surface
Rubber cover designed with raised edge on outer rim
Uniform vibration to all parts of the tray
Fitted with vibration free rubber mountings
SABS plug
Optional items
| RT286-04-003 A | Replacement Rubber cover |  |  | Each
42152518-00014 | RT286-04-004 | Dispenser, Dental Acrylic Powder and Liquid: Orthodontic |  |  | Each
SPECIFICATION A benchtop, dispensing mechanism for the exact proportioning of acrylic powder and liquid in the removeable orthodontic application. The unit shall have the following features:
Operate without compressed air or electricity
Dispensing Unit consists of:
Powder dispenser for approximately 500g
Liquid dispenser for approximately 500ml
Dispensers positioned above working area for gravity feed
Two pedal switches for operating dispensers
Metal nozzles
Extraction system can be fitted
All sides of unit enclosed except front
Carbon filter to eliminate monomer vapours
Integration of at least 4 dispensers
Optional items
| RT286-04-004 A | Replacement metal nozzles |  |  | Each
| RT286-04-004 B | Replacement flexible pipes |  |  | Each
| RT286-04-004 C | Replacement carbon filter |  |  | Each
| RT286-04-004 D | Replacement pedal switches |  |  | Each
42152109-00000 | RT286-04-005 | Polymerisation Bath – Electric |  |  | Each
SPECIFICATION A benchtop, fully automatic, electrically heated water or steam bath utilised as a polymerisation apparatus fitted with a thermostat and designed for long duration polymerisation of any resin in a flask. The unit shall have the following features:
Capacity minimum of 10 flasks
Protected heating element
Overload protection
Optical display of functions
ON / OFF power switch on control panel
Housing of stainless steel
Integrated hinged lid
Heat insulated lid handle
Entire polymerisation process fully automatic and programmable
Progammable two-stage polymerisation
Programming options:
Preselection of pre-polymerisation temperature
Preselection of duration of pre-polymerisation up to 12 hours
Preselection of final polymerisation temperature
Preselection of duration of final polymerisation up to 12 hours
Unit to switch off automatically at end of programme
SABS plug
42152109-00001 | RT286-04-006 | Polymerisation Vessel – Pneumatic |  |  | Each
SPECIFICATION A benchtop, manually operated pressure polymerisation apparatus utilised for polymerisation of any cold curing resins. The unit shall have the following features:
Unit consists of:
Vessel
Lid
Hose & fitting to pressurise vessel
Capacity: minimum of 4 dentures
Operating pressure approx 2 bar
Inside of vessel coated with anti-corrosion teflon coating
The lid shall have the following features:
Safety lock to prevent accidental opening
Locking & pressure release switch
Pressure gauge
Inlet valve for air
Pressure over-load safety valve
Optional items
| RT286-04-006 A | Replacement lid seal |  |  | Each
42152109-00002 | RT286-04-007 | Polymerisation Vessel - Pneumatic & Heat |  |  | Each
SPECIFICATION A benchtop, fully automatic, thermo-regulated water heated polymerisation apparatus fitted with thermostat and designed for polymerisation of any cold curing resins. The unit shall have the following features:
Unit consists of:
Casing with pressure vessel
Integrated hinged lid with locking mechanism
Integrated safety features
Permanent connection to compressed air supply
Casing, chamber & lid of corrosion resistant material
Capacity minimum 2 models
Operating pressure maximum 5 bar
Automatically sealing system for lid
Thermostatically regulated temperature
Temperature manually adjustable
Maximum temperature 95°C
Unit fitted with temperature indicator light
Unit fitted with pressure gauge
Unit fitted with a water drain valve
Unit fitted with separate ON / OFF switch
SABS plug
Optional item
| RT286-04-007 A | Replacement lid seal/O-ring |  |  | Each
| RT286-04-007 B | Replacement ring clamp |  |  | Each
42152209-00000 | RT286-05-001 | Pre-Heating & Burnout Furnace – Small | Pre-Heating & Casting |  | Each
SPECIFICATION A benchtop, pre-heating and burnout furnace with electrical heating elements and electronic temperature controls, fitted with an extraction and catalyser system, utilised for the heating and burning out of casting rings for crowns, bridges and metal denture frames. The unit shall have the following features:
Maximum temperature: 1100°C
Furnace chamber volume: 3 litre
Furnace chamber surface area: approx. 150mm x 150mm
Stainless steel furnace casing
Solvent-free high grade enamel paint
Touch-safe outer wall temperature
Insulation by means of a fibre module
Furnace door to hinge downwards to horizontal position
Observation hole in door
Heating elements encased in ceramic plates
Heating from at least two sides
Exhaust duct in rear wall of furnace
External mounting device and connection for vapour vent with fan or catalyser by exhaust duct
Furnace fitted vapour vent and fan
Fan connection to program controller
Integrated program controller underneath furnace chamber
Features of integrated program controller:
Microprocessor controlled
Storage for at least 5 programmes
At least 3 heating ramps & holding times per program
Program steps: 1°C in 1 minute
Audible alarm
Real-time clock for programming
LED temperature display
Touch control keypad
Keypad protected from dust
Unit supplied with:
Steel collecting tray x 1
Ceramic collecting tray x 1
Ready for use with 2 m electrical cable & plug
SABS plug
Optional items
| RT286-05-001 A | Replacable heating elements |  |  | Each
| RT286-05-001 B | Replacement thermocouple |  |  | Each
| RT286-05-001 C | Replacement steel and ceramic collecting trays |  |  | Each
42152209-00001 | RT286-05-002 | Pre-Heating & Burnout Furnace – Medium |  |  | Each
SPECIFICATION A benchtop, pre-heating and burnout furnace with electrical heating elements and electronic temperature controls, fitted with an extraction and catalyser system, utilised for the heating and burning out of casting rings for crowns, bridges and metal denture frames. The unit shall have the following features:
Maximum temperature: 1100°C
Furnace chamber volume: 5 litre
Furnace chamber surface area: approx. 220mm x 180mm
Stainless steel furnace casing
Solvent-free high grade enamel paint
Touch-safe outer wall temperature
Insulation by means of a fibre module
Furnace door to hinge downwards to horizontal position
Observation hole in door
Heating elements encased in ceramic plates
Heating from at least two sides
Exhaust duct in rear wall of furnace
External mounting device and connection for vapour vent with fan or catalyser by exhaust duct
Furnace fitted vapour vent and fan
Fan connection to program controller
Integrated program controller underneath furnace chamber
Features of integrated program controller:
Microprocessor controlled
Storage for at least 5 programmes
At least 3 heating ramps & holding times per program
Program steps: 1°C in 1 minute
Audible alarm
Real-time clock for programming
LED temperature display
Touch control keypad
Keypad protected from dust
Unit supplied with:
Steel collecting tray x 1
Ceramic collecting tray x 1
Ready for use with 2 m electrical cable & plug
SABS plug
Optional items
| RT286-05-002 A | Replacable heating elements |  |  | Each
| RT286-05-002 B | Replacement thermocouple |  |  | Each
| RT286-05-002 C | Replacement steel and ceramic collecting trays |  |  | Each
42152209-00002 | RT286-05-003 | Pre-Heating & Burnout Furnace – Large |  |  | Each
SPECIFICATION A benchtop, pre-heating and burnout furnace with electrical heating elements and electronic temperature controls, fitted with an extraction and catalyser system, utilised for the heating and burning out of casting rings for crowns, bridges and metal denture frames. The unit shall have the following features:
Maximum temperature: 1100°C
Furnace chamber volume: 9 litre
Furnace chamber surface area: approx. 250mm x 250mm
Stainless steel furnace casing
Solvent-free high grade enamel paint
Touch-safe outer wall temperature
Insulation by means of a fibre module
Furnace door to hinge downwards to horisontal position
Observation hole in door
Heating elements encased in ceramic plates
Heating from at least two sides
Exhaust duct in rear wall of furnace
External mounting device and connection for vapour vent with fan or catalyser by exhaust duct
Furnace fitted vapour vent and fan
Fan connection to program controller
Integrated program controller underneath furnace chamber
Features of integrated program controller:
Microprocessor controlled
Storage for at least 5 programmes
At least 3 heating ramps & holding times per program
Program steps: 1°C in 1 minute
Audible alarm
Real-time clock for programming
LED temperature display
Touch control keypad
Keypad protected from dust
Unit supplied with:
Steel collecting tray x 1
Ceramic collecting tray x 1
Ready for use with 2 m electrical cable & plug
SABS plug
Optional items
| RT286-05-003 A | Replacable heating elements |  |  | Each
| RT286-05-003 B | Replacement thermocouple |  |  | Each
| RT286-05-003 C | Replacement steel and ceramic collecting trays |  |  | Each
42152226-00000 | RT286-05-004 | Casting Machine Centrifugal – Mechanical: Bench-Top |  |  | Each
SPECIFICATION A bench top mounted, mechanically activated centrifugal casting machine with articulated casting arm, utilised for flame casting of all precious and non-precious dental alloys in the dental laboratory. The unit shall have the following features:
Central base housing with integrated spiral spring
Drive shaft with articulated casting arm
Adjustable counterweight
Adjustable crucible and muffle holder
Slideable and Lockable crucible holder
Mechanical activation mechanism for centrifuge
Optional items
| RT286-05-004 | Replacement spiral spring |  |  | Each
| RT286-05-004 | Replacement crucible |  |  | Each
| RT286-05-004 | Replacement crucible insert |  |  | Each
42152226-00001 | RT286-05-005 | Casting Machine Centrifugal – Motorised |  |  | Each
SPECIFICATION A free standing unit, electrically activated motorised centrifugal casting machine with articulated casting arm, utilised for flame casting of all precious and non-precious dental alloys in the dental laboratory. The unit shall have the following features:
Centrifuge housed in free standing unit
Drive shaft with articulated casting arm
Adjustable counterweight
Adjustable crucible and muffle holder
Crucible holder lockable
Electrical activation of motorised mechanism for centrifuge
Motor housed below centrifuge
Rear hinged safety cover on top of unit
SABS plug
Optional items
| RT286-05-005 A | Replacement crucible |  |  | Each
| RT286-05-005 B | Replacement crucible insert |  |  | Each
| RT286-05-005 C | Replacement ring/muffle cradle |  |  | Each
42152226-00002 | RT286-05-006 | Casting Machine Centrifugal – Induction (Bench-Top Model) |  |  | Each
SPECIFICATION A bench top mounted, high frequency induction casting machine for casting of all precious and non-precious dental alloys in the dental laboratory. The unit shall have the following features:
Central base housing with integrated electric motor drive
Drive shaft with casting arm
Adjustable counterweight
Adjustable crucible and muffle holder
Crucible holder lockable
Tinted inspection port on lid of unit
Electrical activation of motorised mechanism for centrifuge
Induction coil surrounding crucible
Water cooled unit
SABS plug
Optional items
| RT286-05-006 A | Replacement spiral heating coil/element |  |  | Each
| RT286-05-006 B | Replacement crucible |  |  | Each
| RT286-05-006 C | Replacement crucible insert |  |  | Each
| RT286-05-006 D | Replacement muffle/ring cradle |  |  | Each
| RT286-05-006 E | Replacement coloured observation port glass |  |  | Each
42152226-00003 | RT286-05-007 | Induction Casting Machine - Vacuum Pressure and High Frequency Melting |  |  | Each
SPECIFICATION Complete with dual optic touch free sensor to read temperature accurately to prevent overheating of alloy.
Automatic and manual operation.
Must store up to 640 casting logs
Rated voltage 220V 50/60 Hz
Power consumption 14 amps
Melting power 3.6kva 65khz
Compressed air minimum 75psi
Water consumption approximately 1.5 gal.min
SABS plug
Optional items
| RT286-05-007 A | Replacement spiral heating coil/element |  |  | Each
| RT286-05-007 B | Replacement crucible |  |  | Each
| RT286-05-007 C | Replacement crucible insert |  |  | Each
| RT286-05-007 D | Replacement muffle/ring cradle |  |  | Each
| RT286-05-007 E | Replacement coloured observation port glass |  |  | Each
42152215-00000 | RT286-06-001 | Blast Cleaning Machine Cobalt–Chrome (Automatic) | Blasting & Extraction |  | Each
SPECIFICATION A benchtop, cobalt-chrome sandblaster unit for automatic blasting of up to 6 castings or manual blasting of individual castings, with integrated timer, regulator, gauges and multi-stage filtering capability. The unit shall have the following features:
Integrated multi-stage filtering capability (at least four stages)
On board filter module for self contained / independent operation
Pre-filter collection capacity approx. 1 kg
Filter module operates automatically when blaster is switched on
Warning signal when filter is full
Safety cut-out activated until full pre-filter bag is replaced
Tank capacity no less than 4,5 kg
Re-circulating system for blasting material
Utilise different grades (micron) blasting material
Air consumption no less than approx. 100 l/min
Compressed air connection between 4-6 bar
Lighting: At least one lighting unit
Lighting unit/s: No less than 15 Watt
Housing of blaster:
Plastic, CFC free
Corrosion resistant
Angled viewing window
One operating apertures per side of unit
Apertures sealed with resistant stretch material sleeve
Rotating sandblasing drum for at least 6 cobalt chrome frames
Drum mounted in-line with with sandblasting nozzle
ON / OFF switch for rotating drum
Timer for rotating drum
SABS plug
Optional items
| RT286-06-001 A | Replacement rotating drum |  |  | Each
| RT286-06-001 B | Replacement Filter bags |  |  | Each
| RT286-06-001 C | Replacement sleeves |  |  | Each
| RT286-06-001 D | Replacement gloves |  |  | Each
| RT286-06-001 E | Replacement Protective glass |  |  | Each
| RT286-06-001 F | Replacement light bulbs |  |  | Each
42152215-00001 | RT286-06-002 | Blast Cleaning Machine Crown, Bridge & Cobalt - Chrome (Manual) |  |  | Each
SPECIFICATION A benchtop, sandblaster unit for manual blasting of investment material in crown, bridge and partial denture work. The unit shall have the following features:
Must have an extraction system for unusable fine dust to be extracted.
Must have a minimum of two blasting tanks.
One Blasting tank for 20 – 70 micron aluminium oxide.
One Blasting tank for 70 – 250 micron aluminium oxide.
Must have a minimum of two Sandblasting nozzles:
One Sandblasting nozzle 0.8mm for 20 – 70 micron aluminium oxide.
One Sandblasting nozzle 1.2mm for 70 – 250 micron aluminium oxide.
SABS plug
Optional items
| RT286-06-002 A | Replacement fixed nozzles |  |  | Each
| RT286-06-002 B | Replacement sleeves |  |  | Each
| RT286-06-002 C | Replacement gloves |  |  | Each
| RT286-06-002 D | Filter bags for extraction unit |  |  | Each
42152215-00002 | RT286-06-003 | Blast Cleaning Machine Crown & Bridge – Micro Sandblaster (Manual) |  |  | Each
SPECIFICATION A benchtop, micro sandblaster with blasting nozzles / pencil jets for manual blasting and polishing of crown & bridge prosthetic units. The unit shall have the following features:
Compact size: No more than 400 x 300 x 560 mm
Casing exterior & interior corrosion proof & wear resistant
Minimum two sizes blasting nozzles / pencil jets with sand reservoirs
Deliver 25-50 micron and 90-250 micron sand
Capability to expand unit with two more sets of sand reservoirs and blasting nozzles
Foot switch controlled pneumatic valve
Immediate Stop (IS) function
Blasting chamber internal features:
Abrasion coated
Full tank width sieve above contaminated sand accumulation area
Tank selection switch
Compressed air regulator
Full tank width observation window
Protection screen for observation window
Built-in illumination
Illumination minimum 2400 LUX
Blasting nozzles & hoses
Colour coded blasting unit, tank selector switch & blasting handles
Constant air pressure in reservoirs and hoses
Reservoir maintain pressure after operation
Valve per reservoir for manual depressurising before opening
Pressure gauge on top of casing
Two operating apertures
Sturdy gloves fitted to apertures
Ventilation slits
Noise Level: Less than 70 dB(A)
Two sizes spare nozzles supplied with each unit
SABS plug
Optional items
| RT286-06-003 A | Replacement light bulbs |  |  | Each
| RT286-06-003 B | Protective gloves |  |  | Each
| RT286-06-003 C | Replacement fixed nozzles |  |  | Each
| RT286-06-003 D | Replacement protective glass |  |  | Each
| RT286-06-003 E | Replacement foot switch for pneumatic valve |  |  | Each
42152207-00000 | RT286-06-004 | Dust Extraction System (Dry) – Portable |  |  | Each
SPECIFICATION A portable dust extractor for one or two stations that can be used when working with handpieces, benchtop grinders or other on-bench appliances to prevent solid particles or dust particles from contaminating the working environment. The unit shall have the following features:
Suction Unit with two speed settings
Tri-layered disposable filter bag for suction unit
Accumulator for solids and most dust
Connecting hose (+-5cm) between accumulator and suction unit
T-Piece for accumulator
2 x Hoses (+-5cm) for work stations
2 x Work station attachments
2 x Remote control for suction unit
SABS plug
Optional items
| RT286-06-004 A | Replacement filter bag |  |  | Each
| RT286-06-004 B | Replacement ring clasp |  |  | Each
42152207-00001 | RT286-06-005 | Dust Extraction System (Dry) – Mobile |  |  | Each
SPECIFICATION A mobile (with castors), dust extractor for lab work stations, blasting units and dry model trimmers, or any other on-bench appliances to prevent solid or dust particles from contaminating the working environment. The unit shall have the following features:
Mobile unit on castors
Brushless motor: EC no-contact energy transfer
Lifespan of motor: Approx 800 hours
Motor noise level: No more than 62 dB (A)
Two stage filter system
Particle capture rate: 99,9%
Filter material to capture health endangering dust of MAK value smaller than 1 mg/ cubic m
Capacity of filter bag: 12kg
On-board automatic activator switch
Each unit shall be supplied with:
Tubing: 3m
Cloth Filter: x1
Replacement filter bags: 1 set of 5
SABS plug
Optional items
| RT286-06-005 A | Replacement filter bag |  |  | Each
| RT286-06-005 B | Replacement cloth filter |  |  | Each
42152207-00002 | RT286-06-006 | Dust Extraction System (Wet) – Fixed |  |  | Each
SPECIFICATION An in-line wet dust extraction system utilised when working with benchtop grinders and other benchtop appliances, for the elimination of gypsum, metal and silicon carbide particles from the air. The unit shall have the following features:
Primary filtering capability by means of water
Silent running of drive motor
T-Piece and hoses to connect to at least two waste particle producing appliances
SABS plug
Optional items
| RT286-06-006 A | Replacement water filter bag |  |  | Each
| RT286-06-006 B | Replacement ring clasp |  |  | Each
42152207-00003 | RT286-06-007 | Dust Extraction Drawer System (Dry) – Fixed |  |  | Each
SPECIFICATION An under-benchtop dust extraction drawer mounted below a lab workbench to suction gypsum, metal, acrylic and silicon carbide particles generated during working procedures of a dental technician. The unit shall have the following features:
Fitted under workbench, without hindering freedom of movement
Maximum dimensions: H180mm, W490mm
Minimum extension: 500mm
Suction hood and handrest
Automatic start-up linked to micromotor
Automatic switch-off after short micromotor pause
Working surface: Wood
Maintenance free induction motor
Filter bag:
Quick change capability
Large dimensional (Full drawer width)
Unit supplied with two spare filter bags
SABS plug
Optional items
| RT286-06-007 A | Replacement filter bag |  |  | Each
42152207-00004 | RT286-06-008 | Suction Hood - Electrical |  |  | Each
SPECIFICATION A benchtop or over-bench mounted hood for use in a furnace room to catch and extract fumes, smoke and heat generated by furnaces. Hood to be connected to existing ventilation. The unit shall have the following features:
Dimensions: At least 1 500mm wide and 700mm deep
Stainless steel construction
Design to expedite removal of all fumes, smoke and heat
Electrical fan mounted in top of hood
Connection to existing ventilation with flexible pipes
Lighting inside hood for improved visibility
ON / OFF switch with pilot light installed on front of hood
Additional carbon filter
SABS plug
Optional items
| RT286-06-008 A | Replacement carbon filter |  |  | Each
| RT286-06-008 B | Replacement light bulb |  |  | Each
42152801-00000 | RT286-07-001 | Pneumatic Chisel / Handpiece for Trimming of Plaster & Investment | Finishing / Surface Treatment |  | Each
SPECIFICATION A pneumatic deflasking chisel / handpiece for the removal of plaster and investment material. The unit shall have the following features:
Operating pressure 4-6 Bar
Piston frequency approximately 130 Hz
Adjustable power on chisel itself
Removable and interchangeable chisels
Heavy duty components made of High-Tech steel with hardcoat layer
Full strike power transferred straight to chisel insert
Non-slip handle
Light weight construction, approximately 400g
Supplied with at least 3 different chisels
Optional items
| RT286-07-001 A | Replacement chisel bits |  |  | Each
| RT286-07-001 B | Replacement ring clasp |  |  | Each
47121805-00000 | RT286-07-002 | Steam Cleaner-Portable |  |  | Each
SPECIFICATION A benchtop, electrically operated portable wet and dry steam cleaner utilised during production in a dental laboratory. The unit shall have the following features:
Water tank capacity no less than 4.5 litres
Portable due to manual filling of tank, no fixed connection to water supply
No less than 4.5Bar of wet or dry steam
Reaching operating pressure no longer than 15 minutes
An adjustable steam pressure knob to accurately select flow level of steam required
Construction of heavy duty stainless steel
The following indicators / switches to be provided:
Low water level light indicator
Water level alarm switch
Automatic pressure gauge protection switch
Automatic thermal safety system heating switch
SABS plug
Optional item
| RT286-07-002 A | One spare pressure cap supplied per unit |  |  | Each
| RT286-07-002 B | Replacement light bulb |  |  | Each
47121805-00001 | RT286-07-003 | Steam Cleaner-Fixed: Wet & Dry (High Performance) |  |  | Each
SPECIFICATION A benchtop, electrically operated fixed wet and dry steam cleaner utilised during production in a dental laboratory. The unit shall have the following features:
The steam cleaner shall be mainly utilised for:
Removing wax & polish residues
Cleaning models & trays
Degreasing metal components for ceramic application
A selection can be made between:
Wet steam
Dry steam (saturated steam)
Mixture dry steam & water
Unit is fixed - permanently connected to main water supply
Steam pressure approx 3 bar
Spray gun:
Push button pressure switch on gun
Pressure switch for thumb use
Insulated gun housing
Flexible insulated hose from gun to unit
Flexible hose at least 1 m long
Unit equipped with demineralisation cartridge
Unit housing of corrosion resistant steel & plastic
Pressure gauge visible on front of housing
Dry/wet/water selector switch on front of housing
Operating indicator LED’s on front of housing
Safety features:
Water leak monitoring sensor built into unit
Leak sensor shuts off water supply automatically
Releasing gun pressure switch cuts off steam
Internal connections shall be copper pipes and threaded joints for safety
Empty demineralisation cartridge shuts off water supply
Ring spanner
SABS plug
Optional items
| RT286-07-003 A | Ring spanner |  |  | Each
| RT286-07-003 B | Replacement full demineralisation cartridge |  |  | Each
| RT286-07-003 C | 2 Inserts |  |  | Each
| RT286-07-003 D | Replacement light bulb |  |  | Each
| RT286-07-003 E | Replacement flexible insulated hose from gun to unit |  |  | Each
| RT286-07-003 F | Replacement water supply hose |  |  | Each
| RT286-07-003 G | Replacement ring clasps |  |  | Each
42152208-00000 | RT286-07-004 | Laboratory Motor Electric (Hanging) |  |  | Each
SPECIFICATION An electric laboratory motor with flexible cable transmission installation comprising of an electric motor, electronic foot control, flexible cable transmission drive and motor stand The unit shall have the following features:
Electric Motor
Power: 220 ‑ 240 V, 50 ‑ 60 Hz
Commutor drive unit with built‑in ventilator for automatic cooling
Motor functioning on variable speeds between approximately 5 000-20 000 rpm
Torque not less than 18 Ncm
Output not less than 180 Watt
Ball bearings maintenance free
Rapid braking when switched off
Radio interference free
Vibration reducing hanging device part of motor assembly
Electronic foot-control
Guaranteed variable speed regulation for ranges of motor
Sealed housing for protection against dust and electrical shock
Detachable from the motor without tools
3 Pin 15 Amp plug with at least 2 meter 3 core flex shall be supplied
Flexible cable drove
Outer cable consist of a special protective flexible spiral metal tube of approximately 12mm in diameter
Cable at motor connection detachable without special tools
Cable at handpiece connection ball bearing Slip-Joint type
Ball bearings at both ends of drive cable
Inner drive shaft not less than 5 mm in diameter
Length not less than 1000 mm
Stand
Over the bench rust protected steel projection
A minimum adjustable vertical length of 1000 mm
G-type screw clamp shall be able to fix to bench tops of 17-61 mm in thickness
SABS plug
Optional items
| RT286-07-004 A | Replacement cable |  |  | Each
| RT286-07-004 B | Replacement hanging motor handpiece catch (end of cable) |  |  | Each
42152208-00001 | RT286-07-005 | Handpiece for Laboratory Motor Electric (Hanging) |  |  | Each
SPECIFICATION A dental laboratory handpiece that is driven by a dental laboratory motor by means of a flexible transmission cable. The unit shall have the following features:
Continuous running with drive speed up to 25 000 rpm.
Exchangeable chucks to accept tools with shafts 2,35 to 3,175 mm in diameter
Chuck to accept 2,35 mm diameter tools shall be supplied
Twist chuck tension device for rapid tool insert
Chuck tension device covered
Handpiece equipped with three sealed bearings
Ball bearings self-lubricating
Connection to cable of the slip-joint type
Anatomically shaped for the average size hand
Metal handpiece housing covered with heat-insulated plastic
Insulation colourfast and not react chemically with dental materials
Weight not exceeding 170 g
Optional item
| RT286-07-005 A | Replacement Chuck |  |  | Each
42152208-00002 | RT286-07-006 | Dental Laboratory Micro Motor (Multi-Purpose) |  |  | Each
SPECIFICATION A dental laboratory micro-motor consisting of a drive system for intermediate and precision work with a hand guided, brushless, handpiece to be used for long periods of time at speeds between 1 000 & 45 000 Rpm’s in the crown & bridge, cobalt-chrome and light acrylic technique. Tabletop and / or foot control unit, forward/reverse, On/Off and foot pedal for speed control function. The unit shall have the following features:
Motor: Brushless - Induction
Speed range: Approximately 1 000 to 45 000 Rpm's
Maximum torque: At least 7.8 Ncm
Concentricity: Approximately 0.02 mm
Motor Cooling: Fan integrated in motor housing
Handpiece:
Design: Ergonomical shape for fitting in the hand
Supplied with 2.35 mm chuck
Tool insert / exchange by means of twist chuck - not lever
Rapid braking on switching off motor
Flexible connect cable with socket from motor to control unit
Control Unit:
Table Top Unit with ON/OFF, Forward/Reverse, speed control & Foot pedal
Foot Unit with ON/OFF, Forward/Reverse & variable speed control foot pedal
Handpiece stand supplied with each unit
SABS plug
Optional item
| RT286-07-006 A | Replacement Chuck |  |  | Each
| RT286-07-006 B | Replacement fuse |  |  | Each
| RT286-07-006 C | Replacement foot pedal |  |  | Each
42152208-00003 | RT286-07-007 | Dental Laboratory Micro Motor (Heavy Duty) |  |  | Each
SPECIFICATION A dental laboratory micro-motor consisting of a drive system for intermediate and precision work with a hand guided, brushless, handpiece to be used for long periods of time at speeds between 3 000 & 65 000 Rpm’s in the heavy acrylic and cobalt-chrome technique. Tabletop and / or foot control unit, forward/reverse, On/Off and foot pedal for speed control function. The unit shall have the following features:
Motor: Brushless - Induction
Speed range: Approximately 3 000 to 50 000 Rpm's
Maximum torque: At least 8.7 Ncm
Concentricity: Approximately 0.02 mm
Motor Cooling: Fan integrated in motor housing
Handpiece:
Design: Ergonomical shape for fitting in the hand
Supplied with 2.35 mm chuck
Tool insert / exchange by means of twist chuck - not lever
Rapid braking on switching off motor
Flexible connect cable with socket from motor to control unit
Control Unit:
Table Top Unit with ON/OFF, Forward/Reverse, speed control & Foot pedal
Foot Unit with ON/OFF, Forward/Reverse & variable speed control foot pedal
Handpiece stand supplied with each unit
SABS plug
Optional item
| RT286-07-007 A | Replacement Chuck |  |  | Each
| RT286-07-007 B | Replacement foot pedal |  |  | Each
42152208-00004 | RT286-07-008 | Air Turbine Non - LED |  |  | Each
SPECIFICATION A benchtop, high-torque laboratory turbine straight handpiece unit, for precise shaping of hard and very hard materials in the production section of a dental laboratory. The unit shall have the following features:
Air consumption: Between 40 & 50 litres / minute
Operating Pressure: Between 2,4 & 3,4 bar
Maximum speed: 300,000 rpm’s
Chuck diameter: 1,6 mm
Wall or under-bench mounting possible
Unit shall consist of:
Turbine handpiece
ON / OFF foot control
Control Unit with filter & pressure gauge
Handpiece rest
Rotor cartridge (Plus spare rotor cartridge)
Water reservoir
Lubrication: Manual or automatically
Handpiece features:
Quick chuck mechanism
Cartridge system
Two-fold precision ball bearings
Extremely flexible connecting hose
Anatomically shaped
Silencer
Air escape after turbine is driven
Water cooling and dispensing
Optional items
| RT286-07-008 A | Replacement rotor cartridge |  |  | Each
| RT286-07-008 B | Replacement foot pedal |  |  | Each
| RT286-07-008 C | Replacement ring clasp |  |  | Each
42152208-00005 | RT286-07-009 | Air Turbine LED |  |  | Each
SPECIFICATION A benchtop, high-torque laboratory turbine straight handpiece unit, for precise shaping of hard and very hard materials in the production section of a dental laboratory. The unit shall have the following features:
Air consumption: Between 40 & 50 litres / minute
Operating Pressure: Between 2,4 & 3,4 bar
Maximum speed: 300,000 rpm’s
Chuck diameter: 1,6 mm
Wall or under-bench mounting possible
Unit shall consist of:
Turbine handpiece
ON / OFF foot control
Control Unit with filter & pressure gauge
Handpiece rest
Rotor cartridge (Plus spare rotor cartridge)
Water reservoir
Lubrication: Manual or automatically
Handpiece features:
Quick chuck mechanism
Cartridge system
Two-fold precision ball bearings
Extremely flexible connecting hose
Anatomically shaped
Silencer
Air escape after turbine is driven
Water cooling and dispensing
LED light
SABS plug
Optional items
| RT286-07-009 A | Replacement rotor cartridge |  |  | Each
| RT286-07-009 B | Replacement LED light |  |  | Each
| RT286-07-009 C | Replacement foot pedal |  |  | Each
| RT286-07-009 D | Replacement ring clasp |  |  | Each
42152208-00006 | RT286-07-010 | Polishing & Grinding Machine with built-in Suction, Lighting & Eye Protection |  |  | Each
SPECIFICATION A polishing and grinding machine used in a dental laboratory, utilised for the rough trim, pumice and polish of plastic and metal dentures. The unit shall have the following features:
Central drive unit with two shafts for mounting/fitting spindles
At least two speed settings
Integrated suction system per working area
Filter bags easily replaceable
Waste drawer for solids built into base of unit
Separate polishing and extraction motor
Separate power and lighting switch
Integrated lighting per working area
LED lights providing lighting
Integrated eye protection shield per working area
Shield of safety glass
Construction of stainless steel, dust tight and splash proof
Working surface anti-corodal aluminium
Rest of unit shall be of lacquered stainless steel and / or high impact plastic
SABS plug
Optional item
| RT286-07-010 A | Replacement filter bag |  |  | Each
| RT286-07-010 B | Replacement LED light |  |  | Each
| RT286-07-010 C | Replacement shield for safety glass |  |  | Each
42152208-00007 | RT286-07-011 | High Speed Grinder with Internal Suction, Lighting & Eye Protection |  |  | Each
SPECIFICATION A benchtop high speed grinding machine used in a dental laboratory, utilised for finishing prosthetic castings made from non-precious and precious metal alloys. The unit shall have the following features:
Drive unit power consumption approximately 250W
Drive system to run on vibration free bearings
High torque at 24 000 rpm’s
No speed reduction under normal work load
No switching off of motor when changing burs/grinding stones/tools
Quick-action collect chuck release for fast tool change
Extraction capacity shall be no less than 3 cubic m/min
Internal lighting shall be provided
Well lit working area
Hinged eye-protection cover shall be provided for the working area
Shield of safety glass
SABS plug
Optional items
| RT286-07-011 A | Replacement chuck |  |  | Each
| RT286-07-011 B | Replacement light bulb |  |  | Each
| RT286-07-011 C | Replacement shield for safety glass |  |  | Each
| RT286-07-011 D | Replacement drive belt |  |  | Each
42152208-00008 | RT286-07-012 | Electrolytic polishing Unit – One Casting |  |  | Each
SPECIFICATION A benchtop, electronically controlled unit utilised for electrolytic polishing of cobalt-chrome castings in a dental laboratory. The unit shall have the following features:
Polishing of single cobalt-chrome castings
Separate immersion polishing bath to control unit
Immersion polishing bath capacity 1 litre
Immersion bath sealed with lid
Features on control panel of housing:
ON / OFF power switch
Current indicated on ammeter
LED operating indicators
Adjustable polishing current by means of a dial
Timer
Constant selected current value
Polishing current electronically controlled
Protected clasp tips
No striations on castings after polishing
Control unit of corrosion resistant steel & plastic
SABS plug
Optional items
| RT286-07-012 A | Replacement clamps and hook |  |  | Each
42152208-00009 | RT286-07-013 | Electrolytic Polishing Unit - Two or More Castings |  |  | Each
SPECIFICATION A benchtop, electronically controlled unit utilised for electrolytic polishing of cobalt-chrome castings in a dental laboratory. The unit shall have the following features:
Simultaneous polishing of at least two cobalt-chrome castings
Integrated immersion polishing tub and control unit
Two laterally moving anode supports
Anode supports fitted with knurled screw clamps
Immersion polishing tub capacity 2 litre
Immersion tub covered by hinged lid
Lid fitted with observation window
Features on control panel of housing:
ON / OFF power indicator
LED current indicators
Adjustable polishing current bmo dial
Timer
Special features:
Thermostatically controlled bath heating
Automatic control of current intensity
Three power settings
Ring cathode
Automatic control of object movement
Polish completion in 2 x 4-6 minutes
Quick action short-circuit cut-out
No striations on castings after polishing
Supplementary cathode for deep palatal sections
Unit made of corrosion resistant steel & plastic
SABS plug
Optional item
| RT286-07-013 A | Replacement clamps and hook |  |  | Each
42152208-00010 | RT286-07-014 | Ultrasonic Cleaning Unit- Laboratory: Large Capacity |  |  | Each
SPECIFICATION A benchtop, electronically controlled, ultrasonic cleaner designed for all types of cleaning of prosthetic appliances and instruments in a dental laboratory. The unit shall have the following features:
Should consist of a hinge cover for easy one hand opening
Stainless steel tank
Tank dimensions: approx 19.7 x 14.6 x 10.5 cm
Tank capacity: approx 2 litre
Insert basket for instruments
Adjustable time cycles
Degas cycle for degassing fresh solution
Selectable run times
Countdown timer
Solution tracking - to see how long solution has been in use
Time cycle: 0 to 30 minutes
Transducer frequency: approx 50 kHz
Thermostatically adjusted heating
SABS plug
Optional item
| RT286-07-014 A | Replacement insert basket |  |  | Each
| RT286-07-014 B | Replacement ultrasonic solution |  |  | Each
42152208-00011 | RT286-07-015 | Ultrasonic Cleaning Unit- Laboratory: Small Capacity |  |  | Each
SPECIFICATION A benchtop, electronically controlled, ultrasonic cleaner designed for cleaning of single dentures (acrylic / metal) as well as crowns and bridges during the production process in a dental laboratory. The unit shall have the following features:
Stainless steel casing & lid
Stainless steel tank
Tank dimensions: approx L:150, W:85, H:65mm
Tank capacity: approx 0,5 litre
Manual or automatic temperature & time control
Time cycle: 0 to 10 minutes
Transducer frequency: approx 60 kHz
Thermostatically adjusted heating
SABS plug
Optional item
| RT286-07-015 A | Replacement ultrasonic solution |  |  | Each
42152209-00003 | RT286-07-016 | Porcelain Firing Furnace - Press/Ceramics |  |  | Each
SPECIFICATION A benchtop dental porcelain firing furnace and vacuum pump, with automated and computerised control for firing of porcelain work with a minimum of 400 definable firing programmes. The unit shall have the following features:
Should be a combination furnace
Both a press and ceramic furnace
Should be a user friendly furnace
Furnace should specially coordinate with IPS e-max press materials
have at least 20 individual press programs
Should contain at least 500 individual firing programs
Should consist of a bottom reflector that ensures homogeneous heat distribution in the firing chamber
Should consist of an integrated Crack Detection System
Should contain Infrared Technology to recognize size and control tempreture of investment rings
Firing status indicated by means of LED lights and audio warning signals
Contain optical status and progress display feature
Electronic press drive with force sensor to produce high-quality results
No compressed air connection required.
Automatic double-range temperature calibration
Vacuum pump maximum power consumption of 250 W
Dimensions of firing chamber at least 90mm diameter and 80 mm high
Maximum firing temperature of at least 1 200°C
Software update via USB flash drive
Vacuum pump integrated with control unit and firing unit
SABS plug
Optional item
| RT286-07-016 A | Replacement heating element |  |  | Each
| RT286-07-016 B | Replacement thermocouple |  |  | Each
| RT286-07-016 C | Firing tray |  |  | Each
42152209-00004 | RT286-07-017 | Porcelain Firing Furnace-VITA |  |  | Each
SPECIFICATION A benchtop dental porcelain firing furnace and vacuum pump, with automated and computerised control for firing of porcelain work with a minimum of 400 definable firing programmes. The unit shall have the following features:
Firing Furnace system consists of firing unit, control unit and vacuum pump
Compatible with all VITA materials and products
Firing unit:
Fully automatic and microprocessor controlled
Firing status indicated by means of LED lights and audio warning signals
Firing muffle of silica quarts
Integrated cooling trays
Safety and convenience features:
Controlled night mode and fast cooling
Automatic cleaning function
Automatic temperature adjustment before every program start
Control Unit:
No smaller than 16 cm
Connection by means of USB
Touch screen with photo viewer
Choice of different firing techniques
Minimum of 400 defineable firing programmes
Able to monitor temperature and vacuum
Vacuum pump integrated with control unit and firing unit
SABS plug
Optional item
| RT286-07-017 A | Replacement heating element |  |  | Each
| RT286-07-017 B | Replacement thermocouple |  |  | Each
| RT286-07-017 C | Replacement USB cable |  |  | Each
| RT286-07-017 D | Firing tray |  |  | Each
42152015-00001 | RT286-07-018 | InLab-InEOS CAD-CAM |  |  | Each
SPECIFICATION A benchtop Multifunctional CAD/CAM system for Dental Laboratories consisting of a Multi Axis Scanner, Milling Unit, Sintering Furnace and Operating System. The unit shall have the following features:
Multi Axis Scanner:
Table Top model with coaxial scanner and milling tools.
Functional and user-friendly design.
Manual and automatic function
Quick image processing at a rate of at least 1mil pixels in 0.140 seconds and capture entire Jaw in less than 23 seconds.
Be able to position the most popular articulators.
Be able to scan models (all popular plasters) and impressions.
The export of scanning data must be in open STL format for processing on third party CAD/CAM systems.
USB interphase
Milling Components:
inLAB MC XL component
Self-calibrating and auto compensation for the wear of the grinding tools ensures utmost precision.
Milling steps at least 7um during grinding of restoration.
Grain roughness of the finished milled restorations must be at least @ 50 um.
The unit must have at least a self contained water and coolant chamber.
The unit must be able to use at least the following materials:
a. Aluminium Oxide
b. Feldspar/glass ceramics
c. Infiltration Ceramics
d. Lithium disilicate
e. Models from polyurethane polymers.
f. Non Precious metals (pre sintered Cobalt Chrome)
g. Resin NANO Ceramics
h. Silicate ceramics
i. Ti Base sets for implants
j. Translucent zirconium oxide for full contour restorations
k. zirconium oxide
Various burs must be available for different materials
The bidder must include all software components needed for Framework, VinCrOn and Wax Up
The unit must at least have the following wide-ranging of clinical indications:
a. Abutments: (Custom hybrid titanium base for most popular implant systems)
b. Attachments: for attachment pins
c. Bars: choice of several geometries, Design tools for customization
d. Crown Coping and bridge frameworks (METAL). Partially or fully reduced frameworks
e. Crown and Bridge frameworks – metal
f. Full Crowns
g. Inlays
h. Veneers. (min thickness 300µm
i. Onlays
j. Scanbodies for transferring exact position of implant to design software
k. Stress breaker attachments
l. Surgical guide
m.Telescopes
inLAB MC X5 component
Must have 2 rotational and 3 translation axes
Accommodate complex geometries such as long span bridges and bars, models, surgical guides
Be able to hold up to 6 tools per magazine
Highly precise detection of the exact position of disks and blocks for processing of blocks with dentine structures and mesostructures
Chamber. At least have a dia of 95mm and height 30mm and able to be used for Zirconium, PMMA, composite and wax
Indications of models, splints, surgical guides, prosthesis etc
Open for disk supplies of all kinds
Multi – block management for blocks up to 40 X 19 X 19 mm and be able to at least used for grinding glass ceramics, hybrid ceramic & lithium disilicate ceramics
External water tank for coolant be supplied
External suction be supplied with automatic communication with the machine
Sintering Furnace
The machine must be programmable with the most common brands of zirconia material
Must have at least four pre-set programme positions for sintering with pre-drying and ventilation
Must have an integrated gas management module
At least be able to sinter up to 50 units at once
Operating System
Must be supplied with all hard- and software to run the system
The software must support all working steps namely scanning, milling of restorations and models, abutments, bars, surgical guides for implant placement
Open system for flexible integration. STL data of all dental CAD software
Intergratable with other systems including Cone beam machines
SABS plug
Optional items
| RT286-07-018 A | Replacement bur chuck |  |  | Each
| RT286-07-018 B | Replacement burs |  |  | Each
| RT286-07-018 C | Replacement water filter |  |  | Each
| RT286-07-018 D | Replacement air filter |  |  | Each
| RT286-07-018 E | Milling fluid |  |  | Each
23271406-00000 | RT286-08-001 | Heavy Duty Dental Spot Welder | Jointing / Soldering |  | Each
SPECIFICATION A benchtop, heavy-duty spot welding and soldering unit with electronic control for precise dose performance within the dental industry. The unit shall have the following features:
Finely adjustable electrode distance and electrode distance to prevent distortion during spot welding
Finely adjustable current to enable addition of fine wires to heavy castings
Spot welding control via optical & acoustic signal
Possible to spot weld and solder without changing electrodes
Foot control for soldering as well as opening the fixed electrodes
Spot welding features:
Possible for wires, clasps, solder
Choice between hand-held or fixed electrodes
Welding time control for overheating protection
Suitable for all dental alloys except titanium
Soldering features:
Freehand soldering with carbon electrodes
Model soldering with copper electrodes
Suitable for all dental alloys except titanium
SABS plug
Optional items
| RT286-08-001 A | Replacement electrodes |  |  | Each
23271408-00000 | RT286-08-002 | Inert Gas Welder |  |  | Each
SPECIFICATION A benchtop, heavy-duty inert gas welding unit and welding microscope with electronic control for applying welds to all common dental alloys and titanium, for work associated with the repairing or manufacturing of dental appliances for patients. The unit shall have the following features:
Consists of control unit, welding handpiece with flexible pipe & welding microscope
Control unit:
Main voltage: 230 V AC
Power input approximately 400 VA
Speed control for welding
Memory slots for for personalised welding parameters
Adjustable Impulse time
Adjustable welding power
Connect to Medical grade Argon gas supply for inert gas welding
Welding handpiece:
Tungsten electrode
Utilise shielding (inert) gas
Oxidation free metal
Low heat generation
Minimum thickness of sheet welding: 0.10 mm
Minimum thickness of wire welding: 0.25 mm
Replaceable tungsten welding electrode
Welds per electrode: at least 1800
Stero microscope:
Mounted on adjustable stand
Stand with supporting arm for welding hadpiece
At least 10x magnification
Eye protection system
Automatic shutter system for eye protection
SABS plug
Optional items
| RT286-08-002 A | Replacement electrodes |  |  | Each
| RT286-08-002 B | Replacement oxidation free metal |  |  | Each
| RT286-08-002 C | Replacement Argon gas regulator |  |  | Each
| RT286-08-002 D | Replacement Argon gas flexible pipe |  |  | Each
| RT286-08-002 E | Replacement ring clasps |  |  | Each
23271405-00000 | RT286-08-003 | Laser Welder |  |  | Each
SPECIFICATION A high performance, inert gas, free standing laser-welding unit with integrated stereo-microscope. The unit shall have the following features:
Microscope magnification: No less than 15-fold
Sighting device consisting of reticle in microscope, plus 2 crossed sighting light diodes
Parameter display at front
Adjustable manipulation openings
Lighting in working chamber: Minimum 2 adjustable cold light reflectors
Integrated cooling system
Integrated argon inert gas system
Inert gas jet: Minimum 1
Integrated fume-gas extraction system
Pulse energy: 50 joules
Pulse length: 0.5 – 20 ms
Mean Power: 20 W
Peak power minimum: 4 kW
SABS plug
Optional items
| RT286-08-003 A | Replacement electrodes |  |  | Each
| RT286-08-003 B | Replacement Argon gas regulator |  |  | Each
| RT286-08-003 C | Replacement Argon gas flexible pipe |  |  | Each
| RT286-08-003 D | Replacement ring clasps |  |  | Each
42151706-00000 | RT286-09-001 | Chair Dental Laboratory upholstered | Dental Laboratory Fixtures |  | Each
SPECIFICATION An anatomically correct, fully adjustable, mobile upholstered chair utilised by dental technicians in the production section of a dental laboratory. The unit shall have the following features:
Frame: Robust metal construction
Metal finish: Chrome/ anodised / powder coated
Base consist of five horisontal legs
One 360° double wheel castor per leg
Seat:
Anatomically contoured
Prevent pressure on thighs
Prevent restriction of blood circulation
Adjustable seat height: 500 - 600mm
Upholstery breathing & perforated
Backrest:
Anatomically contoured
Adjust horisontally & vertically with body position
Prevent tiring of the lower back
Minimum height adjustment: 200mm
Minimum rake adjustment: 100mm
Upholstery breathing & perforated
Foam minimum two layers with:
Approx 100mm heavy density (VP45/ Chip 80 Foam) and
With 10mm light density
42151706-00001 | RT286-09-002 | Chair Dental Laboratory with silicone |  |  | Each
SPECIFICATION An anatomically correct, fully adjustable, mobile silicone upholstered chair utilised by dental technicians in the production section of a dental laboratory. The unit shall have the following features:
Frame: Robust metal construction
Metal finish: Chrome/ anodised / powder coated
Base consist of five horisontal legs
One 360° double wheel castor per leg
Seat:
Anatomically contoured
Prevent pressure on thighs
Prevent restriction of blood circulation
Adjustable seat height: 500 - 600mm
Backrest:
Anatomically contoured
Adjust horisontally & vertically with body position
Prevent tiring of the lower back
Minimum height adjustment: 200mm
Minimum rake adjustment: 100mm
Foam minimum two layers with:
Approx 100mm heavy density (VP45/ Chip 80 Foam) and
With 10mm light density
42152303-00000 | RT286-09-003 | Magnifying Light – Circular |  |  | Each
SPECIFICATION A combination light and magnifier for utilisation in the production section of a dental laboratory. The unit shall have the following features:
Circular in design
No sharp edges to cause accidental injury
Housing of lamp impact resistant plastic
ON / OFF switch on top of lamp housing
Magnifying lens of cut glass - 2 dioptre
Circular lamp positioned around magnifying lens
LED lighting
Lamp fitted into base of lamp housing
Housing to prevent light shining into operator eyes
Mounting / base of lamp by means of:
Bench Clamp, or
Fixing bracket for wall and / or bench, or
Weighted base
Light head three dimensionally adjustable
Adjustment by means of a double articulating arm pivoting in base
Features of lamp:
Create natural light
Consistent light intensity by means of electronic control
Independent of ambient lighting, daylight, or other colour influencing factors
Electric supply to light via base and arm to head
SABS plug
Optional items
| RT286-09-003 A | Replacement circular LED light |  |  | Each
| RT286-09-003 B | Replacement magnifying lens |  |  | Each
42152214-00001 | RT286-09-004 | Plaster Trap - Under Basin |  |  | Each
SPECIFICATION A plaster trap with separating compartments to collect solids and suspended sediments from water from the overhead drain basin, before draining into the regular sewerage system. The unit shall have the following features:
Manufactured of corrosion resistant stainless steel or ABS plastic material
Mounted underneath all commercially available basins and in cabinets
At least 3 seperating compartments
Tight fitting cover / lid over plaster trap to eliminate odours
Plumbing connection to sewerage system according to building regulations
Design to facilitate ease of cleaning
42152008-00001 | RT286-10-001 | Intra-Oral X-Ray Processor-Automatic | Equipment |  | Each
SPECIFICATION An automatic intra-oral x-ray processeor for all sizes of intra-oral x-rays. The unit must have the following features:
Process intra-oral x-ray films up to 65 mm
Processing time approximately 2 minutes for wet endodontic
Processing time approximately 4 minutes for dry
Equipped with integrated drier
Equipped with daylight loader
SABS Plug
42152008-00002 | RT286-10-002 | Intra-Oral X-Ray Processor-Manual |  |  | Each
SPECIFICATION A manual intra-oral x-ray processeor for peri-apical and bite wing dental film, with internal dipping containers for developing, fixing and cleaning (water) liquid. The unit must have the following features:
Process intra-oral x-ray peri-apical and bite wing dental film
Casing manufactured from corosion resistant material
Supplied with 3 internal dipping containers for developing, fixing & cleaning liquid
Equipped with viewing screen
Equipped with sleeved hand access ports
Optional items
| RT286-10-002 A | Replacement sleeves |  |  | Each
42281905-00000 | RT286-10-003 | Heat Sealer |  |  | Each
SPECIFICATION A desktop unit utilised for heat sealing instruments or tools in sterilisation bags (plastic / paper combination prior to sterilisation in dental autoclaves. The unit shall have the following features:
Include a built-in cutter
Must have an automatic temperature control
Must have an automatic Start / Stop control
Optional Items
| RT286-10-003 A | Heating element |  |  | Each
| RT286-10-003 B | Cover |  |  | Each
| RT286-10-003 C | replacement element |  |  | Each
| RT286-10-003 D | Insulator |  |  | Each
| RT286-10-003 E | Cutter Balade |  |  | Each
42152301-00000 | RT286-10-004 | Diode Laser Dental - Portable |  |  | Each
SPECIFICATION A portable dental diode laser for soft tissue surgery and bleaching in the dental surgery. The unit must have the following features:
Application to include soft tissue surgery and bleaching
Operating Voltage to 240 volts
Wavelength between 445 to 980nm
Disposable tips available in a variety of shapes sizes and angles
Touch Screen Menu
Autoclavable handpiece
Handpiece must accept:
a) Standard canulas
b) Multi tips
c) Stripable Fibre
Foot operation
SABS Plug
42293602-00000 | RT286-10-004 | Endodontic Obturator Oven |  |  | Each
SPECIFICATION A specifically designed heat source for heating endodontic obturators. The unit must have the following features:
Give fast and predictable results
Heats obturators in at least 17 seconds
Integrated timer for heating oven
Signal on timer when oven is at optimal temperature
A specifically designed heat source for heating endodontic obturators. The unit must have the following features:
Give fast and predictable results
Heats obturators in at least 17 seconds
Integrated timer for heating oven
Signal on timer when oven is at optimal temperature
41104202-00000 | RT286-10-006 | Water Demineraliser Unit |  |  | Each
SPECIFICATION A desktop water demineraliser unit for dental autoclaves. The unit MUST have the following features:
Wall mounted
No solenoid valves
unpressurized system for tank filling
Monitored by conductivity motor
Acessories
| RT286-10-006 A | Cartriges |  |  | Each
| RT286-10-006 B | SABS Plug |  |  | Each
41104820-00000 | RT286-10-007 | Water Distillation Unit |  |  | Each
SPECIFICATION A desktop water distillation unit for dental autoclaves. The unit MUST have the following features:
Must be able to supply at least 4 Liters of distilled water to different types of autoclaves
Full installation of the water distillation unit on delivery
SABS Plug
42281712-00000 | RT286-10-008 | Ultrasonic Cleaning Unit |  |  | Each
SPECIFICATION An ultrasonic bath utilised for the cleaning of dental instruments prior to sterilisation. The unit shall have the following features:
Minimum 4.5 Litre tank (working capacity)
Supplied with instrument basket
Equipped with an integrated timer
Integrated hinged lid
SABS PLUG
Optional items
| RT286-10-008 A | Basket |  |  | Each
| RT286-10-008 B | Beaker |  |  | Each
| RT286-10-008 C | Bur tray |  |  | Each
42151651-00000 | RT286-10-009 | Amalgam Separator Multiple Surgeries |  |  | Each
SPECIFICATION A fixed installation amalgam separator for up multiple surgeries. The unit MUST have the following features:
Fixed installation
Equipped for connection of multiple treatment units
Serve number of surgeries:2 - 4
Output power: Single-phase, Approximately 210W
Noise level: No more than 60dB(A)
Volume of collector vessel: Approximately 600 cm3
Separation rate: At least 97%
Flow rate at least 15 l/min
Automatic filling level check
Automatic switch on when waste water is available from treatment unit
Centrifugal separator with hydrodynamic self-cleaning function
Optional Items
| RT286-10-009 A | Collection Container with recycle kit |  |  | Each
42151651-00003 | RT286-10-010 | Amalgam Separator 2 Surgeries |  |  | Each
SPECIFICATION A fixed installation amalgam separator for up to 2 surgeries. The unit MUST have the following features:
Fixed installation
Equipped for connection of several treatment units
Serve number of surgeries: Maximum 2
Output power: Single-phase, Approximately 115W
Noise level: No more than 60dB(A)
Volume of collector vessel: Approximately 90 cm3
Separation rate: At least 95%
Flow rate at least 4 l/min
Automatic filling level check
Automatic switch on when waste water is available from treatment unit
Centrifugal separator with hydrodynamic self-cleaning function
Optional Items
| RT286-10-010 A | Collection Container with recycle kit |  |  | Each
42271907-00000 | RT286-10-011 | Suction Wet 2-4 Surgeries |  |  | Each
SPECIFICATION A fixed installation centraised wet aspiration plant for serving 2-4 surgeries. The unit MUST have the following features:
Fixed installation
Serve number of surgeries: 2 - 4
Output power: Single-phase, Approximately 1.7kW - 9A
Maximum flow: No less than 820 l/min
Operating head for continuous service: Maximum 3300mm H2O
Average water consumption: Approximately 6 l/min
Piping pressure: Approximately 4 bar
Dimensions: Approximately 370mm W, 300mm D, 420mm H
42271907-00002 | RT286-10-012 | Suction Dry 4-6 Surgeries |  |  | Each
SPECIFICATION A fixed installation centraised dry aspiration plant for serving 6 surgeries. The unit MUST have the following features:
Fixed installation
Serve number of surgeries: 4 - 6
Output power: Three-phase, Approximately 2kW - 3.5A, 50 Hz
Speed of motor: No less than 2800 rpm
Vacuum setting at approximately 160 mBar
Noise level without housing: No more than 65dB(A)
42271907-00003 | RT286-10-013 | Suction Dry 2-3 Surgeries |  |  | Each
SPECIFICATION A fixed installation centraised dry aspiration plant for serving 3 surgeries. The unit MUST have the following features:
Fixed installation
Serve number of surgeries: 2 - 3
Output power: Single-phase, Approximately 1.8kW - 5A, 50 Hz
Speed of motor: No less than 2800 rpm
Vacuum setting at approximately 170 mBar
Noise level without housing: No more than 65dB(A)
42271907-00004 | RT286-10-014 | Suction Wet 1-2 Surgeries |  |  | Each
SPECIFICATION A fixed installation centraised wet aspiration plant for serving 1-2 surgeries. The unit MUST have the following features:
Fixed installation
Serve number of surgeries: 1 - 2
Output power: Single-phase, Approximately 0.9kW - 5A
Maximum flow: No less than 500 l/min
Operating head for continuous service: Maximum 2700mm H2O
Average water consumption: Approximately 2 l/min
Piping pressure: Approximately 4 bar
Dimensions: Approximately 550mm W, 350mm D, 360mm H
42271907-00005 | RT286-10-015 | Suction Dry 1 Surgery |  |  | Each
SPECIFICATION A fixed installation centraised dry aspiration plant for serving 1 surgery. The unit MUST have the following features:
Fixed installation
Serve number of surgeries: 1
Output power: Single-phase, Approximately 0.4kW - 2A, 50 Hz
Speed of motor: No less than 2800 rpm
Vacuum setting at approximately 170 mBar
Noise level without housing: No more than 65dB(A)
42271907-00006 | RT286-10-016 | Suction Dry Mobile 1 Surgery |  |  | Each
SPECIFICATION A mobile suction unit for continuous utilisation in a dental surgery. The unit MUST have the following features:
Mobile unit to be steered around chair and between surgeries
Equipped with 4 double wheel castor, one per corner
Designed for continuous service
Capacity of canister: minimum 4 Litres
Canister to be emptied manually
Noise level: No more than 65dB(A)
Dimensions: Approximately 550mm W, 350mm D, 750mm H
Net weight: Approximately 35kg
Gross weight: Approximately 40kg
42152009-00000 | RT286-10-017 | Pan X-Ray Viewer |  |  | Each
SPECIFICATION An X-Ray viewer utilised for viewing panoramic and smaller size X-Ray photos. The unit shall have the following features:
Utilised as a Countertop or Wall-mounted item
Viewing area no smaller than 150 x 140mm
Even illumination to be supplied
Viewing panel of shatterproof plexiglass
Rustproof
Easy-to-clean surfaces for infection control
SABS Plug
42292202-00000 | RT286-10-018 | Compressor 1 Surgery - Oil free |  |  | Each
SPECIFICATION A fixed installation dental compressor with dryer (oil-free) for serving 1 surgery. The unit shall have the following features:
Serve number of surgeries: At least 1
Number of cylinders: Minimum 1
Number of dryers: At least 1
Output power: Approximately 1-50Hz, ± 0.5kW, ±3.5A
Air delivery: At least 66N l/min (pressure no less than 5 real bar)
Volume of air tank: No less than 30 Litres
Noise level: No more than 65dB(A)
Dimensions: Approximately 650mm W, 450mm D, 730mm H
Net weight: Approximately 45kg
Gross weight: Approximately 55kg
42292202-00001 | RT286-10-019 | Oil Free air compressor for 4-6 surgery (60DB) |  |  | Each
SPECIFICATION A fixed installation dental compressor with dryer (oil-free) for serving4 -6 surgeries. The unit shall have the following features:
Voltage 220/240 V 1,5 kW with built in thermo overload protection as per user requirement
Delivery capacity not less than 150 1/min maximum pressure 1000 kpa
Storage Capacity at least 35 litres - 75 litres
Fitted with ¼” safety release valve
Manual tank drain valve
Tank pressure gauge
Pressure range control switch, 6 – 10 bar
Noise level not more than 60 DB (A)
Either piston type or turbo type
Reverse pressure exhaust valve for motor protection
Air filter for filtration of 10 micron particles, with automatic drain valve
Oil free with automatic drianage system
42292202-00002 | RT286-10-020 | Oil Free air compressor for 2-4 surgery (60DB) |  |  | Each
SPECIFICATION A fixed installation dental compressor with dryer (oil-free) for serving 2-4 surgeries. The unit shall have the following features:
Voltage 220/240 V 1,5 kW with built in thermo overload protectionas per user requirement
Delivery capacity not less than 150 1/min maximum pressure 1000 kpa
Storage Capacity at least 35 litres - 75 litres
Power 0.75Kw 1HP
Fitted with ¼” safety release valve
Tank pressure gauge
Pressure range control switch, 6 – 10 bar
Noise level maximum 60 DB (A)
Either piston type or turbo type
Reverse pressure exhaust valve for motor protection
Air filter for filtration of 10 micron particles, with automatic drain valve
Oil free with automatic drainage system
42151636-00012 | RT286-10-021 | Ultrasound Scaler Multi-Function |  |  | Each
SPECIFICATION A multi-function ultrasound (piezon) scaler that can also be used specifically for endodontic and post removal. The unit shall have the following features:
Must be supplied with:
Control unit
Handpiece (Fibre Optic)
Handpiece Cord
Water supply tube with water filter
Foot Pedal
Tip Replacement wrench if applicable
Water filter
Unit shall include the following kits or instruments:
Kit for removing supra gingival calculus that include tips for
Work on large areas on all tooth surfaces
Tips for interproximal scaling
Flattened round tip for large areas of calculus
Tip for molar interproximal surfaces
Endodontic revision instrument kit with tips for:
Angled small spherical diamond coated (D3)
Angled instrument with working length up to 10mm with the diamond coated
Angled instrument with working length up to 20mm with the diamond coated
Angled instrument with working length up to 24mm
Instrument tip for crown, bridges and metal post removals
SABS Plug
42151636-00012 | RT286-10-022 | Ultrasound Piezo Scaler |  |  | Each
SPECIFICATION A multi-function ultrasound (piezon) scaler. The unit shall have the following features:
Hand piece
Control unit
Handpiece Cord
Water supply tube with water filter
Foot Pedal
Tip Replacement wrench if applicable
Water filter
Unit shall include the following kits or instruments:
Kit for removing supra gingival calculus that include tips for
Work on large areas on all tooth surfaces
Tips for interproximal scaling
Flattened round tip for large areas of calculus
Tip for molar interproximal surfaces
SABS Plug
42151636-00012 | RT286-10-023 | Ultrasound Scaler Multi-Function |  |  | Each
SPECIFICATION A multi-function ultrasound (piezon) scaler. The unit shall have the following features:
Must be supplied with:
Control unit
Handpiece Cord
Water supply tube with water filter
Foot Pedal
Tip Replacement wrench if applicable
Water filter
Unit shall include the following kits or instruments:
Kit for removing supra gingival calculus that include tips for
Work on large areas on all tooth surfaces
Tips for interproximal scaling
Flattened round tip for large areas of calculus
Tip for molar interproximal surfaces
42151636-00013 | RT286-10-024 | Dental Ultrasonic Scaler with a polishing function (Insert 30K) |  |  | Each
SPECIFICATION A multi-function ultrasonic scaler wit a polishing capability. The unit shall have the following features:
Portable Ultra compact body
Simple and intuitive operation control panel
Water Adjustment
Inserts washabble and autoclavable up to 135ºC
Control unit with hand piece cord and holder
Water filter set
Water Supply connector
Foot control
At least 3 scaling inserts 30K
Must be supplied with a quick connector (Male & Female)
Must be able to perform air polishing
SABS plug
42151636-00014 | RT286-10-025 | Dental Ultrasonic Scaler (Inserts 25K) |  |  | Each
SPECIFICATION A multi-function ultrasonic scaler. The unit shall have the following features:
Portable Ultra compact body
Simple and intuitive operation control panel
Water adjustment switch
Inserts washable and autoclavable up to 135ºC
Control unit wih handpiece cord and Holder
Water filter set
Water supply connector
Foot control
At least 3 scaling insert (25K)
Must be supplied with a quick connector
SABS Plug
42151636-00015 | RT286-10-026 | Dental Air Scaler |  |  | Each
SPECIFICATION A handpiece air scaler to fit on a dental unit:
Delivery unit mountable scaler
Handpiece washable and autoclavable up to 135ºC
At least 3 scaling tips
Tip wrench with Torque limiter
Compitable wih 4/6 hole tubbing
Including accesories
42281508-00006 | RT286-10-027 | Desktop Autoclave S-Type for Dental Instruments |  |  | Each
SPECIFICATION A S-type, casette loading desktop autoclave for sterilising wrapped dental instruments. The unit shall have the following features:
Must comply with ISO Standard (PREN 13060) for S-Type autoclaves
Include water distiller
Include baskets and accessories for sterilising wrapped instruments
Must be a casette type loading autoclave
Casette mesurements no less than 48mm x 17mm x 8mm
Full sterilisation cycle in less than 18 minutes
Capacity minimum 7 litres
SABS Plug
The unit MUST measure chamber pressure with accuracy of ±3% of full scale.
Alarms
Water supply failure
Cycle failure
Safety Features
Shut down when opened.
Minimum capacity of 1.6 litre
42151703-00000 | RT286-10-028 | Mobile Dental Unit non fibre optic without hand pieces |  |  | Each
SPECIFICATION A Complete mobile dental delivery system, non fiberoptic function without hand pieces
Dental cabinet houlsing the following:
3 Handpiece outlets ( 4 - 6 Holes)
3 in1 syringe
Selector switch
Foot control
Master shut off assembly
Compressor (25L - 50L; ≥100L/Min)
Cleanwater system
The unit must operate on 240VAC 5060Hz fitted with 3 prongs plug
Unit to be fitted with smooth running four swivel lockable wheels
Height = 700MM - 900MM Width = 400MM - 500MM Depth = 480MM - 600MM
42151703-00001 | RT286-10-029 | Mobile Dental Unit non fibre optic with hand pieces |  |  | Each
SPECIFICATION A Complete mobile dental delivery sytem, non fiberoptic function witht hand pieces
Dental cabinet housing the following
3 Handpiece outlets ( 4 - 6 Holes)
3 in 1 Syringe
Selector switch
Foot control
Master shut off assembly
Compressor (25L - 50L; ≥100L/Min)
Cleanwater system
Self illumination LED Handpiece with coupling
Internal spray straight handpiece
Contra angle blue handpiece
Prophylaxis handpiece
Air motor 4-hole
The unit must operate on 240VAC 50/60Hz fitted with 3-prongs plug
Unit to be fitted with smooth running four swivel lockable wheels
Height = 700MM - 900MM Width = 400MM - 500MM Depth = 480MM - 600MM
42151703-00002 | RT286-10-030 | Mobile Dental Unit fibre optic with hand pieces |  |  | Each
SPECIFICATION A Complete mobile dental delivery sytem with fiberoptic capeability function comlete with hand pieces
Dental cabinet housing the following:
3 Handpiece outlets (4-6 holes)
3 in 1 Syringe
Selector switch
Foot control
Master shut off assembly
Compressor (25L - 50L; ≥100L/Min)
Cleanwater system
Self illumination handpiece with coupling
Internal spray straight handpiece
Contra angle blue handpiece
Prophylaxis Handpiece
Air motor 4-hole
The unit must operate on 240VAC 50/60Hz fitted with 3-prongs plug
The unit must operate on 240VAC 50/60Hz fitted with 3-prongs plug
The unit to be fitted with smooth running four swivel lockable wheels
Height = 700MM - 900MM Width = 400MM - 500MM Depth = 480MM - 600MM
42281508-00007 | RT286-10-031 | Dental Autoclave class full range |  |  | Each
SPECIFICATION A Table top s class autoclave with a short cycle availble not exceeding 20 min for quicl sterlizing
Table top steam cassette autoclave.
It must include a water distiller
It MUST have programmable cycles : Wrapped & Unwrapped automatic.
Temperature range : 100- 134⁰ with a heat shutoff function.
The controls MUST be sealed against fluid penetration.
Must Include Baskets and accessories for sterilising wrapped instruments
The unit MUST have chamber temperature monitoring, accurate to ± 1⁰C over the operating range.
Must be able to complete a full cycle in less than 18 minutes
The unit MUST measure chamber pressure with accuracy of ±3% of full scale.
Alarms
Water supply failure
Cycle failure
Safety Features
Shut down when opened.
Minimum capacity of 1.6 litre
SABS Plug
42281508-00008 | RT286-10-032 | Dental Autoclave B-Class |  |  | Each
SPECIFICATION A Table top B class autoclave without a PRION Cycle
Mains voltage MUST operate at 220 – 240V.
Current MUST not exceed 15 Amps
MUST have no re-circulation of used water.
Chamber MUST be constructed of stainless steel or copper brass alloy. MUST be ASME or SABS approved.
The unit must have ability to heat the entire chamber efficiently.
MUST have heat insulated casing with built-in condensation coils and tank.
MUST have automatic timer to complete sterilizing cycle.
MUST be fitted with safety pressure valve.
MUST be fitted with temperature and pressure gauges.
Complete cycle after warm up MUST not exceed 30 min.
Chamber dimensions MUST be ±510 mm deep by ±360mm diameter.
Operating pressure MUST be 240 Kpa.
Operating temperature MUST not exceed – 135°C.
Minimum reservoir capacity 4 litres.
16 Amp plug MUST be supplied.
MUST have separate drying cycles which shuts off automatically – 7 – 12 minutes.
The unit MUST be of a horizontal, chamber-type autoclave, and the entire unit MUST be microprocessor controlled.
The unit MUST monitor the water level, and if low, switch off to protect the elements from overheating and damage. The unit MUST switch off when water level low, and not when overheating.
The unit MUST include at least 3 perforated trays with a detacheable handle.
Each tray MUST hold at least 10 dental instruments. The chamber MUST be long enough to sterilize instruments up to 300mm long
All cycles MUST be automatic – no user intervention MUST be required, except to select the correct cycle and start the cycle
End of cycle, and all error messages, MUST be indicated by a clear audible signal
It MUST be possible to terminate any cycle at any stage.
The unit MUST display the operating temperature in an electronic format, and also display pressure.
The unit MUST be automatically controlled by temperature. Control for pressure as an additional measure will serve as a recommendation.
The unit MUST automatically protect itself and the operator against over temperature and overpressure
The unit MUST be able to perform the following cycle:
134º with vacuum with drying
134º with vacuum without drying
134º non-vacuum with drying
134º non-vacuum without drying
121º with vacuum with drying
121º with vacuum without drying
121º non-vacuum with drying
121º non-vacuum without drying
Please provide cycle times for each of the programs for hot and cold cycle
The units MUST be able to sterilize unwrapped solids and hollow instruments, pouched solid and hollow instruments, single wrapped solid and hollow instruments, double wrapped solid and hollow instruments and porous items
The unit must also store data of at least the last 40 cycles. This data must be available in a printable and downloadable format. The unit MUST be able to connect to a computer, via a USB connection, to download the required information. If any software program is needed to perform this, it must be included in the bid price as an accessory.
Capacity of the chamber MUST be 16+ litres.
Water MUST be drained from the chamber after each and every cycle
External elements must be used, please state the wattage of the elements
The unit MUST have an insulated shell to protect operator from accidental heat injuries
Overpressure valve and all other safety valves must vent into a safe area, and not into working area
Unit must monitor the water quality in chamber, and stop cycle if the quality is not satisfactory
Reservoir and chamber to have a manual drain facility that is controlled with a positive on/off valve
Operator controls MUST be by touch-type, and not switches for infection control purposes
All electrical boards/electronic components MUST be well shielded/insulated from heat and potential water spillage
The door mechanism MUST have an auto-lock safety facility to prevent the door being opened during a cycle
It MUST not be possible to start a cycle if the door is not properly closed and sealed
Water distiller/demineraliser to be included for each autoclave. This must meet the required water quality for this autoclave. The price of this unit must be included in the bid price
The autoclave is to display time lapsed and the time remaining on cycle
All error codes to be supplied in digital format
The entire unit must form a self-contained working unit. It must be easy to move the unit to where it’s required
Unit must be supplied with a spare door seal
SABS PLUG
42281508-00009 | RT286-10-033 | Dental Autoclave B-Class with prion cycle |  |  | Each
SPECIFICATION A Table top B class autoclave with a PRION Cycle
Mains voltage MUST operate at 220 – 240V.
Current MUST not exceed 15 Amps
MUST have no re-circulation of used water.
Chamber MUST be constructed of stainless steel or copper brass alloy. MUST be ASME or SABS approved.
The unit must have ability to heat the entire chamber efficiently.
MUST have heat insulated casing with built-in condensation coils and tank.
MUST have automatic timer to complete sterilizing cycle.
MUST be fitted with safety pressure valve.
MUST be fitted with temperature and pressure gauges.
Complete cycle after warm up MUST not exceed 30 min.
Chamber dimensions MUST be ±510 mm deep by ±360mm diameter.
Operating pressure MUST be 240 Kpa.
Operating temperature MUST not exceed – 135°C.
Minimum reservoir capacity 4 litres.
16 Amp plug MUST be supplied.
MUST have separate drying cycles which shuts off automatically – 7 – 12 minutes.
The unit MUST be of a horizontal, chamber-type autoclave, and the entire unit MUST be microprocessor controlled.
The unit MUST monitor the water level, and if low, switch off to protect the elements from overheating and damage. The unit MUST switch off when water level low, and not when overheating.
The unit MUST include at least 3 perforated trays with a detacheable handle.
Each tray MUST hold at least 10 dental instruments. The chamber MUST be long enough to sterilize instruments up to 300mm long
All cycles MUST be automatic – no user intervention MUST be required, except to select the correct cycle and start the cycle
End of cycle, and all error messages, MUST be indicated by a clear audible signal
It MUST be possible to terminate any cycle at any stage.
The unit MUST display the operating temperature in an electronic format, and also display pressure.
The unit MUST be automatically controlled by temperature. Control for pressure as an additional measure will serve as a recommendation.
The unit MUST automatically protect itself and the operator against over temperature and overpressure
The unit MUST be able to perform the following cycle:
134º with vacuum with drying
134º with vacuum without drying
134º non-vacuum with drying
134º non-vacuum without drying
121º with vacuum with drying
121º with vacuum without drying
121º non-vacuum with drying
121º non-vacuum without drying
Unit MUST have a Prion cycle
Please provide cycle times for each of the programs for hot and cold cycle
The units MUST be able to sterilize unwrapped solids and hollow instruments, pouched solid and hollow instruments, single wrapped solid and hollow instruments, double wrapped solid and hollow instruments and porous items
The unit must also store data of at least the last 40 cycles. This data must be available in a printable and downloadable format. The unit MUST be able to connect to a computer, via a USB connection, to download the required information. If any software program is needed to perform this, it must be included in the bid price as an accessory.
Capacity of the chamber MUST be 16+ litres.
Water MUST be drained from the chamber after each and every cycle
External elements must be used, please state the wattage of the elements
The unit MUST have an insulated shell to protect operator from accidental heat injuries
Overpressure valve and all other safety valves must vent into a safe area, and not into working area
Unit must monitor the water quality in chamber, and stop cycle if the quality is not satisfactory
Reservoir and chamber to have a manual drain facility that is controlled with a positive on/off valve
Operator controls MUST be by touch-type, and not switches for infection control purposes
All electrical boards/electronic components MUST be well shielded/insulated from heat and potential water spillage
The door mechanism MUST have an auto-lock safety facility to prevent the door being opened during a cycle
It MUST not be possible to start a cycle if the door is not properly closed and sealed
Water distiller/demineraliser to be included for each autoclave. This must meet the required water quality for this autoclave. The price of this unit must be included in the bid price
The autoclave is to display time lapsed and the time remaining on cycle
All error codes to be supplied in digital format
The entire unit must form a self-contained working unit. It must be easy to move the unit to where it’s required
Unit must be supplied with a spare door seal
42151703-00003 | RT286-10-034 | Portable dental unit |  |  | Each
SPECIFICATION The unit is a fully contained mobile dental apparatus for point of use applications for oral health treatment in community clinics and hospitals. With its compressor, gas tank pressure release filter valve, clean water bottle, saliva gathering bottle, on/off foot switch, it provides a reliable steady dry and uncontaminated air source available for high- speed turbine hand piece. The unit is portable, simply operated, easy to maintain and use widely by dentist, Oral Hygienists, Dental Therapists and doctors.
Easily transported with pull-out handle and wheels.
Weighs 15 to 30kg depending on configuration.
220V compatible, oil less compressor (580W)
7 L tank capacity
Standard Features
Midwest 4 hole hose x2
3- way Syringe with hose
Saliva ejector hose x1
Oil free compressor x1
Portable unit must include the following extras:
1 Piezo ultrasonic scaler.
2 LED curing light.
Fast Hand Piece
Hi speed hand piece 32 PB Borden /Midwest type connection.
Triple spray distribution symmetrical in the direction of the bur tip.
Rotation up to 420000 rpm
Rotors accurately balanced.
Working air pressure between 32 and 35 means 3-bars to 3.5
Slow Hand Piece/ Contra-Angle
Single External Spray:
A silicone hose external to the handpiece is used for passage of the water from the body of the micro motor to the end of the handpiece used.
Rotation:
Maximum of 20,000 RPM
Transmission 1:1.
Ideal Weight:
Nett weight under 71.0 grams
Low work noise
Air Motor
Straight hand piece
42281508-00010 | RT286-10-035 | Autoclave for handpieces |  |  | Each
SPECIFICATION An Autoclave for different dental handpieces with an automatic oiling feuture
The apparatus must be able to clean in all internal channels of the handpieces, contra angles and turbines internally as well as externally.
Must be able to lubricate (oil) the handpieces after internal cleaning but before external hot cleaning and before the sterilization process started.
Must have a drying process.
Must have choice of sterilization programmes that can cater for quick processing (3 minutes) to lower heat at 121 degrees Celsius up to 20 minutes.
Must be a low noise autoclave
Must at least conform to the European or American standard.
A build-in water conductivity sensor that measures the water quality that activates if the quality deteriorates.
The unit must include water distiller/demineraliser.
Alarms
Water supply failure
Cycle failure
Safety Features
Shut down when opened.
SABS Plug
42152003-00000 | RT286-10-036 | Automatic Intra Oral X-Ray developer |  |  | Each
SPECIFICATION An Autonatic intra oral x-ray developer
An automatic processor for all sizes of intra oral –rays (up to 65mm Films)
Processing time approximately 2 min(Wet endodontic) and 4 min(Dry).
Must have a drier
Daylight loader
SABS Plug
42295011-00000 | RT286-10-037 | Digital Intraoral video camera system |  |  | Each
An inra-oral video camara for diagnostic and education purpouses
SPECIFICATION A direct intraoral Video camera shall consist of the following components: and/or the following features:
The software must be compatible with existing software of the user/requestor.
Must be equipped with a high resolution and at least a 3 to 4 point LED light
LED light must be integrated in the handpiece
At least a 80-900 single optic system (Intra and extra Oral)
The camera must
(a) Have intraoral and extraoral views
(b) Image freezing and saving
(c) Direct and mirror vision
(d) Be used on a computer or video monitor.
42152016-00000 | RT286-10-038 | Digital Intraoral imaging plate system with computer |  |  | Each
SPECIFICATION A X-ray sytem utilising plates for quick digilising of intra oral immages. The unit MUST have the following features:
Must be compatible with AC and DC intra oral X-ray Units.
Software including implant module and implant archive. .
Fast Ethernet connection. Must include all cabling and installation from the unit to the surgeries.
4 X imaging plate sets per unit (1 for every surgery) MUST be supplies and includes (per surgery):
a. 2 x 22x33mm plates.
b. 2 X 24x40mm plates.
c. 5 X 31x41 mm Plates
d. 1 X 27x45mm plate
e. 1 X storage box for plates.
f. 100 pieces of hygiene bags per size
The unit must be provided with an up to date desktop/laptop computer with the following requirements:
≥4GB Ram
≥1TB Hard drive
17 inch LED screen
Software for processing all captured images
A keyboard and mouse (cordless/cord)
The unit MUST use a line voltage of 220VAC to 240VAC
The line frequency MUST be 50Hz
at least windows 10
The unit must be free standing with height adjustment
SABS Plug (2 fase)
42152016-00001 | RT286-10-039 | Digital Intraoral imaging plate system without computer |  |  | Each
SPECIFICATION An Intra-oral X-ray system utilising plates for quick digilising of x-ray images. The unit MUST have the following features:
Must be compatible with AC and DC intra oral X-ray Units.
Software including implant module and implant archive. .
Fast Ethernet connection. Must include all cabling and installation from the unit to the surgeries.
4 imaging plate sets per unit (1 for every surgery) MUST be supplies and includes (per surgery):
a. 2 x 22x33mm plates.
b. 2 X 24x40mm plates.
c. 5 X 31x41 mm Plates
d. 1 X 27x45mm plate
e. 1 X storage box for plates.
f. 100 pieces of hygiene bags per size
SABS Plug
42152017-00000 | RT286-10-040 | Cone Beam CT 3D Dental imaging system (small field of view- SFOV) |  |  | Each
SPECIFICATION A Cone Beam CT 3D Dental Imaging System for 3 dimetion image at all planes in vieving field
At least a 8cm X 8cm reconstruction volume.
Standing or seated (wheelchair accessible) Patient position and able to accommodate children.
Focal spot of 0.5 minimum.
X-ray Source – High Frequency, constant potential (DC).
Image detector silicon flat panel to be appropriate with the 8cm X 8cm.
360 degree rotation for Image Acquisition.
Automatically adjust radiation dosage to patient’s age and size.
Must include scanner, cabinet with electronic circuitry.
A Computer Workstation with all the software.
Software must be able to be programmed for Southern Implants.
Must include a computer with LED flat screen to runt the system.
Must be able to handle multiple reconstructions from the same scan.
Software must be able to be converted to DICOM format.
Must be able to produce TMJ images: sagittal, panoramic sagittal and coronal reconstruction.
Must cover the entire range of applications from single implant to complete ENT and Orthogenetic planning and diagnose.
Software must make provision for accurate implant placement and implant guide planning
Must be able to accommodate standing and sitting patients
SABS Pug (2 Phase)
42152017-00001 | RT286-10-041 | Cone Beam CT 3D Dental imaging system(big field of view-FOV) |  |  | Each
SPECIFICATION A Cone Beam CT 3D Dental Imaging System for 3 dimetion image at all planes in vieving field
At least a 15.5cm X 15.5cm reconstruction volume.
Standing or seated (wheelchair accessible) Patient position and able to accommodate children.
Focal spot of 0.5 minimum.
X-ray Source – High Frequency, constant potential (DC).
Image detector silicon flat/cone panel, at least 20cm X 25 cm (field of view).
360 degree rotation for Image Acquisition.
Automatically adjust radiation dosage to patient’s age and size.
Must include scanner, cabinet with electronic circuitry.
A Computer Workstation with all the software.
Software must be able to be programmed for Implants.
Must include a computer with LED flat screen to runt the system.
Must be able to handle multiple recon-structions from the same scan.
Software must be able to be converted to DICOM format.
Must be able to produce TMJ images: sagittal, panoramic sagittal and coronal reconstruction.
Must cover the entire range of applications from single implant to complete ENT and Orthogenetic planning and diagnose.
Software must make provision for accurate implant placement and implant guide planning
Must be able to accommodate standing and sitting patients
SABS Plug (2 Phase)
42152008-00003 | RT286-10-042 | Dental digital panoramic X-Ray |  |  | Each
SPECIFICATION An X-ray unit capable of taking didgital panoramic immages of the mouth to include sinus, TMG and whole jaw
The unit must be able to cater for adult and paediatric patients
The tube output voltage must be between 60kV to 90kV
Focal dimensions must be 0.5mm
The unit must be capable for the following applications:
a. Panoramic
b. Temporomandibular Joint (TMJ) Imaging
The unit must have a maximum current rating of 15mA
The maximum exposure time must be 15 seconds
The unit must be provided with an up to date desktop/laptop computer with the following requirements:
a. ≥4GB Ram
b. ≥1TB Hard drive
c. 17 inch LED screen 2 Mega pixel
d. Software for processing all captured images
e. A keyboard
f. A mouse with a cord/ cordless
The unit MUST use a line voltage of 220VAC to 240VAC
The line frequency MUST be 50Hz
The unit must be free standing with height adjustment
UPS
Dicom system compitable
SABS Plug (2 -phase)
42152008-00004 | RT286-10-043 | Digital Radiography system-sensor without a computer |  |  | Each
SPECIFICATION A Sensor for immediate capturing of digitised X-ray intraoral images
Sensor with USB 2 connection
Mount for sensor storage
Cable length 3 meters / cordless
Hygienic sheaths X50
Twain Compatibility - Sensor can be integrated with any imaging and dental practice management software
Sensor positioning kit
The unit MUST use a line voltage of 220VAC to 240VAC
The line frequency MUST be 50Hz
Software including implant module and implant archive.
42152008-00005 | RT286-10-044 | Digital Radiography system-sensor with a computer |  |  | Each
SPECIFICATION A Sensor for immediate capturing of digitised X-ray intraoral images
Sensor with USB 2 connection
Mount for sensor storage
Cable length 3 meters / cordless
Hygienic sheaths X50
Twain Compatibility - Sensor can be integrated with any imaging and dental practice management software
Sensor positioning kit
The unit must be provided with an up to date desktop/laptop computer with the following requirements:
≥4GB Ram
≥1TB Hard drive
17 inch LCD screen
Software for processing all captured images
A keyboard
The unit MUST use a line voltage of 220VAC to 240VAC
The line frequency MUST be 50Hz
The unit must be free standing with height adjustment
Software including implant module and implant archive.
SABS Plug
42152008-00006 | RT286-10-045 | Intra Oral Dental X-ray (Analogue & Digital) - Mobile |  |  | Each
SPECIFICATION A X-ray unit for capturing intra oral analgue and digital images (WALL Mouted)
The unit must be a mobile type (with an option of a long arm)
The tube head MUST be between 60kV to 70kV
The unit must have a computerized timer with the following functions:
Automatic setting of technique factors in mAs
Object program selection
Tooth type
Patient size
Manual override
Hand switch with 3m cord or cordless
The focal point of the unit MUST be from 0.4mm to 0.8mm
The power supply of the unit must be mounted on the unit
The unit MUST use a line voltage of 220VAC to 240VAC
The line frequency MUST be 50Hz
Analogue and digital
SABS Plug (2 fase)
42152008-00007 | RT286-10-046 | Intra Oral Dental X-Ray (Analogue & Digital) Floor Mounted |  |  | Each
SPECIFICATION A X-ray unit for capturing intra oral analgue and digital images (Floor Mouted)
The unit must be a floor mount type (with an option of a long arm)
The tube head MUST be between 60kV to 70kV
The unit must have a computerized timer with the following functions:
Automatic setting of technique factors in mAs
Object program selection
Tooth type
Patient size
Manual override
Hand switch with 3m cord or cordless
The focal point of the unit MUST be from 0.4mm to 0.8mm
The power supply of the unit must be mounted on the unit
The unit MUST use a line voltage of 220VAC to 240VAC
The line frequency MUST be 50Hz
Analogue and digital
SABS Plug (2 Fase)
42152008-00008 | RT286-10-047 | Intra Oral Dental X-Ray Duarl Function (Analogue & Digital) wall Mounted |  |  | Each
SPECIFICATION A X-ray unit for capturing intra oral analgue and digital images (Floor and wall mouted options)
The unit must be a wall mount type (with an option of a long arm)
The tube head MUST be between 60kV to 70kV
The unit must have a computerized timer with the following functions:
Automatic setting of technique factors in mAs
Object program selection
Tooth type
Patient size
Manual override
Hand switch with 3m cord or cordless
The focal point of the unit MUST be from 0.4mm to 0.8mm
The power supply of the unit must be mounted on the unit
The unit MUST use a line voltage of 220VAC to 240VAC
The line frequency MUST be 50Hz
Analogue and digital
SABS Plug (2 Fase)
42152008-00009 | RT286-10-048 | Dental Digital panoramic X-Ray with cephalometric without a printer |  |  | Each
SPECIFICATION A Panoramic X-ray with the capability to cephalometric photos comlete without a printer
The unit must be able to cater for adult and paediatric patients
The tube output voltage must be between 60kV to 90kV
Focal dimensions must be 0.5mm
The unit must be capable for the following applications:
.Panoramic
.Temporomandibular Joint (TMJ) Imaging
.Cephalometric
The unit must have a maximum current rating of 15mA
The maximum exposure time must be 15 seconds
The unit must be provided with an up to date desktop computer with the following requirements:
.≥4GB Ram
.≥1TB Hard drive
.17 inch LED screen 2 mega pixel
.Software for processing all captured images and dicom campatible
.A keyboard
The unit must use a line voltage of 220VAC to 240VAC
The line frequency must be 50Hz
The unit must be free standing with height adjustment
UPS
SABS Plug (2 Fase)
42152008-00010 | RT286-10-049 | Dental Digital panoramic X-Ray with cephalometric with a printer |  |  | Each
SPECIFICATION A Panoramic X-ray with the capability to cephalometric photos comlete with a printer
The unit must be able to cater for adult and paediatric patients
The tube output voltage must be between 60kV to 90kV
Focal dimensions must be 0.5mm
The unit must be capable for the following applications:
Panoramic
Temporomandibular Joint (TMJ) Imaging
Cephalometric
The unit must have a maximum current rating of 15mA
The maximum exposure time must be 15 seconds
The unit must be provided with an up to date desktop computer with the following requirements:
≥4GB Ram
≥1TB Hard drive
17 inch LED screen 2 mega pixel
Software for processing all captured images and dicom campatible
A keyboard
The unit must use a line voltage of 220VAC to 240VAC
The line frequency must be 50Hz
The unit must be free standing with height adjustment
Printer (capable to print panoramic as well as cephalometric images)
UPS
SABS Plug (2- Phase)
42151624-00014 | RT286-10-050 | Dental Handpieces |  |  | Each
TURBINE FULL RANGE
SPECIFICATION To fit 4-6 hole standard supply tubing.
Speed 350 000 rpm minimum at 2.2 kg per cm.
Ballbearing, output approximately 10 watts
Field replaceable rotor, one spare rotor to be supplied
Metal bur change chuck, pushtype system
Quick action coupling for easy lubrication
Guarantee time of turbine must be indicated – minimum 12 months
Air consumption MUST exceed 36 1/m
Noise level not to exceed 75 db
Autoclavable up to 135°C
With fibre optic function
To come with durable metal casing (stainless steel)
42151624-00015 | RT286-10-051 | Turbine full range |  |  | Each
SPECIFICATION To fit 4-6 hole standard supply tubing.
Speed 350 000 rpm minimum at 2.2 kg per cm.
Ballbearing, output approximately 10 watts
Field replaceable rotor, one spare rotor to be supplied
Metal bur change chuck, pushtype system
Quick action coupling for easy lubrication
Guarantee time of turbine must be indicated – minimum 12 months
Air consumption MUST exceed 36 1/m
Noise level not to exceed 75 db
Autoclavable up to 135°C
With non fibre optic function
To come with durable metal casing (stainless steel)
42151624-00016 | RT286-10-052 | Turbine full range |  |  | Each
SPECIFICATION To fit 4-6 hole standard supply tubing.
Speed 350 000 rpm minimum at 2.2 kg per cm.
Ballbearing, output approximately 10 watts
Field replaceable rotor, one spare rotor to be supplied
Metal bur change chuck, pushtype system
Quick action coupling for easy lubrication
Guarantee time of turbine must be indicated – minimum 12 months
Air consumption MUST exceed 36 1/m
Noise level not to exceed 75 db
Autoclavable up to 135°C
With fibre optic function
To come with durable metal casing (titanium)
42151624-00017 | RT286-10-053 | Turbine full range |  |  | Each
SPECIFICATION To fit 4-6 hole standard supply tubing.
Speed 350 000 rpm minimum at 2.2 kg per cm.
Ballbearing, output approximately 10 watts
Field replaceable rotor, one spare rotor to be supplied
Metal bur change chuck, pushtype system
Quick action coupling for easy lubrication
Guarantee time of turbine must be indicated – minimum 12 months
Air consumption MUST exceed 36 1/m
Noise level not to exceed 75 db
Autoclavable up to 135°C
With non fibre optic function
To come with durable metal casing (titanium)
42151624-00018 | RT286-10-054 | Led straight handpiece for micromotor - Air full range |  |  | Each
SPECIFICATION Standard 1 : 1 transmission
Ballbearing
To take burs Iso type 104 44, 50 mm with diameter of 2,35 mm
Maximum speed 40 000 rpm
E-type quick disconnection
Water spray incorporated internally
Autoclavable up to 135°C
Autoclavable up to 135°C
Quick action coupling for easy lubrication
To come with durable metal casing (titanium)
42151624-00019 | RT286-10-055 | Led straight handpiece for micromotor - Air full range |  |  | Each
SPECIFICATION Standard 1 : 1 transmission
Ballbearing
To take burs Iso type 104 44, 50 mm with diameter of 2,35 mm
Maximum speed 40 000 rpm
E-type quick disconnection
Water spray incorporated internally
Autoclavable up to 135°C
Autoclavable up to 135°C
Quick action coupling for easy lubrication
To come with durable metal casing (stainless steel)
42151624-00020 | RT286-10-056 | LED contra-angle handpiece for micromotor - Air full range |  |  | Each
SPECIFICATION Standard transmission 1 : 1 blue
Push button type system
Ballbearing type
Maximum speed 40 000 rpm
To accommodate ISO type 204 2.35 mm diameter latch type burs
E-type quick disconnection
Water spray incorporated internally
Autoclavable up to 135°C
Quick action coupling for easy lubrication
To come with durable metal casing (stainless steel)
42151624-00021 | RT286-10-057 | LED contra-angle handpiece for micromotor - Air full range |  |  | Each
SPECIFICATION Standard transmission 1 : 1 blue
Push button type system
Ballbearing type
Maximum speed 40 000 rpm
To accommodate ISO type 204 2.35 mm diameter latch type burs
E-type quick disconnection
Water spray incorporated internally
Autoclavable up to 135°C
Quick action coupling for easy lubrication
To come with durable metal casing (titanium)
42151624-00022 | RT286-10-058 | LED contra-angle handpiece for micromotor - Air full range |  |  | Each
SPECIFICATION Reduction hand piece Green
Push button type system
Ballbearing type
To accommodate ISO type 204 2.35 mm diameter latch type burs
E-type quick disconnection
Water spray incorporated internally
Autoclavable up to 135°C
Quick action coupling for easy lubrication
To come with durable metal casing (stainless steel)
42151624-00023 | RT286-10-059 | LED contra-angle handpiece for micromotor - Air full range |  |  | Each
SPECIFICATION Reduction hand piece Green
Push button type system
Ballbearing type
To accommodate ISO type 204 2.35 mm diameter latch type burs
E-type quick disconnection
Water spray incorporated internally
Autoclavable up to 135°C
Quick action coupling for easy lubrication
To come with durable metal casing (titanium)
42151624-00024 | RT286-10-060 | LED contra-angle handpiece for micromotor - Air full range |  |  | Each
SPECIFICATION Speed increasing red
Push button type system
Ballbearing type
To accommodate ISO type 204 2.35 mm diameter latch type burs
E-type quick disconnection
Water spray incorporated internally
Autoclavable up to 135°C
Quick action coupling for easy lubrication
To come with durable metal casing (stainless steel)
42151624-00025 | RT286-10-061 | LED contra-angle handpiece for micromotor - Air full range |  |  | Each
SPECIFICATION Speed increasing red
Push button type system
Ballbearing type
To accommodate ISO type 204 2.35 mm diameter latch type burs
E-type quick disconnection
Water spray incorporated internally
Autoclavable up to 135°C
Quick action coupling for easy lubrication
To come with durable metal casing (titanium)
42151624-00026 | RT286-10-062 | LED prophylaxis handpiece full range |  |  | Each
SPECIFICATION Reduction Green
To fit E-Type quick disconnection
Sealed ballbearing
To accept snap-on and or screw in cups and brushes
Autoclavable up to 135°C
Quick action coupling for easy lubrication
To come with durable metal casing (titanium)
42151624-00027 | RT286-10-063 | LED prophylaxis handpiece full range |  |  | Each
SPECIFICATION Reduction Green
To fit E-Type quick disconnection
Sealed ballbearing
To accept snap-on and or screw in cups and brushes
Autoclavable up to 135°C
Quick action coupling for easy lubrication
To come with durable metal casing (stainless steel)
42152910-00001 | RT286-10-064 | Micromotor - Air full range |  |  | Each
SPECIFICATION Speed range 5 000 – 40 000 rpm
Output 20W Power – min Ncm Torque min
E-Type quick coupling
Forward and reverse rotation
Weight not to exceed 90 gm. Speed control selector by means of control ring
Water spray/air facility to be incorporated internally
Maximum air consumption 45 1/m
Autoclavable up to135°C
Quick action coupling for easy lubrication
To come with durable metal casing (stainless steel)
42152910-00002 | RT286-10-065 | Electric Micro Motors ( stand alone) |  |  | Each
SPECIFICATION E Type fitting
40 000 rpm
Selection 16: 1 , 1: 5 and 1: 1
Fibre Optic
42151702-00000 | RT286-10-066 | Dental operator stool |  |  | Each
SPECIFICATION An ergonomic stool used by the operator
2-way adjustable backrest for lumbar support. Anatomical shaped.
Contoured seat cushions with a “rolled” edge to reduce pressure on the legs.
Easy to reach adjustment paddle for increase lumbar support.
Tear resistant upholstery
Seamless upholstery for easier cleaning.
Upholstary, design and colour to be in line with the patient operator chair
5 leg cast aluminium base
Twin wheel castors.
The unit must have height adjustable
The unit must have active and passive backrest.
42151705-00000 | RT286-10-067 | Dental delivery system |  |  | Each
SPECIFICATION A Dental delivery system, complete with tubing,spitoon, scaler, clean water system and curing light
Control block must have 4-chair position controls (including reset button)
The unit must be connected to the cuspidor/ spitoon/ water unit
Unit must have four/ six (4/6) Handpiece tube holes or more (all modes/functions) with self activations
Automatic selection of cutting H/P
3 in 1 Syringe with pinch valve technique and sterilizable sleeve and nozzle
The unit to have at least two brake handles (one on each side of delivery head)
To have touchpad for controls
Console with water flow control valve for each handpiece, if not incorporated in handpiece: must have a minimum of 4 tubes
Console with 4 control switches for the following:
a) on/off Handpiece water/air spray
b) on/off handpiece air spray coolant
c) on/off tumbler filler
d) on/off spitoon bowel rinse
Individual handpiece pressure adjustment reliable water retraction system required
The heads must be durable, providing many years of reliability
Free hanging straight/ whip arm air outlet hoses are preferred.
The clean water bottle must be standard (1 litre)
It must come with a piezo scaler system with foot switch
The curing light (LED with at least 3 settings levels) must also be supplied with the system
Master on/off switch provided at connection box
SABS PLUG (2 Phase)
42152306-00000 | RT286-10-068 | Dental LED Light |  |  | Each
SPECIFICATION An operator light for ilumination of work field
It MUST be LED light
Touch less sensor ON/OFF and manual control
Voltage 220 – 240, 50 Hz
Mounted to Chair / Ceiling
Rotational axes: Horizontal, vertical, diagonal
Constant colour temperature 5,000K
Temperature increase not more than 5 degrees C above room temperature at maximum intensity.
Light pattern: 85mm - 95mm x 145mm - 160mm at 700 mm
Adjustable light intensity: 15000Lux - 30000Lux
Shadow less illumination
On/Off switch in lamp yoke, and provision for altering light intensity, without changing colour temperature
Power Consumption: 12W - 17W or less
With Mode Selection
The light must have a shield that is scratch and discolouration resistant for protection against bulb shattering
42151702-00001 | RT286-10-069 | Dental assistant stool |  |  | Each
SPECIFICATION An ergonomic stool used by the Dental assistant
2-way adjustable backrest for lumbar support. Anatomical shaped.
Contoured seat cushions with a “rolled” edge to reduce pressure on the legs.
Easy to reach adjustment paddle for increase lumbar support.
Tear resistant upholstery
Seamless upholstery for easier cleaning.
Upholstary, design and colour to be in line with the patient operator chair
5 leg cast aluminium base
Twin wheel castors.
Adjustable torso support that can be converted for left or right handed operators.
Height adjustable foot ring).
The unit must have height adjustable
42152520-00000 | RT286-10-070 | Water Unit / Cuspidor / Spitton |  |  | Each
SPECIFICATION A chairside water and spitoon system
Chair-mounted either left or right
Spittoon must be removable without any tools, for easy cleaning
Ceramic bowls spittoon
Warm water tumbler filler
Spittoon flush facilities
Solids trap for spittoon bowl
Selective holder consisting of:
Solids collector
One HVE suction tube
One saliva ejector tube
Cup fill shut off valve incorporated
Low voltage remote control for suction.
Autoclavable HVE cannulas to be supplied
42151701-00000 | RT286-10-071 | Basic Dental chair |  |  | Each
SPECIFICATION A dental chair facility to do oprative procedures in the oral envioronment
Dental Chair
Anatomically contoured seat and backrest for comfortable full support in all positions suitable for treatment of both adults & children
The chair offered must have double articulating headrest which is reversable (cushion included)
Adjustable backrest - thickness 25mm - 50mm
Upholstery must be seamless durable waterproof material that can be easily cleaned and withstand disinfenctants and cleaning chemicals.
Upholstery must be removable.
Replaceable plastic cover must be provided to protect upholstery in shoe area
Armrest: Positions swing out/ down Arm
Chair must be Programmable
The chair must be floor mount
Accessory foot controls MUST be fitted to chair base to facilitate both height and trendelenburg adjustment
Range of lift height 342.90 mm to 800.10 mm
Recline range +62° to -12°
Total length at maximum extension 1820mm - 2000 mm
Base (H x W x D) : 600mm -650mm x 500mm - 560mm x 490mm - 510mm
Electro mechanical/hydraulic lift system fitted in chair base
The seat surface must be 410mm - 480mm x 400mm - 480mm (LxB)
Weight 200kg SAMHS tto provide (carry capacity of 200kg)
Chair swivel 60° - 70° (≤35° each side of center)
To operate on 240 VAC 50/60Hz, fitted with standard hospital 3-prongs plug.
42151701-00001 | RT286-10-072 | Dental treatment centre (Surgery unit) |  |  | Each
SPECIFICATION A dental surgery chair for either a L-hand or R-hand configuration. The unit must have the following features:
Chair mounted delivery system:
The unit must operate by means of a reliable pneumatic/electronic control block and single diaphragm
The unit must be centrally mounted to facilitate 270 degree movement front or back of the chair for left and right hand dentists. The facility must be immediate without the use of any technical assistance, plumbing and electrical plugs
The design of the unit must be seamless with rounded edges and the carcass in the aerosol zone should be seamless and smooth
The height adjustment must be manual and held firmly in position by a mechanism
The unit must have automatic shut–off controls for all the handpieces
Adjustment of water coolant must be provided for by precise needle valve flow controls knobs. These knobs must be easily be removed for chemical sterilisation.
The unit must be supplied with 3 sets of removable knobs
The controls of the unit must be positioned out of the aerosol zone
Water retraction must be prohibited on all handpieces by means of a reliable non retractable valve system
The dental unit must have an effective tubing and handpiece flushing system that allows increased water volume without running the handpieces
Individual air pressure adjustment must be provided for each handpiece an analogue gauge must be fitted to the dental unit to indicate working air pressure of all handpieces
A disposable cup holder for waste must be fitted to the dental unit
The dental unit must accommodate a standard instrument tray (Approximately size of 250 X 350 mm) and the tray must be easily removable but firmly kept in place
Two removable trays (Approximately size of 250 X 350 mm) must be provided
The dental unit must accommodate the following:
1 x 3 way syringe complete
2 x Fibre optic tubing’s
1 x Electric Motor
The 3 way syringe head and tip must be be quick disconnectable and autoclavable at 135 degrees Celsius
3 x Non-disposable tips must be supplied per syringe
The 3 way syringe head must be able to accommodate disposable tips
The syringe head must be non water retracting
Free hanging straight hoses (Tubing) must withstand at least 100PSI and proof of durability of tubing may be required
The hoses must be smooth for quick and effective wipe down
The power pack for the fibre optics must be incorporated into the dental unit
A programmable touch pad option must be available on the unit
Pneumatic/ electronic pedal type foot control for variable speed regulation of all handpieces to include water spray control and water On/Off switch
An arm control must be provided and all umbilicals must be concealed and protected in the arm
Load carrying capacity of the arm control must not be less than 5 Kg above the weight of the dental unit and must remain fixed
A connection box must be provided to contain air and water filters.
The connection box must supply services to the dental unit, the spittoon and chair without visible tubing
The power supply (220-240 V, 50Hz fused mains input) to be provided from the connection box to the dental unit, spittoon and chair
High performance operating LED lighting mounted and be able to accommodate Left and right handed dentists without any technical assistance, plumbing or electrical modifications
The power supply must be 220 -240 V, 50HZ
Temperature increase must not be more than 5 degrees Celsius above room temperature at a distance of 80 cm when the light is on maximum intensity
Light must have a varying intensity device
The light must be provided with an easy access On/Off switch on lamp head
Operating field distance of 70 centimetres must be approximately 5 X 20 CM with no glare for the patient
The light must have a shield that is scratch and discolouration resistant for protection against bulb shattering.
Light must be rounded for easy cleaning and increases asepsis
Lamp head to be adjustable in 125 degrees vertical and 180 degrees horizontal swivel
The light must move smoothly and all pivots must be permanently lubricated
The light arm must have a single rounded arm without transformer to support adjustable tension spring counter balance
Light handles must be able to accept barrier protection bags
An On/Off master shut-off must be located where itbis visible in the centre.
Cuspidor and assistant arm must have the following
The units must have the ability to be swivel for left en right hand use as must be possible without outside technical support, plumbing or electricity change
Spittoon bowl must be of an one removable piece, easy to clean design that allows easy access to inside of the housing
Bowl flush system and cup filler must be manually controlled from 2 sides
The bowl housing must be mounted to accommodate water cup and to prevent accidental slip of the cup
The bowl housing must have two separate facilities for cup filling and spittoon rinse
The spouts for rinsing and cup filling must be easy removable for autoclaving at 135 degrees Celsius and be constructed of a rust free material
The spittoon bows must provide with a solid trap
The dental assistant side central vacuum and water instrumentation be mounted on the cuspidor to be left and right interchangeable and must include:
Easy clean high volume evacuation
Saliva ejector x 1
Three in on syringe x 1
Easy clean high volume evacuating valve must have positive On/Off controls and be autoclavable at 135 degrees Celsius
The high volume autoclavable valve must be removable and quick disconnectable without the use of tools
High volume autoclavable valves be supplied per dental assistant’s station.
The saliva ejector must have a positive On/Off control and must have stop in full open and closed position and must be autoclavable at 135 degrees Celsius
Saliva ejector must be easy removable without the use of tools
wo saliva ejectors and one high volume suction with no rubber tips must be supplied per dental assistant station
3 x Non-disposable tips must be supplied per dental assistant’s station
The 3 way syringe head must be able to accommodate disposable tips
The syringe head must be non water retracting
Free hanging straight hoses (Tubing) must withstand at least 100PSI and proof of durability of tubing may be required
All suction hoses must be smooth and removable without any tools for easy disinfection and maintenance
Cuspidors must have a factory fitted clean water system with a bottle design that is easy to remove and to clean
Clean water system mus have a minimum capacity of l (one) litre
One additional bottle (in total two bottles of min 1 (one) litre capacity per bottle system ) must be supplied
Connection (umbilical) from cuspidor/spittoon to floor
Box must be supplied; and not lie on the floor
The dental chair must have the following features:
The seat and backrest must be anatomically contoured for full supporting all positions suitable for treatment of adults and children
The backrest must provide adequate support to elbows and lower lumber part of the body, when the backrest is fully declined
Backrest with build in Trendelburg mechanism must have an extendable double articulating anatomically shaped headrest with suitable adjustment for children
The headrest must be easily removable without the use of any tools
The headrest must be easily be reversed to allow for operating on patients outside the chair (for example retarded patient in a wheelchair)
Backrest of the upper section not be thicker than 65 mm and must be rounded in shape for easy access with no obstructions to the operator
Armrest must be fixed and must be movable with bushing pivot points to swivel out of the way
The armrest must not interfere with the operator
The minimum height of the seat must not be more than 450mm in the completely down position
The seat must support a pressure or any load per square inch of a minimum of 100Kg in any area of the seat
The stress points of the seat and backrest must be supported by cast aluminium or metal
The seat and backrest must be upholstered in a durable, washable and cool breathing material with softly rounded upholster edges with o exposed seems at patient contact
The upholster for the seat and backrest must be replaceable in the dental surgery without the use of outside technical assistance with minimal downtime
Replaceable plastic cover must be provided to protect the upholstery in the shoe area and must be easily be changed and fastened under the chair by means of Velcro
The base plate must not interfere with the operator
The back of the chair must have the facility to allow for left and right rotating dental unit in order to ensure comfortable operation for either left or right handed operator. This facility must provide a path for eclectic wiring and must form an integrated part of the complete system
Chair movement must be by means of low voltage switching and foot control must facilitate movement
The foot control must be fitted on the foot control to facilitate manual settings
Adjustment of the foot control memory settings must be without the use of tools and must be user friendly
The foot control must be of an one piece of s durable material like rubber moulded silicon
The foot control must not be fitted to the base of the chair
The foot control must be stable on the floor areas and must not slip away when operated
A lift system must be completely sealed and be able to lift a capacity of minimum 200Kg
The chair movement rate facilitated by the lift system must be adjustable for a minimum of four chair positions, i.e. base up, base down, back up, and back down
The stationary portion of the base must house an integral electric motor
The reservoir must be transparent so that the fluid can be manually inspected and not be pressurised
Fittings and tubings used in the low pressure portion of the system must be attached at all points with barbed fitting and retained upon the barb by a retaining clamp
The stationary part of the base must provide for easy access to mechanical components by removal of a formed secured housing
The system excludes hand-pieces, suction units and compressors
The unit must have safety machanism to stop the movement of the chair when it hits an objecet
The unit must meet the relevant international safety standards
SABS Plug (2 Phase)
42151701-00002 | RT286-10-073 | Portabe Dental chair |  |  | Each
SPECIFICATION A Portable chair to be abled to be carried by an operator
The unit must be able to carry up to 200kg
It must have medium alloy construction
The seat must adjustable from 35 to 60cm
The unit must be able to be used as a sit up chair
The maximum weight must not be more than 12.5kg
The dimensions must not be more than 510 x 270 x1000mm
It must come complete with a back pack
The unit must make provision for a portable dental light, spit funnel and tray
42152434-00000 | RT286-11-001 | Digital Amalgamator | Equipment |  | Each
SPECIFICATION Noise level less than 65BD
Automatic swich off-electronic time control
Operation cycle up to 99S
Body Shock resistant plastoc body
The unit should operate on 240VAC 50/60Hz fitted with 3-prongs plug
Unit to have a sealed and isolated mixing chamber which deactivates on opeing
Replaceable practical and durable capsule holder (both amalgam and glass ionomer capsule)
SABS Plug
42152309-00000 | RT286-11-002 | Corded Laser Caries Detector (Tabletop) |  |  | Each
SPECIFICATION A laser caries detection aid that measures photological changes in tooth structure caused by caries.
Must be able to be used for pit-and -fissure areas, even when outer surface is intact.
LED light integrated in the handpiece
Uses light energy with no danger to patient to operator.
Standard cord connection between handpiece and control unit
Must be portable.
Initial delivery must include the following:
1 X Fiber -optic tubing.
1 X Unit
2 X tips
1 X sterilization cassette.
1 X autoclavable handpiece sheath
1 X calibration disk
1 X operating booklet
Be able to accommodate a perio type tip for detecting subgingival caries
SABS plug
Optional items
| RT286-11-002 A | Replacement tip |  |  | Each
| RT286-11-002 B | Replacement safety glasses |  |  | Each
42153001-00000 | RT286-11-003 | Single-Tooth Anesthesia System |  |  | Each
SPECIFICATION A single tooth anesthesia system that controls the flow and pressure of dental anesthesia. Must be able to be used for all dental application
Must be portable.
Initial delivery must include the following:
1 X Drive Unit
1 X foot control pedal
1 X instruction manual (special conditions)
Box of 50 X handpieces with needles
SABS plug
Optional items
| RT286-11-003 A | Box of 50 X handpieces with needles |  |  | Each
42151655-00000 | RT286-11-004 | Pulp Tester |  |  | Each
SPECIFICATION A pulp tester for testing the vitality of teeth.
Must be able to be used on anterior and posterior teeth.
Must be portable
Digital display
Offers different rates of stimulus
Rechargeable batteries
Initial delivery must include the following:
1 X Pulp tester
Different sizes of autoclavable tip
Optional items
| RT286-11-004 A | Batteries |  |  | Each
| RT286-11-004 B | Tips |  |  | Each
42151691-00000 | RT286-11-005 | Apex Locator |  |  | Each
SPECIFICATION An apex locator for endodontic procedures
The product must have multi frequency technology being able to work in wet and dry canals
Must be portable
Readable display indicating different lengths of the root canal treatment
Initial delivery must include the following:
Rechargeable battery
Battery charger
1 X Drive Unit
2 X Lip hooks
Measuring wire
Fork (Endo file holder)
Optional Items
| RT286-11-005 A | Rechargeable battery |  |  | Each
| RT286-11-005 B | Battery charger |  |  | Each
| RT286-11-005 C | 1 X Drive Unit |  |  | Each
| RT286-11-005 D | 2 X Lip hooks |  |  | Each
| RT286-11-005 E | Measuring wire |  |  | Each
| RT286-11-005 F | Fork (Endo file holder) |  |  | Each
42152302-00000 | RT286-11-006 | Fluorescence visualization unit |  |  | Each
SPECIFICATION A flourescence visualization unit for early detecting mucosal abnormalities , including oral cancer before it can be seen under ordinary white light.
Must be non invasive.
Must be portable.
No rinses or or dyes to be used.
Screening per patient not longer than 3 min
Rechargeable for handheld use
Initial delivery must include the following:
1 X flourecence handpiece
1 X rechargeable battery
1 X charging cradle
1 X external power supply-brick
1 X pair patient safty glasses
1 X examination guide
1 X list of all replacement acessories ( to be put on RT 287)
Optional Items
| RT286-11-006 A | Caps per box to fit unit for individual patients |  |  | Each
| RT286-11-006 B | Sheets |  |  | Each
| RT286-11-006 C | Compatable Camera system for photo documentation |  |  | Each
42151902-00022 | RT286-11-007 | Piezoelectric surgery unit |  |  | Each
SPECIFICATION A piezoelectric unit for bone surgery.
Must be able to deliver about 60 watts.
Must be portable.
Must include autoclavable LED handpiece.
Must have a Large control screen
Must have a pump controlled irrigation system for constant precise flow
Initial delivery must include the following:
1 X Piezotome LED Handpiece
1 X foot control
1 X power cord
1 X Irrigation bracket
1 X Handpiece holder
SABS plug
Optional items
| RT286-11-007 A | 1 X Piezotome LED Handpiece |  |  | Each
| RT286-11-007 B | 1 X foot control |  |  | Each
| RT286-11-007 C | 1 X power cord |  |  | Each
| RT286-11-007 D | 1 X Irrigation bracket |  |  | Each
| RT286-11-007 E | 1 X Handpiece holder |  |  | Each
42151624-00013 | RT286-11-008 | Surgical Contra Angle Handpieces with 45 degree angle ( 1:2.7 with light) |  |  | Each
SPECIFICATION Surgical handpiece with light
Must 1:2.7ratio.
Must be contra angled
Must be push button type
Must be able to hold 1.6mm di FG -type burs
Maximum drive speed 40 000 rpm
Must be able to work with any surgical unit
Must have an external spray
Must be able to be fully dismantled without tools for ease of cleaning prior to sterilization
27111701-00000 | RT286-11-009 | Prosthodontic Electronic Screwdriver |  |  | Each
SPECIFICATION A cordless and electronic prostho screwdriver to work with all latch type burs and torque wrenches and drivers.
Must have 1:2.7ratio.
Torque control can be set from 8 to38 Ncm
to replace all torque wrenches and drivers
Initial delivery must include the following:
1 X Rechargeable battery
1 X Handpiece drive -contrangle
1 X Hexagon chucking system
1 X Charging station
Optional items
| RT286-11-009 A | 1 X Rechargeable battery |  |  | Each
| RT286-11-009 B | 1 X Handpiece drive -contrangle |  |  | Each
| RT286-11-009 C | 1 X Hexagon chucking system |  |  | Each
| RT286-11-009 D | 1 X Charging station |  |  | Each
42152910-00000 | RT286-11-010 | Cleaning, lubrication and purging system for dental handpieces |  |  | Each
SPECIFICATION A cleaning, lubrication and purging system to maintain dental handpieces.
Must have a motor type that turns gears on attachment for maintenance of moving parts.
Must have chuck maintenance section for bur retention.
Must have an internal oil suction that removes oil and debris to minimize cleanup.
Must have provision for an adapter e.g. multiflex for maintaining turbine handpieces.
Must be able to accept other adapters to maintain various brands an handpiece types
SABS plug
Initial delivery must include the following:
1 X Service adapter
1 X Charging cradle
1 X Air hose
1 X Window cover
1 X 500ml of oil
Optional items
| RT286-11-010 A | 1 X Service adapter |  |  | Each
| RT286-11-010 B | 1 X Charging cradle |  |  | Each
| RT286-11-010 C | 1 X Air hose |  |  | Each
| RT286-11-010 D | 1 X Window cover |  |  | Each
| RT286-11-010 E | 1 X 500ml of oil |  |  | Each
42152202-00000 | RT286-11-011 | Hydro Abrasion Cavity Preparation and Polishing System |  |  | Each
SPECIFICATION Hydro abrasion cavity preparation and polishing system for treatment without anesthetic.
Must have a motor type that turns gears on attachment for maintenance of moving parts.
Must have multiple tips for different procedures.
Should have color coded cartridge system .
Must be able to utilize 26 micron to 50 micron Aluminum Oxide Powder an and sodium bicarbonate .
Must be able to utilize single patient nozzle and tip (to be put on RT296)
Foot Control
SABS plug
Initial delivery must include the following:
1 X 29 micron Aluminum Oxide powder (Fine cutting Powder) container.
1 X 53 micron Aluminum Oxide powder(Course Cutting Powder) container.
1 X Stain Removal container (Sodium Bicarbonite)
1 X Polishing and profhylaxis Container1
50 X Disposable tips
Optional items
| RT286-11-011 A | 29 micron Aluminum Oxide powder (Fine cutting Powder) container. |  |  | Each
| RT286-11-011 B | 53 micron Aluminum Oxide powder(Course Cutting Powder) container. |  |  | Each
| RT286-11-011 C | Stain Removal container (Sodium Bicarbonite) |  |  | Each
| RT286-11-011 D | Polishing and profhylaxis Container1 |  |  | Each
| RT286-11-011 E | Disposal tips sizes 0.4/0.6/08mm |  |  | Each
42151693-00000 | RT286-11-012 | Endodontic Irrigation and Debridement Apparatus |  |  | Each
SPECIFICATION Endodontic irrigation and debridement apparatus for disruption of smear layer of biofilm
Must have a contra-angled design for access to posterior teeth.
Must have multiple tips for different procedures.
Should have at least 3 speed motor (sonic) for options ranging from 1800 up to 10200 cycles per minute
Must be cordless.
Tips must be snap-on/snap-off design.
Tip sizes must be color coded
Initial delivery must include the following:
25 X of each type of disposable tips, small, medium and large tips.
100 X protective sleeves for the unit
1 X sequence card
Batteries
Optional items
| RT286-11-012 A | Disposable tips, small, medium and large tips. |  |  | Each
| RT286-11-012 B | Protective sleeves for the unit. |  |  | Each
| RT286-11-012 C | Sequence card. |  |  | Each
42151691-00001 | RT286-11-013 | Rotating Corded Endodontic Motor, Handpiece and Apex Locator Combination |  |  | Each
SPECIFICATION Endodontic motor and handpiece combination for rotating NiTi files
Must have a contra-angled design for access to posterior teeth.
Must have auto reverse function.
Should have adjustable speed motor ranging from 250 up to 1200 rpm
Adjustable tourque ranging from 0.6 - 5.2 N.cm
Integrated apex locator
Self callibration function
Rechargeable batteries
Stand (If separate from the charging unit)
Charging Unit with SABS plug (If separate from the stand)
SABS plug
Optional items
| RT286-11-013 A | Rechargeable batteries |  |  | Each
| RT286-11-013 B | Stand (If separate from the charging unit) |  |  | Each
| RT286-11-013 C | Charging Unit with SABS plug (If separate from the stand) |  |  | Each
| RT286-11-013 D | Contra angle head 2.1 incl conection fork |  |  | Each
| RT286-11-013 E | Handpiece and lead without the head |  |  | Each
| RT286-11-013 F | 1 X mesurement cable and 2 hooks |  |  | Each
| RT286-11-013 G | Duel hook |  |  | Each
| RT286-11-013 H | Lip clips |  |  | Each
| RT286-11-013 I | Fork for apex mesuerments |  |  | Each
42151691-00002 | RT286-11-014 | Rotating Cordless Endodontic Motor and Handpiece Combination |  |  | Each
SPECIFICATION Endodontic motor and handpiece combination for rotating NiTi files- cordles
Must have a contra-angled design for access to posterior teeth.
Must have auto reverse function.
Should have adjustable speed motor ranging from 250 up to 1200 rpm
Adjustable tourque ranging from 0.6 - 5.2 N.cm
Self callibration function
Rechargeable Batteries
Stand (If separate from the charging unit)
Charging Unit with SABS plug (If separate from the stand)
Optional items
| RT286-11-014 A | Rechargeable Batteries |  |  | Each
| RT286-11-014 B | Stand |  |  | Each
| RT286-11-014 C | Charging Unit with SABS plug |  |  | Each
| RT286-11-014 D | Contra angle head |  |  | Each
| RT286-11-014 E | Handpiece without the head |  |  | Each
42151691-00003 | RT286-11-015 | Corded Rotating and reciprocating Endodontic Motor and Handpiece Combination |  |  | Each
SPECIFICATION Endodontic motor and handpiece combination for rotating and reciprocating NiTi files
Must have a contra-angled design for access to posterior teeth.
Must have auto reverse function.
Should have adjustable speed motor ranging from 250 up to 1200 rpm
Adjustable tourque ranging from 0.6 - 5.2 N.cm
Self callibration function
Rechargeable batteries
Stand (If separate from the charging unit)
Charging Unit with SABS plug (If separate from the stand)
SABS plug
Optional items
| RT286-11-015 A | Rechargeable batteries |  |  | Each
| RT286-11-015 B | Stand (If separate from the charging unit) |  |  | Each
| RT286-11-015 C | Charging Unit with SABS plug (If separate from the stand) |  |  | Each
| RT286-11-015 D | Contra angle head |  |  | Each
| RT286-11-015 E | Handpiece without the head |  |  | Each
| RT286-11-015 F | Handpiece Stand |  |  | Each
| RT286-11-015 G | Spray Nozzle (F-type) |  |  | Each
| RT286-11-015 H | WaveOe Contra angle 6:1 ratio |  |  | Each
| RT286-11-015 I | Foot Pedal |  |  | Each
42151691-00004 | RT286-11-016 | Endodontic down Pack and Backfill Root Canal Machine |  |  | Each
SPECIFICATION An Endodontic down Pack and backfill root canal machine
Must have a Down pack capability and have the following characteristic:.
Use of pluggers of various sizes
Pluggers must be able for use of down packing and searing of GP
Must be able to utilise pluggers to be click-on/click of type
Must have a thermal response tip for pulp vitality testing
Must have a Backfill capability and have the following :
Temperature and flow rate to be set individually
Precise release of GP by micro motor
Quick change of cartridges through drop-in/drop out design
Have flexible and extra heat conductive silver alloy cannulas that is suitable for curved canals
One cartridge must be able to fill up to 5 canals
Initial delivery must include the following:
1 X Plugger small ISO 40 with .03 taper
1 X pack of 10 cartridges 20 G
1 X pack of 10 cartridges 23G
SABS plug
Optional items
| RT286-11-016 A | 1 X Plugger small ISO 40 with .03 taper |  |  | Each
| RT286-11-016 B | 1 X Plugger medium ISO 50 with .05 taper |  |  | Each
| RT286-11-016 C | 1 X Plugger large ISO 60 with .06 taper |  |  | Each
| RT286-11-016 D | 1 X pack of 10 cartridges 20 G |  |  | Each
| RT286-11-016 E | 1 X pack of 10 cartridges 23G |  |  | Each
| RT286-11-016 F | Replacement cartridge nuts |  |  | Each
| RT286-11-016 G | Thermanal respose tip |  |  | Each
| RT286-11-016 H | Replacement Cartrige Nut |  |  | Each
| RT286-11-016 I | Cartridge bending tool |  |  | Each
| RT286-11-016 J | Hand Cleaning Brush |  |  | Each
| RT286-11-016 K | Heat shield |  |  | Each
| RT286-11-016 L | GP Cartridges 20G |  |  | Each
| RT286-11-016 M | GP Cartridges 20G |  |  | Each
42152008-00000 | RT286-11-017 | Portable Handheld X Ray Unit |  |  | Each
SPECIFICATION A Portable handheld X-ray for use were conventional mounted X-rays can not be used.
Must be able to work with film, digital sensors and photostimulable phosphor plated
It would be preferable if a attachable sensor and build in computer with a digital screen to see the radiographs "'lives" on the machine can be supplied
The X-ray unit must be wireless, handheld and completely mobile.
Must be for all dental use e.g. diagnostic and endodontics.
Must be able to allow at least 250 images on a single charge.
Must have a Backfill capability and have the following :
Must be supplied with a rechargeable battery.
Backup rechargeable battery
Charging unit
It would be preferred if the unit have the ability to transfer images wireless to a PC.
SABS plug
Optional items
| RT286-11-017 A | Rechargeable battery |  |  | Each
| RT286-11-017 B | Charging unit |  |  | Each
47101529-00000 | RT286-11-018 | Hypodermic Needle incinerator portable |  |  | Each
SPECIFICATION A device that incinerate needles to prevent needle injuries that is portable
A portable hypodemic needle incinirator for dental needles.
Must comply with all health and safety regulations.
Temperature range from 1500 - 2000 degrees celcius
Indicator light
Needle gauge 25mm and 38mm
Dimensions Height 95mm x Width 170mm x Length 300mm
Charging unit
Batteries
SABS plug
47101529-00001 | RT286-11-019 | Hypodermic Needle Incinerator Desktop |  |  | Each
SPECIFICATION A device that incinerate needles to prevent needle injuries that is portable
A desktopn hypodermic needle incinirator for dental needles.
Must comply with all health and safety regulations.
Temperature range from 1500 - 2000 degrees celcius
Indicator light
Needle gauge 25mm and 38mm
Dimensions Height 95mm x Width 170mm x Length 300mm
SABS plug
41103903-00000 | RT286-11-020 | Platelet Rich Plasma Centrifuge (PRP) |  |  | Each
SPECIFICATION Unstainless frame
Microprocessor control
Brushless motor driving
Monitored by conductivity motor
LCD Display
Protection of unbalanced an door cover with function of alarming
SABS Plug
Accessories
| RT286-11-020 A | PRP Tubes |  |  | Each
| RT286-11-020 B | Activator PRP tubes |  |  | Each
42152015-00000 | RT286-11-021 | Intra Oral Wireless Scanner with built in Colour Determination |  |  | Each
SPECIFICATION Intra oral wireless scanner for use in digital impressions with the following features:
Realistic colours, shade measurement, and an intraoral camera
Pen grip Wireless
Portable, lightweight pod and scanner that connects to your laptop via USB
Battery Charger with at least 2 X Batteries
Must be able to include the widest range of treatment options with verified CAD/CAM lab workflows nl.
Crowns
Bridges
Veneers, Inlays & Onlays
Removable Partial Dentures
Posts and Cores
Abutments, Implant Bridges and Bars
Envisaging
Must be able to include the at least the following add on software:
Orthodontics
Implant Studio
Design Studio
SABS Plug
RT286-01-001 | RT286-11-022 | Air purifiers |  |  | Each
SPECIFICATION
Four or five stage filtration system
HEPA filtration system, UVC , activated carbon
Multi humidity settings
With adjustable Timer
Filter replacement indicator
Reduces viruses and bacteria
Multi speed setting
Impregnable seal around filters
Energy efficient
With sleep mode
Filter lifespan more than six months
Low noise levels
Adjustable fan speed
Air quality indicators
RT286-01-002 | RT286-11-023 | Hydroclorous acid generator |  |  | Each
SPECIFICATION
Disinfecting manufacturing unit producing hypochlorous acid and sodium hypochloride.
Capacity – 5lt to 20lt
Working Time – 5 minutes to 1 Hour
PPM – 300 to 1500mg/L
Producing a Chlorine concentration of more than 1000mg/L that is effective against a broad range of microorganisms.
Effective sterilization rate of 99%.
Does not produce carcinogenic, teratogenic, mutagenic and other chlorinated compounds.
Product with shelf life reaching and exceeding 30 days.
Should include minimum of a two-year warranty.
Should include a maintenance contract of 3 to 5 years.
RT286-01-003 | RT286-11-024 | Extra oral vacuum/ extractor |  |  | Each
SPECIFICATION
To remove high volume of droplets and aerosols produced during treatment
Filter efficiency - > 99% elimination of 0.3 micron size particles
For use in closed spaces
Sound - < 100db high speed < or = 60 db low speed
Triple H14 HEPA filter system or quadruple activated carbon and H14 HEPA filter system
Contains UVC lamp working with the HEPA filter to inactivate viruses and bacteria
Adjustable power levels
With Auto HEPA filter replacement reminder
Lightweight – portable
Movable, flexible suction arm
Long working time of approximately 4 hours
RT286-01-004 | RT286-11-025 | Germicidal Ultraviolet Lamp |  |  | Each
SPECIFICATION
· Wall mounted or Ceiling mounted
· Lamp emitting UV-C light that is germicidal
· Covering surface area up to 50 sqm
· Safety Timer
· Designed for consultation rooms, passages, waiting rooms, etc
RT286-01-005 | RT286-12-026 | Hydrocyclone Amalgam IS18 SEPARATOR |  |  | Each
SPECIFICATION
PAL 14, 28, 38, 50
IS 18
VE with Electronic Amalgam Level Control
Components
i) Centrifuge
ii) Hydrocyclone
iii) Pressure relief valve
iv) Control panel
v) Output Power 0.25KW - 0.28 A
vi) Tested by TUV, according to ISO 11143 Standard
vii) W=470mm, D=400mm, H=590mm
RT286-01-006 | RT286-12-027 | Hydrocyclone ISO 6 Amalgam separator |  |  | Each
SPECIFICATION
I) Fixing plates
ii) Re-circulation hoses
iii) TUV Tested according to ISO11143 standard
RT286-01-007 | RT286-12-028 | Hydrocyclone Amalgam separator |  |  | Each
SPECIFICATION
Turbo Smart 2V Before 2017
Components
i)Hydrocyclone ISO 18
ii) Fixing plates
iii) Re-circulation hoses
iv)Tested by TUV according to ISO 11143 Stardard
RT286-01-008 | RT286-12-029 | Hydrocyclone ISO 4 Amalgam separator |  |  | Each
SPECIFICATION
Version for Turbo-Jet 1 Modular without a box
Hydrocyclone
Outlet connector with pressure relief valve
Tested by TUV according to ISO 1143 Standard
RT286-01-009 | RT286-12-030 | Hydrocyclone Amalgam separator |  |  | Each
SPECIFICATION
Multi-surgery version to be installed inside dental unit
Components
i) Mini-separator
ii) Mignon Electro-pneumatic valve
iii)Draining unit
iv) AC 45 PB
v)Ø25 Straight tube holder
vi)Ø30 Straight tube holder
vi)Ø25 bent tube holder
vii)Ø30 bent tube holder
viii) Canister W=195mm, D=135mm, H=340mm
ix) Electrovalve W=120mm, D=105, H=160mm
RT286-01-010 | RT286-12-031 | Hydrocyclone Amalgam separator |  |  | Each
SPECIFICATION
Single-surgery version to be installed inside dental unit
Components
i) Mini-separator
ii) Drainig unit
iii)AC 45 PB
v)Ø25 Straight tube holder
vi)Ø30 Straight tube holder
vi)Ø25 bent tube holder
vii)Ø30 bent tube holder
viii) Canister W=195mm, D=135mm, H=340mm
Note on C2062: Balekile Ngalo: Military clinics in RSA is replaced with at different facilities.

## Sheet2
ITEM NUMBER | GROUP SERIES | ITEM SHORT DESCRIPTION (Please refer to Annexure A for Full spec) | UOM | OFFER # | BIDDER NAME | EQUIPMENT/ ACCESSORY PRICE OFFERIN ZAR INCL. OF VAT | MAMAINTENANCE COST (POST 24 MONTHS WARANTEE PERIOD) |  |  |  |  | MODEL / BRAND NAME | MODEL NUMBER | YEAR WHICH THE MODEL RANGE WAS RELEASED | EXPECTED LIFE SPAN OF THE EQUIPMENT | COUNTRY OF ORIGIN | NAME OF MANUFACTURER | LOCAL CONTENT % | IMPORT CONTENT % | FOREIN CURRENCY NAME
[7 empty] | YEAR 3 | YEAR 4 | YEAR 5 | YEAR 6 | YEAR 7
RT286-01-001 | Group 1 - Equipment | Plaster Dispenser - Electric | Each | Offer 1
RT286-01-002 | Group 2 - Equipment | Model Trimmer - Single Disc / Blade | Each | Offer 1
RT286-01-003 | Group 3 - Equipment | Model Trimmer - Dual Disc / Blade | Each | Offer 1
RT286-01-003 A | Group 3 - Accessory | Replacement disc Coarse for trimming diameter 12" | Each | Offer 1
RT286-01-003 B | Group 3 - Accessory | Replacement disc Fine for Finishing, diameter 12" | Each | Offer 1
RT286-01-003 C | Group 3 - Accessory | Replacement Vertisquare | Each | Offer 1
RT286-01-004 | Group 4 - Equipment | Model Trimmer - Dry System with Partially Coated Disc / Blade | Each | Offer 1
RT286-01-004 A | Group 4 - Accessory | Replacement partially diamond coated disc | Each | Offer 1
RT286-01-005 | Group 5 - Equipment | Pin Hole Drilling Unit | Each | Offer 1
RT286-01-005 A | Group 5 - Accessory | Replacement pin drill | Each | Offer 1
RT286-01-005 B | Group 5 - Accessory | Replacement LED Light | Each | Offer 1
RT286-01-006 | Group 6 - Equipment | Surveyor for Models - Single Arm | Each | Offer 1
RT286-01-006 A | Group 6 - Accessory | Lead marker holder | Each | Offer 1
RT286-01-006 B | Group 6 - Accessory | Lead | Each | Offer 1
RT286-01-006 C | Group 6 - Accessory | Wax knife / scraper | Each | Offer 1
RT286-01-006 D | Group 6 - Accessory | Bracing holder | Each | Offer 1
RT286-01-006 E | Group 6 - Accessory | Analysing pin | Each | Offer 1
RT286-01-006 F | Group 6 - Accessory | Undercut measuring pin: 0,25mm | Each | Offer 1
RT286-01-006 G | Group 6 - Accessory | Undercut measuring pin: 0,50mm | Each | Offer 1
RT286-01-006 H | Group 6 - Accessory | Undercut measuring pin: 0,75mm | Each | Offer 1
RT286-01-007 | Group 7 - Equipment | Surveyor for Models - Double Arm | Each | Offer 1
RT286-01-007 A | Group 7 - Accessory | Lead marker holder | Each | Offer 1
RT286-01-007 B | Group 7 - Accessory | Lead | Each | Offer 1
RT286-01-007 C | Group 7 - Accessory | Wax knife / scraper | Each | Offer 1
RT286-01-007 D | Group 7 - Accessory | Bracing holder | Each | Offer 1
RT286-01-007 E | Group 7 - Accessory | Analysing pin | Each | Offer 1
RT286-01-007 F | Group 7 - Accessory | Undercut measuring pin: 0,25mm | Each | Offer 1
RT286-01-007 G | Group 7 - Accessory | Undercut measuring pin: 0,50mm | Each | Offer 1
RT286-01-007 H | Group 7 - Accessory | Undercut measuring pin: 0,75mm | Each | Offer 1
RT286-01-008 | Group 8 - Equipment | Boiling-Out Unit 6 or 12 (individual pricing) flask capacity | Each | Offer 1
RT286-01-008 A | Group 8 - Accessory | Descaling liquid | Each | Offer 1
RT286-01-008 B | Group 8 - Accessory | Hand shower spray head | Each | Offer 1
RT286-01-008 C | Group 8 - Accessory | Replacement flask baskets | Each | Offer 1
RT286-01-009 | Group 9 - Equipment | Blow-Off Nozzle - Rubber | Each | Offer 1
RT286-01-009 A | Group 9 - Accessory | Replacement ring clamp | Each | Offer 1
RT286-02-001 | Group 10 - Equipment | Duplicating Unit - Reversible Duplicating Material | Each | Offer 1
RT286-02-002 | Group 11 - Equipment | Duplicating Unit - Addition Cured Two-Component Silicone | Each | Offer 1
RT286-02-002 A | Group 11 - Accessory | Pack of mixing nozzles x 100 | Each | Offer 1
RT286-02-003 | Group 12 - Equipment | Model Drying Cabinet | Each | Offer 1
RT286-02-003 A | Group 12 - Accessory | Removable shelf | Each | Offer 1
RT286-02-003 B | Group 12 - Accessory | Replacement globe | Each | Offer 1
RT286-02-003 C | Group 12 - Accessory | Replacement door seal | Each | Offer 1
RT286-02-004 | Group 13 - Equipment | Flask Press - Manual: 1 - 3 Flasks | Each | Offer 1
RT286-02-005 | Group 14 - Equipment | Flask Press - Hydraulic: 1 - 3 Flasks | Each | Offer 1
RT286-02-006 | Group 15 - Equipment | Thermal Moulding Device: Positive Pressure Technique | Each | Offer 1
RT286-02-006 A | Group 15 - Accessory | Replacement set of O-rings | Each | Offer 1
RT286-02-006 B | Group 15 - Accessory | Replacement locking plate for material | Each | Offer 1
RT286-02-006 C | Group 15 - Accessory | Replacement pellets for embedding models in moulding device | Each | Offer 1
RT286-03-001 | Group 16 - Equipment | Electronic Wax Modelling Instrument | Each | Offer 1
RT286-03-001 A | Group 16 - Accessory | Replacement modelling tips | Each | Offer 1
RT286-03-002 B | Group 16 - Accessory | Replacement Heating handle with cable | Each | Offer 1
RT286-03-002 | Group 17 - Equipment | Electronic Wax Heater | Each | Offer 1
RT286-03-003 | Group 18 - Equipment | Bunsen Burner Lab Dual Pipe – Propane Gas | Each | Offer 1
RT286-03-003 A | Group 18 - Accessory | Replacement ring clamp | Each | Offer 1
RT286-03-004 | Group 19 - Equipment | Bunsen Burner Lab Dual Pipe – Natural Gas | Each | Offer 1
RT286-03-004 A | Group 19 - Accessory | Replacement pressure hose for natural gas | Each | Offer 1
RT286-03-004 B | Group 19 - Accessory | Replacement Ring Clamps | Each | Offer 1
RT286-03-005 | Group 20 - Equipment | Burner Outfit, Gas: Fully Automatic Ignition & Cut-Off | Each | Offer 1
RT286-03-005 A | Group 20 - Accessory | Replacement pressure hose for gas | Each | Offer 1
RT286-03-005 B | Group 20 - Accessory | Replacement Ring Clamps | Each | Offer 1
RT286-03-006 | Group 21 - Equipment | Micro Torch – Butane Gas | Each | Offer 1
RT286-03-006 A | Group 21 - Accessory | Bottle Butane gas refill | Each | Offer 1
RT286-03-007 | Group 22 - Equipment | Induction Heater Portable for Hand Instruments | Each | Offer 1
RT286-03-007 A | Group 22 - Accessory | Replacement silicone cup to protect the chamber | Each | Offer 1
RT286-03-008 | Group 23 - Equipment | Soldering and Melting Torch: Microflame – Propane Gas | Each | Offer 1
RT286-03-008 A | Group 23 - Accessory | Replacement needle jets diameter: 0.9, 1.2, 1.6 mm | Each | Offer 1
RT286-03-008 B | Group 23 - Accessory | Replacement pressure hose for oxygen | Each | Offer 1
RT286-03-008 C | Group 23 - Accessory | Replacement pressure hose for propane gas | Each | Offer 1
RT286-03-008 D | Group 23 - Accessory | Replacement Ring Clamps | Each | Offer 1
RT286-03-009 | Group 24 - Equipment | Soldering and Melting Torch: Microflame – Natural Gas | Each | Offer 1
RT286-03-009 A | Group 24 - Accessory | Replacement needle jets diameter: 0.9, 1.2, 1.6 mm | Each | Offer 1
RT286-03-009 B | Group 24 - Accessory | Replacement pressure hose for oxygen | Each | Offer 1
RT286-03-009 C | Group 24 - Accessory | Replacement pressure hose for propane gas | Each | Offer 1
RT286-03-009 D | Group 24 - Accessory | Replacement Ring Clamps | Each | Offer 1
RT286-04-001 | Group 25 - Equipment | Vacuum Mixer: Plaster & Investment Material | Each | Offer 1
RT286-04-001 A | Group 25 - Accessory | Mixing bowls approximately 200ml - 1200ml | Each | Offer 1
RT286-04-001 B | Group 25 - Accessory | Replacement mixing bowl O-Ring | Each | Offer 1
RT286-04-001 C | Group 25 - Accessory | Replacement O-Ring for suction/vacuum nozzle | Each | Offer 1
RT286-04-002 | Group 26 - Equipment | Vacuum Mixer: Plaster, Investment & Silicone Duplicating Material | Each | Offer 1
RT286-04-002 A | Group 26 - Accessory | Mixing bowls approximately 200ml - 1200ml | Each | Offer 1
RT286-04-003 B | Group 26 - Accessory | Replacement mixing bowl O-Ring | Each | Offer 1
RT286-04-004 C | Group 26 - Accessory | Replacement O-Ring for suction/vacuum nozzle | Each | Offer 1
RT286-04-003 | Group 27 - Equipment | Laboratory Vibrator | Each | Offer 1
RT286-04-003 A | Group 27 - Accessory | Replacement Rubber cover | Each | Offer 1
RT286-04-004 | Group 28 - Equipment | Dispenser, Dental Acrylic Powder and Liquid: Orthodontic | Each | Offer 1
RT286-04-004 A | Group 28 - Accessory | Replacement metal nozzles | Each | Offer 1
RT286-04-004 B | Group 28 - Accessory | Replacement flexible pipes | Each | Offer 1
RT286-04-004 C | Group 28 - Accessory | Replacement carbon filter | Each | Offer 1
RT286-04-004 D | Group 28 - Accessory | Replacement pedal switches | Each | Offer 1
RT286-04-005 | Group 29 - Equipment | Polymerisation Bath – Electric | Each | Offer 1
RT286-04-006 | Group 30 - Equipment | Polymerisation Vessel – Pneumatic | Each | Offer 1
RT286-04-006 A | Group 30 - Accessory | Replacement lid seal | Each | Offer 1
RT286-04-007 | Group 31 - Equipment | Polymerisation Vessel - Pneumatic & Heat | Each | Offer 1
RT286-04-007 A | Group 31 - Accessory | Replacement lid seal/O-ring | Each | Offer 1
RT286-04-007 B | Group 31 - Accessory | Replacement ring clamp | Each | Offer 1
RT286-05-001 | Group 32 - Equipment | Pre-Heating & Burnout Furnace – Small | Each | Offer 1
RT286-05-001 A | Group32 - Accessory | Replacable heating elements | Each | Offer 1
RT286-05-001 B | Group32 - Accessory | Replacement thermocouple | Each | Offer 1
RT286-05-001 C | Group32 - Accessory | Replacement steel and ceramic collecting trays | Each | Offer 1
RT286-05-002 | Group 33 - Equipment | Pre-Heating & Burnout Furnace – Medium | Each | Offer 1
RT286-05-002 A | Group33 - Accessory | Replacable heating elements | Each | Offer 1
RT286-05-002 B | Group33 - Accessory | Replacement thermocouple | Each | Offer 1
RT286-05-002 C | Group33 - Accessory | Replacement steel and ceramic collecting trays | Each | Offer 1
RT286-05-003 | Group 34 - Equipment | Pre-Heating & Burnout Furnace – Large | Each | Offer 1
RT286-05-003 A | Group34 - Accessory | Replacable heating elements | Each | Offer 1
RT286-05-003 B | Group34 - Accessory | Replacement thermocouple | Each | Offer 1
RT286-05-003 C | Group34 - Accessory | Replacement steel and ceramic collecting trays | Each | Offer 1
RT286-05-004 | Group 35 - Equipment | Casting Machine Centrifugal – Mechanical: Bench-Top | Each | Offer 1
RT286-05-004 A | Group35 - Accessory | Replacement spiral spring | Each | Offer 1
RT286-05-004 B | Group35 - Accessory | Replacement crucible | Each | Offer 1
RT286-05-004 C | Group35 - Accessory | Replacement crucible insert | Each | Offer 1
RT286-05-005 | Group 36 - Equipment | Casting Machine Centrifugal – Motorised | Each | Offer 1
RT286-05-005 A | Group36 - Accessory | Replacement crucible | Each | Offer 1
RT286-05-005 B | Group36 - Accessory | Replacement crucible insert | Each | Offer 1
RT286-05-005 C | Group36 - Accessory | Replacement ring/muffle cradle | Each | Offer 1
RT286-05-006 | Group37 - Equipment | Casting Machine Centrifugal – Induction (Bench-Top Model) | Each | Offer 1
RT286-05-006 A | Group37 - Accessory | Replacement spiral heating coil/element | Each | Offer 1
RT286-05-006 B | Group37 - Accessory | Replacement crucible | Each | Offer 1
RT286-05-006 C | Group37 - Accessory | Replacement crucible insert | Each | Offer 1
RT286-05-006 D | Group37 - Accessory | Replacement muffle/ring cradle | Each | Offer 1
RT286-05-006 E | Group37 - Accessory | Replacement coloured observation port glass | Each | Offer 1
RT286-05-007 | Group38 - Equipment | Induction Casting Machine - Vacuum Pressure and High Frequency Melting | Each | Offer 1
RT286-05-007 A | Group38 - Accessory | Replacement spiral heating coil/element | Each | Offer 1
RT286-05-007 B | Group38 - Accessory | Replacement crucible | Each | Offer 1
RT286-05-007 C | Group38 - Accessory | Replacement crucible insert | Each | Offer 1
RT286-05-007 D | Group38 - Accessory | Replacement muffle/ring cradle | Each | Offer 1
RT286-05-007 E | Group38 - Accessory | Replacement coloured observation port glass | Each | Offer 1
RT286-06-001 | Group 39 - Equipment | Blast Cleaning Machine Cobalt–Chrome (Automatic) | Each | Offer 1
RT286-06-001 A | Group39 - Accessory | Replacement rotating drum | Each | Offer 1
RT286-06-001 B | Group39 - Accessory | Replacement Filter bags | Each | Offer 1
RT286-06-001 C | Group39 - Accessory | Replacement sleeves | Each | Offer 1
RT286-06-001 D | Group39 - Accessory | Replacement gloves | Each | Offer 1
RT286-06-001 E | Group39 - Accessory | Replacement Protective glass | Each | Offer 1
RT286-06-001 F | Group39 - Accessory | Replacement light bulbs | Each | Offer 1
RT286-06-002 | Group 40 - Equipment | Blast Cleaning Machine Crown, Bridge & Cobalt - Chrome (Manual) | Each | Offer 1
RT286-06-002 A | Group40 - Accessory | Replacement fixed nozzles | Each | Offer 1
RT286-06-002 B | Group40 - Accessory | Replacement sleeves | Each | Offer 1
RT286-06-002 C | Group40 - Accessory | Replacement gloves | Each | Offer 1
RT286-06-002 D | Group40 - Accessory | Filter bags for extraction unit | Each | Offer 1
RT286-06-003 | Group41 - Equipment | Blast Cleaning Machine Crown & Bridge – Micro Sandblaster (Manual) | Each | Offer 1
RT286-06-003 A | Group41 - Accessory | Replacement light bulbs | Each | Offer 1
RT286-06-003 B | Group41 - Accessory | Protective gloves | Each | Offer 1
RT286-06-003 C | Group41 - Accessory | Replacement fixed nozzles | Each | Offer 1
RT286-06-003 D | Group41 - Accessory | Replacement protective glass | Each | Offer 1
RT286-06-003 E | Group41 - Accessory | Replacement foot switch for pneumatic valve | Each | Offer 1
RT286-06-004 | Group42 - Equipment | Dust Extraction System (Dry) – Portable | Each | Offer 1
RT286-06-004 A | Group42 - Accessory | Replacement filter bag | Each | Offer 1
RT286-06-004 B | Group42 - Accessory | Replacement ring clasp | Each | Offer 1
RT286-06-005 | Group43 - Equipment | Dust Extraction System (Dry) – Mobile | Each | Offer 1
RT286-06-005 A | Group43 - Accessory | Replacement filter bag | Each | Offer 1
RT286-06-005 B | Group43 - Accessory | Replacement cloth filter | Each | Offer 1
RT286-06-006 | Group44 - Equipment | Dust Extraction System (Wet) – Fixed | Each | Offer 1
RT286-06-006 A | Group44 - Accessory | Replacement water filter bag | Each | Offer 1
RT286-06-006 B | Group44 - Accessory | Replacement ring clasp | Each | Offer 1
RT286-06-007 | Group45 - Equipment | Dust Extraction Drawer System (Dry) – Fixed | Each | Offer 1
RT286-06-007 A | Group45 - Accessory | Replacement filter bag | Each | Offer 1
RT286-06-008 | Group46 - Equipment | Suction Hood - Electrical | Each | Offer 1
RT286-06-008 A | Group46 - Accessory | Replacement carbon filter | Each | Offer 1
RT286-06-008 B | Group46 - Accessory | Replacement light bulb | Each | Offer 1
RT286-07-001 | Group47 - Equipment | Pneumatic Chisel / Handpiece for Trimming of Plaster & Investment | Each | Offer 1
RT286-07-001 A | Group47 - Accessory | Replacement chisel bits | Each | Offer 1
RT286-07-001 B | Group47 - Accessory | Replacement ring clasp | Each | Offer 1
RT286-07-002 | Group48 - Equipment | Steam Cleaner-Portable | Each | Offer 1
RT286-07-002 A | Group48 - Accessory | One spare pressure cap supplied per unit | Each | Offer 1
RT286-07-002 B | Group48 - Accessory | Replacement light bulb | Each | Offer 1
RT286-07-003 | Group 49 - Equipment | Steam Cleaner-Fixed: Wet & Dry (High Performance) | Each | Offer 1
RT286-07-003 A | Group49 - Accessory | Ring spanner | Each | Offer 1
RT286-07-003 B | Group49 - Accessory | Replacement full demineralisation cartridge | Each | Offer 1
RT286-07-003 C | Group49 - Accessory | 2 Inserts | Each | Offer 1
RT286-07-003 D | Group49 - Accessory | Replacement light bulb | Each | Offer 1
RT286-07-003 E | Group49 - Accessory | Replacement flexible insulated hose from gun to unit | Each | Offer 1
RT286-07-003 F | Group49 - Accessory | Replacement water supply hose | Each | Offer 1
RT286-07-003 G | Group49 - Accessory | Replacement ring clasps | Each | Offer 1
RT286-07-004 | Group 50 - Equipment | Laboratory Motor Electric (Hanging) | Each | Offer 1
RT286-07-004 A | Group50 - Accessory | Replacement cable | Each | Offer 1
RT286-07-004 B | Group50 - Accessory | Replacement hanging motor handpiece catch (end of cable) | Each | Offer 1
RT286-07-005 | Group51 - Equipment | Handpiece for Laboratory Motor Electric (Hanging) | Each | Offer 1
RT286-07-005 A | Group51 - Accessory | Replacement Chuck | Each | Offer 1
RT286-07-006 | Group52 - Equipment | Dental Laboratory Micro Motor (Multi-Purpose) | Each | Offer 1
RT286-07-006 A | Group52 - Accessory | Replacement Chuck | Each | Offer 1
RT286-07-006 B | Group52 - Accessory | Replacement fuse | Each | Offer 1
RT286-07-006 C | Group52 - Accessory | Replacement foot pedal | Each | Offer 1
RT286-07-007 | Group53 - Equipment | Dental Laboratory Micro Motor (Heavy Duty) | Each | Offer 1
RT286-07-007 A | Group53 - Accessory | Replacement Chuck | Each | Offer 1
RT286-07-007 B | Group53 - Accessory | Replacement foot pedal | Each | Offer 1
RT286-07-008 | Group54 - Equipment | Air Turbine Non - LED | Each | Offer 1
RT286-07-008 A | Group54 - Accessory | Replacement rotor cartridge | Each | Offer 1
RT286-07-008 B | Group54 - Accessory | Replacement foot pedal | Each | Offer 1
RT286-07-008 C | Group54 - Accessory | Replacement ring clasp | Each | Offer 1
RT286-07-009 | Group55 - Equipment | Air Turbine LED | Each | Offer 1
RT286-07-009 A | Group55 - Accessory | Replacement rotor cartridge | Each | Offer 1
RT286-07-009 B | Group55 - Accessory | Replacement LED light | Each | Offer 1
RT286-07-009 C | Group55 - Accessory | Replacement foot pedal | Each | Offer 1
RT286-07-009 D | Group55 - Accessory | Replacement ring clasp | Each | Offer 1
RT286-07-010 | Group56 - Equipment | Polishing & Grinding Machine with built-in Suction, Lighting & Eye Protection | Each | Offer 1
RT286-07-010 A | Group56 - Accessory | Replacement filter bag | Each | Offer 1
RT286-07-010 B | Group56 - Accessory | Replacement LED light | Each | Offer 1
RT286-07-010 C | Group56 - Accessory | Replacement shield for safety glass | Each | Offer 1
RT286-07-011 | Group57 - Equipment | High Speed Grinder with Internal Suction, Lighting & Eye Protection | Each | Offer 1
RT286-07-011 A | Group57 - Accessory | Replacement chuck | Each | Offer 1
RT286-07-011 B | Group57 - Accessory | Replacement light bulb | Each | Offer 1
RT286-07-011 C | Group57 - Accessory | Replacement shield for safety glass | Each | Offer 1
RT286-07-011 D | Group57 - Accessory | Replacement drive belt | Each | Offer 1
RT286-07-012 | Group58 - Equipment | Electrolytic polishing Unit – One Casting | Each | Offer 1
RT286-07-012 A | Group58 - Accessory | Replacement clamps and hook | Each | Offer 1
RT286-07-013 | Group59 - Equipment | Electrolytic Polishing Unit - Two or More Castings | Each | Offer 1
RT286-07-013 A | Group59 - Accessory | Replacement clamps and hook | Each | Offer 1
RT286-07-014 | Group60 - Equipment | Ultrasonic Cleaning Unit- Laboratory: Large Capacity | Each | Offer 1
RT286-07-014 A | Group60 - Accessory | Replacement insert basket | Each | Offer 1
RT286-07-014 B | Group60 - Accessory | Replacement ultrasonic solution | Each | Offer 1
RT286-07-015 | Group61 - Equipment | Ultrasonic Cleaning Unit- Laboratory: Small Capacity | Each | Offer 1
RT286-07-015 A | Group61 - Accessory | Replacement ultrasonic solution | Each | Offer 1
RT286-07-016 | Group62 - Equipment | Porcelain Firing Furnace - Press/Ceramics | Each | Offer 1
RT286-07-016 A | Group62 - Accessory | Replacement heating element | Each | Offer 1
RT286-07-016 B | Group62 - Accessory | Replacement thermocouple | Each | Offer 1
RT286-07-016 C | Group62 - Accessory | Firing tray | Each | Offer 1
RT286-07-017 | Group63 - Equipment | Porcelain Firing Furnace-VITA | Each | Offer 1
RT286-07-017 A | Group63 - Accessory | Replacement heating element | Each | Offer 1
RT286-07-017 B | Group63 - Accessory | Replacement thermocouple | Each | Offer 1
RT286-07-017 C | Group63 - Accessory | Replacement USB cable | Each | Offer 1
RT286-07-017 D | Group63 - Accessory | Firing tray | Each | Offer 1
RT286-07-018 | Group64 - Equipment | InLab-InEOS CAD-CAM | Each | Offer 1
RT286-07-018 A | Group64 - Accessory | Replacement bur chuck | Each | Offer 1
RT286-07-018 B | Group64 - Accessory | Replacement burs | Each | Offer 1
RT286-07-018 C | Group64 - Accessory | Replacement water filter | Each | Offer 1
RT286-07-018 D | Group64 - Accessory | Replacement air filter | Each | Offer 1
RT286-07-018 E | Group64 - Accessory | Milling fluid | Each | Offer 1
RT286-08-001 | Group65 - Equipment | Heavy Duty Dental Spot Welder | Each | Offer 1
RT286-08-001 A | Group65 - Accessory | Replacement electrodes | Each | Offer 1
RT286-08-002 | Group66 - Equipment | Inert Gas Welder | Each | Offer 1
RT286-08-002 A | Group66 - Accessory | Replacement electrodes | Each | Offer 1
RT286-08-002 B | Group66 - Accessory | Replacement oxidation free metal | Each | Offer 1
RT286-08-002 C | Group66 - Accessory | Replacement Argon gas regulator | Each | Offer 1
RT286-08-002 D | Group66 - Accessory | Replacement Argon gas flexible pipe | Each | Offer 1
RT286-08-002 E | Group66 - Accessory | Replacement ring clasps | Each | Offer 1
RT286-08-003 | Group67 - Equipment | Laser Welder | Each | Offer 1
RT286-08-003 A | Group67 - Accessory | Replacement electrodes | Each | Offer 1
RT286-08-003 B | Group67 - Accessory | Replacement Argon gas regulator | Each | Offer 1
RT286-08-003 C | Group67 - Accessory | Replacement Argon gas flexible pipe | Each | Offer 1
RT286-08-003 D | Group67 - Accessory | Replacement ring clasps | Each | Offer 1
RT286-09-001 | Group68 - Equipment | Chair Dental Laboratory upholstered | Each | Offer 1
RT286-09-002 | Group69 - Equipment | Chair Dental Laboratory with silicone | Each | Offer 1
RT286-09-003 | Group70 - Equipment | Magnifying Light – Circular | Each | Offer 1
RT286-09-003 A | Group70 - Accessory | Replacement circular LED light | Each | Offer 1
RT286-09-003 B | Group70 - Accessory | Replacement magnifying lens | Each | Offer 1
RT286-09-004 | Group71 - Equipment | Plaster Trap - Under Basin | Each | Offer 1
RT286-10-001 | Group72 - Equipment | Intra-Oral X-Ray Processor-Automatic | Each | Offer 1
RT286-10-002 | Group73 - Equipment | Intra-Oral X-Ray Processor-Manual | Each | Offer 1
RT286-10-002 A | Group73 - Accessory | Replacement sleeves | Each | Offer 1
RT286-10-003 | Group74 - Equipment | Heat Sealer | Each | Offer 1
RT286-10-003 A | Group74 - Accessory | Heating element | Each | Offer 1
RT286-10-003 B | Group74 - Accessory | Cover | Each | Offer 1
RT286-10-003 C | Group74 - Accessory | replacement element | Each | Offer 1
RT286-10-003 D | Group74 - Accessory | Insulator | Each | Offer 1
RT286-10-003 E | Group74 - Accessory | Cutter Balade | Each | Offer 1
RT286-10-004 | Group75 - Equipment | Diode Laser Dental - Portable | Each | Offer 1
RT286-10-005 | Group76 - Equipment | Endodontic Obturator Oven | Each | Offer 1
RT286-10-006 | Group77 - Equipment | Water Demineraliser Unit | Each | Offer 1
RT286-10-006 A | Group77 - Accessory | Cartriges | Each | Offer 1
RT286-10-006 B | Group77 - Accessory | SABS Plug | Each | Offer 1
RT286-10-007 | Group78 - Equipment | Water Distillation Unit | Each | Offer 1
RT286-10-008 | Group79 - Equipment | Ultrasonic Cleaning Unit | Each | Offer 1
RT286-10-008 A | Group79 - Accessory | Basket | Each | Offer 1
RT286-10-008 B | Group79 - Accessory | Beaker | Each | Offer 1
RT286-10-008 C | Group79 - Accessory | Bur tray | Each | Offer 1
RT286-10-009 | Group80 - Equipment | Amalgam Separator Multiple Surgeries | Each | Offer 1
RT286-10-009 A | Group80 - Accessory | Collection Container with recycle kit | Each | Offer 1
RT286-10-010 | Group81 - Equipment | Amalgam Separator 2 Surgeries | Each | Offer 1
RT286-10-010 A | Group81- Accessory | Collection Container with recycle kit | Each | Offer 1
RT286-10-011 | Group82 - Equipment | Suction Wet 2-4 Surgeries | Each | Offer 1
RT286-10-012 | Group83 - Equipment | Suction Dry 4-6 Surgeries | Each | Offer 1
RT286-10-013 | Group84 - Equipment | Suction Dry 2-3 Surgeries | Each | Offer 1
RT286-10-014 | Group85 - Equipment | Suction Wet 1-2 Surgeries | Each | Offer 1
RT286-10-015 | Group86 - Equipment | Suction Dry 1 Surgery | Each | Offer 1
RT286-10-016 | Group87 - Equipment | Suction Dry Mobile 1 Surgery | Each | Offer 1
RT286-10-017 | Group89 - Equipment | Pan X-Ray Viewer | Each | Offer 1
RT286-10-018 | Group90 - Equipment | Compressor 1 Surgery - Oil free | Each | Offer 1
RT286-10-019 | Group91 - Equipment | Oil Free air compressor for 4-6 surgery (60DB) | Each | Offer 1
RT286-10-020 | Group92 - Equipment | Oil Free air compressor for 2-4 surgery (60DB) | Each | Offer 1
RT286-10-021 | Group93 - Equipment | Ultrasound Scaler Multi-Function | Each | Offer 1
RT286-10-022 | Group94 - Equipment | Ultrasound Piezo Scaler | Each | Offer 1
RT286-10-023 | Group95 - Equipment | Ultrasound Scaler Multi-Function | Each | Offer 1
RT286-10-024 | Group96 - Equipment | Dental Ultrasonic Scaler with a polishing function (Insert 30K) | Each | Offer 1
RT286-10-025 | Group97 - Equipment | Dental Ultrasonic Scaler (Inserts 25K) | Each | Offer 1
RT286-10-026 | Group98 - Equipment | Dental Air Scaler | Each | Offer 1
RT286-10-027 | Group99 - Equipment | Desktop Autoclave S-Type for Dental Instruments | Each | Offer 1
RT286-10-028 | Group100 - Equipment | Mobile Dental Unit non fibre optic without hand pieces | Each | Offer 1
RT286-10-029 | Group101 - Equipment | Mobile Dental Unit non fibre optic with hand pieces | Each | Offer 1
RT286-10-030 | Group102 - Equipment | Mobile Dental Unit fibre optic with hand pieces | Each | Offer 1
RT286-10-031 | Group103 - Equipment | Dental Autoclave class full range | Each | Offer 1
RT286-10-032 | Group104 - Equipment | Dental Autoclave B-Class | Each | Offer 1
RT286-10-033 | Group105 - Equipment | Dental Autoclave B-Class with prion cycle | Each | Offer 1
RT286-10-034 | Group106 - Equipment | Portable dental unit | Each | Offer 1
RT286-10-035 | Group107 - Equipment | Autoclave for handpieces | Each | Offer 1
RT286-10-036 | Group108 - Equipment | Automatic Intra Oral X-Ray developer | Each | Offer 1
RT286-10-037 | Group109 - Equipment | Digital Intraoral video camera system | Each | Offer 1
RT286-10-038 | Group110 - Equipment | Digital Intraoral imaging plate system with computer | Each | Offer 1
RT286-10-039 | Group111 - Equipment | Digital Intraoral imaging plate system without computer | Each | Offer 1
RT286-10-040 | Group112 - Equipment | Cone Beam CT 3D Dental imaging system (small field of view- SFOV) | Each | Offer 1
RT286-10-041 | Group113 - Equipment | Cone Beam CT 3D Dental imaging system(big field of view-FOV) | Each | Offer 1
RT286-10-042 | Group114 - Equipment | Dental digital panoramic X-Ray | Each | Offer 1
RT286-10-043 | Group115 - Equipment | Digital Radiography system-sensor without a computer | Each | Offer 1
RT286-10-044 | Group116 - Equipment | Digital Radiography system-sensor with a computer | Each | Offer 1
RT286-10-045 | Group117 - Equipment | Intra Oral Dental X-ray (Analogue & Digital) - Mobile | Each | Offer 1
RT286-10-046 | Group118 - Equipment | Intra Oral Dental X-Ray (Analogue & Digital) Floor Mounted | Each | Offer 1
RT286-10-047 | Group119 - Equipment | Intra Oral Dental X-Ray Duarl Function (Analogue & Digital) wall Mounted | Each | Offer 1
RT286-10-048 | Group120 - Equipment | Dental Digital panoramic X-Ray with cephalometric without a printer | Each | Offer 1
RT286-10-049 | Group121 - Equipment | Dental Digital panoramic X-Ray with cephalometric with a printer | Each | Offer 1
RT286-10-050 | Group122 - Equipment | Dental Handpieces | Each | Offer 1
RT286-10-051 | Group123 - Equipment | Turbine full range | Each | Offer 1
RT286-10-052 | Group124 - Equipment | Turbine full range | Each | Offer 1
RT286-10-053 | Group125 - Equipment | Turbine full range | Each | Offer 1
RT286-10-054 | Group126 - Equipment | Led straight handpiece for micromotor - Air full range | Each | Offer 1
RT286-10-055 | Group127 - Equipment | Led straight handpiece for micromotor - Air full range | Each | Offer 1
RT286-10-056 | Group128 - Equipment | LED contra-angle handpiece for micromotor - Air full range | Each | Offer 1
RT286-10-057 | Group129 - Equipment | LED contra-angle handpiece for micromotor - Air full range | Each | Offer 1
RT286-10-058 | Group130 - Equipment | LED contra-angle handpiece for micromotor - Air full range | Each | Offer 1
RT286-10-059 | Group131 - Equipment | LED contra-angle handpiece for micromotor - Air full range | Each | Offer 1
RT286-10-060 | Group132 - Equipment | LED contra-angle handpiece for micromotor - Air full range | Each | Offer 1
RT286-10-061 | Group133 - Equipment | LED contra-angle handpiece for micromotor - Air full range | Each | Offer 1
RT286-10-062 | Group134 - Equipment | LED prophylaxis handpiece full range | Each | Offer 1
RT286-10-063 | Group135 - Equipment | LED prophylaxis handpiece full range | Each | Offer 1
RT286-10-064 | Group136 - Equipment | Micromotor - Air full range | Each | Offer 1
RT286-10-065 | Group137 - Equipment | Electric Micro Motors ( stand alone) | Each | Offer 1
RT286-10-066 | Group138 - Equipment | Dental operator stool | Each | Offer 1
RT286-10-067 | Group139 - Equipment | Dental delivery system | Each | Offer 1
RT286-10-068 | Group140 - Equipment | Dental LED Light | Each | Offer 1
RT286-10-069 | Group141 - Equipment | Dental assistant stool | Each | Offer 1
RT286-10-070 | Group142 - Equipment | Water Unit / Cuspidor / Spitton | Each | Offer 1
RT286-10-071 | Group143 - Equipment | Basic Dental chair | Each | Offer 1
RT286-10-072 | Group144 - Equipment | Dental treatment centre (Surgery unit) | Each | Offer 1
RT286-10-073 | Group145 - Equipment | Portabe Dental chair | Each | Offer 1
RT286-11-001 | Group146 - Equipment | Digital Amalgamator | Each | Offer 1
RT286-11-002 | Group147 - Equipment | Corded Laser Caries Detector (Tabletop) | Each | Offer 1
RT286-11-002 A | Group147- Accessory | Replacement tip | Each | Offer 1
RT286-11-002 B | Group147- Accessory | Replacement safety glasses | Each | Offer 1
RT286-11-003 | Group148 - Equipment | Single-Tooth Anesthesia System | Each | Offer 1
RT286-11-003 A | Group148- Accessory | Box of 50 X handpieces with needles | Each | Offer 1
RT286-11-004 | Group149 - Equipment | Pulp Tester | Each | Offer 1
RT286-11-004 A | Group149- Accessory | Batteries | Each | Offer 1
RT286-11-004 B | Group149- Accessory | Tips | Each | Offer 1
RT286-11-005 | Group150 - Equipment | Apex Locator | Each | Offer 1
RT286-11-005 A | Group150- Accessory | Rechargeable battery | Each | Offer 1
RT286-11-005 B | Group150- Accessory | Battery charger | Each | Offer 1
RT286-11-005 C | Group150- Accessory | 1 X Drive Unit | Each | Offer 1
RT286-11-005 D | Group150- Accessory | 2 X Lip hooks | Each | Offer 1
RT286-11-005 E | Group150- Accessory | Measuring wire | Each | Offer 1
RT286-11-005 F | Group150- Accessory | Fork (Endo file holder) | Each | Offer 1
RT286-11-006 | Group151 - Equipment | Fluorescence visualization unit | Each | Offer 1
RT286-11-006 A | Group151- Accessory | Caps per box to fit unit for individual patients | Each | Offer 1
RT286-11-006 B | Group151- Accessory | Sheets | Each | Offer 1
RT286-11-006 C | Group151- Accessory | Compatable Camera system for photo documentation | Each | Offer 1
RT286-11-007 | Group152 - Equipment | Piezoelectric surgery unit | Each | Offer 1
RT286-11-007 A | Group152- Accessory | 1 X Piezotome LED Handpiece | Each | Offer 1
RT286-11-007 B | Group152- Accessory | 1 X foot control | Each | Offer 1
RT286-11-007 C | Group152- Accessory | 1 X power cord | Each | Offer 1
RT286-11-007 D | Group152- Accessory | 1 X Irrigation bracket | Each | Offer 1
RT286-11-007 E | Group152- Accessory | 1 X Handpiece holder | Each | Offer 1
RT286-11-008 | Group153 - Equipment | Surgical Contra Angle Handpieces with 45 degree angle ( 1:2.7 with light) | Each | Offer 1
RT286-11-009 | Group154 - Equipment | Prosthodontic Electronic Screwdriver | Each | Offer 1
RT286-11-009 A | Group154- Accessory | 1 X Rechargeable battery | Each | Offer 1
RT286-11-009 B | Group154- Accessory | 1 X Handpiece drive -contrangle | Each | Offer 1
RT286-11-009 C | Group154- Accessory | 1 X Hexagon chucking system | Each | Offer 1
RT286-11-009 D | Group154- Accessory | 1 X Charging station | Each | Offer 1
RT286-11-010 | Group155 - Equipment | Cleaning, lubrication and purging system for dental handpieces | Each | Offer 1
RT286-11-010 A | Group155- Accessory | 1 X Service adapter | Each | Offer 1
RT286-11-010 B | Group155- Accessory | 1 X Charging cradle | Each | Offer 1
RT286-11-010 C | Group155- Accessory | 1 X Air hose | Each | Offer 1
RT286-11-010 D | Group155- Accessory | 1 X Window cover | Each | Offer 1
RT286-11-010 E | Group155- Accessory | 1 X 500ml of oil | Each | Offer 1
RT286-11-011 | Group156 - Equipment | Hydro Abrasion Cavity Preparation and Polishing System | Each | Offer 1
RT286-11-011 A | Group156- Accessory | 29 micron Aluminum Oxide powder (Fine cutting Powder) container. | Each | Offer 1
RT286-11-011 B | Group156- Accessory | 53 micron Aluminum Oxide powder(Course Cutting Powder) container. | Each | Offer 1
RT286-11-011 C | Group156- Accessory | Stain Removal container (Sodium Bicarbonite) | Each | Offer 1
RT286-11-011 D | Group156- Accessory | Polishing and profhylaxis Container1 | Each | Offer 1
RT286-11-011 E | Group156- Accessory | Disposal tips sizes 0.4/0.6/08mm | Each | Offer 1
RT286-11-012 | Group157 - Equipment | Endodontic Irrigation and Debridement Apparatus | Each | Offer 1
RT286-11-012 A | Group157- Accessory | Disposable tips, small, medium and large tips. | Each | Offer 1
RT286-11-012 B | Group157- Accessory | Protective sleeves for the unit. | Each | Offer 1
RT286-11-012 C | Group157- Accessory | Sequence card. | Each | Offer 1
RT286-11-013 | Group158 - Equipment | Rotating Corded Endodontic Motor, Handpiece and Apex Locator Combination | Each | Offer 1
RT286-11-013 A | Group158- Accessory | Rechargeable batteries | Each | Offer 1
RT286-11-013 B | Group158- Accessory | Stand (If separate from the charging unit) | Each | Offer 1
RT286-11-013 C | Group158- Accessory | Charging Unit with SABS plug (If separate from the stand) | Each | Offer 1
RT286-11-013 D | Group158- Accessory | Contra angle head 2.1 incl conection fork | Each | Offer 1
RT286-11-013 E | Group158- Accessory | Handpiece and lead without the head | Each | Offer 1
RT286-11-013 F | Group158- Accessory | 1 X mesurement cable and 2 hooks | Each | Offer 1
RT286-11-013 G | Group158- Accessory | Duel hook | Each | Offer 1
RT286-11-013 H | Group158- Accessory | Lip clips | Each | Offer 1
RT286-11-013 I | Group158- Accessory | Fork for apex mesuerments | Each | Offer 1
RT286-11-014 | Group159 - Equipment | Rotating Cordless Endodontic Motor and Handpiece Combination | Each | Offer 1
RT286-11-014 A | Group159- Accessory | Rechargeable Batteries | Each | Offer 1
RT286-11-014 B | Group159- Accessory | Stand | Each | Offer 1
RT286-11-014 C | Group159- Accessory | Charging Unit with SABS plug | Each | Offer 1
RT286-11-014 D | Group159- Accessory | Contra angle head | Each | Offer 1
RT286-11-014 E | Group159- Accessory | Handpiece without the head | Each | Offer 1
RT286-11-015 | Group160 - Equipment | Corded Rotating and reciprocating Endodontic Motor and Handpiece Combination | Each | Offer 1
RT286-11-015 A | Group161- Accessory | Rechargeable batteries | Each | Offer 1
RT286-11-015 B | Group161- Accessory | Stand (If separate from the charging unit) | Each | Offer 1
RT286-11-015 C | Group161- Accessory | Charging Unit with SABS plug (If separate from the stand) | Each | Offer 1
RT286-11-015 D | Group161- Accessory | Contra angle head | Each | Offer 1
RT286-11-015 E | Group161- Accessory | Handpiece without the head | Each | Offer 1
RT286-11-015 F | Group161- Accessory | Handpiece Stand | Each | Offer 1
RT286-11-015 G | Group161- Accessory | Spray Nozzle (F-type) | Each | Offer 1
RT286-11-015 H | Group161- Accessory | WaveOe Contra angle 6:1 ratio | Each | Offer 1
RT286-11-015 I | Group161- Accessory | Foot Pedal | Each | Offer 1
RT286-11-016 | Group162 - Equipment | Endodontic down Pack and Backfill Root Canal Machine | Each | Offer 1
RT286-11-016 A | Group162- Accessory | 1 X Plugger small ISO 40 with .03 taper | Each | Offer 1
RT286-11-016 B | Group162- Accessory | 1 X Plugger medium ISO 50 with .05 taper | Each | Offer 1
RT286-11-016 C | Group162- Accessory | 1 X Plugger large ISO 60 with .06 taper | Each | Offer 1
RT286-11-016 D | Group162- Accessory | 1 X pack of 10 cartridges 20 G | Each | Offer 1
RT286-11-016 E | Group162- Accessory | 1 X pack of 10 cartridges 23G | Each | Offer 1
RT286-11-016 F | Group162- Accessory | Replacement cartridge nuts | Each | Offer 1
RT286-11-016 G | Group162- Accessory | Thermanal respose tip | Each | Offer 1
RT286-11-016 H | Group162- Accessory | Replacement Cartrige Nut | Each | Offer 1
RT286-11-016 I | Group162- Accessory | Cartridge bending tool | Each | Offer 1
RT286-11-016 J | Group162- Accessory | Hand Cleaning Brush | Each | Offer 1
RT286-11-016 K | Group162- Accessory | Heat shield | Each | Offer 1
RT286-11-016 L | Group162- Accessory | GP Cartridges 20G | Each | Offer 1
RT286-11-016 M | Group162- Accessory | GP Cartridges 20G | Each | Offer 1
RT286-11-017 | Group163 - Equipment | Portable Handheld X Ray Unit | Each | Offer 1
RT286-11-017 A | Group163- Accessory | Rechargeable battery | Each | Offer 1
RT286-11-017 B | Group163- Accessory | Charging unit | Each | Offer 1
RT286-11-018 | Group164 - Equipment | Hypodermic Needle incinerator portable | Each | Offer 1
RT286-11-019 | Group165 - Equipment | Hypodermic Needle Incinerator Desktop | Each | Offer 1
RT286-11-020 | Group166 - Equipment | Platelet Rich Plasma Centrifuge (PRP) | Each | Offer 1
RT286-11-020 A | Group166- Accessory | PRP Tubes | Each | Offer 1
RT286-11-020 B | Group166- Accessory | Activator PRP tubes | Each | Offer 1
RT286-11-021 | Group167 - Equipment | Intra Oral Wireless Scanner with built in Colour Determination | Each | Offer 1
RT286-11-022 | Group168 - Equipment | Air purifiers | Each | Offer 1
RT286-11-023 | Group169 - Equipment | Hydroclorous acid generator | Each | Offer 1
RT286-11-024 | Group170 - Equipment | Extra oral vacuum/ extractor | Each | Offer 1
RT286-11-025 | Group171 - Equipment | Germicidal Ultraviolet Lamp | Each | Offer 1
RT286-12-001 | Group172 - Equipment | Hydrocyclone Amalgam IS18 SEPARATOR | Each | Offer 1
RT286-12-002 | Group173 - Equipment | Hydrocyclone ISO 6 Amalgam separator | Each | Offer 1
RT286-12-003 | Group174 - Equipment | Hydrocyclone Amalgam separator | Each | Offer 1
RT286-12-004 | Group175 - Equipment | Hydrocyclone ISO 4 Amalgam separator | Each | Offer 1
RT286-12-005 | Group176 - Equipment | Hydrocyclone Amalgam separator | Each | Offer 1
RT286-12-006 | Group177 - Equipment | Hydrocyclone Amalgam separator | Each | Offer 1