## RT32 Technical Specifications
RT32 TECHNICAL SPECIFICATIONS
BID DESCRIPTION: SUPPLY AND DELIVERY OF SURGICAL SUNDRIES TO THE STATE FOR A PERIOD OF 60MONTHS TO COMMENCE 1 MARCH 2023 TO 28 FERUARY 2028
Item Number | Product Name | Description | Type | Sizes | NSN Number/ Description | International Standard/s | South African Standard/s | Complies Yes / No | Provide Your Answer In This Column. You Are Advised To Be Straight To The Point
CATEGORY 1: MASKS
RT32-01-001 | Surgical mask Type II or higher | Surgical mask, good breathability; internal and external faces should be clearly identified Type II or higher. Fluid mask without eye shield, single use. The mask should be made from three layers of fabric and pleated horizontally. Have ear loops (preferred) or four tie backs for fastening to head. Fit a wide range of face shapes and sizes to permit easy breathing. Have a nose piece of flexible material at the front which enables the mask to be shaped around the nose and face to ensure a secure fit and good seal. There can also be a strip of foam rubber at the top edge at the back (optional). Packaging Requirements: Bacterial filtration efficiency, latex content, classification and type. Box of 50 Please note: Currently due to constraints testing bacterial filtration efficiency within South Africa, SAHPRA may declare conditional exclusions of and modifications to the tests of SANS 1866-1:2018 for South African manufacturers | Earloop | One Size | Number: 222000992 Description: Surgical mask type 2, pleated, blue 1's | N/A | SANS 1866-1: 2018 Medical devices Part 1: Medical face masks SANS 1866:2008 (Provides alternative tests to the BFE test) SANS 50149:2003 SAHPRA: Class A Non- sterile – exclusion from SAHPRA licence
RT32-01-002 | Surgical mask Type II or higher |  | Four tie back | One Size
RT32-01-003 | Mask surgical face, Type I | Mask surgical face, Type 1: Fluid mask with eye shield. Features of eye shield: Clear PVC 0.1mm thickness, 89% light transmittance, superior anti-fog coating, zero optical distortion. Features of mask: Made of four layers of fabric and pleated horizontally with three pleats. Have ear loops , A strip of foam at the top edge at the back and a nose piece of flexible material at the front which enable the mask to be shaped around the nose and face. Packaging should indicate: Bacterial filtration efficiency, latex content, classification and type. | Earloop | N/A
RT32-01-004 | Mask surgical face, Type I | Mask surgical face, Type 1: Fluid mask with eye shield. Features of eye shield: Clear PVC 0.1mm thickness, 89% light transmittance, superior anti-fog coating, zero optical distortion. Features of mask: Made of four layers of fabric and pleated horizontally with three pleats. Have four tie backs for fastening to head, A strip of foam at the top edge at the back and a nose piece of flexible material at the front which enable the mask to be shaped around the nose and face. Packaging should indicate: Bacterial filtration efficiency, latex content, classification and type. | Four tie back | N/A |  | N/A
RT32-01-005 | Respirators (FFP2) mask | Medical, particulate filtering half mask: Disposable Design requirements: Description: Particulate filtration respirators used for protection against airborne diseases such as tuberculosis. FFP2 or N95 respirator, or higher Good breathability with a design that does not collapse against the mouth (e.g. duckbill, cup shaped). • A particulate filtration half mask must cover the chin, mouth and nose; it may/ may not have an exhalation valve. Note: non-valved respirators are recommended for the COVID 19 pandemic. • The particulate filtration respirator must not disintegrate during continuous or intermittent use in a 8 hour work day. • Parts of the particulate filtration respirator that are more likely to come into contact with the wearer must be hypo-allergenic. • Filter performance must meet standards relevant to the specific respirator classification. • The respirator must have an adjustable/self-adjustable head harness which ensures ease of donning or removal of the respirator. (Please note elastic ear loops are NOT appropriate) • The particulate filtration respirator must have a filtration efficiency of at least 94% for FFP2 respirators and 95% for N95 and KN95 respirators. • The particulate filtration respirator shall have the equivalent of an external hydrophobic fabric layer for droplet protection. • Specific lot and brand testing is mandatory. Labelling Requirements: 1. Name or trademark 2. Filter efficiency/Classification 3. Approval number 4. Standard compliance 5. Size of the respirator A respirator mask, intended for use in a clinical environment, requires pre-approval from the NRCS. The manufacturer, distributor, or wholesaler of said mask, requires licensing approval from SAHPRA as a device establishment. Please read this specification in conjunction with the NDoH respirator quality assurance guideline, as more details are provided. | Duckbill shape | Small | Number: 222000994 (FFP2/ KN95) 222000993 (N95) Description: Mask respirator FFP2/ KN95 1'S OR Mask respirator N95 1'S | N/A | SANS 50149, Respiratory protective devices - Filtering half masks to protect against particles - Requirements, testing, marking SANS 1866, Part 2: Medical respirators SANS 10220, Standard for selection, use and maintenance of respiratory PPE (Guidance standard not tested against) VC8072:2011, as published in Government Notice No. R. 407 (Government Gazette No. 34272) of 13 May 2011 SAHPRA: Class B
RT32-01-006 | Respirators (FFP2) mask |  | Duckbill shape | Medium
RT32-01-007 | Respirators (FFP2) mask |  | Duckbill shape | Large
RT32-01-008 | Respirators (FFP2) mask |  | Flat Shape | Small
RT32-01-009 | Respirators (FFP2) mask |  | Flat Shape | Medium
RT32-01-010 | Respirators (FFP2) mask |  | Flat Shape | Large
RT32-01-011 | Respirators (FFP2) mask |  | Anatomal Shape | Small
RT32-01-012 | Respirators (FFP2) mask |  | Anatomal Shape | Medium
RT32-01-013 | Respirators (FFP2) mask |  | Anatomal Shape | Large
RT32-01-014 | Respirators (N95) mask |  | Duckbill shape | Small
RT32-01-015 | Respirators (N95) mask |  | Duckbill shape | Medium
RT32-01-016 | Respirators (N95) mask |  | Duckbill shape | Large
RT32-01-017 | Respirators (N95) mask |  | Flat Shape | Small
RT32-01-018 | Respirators (N95) mask |  | Flat Shape | Medium
RT32-01-019 | Respirators (N95) mask |  | Flat Shape | Large
RT32-01-020 | Respirators (N95) mask |  | Anatomal Shape | Small
RT32-01-021 | Respirators (N95) mask |  | Anatomal Shape | Medium
RT32-01-022 | Respirators (N95) mask |  | Anatomal Shape | Large
RT32-01-023 | Respirators (KN95) mask |  | Duckbill shape | Small
RT32-01-024 | Respirators (KN95) mask |  | Duckbill shape | Medium
RT32-01-025 | Respirators (KN95) mask |  | Duckbill shape | Large
RT32-01-026 | Respirators (KN95) mask |  | Flat Shape | Small
RT32-01-027 | Respirators (KN95) mask |  | Flat Shape | Medium
RT32-01-028 | Respirators (KN95) mask |  | Flat Shape | Large
RT32-01-029 | Respirators (KN95) mask |  | Anatomal Shape | Small
RT32-01-030 | Respirators (KN95) mask |  | Anatomal Shape | Medium
RT32-01-031 | Respirators (KN95) mask |  | Anatomal Shape | Large
CATEGORY 2: PERFORATORS
RT32-02-001 | Perforator | Perforator, membrane, amniotic, straight, made from Acrylonitrile Butadine Styrine(ABS) or a similar strong material, disposable Sterile Individually peel packed Latex free plastic DEHP free Box of 100 | Straight | N/A | N/A | N/A | N/A
RT32-02-002 | Perforator | Perforator, membrane, amniotic, Curved, made from Acrylonitrile Butadine Styrine(ABS) or a similar strong material, disposable Sterile Individually peel packed Latex free plastic DEHP free Box of 100 | Curved | N/A | N/A | N/A | N/A
CATEGORY 3: THERMOMETERS
RT32-03-001 | Thermometer | Thermometer, non-mercury. For single patient use Clinical thermometer to measure oral, axillary and rectal temperature in degrees celsius. Self-contained without circuits, moving parts or cover. No re-calibration or special storage requirements. Able to withstand exposure to high temperature conditions during storage or transit. Temperature reading to revert back to ambient temperature after use, to allow reuse by same patient after cleaning. Shelf life at least 5 years from date of manufacture. Suitable for use in all environments. | Non-mercury | N/A | N/A | To comply with the latest edition of ASTM 1299 standard or equivalent | N/A
RT32-03-002 | Thermometer | Thermometer, non-mercury, semi-disposable. For single patient use For adherence to the axilla. Safe for use with babies and children, non toxic and latex free. Must not contain mercury, batteries or other hazardous material and be submersible in water To allow continuous reading employing precision phase change technology (liquid crystal) for up to 48 hours after placement in the axilla. | Non-mercury, semi-disposable | N/A | N/A | N/A | N/A
RT32-03-003 | Thermometer | Thermometer, suitable for wall mounting, digital display, minimum/maximum, Range: -40 degrees C to +50 degrees Celsius with push button, must include a working battery for operation Certificate of calibration to be submitted | Suitable for wall mounting, digital display | N/A | N/A | N/A | N/A
RT32-03-004 | Thermometer | Thermometer refrigerator, sturdy, with hanging tab or other suitable mountable device, bi-metal dial, alcohol or digital type thermometer. Certificate of calibration to be submitted Range: -30 to +50 degrees Celsius. Ext dimensions: ±70x50mm or ±70 x 25mm | Refrigerator | N/A | N/A | N/A | N/A
CATEGORY 4: APPLICATORS
RT32-04-001 | Applicator | APPLICATOR (orange stick) Round, smooth hardwood 2mm thickness x 150mm Length Box of 500 | Orange stick | N/A | N/A | N/A | N/A
RT32-04-002 | Applicator | APPLICATOR haemolysis Sealable container of 100 | Haemolysis | N/A | N/A | N/A | N/A
CATEGORY 5: COVERS
RT32-05-001 | Cover | Cover, bedpan, flushable tissue type paper Dimension: 305 mm width x 76 mm height x 406 mm Length Box of 1000 | Bedpan | N/A | N/A | N/A | N/A
RT32-05-002 | Cover | Cover, urinal, flushable tissue type paper Dimension: 152 mm width x 76 mm height x 305 mm Length Box of 1000 | Urinal | N/A | N/A | N/A | N/A
CATEGORY 6: OVERSHOES AND CAPS
RT32-06-001 | Overshoes Fabric | Overshoe: Single use. Fabric to be ± 35g/m². To be made from durable, water-resistant non-slip material. Seam free undersole Blue. Elastic opening to cover shoe. Suitable for all shoe sizes up to size 7 (South African). Overshoe to cover entire shoe. Must be comfortable. Must not be too thin. Box of 100 | Fabric | Suitable for all shoe sizes up to size 7 (South African) | N/A | N/A | N/A
RT32-06-002 | Overshoes Fabric | Overshoe: Single use. Fabric to be ± 35g/m². To be made from durable, water-resistant non-slip material. Seam free undersole Blue. Elastic opening to cover shoe. Suitable for all shoe sizes 8-15 (South African). Overshoe to cover entire shoe. Must be comfortable. Must not be too thin. Box of 100 | Fabric | Suitable for all shoe sizes 8-15 (South African) | N/A | N/A | N/A
RT32-06-003 | Overshoe, plastic, single use | Overshoe, plastic, single use Elasticated opening Forensic and Isolation Units Suitable for all shoes sizes up to size 7 (South African). Thickness of plastic not less 100 Box of 100 | Plastic single use | Suitable for all shoe sizes up to size 7 (South African) | N/A | N/A | N/A
RT32-06-004 | Overshoe, plastic, single use | Overshoe, plastic, single use Elasticated opening Forensic and Isolation Units Suitable for all shoes sizes from 8-15 (South African). Thickness of plastic not less 100 Box of 100 | Plastic single use | Suitable for all shoe sizes 8-15 (South African) | N/A | N/A | N/A
RT32-06-005 | Cap | Cap, beret type, (similar to the mop cap) single use, diameter between 45 and 63cm, non-woven fabric, the frabic weight must be between 9 and 35g/m² with elastic thread all the way around the edge, attached by an overlocked stitch.Must be stretchable. Must not come apart on head. Must last at least 5 hours. Colour: Blue or Green Box of 100 | Beret (similar to the mop cap) | N/A | N/A | N/A | N/A
RT32-06-006 | Cap | Cap, theatre, bonnet type with ties, single use, light weight non-woven fabric ± 25g/m². Back panel to be gathered with an elastic band. Adult one size fit all in order to cover hair. Cap must not be too thin. Must not come apart. Must last at least 5 hours. Colour: Blue or Green Box of 100 | Theatre, bonnet type with ties | N/A | N/A | N/A | N/A
RT32-06-007 | Cap | Cap, theatre, balaclava type, single use. Light weight non-woven fabric between 18 and 22g/m². The hood material shall be fluid repellent and non-linting. Hood to fit over head, Must be made of three pieces of non-woven material, i.e. one oval crown and two side pieces. Must not come apart on head. Must last at least 5 hours. Colour: Blue or Green Box of 100 | Theatre, balaclava type | N/A | N/A | N/A | N/A
CATEGORY 7: SLEEVE PROTECTORS
RT32-07-001 | Sleeve, protector, single use | Sleeve, protector, single use Arm sleeve cover, Surgical disposable low density polythene material, Waterproof Length: 45cm Width: 20cm Double open ended Elasticized both ends Non-sterile | N/A | N/A | N/A | N/A | N/A
RT32-07-002 | Sleeve, surgical, re-useable PVC material, | Sleeve, surgical, re-useable PVC material, Length: 45cm Circumference: 40cm Double open ended Elastisized both sides | N/A | One Size | N/A | N/A | N/A
CATEGORY 8: BRUSHES
RT32-08-001 | Brush | Brush, continuously re-usable without distortion or change in quality | Surgical, nail, autoclavable | N/A | N/A | N/A | N/A
RT32-08-002 | Brush | Brush, single use in theatres with nail cleaner | Surgical, nail, sterile | N/A | N/A | N/A | N/A
CATEGORY 9: IDENTIFICATION BANDS
RT32-09-001 | Band, personal identification of plastic | Band, personal identification of plastic / other suitable non-abrasive material and with the following features: - Non toxic, hypo-allergenic, light-weight, durable - Must allow for indelible printing of patient`s particulars with ordinary ink pen - To be supplied with a locking device which cannot ordinarily be opened after initial closure Box of 100 For premature babies and neonates. To fit ankles with a circumference of 6 cm - 12cm The warning `FOR NEONATES ONLY` must be printed Colour: Pink | Premature babies and neonates | N/A | N/A | N/A | N/A
RT32-09-002 | Band, personal identification of plastic | Band, personal identification of plastic / other suitable non-abrasive material and with the following features: - Non toxic, hypo-allergenic, light-weight, durable - Must allow for indelible printing of patient`s particulars with ordinary ink pen - To be supplied with a locking device which cannot ordinarily be opened after initial closure Box of 100 For children. To fit ankles with circumference 9 cm - 18 cm Total circumference at least 18cm Colour: Pink | Children | N/A | N/A | N/A | N/A
RT32-09-003 | Band, personal identification of plastic | Band, personal identification of plastic / other suitable non-abrasive material and with the following features: - Non toxic, hypo-allergenic, light-weight, durable - Must allow for indelible printing of patient`s particulars with ordinary ink pen - To be supplied with a locking device which cannot ordinarily be opened after initial closure For children To fit ankles with circumference 9 cm - 18 cm Total circumference at least 18cm Colour: White Box of 100 | Children | N/A | N/A | N/A | N/A
RT32-09-004 | Band, personal identification of plastic | Band, personal identification of plastic / other suitable non-abrasive material and with the following features:- Non toxic, hypo-allergenic, light-weight, durable- Must allow for indelible printing of patient`s particulars with ordinary ink pen- To be supplied with a locking device which cannot ordinarily be opened after initial closure. For adults and to fit circumference 12 cm - 28 cmTotal circumference at least 28cm Colour YELLOW Box of 100 | Adult | N/A | N/A | N/A | N/A
CATEGORY 10: EYE COVERS
RT32-10-001 | Eye cover/posy for phototherapy | Eye cover/posy for phototherapy. Protection against eye damage from light in the blue-green spectrum (wavelengths 430-490nm). The material is to be soft, hypo-allergenic. The cover is to be able to stay in place during infant movement. There shall be no pressure on the eyeball. Latex free. A light transmission test will be required |  | Small | N/A | N/A | N/A
RT32-10-002 | Eye cover/posy for phototherapy | Eye cover/posy for phototherapy. Protection against eye damage from light in the blue-green spectrum (wavelengths 430-490nm). The material is to be soft, hypo-allergenic. The cover is to be able to stay in place during infant movement. There shall be no pressure on the eyeball. Latex free. A light transmission test will be required Pack size:each |  | Medium | N/A | N/A | N/A
RT32-10-003 | Eye cover/posy for phototherapy | Eye Cover/Posy for phototherapy. Protection against eye damage from light in the blue-green spectrum (wavelengths 430-490nm). The material is to be soft, hypo-allergenic. The cover is to be able to stay in place during infant movement. There shall be no pressure on the eyeball. Latex free A light transmission test will be required Pack size:each |  | Large | N/A | N/A | N/A
RT32-10-004 | Eye Shield Cartella (to fit both eyes) | Shield, eye, Cartella, universal (to fit both eyes), Make of impact-resistant plastic to be used over eye pad dressing that is applied after surgery of any eye procedure to protect the eye. Shield must have airvent holes Colour: Clear plastic Size: One size fit all Box of 100 |  | One size fit all | N/A | N/A | N/A
CATEGORY 11: CERVICAL COLLARS
RT32-11-001 | Collar | Collar, cervical, made from high density flexible polyether polyurethane foam, 19-20mm thick Covered with cotton woven mesh. Posterior fastening with velcro strip 100mm in length Anatomically shaped edges of foam to be rounded off Width: Under chin: approx. 80mm under chin, 70mm at back Length: 430mm | Cervical | Small | N/A | N/A | N/A
RT32-11-002 | Collar | Collar, cervical, made from high density flexible polyether polyurethane foam, 19-20mm thick Covered with cotton woven mesh. Posterior fastening with velcro strip 100mm in length Anatomically shaped edges of foam to be rounded off Width: Under chin: approx. 90mm under chin, 80mm at back. Length: 500mm | Cervical | Medium | N/A | N/A | N/A
RT32-11-003 | Collar | Collar, cervical, made from high density flexible polyether polyurethane foam, 19-20mm thick Covered with cotton woven mesh. Posterior fastening with velcro strip 100mm in length Anatomically shaped edges of foam to be rounded off Width: Under chin: approx. 90mm under chin, 90mm at back Length: 540mm | Cervical | Large | N/A | N/A | N/A
RT32-11-004 | Collar | Collar, cervical, made from high density flexible polyether polyurethane foam, 19-20mm thick Covered with cotton woven mesh. Posterior fastening with velcro strip 100mm in length Anatomically shaped edges of foam to be rounded off Width: Under chin: approx. 90mm under chin, 90mm at back Length: 600mm | Cervical | X-Large | N/A | N/A | N/A
RT32-11-005 | Collar and cuff | Collar and cuff Made from soft plastic foam 50 mm x 10 mm thick Covered with cotton woven tubular mesh Length: 4 m Bidder to state cord length on the packaging of product |  | XX-Large | N/A | N/A | N/A
CATEGORY 12: VAGINAL SPECULUMS
RT32-12-001 | Speculum | Speculum, vaginal, Cusco, surgical clean or sterile, single use, clear, polystyrene, light-weight, robust, no sharp edges Consists of two duckbill shaped blades, hinged at one extremity forming two handle pieces with a ratchet at an angle to effect opening of the blades. Individually packed Type of sterilization method must be indicated on the packaging Each Size: Small (75mm x 25mm) | Sterile | small (75mm x 25mm) | N/A | N/A | N/A
RT32-12-002 | Speculum | Speculum, vaginal, Speculum, vaginal, Cusco, surgical clean, single use, clear, polystyrene, light-weight, robust, no shar surgical clean or sterile p edges Consists of two duckbill shaped blades, hinged at one extremity forming two handle pieces with a ratchet at an angle to effect opening of the blades. Individually packed Type of sterilization method must be indicated on the packaging | Sterile | Medium (85mm x 30mm) | N/A | N/A | N/A
RT32-12-003 | Speculum | Speculum, vaginal, Speculum, vaginal, Cusco, surgical clean or sterile , single use, clear, polystyrene, light-weight, robust, no sharp edges Consists of two duckbill shaped blades, hinged at one extremity forming two handle pieces with a ratchet at an angle to effect opening of the blades. Individually packed Type of sterilization method must be indicated on the packaging | Sterile | Large (100mm x 35mm) | N/A | N/A | N/A
RT32-12-004 | Speculum | Speculum, vaginal, Speculum, vaginal, Cusco, surgical clean or sterile , single use, clear, polystyrene, light-weight, robust, no sharp edges Consists of two duckbill shaped blades, hinged at one extremity forming two handle pieces with a ratchet at an angle to effect opening of the blades. Individually packed Type of sterilization method must be indicated on the packaging | Sterile | X-Large (110mm x 37mm) | N/A | N/A | N/A
RT32-12-004 | Speculum | Speculum, vaginal, Speculum, vaginal, Cusco, surgical clean or sterile , single use, clear, polystyrene, light-weight, robust, no sharp edges Consists of two duckbill shaped blades, hinged at one extremity forming two handle pieces with a ratchet at an angle to effect opening of the blades. Individually packed Type of sterilization method must be indicated on the packaging | Sterile | X-Large (110mm x 37mm) | N/A | N/A | N/A
CATEGORY 13: SANITARY PADS
RT32-13-001 | Sanitary pad, post delivery, maternity without loops Size width 90mm x length 240mm | Sanitary pad; post delivery, maternity without loops, autoclavable non- woven with wood pulp for absorbency and a polyshield to prevent strike through. Size width 90mm x length 240mm. Non sterile. Pack of 12 | Unscented | N/A | N/A | N/A | N/A
RT32-13-002 | Sanitary pad, maternity with loops. Autoclavable Must comply with the lastest issue of SANS 1043. pack of 12 | Sanitary pad, maternity with loops. autoclavable. non-woven absorbent cotton material in a tubular loose knitted fabric sleeve. it should be retangular and must comply with the lastest issue of SANS 1043. Pack of 12 | Unscented | N/A | N/A | N/A | SANS 1043:2021
RT32-13-003 | Sanitary pad,regular with adhesive back strip and wings Size: ±20cm in length x ±7cm in width with a wing span of ±15cm | Sanitary pad; P regular with adhesive back strip and wings. Size: ±20cm in length x ±7cm in width with a wing span of ±15cm Pad must be • Ultra-thin • soft and absorbent • for light to medium flow • absorbed ±235ml of fluid • covered with a cotton layer for comfort • unscented and • have wings for stay put and must have a leak protection. Pack of 16 |  | N/A | N/A | N/A | N/A
CATEGORY 14: INCONTINENCE PADS
RT32-14-001 | Incontinence, pad; slight urinary incontinence | Incontinence, pad; slight urinary incontinence. outer spunbonded non-woven cover and inner pulp; with mass of not less than 18mg and inner pulp with a mass of not less than 15mg. The pad must have an adhesive must strip on the backing to secure it to the patient`s underwear and must meet the requirements of ISO13485 and ISO 16021. absorbency of 210 to 230ml. | Unisex | Small | N/A | ISO13485 and ISO 16021 | N/A
RT32-14-002 | Incontinence, pad; slight urinary incontinence | INCONTINENCE, PAD; slight urinary incontinence, outer spunbonded non-woven cover and inner pulp; with mass of not less than 18mg and inner pulp with a mass of not less than 15mg, the pad must have an adhesive must strip on the backing to secure it to the patient`s underwear, and must meet the requirements of ISO 13485 AND ISO 16021, absorbency of 210 to 230ml | Unisex | Medium | N/A | ISO13485 and ISO 16021 | N/A
RT32-14-003 | Incontinence, pad; slight urinary incontinence | Sanitary pad; slight urinary incontinence. Outer spunbonded non-woven cover and inner pulp; with mass of not less than 18mg and inner pulp with a mass of not less than 15mg. The pad must have an adhesive strip on the backing to secure it to the patient`s underwear. And must meet the requirements of ISO 13485 and ISO 16021. Absorbency of 210 to 230ml | Unisex | Large | N/A | ISO13485 and ISO 16021 | N/A
RT32-14-004 | Sanitary pad; slight urinary incontinence | Sanitary pad; slight urinary incontinence. Outer spunbonded non-woven cover and inner pulp; with mass of not less than 18mg and inner pulp with a mass of not less than 15mg. The pad must have an adhesive strip on the backing to secure it to the patient`s underwear. And must meet the requirements of ISO 13485 and ISO 16021. Absorbency of 210 to 230ml | Unisex | X-Large | N/A | ISO13485 and ISO 16021 | N/A
RT32-14-005 | Incontinence pad; for moderate urinary and faecal incontinence | Incontinence pad; for moderate urinary and faecal incontinence. fluid distribution layer, ultra absorbent core, absorbent bed, complete non woven covering soft elastic in the crotch area, textile back sheet. Must provide maximum fluid retention. Should be applied with panties/ pants. Fluid protection, absorbency core 3-4 layers. Non woven polypropylene, cuffs, elastics for legs gathers and standing cuffs. Must meet the requirements of ISO 13485 and iso 16021 absorbention Of 1400 to 1650ML. | Unisex | Small | N/A | ISO13485 and ISO 16021 | N/A
RT32-14-006 | Incontinence pad; for moderate urinary and faecal incontinence | Sanitary pad; for moderate urinary and faecal incontinence. fluid distribution layer, ultra absorbent core, absorbent bed, complete non woven covering soft elastic in the crotch area, textile back sheet. Must provide maximum fluid retention. Should be applied with panties/ pants. Fluid protection, absorbency core 3-4 layers. Non woven polypropylene, cuffs, elastics for legs gathers and standing cuffs. Must meet the requirements of ISO 13485 and ISO 16021 absorbention of 1400 to 1650ml. | Unisex | Medium | N/A | ISO13485 and ISO 16021 | N/A
RT32-14-007 | Incontinence pad; for moderate urinary and faecal incontinence | Sanitary pad; for moderate urinary and faecal incontinence, fluid distribution layer, ultra absorbent core, absorbent bed, complete non woven covering soft elastic in the crotch area, textile back sheet, must provide maximum fluid retention, should be applied with panties/ pants, fluid protection, absorbency core 3-4 layers, non woven polypropylene, cuffs, elastics for legs gathers and standing cuffs, must meet the requirements of ISO13485 and ISO 16021 absorbention of 1400 to 1650ml | Unisex | Large | N/A | ISO13485 and ISO 16021 | N/A
RT32-14-008 | Incontinence pad; for moderate urinary and faecal incontinence | Sanitary pad; for moderate urinary and faecal incontinence, fluid distribution layer, ultra absorbent core, absorbent bed, complete non woven covering soft elastic in the crotch area, textile back sheet, must provide maximum fluid retention, should be applied with panties/ pants, fluid protection, absorbency core 3-4 layers, non woven polypropylene, cuffs, elastics for legs gathers and standing cuffs, must meet the requirements of ISO13485 and ISO 16021 absorbention of 1400 to 1650ml | Unisex | X-Large | N/A | ISO13485 and ISO 16021 | N/A
RT32-14-009 | Incontinence pad, men | Sanitary pad; anatomically shaped, men, distribution layer in complete non-woven covering and must meet the requirements of ISO13485 and ISO 16021, absorbency be of 140 to 150ml | Anatomically shaped | One Size |  | ISO13485 and ISO 16021
CATEGORY 15: INCONTINENCE DIAPERS
RT32-15-001 | Incontinence diaper | Incontinence diaper; daytime, unisex. Material non-woven lining and outer plastic material; wrap over type with peel-off adhesive tape at wing extremities, at the centre of the diaper and sandwiched between the non-woven lining and a plastic outer, shall be an absorbent pad consisting of pulp filler with a layer of absorbent compound or a non-woven on one side and encased in a liner to wick away fluid which shall extent beyond the ends of the filler and adequately sealed with longitudinal contrivement flap. The non-woven lining and absorbent bed shall be bonded to the outer plastic material. Must meet the requirements of ISO13485 AND ISO 16021 | Daytime, Unisex | Small | N/A | ISO13485 and ISO 16021 | N/A
RT32-15-002 | Incontinence diaper | Incontinence diaper; DAYTIME, unisex. Material non-woven lining and outer plastic material; wrap over type with peel-off adhesive tape at wing extremities, at the centre of the diaper and sandwiched between the non-woven lining and a plastic outer, shall be an absorbent pad consisting of pulp filler with a layer of absorbent compound or a non-woven on one side and encased in a liner to wick away fluid which shall extent beyond the ends of the filler and adequately sealed with longitudinal contrivement flap. The non-woven lining and absorbent bed shall be bonded to the outer plastic material. must meet the requirements of ISO13485 AND ISO 16021SIZE | Daytime, Unisex | Medium | N/A | ISO13485 and ISO 16021 | N/A
RT32-15-003 | Incontinence diaper | Incontinence diaper; daytime,unisex. Material non-woven lining and outer plastic material; wrap over type with peel-off adhesive tape at wing extremities, at the centre of the diaper and sandwiched between the non-woven lining and a plastic outer, shall be an absorbent pad consisting of pulp filler with a layer of absorbent compound or a non-woven on one side and encased in a liner to wick away fluid which shall extent beyond the ends of the filler and adequately sealed with longitudinal contrivement flap. The non-woven lining and absorbent bed shall be bonded to the outer plastic material. must meet the requirements of ISO13485 AND ISO 16021 | Daytime, Unisex | Large | N/A | ISO13485 and ISO 16021 | N/A
RT32-15-004 | Incontinence diaper | Incontinence diaper; daytime,unisex. Material non-woven lining and outer plastic material; wrap over type with peel-off adhesive tape at wing extremities, at the centre of the diaper and sandwiched between the non-woven lining and a plastic outer, shall be an absorbent pad consisting of pulp filler with a layer of absorbent compound or a non-woven on one side and encased in a liner to wick away fluid which shall extent beyond the ends of the filler and adequately sealed with longitudinal contrivement flap. The non-woven lining and absorbent bed shall be bonded to the outer plastic material. must meet the requirements of ISO13485 AND ISO 16021 | Daytime, Unisex | X-Large | N/A | ISO13485 and ISO 16021 | N/A
RT32-15-005 | Incontinence diaper | Incontinence diaper; daytime,unisex. Material non-woven lining and outer plastic material; wrap over type with peel-off adhesive tape at wing extremities, at the centre of the diaper and sandwiched between the non-woven lining and a plastic outer, shall be an absorbent pad consisting of pulp filler with a layer of absorbent compound or a non-woven on one side and encased in a liner to wick away fluid which shall extent beyond the ends of the filler and adequately sealed with longitudinal contrivement flap. The non-woven lining and absorbent bed shall be bonded to the outer plastic material. must meet the requirements of ISO13485 AND ISO 16021 | Daytime, Unisex | XX-Large | N/A | ISO13485 and ISO 16021 | N/A
RT32-15-006 | Incontinence diaper | Incontinence diaper; overnight, Unisex. Material non-woven lining and outer plastic material; wrap over type with peel-off adhesive tape at wing extremities, at the centre of the diaper and sandwiched between the non-woven lining and a plastic outer, shall be an absorbent pad consisting of pulp filler with a layer of absorbent compound or a non-woven on one side and encased in a liner to wick away fluid which shall extent beyond the ends of the filler and adequately sealed with longitudinal contrivement flap. The non-woven lining and absorbent bed shall be bonded to the outer plastic material. must meet the requirements of ISO13485 AND ISO 16021 | Overnight, Unisex | Small | N/A | ISO13485 and ISO 16021 | N/A
RT32-15-007 | Incontinence diaper | Incontinence diaper; overnight, unisex. Material non-woven lining and outer plastic material; wrap over type with peel-off adhesive tape at wing extremities, at the centre of the diaper and sandwiched between the non-woven lining and a plastic outer, shall be an absorbent pad consisting of pulp filler with a layer of absorbent compound or a non-woven on one side and encased in a liner to wick away fluid which shall extent beyond the ends of the filler and adequately sealed with longitudinal contrivement flap. The non-woven lining and absorbent bed shall be bonded to the outer plastic material. must meet the requirements of ISO13485 AND ISO 16021 | Overnight, Unisex | Medium | N/A | ISO13485 and ISO 16021 | N/A
RT32-15-008 | Incontinence diaper | Incontinence diaper; overnight, unisex. Material non-woven lining and outer plastic material; wrap over type with peel-off adhesive tape at wing extremities, at the centre of the diaper and sandwiched between the non-woven lining and a plastic outer, shall be an absorbent pad consisting of pulp filler with a layer of absorbent compound or a non-woven on one side and encased in a liner to wick away fluid which shall extent beyond the ends of the filler and adequately sealed with longitudinal contrivement flap. The non-woven lining and absorbent bed shall be bonded to the outer plastic material. must meet the requirements of ISO13485 AND ISO 16021 | Overnight, Unisex | Large | N/A | ISO13485 and ISO 16021 | N/A
RT32-15-009 | Incontinence diaper | Incontinence diaper; overnight, unisex. Material non-woven lining and outer plastic material; wrap over type with peel-off adhesive tape at wing extremities, at the centre of the diaper and sandwiched between the non-woven lining and a plastic outer, shall be an absorbent pad consisting of pulp filler with a layer of absorbent compound or a non-woven on one side and encased in a liner to wick away fluid which shall extent beyond the ends of the filler and adequately sealed with longitudinal contrivement flap. The non-woven lining and absorbent bed shall be bonded to the outer plastic material. must meet the requirements of ISO13485 AND ISO 16021 | Overnight, Unisex | X-Large | N/A | ISO13485 and ISO 16021 | N/A
RT32-15-010 | Incontinence diaper | Incontinence diaper; overnight, unisex. Material non-woven lining and outer plastic material; wrap over type with peel-off adhesive tape at wing extremities, at the centre of the diaper and sandwiched between the non-woven lining and a plastic outer, shall be an absorbent pad consisting of pulp filler with a layer of absorbent compound or a non-woven on one side and encased in a liner to wick away fluid which shall extent beyond the ends of the filler and adequately sealed with longitudinal contrivement flap. The non-woven lining and absorbent bed shall be bonded to the outer plastic material. must meet the requirements of ISO13485 AND ISO 16021 | Overnight, Unisex | XX-Large | N/A | ISO13485 and ISO 16021 | N/A
CATEGORY 16: DISPOSABLE DIAPERS
RT32-16-001 | Disposable diaper | Disposable diaper; for use with micro premature babies weight approximately 400 to 700 grams. 2 refastenable tapes, elasticated legs with 2 strands of elastic, absorbent material, super absorbent mass approximately 3 grams. Waterproof outer layer. Latex free | Latex Free | Micro premature baby. Weight approximately 400 to 700 grams | N/A | N/A | SANS 8841:2019
RT32-16-002 | Disposable diaper | Disposable diaper; for use with premature babies weight approximately 800 to 1500 grams. 2 refastenable tapes, elasticated legs with 2 strands of elastic, absorbent material, super absorbent mass approximately 3 grams. Waterproof outer layer | Latex Free | Premature babies. Weight approximately 800 to 1500 grams | N/A | N/A | SANS 8841:2019
RT32-16-003 | Disposable diaper | Disposable diaper; for use with new born babies weight approximately 1.6kg to 2.5kg. 2 refastenable tapes, elasticated legs with 2 strands of elastic, absorbent material, super absorbent mass approximately 3 grams. Waterproof outer layer | Latex Free | New born. Weight approximately 1.6kg to 2.5kg | N/A | N/A | SANS 8841:2019
RT32-16-004 | Disposable diaper | Disposable diaper; for use with small infant weight approximately 3kg to 6kg. 2 refastenable tapes, elasticated legs with 2 strands of elastic, absorbent material, super absorbent mass approximately 3 grams. Waterproof outer layer | Latex Free | Small infant. Weight approximately 3kg to 6kg | N/A | N/A | SANS 8841:2019
RT32-16-005 | Disposable diaper | Disposable diaper; for use with infant weight approximately 7kg to 12kg. 2 refastenable tapes, elasticated legs with 2 strands of elastic, absorbent material, super absorbent mass approximately 3 grams. Waterproof outer layer | Latex Free | Infant. Weight approximately 7kg to 12kg | N/A | N/A | SANS 8841:2019
RT32-16-006 | Disposable diaper | Disposable diaper; for use with toddler weight approximately 13kg to 20kg. 2 refastenable tapes, elasticated legs with 2 strands of elastic, absorbent material, super absorbent mass approximately 3 grams. Waterproof outer layer | Latex Free | Toddler. Weight approximately 13kg to 20kg | N/A | N/A | SANS 8841:2019
RT32-16-007 | Disposable diaper | Disposable diaper; for use with child weight approximately 21kg to 30kg. 2 refastenable tapes, elasticated legs with 2 strands of elastic, absorbent material, super absorbent mass approximately 3 grams. Waterproof outer layer | Latex Free | Child. Weight approximately 21kg to 30kg | N/A | N/A | SANS 8841:2019
CATEGORY 17: GOWNS , APRONS AND SCRUB SUITS
RT32-17-001 | Gown: Disposable | Material Requirements Isolation gowns to be manufactured from protective fabrics with fluid-resistant properties. Design Requirements Design criteria to be consistent for Isolation gowns. Netted or knitted cuffs or thumb/finger loops to anchor sleeves in place, Tape-tab neck closure, Ties for waist. Dimension Requirements Size to be clearly identified on the garment and packaging. Length: (from shoulder to hem) 116cm for large and 127cm for extra-large Colour: Yellow 17.1 A protective 3- layer SMS (Spunbond/Meltblown/Spunbond) polypropylene non-woven product, 17.2 Laminated, coated or calendered materials that comply with the performance criteria as indicated by the stipulated standards | Isolation | Small | Number: 222000998 Description: Gowns - Isolation 1'S | N/A | SANS 53795 as per standard performance (fluid penetration resistance = 20cm) Surgical drapes, gowns and clean air suits, used as medical devices for patients, clinical staff and equipment - General requirements for manufacturers, processors and products, test methods, performance requirements and performance levels SAHPRA: Class A Non- sterile – exclusion from SAHPRA licence
RT32-17-002 | Gown: Disposable |  | Isolation | Medium
RT32-17-003 | Gown: Disposable |  | Isolation | Large
RT32-17-004 | Gown: Disposable |  | Isolation | X-Large
RT32-17-005 | Gown: Disposable |  | Isolation | XX-Large
RT32-17-006 | Gown: Disposable |  | Isolation | XXX-Large
RT32-17-007 | Gown: Disposable |  | Isolation | XXXX-Large
RT32-17-008 | Hospital gowns | Gown, hospital, non-woven construction with a minimum mass of 30 g/m 2, short sleeves with belt and neck closure, no collar and opaque Single use Colour: Any colours excluding white To comply with SANS 53795 standard Individually wrapped | N/A | Child (Age 3yrs - 6yrs) | N/A | N/A | SANS 53795
RT32-17-009 | Hospital gowns | Gown, hospital, non-woven construction with a minimum mass of 30 g/m 2, short sleeves with belt and neck closure, no collar and opaque Single use Colour: Any colours excluding white To comply with SANS 53795 standard Individually wrapped | N/A | Child (Age 6 yrs - 9 yrs) | N/A | N/A | SANS 53795
RT32-17-010 | Hospital gowns | Gown, hospital, non-woven construction with a minimum mass of 30 g/m 2, short sleeves with belt and neck closure, no collar and opaque Single use Colour: Any colours excluding white To comply with SANS 53795 standard Individually wrapped | N/A | Medium | N/A | N/A | SANS 53795
RT32-17-011 | Hospital gowns | Gown, hospital, non-woven construction with a minimum mass of 30 g/m 2, short sleeves with belt and neck closure, no collar and opaque Single use Colour: Any colours excluding white To comply with SANS 53795 standard Individually wrapped | N/A | Large | N/A | N/A | SANS 53795
RT32-17-012 | Hospital gowns | Gown, hospital, non-woven construction with a minimum mass of 30 g/m 2, short sleeves with belt and neck closure, no collar and opaque Single use Colour: Any colours excluding white To comply with SANS 53795 standard Individually wrapped | N/A | X-Large | N/A | N/A | SANS 53795
RT32-17-013 | Hospital gowns | Gown, hospital, non-woven construction with a minimum mass of 30 g/m 2, short sleeves with belt and neck closure, no collar and opaque Single use Colour: Any colours excluding white To comply with SANS 53795 standard Individually wrapped | N/A | XX-Large | N/A | N/A | SANS 53795
RT32-17-014 | Surgical Gown | Gown, surgical, Reinforced, non-woven Body ± 54 g/m², sleeves ± 66 g/m² (Need certification that state the material mass) Long sleeves with cuffs. Reinforced in chest and forearm areas. Water and alcohol repellent. (Need certification that state water and alcohol repellent ) Non-toxic, non-irritant and bacterial migration resistant in wet and dry conditions. Low linting and low flammability qualities. Sterility certification, The content shall be wrapped in a non-woven towel at least 600x600mm which can be opened in a sterile manner. Packed in a peel package which is clearly marked with a list of the contents and expiry date. Individually double peel packed. Hand Towel, Medical X2 to be included in the Gown Pack C Fold, 2 ply, White, Virgin Pulp. Lint free. High absorbency. 40 g/m2 minimum ± 300 X 200 mm folded. Length of the gown must be approximately 3/4 length (between ankle and mid-calf). To comply with SANS 53795 standard Colour: BLUE | Reinforced | Small | N/A | N/A | SANS 53795
RT32-17-015 | Surgical Gown | Gown, surgical, Reinforced, non-woven Body ± 54 g/m², sleeves ± 66 g/m² (Need certification that state the material mass) Long sleeves with cuffs. Reinforced in chest and forearm areas. Water and alcohol repellent. (Need certification that state water and alcohol repellent ) Non-toxic, non-irritant and bacterial migration resistant in wet and dry conditions. Low linting and low flammability qualities. Sterility certification, The content shall be wrapped in a non-woven towel at least 600x600mm which can be opened in a sterile manner. Packed in a peel package which is clearly marked with a list of the contents and expiry date. Individually double peel packed. Hand Towel, Medical X2 to be included in the Gown Pack C Fold, 2 ply, White, Virgin Pulp. Lint free. High absorbency. 40 g/m2 minimum ± 300 X 200 mm folded. Length of the gown must be approximately 3/4 length (between ankle and mid-calf). To comply with SANS 53795 standard Colour: BLUE | Reinforced | Medium | N/A | N/A | SANS 53795
RT32-17-016 | Surgical Gown | Gown, surgical, Reinforced, non-woven Body ± 54 g/m², sleeves ± 66 g/m² (Need certification that state the material mass) Long sleeves with cuffs. Reinforced in chest and forearm areas. Water and alcohol repellent. (Need certification that state water and alcohol repellent ) Non-toxic, non-irritant and bacterial migration resistant in wet and dry conditions. Low linting and low flammability qualities. Sterility certification, The content shall be wrapped in a non-woven towel at least 600x600mm which can be opened in a sterile manner. Packed in a peel package which is clearly marked with a list of the contents and expiry date. Individually double peel packed. Hand Towel, Medical X2 to be included in the Gown Pack C Fold, 2 ply, White, Virgin Pulp. Lint free. High absorbency. 40 g/m2 minimum ± 300 X 200 mm folded. Length of the gown must be approximately 3/4 length (between ankle and mid-calf).To comply with SANS 53795 standard Colour: BLUE | Reinforced | Large |  |  | SANS 53795
RT32-17-017 | Surgical Gown | Gown, surgical, Reinforced, non-woven Body ± 54 g/m², sleeves ± 66 g/m² (Need certification that state the material mass) Long sleeves with cuffs. Reinforced in chest and forearm areas. Water and alcohol repellent. (Need certification that state water and alcohol repellent ) Non-toxic, non-irritant and bacterial migration resistant in wet and dry conditions. Low linting and low flammability qualities. Sterility certification, The content shall be wrapped in a non-woven towel at least 600x600mm which can be opened in a sterile manner. Packed in a peel package which is clearly marked with a list of the contents and expiry date. Individually double peel packed. Hand Towel, Medical X2 to be included in the Gown Pack C Fold, 2 ply, White, Virgin Pulp. Lint free. High absorbency. 40 g/m2 minimum ± 300 X 200 mm folded. Length of the gown must be approximately 3/4 length (between ankle and mid-calf).To comply with SANS 53795 standard Colour: BLUE | Reinforced | X-Large | N/A | N/A | SANS 53795
RT32-17-018 | Surgical Gown | Gown, surgical, Reinforced, non-woven Body ± 54 g/m², sleeves ± 66 g/m² (Need certification that state the material mass) Long sleeves with cuffs. Reinforced in chest and forearm areas. Water and alcohol repellent. (Need certification that state water and alcohol repellent ) Non-toxic, non-irritant and bacterial migration resistant in wet and dry conditions. Low linting and low flammability qualities. Sterility certification, The content shall be wrapped in a non-woven towel at least 600x600mm which can be opened in a sterile manner. Packed in a peel package which is clearly marked with a list of the contents and expiry date. Individually double peel packed. Hand Towel, Medical X2 to be included in the Gown Pack C Fold, 2 ply, White, Virgin Pulp. Lint free. High absorbency. 40 g/m2 minimum ± 300 X 200 mm folded. Length of the gown must be approximately 3/4 length (between ankle and mid-calf). Colour: BLUE | Reinforced | XX-Large | N/A | N/A | SANS 53795
RT32-17-019 | Surgical Gown | Gown, surgical, Reinforced, non-woven Body ± 54 g/m², sleeves ± 66 g/m² (Need certification that state the material mass) Long sleeves with cuffs. Reinforced in chest and forearm areas. Water and alcohol repellent. (Need certification that state water and alcohol repellent ) Non-toxic, non-irritant and bacterial migration resistant in wet and dry conditions. Low linting and low flammability qualities. Sterility certification, The content shall be wrapped in a non-woven towel at least 600x600mm which can be opened in a sterile manner. Packed in a peel package which is clearly marked with a list of the contents and expiry date. Individually double peel packed. Hand Towel, Medical X2 to be included in the Gown Pack C Fold, 2 ply, White, Virgin Pulp. Lint free. High absorbency. 40 g/m2 minimum ± 300 X 200 mm folded. Length of the gown must be approximately 3/4 length (between ankle and mid-calf). Colour: BLUE | Reinforced | XXX-Large | N/A | N/A | SANS 53795
RT32-17-020 | Basic surgical gown | Basic surgical gown: Non-woven, single pack sterile disposable. Material mass: ± 30 g/m². (Need certification that state the material mass) Long sleeves with cuffs. Reinforced in chest and forearm areas. Water and alcohol repellent. (Need certification that state water and alcohol repellent ) Non-toxic, non-irritant and bacterial migration resistant in wet and dry conditions. Low linting and low flammability qualities. Sterility certification, The content shall be wrapped in a non-woven towel at least 600x600mm which can be opened in a sterile manner. Packed in a peel package which is clearly marked with a list of the contents and expiry date. Individually double peel packed. Hand Towel, Medical X2 to be included in the Gown Pack C Fold, 2 ply, White, Virgin Pulp. Lint free. High absorbency. 40 g/m2 minimum ± 300 X 200 mm folded. Length of the gown must be approximately 3/4 length (between ankle and mid-calf).Latex free Colour: BLUE | N/A | Small | N/A | N/A | SANS 53795
RT32-17-021 | Basic surgical gown | BASIC SURGICAL GOWN: Non-woven, single pack sterile disposable. Material mass: ± 30 g/m². (Need certification that state the material mass) Long sleeves with cuffs. Reinforced in chest and forearm areas. Water and alcohol repellent. (Need certification that state water and alcohol repellent ) Non-toxic, non-irritant and bacterial migration resistant in wet and dry conditions. Low linting and low flammability qualities. Sterility certification, The content shall be wrapped in a non-woven towel at least 600x600mm which can be opened in a sterile manner. Packed in a peel package which is clearly marked with a list of the contents and expiry date. Individually double peel packed. Hand Towel, Medical X2 to be included in the Gown Pack C Fold, 2 ply, White, Virgin Pulp. Lint free. High absorbency. 40 g/m2 minimum ± 300 X 200 mm folded. Length of the gown must be approximately 3/4 length (between ankle and mid-calf).Latex free Colour: BLUE | N/A | Medium | N/A | N/A | SANS 53795
RT32-17-022 | Basic surgical gown | BASIC SURGICAL GOWN: Non-woven, single pack sterile disposable. Material mass: ± 30 g/m². (Need certification that state the material mass) Long sleeves with cuffs. Reinforced in chest and forearm areas. Water and alcohol repellent. (Need certification that state water and alcohol repellent ) Non-toxic, non-irritant and bacterial migration resistant in wet and dry conditions. Low linting and low flammability qualities. Sterility certification, The content shall be wrapped in a non-woven towel at least 600x600mm which can be opened in a sterile manner. Packed in a peel package which is clearly marked with a list of the contents and expiry date. Individually double peel packed. Hand Towel, Medical X2 to be included in the Gown Pack C Fold, 2 ply, White, Virgin Pulp. Lint free. High absorbency. 40 g/m2 minimum ± 300 X 200 mm folded. Length of the gown must be approximately 3/4 length (between ankle and mid-calf).Latex free To comply with SANS 53795 standard Colour: BLUE | N/A | Large | N/A | N/A | SANS 53795
RT32-17-023 | Basic surgical gown | BASIC SURGICAL GOWN: Non-woven, single pack sterile disposable. Material mass: ± 30 g/m². (Need certification that state the material mass) Long sleeves with cuffs. Reinforced in chest and forearm areas. Water and alcohol repellent. (Need certification that state water and alcohol repellent ) Non-toxic, non-irritant and bacterial migration resistant in wet and dry conditions. Low linting and low flammability qualities. Sterility certification, The content shall be wrapped in a non-woven towel at least 600x600mm which can be opened in a sterile manner. Packed in a peel package which is clearly marked with a list of the contents and expiry date. Individually double peel packed. Hand Towel, Medical X2 to be included in the Gown Pack C Fold, 2 ply, White, Virgin Pulp. Lint free. High absorbency. 40 g/m2 minimum ± 300 X 200 mm folded. Length of the gown must be approximately 3/4 length (between ankle and mid-calf).Latex free To comply with SANS 53795 standard Colour: BLUE | N/A | X-Large | N/A | N/A | SANS 53795
RT32-17-024 | Basic surgical gown | BASIC SURGICAL GOWN: Non-woven, single pack sterile disposable. Material mass: ± 30 g/m². (Need certification that state the material mass) Long sleeves with cuffs. Reinforced in chest and forearm areas. Water and alcohol repellent. (Need certification that state water and alcohol repellent ) Non-toxic, non-irritant and bacterial migration resistant in wet and dry conditions. Low linting and low flammability qualities. Sterility certification, The content shall be wrapped in a non-woven towel at least 600x600mm which can be opened in a sterile manner. Packed in a peel package which is clearly marked with a list of the contents and expiry date. Individually double peel packed. Hand Towel, Medical X2 to be included in the Gown Pack C Fold, 2 ply, White, Virgin Pulp. Lint free. High absorbency. 40 g/m2 minimum ± 300 X 200 mm folded. Length of the gown must be approximately 3/4 length (between ankle and mid-calf).Latex free To comply with SANS 53795 standard Colour: BLUE | N/A | XX-Large | N/A | N/A | SANS 53795
RT32-17-025 | Basic surgical gown | BASIC SURGICAL GOWN: Non-woven, single pack sterile disposable. Material mass: ± 30 g/m². (Need certification that state the material mass) Long sleeves with cuffs. Reinforced in chest and forearm areas. Water and alcohol repellent. (Need certification that state water and alcohol repellent ) Non-toxic, non-irritant and bacterial migration resistant in wet and dry conditions. Low linting and low flammability qualities. Sterility certification, The content shall be wrapped in a non-woven towel at least 600x600mm which can be opened in a sterile manner. Packed in a peel package which is clearly marked with a list of the contents and expiry date. Individually double peel packed. Hand Towel, Medical X2 to be included in the Gown Pack C Fold, 2 ply, White, Virgin Pulp. Lint free. High absorbency. 40 g/m2 minimum ± 300 X 200 mm folded. Length of the gown must be approximately 3/4 length (between ankle and mid-calf).Latex free To comply with SANS 53795 standard Colour: BLUE | N/A | XXX-Large | N/A | N/A | SANS 53795
RT32-17-026 | Plastic Apron, Colour: Red | Apron, plastic, full body, single use, no-noise smooth plastic material Length from neck: not less than 110cm Width: not less than 65cm Thickness: 25 micron Ties length: not less than 50cm Width: not less than 5cm Each Colour: RED | Plastic | N/A | N/A | N/A | N/A
RT32-17-027 | Plastic Apron, Colour: Blue | Apron, plastic, full body, single use, no-noise smooth plastic material Length from neck: not less than 110cm Width: not less than 65cm Thickness: 25 micron Ties length: not less than 50cm Width: not less than 5cm Each Colour: BLUE | Plastic | N/A | N/A | N/A | N/A
RT32-17-028 | Plastic Apron, Colour: Yellow | Apron, plastic, full body, single use, no-noise smooth plastic material Length from neck: not less than 110cm Width: not less than 65cm Thickness: 25 micron Ties length: not less than 50cm Width: not less than 5cm Each Colour: YELLOW | Plastic | N/A | N/A | N/A | N/A
RT32-17-029 | Plastic Apron, Colour: Yellow | Apron, plastic, full body, single use, no-noise smooth plastic material Length from neck: not less than 110cm Width: not less than 65cm Thickness: 25 micron Ties length: not less than 50cm Width: not less than 5cm Each Colour: White | Plastic | N/A | N/A | N/A | N/A
RT32-17-030 | Suit | Suit, scrub, coloured, non-sterile, type 5, soft, Opaque Polypropylene, non-woven, durable, breathable fabric to comply with SANS 53795 Resistant to liquid penetration. Top and Trousers Set: Top: V-neck or Round-neck Short sleeves Overlocked hems Trousers: Elasticated waist Overlocked hems | Scrub | Small | N/A | N/A | SANS 53795
RT32-17-31 | Suit | Suit, scrub, coloured, non-sterile, type 5, soft, Opaque Polypropylene, non-woven, durable, breathable fabric to comply with SANS 53795 Compliance certificate to be submitted with bid Resistant to liquid penetration. Top and Trousers Set: Top: V-neck or Round-neck Short sleeves Overlocked hems Trousers: Elasticated waist Overlocked hems | Scrub | Medium | N/A | N/A | SANS 53795
RT32-17-32 | Suit | Suit, scrub, coloured, non-sterile, type 5, soft, Opaque Polypropylene, non-woven, durable, breathable fabric to comply with SANS 53795 Compliance certificate to be submitted with bid Resistant to liquid penetration. Top and Trousers Set: Top: V-neck or Round-neck Short sleeves Overlocked hems Trousers: Elasticated waist Overlocked hems | Scrub | Large | N/A | N/A | SANS 53795
RT32-17-033 | Suit | Suit, scrub, coloured, non-sterile, type 5, soft, Opaque Polypropylene, non-woven, durable, breathable fabric to comply with SANS 53795 Compliance certificate to be submitted with bid Resistant to liquid penetration. Top and Trousers Set: Top: V-neck or Round-neck Short sleeves Overlocked hems Trousers: Elasticated waist Overlocked hems | Scrub | X-Large | N/A | N/A | SANS 53795
RT32-17-034 | Suit | Suit, scrub, coloured, non-sterile, type 5, soft, Opaque Polypropylene, non-woven, durable, breathable fabric to comply with SANS 53795 Compliance certificate to be submitted with bid Resistant to liquid penetration. Top and Trousers Set: Top: V-neck or Round-neck Short sleeves Overlocked hems Trousers: Elasticated waist Overlocked hems | Scrub | XX-Large | N/A | N/A | SANS 53795
CATEGORY 18: BODY BAGS AND SHROUDS
RT32-18-001 | Disposable body bags without a windor and without handles | Material Requirements: Plastic: Low density polyethylene, virgin material Colour: White or Green Thickness of bag: Not lees than 150 micron Body bag: The body bag must be double folded Opening: It must be able to open on one of the long sides of the bag must have an overlap on the openning side of not less than 10cm in width. Clear face window must be inserted on top in the middle of the body bag for easy identification of deceased, by heat welding Closure: The body bag must have double sided tape of not less than 36mm in width fitted in length of the- bag on opening side, for closure of bag. The bag must be leak proof at least 3cm / 5cm from the side, by heat welding Packaging: Each | Without a windor and without handles | Infant - 500mmx500 | Number: 222001012 Description: Body bag 1’s | N/A | N/A
RT32-18-002 | Disposable body bags without a windor and without handles |  |  | Child - 700mmx1m
RT32-18-003 | Disposable body bags without a windor and without handles |  |  | Small - 1.3mx1m
RT32-18-004 | Disposable body bags without a windor and without handles |  |  | Medium - 1.7mx1m
RT32-18-005 | Disposable body bags without a windor and without handles |  |  | Large - 2mx1m
RT32-18-006 | Disposable body bags without a windor and without handles |  |  | XL - 2.3x1m
RT32-18-007 | Disposable body bags without a windor and without handles |  |  | XXL - 2.6x1m
RT32-18-008 | Disposable body bags without a windor and without handles |  |  | XXXL - 2.9x1m
RT32-18-009 | Disposable plastic body bag with window, and without handles | Material Requirements: Plastic: Low density polyethylene, virgin material Clear face window sizes: Infant: 100mm x 100 Child: 150mm x 150mm Small : 20mm x 20mm Medium : 25mm x 25mm Large : 25mm x 25mm X-Large : 30mm x 30mm XX-Large : 30mm x 30mm XXX-Large : 30mm x 30mm Colour: White or Green Thickness of bag: Not lees than 150 micron Body bag: The body bag must be double folded Opening: It must be able to open on one of the long sides of the bag must have an overlap on the openning side of not less than 10cm in width. Clear face window must be inserted on top in the middle of the body bag for easy identification of deceased, by heat welding Closure: The body bag must have double sided tape of not less than 36mm in width fitted in length of the- bag on opening side, for closure of bag. The bag must be leak proof at least 3cm / 5cm from the side, by heat welding Packaging: Each | With window, and without handles | Infant - 500mmx500 | Number: 222001012 Description: Body bag 1’s | N/A | N/A
RT32-18-010 | Disposable plastic body bag with window, and without handles |  | With window, and without handles | Child - 700mmx1m
RT32-18-011 | Disposable plastic body bag with window, and without handles |  | With window, and without handles | Small - 1.3mx1m
RT32-18-012 | Disposable plastic body bag with window, and without handles |  | With window, and without handles | Medium - 1.7mx1m
RT32-18-013 | Disposable plastic body bag with window, and without handles |  | With window, and without handles | Large - 2mx1m
RT32-18-014 | Disposable plastic body bag with window, and without handles |  | With window, and without handles | XL - 2.3x1m
RT32-18-015 | Disposable plastic body bag with window, and without handles |  | With window, and without handles | XXL - 2.6x1m
RT32-18-016 | Disposable plastic body bag with window, and without handles |  | With window, and without handles | XXXL - 2.9x1m
RT32-18-017 | Shrouds with tieback | Shrouds with tieback, disposable Material: PVC light weight, non-transparent, waterproof |  | Neonates | N/A | N/A | N/A
RT32-18-018 | Shrouds with tieback | Shrouds with tieback, disposable Material: PVC light weight, non-transparent, waterproof |  | Babies
RT32-18-019 | Shrouds with tieback | Shrouds with tieback, disposable Material: PVC light weight, non-transparent, waterproof |  | Children
RT32-18-020 | Shrouds with tieback | Shrouds with tieback, disposable Material: PVC light weight, non-transparent, waterproof |  | Adults, Small
RT32-18-021 | Shrouds with tieback | Shrouds with tieback, disposable Material: PVC light weight, non-transparent, waterproof |  | Adults, Medium
RT32-18-022 | Shrouds with tieback | Shrouds with tieback, disposable Material: PVC light weight, non-transparent, waterproof |  | Adults, Large
RT32-18-023 | Shrouds with tieback | Shrouds with tieback, disposable Material: PVC light weight, non-transparent, waterproof |  | Adults, X-Large
RT32-18-024 | Shrouds with tieback | Shrouds with tieback, disposable Material: PVC light weight, non-transparent, waterproof |  | Adults, XX-Large
RT32-18-025 | Shrouds with tieback | Shrouds with tieback, disposable Material: PVC light weight, non-transparent, waterproof |  | Adults, XXX-Large
CATEGORY 19: GOGGLES AND FACE SHIELD
RT32-19-001 | Goggles | Goggle frames: Manufactured from Polypropylene or similar material that is pliable enough to seal against the facial contours and maintain rigidity to hold the lens securely in position, even under pressure. The grade of material used in the frame must not cause irritation or allergic reaction when in contact with the skin of the user and allows the penetration of light and vision through the frame. Direct Venting refers to openings/gaps on the frame to allow airflow to prevent lenses from fogging. Indirect Venting is achieved by designing the goggle frames to limit “direct airflow” (each manufacturer has its own method of achieving this). Indirect venting is preferred to prevent fogging. Goggle lenses: Manufactured from a transparent material with fog and scratch resistant treatment that allows one to see without allowing any droplets or liquids to pass through and/or enter. Must be clear with minimal tinting and provide a clear field of vision, be optically correct with no blemishes or inconsistencies in the lens or surfaces. The lenses must be securely mounted in the goggle frames and not easily dislodged by bumps or impact and to be manufactured from impact resistant clear grade polycarbonate or other compliant material. Securing headband: To be a minimum of 15mm width braided elastic to provide secure positioning of the goggle frame.The head band has to be adjustable for the comfort of the wearer. Made from material that is sustainable and non-toxic and does not cause irritation to the user. | Medical | One Size | Number: 222000996 Description: Goggles 1’s | N/A | SANS 1404, Eye-protectors for industrial and non-industrial use SANS 50166 SAHPRA Class A Non- sterile – exclusion from SAHPRA licence
RT32-19-002 | Face Shield , Visor (window/lenses) | Visor (window/lense): • Material: o Manufactured from polycarbonate, propionate, acetate, polyvinyl chloride, and polyethylene terephthalate glycol (PETG). Acetate provides the best clarity. Polycarbonate and propionate offer better, although still somewhat imperfect, the optical quality that aids in reducing eye strain associated with face shield wear. PETG is the most economical and manufactured locally, and therefore mostly used. Should be made with anti-fogging material properties. . • Dimensions: o Width - Sufficient width to reach at least the point of the ear, minimum 28cm. o Length - Full facepiece length that extends to the bottom of the chin (minimum length of 16cm measured from face shield band) OR Face/neck length that also covers the anterior neck area (minimum length of 20cm measured from face shield band). • Thickness of lense: 200 micron • Single use Frame: Manufactured from lightweight, smooth material like polypropylene plastic. Fully or partially encircle the circumference of the head. The headband should be ergonomic in design/material for comfort and durability of use. Suspension system: Suspension system to have either fully or partially circumferential attachment features. Fully circumferential suspension systems include plastic headbands that are adjustable for comfort by a, Velcro;adjustable systems employ elastic straps. A top foam band or plastic injection-molded headpiece that is either fixed or adjusted for depth to ensure a snug fit against the forehead. There must be sufficient space between the visor and the face to wear spectacles and to breathe and talk comfortably. Packaging : Each | Medical | Full facepiece length that extends to the bottom of the chin, Face/neck length that also covers the anterior neck area | N/A | N/A | SANS 1404, Eye-protectors for industrial and non-industrial use amended test methods. SANS 50166 SAHPRA Class A Non- sterile – exclusion from SAHPRA licence
CATEGORY 20: NET PANTS AND DISPOSABLE UNDERWEARS (BRIEFS)
RT32-20-001 | Pants, net | Pants, net; elasticated. pants for safe, hygenic securing of incontinence pads of all types, lights, soft and permeable to air, with perma-elastic threads for stability. to adapt well to body contours without interwoven colour strips for size identification. washable extra long leg section and broad waist bands |  | Small | N/A | N/A | N/A
RT32-20-002 | Pants, net | PANTS, NET; elasticated. pants for safe, hygenic securing of incontinence pads of all types, lights, soft and permeable to air, with perma-elastic threads for stability. to adapt well to body contours without interwoven colour strips for size identification. washable extra long leg section and broad waist bands |  | Medium | N/A | N/A | N/A
RT32-20-003 | Pants, net | PANTS, NET; elasticated. pants for safe, hygenic securing of incontinence pads of all types, lights, soft and permeable to air, with perma-elastic threads for stability. to adapt well to body contours without interwoven colour strips for size identification. washable extra long leg section and broad waist bands |  | Large | N/A | N/A | N/A
RT32-20-004 | Pants, net | PANTS, NET; elasticated. pants for safe, hygenic securing of incontinence pads of all types, lights, soft and permeable to air, with perma-elastic threads for stability. to adapt well to body contours without interwoven colour strips for size identification. washable extra long leg section and broad waist bands |  | X-Large | N/A | N/A | N/A
RT32-20-005 | Pants, net | PANTS, NET; elasticated. pants for safe, hygenic securing of incontinence pads of all types, lights, soft and permeable to air, with perma-elastic threads for stability. to adapt well to body contours without interwoven colour strips for size identification. washable extra long leg section and broad waist bands |  | XX-Large | N/A | N/A | N/A
RT32-20-006 | Pants, net, Size XXX-Large | PANTS, NET; elasticated. pants for safe, hygenic securing of incontinence pads of all types, lights, soft and permeable to air, with perma-elastic threads for stability. to adapt well to body contours without interwoven colour strips for size identification. washable extra long leg section and broad waist bands |  | XXX-Large | N/A | N/A | N/A
RT32-20-007 | Underwear, disposable (Briefs) | Underwear, disposable (Briefs) Manufactured from non-transparent, non-woven, durable fabric with elasticised yarn (or similar material) over-locked around leg and waist opening. Each Size: X-small Waist circumference: 50-80 cm Leg circumference: 24-40cm Waist-crotch: not less than 24 cm |  | X-Small | N/A | N/A | N/A
RT32-20-008 | Underwear, disposable (Briefs) | Underwear, disposable (Briefs) Manufactured from non-transparent, non-woven, durable fabric with elasticised yarn (or similar material) over-locked around leg and waist opening. Each Size: Small Waist circumference: 80-90 cm Leg circumference: 28-50cm Waist-crotch: not less than 28 cm |  | Small | N/A | N/A | N/A
RT32-20-009 | Underwear, disposable (Briefs) | Underwear, disposable (Briefs) Manufactured from non-transparent, non-woven, durable fabric with elasticised yarn (or similar material) over-locked around leg and waist opening. Each Size: MEDIUM Waist circumference: 90-100 cm Leg circumference: 32-64cm Waist-crotch: not less than 32 cm |  | Medium | N/A | N/A | N/A
RT32-20-010 | Underwear, disposable (Briefs) | Underwear, disposable (Briefs) Manufactured from non-transparent, non-woven, durable fabric with elasticised yarn (or similar material) over-locked around leg and waist opening. Each Size: LARGE Waist circumference: 100-120 cm Leg circumference: 36-72 cm Waist-crotch: not less than 35 cm |  | Large | N/A | N/A | N/A
RT32-20-011 | Underwear, disposable (Briefs) | Underwear, disposable (Briefs) Manufactured from non-transparent, non-woven, durable fabric with elasticised yarn (or similar material) over-locked around leg and waist opening. Each Size: X-LARGE Waist circumference: 120-140 cm Leg circumference: 36-76cm Waist-crotch: not less than 40 cm |  | X-Large | N/A | N/A | N/A
RT32-20-012 | Underwear, disposable (Briefs) | Underwear, disposable (Briefs) Manufactured from non-transparent, non-woven, durable fabric with elasticised yarn (or similar material) over-locked around leg and waist opening. Each Size: XX-LARGE Waist circumference: 130-150 cm Leg circumference: 40-80 cm Waist-crotch: not less than 44 cm |  | XX-Large | N/A | N/A | N/A
RT32-20-013 | Underwear, disposable (Briefs) | Underwear, disposable (Briefs) Manufactured from non-transparent, non-woven, durable fabric with elasticised yarn (or similar material) over-locked around leg and waist opening. Each Size: XXX-LARGE Waist circumference: 150-170 cm Leg circumference: 60-100 cm Waist-crotch: not less than 48 cm |  | XXX-Large | N/A | N/A | N/A
CATEGORY 21: DISPOSABLE COVERALLS
RT32-21-001 | Coveralls: disposable | Coveralls: Disposable Single pack, non-sterile, non-woven, multi-layered construction with a minimum mass of 50g / m2 Material Requirements Coveralls to be manufactured from protective fabrics with fluid-resistant properties.water repellent or impermeable. Possible materials include: a protective 3- layer SMS (Spunbond/Meltblown/Spunbond) polypropylene non-woven product, other alternatives are laminated, coated or calendered materials that comply with the performance criteria as indicated by the stipulated standards.Design RequirementsDesign criteria to be consistent for coveralls . full bodysuit , with hood. elsticated cuffs at wrists . legs and waist.With a zipper front and self adhesiive flap and chin closure .All seams must be tape slide Colour: White and must not be transparent.Dimension RequirementsSize to be clearly identified on the garment and packaging. Sizes: Small International Standards : ISO 13982-1: 2004, EN 1149-5:2008 |  | Small | N/A | ISO 13982-1: 2004, EN 1149-5:2008 | N/A
RT32-21-002 | Coveralls: disposable | Coveralls: Disposable Material Requirements Coveralls to be manufactured from protective fabrics with fluid-resistant properties.water repellent or impermeable. Possible materials include: a protective 3- layer, Spunbond/Meltblown/Spunbond (SMS) polypropylene non-woven product, other alternatives are laminated, coated or calendered materials that comply with the performance criteria as indicated by the stipulated standards. Design Requirements Design criteria to be consistent for coveralls . full bodysuit , with hood. elsticated cuffs at wrists . legs and waist.With a zipper front and self adhesiive flap and chin closure.All seams must be tape slide. Colour: White and must not be transparent. Dimension Requirements Size to be clearly identified on the garment and packaging. Sizes:Medium International Standards: CE , ISO 13982-1: 2004, EN 1149-5:2008 |  | Medium | N/A | CE , ISO 13982-1: 2004, EN 1149-5:2008 | N/A
RT32-21-003 | Coveralls: disposable, size: Large | Coveralls: Disposable Material Requirements Coveralls to be manufactured from protective fabrics with fluid-resistant properties.water repellent or impermeable. Possible materials include: a protective 3- layer SMS (Spunbond/Meltblown/Spunbond) polypropylene non-woven product, other alternatives are laminated, coated or calendered materials that comply with the performance criteria as indicated by the stipulated standards. Design Requirements Design criteria to be consistent for coveralls . full bodysuit , with hood. elsticated cuffs at wrists . legs and waist.With a zipper front and self adhesiive flap and chin closure. All seams must be tape slide. Colour: White and must not be transparent. Dimension Requirements Size to be clearly identified on the garment and packaging. Sizes:Large International Standards : CE , ISO 13982-1: 2004, EN 1149-5:2008 |  | Large |  | CE , ISO 13982-1: 2004, EN 1149-5:2008
RT32-21-004 | Coveralls: disposable | Coveralls: Disposable Material RequirementsCoveralls to be manufactured from protective fabrics with fluid-resistant properties.water repellent or impermeable. Possible materials include: a protective 3- layer SMS (Spunbond/Meltblown/Spunbond) polypropylene non-woven product, other alternatives are laminated, coated or calendered materials that comply with the performance criteria as indicated by the stipulated standards.Design RequirementsDesign criteria to be consistent for coveralls . full bodysuit , with hood. elsticated cuffs at wrists . legs and waist.With a zipper front and self adhesiive flap and chin closure.All seams must be tape slide. Colour: White and must not be transparent.Dimension RequirementsSize to be clearly identified on the garment and packaging. Sizes: X-Large International Standards : CE , ISO 13982-1: 2004, EN 1149-5:2008 |  | X-Large | N/A | CE , ISO 13982-1: 2004, EN 1149-5:2008 | N/A
RT32-21-005 | Coveralls: disposable | Coveralls: Disposable Material Requirements Coveralls to be manufactured from protective fabrics with fluid-resistant properties.water repellent or impermeable. Possible materials include: a protective 3- layer SMS (Spunbond/Meltblown/Spunbond) polypropylene non-woven product, other alternatives are laminated, coated or calendered materials that comply with the performance criteria as indicated by the stipulated standards. Design Requirements Design criteria to be consistent for coveralls . full bodysuit , with hood. elsticated cuffs at wrists . legs and waist.With a zipper front and self adhesiive flap and chin closure. All seams must be tape slide. Colour: White and must not be transparent. Dimension Requirements Size to be clearly identified on the garment and packaging. Sizes: XX-Large International Standards : CE , ISO 13982-1: 2004, EN 1149-5:2008 |  | XX-Large | N/A | CE , ISO 13982-1: 2004, EN 1149-5:2008 | N/A
RT32-21-006 | Coveralls: disposable | Coveralls: Disposable Material Requirements Coveralls to be manufactured from protective fabrics with fluid-resistant properties.water repellent or impermeable. Possible materials include: a protective 3- layer SMS (Spunbond/Meltblown/Spunbond) polypropylene non-woven product, other alternatives are laminated, coated or calendered materials that comply with the performance criteria as indicated by the stipulated standards. Design Requirements Design criteria to be consistent for coveralls . full bodysuit , with hood. elasticated cuffs at wrists . legs and waist.With a zipper front and self adhesiive flap and chin closure. All seams must be tape slide. All seams must be tape slide. Colour: White and must not be transparent. Dimension Requirements Size to be clearly identified on the garment and packaging. Sizes: XXX-Large International Standards : CE , ISO 13982-1: 2004, EN 1149-5:2008 |  | XXX-Large | N/A | CE , ISO 13982-1: 2004, EN 1149-5:2008 | N/A
RT32-21-007 | Coveralls: disposable | Coveralls: Disposable Material RequirementsCoveralls to be manufactured from protective fabrics with fluid-resistant properties.water repellent or impermeable. Possible materials include: a protective 3- layer SMS (Spunbond/Meltblown/Spunbond) polypropylene non-woven product, other alternatives are laminated, coated or calendered materials that comply with the performance criteria as indicated by the stipulated standards.Design RequirementsDesign criteria to be consistent for coveralls . full bodysuit , with hood. elsticated cuffs at wrists . legs and waist.With a zipper front and self adhesiive flap and chin closure. All seams must be tape slide. Colour: White and must not be transparent.Dimension RequirementsSize to be clearly identified on the garment and packaging. Sizes: XXXX-Large International Standards : CE , ISO 13982-1: 2004, EN 1149-5:2008 |  | XXXX-Large | N/A | CE , ISO 13982-1: 2004, EN 1149-5:2008 | N/A
CATEGORY 22: PAPER TOWEL, TYPE MEDICAL DRESSING AND MEDICAL UNDER PADS
RT32-22-001 | Paper towel | Paper towel; type medical dressing, 2 ply, size: 425 mm width, 480 mm length, in accordance with the latest issue of SANS 1887-1, edition 1, for steam sterilizing, the paper shall be free from optical bleaching agent to the extent required in SANS 50868-4, white, 100 perforated sheets per roll, | Medical dressing, 2 ply | 425 mm width, 480 mm length | N/A | N/A | SANS 1887-1 & 6
RT32-22-002 | Paper towel | Paper towel; type medical dressing, 2 ply. size: 444 mm width, 508 mm length. in accordance with the latest issue of SANS 1887-1, edition 1, for steam sterilizing, the paper shall be free from optical bleaching agent to the extent required in SANS 50868-4, white, 100 perforated sheets per roll, | Medical dressing, 2 ply | 444 mm width, 508 mm length |  |  | SANS 1887- part 1 & 6
RT32-22-003 | Paper towel | Paper towel; medical dressing, 2 ply, perforated, size: 254 mm width, 444 mm length. in accordance with the latest issue of sans 1887-1, edition 1, for steam sterilizing, the paper shall be free from optical bleaching agent to the extent required in SANS 50868-4, white, 100 perforated sheets per roll, | Medical dressing, 2 ply, perforated | 254 mm width, 444 mm length |  |  | SANS 1887- part 1 & 6
RT32-22-004 | Linen saver/medical underpads | Linen saver/medical underpads; 6 ply, consisting of 6 absorbent tissue inner layers of paper, a polyethylene sheet and a non-woven outer sheet. size: 510 to 530 width and 640 to 650 length. | 6 ply, consisting of 6 absorbent tissue inner layers of paper | 510 to 530 width and 640 to 650 length
RT32-22-005 | Linen saver/medical underpad | Linen saver/medical underpad; 6 ply, consisting of 6 absorbent tissue inner layers of paper, a polyethylene sheet and a non-woven outer sheet. size: 370 to 390 width and 500 to 520 length. | 6 ply, consisting of 6 absorbent tissue inner layers of paper | 370 to 390 width and 500 to 520 length
RT32-22-006 | Linen saver/medical underpad | Linen saver/medical underpad; 4 ply. consisting of 4 absorbent tissue inner layers of paper, a polyethylene sheet and a non-woven outer sheet. size: 510 to 530 width and 640 to 650 length. | 4 ply. consisting of 4 absorbent tissue inner layers of paper | 510 to 530 width and 640 to 650 length
RT32-22-007 | Linen saver/medical underpad | Linen saver/medical underpad; 6 ply, 6 layers of cellulose, absorbent tissue layers with polyethylene film. size: 340 to 380 width by 600 to 610 length. | 6 ply, 6 layers of cellulose | 340 to 380 width by 600 to 610 length
RT32-22-008 | Disposable draw sheet | Disposable draw sheet, white rectangular absorbent sheet, suitable to use in hospitals as proctection sheets for bed linen, stretcher covers, theatre or examination table covers and / or similar protective use. Draw sheet material shall be laminated to crepe tissue paper and plastic film and shall, when tested in accordance with methods prescribed in SANS 1887 part 1 and part 4. Size 1m x 1.56m. | Rectangular absorbent sheet | 1m x 1.56m
CATEGORY 23: BIBS
RT32-23-001 | Bib | Bib for protecting clothes construct of water proof material, unsterile, disposable rustle proof for Paediatrics use, Ties for fastening bib must not be shorter than 15cm Packaging each |  | Paediatrics | N/A | N/A | N/A
RT32-23-002 | Bib | BIB for protecting clothes, unsterile, disposable for Adults Ties for fastening bib must not be shorter than 30cm Packaging each |  | Adults | N/A | N/A | N/A
CATEGORY 24: ELECTROCARDIOGRAMS
RT32-24-001 | Electrode, electrocardiograph, General monitoring | Electrode, electrocardiograph, General monitoring Infant, Adhesive (Must be able to adhere to a patient’s skin for at least 36 hours), Disposable, Pre-gelled, Press stud type, Irregular soft cloth tape, Short and long term use, Packed of 25 electrodes |  | Infact - Pack of 25 electrodes | N/A | N/A | N/A
RT32-24-002 | Electrode, electrocardiograph, General monitoring | Electrode, electrocardiograph, General monitoring Infant, Adhesive (Must be able to adhere to a patient’s skin for at least 36 hours), Disposable, Pre-gelled, Press stud type, Irregular soft cloth tape, Short and long term use, Packed of 50 electrodes |  | Infact - Pack of 50 electrodes | N/A | N/A | N/A
RT32-24-003 | Electrode, electrocardiograph, general monitoring | Electrode, electrocardiograph, general monitoring, Paediatric, Adhesive (Must be able to adhere to a patient’s skin for at least 36 hours), Disposable, Pre-gelled, Foam substrate, Press stud type, Short and long term use, Pack size: Packet with 25 electrodes |  | Paediatric, Pack of 25 electrodes | N/A | N/A | N/A
RT32-24-004 | Electrode, electrocardiograph, general monitoring | ELECTRODE, ELECTROCARDIOGRAPH, General monitoring, Paediatric, Adhesive (Must be able to adhere to a patient’s skin for at least 36 hours), Disposable, Pre-gelled, Foam substrate, Press stud type, Short and long term use, Pack size: Packet with 50 electrodes |  | Paediatric, Pack of 50 electrodes | N/A | N/A | N/A
RT32-24-005 | Electrode, electrocardiograph, general monitoring | Electrode, electrocardiograph, general monitoring, adult, Adhesive (Must be able to adhere to a patient’s skin for at least 36 hours), Disposable, Pre-gelled, Foam substrate, Press stud type, Short and long term use, Pack size: Packet with 25 electrodes |  | Adults, Pack of 25 electrodes | N/A | N/A | N/A
RT32-24-006 | Electrode, electrocardiograph, general monitoring | ELECTRODE, ELECTROCARDIOGRAPH, General monitoring, Adult , Adhesive (Must be able to adhere to a patient’s skin for at least 36 hours), Disposable, Pre-gelled, Foam substrate, Press stud type, Short and long term use, Pack size: Packet with 50 electrodes |  | Adults, Pack of 50 electrodes | N/A | N/A | N/A
CATEGORY 25: CLAMP, UMBILICAL CORD
RT32-25-001 | Clamp,umbilical cord | Clamp,umbilical cord, Clamp must be constructed for easy application in wet, slippery conditions. Serrations and stoppers must prevent cord from slipping. Clamp must remain securely locked after application and must not cut through cord. Single use sterile pack Size: Large. Nominal measurements: Overal length: 55mm Jaw length: 40mm Width of clamping surface: 6mm Colour: Different colours acceptable To comply with the latest issue of CKS 312. Each |  | Large | N/A | N/A | CKS 312
Policy for the Regulation of Quality Respiratory Protective Equipment (RPE) Supply in Healthcare Please find this policy document shared along with these minimum specifications. The policy document on QA for RPE is critical for the procurement of safe, high quality respirators. The respirator specification provided in this minimum specifictions document should be read in conjunction with the Policy for the Regulation of Quality Respiratory Protective Equipment (RPE) Supply in Healthcare Note from the Department of Trade Industry and Competition (DTIC) “I call on every South African to remain strong and steadfast in these most difficult times. Stay safe and protect South Africa.” President Matamela Cyril Ramaphosa, 01 August 2020 This specifications document has been adopted and adapted for the Republic of South Africa’s demographics, taking into cognisance best practices and latest medical research to mitigate health care risks. In the procurement of medical products for the South African population and in response to Covid-19, the South African Specifications referenced in this document must be adhered to. The international standards made reference to in this document serve to assist and guide in referencing such foreign specifications. In addition to adherence to latest medical research and best practices, it must be requested that in the procurement of medical products the Republic of South Africa’s Re-imagined Industrial Strategy , Industrial Policy and Procurement Legislation be given attention. The central tenant of our Republic’s re-imagined industrial strategy and policy is to address our triple challenges of poverty, inequality and unemployment. Within this context it is imperative that the procurement of medical products in response to Covid-19 (and in-general) should seek to support local manufacturers – the procurement (demand) of local manufactured goods/products (supply) has a net-positive impact on our country’s socio-economic realties. Within this context our Republic has enacted Legislation and Regulations that impose strict duties on public sector procurement: The Preferential Procurement Policy Framework Act 05 of 2000 (PPPFA) and the PPPFA Preferential Procurement Regulations , which came into effect on the 07 December 2011 empower the Department of Trade, Industry, and Competition (the dtic) to designate industries, sectors and sub-sectors for local production at a specified level of local content. It must be noted that regulation 8(2) of the Preferential Procurement Regulations requires that organs of state to include local content in the invitation of bids and tenders. Regulation 8(5) states that if bids fail to meet the required local content such an application should not be acceptable. Further, local content rules require organs of state to populate Standard Bidding Documents 6. 2 (SBD 6.2) with a list of designated products and to circulate local content annexures. Local content is applicable on goods designated for local production regardless on the procurement methods used by an organ of state. It is also applicable on turnkey projects where public money is spent and the project owned by the government. The Designated Local Content information can be found here: http://www.thedtic.gov.za/sectors-and-services-2/industrial-development/industrial-procurement/ The dtic remains the custodian of the Designated Local Content legislation. The dtic may and can request all public entities to supply procurement information, especially information pertaining to adherence to local content legislation. Moreover, attention must also be given to the Judgement of 20 July 2020 in the High Court of South Africa, Gauteng Division Pretoria by Judge P. D. Phahlane in regard to Designated Local Content and procurement by organs of state. Judge Phahlane noted in the judgement that the dtic has discretionary power and is enabled by the PPPFA in enforcing policies of our government, particularly when ensuring procurement from local manufacturers using local materials within the Republic of South Africa.

## RT32 Pricing Schedule
RT32 PRICING SCHEDULE
BID DESCRIPTION: SUPPLY AND DELIVERY OF SURGICAL SUNDRIES TO THE STATE FOR A PERIOD OF 60MONTHS TO COMMENCE 1 MARCH 2023 TO 28 FERUARY 2028
If the bidder has more than one offer, additional rows may be added below offer one
ITEM NUMBER | ITEM DESCRIPTION | OFFER OPTION | UNIT OF MEASURE | BIDDER'S NAME | OFFERED PRICE IN ZAR (INCL. VAT) | BRAND NAME | PRODUCT PACK SIZE | NAME OF MANUFACTURER | COUNTRY OF ORIGIN | LEAD TIME DAY/S | FOREIGN CURRENCY | MINIMUM ORDER QUANTITY
CATEGORY 1: MASKS
RT32-01-001 | Surgical mask Type II or higher(Earloop) | Offer 1 | Each
RT32-01-002 | Surgical mask Type II or higher(Four tie back) | Offer 1 | Each
RT32-01-003 | Mask surgical face, Type I (Earloop) | Offer 1 | Each
RT32-01-004 | Mask surgical face, Type I (Four tie back) | Offer 1 | Each
RT32-01-005 | Respirators (FFP2) mask, (Duckbill shape) Size: Small | Offer 1 | Each
RT32-01-006 | Respirators (FFP2) mask, (Duckbill shape) Size: Medium | Offer 1 | Each
RT32-01-007 | Respirators (FFP2) mask, (Duckbill shape) Size: Large | Offer 1 | Each
RT32-01-008 | Respirators (FFP2) mask,(Flat Shape) Size: Small | Offer 1 | Each
RT32-01-009 | Respirators (FFP2) mask, (Flat Shape) Size: Medium | Offer 1 | Each
RT32-01-010 | Respirators (FFP2) mask, (Flat Shape) Size: Large | Offer 1 | Each
RT32-01-011 | Respirators (FFP2) mask, (Anatomal Shape) Size: Small | Offer 1 | Each
RT32-01-012 | Respirators (FFP2) mask, (Anatomical Shape) Size: Medium | Offer 1 | Each
RT32-01-013 | Respirators (FFP2) mask, (Anatomical Shape) Size: Large | Offer 1 | Each
RT32-01-014 | Respirators (N95) mask, (Duckbill shape) Size: Small | Offer 1 | Each
RT32-01-015 | Respirators (N95) mask, (Duckbill shape) Size: Medium | Offer 1 | Each
RT32-01-016 | Respirators (N95) mask, (Duckbill shape) Size: Large | Offer 1 | Each
RT32-01-017 | Respirators (N95) mask, (Flat Shape) Size: Small | Offer 1 | Each
RT32-01-018 | Respirators (N95) mask, (Flat Shape) Size: Medium | Offer 1 | Each
RT32-01-019 | Respirators (N95) mask, (Flat Shape) Size: Large | Offer 1 | Each
RT32-01-020 | Respirators (N95) mask,(Anatomal Shape) Size: Small | Offer 1 | Each
RT32-01-021 | Respirators (N95) mask, (Anatomical Shape) Size: Medium | Offer 1 | Each
RT32-01-022 | Respirators (N95) mask, (Anatomical Shape) Size: Large | Offer 1 | Each
RT32-01-023 | Respirators (KN95) mask, (Duckbill shape) Size: Small | Offer 1 | Each
RT32-01-024 | Respirators (KN95) mask, (Duckbill shape) Size: Medium | Offer 1 | Each
RT32-01-025 | Respirators (KN95) mask, (Duckbill shape) Size: Large | Offer 1 | Each
RT32-01-026 | Respirators (KN95) mask, (Flat Shape) Size: Small | Offer 1 | Each
RT32-01-027 | Respirators (KN95) mask, (Flat Shape) Size: Medium | Offer 1 | Each
RT32-01-028 | Respirators (KN95) mask, (Flat Shape) Size: Large | Offer 1 | Each
RT32-01-029 | Respirators (KN95) mask,(Anatomal Shape) Size: Small | Offer 1 | Each
RT32-01-030 | Respirators (KN95) mask, (Anatomical Shape) Size: Medium | Offer 1 | Each
RT32-01-031 | Respirators (KN95) mask, (Anatomical Shape) Size: Large | Offer 1 | Each
CATEGORY 2: PERFORATOR
RT32-02-001 | Perforator (Straight) | Offer 1 | Each
RT32-02-002 | Perforator (Curved) | Offer 1 | Each
CATEGORY 3: THERMOMETERS
RT32-03-001 | Thermometer non-mercury | Offer 1 | Each
RT32-03-002 | Thermometer, non-mercury, semi-disposable | Offer 1 | Each
RT32-03-003 | Thermometer, suitable for wall mounting, digital display | Offer 1 | Each
RT32-03-004 | Thermometer refrigerator | Offer 1 | Each
CATEGORY 4: APPLICATORS
RT32-04-001 | Applicator(orange stick) | Offer 1 | Each
RT32-04-002 | Applicator(haemolysis) | Offer 1 | Each
CATEGORY 5: COVERS
RT32-05-001 | Cover, bedpan | Offer 1 | Each
RT32-05-002 | Cover, urinal | Offer 1 | Each
CATEGORY 6: OVERSHOES AND CAPS
RT32-06-001 | Overshoes Fabric Suitable for all shoe sizes up to size 7 (South African) | Offer 1 | Each
RT32-06-002 | Overshoes Fabric Suitable for all shoe sizes 8-15 (South African) | Offer 1 | Each
RT32-06-003 | Overshoe, plastic, single use Suitable for all shoe sizes up to size 7 (South African) | Offer 1 | Each
RT32-06-004 | Overshoe, plastic, single use Suitable for all shoe sizes 8-15 (South African) | Offer 1 | Each
RT32-06-005 | Cap, beret type, (similar to the mop cap) | Offer 1 | Each
RT32-06-006 | Cap, theatre, bonnet type with ties | Offer 1 | Each
RT32-06-007 | Cap, theatre, balaclava type | Offer 1 | Each
CATEGORY 7: SLEEVE PROTECTORS
RT32-07-001 | Sleeve, protector, single use | Offer 1 | Each
RT32-07-002 | Sleeve, surgical, re-useable PVC material, | Offer 1 | Each
CATEGORY 8: BRUSHS
RT32-08-001 | Brush, surgical, nail, autoclavable, continuously re-usable without distortion or change in quality | Offer 1 | Each
RT32-08-002 | Brush, surgical, nail, sterile, single use in theatres with nail cleaner | Offer 1 | Each
CATEGORY 9: IDENTIFICATION BANDS
RT32-09-001 | Band, personal identification of plastic For premature babies and neonates,Colour Pink | Offer 1 | Each
RT32-09-002 | Band, personal identification of plastic For children,Colour Pink | Offer 1 | Each
RT32-09-003 | Band, personal identification of plastic For children,Colour White | Offer 1 | Each
RT32-09-004 | Band, personal identification of plastic For adults,Colour Yellow | Offer 1 | Each
CATEGORY 10: EYE COVERS
RT32-10-001 | Eye cover/posy for phototherapy:Size Small | Offer 1 | Each
RT32-10-002 | Eye cover/posy for phototherapy:Size Medium | Offer 1 | Each
RT32-10-003 | Eye cover/posy for phototherapy:Size Large | Offer 1 | Each
RT32-10-004 | Eye Shield Cartella, universal (to fit both eyes) | Offer 1 | Each
CATEGORY 11: CERVICAL COLLARS
RT32-11-001 | Collar cervical, Size Small | Offer 1 | Each
RT32-11-002 | Collar cervical, Size Medium | Offer 1 | Each
RT32-11-003 | Collar cervical, Size Large | Offer 1 | Each
RT32-11-004 | Collar cervical, Size X-Large | Offer 1 | Each
RT32-11-005 | Collar cervical, Size XX-Large | Offer 1 | Each
CATEGORY 12: VAGINAL SPECULUMS
RT32-12-001 | Speculum, vaginal, Size: Small (75mm x 25mm) | Offer 1 | Each
RT32-12-002 | Speculum, vaginal, Medium (85mm x 30mm) | Offer 1 | Each
RT32-12-003 | Speculum, vaginal, Size: Large (100mm x 35mm) | Offer 1 | Each
RT32-12-004 | Speculum, vaginal, Size: X-Large (110 x 37) | Offer 1 | Each
RT32-12-005 | Speculum, vaginal, Size: Small (75mm x 25mm) | Offer 1 | Each
RT32-12-006 | Speculum, vaginal, Medium (85mm x 30mm) | Offer 1 | Each
RT32-12-007 | Speculum, vaginal, Size: Large (100mm x 35mm) | Offer 1 | Each
RT32-12-008 | Speculum, vaginal, Size: X-Large (110 x 37) | Offer 1 | Each
CATEGORY 13: SANITARY PADS
RT32-13-001 | Sanitary pad, post delivery, maternity without loops Size width 90mm x length 240mm | Offer 1 | Each
RT32-13-002 | Sanitary pad, maternity with loops. Autoclavable Must comply with the lastest issue of SANS 1043. pack of 12 | Offer 1 | Each
RT32-13-003 | Sanitary pad,regular with adhesive back strip and wings Size: ±20cm in length x ±7cm in width with a wing span of ±15cm | Offer 1 | Each
CATEGORY 14: INCONTINENCE PADS
RT32-14-001 | Incontinence, pad; slight urinary incontinence, Size Small | Offer 1 | Each
RT32-14-002 | Incontinence, pad; slight urinary incontinence, Size Medium | Offer 1 | Each
RT32-14-003 | Incontinence, pad; slight urinary incontinence, Size Large | Offer 1 | Each
RT32-14-005 | Incontinence pad; for moderate urinary and faecal incontinence, Size Small | Offer 1 | Each
RT32-14-006 | Incontinence pad; for moderate urinary and faecal incontinence, Size Medium | Offer 1 | Each
RT32-14-007 | Incontinence pad; for moderate urinary and faecal incontinence, Size Large | Offer 1 | Each
RT32-14-008 | Incontinence pad; for moderate urinary and faecal incontinence, Size X-Large | Offer 1 | Each
RT32-14-009 | Incontinence pad; anatomically shaped, men,Size One size fit all | Offer 1 | Each
CATEGORY 15: INCONTINENCE DIAPER
RT32-15-001 | Incontinence diaper; daytime, unisex,Size: Small | Offer 1 | Each
RT32-15-002 | Incontinence diaper; daytime, unisex,Size: Medium | Offer 1 | Each
RT32-15-003 | Incontinence diaper; daytime, unisex,Size: Large | Offer 1 | Each
RT32-15-004 | Incontinence diaper; daytime, unisex,Size: X-Large | Offer 1 | Each
RT32-15-005 | Incontinence diaper; daytime, unisex,Size: XX-Large | Offer 1 | Each
RT32-15-006 | Incontinence diaper; overnight, Unisex,Size: Small | Offer 1 | Each
RT32-15-007 | Incontinence diaper; overnight, Unisex,Size: Medium | Offer 1 | Each
RT32-15-008 | Incontinence diaper; overnight, Unisex,Size: Large | Offer 1 | Each
RT32-15-009 | Incontinence diaper; overnight, Unisex,Size: X-Large | Offer 1 | Each
RT32-15-010 | Incontinence diaper; overnight, Unisex,Size: XX-Large | Offer 1 | Each
CATEGORY 16: DISPOSABLE DIAPERS
RT32-16-001 | Disposable diaper: for use with micro premature babies weight approximately 400 to 700 grams, size: Size: Micro premature baby. | Offer 1 | Each
RT32-16-002 | Disposable diaper: for use with premature babies weight approximately 800 to 1500 grams, size: Size: Premature babies | Offer 1 | Each
RT32-16-003 | Disposable diaper: for use with new born babies weight approximately 1.6kg to 2.5kg, size: Size: New born | Offer 1 | Each
RT32-16-004 | Disposable diaper: ffor use with small infant weight approximately 3kg to 6kg, size: Size: Small infant | Offer 1 | Each
RT32-16-005 | Disposable diaper: for use with infant weight approximately 7kg to 12kg, size: Size: Infant | Offer 1 | Each
RT32-16-006 | Disposable diaper: for use with toddler weight approximately 13kg to 20kg, size: Size: Toddler | Offer 1 | Each
RT32-16-007 | Disposable diaper: ffor use with child weight approximately 21kg to 30kg, size: Size: Child | Offer 1 | Each
CATEGORY 17: GOWNS , APRONS AND SCRUB SUIT
RT32-17-001 | Gown: Disposable (isolation), Size: Small | Offer 1 | Roll
RT32-17-002 | Gown: Disposable (isolation), Size: Medium | Offer 1 | Each
RT32-17-003 | Gown: Disposable (isolation), Size: Large | Offer 1 | Each
RT32-17-004 | Gown: Disposable (isolation), Size: X-Large | Offer 1 | Each
RT32-17-005 | Gown: Disposable (isolation), Size: XX-Large | Offer 1 | Each
RT32-17-006 | Gown: Disposable (isolation), Size: XXX-Large | Offer 1 | Each
RT32-17-007 | Gown: Disposable (isolation), Size: XXXX-Large | Offer 1 | Each
RT32-17-008 | Hospital gowns, Size: Child (Age 3 yrs - 6 yrs) | Offer 1 | Each
RT32-17-009 | Hospital gowns, Size: Child (Age 6 yrs - 9 yrs) | Offer 1 | Each
RT32-17-010 | Hospital gowns, Size: Medium | Offer 1 | Each
RT32-17-011 | Hospital gowns, Size: Large | Offer 1 | Each
RT32-17-012 | Hospital gowns, Size: X-Large | Offer 1 | Each
RT32-17-013 | Hospital gowns, Size: XX-Large
RT32-17-014 | Surgical Gown, Reinforced, Size: Small | Offer 1 | Each
RT32-17-015 | Surgical Gown, Reinforced, Size: Medium | Offer 1 | Each
RT32-17-016 | Surgical Gown, Reinforced, Size: Large | Offer 1 | Each
RT32-17-017 | Surgical Gown, Reinforced, Size: X-Large | Offer 1 | Each
RT32-17-018 | Surgical Gown, Reinforced, Size: XX-Large | Offer 1 | Each
RT32-17-019 | Surgical Gown, Reinforced, Size: XXX-Large | Offer 1 | Each
RT32-17-020 | Basic surgical gown, Size: Small | Offer 1 | Each
RT32-17-021 | Basic surgical gown, Size: Medium | Offer 1 | Each
RT32-17-022 | Basic surgical gown, Size: Large | Offer 1 | Each
RT32-17-023 | Basic surgical gown, Size: X-Large | Offer 1 | Each
RT32-17-024 | Basic surgical gown, Size: XX-Large | Offer 1 | Each
RT32-17-025 | Basic surgical gown, Size: XXX-Large | Offer 1 | Each
RT32-17-026 | Plastic Apron, Colour: Red | Offer 1 | Each
RT32-17-027 | Plastic Apron, Colour: Blue | Offer 1 | Each
RT32-17-028 | Plastic Apron, Colour: Yellow | Offer 1 | Each
RT32-17-029 | Plastic Apron, Colour: White | Offer 2 | Each
RT32-17-030 | Suit, scrub, Size: Small | Offer 1 | Each
RT32-17-031 | Suit, scrub, Size: Medium | Offer 1 | Each
RT32-17-032 | Suit, scrub, Size: Large | Offer 1 | Each
RT32-17-033 | Suit, scrub, Size: X-Large | Offer 1 | Each
RT32-17-034 | Suit, scrub, Size: XX-Large | Offer 1 | Each
CATEGORY 18: BODY BAGS AND SHROUDS
RT32-18-001 | Disposable plastic body bags without a windor and without handles, Size: Infant 500mm x 500 | Offer 1 | Each
RT32-18-002 | Disposable plastic body bags without a windor and without handles, Size: Child 700mm x 1m | Offer 1 | Each
RT32-18-003 | Disposable plastic body bags without a windor and without handles, Size: Small - 1.3mx1m | Offer 1 | Each
RT32-18-004 | Disposable plastic body bags without a windor and without handles, Size: Medium - 1.7mx1m | Offer 1 | Each
RT32-18-005 | Disposable plastic body bags without a windor and without handles, Size: Large - 2mx1m | Offer 1 | Each
RT32-18-006 | Disposable plastic body bags without a windor and without handles, Size: X-Large - 2.3x1m | Offer 1 | Each
RT32-18-007 | Disposable plastic body bags without a windor and without handles, Size: XX-Large - 2.6x1m | Offer 1 | Each
RT32-18-008 | Disposable plastic body bags without a windor and without handles, Size: XXX-Large - 2.9x1m | Offer 1 | Each
RT32-18-009 | Disposable plastic body bag with window, and without handles, Size: Infant 100mm x 100 | Offer 1 | Each
RT32-18-010 | Disposable plastic body bag with window, and without handles, Size: Child 150mm x 150mm | Offer 1 | Each
RT32-18-011 | Disposable plastic body bag with window, and without handles, Size: Small - 20mm x 20mm | Offer 1 | Each
RT32-18-012 | Body bags, Size: Medium - 25mm x 25mm | Offer 1 | Each
RT32-18-013 | Body bags, Size: Large - 25mm x 25mm | Offer 1 | Each
RT32-18-014 | Body bags, Size: X-Large - 30mm x 30mm | Offer 1 | Each
RT32-18-015 | Body bags, Size: XX-Large - 30mm x 30mm | Offer 1 | Each
RT32-18-016 | Body bags, Size: XXX-Large - 30mm x 30mm | Offer 1 | Each
RT32-18-017 | Shrouds with tieback, disposable Material: PVC light weight, non-transparent, waterproof Each Size: Neonates | Offer 1 | Each
RT32-18-018 | Shrouds with tieback, disposable Material: PVC light weight, non-transparent, waterproof Each Size: For babies | Offer 1 | Each
RT32-18-019 | Shrouds with tieback, disposable Material: PVC light weight, non-transparent, waterproof Each Size: For children | Offer 1 | Each
RT32-18-020 | Shrouds with tieback, disposable Material: PVC light weight, non-transparent, waterproof Each Size: Adults: Small | Offer 1 | Each
RT32-18-021 | Shrouds with tieback, disposable Material: PVC light weight, non-transparent, waterproof Each Size: Adults: Medium | Offer 1 | Each
RT32-18-022 | Shrouds with tieback, disposable Material: PVC light weight, non-transparent, waterproof Each Size: Adults Large | Offer 1 | Each
RT32-18-023 | Shrouds with tieback, disposable Material: PVC light weight, non-transparent, waterproof Each Size: Adults: X Large | Offer 1 | Each
RT32-18-024 | Shrouds with tieback, disposable Material: PVC light weight, non-transparent, waterproof Each Size: Adults: XX Large | Offer 1 | Each
RT32-18-025 | Shrouds with tieback, disposable Material: PVC light weight, non-transparent, waterproof Each Size: Adults: XXX Large | Offer 1 | Each
CATEGORY 19: GOOGLES AND FACE SHIELD
RT32-19-001 | Goggles,Medical, Size: One size fit all | Offer 1 | Each
RT32-19-002 | Face Shield - Medical, Visor (window/lenses) | Offer 1 | Each
CATEGORY 20: NET PANTS AND DISPOSABLE UNDERWEAR (BRIEFS)
RT32-20-001 | Pants, net, Size Small | Offer 1 | Each
RT32-20-002 | Pants, net, Size Medium | Offer 1 | Each
RT32-20-003 | Pants, net, Size Large | Offer 1 | Each
RT32-20-004 | Pants, net, Size X-Large | Offer 1 | Each
RT32-20-005 | Pants, net, Size XX-Large | Offer 1 | Each
RT32-20-006 | Pants, net, Size XXX-Large | Offer 1 | Each
RT32-20-007 | Underwear, disposable (Briefs)Size X-Small | Offer 1 | Each
RT32-20-008 | Underwear, disposable (Briefs)Size Small | Offer 1 | Each
RT32-20-009 | Underwear, disposable (Briefs)Size Medium | Offer 1 | Each
RT32-20-010 | Underwear, disposable (Briefs)Size Large | Offer 1 | Each
RT32-20-011 | Underwear, disposable (Briefs)Size X-Large | Offer 1 | Each
RT32-20-012 | Underwear, disposable (Briefs)Size XX-Large | Offer 1 | Each
RT32-20-013 | Underwear, disposable (Briefs)Size XXX-Large | Offer 1 | Each
CATEGORY 21: DISPOSABLE COVERALLS
RT32-21-001 | Coveralls: disposable, size: Small | Offer 1 | Each
RT32-21-002 | Coveralls: disposable, size: Medium | Offer 1 | Each
RT32-21-003 | Coveralls: disposable, size: Large | Offer 1 | Each
RT32-21-004 | Coveralls: disposable, size: X- Large | Offer 1 | Each
RT32-21-005 | Coveralls: disposable, size: XX -Large | Offer 1 | Each
RT32-21-006 | Coveralls: disposable, size: XXX - Large | Offer 1 | Each
RT32-21-007 | Coveralls: disposable, size: XXXX - Large | Offer 1 | Each
CATEGORY 22: PAPER TOWEL, TYPE MEDICAL DRESSING AND MEDICAL UNDER PADS
RT32-22-001 | Paper towel; type medical dressing, 2 ply, size: 425 mm width, 480 mm length | Offer 1 | Each
RT32-22-002 | Paper towel; type medical dressing, 2 ply. size: 444 mm width, 508 mm length | Offer 1 | Each
RT32-22-003 | Paper towel; medical dressing, 2 ply, perforated, size: 254 mm width, 444 mm length | Offer 1 | Each
RT32-22-004 | Linen saver/medical underpads; 6 ply, consisting of 6 absorbent tissue inner layers of paper,size: 510 to 530 width and 640 to 650 length. | Offer 1 | Each
RT32-22-005 | Linen saver/medical underpad; 6 ply, consisting of 6 absorbent tissue inner layers of paper, size: 370 to 390 width and 500 to 520 length. | Offer 1 | Each
RT32-22-006 | Linen saver/medical underpad; 4 ply. consisting of 4 absorbent tissue inner layers of paper, size: 510 to 530 width and 640 to 650 length. | Offer 1 | Each
RT32-22-007 | Linen saver/medical underpad; 6 ply, 6 layers of cellulose, . size: 340 to 380 width by 600 to 610 length. | Offer 1 | Each
RT32-22-008 | Disposable draw sheet, Size 1m x 1.56m | Offer 1 | Each
CATEGORY 23: BIB
RT32-23-001 | Bib, paediatrics | Offer 1 | Each
RT32-23-002 | Bib, adults | Offer 1 | Each
CATEGORY 24:ELECTROCARDIOGRAM
RT32-24-001 | Electrode, electrocardiograph, General monitoring,Infant,Packed of 25 electrodes | Offer 1 | Each
RT32-24-002 | Electrode, electrocardiograph, General monitoring,Infant,Packed of 50 electrodes | Offer 1 | Each
RT32-24-003 | Electrode, electrocardiograph, general monitoring, Paediatric, Packed of 25 electrodes | Offer 1 | Each
RT32-24-004 | Electrode, electrocardiograph, general monitoring, Paediatric, Packed of 50 electrodes | Offer 1 | Each
RT32-24-005 | Electrode, electrocardiograph, general monitoring, adult, Packed of 25 electrodes | Offer 1 | Each
RT32-24-006 | Electrode, electrocardiograph, general monitoring, adult, Packed of 50 electrodes | Offer 1 | Each
CATEGORY 25: CLAMP, UMBILICAL CORD
RT32-25-001 | Clamp,umbilical cord | Offer 1 | Each

## RT32 Pricing Schedule
RT32 PRICING SCHEDULE
BID DESCRIPTION: SUPPLY AND DELIVERY OF SURGICAL SUNDRIES TO THE STATE FOR A PERIOD OF 60MONTHS TO COMMENCE 1 MARCH 2023 TO 28 FERUARY 2028
If the bidder has more than one offer, additional rows may be added below offer one
ITEM NUMBER | ITEM DESCRIPTION | OFFER OPTION | UNIT OF MEASURE | BIDDER'S NAME | OFFERED PRICE IN ZAR (EXCL. VAT) | BRAND NAME | PRODUCT PACK SIZE | NAME OF MANUFACTURER | COUNTRY OF ORIGIN | LEAD TIME DAY/S | FOREIGN CURRENCY | MINIMUM ORDER QUANTITY
CATEGORY 13: SANITARY PADS
RT32-13-001 | Sanitary pad, post delivery, maternity without loops Size width 90mm x length 240mm | Offer 1 | Each
RT32-13-002 | Sanitary pad, maternity with loops. Autoclavable Must comply with the lastest issue of SANS 1043. pack of 12 | Offer 1 | Each
RT32-13-003 | Sanitary pad,regular with adhesive back strip and wings Size: ±20cm in length x ±7cm in width with a wing span of ±15cm | Offer 1 | Each