## RT13-1-2024 Items
BID NUMBER: RT13-1-2024 ANNEXURE A TECHNICAL SPECIFICATION
# | ITEM NO. | ITEM CATEGORY | ITEM SUB-CATEGORY | DESCRIPTION | SIZE DESCRIPTION | DID YOU SUBMIT THE ISO CERTIFICATE WHERE REQUIRED? | STATE ANY TECHINICAL SPECIFICATION DEVIATION
1 | RT13-01-006 | Intubation Devices | Nasal Cannula | Nasal Cannula with integrated C02 sampling for use in an Pre mature Neonates Prong diameter: 1.5mm Prong length: 5.5mm Distance between medial surfaces of the prongs: 4mm NASAL CANNULA WITH INTEGRATED C02 SAMPLING Purpose: For low flow nasal oxygen delivery in a spontaneously breathing patient, with simultaneous C02 sampling Clear or colour tinted non-kinking, non-crush PVC tubing at least 2m long Over-the-ear style with under-chin toggle or sleeve. Nasal cannula tips must be soft, anatomically shaped (curved) and tapered at the ends Any portion of the cannula which comes in contact with facial skin must be soft and uniformly contoured without protruding edges or ridges or flaps Additional C02 sampling line with a male luer connector/ attached to the PVC tubing All components must be latex-free clinically clean and individually packed in peel pouch that is easy to open - For SINGLE use only Information brochure or the outer package of the item should clearly indicate latex free | Nasal Cannula with integrated C02 sampling for use in an Pre mature Neonates Prong diameter: 1.5mm Prong length: 5.5mm Distance between medial surfaces of the prongs: 4mm
2 | RT13-01-016 | Intubation Devices | Nasal Cannula | Nasal Cannula with integrated C02 sampling for use in an pre mature neonates - Hypoallergenic Prong diameter: 1.5mm Prong length: 5.5mm Distance between medial surfaces of the prongs: 4mm NASAL CANNULA WITH INTEGRATED C02 SAMPLING - HYPOALLERGENIC Purpose: For low flow nasal oxygen delivery in a spontaneously breathing patient, with simultaneous C02 sampling Clear or colour tinted non-kinking, non-crush PVC tubing at least 2m long Over-the-ear style with under-chin toggle or sleeve. Nasal cannula tips must be soft, anatomically shaped (curved) and tapered at the ends Any portion of the cannula which comes in contact with facial skin must be soft and uniformly contoured without protruding edges or ridges or flaps Additional C02 sampling line with a male luer connector/ attached to the PVC tubing All components must be latex-free, DEHP free and safe for use in hypoallergenic patients Clinically clean and individually packed in peel pouch that is easy to open - For SINGLE use only Information brochure or the outer package of the item should clearly indicate latex free and DEHP free | Nasal Cannula with integrated C02 sampling for use in an pre mature neonates - Hypoallergenic Prong diameter: 1.5mm Prong length: 5.5mm Distance between medial surfaces of the prongs: 4mm
3 | RT13-01-018 | Intubation Devices | Nasal Cannula | NASAL CANNULA WITH INTEGRATED C02 SAMPLING - HYPOALLERGENIC Purpose: For low flow nasal oxygen delivery in a spontaneously breathing patient, with simultaneous C02 sampling Clear or colour tinted non-kinking, non-crush PVC tubing at least 2m long Over-the-ear style with under-chin toggle or sleeve. Nasal cannula tips must be soft, anatomically shaped (curved) and tapered at the ends Any portion of the cannula which comes in contact with facial skin must be soft and uniformly contoured without protruding edges or ridges or flaps Additional C02 sampling line with a male luer connector/ attached to the PVC tubing All components must be latex-free, DEHP free and safe for use in hypoallergenic patients Clinically clean and individually packed in peel pouch that is easy to open - For SINGLE use only Information brochure or the outer package of the item should clearly indicate latex free and DEHP free | Nasal Cannula with integrated C02 sampling for use in Infants - Hypoallergenic Prong diameter: 4mm Prong length: 10mm Distance between medial surfaces of the prongs: 7mm
4 | RT13-02-028 | Oxygen Mask | Oxygen Mask | VENTURI OXYGEN MASK - MRI COMPATIBLE Purpose: Delivery of oxygen and entrained room air to a spontaneously breathing patient. Suitable for the MRI Suite Modular design correctly sized and comfortable anatomically shaped preformed mask shaped to fit over the patient’s nose and mouth with under the chin cover for a better seal Atraumatic to the nose. 2 exhalation holes in the side of the mask, a secure retaining strap. All components to be MRI compatible. Flexible corrugated tube: Length 10 - 20cm with a 22mm connector Must have clear kink resistant, crush proof delivery tubing > 2 m Colour coded Venturi device with the Oxygen concentration in % and required flow rate must be moulded or clearly indicated on the device The components must be securely attached to each other and not separate accidentally in normal use. Manufactured from latex-free medical grade plastic - must comply with BS EN 13544-3:2001 standards Individually packed, easy to open without tubing getting tangled, clinically clean - SINGLE use Information brochure or the outer package of the item should clearly indicate latex free | Venturi Oxygen Mask MRI COMPATIBLE 28% - Adult ,Colour coded venturi: White
5 | RT13-02-029 | Oxygen Mask | Oxygen Mask | VENTURI OXYGEN MASK - MRI COMPATIBLE Purpose: Delivery of oxygen and entrained room air to a spontaneously breathing patient. Suitable for the MRI Suite Modular design correctly sized and comfortable anatomically shaped preformed mask shaped to fit over the patient’s nose and mouth with under the chin cover for a better seal Atraumatic to the nose. 2 exhalation holes in the side of the mask, a secure retaining strap. All components to be MRI compatible. Flexible corrugated tube: Length 10 - 20cm with a 22mm connector Must have clear kink resistant, crush proof delivery tubing > 2 m Colour coded Venturi device with the Oxygen concentration in % and required flow rate must be moulded or clearly indicated on the device The components must be securely attached to each other and not separate accidentally in normal use. Manufactured from latex-free medical grade plastic - must comply with BS EN 13544-3:2001 standards Individually packed, easy to open without tubing getting tangled, clinically clean - SINGLE use Information brochure or the outer package of the item should clearly indicate latex free | Venturi Oxygen Mask MRI COMPATIBLE 40% - Adult, Colour coded venturi: Red
6 | RT13-02-030 | Oxygen Mask | Oxygen Mask | VENTURI OXYGEN MASK - MRI COMPATIBLE Purpose: Delivery of oxygen and entrained room air to a spontaneously breathing patient. Suitable for the MRI Suite Modular design correctly sized and comfortable anatomically shaped preformed mask shaped to fit over the patient’s nose and mouth with under the chin cover for a better seal Atraumatic to the nose. 2 exhalation holes in the side of the mask, a secure retaining strap. All components to be MRI compatible. Flexible corrugated tube: Length 10 - 20cm with a 22mm connector Must have clear kink resistant, crush proof delivery tubing > 2 m Colour coded Venturi device with the Oxygen concentration in % and required flow rate must be moulded or clearly indicated on the device The components must be securely attached to each other and not separate accidentally in normal use. Manufactured from latex-free medical grade plastic - must comply with BS EN 13544-3:2001 standards Individually packed, easy to open without tubing getting tangled, clinically clean - SINGLE use Information brochure or the outer package of the item should clearly indicate latex free | Venturi Oxygen Mask MRI COMPATIBLE 60% - Adult ,Colour coded venturi: Green
7 | RT13-02-031 | Oxygen Mask | VENTURI OXYGEN MASK - MRI COMPATIBLE | VENTURI OXYGEN MASK - MRI COMPATIBLE - 24% - Paediatric, Colour coded venturi: Blue Purpose: Delivery of oxygen and entrained room air to a spontaneously breathing patient. Suitable for the MRI Suite Modular design correctly sized and comfortable anatomically shaped preformed mask shaped to fit over the patient’s nose and mouth with under the chin cover for a better seal Atraumatic to the nose. 2 exhalation holes in the side of the mask, a secure retaining strap. All components to be MRI compatible. Must have clear kink resistant, crush proof delivery tubing > 2 m Colour coded Venturi device with the Oxygen concentration in % and required flow rate must be moulded or clearly indicated on the device The components must be securely attached to each other and not separate accidentally in normal use. Manufactured from latex-free medical grade plastic - must comply with BS EN 13544-3:2001 standards Individually packed, easy to open without tubing getting tangled, clinically clean - SINGLE use Information brochure or the outer package of the item should clearly indicate latex free | Venturi Oxygen Mask MRI COMPATIBLE 24% - Paediatric, Colour coded venturi: Blue
8 | RT13-02-032 | Oxygen Mask | VENTURI OXYGEN MASK - MRI COMPATIBLE | VENTURI OXYGEN MASK - MRI COMPATIBLE - 28% - Paediatric, Colour coded venturi: White Purpose: Delivery of oxygen and entrained room air to a spontaneously breathing patient. Suitable for the MRI Suite Modular design correctly sized and comfortable anatomically shaped preformed mask shaped to fit over the patient’s nose and mouth with under the chin cover for a better seal Atraumatic to the nose. 2 exhalation holes in the side of the mask, a secure retaining strap. All components to be MRI compatible. Must have clear kink resistant, crush proof delivery tubing > 2 m Colour coded Venturi device with the Oxygen concentration in % and required flow rate must be moulded or clearly indicated on the device The components must be securely attached to each other and not separate accidentally in normal use. Manufactured from latex-free medical grade plastic - must comply with BS EN 13544-3:2001 standards Individually packed, easy to open without tubing getting tangled, clinically clean - SINGLE use Information brochure or the outer package of the item should clearly indicate latex free | Venturi Oxygen Mask MRI COMPATIBLE 28% - Paediatric, Colour coded venturi: White
9 | RT13-02-033 | Oxygen Mask | VENTURI OXYGEN MASK - MRI COMPATIBLE | VENTURI OXYGEN MASK - MRI COMPATIBLE - 40% - Paediatric, Colour coded venturi: Red Purpose: Delivery of oxygen and entrained room air to a spontaneously breathing patient. Suitable for the MRI Suite Modular design correctly sized and comfortable anatomically shaped preformed mask shaped to fit over the patient’s nose and mouth with under the chin cover for a better seal Atraumatic to the nose. 2 exhalation holes in the side of the mask, a secure retaining strap. All components to be MRI compatible. Must have clear kink resistant, crush proof delivery tubing > 2 m Colour coded Venturi device with the Oxygen concentration in % and required flow rate must be moulded or clearly indicated on the device The components must be securely attached to each other and not separate accidentally in normal use. Manufactured from latex-free medical grade plastic - must comply with BS EN 13544-3:2001 standards Individually packed, easy to open without tubing getting tangled, clinically clean - SINGLE use Information brochure or the outer package of the item should clearly indicate latex free | Venturi Oxygen Mask MRI COMPATIBLE 40% - Paediatric, Colour coded venturi: Red
10 | RT13-02-034 | Oxygen Mask | VENTURI OXYGEN MASK - MRI COMPATIBLE | VENTURI OXYGEN MASK - MRI COMPATIBLE Purpose: Delivery of oxygen and entrained room air to a spontaneously breathing patient. Suitable for the MRI Suite Modular design correctly sized and comfortable anatomically shaped preformed mask shaped to fit over the patient’s nose and mouth with under the chin cover for a better seal Atraumatic to the nose. 2 exhalation holes in the side of the mask, a secure retaining strap. All components to be MRI compatible. Must have clear kink resistant, crush proof delivery tubing > 2 m Colour coded Venturi device with the Oxygen concentration in % and required flow rate must be moulded or clearly indicated on the device The components must be securely attached to each other and not separate accidentally in normal use. Manufactured from latex-free medical grade plastic - must comply with BS EN 13544-3:2001 standards Individually packed, easy to open without tubing getting tangled, clinically clean - SINGLE use Information brochure or the outer package of the item should clearly indicate latex free | Venturi Oxygen Mask MRI COMPATIBLE 60% - Paediatric ,Colour coded venturi: Green
11 | RT13-02-035 | Oxygen Mask | VENTURI OXYGEN MASK - MRI COMPATIBLE | VENTURI OXYGEN MASK - MRI COMPATIBLE 24% - Adult, Colour coded venturi: Blue Purpose: Delivery of oxygen and entrained room air to a spontaneously breathing patient. Suitable for the MRI Suite Modular design correctly sized and comfortable anatomically shaped preformed mask shaped to fit over the patient’s nose and mouth with under the chin cover for a better seal Atraumatic to the nose. 2 exhalation holes in the side of the mask, a secure retaining strap. All components to be MRI compatible. Must have clear kink resistant, crush proof delivery tubing > 2 m Colour coded Venturi device with the Oxygen concentration in % and required flow rate must be moulded or clearly indicated on the device The components must be securely attached to each other and not separate accidentally in normal use. Manufactured from latex-free medical grade plastic - must comply with BS EN 13544-3:2001 standards Individually packed, easy to open without tubing getting tangled, clinically clean - SINGLE use Information brochure or the outer package of the item should clearly indicate latex free | Venturi Oxygen Mask MRI COMPATIBLE 24% - Adult, Colour coded venturi: Blue
12 | RT13-02-036 | Oxygen Mask | VENTURI OXYGEN MASK - MRI COMPATIBLE | VENTURI OXYGEN MASK - MRI COMPATIBLE 28% - Adult, Colour coded venturi: White Purpose: Delivery of oxygen and entrained room air to a spontaneously breathing patient. Suitable for the MRI Suite Modular design correctly sized and comfortable anatomically shaped preformed mask shaped to fit over the patient’s nose and mouth with under the chin cover for a better seal Atraumatic to the nose. 2 exhalation holes in the side of the mask, a secure retaining strap. All components to be MRI compatible. Must have clear kink resistant, crush proof delivery tubing > 2 m Colour coded Venturi device with the Oxygen concentration in % and required flow rate must be moulded or clearly indicated on the device The components must be securely attached to each other and not separate accidentally in normal use. Manufactured from latex-free medical grade plastic - must comply with BS EN 13544-3:2001 standards Individually packed, easy to open without tubing getting tangled, clinically clean - SINGLE use Information brochure or the outer package of the item should clearly indicate latex free | Venturi Oxygen Mask MRI COMPATIBLE 28% - Adult, Colour coded venturi: White
13 | RT13-02-037 | Oxygen Mask | VENTURI OXYGEN MASK - MRI COMPATIBLE | VENTURI OXYGEN MASK - MRI COMPATIBLE 40% - Adult, Colour coded venturi: Red Purpose: Delivery of oxygen and entrained room air to a spontaneously breathing patient. Suitable for the MRI Suite Modular design correctly sized and comfortable anatomically shaped preformed mask shaped to fit over the patient’s nose and mouth with under the chin cover for a better seal Atraumatic to the nose. 2 exhalation holes in the side of the mask, a secure retaining strap. All components to be MRI compatible. Must have clear kink resistant, crush proof delivery tubing > 2 m Colour coded Venturi device with the Oxygen concentration in % and required flow rate must be moulded or clearly indicated on the device The components must be securely attached to each other and not separate accidentally in normal use. Manufactured from latex-free medical grade plastic - must comply with BS EN 13544-3:2001 standards Individually packed, easy to open without tubing getting tangled, clinically clean - SINGLE use Information brochure or the outer package of the item should clearly indicate latex free | Venturi Oxygen Mask MRI COMPATIBLE 40% - Adult, Colour coded venturi: Red
14 | RT13-02-038 | Oxygen Mask | VENTURI OXYGEN MASK - MRI COMPATIBLE | VENTURI OXYGEN MASK - MRI COMPATIBLE 60% - Adult ,Colour coded venturi: Green Purpose: Delivery of oxygen and entrained room air to a spontaneously breathing patient. Suitable for the MRI Suite Modular design correctly sized and comfortable anatomically shaped preformed mask shaped to fit over the patient’s nose and mouth with under the chin cover for a better seal Atraumatic to the nose. 2 exhalation holes in the side of the mask, a secure retaining strap. All components to be MRI compatible. Must have clear kink resistant, crush proof delivery tubing > 2 m Colour coded Venturi device with the Oxygen concentration in % and required flow rate must be moulded or clearly indicated on the device The components must be securely attached to each other and not separate accidentally in normal use. Manufactured from latex-free medical grade plastic - must comply with BS EN 13544-3:2001 standards Individually packed, easy to open without tubing getting tangled, clinically clean - SINGLE use Information brochure or the outer package of the item should clearly indicate latex free | Venturi Oxygen Mask MRI COMPATIBLE 60% - Adult ,Colour coded venturi: Green
15 | RT13-05-003 | AIRWAY DEVICES | OROPHARYNGEAL AIRWAYS | OROPHARYNGEAL AIRWAYS Purpose: Used to maintain oral airway patency Anatomically shaped PVA or PVC airway with an appropriate contour that has a flanged buccal end with the size clearly moulded onto the plastic. Must have a robust bite block to prevent occlusion, smooth rounded distal end to minimise tissue trauma. The edges or corners that come into contact with patient tissue, must be smooth with a minimal radius of curvature as per ISO 5364:2016 One-piece design or reinforcement components for that do not separate easily or could be an aspiration risk Distance measured from buccal flange to tip of OPA with tolerance as per ISO 5364:2016 Validation of measurements must be provided. Brochures will be acceptable. All components must be latex-free, clinically clean and individually packed in peel pouch that is easy to open - For SINGLE use only Information brochure or the outer package of the item should clearly indicate latex free | Oropharyngeal airway (guedels) 40-45mm Size 00 (Pink/Cream) Minimum inside dimension: 3.0mm
16 | RT13-05-004 | AIRWAY DEVICES | OROPHARYNGEAL AIRWAYS | OROPHARYNGEAL AIRWAYS Purpose: Used to maintain oral airway patency Anatomically shaped PVA or PVC airway with an appropriate contour that has a flanged buccal end with the size clearly moulded onto the plastic. Must have a robust bite block to prevent occlusion, smooth rounded distal end to minimise tissue trauma. The edges or corners that come into contact with patient tissue, must be smooth with a minimal radius of curvature as per ISO 5364:2016 One-piece design or reinforcement components for that do not separate easily or could be an aspiration risk Distance measured from buccal flange to tip of OPA with tolerance as per ISO 5364:2016 Validation of measurements must be provided. Brochures will be acceptable. All components must be latex-free, clinically clean and individually packed in peel pouch that is easy to open - For SINGLE use only Information brochure or the outer package of the item should clearly indicate latex free | Oropharyngeal airway (guedels) 50-55mm Size 0 (Light Blue/Grey) Minimum inside dimension: 3.5mm
17 | RT13-05-005 | AIRWAY DEVICES | OROPHARYNGEAL AIRWAYS | OROPHARYNGEAL AIRWAYS Purpose: Used to maintain oral airway patency Anatomically shaped PVA or PVC airway with an appropriate contour that has a flanged buccal end with the size clearly moulded onto the plastic. Must have a robust bite block to prevent occlusion, smooth rounded distal end to minimise tissue trauma. The edges or corners that come into contact with patient tissue, must be smooth with a minimal radius of curvature as per ISO 5364:2016 One-piece design or reinforcement components for that do not separate easily or could be an aspiration risk Distance measured from buccal flange to tip of OPA with tolerance as per ISO 5364:2016 Validation of measurements must be provided. Brochures will be acceptable. All components must be latex-free, clinically clean and individually packed in peel pouch that is easy to open - For SINGLE use only Information brochure or the outer package of the item should clearly indicate latex free | Oropharyngeal airway (guedels) 60 - 65mm 6.0 Size 1 (Black/Brown) Minimum inside dimension: 4.0mm
18 | RT13-05-006 | AIRWAY DEVICES | OROPHARYNGEAL AIRWAYS | OROPHARYNGEAL AIRWAYS Purpose: Used to maintain oral airway patency Anatomically shaped PVA or PVC airway with an appropriate contour that has a flanged buccal end with the size clearly moulded onto the plastic. Must have a robust bite block to prevent occlusion, smooth rounded distal end to minimise tissue trauma. The edges or corners that come into contact with patient tissue, must be smooth with a minimal radius of curvature as per ISO 5364:2016 One-piece design or reinforcement components for that do not separate easily or could be an aspiration risk Distance measured from buccal flange to tip of OPA with tolerance as per ISO 5364:2016 Validation of measurements must be provided. Brochures will be acceptable. All components must be latex-free, clinically clean and individually packed in peel pouch that is easy to open - For SINGLE use only Information brochure or the outer package of the item should clearly indicate latex free | Oropharyngeal airway (guedels) 70mm Size 2 (White) Minimum inside dimension: 4.0mm
19 | RT13-05-007 | AIRWAY DEVICES | OROPHARYNGEAL AIRWAYS | OROPHARYNGEAL AIRWAYS Purpose: Used to maintain oral airway patency Anatomically shaped PVA or PVC airway with an appropriate contour that has a flanged buccal end with the size clearly moulded onto the plastic. Must have a robust bite block to prevent occlusion, smooth rounded distal end to minimise tissue trauma. The edges or corners that come into contact with patient tissue, must be smooth with a minimal radius of curvature as per ISO 5364:2016 One-piece design or reinforcement components for that do not separate easily or could be an aspiration risk Distance measured from buccal flange to tip of OPA with tolerance as per ISO 5364:2016 Validation of measurements must be provided. Brochures will be acceptable. All components must be latex-free, clinically clean and individually packed in peel pouch that is easy to open - For SINGLE use only Information brochure or the outer package of the item should clearly indicate latex free | Oropharyngeal airway (guedels) 80mm ID: 3 (Green) Minimum inside dimension: 4.5mm
20 | RT13-05-008 | AIRWAY DEVICES | OROPHARYNGEAL AIRWAYS | OROPHARYNGEAL AIRWAYS Purpose: Used to maintain oral airway patency Anatomically shaped PVA or PVC airway with an appropriate contour that has a flanged buccal end with the size clearly moulded onto the plastic. Must have a robust bite block to prevent occlusion, smooth rounded distal end to minimise tissue trauma. The edges or corners that come into contact with patient tissue, must be smooth with a minimal radius of curvature as per ISO 5364:2016 One-piece design or reinforcement components for that do not separate easily or could be an aspiration risk Distance measured from buccal flange to tip of OPA with tolerance as per ISO 5364:2016 Validation of measurements must be provided. Brochures will be acceptable. All components must be latex-free, clinically clean and individually packed in peel pouch that is easy to open - For SINGLE use only Information brochure or the outer package of the item should clearly indicate latex free | Oropharyngeal airway (guedels) 90mm Size 4 (Yellow) Minimum inside dimension: 4.5mm
21 | RT13-05-009 | AIRWAY DEVICES | OROPHARYNGEAL AIRWAYS | OROPHARYNGEAL AIRWAYS Purpose: Used to maintain oral airway patency Anatomically shaped PVA or PVC airway with an appropriate contour that has a flanged buccal end with the size clearly moulded onto the plastic. Must have a robust bite block to prevent occlusion, smooth rounded distal end to minimise tissue trauma. The edges or corners that come into contact with patient tissue, must be smooth with a minimal radius of curvature as per ISO 5364:2016 One-piece design or reinforcement components for that do not separate easily or could be an aspiration risk Distance measured from buccal flange to tip of OPA with tolerance as per ISO 5364:2016 Validation of measurements must be provided. Brochures will be acceptable. All components must be latex-free, clinically clean and individually packed in peel pouch that is easy to open - For SINGLE use only Information brochure or the outer package of the item should clearly indicate latex free | Oropharyngeal airway (guedels) 100mm Size 5 (Red) Minimum inside dimension: 5.0mm
22 | RT13-05-010 | AIRWAY DEVICES | OROPHARYNGEAL AIRWAYS | OROPHARYNGEAL AIRWAYS Purpose: Used to maintain oral airway patency Anatomically shaped PVA or PVC airway with an appropriate contour that has a flanged buccal end with the size clearly moulded onto the plastic. Must have a robust bite block to prevent occlusion, smooth rounded distal end to minimise tissue trauma. The edges or corners that come into contact with patient tissue, must be smooth with a minimal radius of curvature as per ISO 5364:2016 One-piece design or reinforcement components for that do not separate easily or could be an aspiration risk Distance measured from buccal flange to tip of OPA with tolerance as per ISO 5364:2016 Validation of measurements must be provided. Brochures will be acceptable. All components must be latex-free, clinically clean and individually packed in peel pouch that is easy to open - For SINGLE use only Information brochure or the outer package of the item should clearly indicate latex free | Oropharyngeal airway (guedels) 110mm (Orange) Minimum inside dimension: 5.5mm
23 | RT13-05-011 | AIRWAY DEVICES | OROPHARYNGEAL AIRWAYS | OROPHARYNGEAL AIRWAYS Purpose: Used to maintain oral airway patency Anatomically shaped PVA or PVC airway with an appropriate contour that has a flanged buccal end with the size clearly moulded onto the plastic. Must have a robust bite block to prevent occlusion, smooth rounded distal end to minimise tissue trauma. The edges or corners that come into contact with patient tissue, must be smooth with a minimal radius of curvature as per ISO 5364:2016 One-piece design or reinforcement components for that do not separate easily or could be an aspiration risk Distance measured from buccal flange to tip of OPA with tolerance as per ISO 5364:2016 Validation of measurements must be provided. Brochures will be acceptable. All components must be latex-free, clinically clean and individually packed in peel pouch that is easy to open - For SINGLE use only Information brochure or the outer package of the item should clearly indicate latex free | Oropharyngeal airway (guedels) 120mm (Purple) Minimum inside dimension: 5.5mm
24 | RT13-06-001 | Oral Fibre-Optic Guide | Oral Fibre-Optic Guide | ORAL FIBRE OPTIC GUIDES Purpose: Modified oral airway specifically designed for oral fibre-optic intubation. Used to guide the tracheal tube and scope around the tongue and protect the scope should the patient bite down Straight with groove to allow removal of scope upon intubation. Latex free medical grade plastic SINGLE use, clinically clean. Individually packed in peel pouch Information brochure or the outer package of the item should clearly indicate latex free | Oral Fibre-Optic Guide – Size 000 Length: 40mm
25 | RT13-06-002 | Oral Fibre-Optic Guide | Oral Fibre-Optic Guide | Oral Fibre-Optic Guide Size 00 Length: 50mm ORAL FIBRE OPTIC GUIDES Purpose: Modified oral airway specifically designed for oral fibre-optic intubation. Used to guide the tracheal tube and scope around the tongue and protect the scope should the patient bite down Straight with groove to allow removal of scope upon intubation. Latex free medical grade plastic SINGLE use, clinically clean. Individually packed in peel pouch Information brochure or the outer package of the item should clearly indicate latex free | Oral Fibre-Optic Guide – Size 00 Length: 50mm
26 | RT13-06-003 | Oral Fibre-Optic Guide | Oral Fibre-Optic Guide | Oral Fibre-Optic Guide Size 0 Length: 60mm ORAL FIBRE OPTIC GUIDES Purpose: Modified oral airway specifically designed for oral fibre-optic intubation. Used to guide the tracheal tube and scope around the tongue and protect the scope should the patient bite down Straight with groove to allow removal of scope upon intubation. Latex free medical grade plastic SINGLE use, clinically clean. Individually packed in peel pouch | Oral Fibre-Optic Guide – Size 0 Length: 60mm
27 | RT13-06-004 | Oral Fibre-Optic Guide | Oral Fibre-Optic Guide | Oral Fibre-Optic Guide Size 1 Length: 70mm ORAL FIBRE OPTIC GUIDES Purpose: Modified oral airway specifically designed for oral fibre-optic intubation. Used to guide the tracheal tube and scope around the tongue and protect the scope should the patient bite down Straight with groove to allow removal of scope upon intubation. Latex free medical grade plastic SINGLE use, clinically clean. Individually packed in peel pouch Information brochure or the outer package of the item should clearly indicate latex free | Oral Fibre-Optic Guide – Size 1 Length: 70mm
28 | RT13-06-005 | Oral Fibre-Optic Guide | Oral Fibre-Optic Guide | Oral Fibre-Optic Guide Size 2 Length: 80mm ORAL FIBRE OPTIC GUIDES Purpose: Modified oral airway specifically designed for oral fibre-optic intubation. Used to guide the tracheal tube and scope around the tongue and protect the scope should the patient bite down Straight with groove to allow removal of scope upon intubation. Latex free medical grade plastic SINGLE use, clinically clean. Individually packed in peel pouch Information brochure or the outer package of the item should clearly indicate latex free | Oral Fibre-Optic Guide – Size 2 Length: 80mm
29 | RT13-06-006 | Oral Fibre-Optic Guide | Oral Fibre-Optic Guide | Oral Fibre-Optic Guide Size 3 Length: 90mm ORAL FIBRE OPTIC GUIDES Purpose: Modified oral airway specifically designed for oral fibre-optic intubation. Used to guide the tracheal tube and scope around the tongue and protect the scope should the patient bite down Straight with groove to allow removal of scope upon intubation. Latex free medical grade plastic SINGLE use, clinically clean. Individually packed in peel pouch Information brochure or the outer package of the item should clearly indicate latex free | Oral Fibre-Optic Guide – Size 3 Length: 90mm
30 | RT13-06-007 | Oral Fibre-Optic Guide | Oral Fibre-Optic Guide | Oral Fibre-Optic Guide Size 4 Length: 100mm ORAL FIBRE OPTIC GUIDES Purpose: Modified oral airway specifically designed for oral fibre-optic intubation. Used to guide the tracheal tube and scope around the tongue and protect the scope should the patient bite down Straight with groove to allow removal of scope upon intubation. Latex free medical grade plastic SINGLE use, clinically clean. Individually packed in peel pouch Information brochure or the outer package of the item should clearly indicate latex free | Oral Fibre-Optic Guide – Size 4 Length: 100mm
31 | RT13-06-008 | Oral Fibre-Optic Guide | Oral Fibre-Optic Guide | Oral Fibre-Optic Guide Size 5 Length: 110mm ORAL FIBRE OPTIC GUIDES Purpose: Modified oral airway specifically designed for oral fibre-optic intubation. Used to guide the tracheal tube and scope around the tongue and protect the scope should the patient bite down Straight with groove to allow removal of scope upon intubation. Latex free medical grade plastic SINGLE use, clinically clean. Individually packed in peel pouch Information brochure or the outer package of the item should clearly indicate latex free | Oral Fibre-Optic Guide – Size 5 Length: 110mm
32 | RT13-06-009 | Oral Fibre-Optic Guide | Oral Fibre-Optic Guide | Oral Fibre-optic Guide - Ovassapian type Size 2 ORAL FIBRE OPTIC GUIDES - Ovassapian type Adult Purpose: Modified oral airways specifically designed for oral fibre-optic intubation. Anatomically shaped EVA or PVC airway with an appropriate atraumatic contour The edges or corners that come into contact with patient tissue, must be smooth One-piece design with a strong bite-block to protect the fibre-optic should the patient bite down Must split longitudinally to facilitate the placing of an endotracheal tube over the fibre optic scope All components must be latex-free, clinically clean and individually packed in peel pouch that is easy to open - For SINGLE use only Information brochure or the outer package of the item should clearly indicate latex free | Oral Fibre-optic Guide - Ovassapian type Size 2
33 | RT13-06-010 | Oral Fibre-Optic Guide | Oral Fibre-Optic Guide | Oral Fibre-optic Guide - Ovassapian type Size 3 ORAL FIBRE OPTIC GUIDES - Ovassapian type Adult Purpose: Modified oral airways specifically designed for oral fibre-optic intubation. Anatomically shaped EVA or PVC airway with an appropriate atraumatic contour The edges or corners that come into contact with patient tissue, must be smooth One-piece design with a strong bite-block to protect the fibre-optic should the patient bite down Must split longitudinally to facilitate the placing of an endotracheal tube over the fibre optic scope All components must be latex-free, clinically clean and individually packed in peel pouch that is easy to open - For SINGLE use only Information brochure or the outer package of the item should clearly indicate latex free | Oral Fibre-optic Guide - Ovassapian type Size 3
34 | RT13-06-011 | Oral Fibre-Optic Guide | Oral Fibre-Optic Guide | Oral Fibre-optic Guide - Ovassapian type Size 4 ORAL FIBRE OPTIC GUIDES - Ovassapian type Adult Purpose: Modified oral airways specifically designed for oral fibre-optic intubation. Anatomically shaped EVA or PVC airway with an appropriate atraumatic contour The edges or corners that come into contact with patient tissue, must be smooth One-piece design with a strong bite-block to protect the fibre-optic should the patient bite down Must split longitudinally to facilitate the placing of an endotracheal tube over the fibre optic scope All components must be latex-free, clinically clean and individually packed in peel pouch that is easy to open - For SINGLE use only Information brochure or the outer package of the item should clearly indicate latex free | Oral Fibre-optic Guide - Ovassapian type Size 4
35 | RT13-07-003 | AIRWAY DEVICES | Nasopharyngeal airway | Nasopharyngeal airway with Oxygen Port – 25Fr ID: 6.0 NASOPHARYNGEAL AIRWAYS with OXYGEN PORTS Purpose: Nasally placed airway management Anatomically shaped soft silicone airway with an appropriate contour that maintains the pharyngeal curvature without collapse of the lumen Proximal end: Trumpeted end with integrated 15mm oxygen adaptor/ flanged end with a15mm oxygen adaptor. Separate oxygen elbow to allow for oxygen delivery Distal end: Soft atraumatic bevelled tip Must have sizing and dimension clearly marked on the device All components must be MRI compatible, latex-free, clinically clean and individually packed in peel pouch that is easy to open -For SINGLE use only Information brochure or the outer package of the item should clearly indicate latex free | Nasopharyngeal airway with Oxygen Port – 25Fr ID: 6.0
36 | RT13-07-004 | AIRWAY DEVICES | Nasopharyngeal airway | Nasopharyngeal airway with Oxygen Port – 29Fr ID: 7.0 NASOPHARYNGEAL AIRWAYS with OXYGEN PORTS Purpose: Nasally placed airway management Anatomically shaped soft silicone airway with an appropriate contour that maintains the pharyngeal curvature without collapse of the lumen Proximal end: Trumpeted end with integrated 15mm oxygen adaptor/ flanged end with a15mm oxygen adaptor. Separate oxygen elbow to allow for oxygen delivery Distal end: Soft atraumatic bevelled tip Must have sizing and dimension clearly marked on the device All components must be MRI compatible, latex-free, clinically clean and individually packed in peel pouch that is easy to open -For SINGLE use only Information brochure or the outer package of the item should clearly indicate latex free | Nasopharyngeal airway with Oxygen Port – 29Fr ID: 7.0
37 | RT13-07-005 | AIRWAY DEVICES | Nasopharyngeal airway | Nasopharyngeal airway with Oxygen Port – 32Fr ID: 8.0 NASOPHARYNGEAL AIRWAYS with OXYGEN PORTS Purpose: Nasally placed airway management Anatomically shaped soft silicone airway with an appropriate contour that maintains the pharyngeal curvature without collapse of the lumen Proximal end: Trumpeted end with integrated 15mm oxygen adaptor/ flanged end with a15mm oxygen adaptor. Separate oxygen elbow to allow for oxygen delivery Distal end: Soft atraumatic bevelled tip Must have sizing and dimension clearly marked on the device All components must be MRI compatible, latex-free, clinically clean and individually packed in peel pouch that is easy to open -For SINGLE use only Information brochure or the outer package of the item should clearly indicate latex free | Nasopharyngeal airway with Oxygen Port – 32Fr ID: 8.0
38 | RT13-07-006 | AIRWAY DEVICES | Nasopharyngeal airway | Nasopharyngeal airway with Oxygen Port and Capnography Line – 18Fr ID: 4.0 NASOPHARYNGEAL AIRWAYS with OXYGEN AND CAPNOGRAPHY PORTS Purpose: Nasally placed airway management Anatomically shaped soft silicone airway with an appropriate contour that maintains the pharyngeal curvature without collapse of the lumen Proximal end: Trumpeted end with integrated 15mm oxygen adaptor/ flanged end with a15mm oxygen adaptor. Separate oxygen elbow with direct side stream sampling line to allow for oxygen delivery and ETCO2 monitoring. Distal end: Soft atraumatic bevelled tip Must have sizing and dimension clearly marked on the device All components must be MRI compatible, latex-free, clinically clean and individually packed in peel pouch that is easy to open - For SINGLE use only Information brochure or the outer package of the item should clearly indicate latex free Information brochure or the outer package of the item should clearly indicate latex free | Nasopharyngeal airway with Oxygen Port and Capnography Line – 18Fr ID: 4.0
39 | RT13-07-007 | AIRWAY DEVICES | Nasopharyngeal airway | Nasopharyngeal airway with Oxygen Port and Capnography Line – 22Fr ID: 5.0 NASOPHARYNGEAL AIRWAYS with OXYGEN AND CAPNOGRAPHY PORTS Purpose: Nasally placed airway management Anatomically shaped soft silicone airway with an appropriate contour that maintains the pharyngeal curvature without collapse of the lumen Proximal end: Trumpeted end with integrated 15mm oxygen adaptor/ flanged end with a15mm oxygen adaptor. Separate oxygen elbow with direct side stream sampling line to allow for oxygen delivery and ETCO2 monitoring. Distal end: Soft atraumatic bevelled tip Must have sizing and dimension clearly marked on the device All components must be MRI compatible, latex-free, clinically clean and individually packed in peel pouch that is easy to open - For SINGLE use only Information brochure or the outer package of the item should clearly indicate latex free | Nasopharyngeal airway with Oxygen Port and Capnography Line – 22Fr ID: 5.0
40 | RT13-07-008 | AIRWAY DEVICES | Nasopharyngeal airway | Nasopharyngeal airway with Oxygen Port and Capnography Line – 25Fr ID: 6.0 NASOPHARYNGEAL AIRWAYS with OXYGEN AND CAPNOGRAPHY PORTS Purpose: Nasally placed airway management Anatomically shaped soft silicone airway with an appropriate contour that maintains the pharyngeal curvature without collapse of the lumen Proximal end: Trumpeted end with integrated 15mm oxygen adaptor/ flanged end with a15mm oxygen adaptor. Separate oxygen elbow with direct side stream sampling line to allow for oxygen delivery and ETCO2 monitoring. Distal end: Soft atraumatic bevelled tip Must have sizing and dimension clearly marked on the device All components must be MRI compatible, latex-free, clinically clean and individually packed in peel pouch that is easy to open - For SINGLE use only Information brochure or the outer package of the item should clearly indicate latex free | Nasopharyngeal airway with Oxygen Port and Capnography Line – 25Fr ID: 6.0
41 | RT13-07-009 | AIRWAY DEVICES | Nasopharyngeal airway | Nasopharyngeal airway with Oxygen Port and Capnography Line – 29Fr ID: 7.0 NASOPHARYNGEAL AIRWAYS with OXYGEN AND CAPNOGRAPHY PORTS Purpose: Nasally placed airway management Anatomically shaped soft silicone airway with an appropriate contour that maintains the pharyngeal curvature without collapse of the lumen Proximal end: Trumpeted end with integrated 15mm oxygen adaptor/ flanged end with a15mm oxygen adaptor. Separate oxygen elbow with direct side stream sampling line to allow for oxygen delivery and ETCO2 monitoring. Distal end: Soft atraumatic bevelled tip Must have sizing and dimension clearly marked on the device All components must be MRI compatible, latex-free, clinically clean and individually packed in peel pouch that is easy to open - For SINGLE use only Information brochure or the outer package of the item should clearly indicate latex free | Nasopharyngeal airway with Oxygen Port and Capnography Line – 29Fr ID: 7.0
42 | RT13-07-010 | AIRWAY DEVICES | Nasopharyngeal airway | Nasopharyngeal airway with Oxygen Port and Capnography Line – 32Fr ID: 8.0 NASOPHARYNGEAL AIRWAYS with OXYGEN AND CAPNOGRAPHY PORTS Purpose: Nasally placed airway management Anatomically shaped soft silicone airway with an appropriate contour that maintains the pharyngeal curvature without collapse of the lumen Proximal end: Trumpeted end with integrated 15mm oxygen adaptor/ flanged end with a15mm oxygen adaptor. Separate oxygen elbow with direct side stream sampling line to allow for oxygen delivery and ETCO2 monitoring. Distal end: Soft atraumatic bevelled tip Must have sizing and dimension clearly marked on the device All components must be MRI compatible, latex-free, clinically clean and individually packed in peel pouch that is easy to open - For SINGLE use only Information brochure or the outer package of the item should clearly indicate latex free | Nasopharyngeal airway with Oxygen Port and Capnography Line – 32Fr ID: 8.0
43 | RT13-09-004 | Oxygen Mask | CPAP Mask | CPAP MASK - x Large Purpose: For non-invasive continuous positive airway pressure delivery The double-ported mask must connect to the PEEP valve and the oxygen source. The mask must have a hard clear durable anatomically moulded face plate with inflatable rim with hooks to secure harness. The system must be able to create a good seal between the mask and patient's face, must be atraumatic to the nasal bridge and must not cause skin trauma. There must be hook rings for the attachment of the harness Latex-free, single use, individually packed and surgical clean. Information brochure or the outer package of the item should clearly indicate latex free | CPAP Mask - X-Large
44 | RT13-09-005 | Oxygen Mask | CPAP Mask | CPAP MASK - X Large Purpose: For non-invasive continuous positive airway pressure delivery The double-ported mask must connect to the PEEP valve and the oxygen source. The mask must have a hard clear durable anatomically moulded face plate with inflatable rim with hooks to secure harness. The system must be able to create a good seal between the mask and patient's face, must be atraumatic to the nasal bridge and must not cause skin trauma. There must be hook rings for the attachment of the harness Latex-free, single use, individually packed and surgical clean. Information brochure or the outer package of the item should clearly indicate latex free | CPAP System - X-Large
45 | RT13-09-006 | Oxygen Mask | CPAP Mask | CONSTANT POSITIVE AIRWAY PRESSURE SYSTEM (CPAP) Purpose: To administer fresh gas to the patient through the entire respiratory cycle. The system must include a mask with a pressure relief valve, harness, oxygen tubing and a PEEP valve. The double-ported mask must connect to the PEEP valve and the oxygen source. The mask must have a hard clear durable anatomically moulded face plate with inflatable rim with hooks to secure harness. The system must be able to create a good seal between the mask and patient's face, must be atraumatic to the nasal bridge and must not cause skin trauma. There must be hook rings for the attachment of the harness The harness must be comfortable, adjustable, easy to release and non-irritant. An adjustable PEEP valve or different sized PEEP valves must be included. Must be supplied with kink resistant O2 delivery tubing ± 2m. A standard flow meter should achieve 5-20cmH2O of CPAP. Additional adapters must be included to fit PEEP valves and oxygen tubing Latex-free, single use, individually packed and surgical clean. Information brochure or the outer package of the item should clearly indicate latex free | CPAP Mask - Paediatric
46 | RT13-09-010 | Oxygen Mask | CPAP Mask | CONSTANT POSITIVE AIRWAY PRESSURE SYSTEM (CPAP) Purpose: To administer fresh gas to the patient through the entire respiratory cycle. The system must include a mask with a pressure relief valve, harness, oxygen tubing and a PEEP valve. The double-ported mask must connect to the PEEP valve and the oxygen source. The mask must have a hard clear durable anatomically moulded face plate with inflatable rim with hooks to secure harness. The system must be able to create a good seal between the mask and patient's face, must be atraumatic to the nasal bridge and must not cause skin trauma. There must be hook rings for the attachment of the harness The harness must be comfortable, adjustable, easy to release and non-irritant. An adjustable PEEP valve or different sized PEEP valves must be included. Must be supplied with kink resistant O2 delivery tubing ± 2m. A standard flow meter should achieve 5-20cmH2O of CPAP. Additional adapters must be included to fit PEEP valves and oxygen tubing Latex-free, single use, individually packed and surgical clean. Information brochure or the outer package of the item should clearly indicate latex free | CPAP System - Paediatric
47 | RT13-11-001 | Resuscitation and Re-breathing bags | resuscitation bag | BAG VALVE MASK RESUSCITATION BAG WITH PEEP VALVE - AUTOCLAVABLE Purpose: Used to provide positive pressure ventilation to a patient who is not breathing or who is breathing inadequately Ventilation bag: Clear, robust silicone. Self-inflating with a fast recoil time, and thin-walled bag to allow lung compliance assessment. Must have an ergonomic grip, Oxygen reservoir bag: Reusable PVC One-way Intake / Reservoir valve between ventilation bag and reservoir bag that only allows one way flow into the ventilation bag Oxygen from reservoir bag and room air must enter the ventilation bag if fresh gas flow is inadequate. Secure connector for oxygen tubing. Oxygen tubing: >1 m Patient valves between ventilation bag and mask. Comprising of a lip valve (duck-billed valve), expiratory valve to prevent entrainment of room air, pressure release valve – 35cmH20 (must be printed on valve) Standard connector 15mm ID, 22mm OD. Must have a PEEP valve Masks with comfortable patient fit and user interface. Malleable, non-collapsible with smooth surface or ridges. 1 piece modular face mask Able to create a tight seal between the mask and patient's’ face and must be transparent. There must be a secure fit between mask connector and breathing circuit All connections between components must be easily assembled and disassembled with tight fit and no leak. Must not disassemble during normal use Manufactured from latex-free, clear silicone, PVC, and medical grade plastic - All components Autoclavable Information brochure or the outer package of the item should clearly indicate latex free | Bag valve mask resuscitation bag with PEEP valve. Preterm Patient size: Up to 2.5kg Ventilation bag: Volume: 240ml , Oxygen reservoir bag: Volume: 600ml Peep valve adjustable to ± 1.5-10 cmH2O Must be supplied with 2 round masks – size 00 and 0/1
48 | RT13-11-002 | Resuscitation and Re-breathing bags | resuscitation bag | BAG VALVE MASK RESUSCITATION BAG WITH PEEP VALVE - AUTOCLAVABLE Purpose: Used to provide positive pressure ventilation to a patient who is not breathing or who is breathing inadequately Ventilation bag: Clear, robust silicone. Self-inflating with a fast recoil time, and thin-walled bag to allow lung compliance assessment. Must have an ergonomic grip, Oxygen reservoir bag: Reusable PVC One-way Intake / Reservoir valve between ventilation bag and reservoir bag that only allows one way flow into the ventilation bag Oxygen from reservoir bag and room air must enter the ventilation bag if fresh gas flow is inadequate. Secure connector for oxygen tubing. Oxygen tubing: >1 m Patient valves between ventilation bag and mask. Comprising of a lip valve (duck-billed valve), expiratory valve to prevent entrainment of room air, pressure release valve – 35cmH20 (must be printed on valve) Standard connector 15mm ID, 22mm OD. Must have a PEEP valve Masks with comfortable patient fit and user interface. Malleable, non-collapsible with smooth surface or ridges. 1 piece modular face mask Able to create a tight seal between the mask and patient's’ face and must be transparent. There must be a secure fit between mask connector and breathing circuit All connections between components must be easily assembled and disassembled with tight fit and no leak. Must not disassemble during normal use Manufactured from latex-free, clear silicone, PVC, and medical grade plastic - All components Autoclavable Information brochure or the outer package of the item should clearly indicate latex free | Bag valve mask resuscitation bag with PEEP valve. Paediatric Patient size: 2.5kg – 20kg Ventilation bag: Volume: 500ml Oxygen reservoir bag: Volume: 600ml Peep valve adjustable to ± 1.5-10 cmH2O Must be supplied with 2 round masks – size 0/1 and 2
49 | RT13-11-003 | Resuscitation and Re-breathing bags | resuscitation bag | BAG VALVE MASK RESUSCITATION BAG WITH PEEP VALVE - AUTOCLAVABLE Purpose: Used to provide positive pressure ventilation to a patient who is not breathing or who is breathing inadequately Ventilation bag: Clear, robust silicone. Self-inflating with a fast recoil time, and thin-walled bag to allow lung compliance assessment. Must have an ergonomic grip, Oxygen reservoir bag: Reusable PVC One-way Intake / Reservoir valve between ventilation bag and reservoir bag that only allows one way flow into the ventilation bag Oxygen from reservoir bag and room air must enter the ventilation bag if fresh gas flow is inadequate. Secure connector for oxygen tubing. Oxygen tubing: >1 m Patient valves between ventilation bag and mask. Comprising of a lip valve (duck-billed valve), expiratory valve to prevent entrainment of room air, pressure release valve – 35cmH20 (must be printed on valve) Standard connector 15mm ID, 22mm OD. Must have a PEEP valve Masks with comfortable patient fit and user interface. Malleable, non-collapsible with smooth surface or ridges. 1 piece modular face mask Able to create a tight seal between the mask and patient's’ face and must be transparent. There must be a secure fit between mask connector and breathing circuit All connections between components must be easily assembled and disassembled with tight fit and no leak. Must not disassemble during normal use Manufactured from latex-free, clear silicone, PVC, and medical grade plastic - All components Autoclavable Information brochure or the outer package of the item should clearly indicate latex free | Bag valve mask resuscitation bag with PEEP valve.Adult Patient size: > 40kg Ventilation bag: Volume: 1000- 1600ml Oxygen reservoir bag: Volume: 2600ml Peep valve adjustable to 5-20 cm/H20 Must be supplied with 2 anatomically correct silicone masks – size 4 and 5
50 | RT13-11-004 | Resuscitation and Re-breathing bags | resuscitation bag | BAG VALVE MASK RESUSCITATION BAG WITH PEEP VALVE - SINGLE USE ONLY Purpose: Used to provide positive pressure ventilation to a patient who is not breathing or who is breathing inadequately Ventilation bag: Clear, robust silicone. Self-inflating with a fast recoil time, and thin-walled bag to allow lung compliance assessment. Must have an ergonomic grip, Oxygen reservoir bag: PVC One-way Intake / Reservoir valve between ventilation bag and reservoir bag that only allows one way flow into the ventilation bag Oxygen from reservoir bag and room air must enter the ventilation bag if fresh gas flow is inadequate. Secure connector for oxygen tubing. Oxygen tubing: >1 m Patient valves between ventilation bag and mask. Comprising of a lip valve (duck-billed valve), expiratory valve to prevent entrainment of room air, pressure release valve – 35cmH20 (must be printed on valve) Standard connector 15mm ID, 22mm OD. Must have a PEEP valve Masks with comfortable patient fit and user interface. Malleable, non-collapsible with smooth surface or ridges. 1 piece modular face mask Able to create a tight seal between the mask and patient's’ face and must be transparent. There must be a secure fit between mask connector and breathing circuit All connections between components must be easily assembled and disassembled with tight fit and no leak. Must not disassemble during normal use Manufactured from latex-free, clear silicone, PVC, and medical grade plastic Information brochure or the outer package of the item should clearly indicate latex free | Bag valve mask resuscitation bag with PEEP valve. Infant Patient Ventilation bag: Volume: 240ml , Oxygen reservoir bag: Volume: 600ml Peep valve adjustable to ± 1.5-10 cmH2O Must be supplied with 2 round masks – size 00 and 0/1
51 | RT13-11-005 | Resuscitation and Re-breathing bags | resuscitation bag | BAG VALVE MASK RESUSCITATION BAG WITH PEEP VALVE - SINGLE USE ONLY Purpose: Used to provide positive pressure ventilation to a patient who is not breathing or who is breathing inadequately Ventilation bag: Clear, robust silicone. Self-inflating with a fast recoil time, and thin-walled bag to allow lung compliance assessment. Must have an ergonomic grip, Oxygen reservoir bag: PVC One-way Intake / Reservoir valve between ventilation bag and reservoir bag that only allows one way flow into the ventilation bag Oxygen from reservoir bag and room air must enter the ventilation bag if fresh gas flow is inadequate. Secure connector for oxygen tubing. Oxygen tubing: >1 m Patient valves between ventilation bag and mask. Comprising of a lip valve (duck-billed valve), expiratory valve to prevent entrainment of room air, pressure release valve – 35cmH20 (must be printed on valve) Standard connector 15mm ID, 22mm OD. Must have a PEEP valve Masks with comfortable patient fit and user interface. Malleable, non-collapsible with smooth surface or ridges. 1 piece modular face mask Able to create a tight seal between the mask and patient's’ face and must be transparent. There must be a secure fit between mask connector and breathing circuit All connections between components must be easily assembled and disassembled with tight fit and no leak. Must not disassemble during normal use Manufactured from latex-free, clear silicone, PVC, and medical grade plastic Information brochure or the outer package of the item should clearly indicate latex free | Bag valve mask resuscitation bag with PEEP valve. Paediatric Patient Ventilation bag: Volume: 500ml Oxygen reservoir bag: Volume: 1600ml Peep valve adjustable to ± 1.5-10 cmH2O Must be supplied with 2 round masks – size 0/1 and 2
52 | RT13-11-006 | Resuscitation and Re-breathing bags | resuscitation bag | BAG VALVE MASK RESUSCITATION BAG WITH PEEP VALVE - SINGLE USE ONLY Purpose: Used to provide positive pressure ventilation to a patient who is not breathing or who is breathing inadequately Ventilation bag: Clear, robust silicone. Self-inflating with a fast recoil time, and thin-walled bag to allow lung compliance assessment. Must have an ergonomic grip, Oxygen reservoir bag: PVC One-way Intake / Reservoir valve between ventilation bag and reservoir bag that only allows one way flow into the ventilation bag Oxygen from reservoir bag and room air must enter the ventilation bag if fresh gas flow is inadequate. Secure connector for oxygen tubing. Oxygen tubing: >1 m Patient valves between ventilation bag and mask. Comprising of a lip valve (duck-billed valve), expiratory valve to prevent entrainment of room air, pressure release valve – 35cmH20 (must be printed on valve) Standard connector 15mm ID, 22mm OD. Must have a PEEP valve Masks with comfortable patient fit and user interface. Malleable, non-collapsible with smooth surface or ridges. 1 piece modular face mask Able to create a tight seal between the mask and patient's’ face and must be transparent. There must be a secure fit between mask connector and breathing circuit All connections between components must be easily assembled and disassembled with tight fit and no leak. Must not disassemble during normal use Manufactured from latex-free, clear silicone, PVC, and medical grade plastic Information brochure or the outer package of the item should clearly indicate latex free | Bag valve mask resuscitation bag with PEEP valve.Adult Patient Ventilation bag: Volume: 1000- 1600ml Oxygen reservoir bag: Volume: 2600ml Peep valve adjustable to 5-20 cm/H20 Must be supplied with 2 anatomically correct silicone masks – size 4 and 5
53 | RT13-11-007 | Resuscitation and Re-breathing bags | resuscitation bag | Bag valve mask resuscitation bag with PEEP valve. Preterm Patient size: Up to 2.5kg Ventilation bag: Volume: 240ml , Oxygen reservoir bag: Volume: 600ml Peep valve adjustable to ± 1.5-10 cmH2O Must be supplied with 2 round masks – size 00 and 0/1 BAG VALVE MASK RESUSCITATION BAG WITH PEEP VALVE - AUTOCLAVABLE Purpose: Used to provide positive pressure ventilation to a patient who is not breathing or who is breathing inadequately Ventilation bag: Clear, robust silicone. Self-inflating with a fast recoil time, and thin-walled bag to allow lung compliance assessment. Must have an ergonomic grip, Oxygen reservoir bag: Reusable PVC One-way Intake / Reservoir valve between ventilation bag and reservoir bag that only allows one way flow into the ventilation bag Oxygen from reservoir bag and room air must enter the ventilation bag if fresh gas flow is inadequate. Secure connector for oxygen tubing. Oxygen tubing: >1 m Patient valves between ventilation bag and mask. Comprising of a lip valve (duck-billed valve), expiratory valve to prevent entrainment of room air, pressure release valve – 35cmH20 (must be printed on valve) Standard connector 15mm ID, 22mm OD. Must have a PEEP valve Masks with comfortable patient fit and user interface. Malleable, non-collapsible with smooth surface or ridges. 1 piece modular face mask Able to create a tight seal between the mask and patient's’ face and must be transparent. There must be a secure fit between mask connector and breathing circuit All connections between components must be easily assembled and disassembled with tight fit and no leak. Must not disassemble during normal use Manufactured from latex-free, clear silicone, PVC, and medical grade plastic - All components Autoclavable | Bag valve mask resuscitation bag with PEEP valve. Preterm Patient size: Up to 2.5kg Ventilation bag: Volume: 240ml , Oxygen reservoir bag: Volume: 600ml Peep valve adjustable to ± 1.5-10 cmH2O Must be supplied with 2 round masks – size 00 and 0/1
54 | RT13-13-001 | MOUTHPIECE | MOUTHPIECE | MOUTHPIECE FOR IPPB MACHINE ID: 15mm OD: 22mm Consists of a mouthpiece made of translucent molded-plastic . Must have a comfortable fit and be compatible with IPPB machines. Single use, individually packed and surgical clean | MOUTHPIECE FOR IPPB MACHINE ID: 15mm OD: 22mm
55 | RT13-16-019 | Cuffed and Uncuffed Tubes | UNCUFFED EXTRA-SOFT PLASTICIZED PVC | UNCUFFED EXTRA-SOFT PLASTICIZED PVC ORAL/NASAL TRACHEAL TUBE FOR LONG TERM INTUBATION Purpose: Long term oral or nasal intubation in paediatric patient's Consists of a firm and kink resistant extra-soft plasticized PVC with a full Magill curvature that softens at body temperature to conform to anatomy Must have a radiopaque line and must be marked with numbered depth markings in 50mm or 1cm intervals and the side of the tube with tube size, length, ID, OD, trade name. The markings on tube must be clear and must not rub off Must have a precision band mark to aid with depth of intubation and guide proper placement. Solid or circumferential precision band markings accepted and must be no more than 20mm from the tip Proximal end - secure 15mm OD circuit connector made from medical grade plastic Distal end - Atraumatic left-facing bevelled tip without Murphy eye All components must be latex and DEHP free, sterile and individually packed in peel pouch that is easy to open - For SINGLE use only To comply with the latest ISO 5361 Information brochure or the outer package of the item should clearly indicate latex free and DEHP free | Uncuffed Extra-soft Plasticized PVC Oral/Nasal tracheal tube for long-term intubation. ID: 2.0mm OD: 3.4mm Length: 165mm Depth markings in 0.5cm intervals, starting at 3cm.
56 | RT13-16-020 | Cuffed and Uncuffed Tubes | UNCUFFED EXTRA-SOFT PLASTICIZED PVC | UNCUFFED EXTRA-SOFT PLASTICIZED PVC ORAL/NASAL TRACHEAL TUBE FOR LONG TERM INTUBATION Purpose: Long term oral or nasal intubation in paediatric patient's Consists of a firm and kink resistant extra-soft plasticized PVC with a full Magill curvature that softens at body temperature to conform to anatomy Must have a radiopaque line and must be marked with numbered depth markings in 50mm or 1cm intervals and the side of the tube with tube size, length, ID, OD, trade name. The markings on tube must be clear and must not rub off Must have a precision band mark to aid with depth of intubation and guide proper placement. Solid or circumferential precision band markings accepted and must be no more than 20mm from the tip Proximal end - secure 15mm OD circuit connector made from medical grade plastic Distal end - Atraumatic left-facing bevelled tip without Murphy eye All components must be latex and DEHP free, sterile and individually packed in peel pouch that is easy to open - For SINGLE use only To comply with the latest ISO 5361 Information brochure or the outer package of the item should clearly indicate latex free and DEHP free | Uncuffed Extra-soft Plasticized PVC Oral/Nasal tracheal tube for long-term intubation. ID: 2.5mm OD: 4.1mm Length: 165mm Depth markings in 0.5cm intervals, starting at 3cm.
57 | RT13-16-021 | Cuffed and Uncuffed Tubes | UNCUFFED EXTRA-SOFT PLASTICIZED PVC | UNCUFFED EXTRA-SOFT PLASTICIZED PVC ORAL/NASAL TRACHEAL TUBE FOR LONG TERM INTUBATION Purpose: Long term oral or nasal intubation in paediatric patient's Consists of a firm and kink resistant extra-soft plasticized PVC with a full Magill curvature that softens at body temperature to conform to anatomy Must have a radiopaque line and must be marked with numbered depth markings in 50mm or 1cm intervals and the side of the tube with tube size, length, ID, OD, trade name. The markings on tube must be clear and must not rub off Must have a precision band mark to aid with depth of intubation and guide proper placement. Solid or circumferential precision band markings accepted and must be no more than 20mm from the tip Proximal end - secure 15mm OD circuit connector made from medical grade plastic Distal end - Atraumatic left-facing bevelled tip without Murphy eye All components must be latex and DEHP free, sterile and individually packed in peel pouch that is easy to open - For SINGLE use only To comply with the latest ISO 5361 Information brochure or the outer package of the item should clearly indicate latex free and DEHP free | Uncuffed Extra-soft Plasticized PVC Oral/Nasal tracheal tube for long-term intubation. ID: 3.0mm OD: 4.6mm Length: 165mm Depth markings in 0.5cm intervals, starting at 3cm.
58 | RT13-16-022 | Cuffed and Uncuffed Tubes | UNCUFFED EXTRA-SOFT PLASTICIZED PVC | UNCUFFED EXTRA-SOFT PLASTICIZED PVC ORAL/NASAL TRACHEAL TUBE FOR LONG TERM INTUBATION Purpose: Long term oral or nasal intubation in paediatric patient's Consists of a firm and kink resistant extra-soft plasticized PVC with a full Magill curvature that softens at body temperature to conform to anatomy Must have a radiopaque line and must be marked with numbered depth markings in 50mm or 1cm intervals and the side of the tube with tube size, length, ID, OD, trade name. The markings on tube must be clear and must not rub off Must have a precision band mark to aid with depth of intubation and guide proper placement. Solid or circumferential precision band markings accepted and must be no more than 20mm from the tip Proximal end - secure 15mm OD circuit connector made from medical grade plastic Distal end - Atraumatic left-facing bevelled tip without Murphy eye All components must be latex and DEHP free, sterile and individually packed in peel pouch that is easy to open - For SINGLE use only To comply with the latest ISO 5361 Information brochure or the outer package of the item should clearly indicate latex free and DEHP free | Uncuffed Extra-soft Plasticized PVC Oral/Nasal tracheal tube for long-term intubation. ID: 3.5mm OD: 5.2mm Length: 165mm Depth markings in 0.5cm intervals starting at 3cm.
59 | RT13-16-023 | Cuffed and Uncuffed Tubes | UNCUFFED EXTRA-SOFT PLASTICIZED PVC | UNCUFFED EXTRA-SOFT PLASTICIZED PVC ORAL/NASAL TRACHEAL TUBE FOR LONG TERM INTUBATION Purpose: Long term oral or nasal intubation in paediatric patient's Consists of a firm and kink resistant extra-soft plasticized PVC with a full Magill curvature that softens at body temperature to conform to anatomy Must have a radiopaque line and must be marked with numbered depth markings in 50mm or 1cm intervals and the side of the tube with tube size, length, ID, OD, trade name. The markings on tube must be clear and must not rub off Must have a precision band mark to aid with depth of intubation and guide proper placement. Solid or circumferential precision band markings accepted and must be no more than 20mm from the tip Proximal end - secure 15mm OD circuit connector made from medical grade plastic Distal end - Atraumatic left-facing bevelled tip without Murphy eye All components must be latex and DEHP free, sterile and individually packed in peel pouch that is easy to open - For SINGLE use only To comply with the latest ISO 5361 Information brochure or the outer package of the item should clearly indicate latex free and DEHP free | Uncuffed Extra-soft Plasticized PVC Oral/Nasal tracheal tube for long-term intubation. ID: 4.0mm OD: 5.7mm Length: 230mm Depth markings in 1cm intervals from 9 - 21cm.
60 | RT13-16-024 | Cuffed and Uncuffed Tubes | UNCUFFED EXTRA-SOFT PLASTICIZED PVC | UNCUFFED EXTRA-SOFT PLASTICIZED PVC ORAL/NASAL TRACHEAL TUBE FOR LONG TERM INTUBATION Purpose: Long term oral or nasal intubation in paediatric patient's Consists of a firm and kink resistant extra-soft plasticized PVC with a full Magill curvature that softens at body temperature to conform to anatomy Must have a radiopaque line and must be marked with numbered depth markings in 50mm or 1cm intervals and the side of the tube with tube size, length, ID, OD, trade name. The markings on tube must be clear and must not rub off Must have a precision band mark to aid with depth of intubation and guide proper placement. Solid or circumferential precision band markings accepted and must be no more than 20mm from the tip Proximal end - secure 15mm OD circuit connector made from medical grade plastic Distal end - Atraumatic left-facing bevelled tip without Murphy eye All components must be latex and DEHP free, sterile and individually packed in peel pouch that is easy to open - For SINGLE use only To comply with the latest ISO 5361 Information brochure or the outer package of the item should clearly indicate latex free and DEHP free | Uncuffed Extra-soft Plasticized PVC Oral/Nasal tracheal tube for long-term intubation. ID: 4.5mm OD: 6.2mm Length: 230mm Depth markings in 1cm intervals from 9 - 21cm
61 | RT13-16-025 | Cuffed and Uncuffed Tubes | UNCUFFED EXTRA-SOFT PLASTICIZED PVC | UNCUFFED EXTRA-SOFT PLASTICIZED PVC ORAL/NASAL TRACHEAL TUBE FOR LONG TERM INTUBATION Purpose: Long term oral or nasal intubation in paediatric patient's Consists of a firm and kink resistant extra-soft plasticized PVC with a full Magill curvature that softens at body temperature to conform to anatomy Must have a radiopaque line and must be marked with numbered depth markings in 50mm or 1cm intervals and the side of the tube with tube size, length, ID, OD, trade name. The markings on tube must be clear and must not rub off Must have a precision band mark to aid with depth of intubation and guide proper placement. Solid or circumferential precision band markings accepted and must be no more than 20mm from the tip Proximal end - secure 15mm OD circuit connector made from medical grade plastic Distal end - Atraumatic left-facing bevelled tip without Murphy eye All components must be latex and DEHP free, sterile and individually packed in peel pouch that is easy to open - For SINGLE use only To comply with the latest ISO 5361 Information brochure or the outer package of the item should clearly indicate latex free and DEHP free | Uncuffed Extra-soft Plasticized PVC Oral/Nasal tracheal tube for long-term intubation. ID: 5.0mm OD: 7.0mm Length: 230mm Depth markings in 1cm intervals from 13 - 21cm
62 | RT13-16-026 | Cuffed and Uncuffed Tubes | UNCUFFED EXTRA-SOFT PLASTICIZED PVC | UNCUFFED EXTRA-SOFT PLASTICIZED PVC ORAL/NASAL TRACHEAL TUBE FOR LONG TERM INTUBATION Purpose: Long term oral or nasal intubation in paediatric patient's Consists of a firm and kink resistant extra-soft plasticized PVC with a full Magill curvature that softens at body temperature to conform to anatomy Must have a radiopaque line and must be marked with numbered depth markings in 50mm or 1cm intervals and the side of the tube with tube size, length, ID, OD, trade name. The markings on tube must be clear and must not rub off Must have a precision band mark to aid with depth of intubation and guide proper placement. Solid or circumferential precision band markings accepted and must be no more than 20mm from the tip Proximal end - secure 15mm OD circuit connector made from medical grade plastic Distal end - Atraumatic left-facing bevelled tip without Murphy eye All components must be latex and DEHP free, sterile and individually packed in peel pouch that is easy to open - For SINGLE use only To comply with the latest ISO 5361 Information brochure or the outer package of the item should clearly indicate latex free and DEHP free | Uncuffed Extra-soft Plasticized PVC Oral/Nasal tracheal tube for long-term intubation. ID: 5.5mm OD: 8.0mm Length: 270mm Depth markings in 1cm intervals from 15 - 25cm
63 | RT13-16-027 | Cuffed and Uncuffed Tubes | UNCUFFED EXTRA-SOFT PLASTICIZED PVC | UNCUFFED EXTRA-SOFT PLASTICIZED PVC ORAL/NASAL TRACHEAL TUBE FOR LONG TERM INTUBATION Purpose: Long term oral or nasal intubation in paediatric patient's Consists of a firm and kink resistant extra-soft plasticized PVC with a full Magill curvature that softens at body temperature to conform to anatomy Must have a radiopaque line and must be marked with numbered depth markings in 50mm or 1cm intervals and the side of the tube with tube size, length, ID, OD, trade name. The markings on tube must be clear and must not rub off Must have a precision band mark to aid with depth of intubation and guide proper placement. Solid or circumferential precision band markings accepted and must be no more than 20mm from the tip Proximal end - secure 15mm OD circuit connector made from medical grade plastic Distal end - Atraumatic left-facing bevelled tip without Murphy eye All components must be latex and DEHP free, sterile and individually packed in peel pouch that is easy to open - For SINGLE use only To comply with the latest ISO 5361 Information brochure or the outer package of the item should clearly indicate latex free and DEHP free | Uncuffed Extra-soft Plasticized PVC Oral/Nasal tracheal tube for long-term intubation. ID: 6.0mm OD: 8.5mm Length: 270mm Depth markings in 1cm intervals from 15 - 25cm
64 | RT13-17-001 | Cuffed and Uncuffed Tubes | Cuffed Oral/Nasal | Cuffed Oral/Nasal tracheal tube with a Murphy eye ID: 2.0 Age: Newborn (< 2kg) CUFFED ORAL/NASAL TRACHEAL TUBE WITH/ WITHOUT A MURPHY EYE - NEWBORN Purpose: Short term oral or nasal intubation Consists of a clear thermo-sensitive PVC tube that is kink resistant, atraumatic to airways and has a full Magill curvature. Must be radiopaque or have a radiopaque line Must be marked with numbered depth markings in 1cm intervals, tube size, length, ID, OD, trade name. The markings on tube must be clear and must not rub off when wiped with alcohol. Must have a precision band mark just above the cuff to aid with depth of intubation and guide proper placement. Preference will be given to tubes with a second circumferential band (not more than 1cm from each other) to guide proper placement. Proximal end - secure 15mm OD circuit connector made from medical grade plastic Distal end - Atraumatic left-facing bevelled tip without or 1 Murphy eye on the opposite side Must have a high-volume-low-pressure barrel-shaped polyvinyl chloride cuff with a pilot cuff. Cuff thickness ± 50 microns The pilot cuff must have a one-way valve and must indicate the degree of inflation accurately. The cuff must inflate and deflate easily, and when deflated must conform closely to tube to minimise the risk of trauma during intubation, and extubation and not increase the OD of the tube. Must be appropriately sized in relation to the tube size and appropriately placed with regards to the distance from the distal tip. The cuff must be resistant to Nitrous Oxide effects All components must be latex-free sterile and individually packed in peel pouch that is easy to open; For SINGLE use only To comply with the latest ISO 5361 Information brochure or the outer package of the item should clearly indicate latex free | Cuffed Oral/Nasal tracheal tube with a Murphy eye ID: 2.0 Age: Newborn (< 2kg)
65 | RT13-17-016 | Cuffed and Uncuffed Tubes | CUFFED SILICONIZED ENDOTRACHEAL TUBE | CUFFED SILICONIZED ENDOTRACHEAL TUBE WITH A MURPHY EYE - PAEDIATRIC Purpose: Short term oral or nasal intubation Medically safe, thermosensitive, soft, pliable and kink resistant, SILICONIZED PVC tube with a radiopaque line and a full Magill curvature. Must be radiopaque or have a radiopaque line Must be marked with numbered depth markings in 1cm intervals, tube size, length, ID, OD, trade name. The markings on tube must be clear and must not rub off when wiped with alcohol. Must have a precision band mark just above the cuff to aid with depth of intubation and guide proper placement. Preference will be given to tubes with a second circumferential band (not more than 1cm from each other) to guide proper placement. Proximal end - secure 15mm OD circuit connector made from medical grade plastic Distal end - Atraumatic left-facing bevelled tip with 1 or 2 Murphy eyes on the opposite side Must have a high-volume-low-pressure barrel-shaped polyvinyl chloride cuff with a pilot cuff. Cuff thickness ± 50 microns The pilot cuff must have a one-way valve and must indicate the degree of inflation accurately. The cuff must inflate and deflate easily, and when deflated must conform closely to tube to minimise the risk of trauma during intubation, and extubation and not increase the OD of the tube. Must be appropriately sized in relation to the tube size and appropriately placed with regards to the distance from the distal tip. The cuff must be resistant to Nitrous Oxide effects All components must be latex-free; DEHP free; sterile and individually packed in peel pouch that is easy to open; For SINGLE use only To comply with the latest ISO 5361 Information brochure or the outer package of the item should clearly indicate latex free and DEHP free | Cuffed Siliconized Oral/Nasal endotracheal tube with Murphy Eye: Size ID: 3.0mm
66 | RT13-17-017 | Cuffed and Uncuffed Tubes | CUFFED SILICONIZED ENDOTRACHEAL TUBE | CUFFED SILICONIZED ENDOTRACHEAL TUBE WITH A MURPHY EYE - PAEDIATRIC Purpose: Short term oral or nasal intubation Medically safe, thermosensitive, soft, pliable and kink resistant, SILICONIZED PVC tube with a radiopaque line and a full Magill curvature. Must be radiopaque or have a radiopaque line Must be marked with numbered depth markings in 1cm intervals, tube size, length, ID, OD, trade name. The markings on tube must be clear and must not rub off when wiped with alcohol. Must have a precision band mark just above the cuff to aid with depth of intubation and guide proper placement. Preference will be given to tubes with a second circumferential band (not more than 1cm from each other) to guide proper placement. Proximal end - secure 15mm OD circuit connector made from medical grade plastic Distal end - Atraumatic left-facing bevelled tip with 1 or 2 Murphy eyes on the opposite side Must have a high-volume-low-pressure barrel-shaped polyvinyl chloride cuff with a pilot cuff. Cuff thickness ± 50 microns The pilot cuff must have a one-way valve and must indicate the degree of inflation accurately. The cuff must inflate and deflate easily, and when deflated must conform closely to tube to minimise the risk of trauma during intubation, and extubation and not increase the OD of the tube. Must be appropriately sized in relation to the tube size and appropriately placed with regards to the distance from the distal tip. The cuff must be resistant to Nitrous Oxide effects All components must be latex-free; DEHP free; sterile and individually packed in peel pouch that is easy to open; For SINGLE use only To comply with the latest ISO 5361 Information brochure or the outer package of the item should clearly indicate latex free and DEHP free | Cuffed Siliconized Oral/Nasal endotracheal tube with Murphy Eye: Size ID: 3.5mm
67 | RT13-17-018 | Cuffed and Uncuffed Tubes | CUFFED SILICONIZED ENDOTRACHEAL TUBE | CUFFED SILICONIZED ENDOTRACHEAL TUBE WITH A MURPHY EYE - PAEDIATRIC Purpose: Short term oral or nasal intubation Medically safe, thermosensitive, soft, pliable and kink resistant, SILICONIZED PVC tube with a radiopaque line and a full Magill curvature. Must be radiopaque or have a radiopaque line Must be marked with numbered depth markings in 1cm intervals, tube size, length, ID, OD, trade name. The markings on tube must be clear and must not rub off when wiped with alcohol. Must have a precision band mark just above the cuff to aid with depth of intubation and guide proper placement. Preference will be given to tubes with a second circumferential band (not more than 1cm from each other) to guide proper placement. Proximal end - secure 15mm OD circuit connector made from medical grade plastic Distal end - Atraumatic left-facing bevelled tip with 1 or 2 Murphy eyes on the opposite side Must have a high-volume-low-pressure barrel-shaped polyvinyl chloride cuff with a pilot cuff. Cuff thickness ± 50 microns The pilot cuff must have a one-way valve and must indicate the degree of inflation accurately. The cuff must inflate and deflate easily, and when deflated must conform closely to tube to minimise the risk of trauma during intubation, and extubation and not increase the OD of the tube. Must be appropriately sized in relation to the tube size and appropriately placed with regards to the distance from the distal tip. The cuff must be resistant to Nitrous Oxide effects All components must be latex-free; DEHP free; sterile and individually packed in peel pouch that is easy to open; For SINGLE use only To comply with the latest ISO 5361 Information brochure or the outer package of the item should clearly indicate latex free and DEHP free | Cuffed Siliconized Oral/Nasal endotracheal tube with Murphy Eye: Size ID: 4.0mm
68 | RT13-17-019 | Cuffed and Uncuffed Tubes | CUFFED SILICONIZED ENDOTRACHEAL TUBE | CUFFED SILICONIZED ENDOTRACHEAL TUBE WITH A MURPHY EYE - PAEDIATRIC Purpose: Short term oral or nasal intubation Medically safe, thermosensitive, soft, pliable and kink resistant, SILICONIZED PVC tube with a radiopaque line and a full Magill curvature. Must be radiopaque or have a radiopaque line Must be marked with numbered depth markings in 1cm intervals, tube size, length, ID, OD, trade name. The markings on tube must be clear and must not rub off when wiped with alcohol. Must have a precision band mark just above the cuff to aid with depth of intubation and guide proper placement. Preference will be given to tubes with a second circumferential band (not more than 1cm from each other) to guide proper placement. Proximal end - secure 15mm OD circuit connector made from medical grade plastic Distal end - Atraumatic left-facing bevelled tip with 1 or 2 Murphy eyes on the opposite side Must have a high-volume-low-pressure barrel-shaped polyvinyl chloride cuff with a pilot cuff. Cuff thickness ± 50 microns The pilot cuff must have a one-way valve and must indicate the degree of inflation accurately. The cuff must inflate and deflate easily, and when deflated must conform closely to tube to minimise the risk of trauma during intubation, and extubation and not increase the OD of the tube. Must be appropriately sized in relation to the tube size and appropriately placed with regards to the distance from the distal tip. The cuff must be resistant to Nitrous Oxide effects All components must be latex-free; DEHP free; sterile and individually packed in peel pouch that is easy to open; For SINGLE use only To comply with the latest ISO 5361 Information brochure or the outer package of the item should clearly indicate latex free and DEHP free | Cuffed Siliconized Oral/Nasal endotracheal tube with Murphy Eye: Size ID: 4.5mm
69 | RT13-17-020 | Cuffed and Uncuffed Tubes | CUFFED SILICONIZED ENDOTRACHEAL TUBE | CUFFED SILICONIZED ENDOTRACHEAL TUBE WITH A MURPHY EYE - PAEDIATRIC Purpose: Short term oral or nasal intubation Medically safe, thermosensitive, soft, pliable and kink resistant, SILICONIZED PVC tube with a radiopaque line and a full Magill curvature. Must be radiopaque or have a radiopaque line Must be marked with numbered depth markings in 1cm intervals, tube size, length, ID, OD, trade name. The markings on tube must be clear and must not rub off when wiped with alcohol. Must have a precision band mark just above the cuff to aid with depth of intubation and guide proper placement. Preference will be given to tubes with a second circumferential band (not more than 1cm from each other) to guide proper placement. Proximal end - secure 15mm OD circuit connector made from medical grade plastic Distal end - Atraumatic left-facing bevelled tip with 1 or 2 Murphy eyes on the opposite side Must have a high-volume-low-pressure barrel-shaped polyvinyl chloride cuff with a pilot cuff. Cuff thickness ± 50 microns The pilot cuff must have a one-way valve and must indicate the degree of inflation accurately. The cuff must inflate and deflate easily, and when deflated must conform closely to tube to minimise the risk of trauma during intubation, and extubation and not increase the OD of the tube. Must be appropriately sized in relation to the tube size and appropriately placed with regards to the distance from the distal tip. The cuff must be resistant to Nitrous Oxide effects All components must be latex-free; DEHP free; sterile and individually packed in peel pouch that is easy to open; For SINGLE use only To comply with the latest ISO 5361 Information brochure or the outer package of the item should clearly indicate latex free and DEHP free | Cuffed Siliconized Oral/Nasal endotracheal tube with Murphy Eye: Size ID: 5.0mm
70 | RT13-17-021 | Cuffed and Uncuffed Tubes | CUFFED SILICONIZED ENDOTRACHEAL TUBE | CUFFED SILICONIZED ENDOTRACHEAL TUBE WITH A MURPHY EYE - PAEDIATRIC Purpose: Short term oral or nasal intubation Medically safe, thermosensitive, soft, pliable and kink resistant, SILICONIZED PVC tube with a radiopaque line and a full Magill curvature. Must be radiopaque or have a radiopaque line Must be marked with numbered depth markings in 1cm intervals, tube size, length, ID, OD, trade name. The markings on tube must be clear and must not rub off when wiped with alcohol. Must have a precision band mark just above the cuff to aid with depth of intubation and guide proper placement. Preference will be given to tubes with a second circumferential band (not more than 1cm from each other) to guide proper placement. Proximal end - secure 15mm OD circuit connector made from medical grade plastic Distal end - Atraumatic left-facing bevelled tip with 1 or 2 Murphy eyes on the opposite side Must have a high-volume-low-pressure barrel-shaped polyvinyl chloride cuff with a pilot cuff. Cuff thickness ± 50 microns The pilot cuff must have a one-way valve and must indicate the degree of inflation accurately. The cuff must inflate and deflate easily, and when deflated must conform closely to tube to minimise the risk of trauma during intubation, and extubation and not increase the OD of the tube. Must be appropriately sized in relation to the tube size and appropriately placed with regards to the distance from the distal tip. The cuff must be resistant to Nitrous Oxide effects All components must be latex-free; DEHP free; sterile and individually packed in peel pouch that is easy to open; For SINGLE use only To comply with the latest ISO 5361 Information brochure or the outer package of the item should clearly indicate latex free and DEHP free | Cuffed Siliconized Oral/Nasal endotracheal tube with Murphy Eye: Size ID: 5.5mm
71 | RT13-17-022 | Cuffed and Uncuffed Tubes | CUFFED SILICONIZED ENDOTRACHEAL TUBE | CUFFED SILICONIZED ENDOTRACHEAL TUBE WITH A MURPHY EYE - PAEDIATRIC Purpose: Short term oral or nasal intubation Medically safe, thermosensitive, soft, pliable and kink resistant, SILICONIZED PVC tube with a radiopaque line and a full Magill curvature. Must be radiopaque or have a radiopaque line Must be marked with numbered depth markings in 1cm intervals, tube size, length, ID, OD, trade name. The markings on tube must be clear and must not rub off when wiped with alcohol. Must have a precision band mark just above the cuff to aid with depth of intubation and guide proper placement. Preference will be given to tubes with a second circumferential band (not more than 1cm from each other) to guide proper placement. Proximal end - secure 15mm OD circuit connector made from medical grade plastic Distal end - Atraumatic left-facing bevelled tip with 1 or 2 Murphy eyes on the opposite side Must have a high-volume-low-pressure barrel-shaped polyvinyl chloride cuff with a pilot cuff. Cuff thickness ± 50 microns The pilot cuff must have a one-way valve and must indicate the degree of inflation accurately. The cuff must inflate and deflate easily, and when deflated must conform closely to tube to minimise the risk of trauma during intubation, and extubation and not increase the OD of the tube. Must be appropriately sized in relation to the tube size and appropriately placed with regards to the distance from the distal tip. The cuff must be resistant to Nitrous Oxide effects All components must be latex-free; DEHP free; sterile and individually packed in peel pouch that is easy to open; For SINGLE use only To comply with the latest ISO 5361 Information brochure or the outer package of the item should clearly indicate latex free and DEHP free | Cuffed Siliconized Oral/Nasal endotracheal tube with Murphy Eye: Size ID: 6.0mm
72 | RT13-17-023 | Cuffed and Uncuffed Tubes | CUFFED SILICONIZED ENDOTRACHEAL TUBE | CUFFED SILICONIZED ENDOTRACHEAL TUBE WITH A MURPHY EYE - ADULT Purpose: Short term oral or nasal intubation Medically safe, thermosensitive, soft, pliable and kink resistant, SILICONIZED PVC tube with a radiopaque line and a full Magill curvature. Must be radiopaque or have a radiopaque line Must be marked with numbered depth markings in 1cm intervals, tube size, length, ID, OD, trade name. The markings on tube must be clear and must not rub off when wiped with alcohol. Must have a precision band mark just above the cuff to aid with depth of intubation and guide proper placement. Preference will be given to tubes with a second circumferential band (not more than 2cm from each other) to guide proper placement. Proximal end - secure 15mm OD circuit connector made from medical grade plastic Distal end - Atraumatic left-facing bevelled tip with 1 or 2 Murphy eyes on the opposite side Must have a high-volume-low-pressure barrel-shaped polyvinyl chloride cuff with a pilot cuff. Cuff thickness ± 50 microns The pilot cuff must have a one-way valve and must indicate the degree of inflation accurately. The cuff must inflate and deflate easily, and when deflated must conform closely to tube to minimise the risk of trauma during intubation, and extubation and not increase the OD of the tube. Must be appropriately sized in relation to the tube size and appropriately placed with regards to the distance from the distal tip. The cuff must be resistant to Nitrous Oxide effects All components must be latex-free; DEHP free; sterile and individually packed in peel pouch that is easy to open; For SINGLE use only To comply with the latest ISO 5361 Information brochure or the outer package of the item should clearly indicate latex free and DEHP free | Cuffed Siliconized Oral/Nasal endotracheal tube with a Murphy Eye Size ID: 6.5mm
73 | RT13-17-024 | Cuffed and Uncuffed Tubes | CUFFED SILICONIZED ENDOTRACHEAL TUBE | CUFFED SILICONIZED ENDOTRACHEAL TUBE WITH A MURPHY EYE - ADULT Purpose: Short term oral or nasal intubation Medically safe, thermosensitive, soft, pliable and kink resistant, SILICONIZED PVC tube with a radiopaque line and a full Magill curvature. Must be radiopaque or have a radiopaque line Must be marked with numbered depth markings in 1cm intervals, tube size, length, ID, OD, trade name. The markings on tube must be clear and must not rub off when wiped with alcohol. Must have a precision band mark just above the cuff to aid with depth of intubation and guide proper placement. Preference will be given to tubes with a second circumferential band (not more than 2cm from each other) to guide proper placement. Proximal end - secure 15mm OD circuit connector made from medical grade plastic Distal end - Atraumatic left-facing bevelled tip with 1 or 2 Murphy eyes on the opposite side Must have a high-volume-low-pressure barrel-shaped polyvinyl chloride cuff with a pilot cuff. Cuff thickness ± 50 microns The pilot cuff must have a one-way valve and must indicate the degree of inflation accurately. The cuff must inflate and deflate easily, and when deflated must conform closely to tube to minimise the risk of trauma during intubation, and extubation and not increase the OD of the tube. Must be appropriately sized in relation to the tube size and appropriately placed with regards to the distance from the distal tip. The cuff must be resistant to Nitrous Oxide effects All components must be latex-free; DEHP free; sterile and individually packed in peel pouch that is easy to open; For SINGLE use only To comply with the latest ISO 5361 Information brochure or the outer package of the item should clearly indicate latex free and DEHP free | Cuffed Siliconized Oral/Nasal endotracheal tube with a Murphy Eye Size ID: 7.0mm
74 | RT13-17-025 | Cuffed and Uncuffed Tubes | CUFFED SILICONIZED ENDOTRACHEAL TUBE | CUFFED SILICONIZED ENDOTRACHEAL TUBE WITH A MURPHY EYE - ADULT Purpose: Short term oral or nasal intubation Medically safe, thermosensitive, soft, pliable and kink resistant, SILICONIZED PVC tube with a radiopaque line and a full Magill curvature. Must be radiopaque or have a radiopaque line Must be marked with numbered depth markings in 1cm intervals, tube size, length, ID, OD, trade name. The markings on tube must be clear and must not rub off when wiped with alcohol. Must have a precision band mark just above the cuff to aid with depth of intubation and guide proper placement. Preference will be given to tubes with a second circumferential band (not more than 2cm from each other) to guide proper placement. Proximal end - secure 15mm OD circuit connector made from medical grade plastic Distal end - Atraumatic left-facing bevelled tip with 1 or 2 Murphy eyes on the opposite side Must have a high-volume-low-pressure barrel-shaped polyvinyl chloride cuff with a pilot cuff. Cuff thickness ± 50 microns The pilot cuff must have a one-way valve and must indicate the degree of inflation accurately. The cuff must inflate and deflate easily, and when deflated must conform closely to tube to minimise the risk of trauma during intubation, and extubation and not increase the OD of the tube. Must be appropriately sized in relation to the tube size and appropriately placed with regards to the distance from the distal tip. The cuff must be resistant to Nitrous Oxide effects All components must be latex-free; DEHP free; sterile and individually packed in peel pouch that is easy to open; For SINGLE use only To comply with the latest ISO 5361 Information brochure or the outer package of the item should clearly indicate latex free and DEHP free | Cuffed Siliconized Oral/Nasal endotracheal tube with a Murphy Eye Size ID: 7.5mm
75 | RT13-17-026 | Cuffed and Uncuffed Tubes | CUFFED SILICONIZED ENDOTRACHEAL TUBE | CUFFED SILICONIZED ENDOTRACHEAL TUBE WITH A MURPHY EYE - ADULT Purpose: Short term oral or nasal intubation Medically safe, thermosensitive, soft, pliable and kink resistant, SILICONIZED PVC tube with a radiopaque line and a full Magill curvature. Must be radiopaque or have a radiopaque line Must be marked with numbered depth markings in 1cm intervals, tube size, length, ID, OD, trade name. The markings on tube must be clear and must not rub off when wiped with alcohol. Must have a precision band mark just above the cuff to aid with depth of intubation and guide proper placement. Preference will be given to tubes with a second circumferential band (not more than 2cm from each other) to guide proper placement. Proximal end - secure 15mm OD circuit connector made from medical grade plastic Distal end - Atraumatic left-facing bevelled tip with 1 or 2 Murphy eyes on the opposite side Must have a high-volume-low-pressure barrel-shaped polyvinyl chloride cuff with a pilot cuff. Cuff thickness ± 50 microns The pilot cuff must have a one-way valve and must indicate the degree of inflation accurately. The cuff must inflate and deflate easily, and when deflated must conform closely to tube to minimise the risk of trauma during intubation, and extubation and not increase the OD of the tube. Must be appropriately sized in relation to the tube size and appropriately placed with regards to the distance from the distal tip. The cuff must be resistant to Nitrous Oxide effects All components must be latex-free; DEHP free; sterile and individually packed in peel pouch that is easy to open; For SINGLE use only To comply with the latest ISO 5361 Information brochure or the outer package of the item should clearly indicate latex free and DEHP free | Cuffed Siliconized Oral/Nasal endotracheal tube with a Murphy Eye Size ID: 8.0mm
76 | RT13-17-027 | Cuffed and Uncuffed Tubes | CUFFED SILICONIZED ENDOTRACHEAL TUBE | CUFFED SILICONIZED ENDOTRACHEAL TUBE WITH A MURPHY EYE - ADULT Purpose: Short term oral or nasal intubation Medically safe, thermosensitive, soft, pliable and kink resistant, SILICONIZED PVC tube with a radiopaque line and a full Magill curvature. Must be radiopaque or have a radiopaque line Must be marked with numbered depth markings in 1cm intervals, tube size, length, ID, OD, trade name. The markings on tube must be clear and must not rub off when wiped with alcohol. Must have a precision band mark just above the cuff to aid with depth of intubation and guide proper placement. Preference will be given to tubes with a second circumferential band (not more than 2cm from each other) to guide proper placement. Proximal end - secure 15mm OD circuit connector made from medical grade plastic Distal end - Atraumatic left-facing bevelled tip with 1 or 2 Murphy eyes on the opposite side Must have a high-volume-low-pressure barrel-shaped polyvinyl chloride cuff with a pilot cuff. Cuff thickness ± 50 microns The pilot cuff must have a one-way valve and must indicate the degree of inflation accurately. The cuff must inflate and deflate easily, and when deflated must conform closely to tube to minimise the risk of trauma during intubation, and extubation and not increase the OD of the tube. Must be appropriately sized in relation to the tube size and appropriately placed with regards to the distance from the distal tip. The cuff must be resistant to Nitrous Oxide effects All components must be latex-free; DEHP free; sterile and individually packed in peel pouch that is easy to open; For SINGLE use only To comply with the latest ISO 5361 Information brochure or the outer package of the item should clearly indicate latex free and DEHP free | Cuffed Siliconized Oral/Nasal endotracheal tube with a Murphy Eye Size ID: 8.5mm
77 | RT13-17-028 | Cuffed and Uncuffed Tubes | CUFFED SILICONIZED ENDOTRACHEAL TUBE | CUFFED SILICONIZED ENDOTRACHEAL TUBE WITH A MURPHY EYE - ADULT Purpose: Short term oral or nasal intubation Medically safe, thermosensitive, soft, pliable and kink resistant, SILICONIZED PVC tube with a radiopaque line and a full Magill curvature. Must be radiopaque or have a radiopaque line Must be marked with numbered depth markings in 1cm intervals, tube size, length, ID, OD, trade name. The markings on tube must be clear and must not rub off when wiped with alcohol. Must have a precision band mark just above the cuff to aid with depth of intubation and guide proper placement. Preference will be given to tubes with a second circumferential band (not more than 2cm from each other) to guide proper placement. Proximal end - secure 15mm OD circuit connector made from medical grade plastic Distal end - Atraumatic left-facing bevelled tip with 1 or 2 Murphy eyes on the opposite side Must have a high-volume-low-pressure barrel-shaped polyvinyl chloride cuff with a pilot cuff. Cuff thickness ± 50 microns The pilot cuff must have a one-way valve and must indicate the degree of inflation accurately. The cuff must inflate and deflate easily, and when deflated must conform closely to tube to minimise the risk of trauma during intubation, and extubation and not increase the OD of the tube. Must be appropriately sized in relation to the tube size and appropriately placed with regards to the distance from the distal tip. The cuff must be resistant to Nitrous Oxide effects All components must be latex-free; DEHP free; sterile and individually packed in peel pouch that is easy to open; For SINGLE use only To comply with the latest ISO 5361 Information brochure or the outer package of the item should clearly indicate latex free and DEHP free | Cuffed Siliconized Oral/Nasal endotracheal tube with a Murphy Eye Size ID: 9.0mm
78 | RT13-19-001 | Cuffed and Uncuffed Tubes | UNCUFFED ORAL/NASAL TRACHEAL TUBE | UNCUFFED ORAL/NASAL TRACHEAL TUBE WITH REDUCED TRAUMA TIP Purpose: Short term oral or nasal intubation with tip designed to reduce airway trauma Consists of a clear thermo-sensitive PVC tube that is kink resistant, atraumatic to airways and has a full Magill curvature and two Murphy eyes. Must be radiopaque or have a radiopaque line Must be marked with numbered depth markings in 1cm intervals, tube size, length, ID, OD, trade name. The markings on tube must be clear and must not rub off when wiped with alcohol. Must have a precision band mark to aid with depth of intubation and guide proper placement. Solid or circumferential precision band markings accepted Preference will be given to tubes with a second circumferential band (not more than 1cm from each other) to guide proper placement. Proximal end - secure 15mm OD circuit connector made from medical grade plastic Distal end - Tapered, soft flexible pointed tip with an upturned leading edge , configured with the tip point in the midline so as to pass easily between the vocal cords without touching them. The tube must be able to demonstrate a reduction in intubation induced airway trauma. All components must be latex-free sterile and individually packed in peel pouch that is easy to open - For SINGLE use only To comply with the latest ISO 5361 Information brochure or the outer package of the item should clearly indicate latex free | Uncuffed Oral/Nasal tracheal tube with Reduced Trauma Tip: Size 2.5
79 | RT13-19-002 | Cuffed and Uncuffed Tubes | UNCUFFED ORAL/NASAL TRACHEAL TUBE | UNCUFFED ORAL/NASAL TRACHEAL TUBE WITH REDUCED TRAUMA TIP Purpose: Short term oral or nasal intubation with tip designed to reduce airway trauma Consists of a clear thermo-sensitive PVC tube that is kink resistant, atraumatic to airways and has a full Magill curvature and two Murphy eyes. Must be radiopaque or have a radiopaque line Must be marked with numbered depth markings in 1cm intervals, tube size, length, ID, OD, trade name. The markings on tube must be clear and must not rub off when wiped with alcohol. Must have a precision band mark to aid with depth of intubation and guide proper placement. Solid or circumferential precision band markings accepted Preference will be given to tubes with a second circumferential band (not more than 1cm from each other) to guide proper placement. Proximal end - secure 15mm OD circuit connector made from medical grade plastic Distal end - Tapered, soft flexible pointed tip with an upturned leading edge , configured with the tip point in the midline so as to pass easily between the vocal cords without touching them. The tube must be able to demonstrate a reduction in intubation induced airway trauma. All components must be latex-free sterile and individually packed in peel pouch that is easy to open - For SINGLE use only To comply with the latest ISO 5361 Information brochure or the outer package of the item should clearly indicate latex free | Uncuffed Oral/Nasal tracheal tube with Reduced Trauma Tip: Size 3.0
80 | RT13-19-003 | Cuffed and Uncuffed Tubes | UNCUFFED ORAL/NASAL TRACHEAL TUBE | UNCUFFED ORAL/NASAL TRACHEAL TUBE WITH REDUCED TRAUMA TIP Purpose: Short term oral or nasal intubation with tip designed to reduce airway trauma Consists of a clear thermo-sensitive PVC tube that is kink resistant, atraumatic to airways and has a full Magill curvature and two Murphy eyes. Must be radiopaque or have a radiopaque line Must be marked with numbered depth markings in 1cm intervals, tube size, length, ID, OD, trade name. The markings on tube must be clear and must not rub off when wiped with alcohol. Must have a precision band mark to aid with depth of intubation and guide proper placement. Solid or circumferential precision band markings accepted Preference will be given to tubes with a second circumferential band (not more than 1cm from each other) to guide proper placement. Proximal end - secure 15mm OD circuit connector made from medical grade plastic Distal end - Tapered, soft flexible pointed tip with an upturned leading edge , configured with the tip point in the midline so as to pass easily between the vocal cords without touching them. The tube must be able to demonstrate a reduction in intubation induced airway trauma. All components must be latex-free sterile and individually packed in peel pouch that is easy to open - For SINGLE use only To comply with the latest ISO 5361 Information brochure or the outer package of the item should clearly indicate latex free | Uncuffed Oral/Nasal tracheal tube with Reduced Trauma Tip: Size 3.5
81 | RT13-19-004 | Cuffed and Uncuffed Tubes | UNCUFFED ORAL/NASAL TRACHEAL TUBE | UNCUFFED ORAL/NASAL TRACHEAL TUBE WITH REDUCED TRAUMA TIP Purpose: Short term oral or nasal intubation with tip designed to reduce airway trauma Consists of a clear thermo-sensitive PVC tube that is kink resistant, atraumatic to airways and has a full Magill curvature and two Murphy eyes. Must be radiopaque or have a radiopaque line Must be marked with numbered depth markings in 1cm intervals, tube size, length, ID, OD, trade name. The markings on tube must be clear and must not rub off when wiped with alcohol. Must have a precision band mark to aid with depth of intubation and guide proper placement. Solid or circumferential precision band markings accepted Preference will be given to tubes with a second circumferential band (not more than 1cm from each other) to guide proper placement. Proximal end - secure 15mm OD circuit connector made from medical grade plastic Distal end - Tapered, soft flexible pointed tip with an upturned leading edge , configured with the tip point in the midline so as to pass easily between the vocal cords without touching them. The tube must be able to demonstrate a reduction in intubation induced airway trauma. All components must be latex-free sterile and individually packed in peel pouch that is easy to open - For SINGLE use only To comply with the latest ISO 5361 Information brochure or the outer package of the item should clearly indicate latex free | Uncuffed Oral/Nasal tracheal tube with Reduced Trauma Tip: Size 4.0
82 | RT13-19-005 | Cuffed and Uncuffed Tubes | UNCUFFED ORAL/NASAL TRACHEAL TUBE | UNCUFFED ORAL/NASAL TRACHEAL TUBE WITH REDUCED TRAUMA TIP Purpose: Short term oral or nasal intubation with tip designed to reduce airway trauma Consists of a clear thermo-sensitive PVC tube that is kink resistant, atraumatic to airways and has a full Magill curvature and two Murphy eyes. Must be radiopaque or have a radiopaque line Must be marked with numbered depth markings in 1cm intervals, tube size, length, ID, OD, trade name. The markings on tube must be clear and must not rub off when wiped with alcohol. Must have a precision band mark to aid with depth of intubation and guide proper placement. Solid or circumferential precision band markings accepted Preference will be given to tubes with a second circumferential band (not more than 1cm from each other) to guide proper placement. Proximal end - secure 15mm OD circuit connector made from medical grade plastic Distal end - Tapered, soft flexible pointed tip with an upturned leading edge , configured with the tip point in the midline so as to pass easily between the vocal cords without touching them. The tube must be able to demonstrate a reduction in intubation induced airway trauma. All components must be latex-free sterile and individually packed in peel pouch that is easy to open - For SINGLE use only To comply with the latest ISO 5361 Information brochure or the outer package of the item should clearly indicate latex free | Uncuffed Oral/Nasal tracheal tube with Reduced Trauma Tip: Size 4.5
83 | RT13-19-006 | Cuffed and Uncuffed Tubes | UNCUFFED ORAL/NASAL TRACHEAL TUBE | UNCUFFED ORAL/NASAL TRACHEAL TUBE WITH REDUCED TRAUMA TIP Purpose: Short term oral or nasal intubation with tip designed to reduce airway trauma Consists of a clear thermo-sensitive PVC tube that is kink resistant, atraumatic to airways and has a full Magill curvature and two Murphy eyes. Must be radiopaque or have a radiopaque line Must be marked with numbered depth markings in 1cm intervals, tube size, length, ID, OD, trade name. The markings on tube must be clear and must not rub off when wiped with alcohol. Must have a precision band mark to aid with depth of intubation and guide proper placement. Solid or circumferential precision band markings accepted Preference will be given to tubes with a second circumferential band (not more than 1cm from each other) to guide proper placement. Proximal end - secure 15mm OD circuit connector made from medical grade plastic Distal end - Tapered, soft flexible pointed tip with an upturned leading edge , configured with the tip point in the midline so as to pass easily between the vocal cords without touching them. The tube must be able to demonstrate a reduction in intubation induced airway trauma. All components must be latex-free sterile and individually packed in peel pouch that is easy to open - For SINGLE use only To comply with the latest ISO 5361 Information brochure or the outer package of the item should clearly indicate latex free | Uncuffed Oral/Nasal tracheal tube with Reduced Trauma Tip: Size 5.0
84 | RT13-19-007 | Cuffed and Uncuffed Tubes | UNCUFFED ORAL/NASAL TRACHEAL TUBE | UNCUFFED ORAL/NASAL TRACHEAL TUBE WITH REDUCED TRAUMA TIP Purpose: Short term oral or nasal intubation with tip designed to reduce airway trauma Consists of a clear thermo-sensitive PVC tube that is kink resistant, atraumatic to airways and has a full Magill curvature and two Murphy eyes. Must be radiopaque or have a radiopaque line Must be marked with numbered depth markings in 1cm intervals, tube size, length, ID, OD, trade name. The markings on tube must be clear and must not rub off when wiped with alcohol. Must have a precision band mark to aid with depth of intubation and guide proper placement. Solid or circumferential precision band markings accepted Preference will be given to tubes with a second circumferential band (not more than 1cm from each other) to guide proper placement. Proximal end - secure 15mm OD circuit connector made from medical grade plastic Distal end - Tapered, soft flexible pointed tip with an upturned leading edge , configured with the tip point in the midline so as to pass easily between the vocal cords without touching them. The tube must be able to demonstrate a reduction in intubation induced airway trauma. All components must be latex-free sterile and individually packed in peel pouch that is easy to open - For SINGLE use only To comply with the latest ISO 5361 Information brochure or the outer package of the item should clearly indicate latex free | Uncuffed Oral/Nasal tracheal tube with Reduced Trauma Tip: Size 5.5
85 | RT13-19-008 | Cuffed and Uncuffed Tubes | UNCUFFED ORAL/NASAL TRACHEAL TUBE | UNCUFFED ORAL/NASAL TRACHEAL TUBE WITH REDUCED TRAUMA TIP Purpose: Short term oral or nasal intubation with tip designed to reduce airway trauma Consists of a clear thermo-sensitive PVC tube that is kink resistant, atraumatic to airways and has a full Magill curvature and two Murphy eyes. Must be radiopaque or have a radiopaque line Must be marked with numbered depth markings in 1cm intervals, tube size, length, ID, OD, trade name. The markings on tube must be clear and must not rub off when wiped with alcohol. Must have a precision band mark to aid with depth of intubation and guide proper placement. Solid or circumferential precision band markings accepted Preference will be given to tubes with a second circumferential band (not more than 1cm from each other) to guide proper placement. Proximal end - secure 15mm OD circuit connector made from medical grade plastic Distal end - Tapered, soft flexible pointed tip with an upturned leading edge , configured with the tip point in the midline so as to pass easily between the vocal cords without touching them. The tube must be able to demonstrate a reduction in intubation induced airway trauma. All components must be latex-free sterile and individually packed in peel pouch that is easy to open - For SINGLE use only To comply with the latest ISO 5361 Information brochure or the outer package of the item should clearly indicate latex free | Uncuffed Oral/Nasal tracheal tube with Reduced Trauma Tip: Size 6.0
86 | RT13-20-001 | Cuffed and Uncuffed Tubes | Cuffed Oral/Nasal | CUFFED ORAL/NASAL ENDOTRACHEAL TUBE WITH REDUCED TRAUMA TIP - PAEDIATRIC Purpose: Short term oral or nasal intubation with tip designed to reduce airway trauma Consists of a clear thermo-sensitive PVC tube that is kink resistant, atraumatic to airways and has a full Magill curvature and two Murphy eyes Must be radiopaque or have a radiopaque line Must be marked with numbered depth markings in 1cm intervals, tube size, length, ID, OD, trade name. The markings on tube must be clear and must not rub off when wiped with alcohol. Must have a precision band mark just above the cuff to aid with depth of intubation and guide proper placement. Preference will be given to tubes with a second circumferential band (not more than 1cm from each other) to guide proper placement. Proximal end - secure 15mm OD circuit connector made from medical grade plastic Distal end - Tapered, soft flexible pointed tip with an upturned leading edge , configured with the tip point in the midline so as to pass easily between the vocal cords without touching them. The tube must be able to demonstrate a reduction in intubation induced airway trauma. Must have a high-volume-low-pressure polyurethane barrel shaped cuff with pilot cuff. Cuff thickness ± 50 microns a shore hardness of no more than 86 microns The pilot cuff must have a one-way valve and must indicate the degree of inflation accurately. The cuff must inflate and deflate easily and when deflated must conform closely to tube to minimise the risk of trauma during intubation and extubation and not increase the OD of the tube. Cuff must be barrel shaped to facilitate optimal pressure dispersion against the tracheal wall. The cuff must be resistant to Nitrous Oxide effects All components must be latex-free sterile and individually packed in peel pouch that is easy to open - For SINGLE use only To comply with the latest ISO 5361 Information brochure or the outer package of the item should clearly indicate latex free | Cuffed Oral/Nasal endotracheal tube with Reduced Trauma Tip: ID: 3.0mm
87 | RT13-20-002 | Cuffed and Uncuffed Tubes | Cuffed Oral/Nasal | CUFFED ORAL/NASAL ENDOTRACHEAL TUBE WITH REDUCED TRAUMA TIP - PAEDIATRIC Purpose: Short term oral or nasal intubation with tip designed to reduce airway trauma Consists of a clear thermo-sensitive PVC tube that is kink resistant, atraumatic to airways and has a full Magill curvature and two Murphy eyes Must be radiopaque or have a radiopaque line Must be marked with numbered depth markings in 1cm intervals, tube size, length, ID, OD, trade name. The markings on tube must be clear and must not rub off when wiped with alcohol. Must have a precision band mark just above the cuff to aid with depth of intubation and guide proper placement. Preference will be given to tubes with a second circumferential band (not more than 1cm from each other) to guide proper placement. Proximal end - secure 15mm OD circuit connector made from medical grade plastic Distal end - Tapered, soft flexible pointed tip with an upturned leading edge , configured with the tip point in the midline so as to pass easily between the vocal cords without touching them. The tube must be able to demonstrate a reduction in intubation induced airway trauma. Must have a high-volume-low-pressure polyurethane barrel shaped cuff with pilot cuff. Cuff thickness ± 50 microns a shore hardness of no more than 86 microns The pilot cuff must have a one-way valve and must indicate the degree of inflation accurately. The cuff must inflate and deflate easily and when deflated must conform closely to tube to minimise the risk of trauma during intubation and extubation and not increase the OD of the tube. Cuff must be barrel shaped to facilitate optimal pressure dispersion against the tracheal wall. The cuff must be resistant to Nitrous Oxide effects All components must be latex-free sterile and individually packed in peel pouch that is easy to open - For SINGLE use only To comply with the latest ISO 5361 Information brochure or the outer package of the item should clearly indicate latex free | Cuffed Oral/Nasal endotracheal tube with Reduced Trauma Tip: ID: 3.5mm
88 | RT13-20-003 | Cuffed and Uncuffed Tubes | Cuffed Oral/Nasal | CUFFED ORAL/NASAL ENDOTRACHEAL TUBE WITH REDUCED TRAUMA TIP - PAEDIATRIC Purpose: Short term oral or nasal intubation with tip designed to reduce airway trauma Consists of a clear thermo-sensitive PVC tube that is kink resistant, atraumatic to airways and has a full Magill curvature and two Murphy eyes Must be radiopaque or have a radiopaque line Must be marked with numbered depth markings in 1cm intervals, tube size, length, ID, OD, trade name. The markings on tube must be clear and must not rub off when wiped with alcohol. Must have a precision band mark just above the cuff to aid with depth of intubation and guide proper placement. Preference will be given to tubes with a second circumferential band (not more than 1cm from each other) to guide proper placement. Proximal end - secure 15mm OD circuit connector made from medical grade plastic Distal end - Tapered, soft flexible pointed tip with an upturned leading edge , configured with the tip point in the midline so as to pass easily between the vocal cords without touching them. The tube must be able to demonstrate a reduction in intubation induced airway trauma. Must have a high-volume-low-pressure polyurethane barrel shaped cuff with pilot cuff. Cuff thickness ± 50 microns a shore hardness of no more than 86 microns The pilot cuff must have a one-way valve and must indicate the degree of inflation accurately. The cuff must inflate and deflate easily and when deflated must conform closely to tube to minimise the risk of trauma during intubation and extubation and not increase the OD of the tube. Cuff must be barrel shaped to facilitate optimal pressure dispersion against the tracheal wall. The cuff must be resistant to Nitrous Oxide effects All components must be latex-free sterile and individually packed in peel pouch that is easy to open - For SINGLE use only To comply with the latest ISO 5361 Information brochure or the outer package of the item should clearly indicate latex free | Cuffed Oral/Nasal endotracheal tube with Reduced Trauma Tip: ID: 4.0mm
89 | RT13-20-004 | Cuffed and Uncuffed Tubes | Cuffed Oral/Nasal | CUFFED ORAL/NASAL ENDOTRACHEAL TUBE WITH REDUCED TRAUMA TIP - PAEDIATRIC Purpose: Short term oral or nasal intubation with tip designed to reduce airway trauma Consists of a clear thermo-sensitive PVC tube that is kink resistant, atraumatic to airways and has a full Magill curvature and two Murphy eyes Must be radiopaque or have a radiopaque line Must be marked with numbered depth markings in 1cm intervals, tube size, length, ID, OD, trade name. The markings on tube must be clear and must not rub off when wiped with alcohol. Must have a precision band mark just above the cuff to aid with depth of intubation and guide proper placement. Preference will be given to tubes with a second circumferential band (not more than 1cm from each other) to guide proper placement. Proximal end - secure 15mm OD circuit connector made from medical grade plastic Distal end - Tapered, soft flexible pointed tip with an upturned leading edge , configured with the tip point in the midline so as to pass easily between the vocal cords without touching them. The tube must be able to demonstrate a reduction in intubation induced airway trauma. Must have a high-volume-low-pressure polyurethane barrel shaped cuff with pilot cuff. Cuff thickness ± 50 microns a shore hardness of no more than 86 microns The pilot cuff must have a one-way valve and must indicate the degree of inflation accurately. The cuff must inflate and deflate easily and when deflated must conform closely to tube to minimise the risk of trauma during intubation and extubation and not increase the OD of the tube. Cuff must be barrel shaped to facilitate optimal pressure dispersion against the tracheal wall. The cuff must be resistant to Nitrous Oxide effects All components must be latex-free sterile and individually packed in peel pouch that is easy to open - For SINGLE use only To comply with the latest ISO 5361 Information brochure or the outer package of the item should clearly indicate latex free | Cuffed Oral/Nasal endotracheal tube with Reduced Trauma Tip: ID: 4.5mm
90 | RT13-20-005 | Cuffed and Uncuffed Tubes | Cuffed Oral/Nasal | CUFFED ORAL/NASAL ENDOTRACHEAL TUBE WITH REDUCED TRAUMA TIP - PAEDIATRIC Purpose: Short term oral or nasal intubation with tip designed to reduce airway trauma Consists of a clear thermo-sensitive PVC tube that is kink resistant, atraumatic to airways and has a full Magill curvature and two Murphy eyes Must be radiopaque or have a radiopaque line Must be marked with numbered depth markings in 1cm intervals, tube size, length, ID, OD, trade name. The markings on tube must be clear and must not rub off when wiped with alcohol. Must have a precision band mark just above the cuff to aid with depth of intubation and guide proper placement. Preference will be given to tubes with a second circumferential band (not more than 1cm from each other) to guide proper placement. Proximal end - secure 15mm OD circuit connector made from medical grade plastic Distal end - Tapered, soft flexible pointed tip with an upturned leading edge , configured with the tip point in the midline so as to pass easily between the vocal cords without touching them. The tube must be able to demonstrate a reduction in intubation induced airway trauma. Must have a high-volume-low-pressure polyurethane barrel shaped cuff with pilot cuff. Cuff thickness ± 50 microns a shore hardness of no more than 86 microns The pilot cuff must have a one-way valve and must indicate the degree of inflation accurately. The cuff must inflate and deflate easily and when deflated must conform closely to tube to minimise the risk of trauma during intubation and extubation and not increase the OD of the tube. Cuff must be barrel shaped to facilitate optimal pressure dispersion against the tracheal wall. The cuff must be resistant to Nitrous Oxide effects All components must be latex-free sterile and individually packed in peel pouch that is easy to open - For SINGLE use only To comply with the latest ISO 5361 Information brochure or the outer package of the item should clearly indicate latex free | Cuffed Oral/Nasal endotracheal tube with Reduced Trauma Tip: ID: 5.0mm
91 | RT13-20-006 | Cuffed and Uncuffed Tubes | Cuffed Oral/Nasal | CUFFED ORAL/NASAL ENDOTRACHEAL TUBE WITH REDUCED TRAUMA TIP - PAEDIATRIC Purpose: Short term oral or nasal intubation with tip designed to reduce airway trauma Consists of a clear thermo-sensitive PVC tube that is kink resistant, atraumatic to airways and has a full Magill curvature and two Murphy eyes Must be radiopaque or have a radiopaque line Must be marked with numbered depth markings in 1cm intervals, tube size, length, ID, OD, trade name. The markings on tube must be clear and must not rub off when wiped with alcohol. Must have a precision band mark just above the cuff to aid with depth of intubation and guide proper placement. Preference will be given to tubes with a second circumferential band (not more than 1cm from each other) to guide proper placement. Proximal end - secure 15mm OD circuit connector made from medical grade plastic Distal end - Tapered, soft flexible pointed tip with an upturned leading edge , configured with the tip point in the midline so as to pass easily between the vocal cords without touching them. The tube must be able to demonstrate a reduction in intubation induced airway trauma. Must have a high-volume-low-pressure polyurethane barrel shaped cuff with pilot cuff. Cuff thickness ± 50 microns a shore hardness of no more than 86 microns The pilot cuff must have a one-way valve and must indicate the degree of inflation accurately. The cuff must inflate and deflate easily and when deflated must conform closely to tube to minimise the risk of trauma during intubation and extubation and not increase the OD of the tube. Cuff must be barrel shaped to facilitate optimal pressure dispersion against the tracheal wall. The cuff must be resistant to Nitrous Oxide effects All components must be latex-free sterile and individually packed in peel pouch that is easy to open - For SINGLE use only To comply with the latest ISO 5361 Information brochure or the outer package of the item should clearly indicate latex free | Cuffed Oral/Nasal endotracheal tube with Reduced Trauma Tip: ID: 5.5mm
92 | RT13-20-007 | Cuffed and Uncuffed Tubes | Cuffed Oral/Nasal | CUFFED ORAL/NASAL ENDOTRACHEAL TUBE WITH REDUCED TRAUMA TIP - PAEDIATRIC Purpose: Short term oral or nasal intubation with tip designed to reduce airway trauma Consists of a clear thermo-sensitive PVC tube that is kink resistant, atraumatic to airways and has a full Magill curvature and two Murphy eyes Must be radiopaque or have a radiopaque line Must be marked with numbered depth markings in 1cm intervals, tube size, length, ID, OD, trade name. The markings on tube must be clear and must not rub off when wiped with alcohol. Must have a precision band mark just above the cuff to aid with depth of intubation and guide proper placement. Preference will be given to tubes with a second circumferential band (not more than 1cm from each other) to guide proper placement. Proximal end - secure 15mm OD circuit connector made from medical grade plastic Distal end - Tapered, soft flexible pointed tip with an upturned leading edge , configured with the tip point in the midline so as to pass easily between the vocal cords without touching them. The tube must be able to demonstrate a reduction in intubation induced airway trauma. Must have a high-volume-low-pressure polyurethane barrel shaped cuff with pilot cuff. Cuff thickness ± 50 microns a shore hardness of no more than 86 microns The pilot cuff must have a one-way valve and must indicate the degree of inflation accurately. The cuff must inflate and deflate easily and when deflated must conform closely to tube to minimise the risk of trauma during intubation and extubation and not increase the OD of the tube. Cuff must be barrel shaped to facilitate optimal pressure dispersion against the tracheal wall. The cuff must be resistant to Nitrous Oxide effects All components must be latex-free sterile and individually packed in peel pouch that is easy to open - For SINGLE use only To comply with the latest ISO 5361 Information brochure or the outer package of the item should clearly indicate latex free | Cuffed Oral/Nasal endotracheal tube with Reduced Trauma Tip: ID: 6.0mm
93 | RT13-20-008 | Cuffed and Uncuffed Tubes | CUFFED ORAL/NASAL ENDOTRACHEAL TUBE | CUFFED ORAL/NASAL ENDOTRACHEAL TUBE WITH REDUCED TRAUMA TIP - ADULT Purpose: Short term oral or nasal intubation with tip designed to reduce airway trauma Consists of a clear thermo-sensitive PVC tube that is kink resistant, atraumatic to airways and has a full Magill curvature and two Murphy eyes Must be radiopaque or have a radiopaque line Must be marked with numbered depth markings in 1cm intervals, tube size, length, ID, OD, trade name. The markings on tube must be clear and must not rub off when wiped with alcohol. Must have a precision band mark just above the cuff to aid with depth of intubation and guide proper placement. Preference will be given to tubes with a second circumferential band (not more than 2cm from each other) to guide proper placement. Proximal end - secure 15mm OD circuit connector made from medical grade plastic Distal end - Tapered, soft flexible pointed tip with an upturned leading edge , configured with the tip point in the midline so as to pass easily between the vocal cords without touching them. The tube must be able to demonstrate a reduction in intubation induced airway trauma. Must have a high-volume-low-pressure polyurethane barrel shaped cuff with pilot cuff. Cuff thickness ± 50 microns a shore hardness of no more than 86 microns The pilot cuff must have a one-way valve and must indicate the degree of inflation accurately. The cuff must inflate and deflate easily and when deflated must conform closely to tube to minimise the risk of trauma during intubation and extubation and not increase the OD of the tube. Cuff must be barrel shaped to facilitate optimal pressure dispersion against the tracheal wall. The cuff must be resistant to Nitrous Oxide effects All components must be latex-free sterile and individually packed in peel pouch that is easy to open - For SINGLE use only To comply with the latest ISO 5361 Information brochure or the outer package of the item should clearly indicate latex free | Cuffed Oral/Nasal endotracheal tube with Reduced Trauma Tip: ID: 6.5mm
94 | RT13-20-009 | Cuffed and Uncuffed Tubes | CUFFED ORAL/NASAL ENDOTRACHEAL TUBE | CUFFED ORAL/NASAL ENDOTRACHEAL TUBE WITH REDUCED TRAUMA TIP - ADULT Purpose: Short term oral or nasal intubation with tip designed to reduce airway trauma Consists of a clear thermo-sensitive PVC tube that is kink resistant, atraumatic to airways and has a full Magill curvature and two Murphy eyes Must be radiopaque or have a radiopaque line Must be marked with numbered depth markings in 1cm intervals, tube size, length, ID, OD, trade name. The markings on tube must be clear and must not rub off when wiped with alcohol. Must have a precision band mark just above the cuff to aid with depth of intubation and guide proper placement. Preference will be given to tubes with a second circumferential band (not more than 2cm from each other) to guide proper placement. Proximal end - secure 15mm OD circuit connector made from medical grade plastic Distal end - Tapered, soft flexible pointed tip with an upturned leading edge , configured with the tip point in the midline so as to pass easily between the vocal cords without touching them. The tube must be able to demonstrate a reduction in intubation induced airway trauma. Must have a high-volume-low-pressure polyurethane barrel shaped cuff with pilot cuff. Cuff thickness ± 50 microns a shore hardness of no more than 86 microns The pilot cuff must have a one-way valve and must indicate the degree of inflation accurately. The cuff must inflate and deflate easily and when deflated must conform closely to tube to minimise the risk of trauma during intubation and extubation and not increase the OD of the tube. Cuff must be barrel shaped to facilitate optimal pressure dispersion against the tracheal wall. The cuff must be resistant to Nitrous Oxide effects All components must be latex-free sterile and individually packed in peel pouch that is easy to open - For SINGLE use only To comply with the latest ISO 5361 Information brochure or the outer package of the item should clearly indicate latex free | Cuffed Oral/Nasal endotracheal tube with Reduced Trauma Tip: ID: 7.0mm
95 | RT13-20-010 | Cuffed and Uncuffed Tubes | CUFFED ORAL/NASAL ENDOTRACHEAL TUBE | CUFFED ORAL/NASAL ENDOTRACHEAL TUBE WITH REDUCED TRAUMA TIP - ADULT Purpose: Short term oral or nasal intubation with tip designed to reduce airway trauma Consists of a clear thermo-sensitive PVC tube that is kink resistant, atraumatic to airways and has a full Magill curvature and two Murphy eyes Must be radiopaque or have a radiopaque line Must be marked with numbered depth markings in 1cm intervals, tube size, length, ID, OD, trade name. The markings on tube must be clear and must not rub off when wiped with alcohol. Must have a precision band mark just above the cuff to aid with depth of intubation and guide proper placement. Preference will be given to tubes with a second circumferential band (not more than 2cm from each other) to guide proper placement. Proximal end - secure 15mm OD circuit connector made from medical grade plastic Distal end - Tapered, soft flexible pointed tip with an upturned leading edge , configured with the tip point in the midline so as to pass easily between the vocal cords without touching them. The tube must be able to demonstrate a reduction in intubation induced airway trauma. Must have a high-volume-low-pressure polyurethane barrel shaped cuff with pilot cuff. Cuff thickness ± 50 microns a shore hardness of no more than 86 microns The pilot cuff must have a one-way valve and must indicate the degree of inflation accurately. The cuff must inflate and deflate easily and when deflated must conform closely to tube to minimise the risk of trauma during intubation and extubation and not increase the OD of the tube. Cuff must be barrel shaped to facilitate optimal pressure dispersion against the tracheal wall. The cuff must be resistant to Nitrous Oxide effects All components must be latex-free sterile and individually packed in peel pouch that is easy to open - For SINGLE use only To comply with the latest ISO 5361 Information brochure or the outer package of the item should clearly indicate latex free | Cuffed Oral/Nasal endotracheal tube with Reduced Trauma Tip: ID: 7.5mm
96 | RT13-20-011 | Cuffed and Uncuffed Tubes | CUFFED ORAL/NASAL ENDOTRACHEAL TUBE | CUFFED ORAL/NASAL ENDOTRACHEAL TUBE WITH REDUCED TRAUMA TIP - ADULT Purpose: Short term oral or nasal intubation with tip designed to reduce airway trauma Consists of a clear thermo-sensitive PVC tube that is kink resistant, atraumatic to airways and has a full Magill curvature and two Murphy eyes Must be radiopaque or have a radiopaque line Must be marked with numbered depth markings in 1cm intervals, tube size, length, ID, OD, trade name. The markings on tube must be clear and must not rub off when wiped with alcohol. Must have a precision band mark just above the cuff to aid with depth of intubation and guide proper placement. Preference will be given to tubes with a second circumferential band (not more than 2cm from each other) to guide proper placement. Proximal end - secure 15mm OD circuit connector made from medical grade plastic Distal end - Tapered, soft flexible pointed tip with an upturned leading edge , configured with the tip point in the midline so as to pass easily between the vocal cords without touching them. The tube must be able to demonstrate a reduction in intubation induced airway trauma. Must have a high-volume-low-pressure polyurethane barrel shaped cuff with pilot cuff. Cuff thickness ± 50 microns a shore hardness of no more than 86 microns The pilot cuff must have a one-way valve and must indicate the degree of inflation accurately. The cuff must inflate and deflate easily and when deflated must conform closely to tube to minimise the risk of trauma during intubation and extubation and not increase the OD of the tube. Cuff must be barrel shaped to facilitate optimal pressure dispersion against the tracheal wall. The cuff must be resistant to Nitrous Oxide effects All components must be latex-free sterile and individually packed in peel pouch that is easy to open - For SINGLE use only To comply with the latest ISO 5361 Information brochure or the outer package of the item should clearly indicate latex free | Cuffed Oral/Nasal endotracheal tube with Reduced Trauma Tip: ID: 8.0mm
97 | RT13-20-012 | Cuffed and Uncuffed Tubes | CUFFED ORAL/NASAL ENDOTRACHEAL TUBE | CUFFED ORAL/NASAL ENDOTRACHEAL TUBE WITH REDUCED TRAUMA TIP - ADULT Purpose: Short term oral or nasal intubation with tip designed to reduce airway trauma Consists of a clear thermo-sensitive PVC tube that is kink resistant, atraumatic to airways and has a full Magill curvature and two Murphy eyes Must be radiopaque or have a radiopaque line Must be marked with numbered depth markings in 1cm intervals, tube size, length, ID, OD, trade name. The markings on tube must be clear and must not rub off when wiped with alcohol. Must have a precision band mark just above the cuff to aid with depth of intubation and guide proper placement. Preference will be given to tubes with a second circumferential band (not more than 2cm from each other) to guide proper placement. Proximal end - secure 15mm OD circuit connector made from medical grade plastic Distal end - Tapered, soft flexible pointed tip with an upturned leading edge , configured with the tip point in the midline so as to pass easily between the vocal cords without touching them. The tube must be able to demonstrate a reduction in intubation induced airway trauma. Must have a high-volume-low-pressure polyurethane barrel shaped cuff with pilot cuff. Cuff thickness ± 50 microns a shore hardness of no more than 86 microns The pilot cuff must have a one-way valve and must indicate the degree of inflation accurately. The cuff must inflate and deflate easily and when deflated must conform closely to tube to minimise the risk of trauma during intubation and extubation and not increase the OD of the tube. Cuff must be barrel shaped to facilitate optimal pressure dispersion against the tracheal wall. The cuff must be resistant to Nitrous Oxide effects All components must be latex-free sterile and individually packed in peel pouch that is easy to open - For SINGLE use only To comply with the latest ISO 5361 Information brochure or the outer package of the item should clearly indicate latex free | Cuffed Oral/Nasal endotracheal tube with Reduced Trauma Tip: ID: 8.5mm
98 | RT13-20-013 | Cuffed and Uncuffed Tubes | Cuffed Oral South-Facing Preformed | CUFFED ORAL SOUTH-FACING PREFORMED (RAE) ENDOTRACHEAL TUBE WITH REDUCED TRAUMA TIP Purpose: Short term oral intubation where tube is South facing and removed from surgical field Consists of a clear thermo-sensitive PVC tube that is kink resistant, with a South polar preformed curve with a soft flexible pointed tip and two Murphy eyes Must be able to temporarily straightened the tube to allow the passage of a suction catheter Must have a radiopaque line and must be marked with numbered depth markings in 1cm intervals, tube size, length, ID, OD, trade name. The markings on tube must be clear and must not rub off. Must also have marks/reference lines proximal to cuff to aid with depth of intubation and guide proper placement. Preference will be given to tubes with a second circumferential band (not more than 1cm from each other) to guide proper placement. Must also have a distinct mark at preformed curve to aid positioning Proximal end - secure 15mm OD circuit connector made from medical grade plastic Distal end - soft flexible pointed tip with an upturned leading edge , configured with the tip point in the midline so as to pass easily between the vocal cords without touching them. The tube must be able to demonstrate a reduction in intubation induced airway trauma. All components must be latex-free sterile and individually packed in peel pouch that is easy to open - For SINGLE use only To comply with the latest ISO 5361 Information brochure or the outer package of the item should clearly indicate latex free | Cuffed Oral South-Facing Preformed (RAE) Endotracheal Tube with Reduced Trauma Tip. ID: 3.0mm OD 4.0mm Length ± 105mm
99 | RT13-20-014 | Cuffed and Uncuffed Tubes | Cuffed Oral South-Facing Preformed | CUFFED ORAL SOUTH-FACING PREFORMED (RAE) ENDOTRACHEAL TUBE WITH REDUCED TRAUMA TIP Purpose: Short term oral intubation where tube is South facing and removed from surgical field Consists of a clear thermo-sensitive PVC tube that is kink resistant, with a South polar preformed curve with a soft flexible pointed tip and two Murphy eyes Must be able to temporarily straightened the tube to allow the passage of a suction catheter Must have a radiopaque line and must be marked with numbered depth markings in 1cm intervals, tube size, length, ID, OD, trade name. The markings on tube must be clear and must not rub off. Must also have marks/reference lines proximal to cuff to aid with depth of intubation and guide proper placement. Preference will be given to tubes with a second circumferential band (not more than 1cm from each other) to guide proper placement. Must also have a distinct mark at preformed curve to aid positioning Proximal end - secure 15mm OD circuit connector made from medical grade plastic Distal end - soft flexible pointed tip with an upturned leading edge , configured with the tip point in the midline so as to pass easily between the vocal cords without touching them. The tube must be able to demonstrate a reduction in intubation induced airway trauma. All components must be latex-free sterile and individually packed in peel pouch that is easy to open - For SINGLE use only To comply with the latest ISO 5361 Information brochure or the outer package of the item should clearly indicate latex free | Cuffed Oral South-Facing Preformed (RAE) Endotracheal Tube with Reduced Trauma Tip. ID: 3.0mm OD 4.0mm Length ± 105mm
100 | RT13-20-015 | Cuffed and Uncuffed Tubes | Cuffed Oral South-Facing Preformed | CUFFED ORAL SOUTH-FACING PREFORMED (RAE) ENDOTRACHEAL TUBE WITH REDUCED TRAUMA TIP Purpose: Short term oral intubation where tube is South facing and removed from surgical field Consists of a clear thermo-sensitive PVC tube that is kink resistant, with a South polar preformed curve with a soft flexible pointed tip and two Murphy eyes Must be able to temporarily straightened the tube to allow the passage of a suction catheter Must have a radiopaque line and must be marked with numbered depth markings in 1cm intervals, tube size, length, ID, OD, trade name. The markings on tube must be clear and must not rub off. Must also have marks/reference lines proximal to cuff to aid with depth of intubation and guide proper placement. Preference will be given to tubes with a second circumferential band (not more than 1cm from each other) to guide proper placement. Must also have a distinct mark at preformed curve to aid positioning Proximal end - secure 15mm OD circuit connector made from medical grade plastic Distal end - soft flexible pointed tip with an upturned leading edge , configured with the tip point in the midline so as to pass easily between the vocal cords without touching them. The tube must be able to demonstrate a reduction in intubation induced airway trauma. All components must be latex-free sterile and individually packed in peel pouch that is easy to open - For SINGLE use only To comply with the latest ISO 5361 Information brochure or the outer package of the item should clearly indicate latex free | Cuffed Oral South-Facing Preformed (RAE) Endotracheal Tube with Reduced Trauma Tip. ID: 4.0mm OD 5.3mm Length ± 130mm
101 | RT13-20-016 | Cuffed and Uncuffed Tubes | Cuffed Oral South-Facing Preformed | CUFFED ORAL SOUTH-FACING PREFORMED (RAE) ENDOTRACHEAL TUBE WITH REDUCED TRAUMA TIP Purpose: Short term oral intubation where tube is South facing and removed from surgical field Consists of a clear thermo-sensitive PVC tube that is kink resistant, with a South polar preformed curve with a soft flexible pointed tip and two Murphy eyes Must be able to temporarily straightened the tube to allow the passage of a suction catheter Must have a radiopaque line and must be marked with numbered depth markings in 1cm intervals, tube size, length, ID, OD, trade name. The markings on tube must be clear and must not rub off. Must also have marks/reference lines proximal to cuff to aid with depth of intubation and guide proper placement. Preference will be given to tubes with a second circumferential band (not more than 1cm from each other) to guide proper placement. Must also have a distinct mark at preformed curve to aid positioning Proximal end - secure 15mm OD circuit connector made from medical grade plastic Distal end - soft flexible pointed tip with an upturned leading edge , configured with the tip point in the midline so as to pass easily between the vocal cords without touching them. The tube must be able to demonstrate a reduction in intubation induced airway trauma. All components must be latex-free sterile and individually packed in peel pouch that is easy to open - For SINGLE use only To comply with the latest ISO 5361 Information brochure or the outer package of the item should clearly indicate latex free | Cuffed Oral South-Facing Preformed (RAE) Endotracheal Tube with Reduced Trauma Tip ID: 4.5mm OD 6.0mm Length ± 145mm
102 | RT13-20-017 | Cuffed and Uncuffed Tubes | Cuffed Oral South-Facing Preformed | CUFFED ORAL SOUTH-FACING PREFORMED (RAE) ENDOTRACHEAL TUBE WITH REDUCED TRAUMA TIP Purpose: Short term oral intubation where tube is South facing and removed from surgical field Consists of a clear thermo-sensitive PVC tube that is kink resistant, with a South polar preformed curve with a soft flexible pointed tip and two Murphy eyes Must be able to temporarily straightened the tube to allow the passage of a suction catheter Must have a radiopaque line and must be marked with numbered depth markings in 1cm intervals, tube size, length, ID, OD, trade name. The markings on tube must be clear and must not rub off. Must also have marks/reference lines proximal to cuff to aid with depth of intubation and guide proper placement. Preference will be given to tubes with a second circumferential band (not more than 1cm from each other) to guide proper placement. Must also have a distinct mark at preformed curve to aid positioning Proximal end - secure 15mm OD circuit connector made from medical grade plastic Distal end - soft flexible pointed tip with an upturned leading edge , configured with the tip point in the midline so as to pass easily between the vocal cords without touching them. The tube must be able to demonstrate a reduction in intubation induced airway trauma. All components must be latex-free sterile and individually packed in peel pouch that is easy to open - For SINGLE use only To comply with the latest ISO 5361 Information brochure or the outer package of the item should clearly indicate latex free | Cuffed Oral South-Facing Preformed (RAE) Endotracheal Tube with Reduced Trauma Tip. ID: 5.0mm OD 6.7mm Length ± 155mm
103 | RT13-20-018 | Cuffed and Uncuffed Tubes | Cuffed Oral South-Facing Preformed | CUFFED ORAL SOUTH-FACING PREFORMED (RAE) ENDOTRACHEAL TUBE WITH REDUCED TRAUMA TIP Purpose: Short term oral intubation where tube is South facing and removed from surgical field Consists of a clear thermo-sensitive PVC tube that is kink resistant, with a South polar preformed curve with a soft flexible pointed tip and two Murphy eyes Must be able to temporarily straightened the tube to allow the passage of a suction catheter Must have a radiopaque line and must be marked with numbered depth markings in 1cm intervals, tube size, length, ID, OD, trade name. The markings on tube must be clear and must not rub off. Must also have marks/reference lines proximal to cuff to aid with depth of intubation and guide proper placement. Preference will be given to tubes with a second circumferential band (not more than 1cm from each other) to guide proper placement. Must also have a distinct mark at preformed curve to aid positioning Proximal end - secure 15mm OD circuit connector made from medical grade plastic Distal end - soft flexible pointed tip with an upturned leading edge , configured with the tip point in the midline so as to pass easily between the vocal cords without touching them. The tube must be able to demonstrate a reduction in intubation induced airway trauma. All components must be latex-free sterile and individually packed in peel pouch that is easy to open - For SINGLE use only To comply with the latest ISO 5361 Information brochure or the outer package of the item should clearly indicate latex free | Cuffed Oral South-Facing Preformed (RAE) Endotracheal Tube with Reduced Trauma Tip ID: 5.5mm OD 7.3mm Length ± 165mm
104 | RT13-20-019 | Cuffed and Uncuffed Tubes | Cuffed Oral South-Facing Preformed | CUFFED ORAL SOUTH-FACING PREFORMED (RAE) ENDOTRACHEAL TUBE WITH REDUCED TRAUMA TIP Purpose: Short term oral intubation where tube is South facing and removed from surgical field Consists of a clear thermo-sensitive PVC tube that is kink resistant, with a South polar preformed curve with a soft flexible pointed tip and two Murphy eyes Must be able to temporarily straightened the tube to allow the passage of a suction catheter Must have a radiopaque line and must be marked with numbered depth markings in 1cm intervals, tube size, length, ID, OD, trade name. The markings on tube must be clear and must not rub off. Must also have marks/reference lines proximal to cuff to aid with depth of intubation and guide proper placement. Preference will be given to tubes with a second circumferential band (not more than 1cm from each other) to guide proper placement. Must also have a distinct mark at preformed curve to aid positioning Proximal end - secure 15mm OD circuit connector made from medical grade plastic Distal end - soft flexible pointed tip with an upturned leading edge , configured with the tip point in the midline so as to pass easily between the vocal cords without touching them. The tube must be able to demonstrate a reduction in intubation induced airway trauma. All components must be latex-free sterile and individually packed in peel pouch that is easy to open - For SINGLE use only To comply with the latest ISO 5361 Information brochure or the outer package of the item should clearly indicate latex free | Cuffed Oral South-Facing Preformed (RAE) Endotracheal Tube with Reduced Trauma Tip ID: 6.0mm OD 8.0mm Length ± 180mm
105 | RT13-20-020 | Cuffed and Uncuffed Tubes | Cuffed Oral South-Facing Preformed | CUFFED ORAL SOUTH-FACING PREFORMED (RAE) ENDOTRACHEAL TUBE WITH REDUCED TRAUMA TIP Purpose: Short term oral intubation where tube is South facing and removed from surgical field Consists of a clear thermo-sensitive PVC tube that is kink resistant, with a South polar preformed curve with a soft flexible pointed tip and two Murphy eyes Must be able to temporarily straightened the tube to allow the passage of a suction catheter Must have a radiopaque line and must be marked with numbered depth markings in 1cm intervals, tube size, length, ID, OD, trade name. The markings on tube must be clear and must not rub off. Must also have marks/reference lines proximal to cuff to aid with depth of intubation and guide proper placement. Preference will be given to tubes with a second circumferential band (not more than 1cm from each other) to guide proper placement. Must also have a distinct mark at preformed curve to aid positioning Proximal end - secure 15mm OD circuit connector made from medical grade plastic Distal end - soft flexible pointed tip with an upturned leading edge , configured with the tip point in the midline so as to pass easily between the vocal cords without touching them. The tube must be able to demonstrate a reduction in intubation induced airway trauma. All components must be latex-free sterile and individually packed in peel pouch that is easy to open - For SINGLE use only To comply with the latest ISO 5361 Information brochure or the outer package of the item should clearly indicate latex free | Cuffed Oral South-Facing Preformed (RAE) Endotracheal Tube with Reduced Trauma Tip ID: 6.5mm OD 8.2mm Length ± 180mm
106 | RT13-20-021 | Cuffed and Uncuffed Tubes | UNCUFFED NASAL NORTH-FACING | UNCUFFED NASAL NORTH-FACING PREFORMED (RAE) ENDOTRACHEAL TUBE WITH REDUCED TRAUMA TIP Purpose: Short term oral intubation where tube is North facing and removed from surgical field Consists of a clear thermo-sensitive PVC tube that is kink resistant, with a North polar preformed curve with a soft flexible pointed tip and two Murphy eyes Must be able to temporarily straightened the tube to allow the passage of a suction catheter Must have a radiopaque line and must be marked with numbered depth markings in 1cm intervals, tube size, length, ID, OD, trade name. The markings on tube must be clear and must not rub off. Must also have marks/reference lines proximal to cuff to aid with depth of intubation and guide proper placement. Preference will be given to tubes with a second circumferential band (not more than 1cm from each other) to guide proper placement. Must also have a distinct mark at preformed curve to aid positioning Proximal end - secure 15mm OD circuit connector made from medical grade plastic Distal end - Tapered soft flexible pointed tip with an upturned leading edge , configured with the tip point in the midline so as to pass easily between the vocal cords without touching them. The tube must be able to demonstrate a reduction in intubation induced airway trauma. All components must be latex-free sterile and individually packed in peel pouch that is easy to open - For SINGLE use only To comply with the latest ISO 5361 Information brochure or the outer package of the item should clearly indicate latex free | UnCuffed Oral South-Facing Preformed (RAE) Endotracheal Tube with Reduced Trauma Tip. ID: 3.5mm OD 4.0mm Length ± 105mm
107 | RT13-20-022 | Cuffed and Uncuffed Tubes | UNCUFFED NASAL NORTH-FACING | UNCUFFED NASAL NORTH-FACING PREFORMED (RAE) ENDOTRACHEAL TUBE WITH REDUCED TRAUMA TIP Purpose: Short term oral intubation where tube is North facing and removed from surgical field Consists of a clear thermo-sensitive PVC tube that is kink resistant, with a North polar preformed curve with a soft flexible pointed tip and two Murphy eyes Must be able to temporarily straightened the tube to allow the passage of a suction catheter Must have a radiopaque line and must be marked with numbered depth markings in 1cm intervals, tube size, length, ID, OD, trade name. The markings on tube must be clear and must not rub off. Must also have marks/reference lines proximal to cuff to aid with depth of intubation and guide proper placement. Preference will be given to tubes with a second circumferential band (not more than 1cm from each other) to guide proper placement. Must also have a distinct mark at preformed curve to aid positioning Proximal end - secure 15mm OD circuit connector made from medical grade plastic Distal end - Tapered soft flexible pointed tip with an upturned leading edge , configured with the tip point in the midline so as to pass easily between the vocal cords without touching them. The tube must be able to demonstrate a reduction in intubation induced airway trauma. All components must be latex-free sterile and individually packed in peel pouch that is easy to open - For SINGLE use only To comply with the latest ISO 5361 Information brochure or the outer package of the item should clearly indicate latex free | UnCuffed Oral South-Facing Preformed (RAE) Endotracheal Tube with Reduced Trauma Tip. ID: 4.0mm OD 5.3mm Length ± 130mm
108 | RT13-20-023 | Cuffed and Uncuffed Tubes | UNCUFFED NASAL NORTH-FACING | UNCUFFED NASAL NORTH-FACING PREFORMED (RAE) ENDOTRACHEAL TUBE WITH REDUCED TRAUMA TIP Purpose: Short term oral intubation where tube is North facing and removed from surgical field Consists of a clear thermo-sensitive PVC tube that is kink resistant, with a North polar preformed curve with a soft flexible pointed tip and two Murphy eyes Must be able to temporarily straightened the tube to allow the passage of a suction catheter Must have a radiopaque line and must be marked with numbered depth markings in 1cm intervals, tube size, length, ID, OD, trade name. The markings on tube must be clear and must not rub off. Must also have marks/reference lines proximal to cuff to aid with depth of intubation and guide proper placement. Preference will be given to tubes with a second circumferential band (not more than 1cm from each other) to guide proper placement. Must also have a distinct mark at preformed curve to aid positioning Proximal end - secure 15mm OD circuit connector made from medical grade plastic Distal end - Tapered soft flexible pointed tip with an upturned leading edge , configured with the tip point in the midline so as to pass easily between the vocal cords without touching them. The tube must be able to demonstrate a reduction in intubation induced airway trauma. All components must be latex-free sterile and individually packed in peel pouch that is easy to open - For SINGLE use only To comply with the latest ISO 5361 Information brochure or the outer package of the item should clearly indicate latex free | UnCuffed Oral South-Facing Preformed (RAE) Endotracheal Tube with Reduced Trauma Tip ID: 4.5mm OD 6.0mm Length ± 145mm
109 | RT13-20-024 | Cuffed and Uncuffed Tubes | UNCUFFED NASAL NORTH-FACING | UNCUFFED NASAL NORTH-FACING PREFORMED (RAE) ENDOTRACHEAL TUBE WITH REDUCED TRAUMA TIP Purpose: Short term oral intubation where tube is North facing and removed from surgical field Consists of a clear thermo-sensitive PVC tube that is kink resistant, with a North polar preformed curve with a soft flexible pointed tip and two Murphy eyes Must be able to temporarily straightened the tube to allow the passage of a suction catheter Must have a radiopaque line and must be marked with numbered depth markings in 1cm intervals, tube size, length, ID, OD, trade name. The markings on tube must be clear and must not rub off. Must also have marks/reference lines proximal to cuff to aid with depth of intubation and guide proper placement. Preference will be given to tubes with a second circumferential band (not more than 1cm from each other) to guide proper placement. Must also have a distinct mark at preformed curve to aid positioning Proximal end - secure 15mm OD circuit connector made from medical grade plastic Distal end - Tapered soft flexible pointed tip with an upturned leading edge , configured with the tip point in the midline so as to pass easily between the vocal cords without touching them. The tube must be able to demonstrate a reduction in intubation induced airway trauma. All components must be latex-free sterile and individually packed in peel pouch that is easy to open - For SINGLE use only To comply with the latest ISO 5361 Information brochure or the outer package of the item should clearly indicate latex free | UnCuffed Oral South-Facing Preformed (RAE) Endotracheal Tube with Reduced Trauma Tip. ID: 5.0mm OD 6.7mm Length ± 155mm
110 | RT13-20-025 | Cuffed and Uncuffed Tubes | UNCUFFED NASAL NORTH-FACING | UNCUFFED NASAL NORTH-FACING PREFORMED (RAE) ENDOTRACHEAL TUBE WITH REDUCED TRAUMA TIP Purpose: Short term oral intubation where tube is North facing and removed from surgical field Consists of a clear thermo-sensitive PVC tube that is kink resistant, with a North polar preformed curve with a soft flexible pointed tip and two Murphy eyes Must be able to temporarily straightened the tube to allow the passage of a suction catheter Must have a radiopaque line and must be marked with numbered depth markings in 1cm intervals, tube size, length, ID, OD, trade name. The markings on tube must be clear and must not rub off. Must also have marks/reference lines proximal to cuff to aid with depth of intubation and guide proper placement. Preference will be given to tubes with a second circumferential band (not more than 1cm from each other) to guide proper placement. Must also have a distinct mark at preformed curve to aid positioning Proximal end - secure 15mm OD circuit connector made from medical grade plastic Distal end - Tapered soft flexible pointed tip with an upturned leading edge , configured with the tip point in the midline so as to pass easily between the vocal cords without touching them. The tube must be able to demonstrate a reduction in intubation induced airway trauma. All components must be latex-free sterile and individually packed in peel pouch that is easy to open - For SINGLE use only To comply with the latest ISO 5361 Information brochure or the outer package of the item should clearly indicate latex free | UnCuffed Oral South-Facing Preformed (RAE) Endotracheal Tube with Reduced Trauma Tip ID: 5.5mm OD 7.3mm Length ± 165mm
111 | RT13-20-026 | Cuffed and Uncuffed Tubes | UNCUFFED NASAL NORTH-FACING | UNCUFFED NASAL NORTH-FACING PREFORMED (RAE) ENDOTRACHEAL TUBE WITH REDUCED TRAUMA TIP Purpose: Short term oral intubation where tube is North facing and removed from surgical field Consists of a clear thermo-sensitive PVC tube that is kink resistant, with a North polar preformed curve with a soft flexible pointed tip and two Murphy eyes Must be able to temporarily straightened the tube to allow the passage of a suction catheter Must have a radiopaque line and must be marked with numbered depth markings in 1cm intervals, tube size, length, ID, OD, trade name. The markings on tube must be clear and must not rub off. Must also have marks/reference lines proximal to cuff to aid with depth of intubation and guide proper placement. Preference will be given to tubes with a second circumferential band (not more than 1cm from each other) to guide proper placement. Must also have a distinct mark at preformed curve to aid positioning Proximal end - secure 15mm OD circuit connector made from medical grade plastic Distal end - Tapered soft flexible pointed tip with an upturned leading edge , configured with the tip point in the midline so as to pass easily between the vocal cords without touching them. The tube must be able to demonstrate a reduction in intubation induced airway trauma. All components must be latex-free sterile and individually packed in peel pouch that is easy to open - For SINGLE use only To comply with the latest ISO 5361 Information brochure or the outer package of the item should clearly indicate latex free | UnCuffed Oral South-Facing Preformed (RAE) Endotracheal Tube with Reduced Trauma Tip ID: 6.0mm OD 8.0mm Length ± 180mm
112 | RT13-20-027 | Cuffed and Uncuffed Tubes | UNCUFFED NASAL NORTH-FACING | UNCUFFED NASAL NORTH-FACING PREFORMED (RAE) ENDOTRACHEAL TUBE WITH REDUCED TRAUMA TIP Purpose: Short term oral intubation where tube is North facing and removed from surgical field Consists of a clear thermo-sensitive PVC tube that is kink resistant, with a North polar preformed curve with a soft flexible pointed tip and two Murphy eyes Must be able to temporarily straightened the tube to allow the passage of a suction catheter Must have a radiopaque line and must be marked with numbered depth markings in 1cm intervals, tube size, length, ID, OD, trade name. The markings on tube must be clear and must not rub off. Must also have marks/reference lines proximal to cuff to aid with depth of intubation and guide proper placement. Preference will be given to tubes with a second circumferential band (not more than 1cm from each other) to guide proper placement. Must also have a distinct mark at preformed curve to aid positioning Proximal end - secure 15mm OD circuit connector made from medical grade plastic Distal end - Tapered soft flexible pointed tip with an upturned leading edge , configured with the tip point in the midline so as to pass easily between the vocal cords without touching them. The tube must be able to demonstrate a reduction in intubation induced airway trauma. All components must be latex-free sterile and individually packed in peel pouch that is easy to open - For SINGLE use only To comply with the latest ISO 5361 Information brochure or the outer package of the item should clearly indicate latex free | UnCuffed Oral South-Facing Preformed (RAE) Endotracheal Tube with Reduced Trauma Tip ID: 6.5mm OD 8.2mm Length ± 180mm
113 | RT13-20-028 | Cuffed and Uncuffed Tubes | UNCUFFED NASAL NORTH-FACING | UNCUFFED NASAL NORTH-FACING PREFORMED (RAE) ENDOTRACHEAL TUBE WITH REDUCED TRAUMA TIP Purpose: Short term oral intubation where tube is North facing and removed from surgical field Consists of a clear thermo-sensitive PVC tube that is kink resistant, with a North polar preformed curve with a soft flexible pointed tip and two Murphy eyes Must be able to temporarily straightened the tube to allow the passage of a suction catheter Must have a radiopaque line and must be marked with numbered depth markings in 1cm intervals, tube size, length, ID, OD, trade name. The markings on tube must be clear and must not rub off. Must also have marks/reference lines proximal to cuff to aid with depth of intubation and guide proper placement. Preference will be given to tubes with a second circumferential band (not more than 1cm from each other) to guide proper placement. Must also have a distinct mark at preformed curve to aid positioning Proximal end - secure 15mm OD circuit connector made from medical grade plastic Distal end - Tapered soft flexible pointed tip with an upturned leading edge , configured with the tip point in the midline so as to pass easily between the vocal cords without touching them. The tube must be able to demonstrate a reduction in intubation induced airway trauma. All components must be latex-free sterile and individually packed in peel pouch that is easy to open - For SINGLE use only To comply with the latest ISO 5361 Information brochure or the outer package of the item should clearly indicate latex free | Uncuffed Nasal North-Facing Preformed (RAE) Endotracheal Tube with Reduced Trauma Tip ID: 3.0mm OD 4.0mm Length at nasal curve ± 155mm
114 | RT13-20-029 | Cuffed and Uncuffed Tubes | UNCUFFED REINFORCED ORAL/NASAL | UNCUFFED REINFORCED ORAL/NASAL ENDOTRACHEAL TUBE Purpose: Oral or nasal intubation where tube patency and flexibility is essential Consists of a soft flexible PVC tube with spiral-wound medical grade reinforcing wire that is sealed tightly against the bonded 15mm connector. Must be kink resistant and atraumatic to airways. Must be radiopaque or have a radiopaque line Must be marked with numbered depth markings in 1cm intervals, tube size, length, ID, OD, trade name. The markings on tube must be clear and must not rub off when wiped with alcohol. Must have a precision band mark to aid with depth of intubation and guide proper placement. Solid or circumferential precision band markings accepted Preference will be given to tubes with a second circumferential band (not more than 1cm from each other) to guide proper placement. Proximal end - secure 15mm OD circuit connector made from medical grade plastic Distal end - Atraumatic left-facing bevelled tip with Murphy eye on the opposite side All components must be latex-free, sterile and individually packed in peel pouch that is easy to open - For SINGLE use only To comply with the latest ISO 5361 Information brochure or the outer package of the item should clearly indicate latex free | Uncuffed Reinforced Oral/Nasal endotracheal tube: Size ID: 3.0mm OD 4.8mm Length ±184mm
115 | RT13-20-030 | Cuffed and Uncuffed Tubes | UNCUFFED REINFORCED ORAL/NASAL | UNCUFFED REINFORCED ORAL/NASAL ENDOTRACHEAL TUBE Purpose: Oral or nasal intubation where tube patency and flexibility is essential Consists of a soft flexible PVC tube with spiral-wound medical grade reinforcing wire that is sealed tightly against the bonded 15mm connector. Must be kink resistant and atraumatic to airways. Must be radiopaque or have a radiopaque line Must be marked with numbered depth markings in 1cm intervals, tube size, length, ID, OD, trade name. The markings on tube must be clear and must not rub off when wiped with alcohol. Must have a precision band mark to aid with depth of intubation and guide proper placement. Solid or circumferential precision band markings accepted Preference will be given to tubes with a second circumferential band (not more than 1cm from each other) to guide proper placement. Proximal end - secure 15mm OD circuit connector made from medical grade plastic Distal end - Atraumatic left-facing bevelled tip with Murphy eye on the opposite side All components must be latex-free, sterile and individually packed in peel pouch that is easy to open - For SINGLE use only To comply with the latest ISO 5361 Information brochure or the outer package of the item should clearly indicate latex free | Uncuffed Reinforced Oral/Nasal endotracheal tube: Size ID: 3.5mm OD 5.5mm Length ±185mm
116 | RT13-20-031 | Cuffed and Uncuffed Tubes | UNCUFFED REINFORCED ORAL/NASAL | UNCUFFED REINFORCED ORAL/NASAL ENDOTRACHEAL TUBE Purpose: Oral or nasal intubation where tube patency and flexibility is essential Consists of a soft flexible PVC tube with spiral-wound medical grade reinforcing wire that is sealed tightly against the bonded 15mm connector. Must be kink resistant and atraumatic to airways. Must be radiopaque or have a radiopaque line Must be marked with numbered depth markings in 1cm intervals, tube size, length, ID, OD, trade name. The markings on tube must be clear and must not rub off when wiped with alcohol. Must have a precision band mark to aid with depth of intubation and guide proper placement. Solid or circumferential precision band markings accepted Preference will be given to tubes with a second circumferential band (not more than 1cm from each other) to guide proper placement. Proximal end - secure 15mm OD circuit connector made from medical grade plastic Distal end - Atraumatic left-facing bevelled tip with Murphy eye on the opposite side All components must be latex-free, sterile and individually packed in peel pouch that is easy to open - For SINGLE use only To comply with the latest ISO 5361 Information brochure or the outer package of the item should clearly indicate latex free | Uncuffed Reinforced Oral/Nasal endotracheal tube: Size ID: 4.0mm OD 6.0mm Length ±195mm
117 | RT13-20-032 | Cuffed and Uncuffed Tubes | UNCUFFED REINFORCED ORAL/NASAL | UNCUFFED REINFORCED ORAL/NASAL ENDOTRACHEAL TUBE Purpose: Oral or nasal intubation where tube patency and flexibility is essential Consists of a soft flexible PVC tube with spiral-wound medical grade reinforcing wire that is sealed tightly against the bonded 15mm connector. Must be kink resistant and atraumatic to airways. Must be radiopaque or have a radiopaque line Must be marked with numbered depth markings in 1cm intervals, tube size, length, ID, OD, trade name. The markings on tube must be clear and must not rub off when wiped with alcohol. Must have a precision band mark to aid with depth of intubation and guide proper placement. Solid or circumferential precision band markings accepted Preference will be given to tubes with a second circumferential band (not more than 1cm from each other) to guide proper placement. Proximal end - secure 15mm OD circuit connector made from medical grade plastic Distal end - Atraumatic left-facing bevelled tip with Murphy eye on the opposite side All components must be latex-free, sterile and individually packed in peel pouch that is easy to open - For SINGLE use only To comply with the latest ISO 5361 Information brochure or the outer package of the item should clearly indicate latex free | Uncuffed Reinforced Oral/Nasal endotracheal tube: Size ID: 4.5mm OD 6.5mm Length ±198mm
118 | RT13-20-033 | Cuffed and Uncuffed Tubes | UNCUFFED REINFORCED ORAL/NASAL | UNCUFFED REINFORCED ORAL/NASAL ENDOTRACHEAL TUBE Purpose: Oral or nasal intubation where tube patency and flexibility is essential Consists of a soft flexible PVC tube with spiral-wound medical grade reinforcing wire that is sealed tightly against the bonded 15mm connector. Must be kink resistant and atraumatic to airways. Must be radiopaque or have a radiopaque line Must be marked with numbered depth markings in 1cm intervals, tube size, length, ID, OD, trade name. The markings on tube must be clear and must not rub off when wiped with alcohol. Must have a precision band mark to aid with depth of intubation and guide proper placement. Solid or circumferential precision band markings accepted Preference will be given to tubes with a second circumferential band (not more than 1cm from each other) to guide proper placement. Proximal end - secure 15mm OD circuit connector made from medical grade plastic Distal end - Atraumatic left-facing bevelled tip with Murphy eye on the opposite side All components must be latex-free, sterile and individually packed in peel pouch that is easy to open - For SINGLE use only To comply with the latest ISO 5361 Information brochure or the outer package of the item should clearly indicate latex free | Uncuffed Reinforced Oral/Nasal endotracheal tube: Size ID: 5.0mm OD 7.2mm Length ±208mm
119 | RT13-20-034 | Cuffed and Uncuffed Tubes | UNCUFFED REINFORCED ORAL/NASAL | UNCUFFED REINFORCED ORAL/NASAL ENDOTRACHEAL TUBE Purpose: Oral or nasal intubation where tube patency and flexibility is essential Consists of a soft flexible PVC tube with spiral-wound medical grade reinforcing wire that is sealed tightly against the bonded 15mm connector. Must be kink resistant and atraumatic to airways. Must be radiopaque or have a radiopaque line Must be marked with numbered depth markings in 1cm intervals, tube size, length, ID, OD, trade name. The markings on tube must be clear and must not rub off when wiped with alcohol. Must have a precision band mark to aid with depth of intubation and guide proper placement. Solid or circumferential precision band markings accepted Preference will be given to tubes with a second circumferential band (not more than 1cm from each other) to guide proper placement. Proximal end - secure 15mm OD circuit connector made from medical grade plastic Distal end - Atraumatic left-facing bevelled tip with Murphy eye on the opposite side All components must be latex-free, sterile and individually packed in peel pouch that is easy to open - For SINGLE use only To comply with the latest ISO 5361 Information brochure or the outer package of the item should clearly indicate latex free | Uncuffed Reinforced Oral/Nasal endotracheal tube: Size ID: 5.5mm OD 7.7mm Length ±208mm
120 | RT13-20-035 | Cuffed and Uncuffed Tubes | UNCUFFED REINFORCED ORAL/NASAL | UNCUFFED REINFORCED ORAL/NASAL ENDOTRACHEAL TUBE Purpose: Oral or nasal intubation where tube patency and flexibility is essential Consists of a soft flexible PVC tube with spiral-wound medical grade reinforcing wire that is sealed tightly against the bonded 15mm connector. Must be kink resistant and atraumatic to airways. Must be radiopaque or have a radiopaque line Must be marked with numbered depth markings in 1cm intervals, tube size, length, ID, OD, trade name. The markings on tube must be clear and must not rub off when wiped with alcohol. Must have a precision band mark to aid with depth of intubation and guide proper placement. Solid or circumferential precision band markings accepted Preference will be given to tubes with a second circumferential band (not more than 1cm from each other) to guide proper placement. Proximal end - secure 15mm OD circuit connector made from medical grade plastic Distal end - Atraumatic left-facing bevelled tip with Murphy eye on the opposite side All components must be latex-free, sterile and individually packed in peel pouch that is easy to open - For SINGLE use only To comply with the latest ISO 5361 Information brochure or the outer package of the item should clearly indicate latex free | Uncuffed Reinforced Oral/Nasal endotracheal tube: Size ID: 6.0mm OD 8.4mm Length ± 329mm
121 | RT13-20-036 | Cuffed and Uncuffed Tubes | UNCUFFED REINFORCED ORAL/NASAL | CUFFED REINFORCED ORAL/NASAL ENDOTRACHEAL TUBE - PAEDIATRIC Purpose: Oral or nasal intubation where tube patency and flexibility is essential Consists of a soft flexible PVC tube with spiral-wound medical grade reinforcing wire that is sealed tightly against the bonded 15mm connector. Must be kink resistant and atraumatic to airways. Must be radiopaque or have a radiopaque line Must be marked with numbered depth markings in 1cm intervals, tube size, length, ID, OD, trade name. The markings on tube must be clear and must not rub off when wiped with alcohol. Must have a precision band mark to aid with depth of intubation and guide proper placement. Solid or circumferential precision band markings accepted Preference will be given to tubes with a second circumferential band (not more than 1cm from each other) to guide proper placement. Proximal end - secure 15mm OD circuit connector made from medical grade plastic Distal end - Atraumatic left-facing bevelled tip with Murphy eye on the opposite side Must have a high-volume-low-pressure polyurethane barrel shaped cuff with pilot cuff. Cuff thickness ± 50 microns. The pilot cuff must have a one-way valve and must indicate the degree of inflation accurately. The cuff must inflate and deflate easily and when deflated must conform closely to tube to minimise the risk of trauma during intubation and extubation and not increase the OD of the tube. The cuff must be resistant to Nitrous Oxide effects All components must be latex-free, sterile and individually packed in peel pouch that is easy to open - For SINGLE use only To comply with the latest ISO 5361 Information brochure or the outer package of the item should clearly indicate latex free | Cuffed Reinforced Oral/Nasal endotracheal tube: ID: 3.0mm OD 4.8mm Length ±184mm
122 | RT13-20-037 | Cuffed and Uncuffed Tubes | UNCUFFED REINFORCED ORAL/NASAL | CUFFED REINFORCED ORAL/NASAL ENDOTRACHEAL TUBE - PAEDIATRIC Purpose: Oral or nasal intubation where tube patency and flexibility is essential Consists of a soft flexible PVC tube with spiral-wound medical grade reinforcing wire that is sealed tightly against the bonded 15mm connector. Must be kink resistant and atraumatic to airways. Must be radiopaque or have a radiopaque line Must be marked with numbered depth markings in 1cm intervals, tube size, length, ID, OD, trade name. The markings on tube must be clear and must not rub off when wiped with alcohol. Must have a precision band mark to aid with depth of intubation and guide proper placement. Solid or circumferential precision band markings accepted Preference will be given to tubes with a second circumferential band (not more than 1cm from each other) to guide proper placement. Proximal end - secure 15mm OD circuit connector made from medical grade plastic Distal end - Atraumatic left-facing bevelled tip with Murphy eye on the opposite side Must have a high-volume-low-pressure polyurethane barrel shaped cuff with pilot cuff. Cuff thickness ± 50 microns. The pilot cuff must have a one-way valve and must indicate the degree of inflation accurately. The cuff must inflate and deflate easily and when deflated must conform closely to tube to minimise the risk of trauma during intubation and extubation and not increase the OD of the tube. The cuff must be resistant to Nitrous Oxide effects All components must be latex-free, sterile and individually packed in peel pouch that is easy to open - For SINGLE use only To comply with the latest ISO 5361 Information brochure or the outer package of the item should clearly indicate latex free | Cuffed Reinforced Oral/Nasal endotracheal tube: ID: 3.5mm OD 5.5mm Length ±185mm
123 | RT13-20-038 | Cuffed and Uncuffed Tubes | UNCUFFED REINFORCED ORAL/NASAL | CUFFED REINFORCED ORAL/NASAL ENDOTRACHEAL TUBE - PAEDIATRIC Purpose: Oral or nasal intubation where tube patency and flexibility is essential Consists of a soft flexible PVC tube with spiral-wound medical grade reinforcing wire that is sealed tightly against the bonded 15mm connector. Must be kink resistant and atraumatic to airways. Must be radiopaque or have a radiopaque line Must be marked with numbered depth markings in 1cm intervals, tube size, length, ID, OD, trade name. The markings on tube must be clear and must not rub off when wiped with alcohol. Must have a precision band mark to aid with depth of intubation and guide proper placement. Solid or circumferential precision band markings accepted Preference will be given to tubes with a second circumferential band (not more than 1cm from each other) to guide proper placement. Proximal end - secure 15mm OD circuit connector made from medical grade plastic Distal end - Atraumatic left-facing bevelled tip with Murphy eye on the opposite side Must have a high-volume-low-pressure polyurethane barrel shaped cuff with pilot cuff. Cuff thickness ± 50 microns. The pilot cuff must have a one-way valve and must indicate the degree of inflation accurately. The cuff must inflate and deflate easily and when deflated must conform closely to tube to minimise the risk of trauma during intubation and extubation and not increase the OD of the tube. The cuff must be resistant to Nitrous Oxide effects All components must be latex-free, sterile and individually packed in peel pouch that is easy to open - For SINGLE use only To comply with the latest ISO 5361 Information brochure or the outer package of the item should clearly indicate latex free | Cuffed Reinforced Oral/Nasal endotracheal tube: ID: 4.0mm OD 6.0mm Length ±195mm
124 | RT13-20-039 | Cuffed and Uncuffed Tubes | UNCUFFED REINFORCED ORAL/NASAL | CUFFED REINFORCED ORAL/NASAL ENDOTRACHEAL TUBE - PAEDIATRIC Purpose: Oral or nasal intubation where tube patency and flexibility is essential Consists of a soft flexible PVC tube with spiral-wound medical grade reinforcing wire that is sealed tightly against the bonded 15mm connector. Must be kink resistant and atraumatic to airways. Must be radiopaque or have a radiopaque line Must be marked with numbered depth markings in 1cm intervals, tube size, length, ID, OD, trade name. The markings on tube must be clear and must not rub off when wiped with alcohol. Must have a precision band mark to aid with depth of intubation and guide proper placement. Solid or circumferential precision band markings accepted Preference will be given to tubes with a second circumferential band (not more than 1cm from each other) to guide proper placement. Proximal end - secure 15mm OD circuit connector made from medical grade plastic Distal end - Atraumatic left-facing bevelled tip with Murphy eye on the opposite side Must have a high-volume-low-pressure polyurethane barrel shaped cuff with pilot cuff. Cuff thickness ± 50 microns. The pilot cuff must have a one-way valve and must indicate the degree of inflation accurately. The cuff must inflate and deflate easily and when deflated must conform closely to tube to minimise the risk of trauma during intubation and extubation and not increase the OD of the tube. The cuff must be resistant to Nitrous Oxide effects All components must be latex-free, sterile and individually packed in peel pouch that is easy to open - For SINGLE use only To comply with the latest ISO 5361 Information brochure or the outer package of the item should clearly indicate latex free | Cuffed Reinforced Oral/Nasal endotracheal tube: ID: 4.5mm OD 6.5mm Length ±198mm
125 | RT13-20-040 | Cuffed and Uncuffed Tubes | UNCUFFED REINFORCED ORAL/NASAL | CUFFED REINFORCED ORAL/NASAL ENDOTRACHEAL TUBE - PAEDIATRIC Purpose: Oral or nasal intubation where tube patency and flexibility is essential Consists of a soft flexible PVC tube with spiral-wound medical grade reinforcing wire that is sealed tightly against the bonded 15mm connector. Must be kink resistant and atraumatic to airways. Must be radiopaque or have a radiopaque line Must be marked with numbered depth markings in 1cm intervals, tube size, length, ID, OD, trade name. The markings on tube must be clear and must not rub off when wiped with alcohol. Must have a precision band mark to aid with depth of intubation and guide proper placement. Solid or circumferential precision band markings accepted Preference will be given to tubes with a second circumferential band (not more than 1cm from each other) to guide proper placement. Proximal end - secure 15mm OD circuit connector made from medical grade plastic Distal end - Atraumatic left-facing bevelled tip with Murphy eye on the opposite side Must have a high-volume-low-pressure polyurethane barrel shaped cuff with pilot cuff. Cuff thickness ± 50 microns. The pilot cuff must have a one-way valve and must indicate the degree of inflation accurately. The cuff must inflate and deflate easily and when deflated must conform closely to tube to minimise the risk of trauma during intubation and extubation and not increase the OD of the tube. The cuff must be resistant to Nitrous Oxide effects All components must be latex-free, sterile and individually packed in peel pouch that is easy to open - For SINGLE use only To comply with the latest ISO 5361 Information brochure or the outer package of the item should clearly indicate latex free | Cuffed Reinforced Oral/Nasal endotracheal tube: ID: 5.0mm OD 7.3mm Length ± 321mm
126 | RT13-20-041 | Cuffed and Uncuffed Tubes | UNCUFFED REINFORCED ORAL/NASAL | CUFFED REINFORCED ORAL/NASAL ENDOTRACHEAL TUBE - PAEDIATRIC Purpose: Oral or nasal intubation where tube patency and flexibility is essential Consists of a soft flexible PVC tube with spiral-wound medical grade reinforcing wire that is sealed tightly against the bonded 15mm connector. Must be kink resistant and atraumatic to airways. Must be radiopaque or have a radiopaque line Must be marked with numbered depth markings in 1cm intervals, tube size, length, ID, OD, trade name. The markings on tube must be clear and must not rub off when wiped with alcohol. Must have a precision band mark to aid with depth of intubation and guide proper placement. Solid or circumferential precision band markings accepted Preference will be given to tubes with a second circumferential band (not more than 1cm from each other) to guide proper placement. Proximal end - secure 15mm OD circuit connector made from medical grade plastic Distal end - Atraumatic left-facing bevelled tip with Murphy eye on the opposite side Must have a high-volume-low-pressure polyurethane barrel shaped cuff with pilot cuff. Cuff thickness ± 50 microns. The pilot cuff must have a one-way valve and must indicate the degree of inflation accurately. The cuff must inflate and deflate easily and when deflated must conform closely to tube to minimise the risk of trauma during intubation and extubation and not increase the OD of the tube. The cuff must be resistant to Nitrous Oxide effects All components must be latex-free, sterile and individually packed in peel pouch that is easy to open - For SINGLE use only To comply with the latest ISO 5361 Information brochure or the outer package of the item should clearly indicate latex free | Cuffed Reinforced Oral/Nasal endotracheal tube: ID: 5.5mm OD 7.8mm Length ± 322mm
127 | RT13-20-042 | Cuffed and Uncuffed Tubes | UNCUFFED REINFORCED ORAL/NASAL | CUFFED REINFORCED ORAL/NASAL ENDOTRACHEAL TUBE - PAEDIATRIC Purpose: Oral or nasal intubation where tube patency and flexibility is essential Consists of a soft flexible PVC tube with spiral-wound medical grade reinforcing wire that is sealed tightly against the bonded 15mm connector. Must be kink resistant and atraumatic to airways. Must be radiopaque or have a radiopaque line Must be marked with numbered depth markings in 1cm intervals, tube size, length, ID, OD, trade name. The markings on tube must be clear and must not rub off when wiped with alcohol. Must have a precision band mark to aid with depth of intubation and guide proper placement. Solid or circumferential precision band markings accepted Preference will be given to tubes with a second circumferential band (not more than 1cm from each other) to guide proper placement. Proximal end - secure 15mm OD circuit connector made from medical grade plastic Distal end - Atraumatic left-facing bevelled tip with Murphy eye on the opposite side Must have a high-volume-low-pressure polyurethane barrel shaped cuff with pilot cuff. Cuff thickness ± 50 microns. The pilot cuff must have a one-way valve and must indicate the degree of inflation accurately. The cuff must inflate and deflate easily and when deflated must conform closely to tube to minimise the risk of trauma during intubation and extubation and not increase the OD of the tube. The cuff must be resistant to Nitrous Oxide effects All components must be latex-free, sterile and individually packed in peel pouch that is easy to open - For SINGLE use only To comply with the latest ISO 5361 Information brochure or the outer package of the item should clearly indicate latex free | Cuffed Reinforced Oral/Nasal endotracheal tube: ID: 6.0mm OD 8.4mm Length ± 329mm
128 | RT13-21-001 | ENDOTRACHEAL TUBE | ENDOTRACHEAL TUBE | TAPER-CUFFED ENDOTRACHEAL TUBE Purpose: endotracheal tube with a tapered cuff for long stay oral/nasal intubation aimed to reduce micro-aspiration Consists of a clear thermo-sensitive PVC tube that is kink resistant, atraumatic to airways and have a full Magill curvature. Must have a radiopaque line and must be marked with numbered depth markings in 1cm intervals, tube size, length, ID, OD, trade name. The markings on tube must be clear and must not rub off. Must have a precision band mark just above the cuff to aid with depth of intubation and guide proper placement. Preference will be given to tubes with a second circumferential band (not more than 2cm from each other) to guide proper placement. Proximal end - secure 15mm OD circuit connector made from medical grade plastic Distal end - Atraumatic left-facing bevelled tip with Murphy eye on the opposite side Must have a high-volume-low-pressure polyurethane taper-shaped cuff with pilot cuff. Cuff thickness ± 50 microns The cuff must inflate and deflate easily and when deflated must conform closely to tube to minimise the risk of trauma during intubation and extubation. Cuff must be barrel shaped to facilitate optimal pressure dispersion against the tracheal wall The cuff must be resistant to Nitrous Oxide effects The pilot cuff must have a one-way valve and must indicate the degree of inflation accurately. All components must be latex-free sterile and individually packed in peel pouch that is easy to open - For SINGLE use only To comply with the latest ISO 5361:2016 Information brochure or the outer package of the item should clearly indicate latex free | Taper-Cuffed Oral/Nasal endotracheal tube: ID: 5.0mm OD: 6.9mm
129 | RT13-21-002 | ENDOTRACHEAL TUBE | ENDOTRACHEAL TUBE | TAPER-CUFFED ENDOTRACHEAL TUBE Purpose: endotracheal tube with a tapered cuff for long stay oral/nasal intubation aimed to reduce micro-aspiration Consists of a clear thermo-sensitive PVC tube that is kink resistant, atraumatic to airways and have a full Magill curvature. Must have a radiopaque line and must be marked with numbered depth markings in 1cm intervals, tube size, length, ID, OD, trade name. The markings on tube must be clear and must not rub off. Must have a precision band mark just above the cuff to aid with depth of intubation and guide proper placement. Preference will be given to tubes with a second circumferential band (not more than 2cm from each other) to guide proper placement. Proximal end - secure 15mm OD circuit connector made from medical grade plastic Distal end - Atraumatic left-facing bevelled tip with Murphy eye on the opposite side Must have a high-volume-low-pressure polyurethane taper-shaped cuff with pilot cuff. Cuff thickness ± 50 microns The cuff must inflate and deflate easily and when deflated must conform closely to tube to minimise the risk of trauma during intubation and extubation. Cuff must be barrel shaped to facilitate optimal pressure dispersion against the tracheal wall The cuff must be resistant to Nitrous Oxide effects The pilot cuff must have a one-way valve and must indicate the degree of inflation accurately. All components must be latex-free sterile and individually packed in peel pouch that is easy to open - For SINGLE use only To comply with the latest ISO 5361:2016 Information brochure or the outer package of the item should clearly indicate latex free | Taper-Cuffed Oral/Nasal endotracheal tube: ID: 5.5mm OD: 7.5mm
130 | RT13-21-003 | ENDOTRACHEAL TUBE | ENDOTRACHEAL TUBE | TAPER-CUFFED ENDOTRACHEAL TUBE Purpose: endotracheal tube with a tapered cuff for long stay oral/nasal intubation aimed to reduce micro-aspiration Consists of a clear thermo-sensitive PVC tube that is kink resistant, atraumatic to airways and have a full Magill curvature. Must have a radiopaque line and must be marked with numbered depth markings in 1cm intervals, tube size, length, ID, OD, trade name. The markings on tube must be clear and must not rub off. Must have a precision band mark just above the cuff to aid with depth of intubation and guide proper placement. Preference will be given to tubes with a second circumferential band (not more than 2cm from each other) to guide proper placement. Proximal end - secure 15mm OD circuit connector made from medical grade plastic Distal end - Atraumatic left-facing bevelled tip with Murphy eye on the opposite side Must have a high-volume-low-pressure polyurethane taper-shaped cuff with pilot cuff. Cuff thickness ± 50 microns The cuff must inflate and deflate easily and when deflated must conform closely to tube to minimise the risk of trauma during intubation and extubation. Cuff must be barrel shaped to facilitate optimal pressure dispersion against the tracheal wall The cuff must be resistant to Nitrous Oxide effects The pilot cuff must have a one-way valve and must indicate the degree of inflation accurately. All components must be latex-free sterile and individually packed in peel pouch that is easy to open - For SINGLE use only To comply with the latest ISO 5361:2016 Information brochure or the outer package of the item should clearly indicate latex free | Taper-Cuffed Oral/Nasal endotracheal tube: ID: 6.0mm OD: 8.2mm
131 | RT13-21-004 | ENDOTRACHEAL TUBE | ENDOTRACHEAL TUBE | TAPER-CUFFED ENDOTRACHEAL TUBE Purpose: endotracheal tube with a tapered cuff for long stay oral/nasal intubation aimed to reduce micro-aspiration Consists of a clear thermo-sensitive PVC tube that is kink resistant, atraumatic to airways and have a full Magill curvature. Must have a radiopaque line and must be marked with numbered depth markings in 1cm intervals, tube size, length, ID, OD, trade name. The markings on tube must be clear and must not rub off. Must have a precision band mark just above the cuff to aid with depth of intubation and guide proper placement. Preference will be given to tubes with a second circumferential band (not more than 2cm from each other) to guide proper placement. Proximal end - secure 15mm OD circuit connector made from medical grade plastic Distal end - Atraumatic left-facing bevelled tip with Murphy eye on the opposite side Must have a high-volume-low-pressure polyurethane taper-shaped cuff with pilot cuff. Cuff thickness ± 50 microns The cuff must inflate and deflate easily and when deflated must conform closely to tube to minimise the risk of trauma during intubation and extubation. Cuff must be barrel shaped to facilitate optimal pressure dispersion against the tracheal wall The cuff must be resistant to Nitrous Oxide effects The pilot cuff must have a one-way valve and must indicate the degree of inflation accurately. All components must be latex-free sterile and individually packed in peel pouch that is easy to open - For SINGLE use only To comply with the latest ISO 5361:2016 Information brochure or the outer package of the item should clearly indicate latex free | Taper-Cuffed Oral/Nasal endotracheal tube: ID: 6.5mm OD: 8.9mm
132 | RT13-21-005 | ENDOTRACHEAL TUBE | ENDOTRACHEAL TUBE | TAPER-CUFFED ENDOTRACHEAL TUBE Purpose: endotracheal tube with a tapered cuff for long stay oral/nasal intubation aimed to reduce micro-aspiration Consists of a clear thermo-sensitive PVC tube that is kink resistant, atraumatic to airways and have a full Magill curvature. Must have a radiopaque line and must be marked with numbered depth markings in 1cm intervals, tube size, length, ID, OD, trade name. The markings on tube must be clear and must not rub off. Must have a precision band mark just above the cuff to aid with depth of intubation and guide proper placement. Preference will be given to tubes with a second circumferential band (not more than 2cm from each other) to guide proper placement. Proximal end - secure 15mm OD circuit connector made from medical grade plastic Distal end - Atraumatic left-facing bevelled tip with Murphy eye on the opposite side Must have a high-volume-low-pressure polyurethane taper-shaped cuff with pilot cuff. Cuff thickness ± 50 microns The cuff must inflate and deflate easily and when deflated must conform closely to tube to minimise the risk of trauma during intubation and extubation. Cuff must be barrel shaped to facilitate optimal pressure dispersion against the tracheal wall The cuff must be resistant to Nitrous Oxide effects The pilot cuff must have a one-way valve and must indicate the degree of inflation accurately. All components must be latex-free sterile and individually packed in peel pouch that is easy to open - For SINGLE use only To comply with the latest ISO 5361:2016 Information brochure or the outer package of the item should clearly indicate latex free | Taper-Cuffed Oral/Nasal endotracheal tube: ID: 7.0mm OD: 9.5mm
133 | RT13-21-006 | ENDOTRACHEAL TUBE | ENDOTRACHEAL TUBE | TAPER-CUFFED ENDOTRACHEAL TUBE Purpose: endotracheal tube with a tapered cuff for long stay oral/nasal intubation aimed to reduce micro-aspiration Consists of a clear thermo-sensitive PVC tube that is kink resistant, atraumatic to airways and have a full Magill curvature. Must have a radiopaque line and must be marked with numbered depth markings in 1cm intervals, tube size, length, ID, OD, trade name. The markings on tube must be clear and must not rub off. Must have a precision band mark just above the cuff to aid with depth of intubation and guide proper placement. Preference will be given to tubes with a second circumferential band (not more than 2cm from each other) to guide proper placement. Proximal end - secure 15mm OD circuit connector made from medical grade plastic Distal end - Atraumatic left-facing bevelled tip with Murphy eye on the opposite side Must have a high-volume-low-pressure polyurethane taper-shaped cuff with pilot cuff. Cuff thickness ± 50 microns The cuff must inflate and deflate easily and when deflated must conform closely to tube to minimise the risk of trauma during intubation and extubation. Cuff must be barrel shaped to facilitate optimal pressure dispersion against the tracheal wall The cuff must be resistant to Nitrous Oxide effects The pilot cuff must have a one-way valve and must indicate the degree of inflation accurately. All components must be latex-free sterile and individually packed in peel pouch that is easy to open - For SINGLE use only To comply with the latest ISO 5361:2016 Information brochure or the outer package of the item should clearly indicate latex free | Taper-Cuffed Oral/Nasal endotracheal tube: ID: 7.5mm OD: 10.2mm
134 | RT13-21-007 | ENDOTRACHEAL TUBE | ENDOTRACHEAL TUBE | TAPER-CUFFED ENDOTRACHEAL TUBE Purpose: endotracheal tube with a tapered cuff for long stay oral/nasal intubation aimed to reduce micro-aspiration Consists of a clear thermo-sensitive PVC tube that is kink resistant, atraumatic to airways and have a full Magill curvature. Must have a radiopaque line and must be marked with numbered depth markings in 1cm intervals, tube size, length, ID, OD, trade name. The markings on tube must be clear and must not rub off. Must have a precision band mark just above the cuff to aid with depth of intubation and guide proper placement. Preference will be given to tubes with a second circumferential band (not more than 2cm from each other) to guide proper placement. Proximal end - secure 15mm OD circuit connector made from medical grade plastic Distal end - Atraumatic left-facing bevelled tip with Murphy eye on the opposite side Must have a high-volume-low-pressure polyurethane taper-shaped cuff with pilot cuff. Cuff thickness ± 50 microns The cuff must inflate and deflate easily and when deflated must conform closely to tube to minimise the risk of trauma during intubation and extubation. Cuff must be barrel shaped to facilitate optimal pressure dispersion against the tracheal wall The cuff must be resistant to Nitrous Oxide effects The pilot cuff must have a one-way valve and must indicate the degree of inflation accurately. All components must be latex-free sterile and individually packed in peel pouch that is easy to open - For SINGLE use only To comply with the latest ISO 5361:2016 Information brochure or the outer package of the item should clearly indicate latex free | Taper-Cuffed Oral/Nasal endotracheal tube: ID: 8.0mm OD: 10.8mm
135 | RT13-21-008 | ENDOTRACHEAL TUBE | ENDOTRACHEAL TUBE | TAPER-CUFFED ENDOTRACHEAL TUBE Purpose: endotracheal tube with a tapered cuff for long stay oral/nasal intubation aimed to reduce micro-aspiration Consists of a clear thermo-sensitive PVC tube that is kink resistant, atraumatic to airways and have a full Magill curvature. Must have a radiopaque line and must be marked with numbered depth markings in 1cm intervals, tube size, length, ID, OD, trade name. The markings on tube must be clear and must not rub off. Must have a precision band mark just above the cuff to aid with depth of intubation and guide proper placement. Preference will be given to tubes with a second circumferential band (not more than 2cm from each other) to guide proper placement. Proximal end - secure 15mm OD circuit connector made from medical grade plastic Distal end - Atraumatic left-facing bevelled tip with Murphy eye on the opposite side Must have a high-volume-low-pressure polyurethane taper-shaped cuff with pilot cuff. Cuff thickness ± 50 microns The cuff must inflate and deflate easily and when deflated must conform closely to tube to minimise the risk of trauma during intubation and extubation. Cuff must be barrel shaped to facilitate optimal pressure dispersion against the tracheal wall The cuff must be resistant to Nitrous Oxide effects The pilot cuff must have a one-way valve and must indicate the degree of inflation accurately. All components must be latex-free sterile and individually packed in peel pouch that is easy to open - For SINGLE use only To comply with the latest ISO 5361:2016 Information brochure or the outer package of the item should clearly indicate latex free | Taper-Cuffed Oral/Nasal endotracheal tube: ID: 8.5mm OD: 11.4mm
136 | RT13-21-009 | ENDOTRACHEAL TUBE | ENDOTRACHEAL TUBE | TAPER-CUFFED ENDOTRACHEAL TUBE Purpose: endotracheal tube with a tapered cuff for long stay oral/nasal intubation aimed to reduce micro-aspiration Consists of a clear thermo-sensitive PVC tube that is kink resistant, atraumatic to airways and have a full Magill curvature. Must have a radiopaque line and must be marked with numbered depth markings in 1cm intervals, tube size, length, ID, OD, trade name. The markings on tube must be clear and must not rub off. Must have a precision band mark just above the cuff to aid with depth of intubation and guide proper placement. Preference will be given to tubes with a second circumferential band (not more than 2cm from each other) to guide proper placement. Proximal end - secure 15mm OD circuit connector made from medical grade plastic Distal end - Atraumatic left-facing bevelled tip with Murphy eye on the opposite side Must have a high-volume-low-pressure polyurethane taper-shaped cuff with pilot cuff. Cuff thickness ± 50 microns The cuff must inflate and deflate easily and when deflated must conform closely to tube to minimise the risk of trauma during intubation and extubation. Cuff must be barrel shaped to facilitate optimal pressure dispersion against the tracheal wall The cuff must be resistant to Nitrous Oxide effects The pilot cuff must have a one-way valve and must indicate the degree of inflation accurately. All components must be latex-free sterile and individually packed in peel pouch that is easy to open - For SINGLE use only To comply with the latest ISO 5361:2016 Information brochure or the outer package of the item should clearly indicate latex free | Taper-Cuffed Oral/Nasal endotracheal tube: ID: 9.0mm OD: 12.1mm
137 | RT13-21-010 | ENDOTRACHEAL TUBE | TAPER-CUFFED ORAL/NASAL ENDOTRACHEAL TUBE WITH EVACUATION LUMEN AT THE SUBGLOTTIC SPACE | TAPER-CUFFED ORAL/NASAL ENDOTRACHEAL TUBE WITH EVACUATION LUMEN AT THE SUBGLOTTIC SPACE Purpose: endotracheal tube with a tapered cuff for long stay oral/nasal intubation aimed to reduce micro-aspiration Consists of a clear thermo-sensitive PVC tube that is kink resistant, atraumatic to airways and have a full Magill curvature. Must have a radiopaque line and must be marked with numbered depth markings in 1cm intervals, tube size, length, ID, OD, trade name. The markings on tube must be clear and must not rub off. Must have a precision band mark just above the cuff to aid with depth of intubation and guide proper placement. Preference will be given to tubes with a second circumferential band (not more than 2cm from each other) to guide proper placement. Proximal end - secure 15mm OD circuit connector made from medical grade plastic Distal end - Atraumatic left-facing bevelled tip with Murphy eye on the opposite side Must have a high-volume-low-pressure polyurethane taper-shaped cuff with pilot cuff. Cuff thickness ± 50 microns The integrated evacuation lumen above the tapered cuff should be on the posterior aspect of the tube for the suction of subglottic secretions. The suction lumen must have a toggled capped The cuff must inflate and deflate easily and when deflated must conform closely to tube to minimise the risk of trauma during intubation and extubation. Cuff must be barrel shaped to facilitate optimal pressure dispersion against the tracheal wall The cuff must be resistant to Nitrous Oxide effects The pilot cuff must have a one-way valve and must indicate the degree of inflation accurately. All components must be latex-free sterile and individually packed in peel pouch that is easy to open - For SINGLE use only To comply with the latest ISO 5361:2016 Information brochure or the outer package of the item should clearly indicate latex free | Taper-Cuffed Oral/Nasal endotracheal tube with Evacuation Lumen: ID: 6.0mm OD: 9.0mm
138 | RT13-21-011 | ENDOTRACHEAL TUBE | TAPER-CUFFED ORAL/NASAL ENDOTRACHEAL TUBE WITH EVACUATION LUMEN AT THE SUBGLOTTIC SPACE | TAPER-CUFFED ORAL/NASAL ENDOTRACHEAL TUBE WITH EVACUATION LUMEN AT THE SUBGLOTTIC SPACE Purpose: endotracheal tube with a tapered cuff for long stay oral/nasal intubation aimed to reduce micro-aspiration Consists of a clear thermo-sensitive PVC tube that is kink resistant, atraumatic to airways and have a full Magill curvature. Must have a radiopaque line and must be marked with numbered depth markings in 1cm intervals, tube size, length, ID, OD, trade name. The markings on tube must be clear and must not rub off. Must have a precision band mark just above the cuff to aid with depth of intubation and guide proper placement. Preference will be given to tubes with a second circumferential band (not more than 2cm from each other) to guide proper placement. Proximal end - secure 15mm OD circuit connector made from medical grade plastic Distal end - Atraumatic left-facing bevelled tip with Murphy eye on the opposite side Must have a high-volume-low-pressure polyurethane taper-shaped cuff with pilot cuff. Cuff thickness ± 50 microns The integrated evacuation lumen above the tapered cuff should be on the posterior aspect of the tube for the suction of subglottic secretions. The suction lumen must have a toggled capped The cuff must inflate and deflate easily and when deflated must conform closely to tube to minimise the risk of trauma during intubation and extubation. Cuff must be barrel shaped to facilitate optimal pressure dispersion against the tracheal wall The cuff must be resistant to Nitrous Oxide effects The pilot cuff must have a one-way valve and must indicate the degree of inflation accurately. All components must be latex-free sterile and individually packed in peel pouch that is easy to open - For SINGLE use only To comply with the latest ISO 5361 Size 6mm - 8mm considered in series Size 8.0mm, 8.5mm and 9.0mm considered individually | Taper-Cuffed Oral/Nasal endotracheal tube with Evacuation Lumen: ID: 6.5mm OD: 9.8mm
139 | RT13-21-012 | ENDOTRACHEAL TUBE | TAPER-CUFFED ORAL/NASAL ENDOTRACHEAL TUBE WITH EVACUATION LUMEN AT THE SUBGLOTTIC SPACE | TAPER-CUFFED ORAL/NASAL ENDOTRACHEAL TUBE WITH EVACUATION LUMEN AT THE SUBGLOTTIC SPACE Purpose: endotracheal tube with a tapered cuff for long stay oral/nasal intubation aimed to reduce micro-aspiration Consists of a clear thermo-sensitive PVC tube that is kink resistant, atraumatic to airways and have a full Magill curvature. Must have a radiopaque line and must be marked with numbered depth markings in 1cm intervals, tube size, length, ID, OD, trade name. The markings on tube must be clear and must not rub off. Must have a precision band mark just above the cuff to aid with depth of intubation and guide proper placement. Preference will be given to tubes with a second circumferential band (not more than 2cm from each other) to guide proper placement. Proximal end - secure 15mm OD circuit connector made from medical grade plastic Distal end - Atraumatic left-facing bevelled tip with Murphy eye on the opposite side Must have a high-volume-low-pressure polyurethane taper-shaped cuff with pilot cuff. Cuff thickness ± 50 microns The integrated evacuation lumen above the tapered cuff should be on the posterior aspect of the tube for the suction of subglottic secretions. The suction lumen must have a toggled capped The cuff must inflate and deflate easily and when deflated must conform closely to tube to minimise the risk of trauma during intubation and extubation. Cuff must be barrel shaped to facilitate optimal pressure dispersion against the tracheal wall The cuff must be resistant to Nitrous Oxide effects The pilot cuff must have a one-way valve and must indicate the degree of inflation accurately. All components must be latex-free sterile and individually packed in peel pouch that is easy to open - For SINGLE use only To comply with the latest ISO 5361 Information brochure or the outer package of the item should clearly indicate latex free | Taper-Cuffed Oral/Nasal endotracheal tube with Evacuation Lumen: ID: 7.0mm OD: 10.4mm
140 | RT13-21-013 | ENDOTRACHEAL TUBE | TAPER-CUFFED ORAL/NASAL ENDOTRACHEAL TUBE WITH EVACUATION LUMEN AT THE SUBGLOTTIC SPACE | TAPER-CUFFED ORAL/NASAL ENDOTRACHEAL TUBE WITH EVACUATION LUMEN AT THE SUBGLOTTIC SPACE Purpose: endotracheal tube with a tapered cuff for long stay oral/nasal intubation aimed to reduce micro-aspiration Consists of a clear thermo-sensitive PVC tube that is kink resistant, atraumatic to airways and have a full Magill curvature. Must have a radiopaque line and must be marked with numbered depth markings in 1cm intervals, tube size, length, ID, OD, trade name. The markings on tube must be clear and must not rub off. Must have a precision band mark just above the cuff to aid with depth of intubation and guide proper placement. Preference will be given to tubes with a second circumferential band (not more than 2cm from each other) to guide proper placement. Proximal end - secure 15mm OD circuit connector made from medical grade plastic Distal end - Atraumatic left-facing bevelled tip with Murphy eye on the opposite side Must have a high-volume-low-pressure polyurethane taper-shaped cuff with pilot cuff. Cuff thickness ± 50 microns The integrated evacuation lumen above the tapered cuff should be on the posterior aspect of the tube for the suction of subglottic secretions. The suction lumen must have a toggled capped The cuff must inflate and deflate easily and when deflated must conform closely to tube to minimise the risk of trauma during intubation and extubation. Cuff must be barrel shaped to facilitate optimal pressure dispersion against the tracheal wall The cuff must be resistant to Nitrous Oxide effects The pilot cuff must have a one-way valve and must indicate the degree of inflation accurately. All components must be latex-free sterile and individually packed in peel pouch that is easy to open - For SINGLE use only To comply with the latest ISO 5361:2016 Information brochure or the outer package of the item should clearly indicate latex free | Taper-Cuffed Oral/Nasal endotracheal tube with Evacuation Lumen: ID: 7.5mm OD: 11.2mm
141 | RT13-21-014 | ENDOTRACHEAL TUBE | TAPER-CUFFED ORAL/NASAL ENDOTRACHEAL TUBE WITH EVACUATION LUMEN AT THE SUBGLOTTIC SPACE | TAPER-CUFFED ORAL/NASAL ENDOTRACHEAL TUBE WITH EVACUATION LUMEN AT THE SUBGLOTTIC SPACE Purpose: endotracheal tube with a tapered cuff for long stay oral/nasal intubation aimed to reduce micro-aspiration Consists of a clear thermo-sensitive PVC tube that is kink resistant, atraumatic to airways and have a full Magill curvature. Must have a radiopaque line and must be marked with numbered depth markings in 1cm intervals, tube size, length, ID, OD, trade name. The markings on tube must be clear and must not rub off. Must have a precision band mark just above the cuff to aid with depth of intubation and guide proper placement. Preference will be given to tubes with a second circumferential band (not more than 2cm from each other) to guide proper placement. Proximal end - secure 15mm OD circuit connector made from medical grade plastic Distal end - Atraumatic left-facing bevelled tip with Murphy eye on the opposite side Must have a high-volume-low-pressure polyurethane taper-shaped cuff with pilot cuff. Cuff thickness ± 50 microns The integrated evacuation lumen above the tapered cuff should be on the posterior aspect of the tube for the suction of subglottic secretions. The suction lumen must have a toggled capped The cuff must inflate and deflate easily and when deflated must conform closely to tube to minimise the risk of trauma during intubation and extubation. Cuff must be barrel shaped to facilitate optimal pressure dispersion against the tracheal wall The cuff must be resistant to Nitrous Oxide effects The pilot cuff must have a one-way valve and must indicate the degree of inflation accurately. All components must be latex-free sterile and individually packed in peel pouch that is easy to open - For SINGLE use only To comply with the latest ISO 5361 Information brochure or the outer package of the item should clearly indicate latex free | Taper-Cuffed Oral/Nasal endotracheal tube with Evacuation Lumen: ID: 8.0mm OD: 11.8mm
142 | RT13-22-001 | Cuffed and Uncuffed Tubes | UNCUFFED NASAL NORTH-FACING | UNCUFFED NASAL NORTH-FACING PREFORMED (RAE) ENDOTRACHEAL TUBE - NEWBORN Purpose: Short term oral intubation where tube is North facing and removed from surgical field Consists of a clear thermo-sensitive PVC tube that is kink resistant, with a North polar preformed curve with a soft flexible pointed tip and two Murphy eyes Must be able to temporarily straightened the tube to allow the passage of a suction catheter Must have a radiopaque line and must be marked with numbered depth markings in 1cm intervals, tube size, length, ID, OD, trade name. The markings on tube must be clear and must not rub off. Must also have marks/reference lines proximal to cuff to aid with depth of intubation and guide proper placement. Solid or circumferential precision band markings accepted Preference will be given to tubes with a second circumferential band (not more than 1cm from each other) to guide proper placement. Must also have a distinct mark at preformed curve to aid positioning Proximal end - secure 15mm OD: circuit connector made from medical grade plastic Distal end - Atraumatic left-facing bevelled tip NO or 1 Murphy eyes on the opposite side All components must be latex-free sterile and individually packed in peel pouch that is easy to open - For SINGLE use only To comply with the latest ISO 5361 Information brochure or the outer package of the item should clearly indicate latex free | Uncuffed Nasal North-Facing Preformed (RAE) Endotracheal Tube - Newborn ID: 2.0mm OD 4.0mm Length ± 115mm
143 | RT13-22-002 | Cuffed and Uncuffed Tubes | UNCUFFED NASAL NORTH-FACING | UNCUFFED NASAL NORTH-FACING PREFORMED (RAE) ENDOTRACHEAL TUBE - NEWBORN Purpose: Short term oral intubation where tube is North facing and removed from surgical field Consists of a clear thermo-sensitive PVC tube that is kink resistant, with a North polar preformed curve with a soft flexible pointed tip and two Murphy eyes Must be able to temporarily straightened the tube to allow the passage of a suction catheter Must have a radiopaque line and must be marked with numbered depth markings in 1cm intervals, tube size, length, ID, OD, trade name. The markings on tube must be clear and must not rub off. Must also have marks/reference lines proximal to cuff to aid with depth of intubation and guide proper placement. Solid or circumferential precision band markings accepted Preference will be given to tubes with a second circumferential band (not more than 1cm from each other) to guide proper placement. Must also have a distinct mark at preformed curve to aid positioning Proximal end - secure 15mm OD: circuit connector made from medical grade plastic Distal end - Atraumatic left-facing bevelled tip NO or 1 Murphy eyes on the opposite side All components must be latex-free sterile and individually packed in peel pouch that is easy to open - For SINGLE use only To comply with the latest ISO 5361 Information brochure or the outer package of the item should clearly indicate latex free | Uncuffed Nasal North-Facing Preformed (RAE) Endotracheal Tube - Newborn ID: 2.5mm OD 4.0mm Length ± 115mm
144 | RT13-22-003 | Cuffed and Uncuffed Tubes | UNCUFFED NASAL NORTH-FACING | UNCUFFED NASAL NORTH-FACING PREFORMED (RAE) ENDOTRACHEAL TUBE - PAEDIATRIC Purpose: Short term oral intubation where tube is North facing and removed from surgical field Consists of a clear thermo-sensitive PVC tube that is kink resistant, with a North polar preformed curve with a soft flexible pointed tip and two Murphy eyes Must be able to temporarily straightened the tube to allow the passage of a suction catheter Must have a radiopaque line and must be marked with numbered depth markings in 1cm intervals, tube size, length, ID, OD, trade name. The markings on tube must be clear and must not rub off. Must also have marks/reference lines proximal to cuff to aid with depth of intubation and guide proper placement. Solid or circumferential precision band markings accepted Preference will be given to tubes with a second circumferential band (not more than 1cm from each other) to guide proper placement. Must also have a distinct mark at preformed curve to aid positioning Proximal end - secure 15mm OD: circuit connector made from medical grade plastic Distal end - Atraumatic left-facing bevelled tip 1 Murphy eye on the opposite side All components must be latex-free sterile and individually packed in peel pouch that is easy to open - For SINGLE use only To comply with the latest ISO 5361 Information brochure or the outer package of the item should clearly indicate latex free | Uncuffed Nasal North-Facing Preformed (RAE) Endotracheal Tube - Paediatric ID: 3.0mm OD 4.0mm Length ± 115mm
145 | RT13-22-004 | Cuffed and Uncuffed Tubes | UNCUFFED NASAL NORTH-FACING | UNCUFFED NASAL NORTH-FACING PREFORMED (RAE) ENDOTRACHEAL TUBE - PAEDIATRIC Purpose: Short term oral intubation where tube is North facing and removed from surgical field Consists of a clear thermo-sensitive PVC tube that is kink resistant, with a North polar preformed curve with a soft flexible pointed tip and two Murphy eyes Must be able to temporarily straightened the tube to allow the passage of a suction catheter Must have a radiopaque line and must be marked with numbered depth markings in 1cm intervals, tube size, length, ID, OD, trade name. The markings on tube must be clear and must not rub off. Must also have marks/reference lines proximal to cuff to aid with depth of intubation and guide proper placement. Solid or circumferential precision band markings accepted Preference will be given to tubes with a second circumferential band (not more than 1cm from each other) to guide proper placement. Must also have a distinct mark at preformed curve to aid positioning Proximal end - secure 15mm OD: circuit connector made from medical grade plastic Distal end - Atraumatic left-facing bevelled tip 1 Murphy eye on the opposite side All components must be latex-free sterile and individually packed in peel pouch that is easy to open - For SINGLE use only To comply with the latest ISO 5361 Information brochure or the outer package of the item should clearly indicate latex free | Uncuffed Nasal North-Facing Preformed (RAE) Endotracheal Tube - Paediatric ID: 3.5mm OD 5.3mm Length ± 155mm
146 | RT13-22-005 | Cuffed and Uncuffed Tubes | UNCUFFED NASAL NORTH-FACING | UNCUFFED NASAL NORTH-FACING PREFORMED (RAE) ENDOTRACHEAL TUBE - PAEDIATRIC Purpose: Short term oral intubation where tube is North facing and removed from surgical field Consists of a clear thermo-sensitive PVC tube that is kink resistant, with a North polar preformed curve with a soft flexible pointed tip and two Murphy eyes Must be able to temporarily straightened the tube to allow the passage of a suction catheter Must have a radiopaque line and must be marked with numbered depth markings in 1cm intervals, tube size, length, ID, OD, trade name. The markings on tube must be clear and must not rub off. Must also have marks/reference lines proximal to cuff to aid with depth of intubation and guide proper placement. Solid or circumferential precision band markings accepted Preference will be given to tubes with a second circumferential band (not more than 1cm from each other) to guide proper placement. Must also have a distinct mark at preformed curve to aid positioning Proximal end - secure 15mm OD: circuit connector made from medical grade plastic Distal end - Atraumatic left-facing bevelled tip 1 Murphy eye on the opposite side All components must be latex-free sterile and individually packed in peel pouch that is easy to open - For SINGLE use only To comply with the latest ISO 5361 Information brochure or the outer package of the item should clearly indicate latex free | Uncuffed Nasal North-Facing Preformed (RAE) Endotracheal Tube - Paediatric ID: 4.0mm OD 5.3mm Length ± 155mm
147 | RT13-22-006 | Cuffed and Uncuffed Tubes | UNCUFFED NASAL NORTH-FACING | UNCUFFED NASAL NORTH-FACING PREFORMED (RAE) ENDOTRACHEAL TUBE - PAEDIATRIC Purpose: Short term oral intubation where tube is North facing and removed from surgical field Consists of a clear thermo-sensitive PVC tube that is kink resistant, with a North polar preformed curve with a soft flexible pointed tip and two Murphy eyes Must be able to temporarily straightened the tube to allow the passage of a suction catheter Must have a radiopaque line and must be marked with numbered depth markings in 1cm intervals, tube size, length, ID, OD, trade name. The markings on tube must be clear and must not rub off. Must also have marks/reference lines proximal to cuff to aid with depth of intubation and guide proper placement. Solid or circumferential precision band markings accepted Preference will be given to tubes with a second circumferential band (not more than 1cm from each other) to guide proper placement. Must also have a distinct mark at preformed curve to aid positioning Proximal end - secure 15mm OD: circuit connector made from medical grade plastic Distal end - Atraumatic left-facing bevelled tip 1 Murphy eye on the opposite side All components must be latex-free sterile and individually packed in peel pouch that is easy to open - For SINGLE use only To comply with the latest ISO 5361 Information brochure or the outer package of the item should clearly indicate latex free | Uncuffed Nasal North-Facing Preformed (RAE) Endotracheal Tube - Paediatric ID: 4.5mm OD 6.0mm Length ± 180mm
148 | RT13-22-007 | Cuffed and Uncuffed Tubes | UNCUFFED NASAL NORTH-FACING | UNCUFFED NASAL NORTH-FACING PREFORMED (RAE) ENDOTRACHEAL TUBE - PAEDIATRIC Purpose: Short term oral intubation where tube is North facing and removed from surgical field Consists of a clear thermo-sensitive PVC tube that is kink resistant, with a North polar preformed curve with a soft flexible pointed tip and two Murphy eyes Must be able to temporarily straightened the tube to allow the passage of a suction catheter Must have a radiopaque line and must be marked with numbered depth markings in 1cm intervals, tube size, length, ID, OD, trade name. The markings on tube must be clear and must not rub off. Must also have marks/reference lines proximal to cuff to aid with depth of intubation and guide proper placement. Solid or circumferential precision band markings accepted Preference will be given to tubes with a second circumferential band (not more than 1cm from each other) to guide proper placement. Must also have a distinct mark at preformed curve to aid positioning Proximal end - secure 15mm OD: circuit connector made from medical grade plastic Distal end - Atraumatic left-facing bevelled tip 1 Murphy eye on the opposite side All components must be latex-free sterile and individually packed in peel pouch that is easy to open - For SINGLE use only To comply with the latest ISO 5361 Information brochure or the outer package of the item should clearly indicate latex free | Uncuffed Nasal North-Facing Preformed (RAE) Endotracheal Tube - Paediatric ID: 5.0mm OD 6.7mm Length ± 190mm
149 | RT13-22-010 | Cuffed and Uncuffed Tubes | CUFFED NASAL NORTH-FACING | CUFFED NASAL NORTH-FACING PREFORMED (RAE) ENDOTRACHEAL TUBE - NEWBORN Purpose: Short term oral intubation where tube is North facing and removed from surgical field Consists of a clear thermo-sensitive PVC tube that is kink resistant, with a North polar preformed curve with a soft flexible pointed tip and two Murphy eyes Must be able to temporarily straightened the tube to allow the passage of a suction catheter Must have a radiopaque line and must be marked with numbered depth markings in 1cm intervals, tube size, length, ID, OD, trade name. The markings on tube must be clear and must not rub off. Must also have marks/reference lines proximal to cuff to aid with depth of intubation and guide proper placement. Solid or circumferential precision band markings accepted Preference will be given to tubes with a second circumferential band (not more than 1cm from each other) to guide proper placement. Must also have a distinct mark at preformed curve to aid positioning Proximal end - secure 15mm OD: circuit connector made from medical grade plastic Distal end - Atraumatic left-facing bevelled tip with 1 Murphy eye on the opposite side Must have a high-volume-low-pressure polyurethane barrel shaped cuff with pilot cuff. Cuff thickness ± 50 microns. The pilot cuff must have a one-way valve and must indicate the degree of inflation accurately. The cuff must inflate and deflate easily and when deflated must conform closely to tube to minimise the risk of trauma during intubation and extubation and not increase the OD of the tube. The cuff must be resistant to Nitrous Oxide effects All components must be latex-free sterile and individually packed in peel pouch that is easy to open - For SINGLE use only To comply with the latest ISO 5361 Information brochure or the outer package of the item should clearly indicate latex free | Cuffed Nasal North-Facing Preformed (RAE) Endotracheal Tube - Newborn ID: 2.0mm
150 | RT13-22-011 | Cuffed and Uncuffed Tubes | CUFFED NASAL NORTH-FACING | CUFFED NASAL NORTH-FACING PREFORMED (RAE) ENDOTRACHEAL TUBE - NEWBORN Purpose: Short term oral intubation where tube is North facing and removed from surgical field Consists of a clear thermo-sensitive PVC tube that is kink resistant, with a North polar preformed curve with a soft flexible pointed tip and two Murphy eyes Must be able to temporarily straightened the tube to allow the passage of a suction catheter Must have a radiopaque line and must be marked with numbered depth markings in 1cm intervals, tube size, length, ID, OD, trade name. The markings on tube must be clear and must not rub off. Must also have marks/reference lines proximal to cuff to aid with depth of intubation and guide proper placement. Solid or circumferential precision band markings accepted Preference will be given to tubes with a second circumferential band (not more than 1cm from each other) to guide proper placement. Must also have a distinct mark at preformed curve to aid positioning Proximal end - secure 15mm OD: circuit connector made from medical grade plastic Distal end - Atraumatic left-facing bevelled tip with 1 Murphy eye on the opposite side Must have a high-volume-low-pressure polyurethane barrel shaped cuff with pilot cuff. Cuff thickness ± 50 microns. The pilot cuff must have a one-way valve and must indicate the degree of inflation accurately. The cuff must inflate and deflate easily and when deflated must conform closely to tube to minimise the risk of trauma during intubation and extubation and not increase the OD of the tube. The cuff must be resistant to Nitrous Oxide effects All components must be latex-free sterile and individually packed in peel pouch that is easy to open - For SINGLE use only To comply with the latest ISO 5361 Information brochure or the outer package of the item should clearly indicate latex free | Cuffed Nasal North-Facing Preformed (RAE) Endotracheal Tube - Newborn ID: 2.5mm
151 | RT13-22-012 | Cuffed and Uncuffed Tubes | CUFFED NASAL NORTH-FACING | CUFFED NASAL NORTH-FACING PREFORMED (RAE) ENDOTRACHEAL TUBE - PAEDIATRIC Purpose: Short term oral intubation where tube is North facing and removed from surgical field Consists of a clear thermo-sensitive PVC tube that is kink resistant, with a North polar preformed curve with a soft flexible pointed tip and two Murphy eyes Must be able to temporarily straightened the tube to allow the passage of a suction catheter Must have a radiopaque line and must be marked with numbered depth markings in 1cm intervals, tube size, length, ID, OD, trade name. The markings on tube must be clear and must not rub off. Must also have marks/reference lines proximal to cuff to aid with depth of intubation and guide proper placement. Solid or circumferential precision band markings accepted Preference will be given to tubes with a second circumferential band (not more than 1cm from each other) to guide proper placement. Must also have a distinct mark at preformed curve to aid positioning Proximal end - secure 15mm OD: circuit connector made from medical grade plastic Distal end - Atraumatic left-facing bevelled tip with 2 Murphy eyes on the opposite side Must have a high-volume-low-pressure polyurethane barrel shaped cuff with pilot cuff. Cuff thickness ± 50 microns. The pilot cuff must have a one-way valve and must indicate the degree of inflation accurately. The cuff must inflate and deflate easily and when deflated must conform closely to tube to minimise the risk of trauma during intubation and extubation and not increase the OD of the tube. The cuff must be resistant to Nitrous Oxide effects All components must be latex-free sterile and individually packed in peel pouch that is easy to open - For SINGLE use only To comply with the latest ISO 5361 Information brochure or the outer package of the item should clearly indicate latex free | Cuffed Nasal North-Facing Preformed (RAE) Endotracheal Tube - Paediatric ID: 3.0mm OD 4.0mm Length ± 115mm
152 | RT13-22-013 | Cuffed and Uncuffed Tubes | CUFFED NASAL NORTH-FACING | CUFFED NASAL NORTH-FACING PREFORMED (RAE) ENDOTRACHEAL TUBE - PAEDIATRIC Purpose: Short term oral intubation where tube is North facing and removed from surgical field Consists of a clear thermo-sensitive PVC tube that is kink resistant, with a North polar preformed curve with a soft flexible pointed tip and two Murphy eyes Must be able to temporarily straightened the tube to allow the passage of a suction catheter Must have a radiopaque line and must be marked with numbered depth markings in 1cm intervals, tube size, length, ID, OD, trade name. The markings on tube must be clear and must not rub off. Must also have marks/reference lines proximal to cuff to aid with depth of intubation and guide proper placement. Solid or circumferential precision band markings accepted Preference will be given to tubes with a second circumferential band (not more than 1cm from each other) to guide proper placement. Must also have a distinct mark at preformed curve to aid positioning Proximal end - secure 15mm OD: circuit connector made from medical grade plastic Distal end - Atraumatic left-facing bevelled tip with 2 Murphy eyes on the opposite side Must have a high-volume-low-pressure polyurethane barrel shaped cuff with pilot cuff. Cuff thickness ± 50 microns. The pilot cuff must have a one-way valve and must indicate the degree of inflation accurately. The cuff must inflate and deflate easily and when deflated must conform closely to tube to minimise the risk of trauma during intubation and extubation and not increase the OD of the tube. The cuff must be resistant to Nitrous Oxide effects All components must be latex-free sterile and individually packed in peel pouch that is easy to open - For SINGLE use only To comply with the latest ISO 5361 Information brochure or the outer package of the item should clearly indicate latex free | Cuffed Nasal North-Facing Preformed (RAE) Endotracheal Tube - Paediatric ID: 3.5mm OD 5.3mm Length ± 155mm
153 | RT13-22-014 | Cuffed and Uncuffed Tubes | CUFFED NASAL NORTH-FACING | CUFFED NASAL NORTH-FACING PREFORMED (RAE) ENDOTRACHEAL TUBE - PAEDIATRIC Purpose: Short term oral intubation where tube is North facing and removed from surgical field Consists of a clear thermo-sensitive PVC tube that is kink resistant, with a North polar preformed curve with a soft flexible pointed tip and two Murphy eyes Must be able to temporarily straightened the tube to allow the passage of a suction catheter Must have a radiopaque line and must be marked with numbered depth markings in 1cm intervals, tube size, length, ID, OD, trade name. The markings on tube must be clear and must not rub off. Must also have marks/reference lines proximal to cuff to aid with depth of intubation and guide proper placement. Solid or circumferential precision band markings accepted Preference will be given to tubes with a second circumferential band (not more than 1cm from each other) to guide proper placement. Must also have a distinct mark at preformed curve to aid positioning Proximal end - secure 15mm OD: circuit connector made from medical grade plastic Distal end - Atraumatic left-facing bevelled tip with 2 Murphy eyes on the opposite side Must have a high-volume-low-pressure polyurethane barrel shaped cuff with pilot cuff. Cuff thickness ± 50 microns. The pilot cuff must have a one-way valve and must indicate the degree of inflation accurately. The cuff must inflate and deflate easily and when deflated must conform closely to tube to minimise the risk of trauma during intubation and extubation and not increase the OD of the tube. The cuff must be resistant to Nitrous Oxide effects All components must be latex-free sterile and individually packed in peel pouch that is easy to open - For SINGLE use only To comply with the latest ISO 5361 Information brochure or the outer package of the item should clearly indicate latex free | Cuffed Nasal North-Facing Preformed (RAE) Endotracheal Tube - Paediatric ID: 4.0mm OD 5.3mm Length ± 155mm
154 | RT13-22-015 | Cuffed and Uncuffed Tubes | CUFFED NASAL NORTH-FACING | CUFFED NASAL NORTH-FACING PREFORMED (RAE) ENDOTRACHEAL TUBE - PAEDIATRIC Purpose: Short term oral intubation where tube is North facing and removed from surgical field Consists of a clear thermo-sensitive PVC tube that is kink resistant, with a North polar preformed curve with a soft flexible pointed tip and two Murphy eyes Must be able to temporarily straightened the tube to allow the passage of a suction catheter Must have a radiopaque line and must be marked with numbered depth markings in 1cm intervals, tube size, length, ID, OD, trade name. The markings on tube must be clear and must not rub off. Must also have marks/reference lines proximal to cuff to aid with depth of intubation and guide proper placement. Solid or circumferential precision band markings accepted Preference will be given to tubes with a second circumferential band (not more than 1cm from each other) to guide proper placement. Must also have a distinct mark at preformed curve to aid positioning Proximal end - secure 15mm OD: circuit connector made from medical grade plastic Distal end - Atraumatic left-facing bevelled tip with 2 Murphy eyes on the opposite side Must have a high-volume-low-pressure polyurethane barrel shaped cuff with pilot cuff. Cuff thickness ± 50 microns. The pilot cuff must have a one-way valve and must indicate the degree of inflation accurately. The cuff must inflate and deflate easily and when deflated must conform closely to tube to minimise the risk of trauma during intubation and extubation and not increase the OD of the tube. The cuff must be resistant to Nitrous Oxide effects All components must be latex-free sterile and individually packed in peel pouch that is easy to open - For SINGLE use only To comply with the latest ISO 5361 Information brochure or the outer package of the item should clearly indicate latex free | Cuffed Nasal North-Facing Preformed (RAE) Endotracheal Tube - Paediatric ID: 4.5mm OD 6.0mm Length ± 180mm
155 | RT13-22-016 | Cuffed and Uncuffed Tubes | CUFFED NASAL NORTH-FACING | CUFFED NASAL NORTH-FACING PREFORMED (RAE) ENDOTRACHEAL TUBE - PAEDIATRIC Purpose: Short term oral intubation where tube is North facing and removed from surgical field Consists of a clear thermo-sensitive PVC tube that is kink resistant, with a North polar preformed curve with a soft flexible pointed tip and two Murphy eyes Must be able to temporarily straightened the tube to allow the passage of a suction catheter Must have a radiopaque line and must be marked with numbered depth markings in 1cm intervals, tube size, length, ID, OD, trade name. The markings on tube must be clear and must not rub off. Must also have marks/reference lines proximal to cuff to aid with depth of intubation and guide proper placement. Solid or circumferential precision band markings accepted Preference will be given to tubes with a second circumferential band (not more than 1cm from each other) to guide proper placement. Must also have a distinct mark at preformed curve to aid positioning Proximal end - secure 15mm OD: circuit connector made from medical grade plastic Distal end - Atraumatic left-facing bevelled tip with 2 Murphy eyes on the opposite side Must have a high-volume-low-pressure polyurethane barrel shaped cuff with pilot cuff. Cuff thickness ± 50 microns. The pilot cuff must have a one-way valve and must indicate the degree of inflation accurately. The cuff must inflate and deflate easily and when deflated must conform closely to tube to minimise the risk of trauma during intubation and extubation and not increase the OD of the tube. The cuff must be resistant to Nitrous Oxide effects All components must be latex-free sterile and individually packed in peel pouch that is easy to open - For SINGLE use only To comply with the latest ISO 5361 Information brochure or the outer package of the item should clearly indicate latex free | Cuffed Nasal North-Facing Preformed (RAE) Endotracheal Tube - Paediatric ID: 5.0mm OD 6.7mm Length ± 190mm
156 | RT13-22-017 | Cuffed and Uncuffed Tubes | CUFFED NASAL NORTH-FACING | CUFFED NASAL NORTH-FACING PREFORMED (RAE) ENDOTRACHEAL TUBE - PAEDIATRIC Purpose: Short term oral intubation where tube is North facing and removed from surgical field Consists of a clear thermo-sensitive PVC tube that is kink resistant, with a North polar preformed curve with a soft flexible pointed tip and two Murphy eyes Must be able to temporarily straightened the tube to allow the passage of a suction catheter Must have a radiopaque line and must be marked with numbered depth markings in 1cm intervals, tube size, length, ID, OD, trade name. The markings on tube must be clear and must not rub off. Must also have marks/reference lines proximal to cuff to aid with depth of intubation and guide proper placement. Solid or circumferential precision band markings accepted Preference will be given to tubes with a second circumferential band (not more than 1cm from each other) to guide proper placement. Must also have a distinct mark at preformed curve to aid positioning Proximal end - secure 15mm OD: circuit connector made from medical grade plastic Distal end - Atraumatic left-facing bevelled tip with 2 Murphy eyes on the opposite side Must have a high-volume-low-pressure polyurethane barrel shaped cuff with pilot cuff. Cuff thickness ± 50 microns. The pilot cuff must have a one-way valve and must indicate the degree of inflation accurately. The cuff must inflate and deflate easily and when deflated must conform closely to tube to minimise the risk of trauma during intubation and extubation and not increase the OD of the tube. The cuff must be resistant to Nitrous Oxide effects All components must be latex-free sterile and individually packed in peel pouch that is easy to open - For SINGLE use only To comply with the latest ISO 5361 Information brochure or the outer package of the item should clearly indicate latex free | Cuffed Nasal North-Facing Preformed (RAE) Endotracheal Tube - Paediatric ID: 5.5mm OD 7.3mm Length ± 195mm
157 | RT13-22-018 | Cuffed and Uncuffed Tubes | CUFFED NASAL NORTH-FACING | CUFFED NASAL NORTH-FACING PREFORMED (RAE) ENDOTRACHEAL TUBE - PAEDIATRIC Purpose: Short term oral intubation where tube is North facing and removed from surgical field Consists of a clear thermo-sensitive PVC tube that is kink resistant, with a North polar preformed curve with a soft flexible pointed tip and two Murphy eyes Must be able to temporarily straightened the tube to allow the passage of a suction catheter Must have a radiopaque line and must be marked with numbered depth markings in 1cm intervals, tube size, length, ID, OD, trade name. The markings on tube must be clear and must not rub off. Must also have marks/reference lines proximal to cuff to aid with depth of intubation and guide proper placement. Solid or circumferential precision band markings accepted Preference will be given to tubes with a second circumferential band (not more than 1cm from each other) to guide proper placement. Must also have a distinct mark at preformed curve to aid positioning Proximal end - secure 15mm OD: circuit connector made from medical grade plastic Distal end - Atraumatic left-facing bevelled tip with 2 Murphy eyes on the opposite side Must have a high-volume-low-pressure polyurethane barrel shaped cuff with pilot cuff. Cuff thickness ± 50 microns. The pilot cuff must have a one-way valve and must indicate the degree of inflation accurately. The cuff must inflate and deflate easily and when deflated must conform closely to tube to minimise the risk of trauma during intubation and extubation and not increase the OD of the tube. The cuff must be resistant to Nitrous Oxide effects All components must be latex-free sterile and individually packed in peel pouch that is easy to open - For SINGLE use only To comply with the latest ISO 5361 Information brochure or the outer package of the item should clearly indicate latex free | Cuffed Nasal North-Facing Preformed (RAE) Endotracheal Tube - Paediatric ID: 6.0mm OD 8.0mm Length ± 205mm
158 | RT13-22-019 | Cuffed and Uncuffed Tubes | CUFFED NASAL NORTH-FACING | CUFFED NASAL NORTH-FACING PREFORMED (RAE) ENDOTRACHEAL TUBE - ADULT Purpose: Short term oral intubation where tube is North facing and removed from surgical field Consists of a clear thermo-sensitive PVC tube that is kink resistant, with a North polar preformed curve with a soft flexible pointed tip and two Murphy eyes Must be able to temporarily straightened the tube to allow the passage of a suction catheter Must have a radiopaque line and must be marked with numbered depth markings in 1cm intervals, tube size, length, ID, OD, trade name. The markings on tube must be clear and must not rub off. Must also have marks/reference lines proximal to cuff to aid with depth of intubation and guide proper placement. Solid or circumferential precision band markings accepted Preference will be given to tubes with a second circumferential band (not more than 2cm from each other) to guide proper placement. Must also have a distinct mark at preformed curve to aid positioning Proximal end - secure 15mm OD: circuit connector made from medical grade plastic Distal end - Atraumatic left-facing bevelled tip with 2 Murphy eyes on the opposite side Must have a high-volume-low-pressure polyurethane barrel shaped cuff with pilot cuff. Cuff thickness ± 50 microns. The pilot cuff must have a one-way valve and must indicate the degree of inflation accurately. The cuff must inflate and deflate easily and when deflated must conform closely to tube to minimise the risk of trauma during intubation and extubation and not increase the OD of the tube. The cuff must be resistant to Nitrous Oxide effects All components must be latex-free sterile and individually packed in peel pouch that is easy to open - For SINGLE use only To comply with the latest ISO 5361 Information brochure or the outer package of the item should clearly indicate latex free | Cuffed Nasal North-Facing Preformed (RAE) Endotracheal Tube - Adult ID: 6.5mm
159 | RT13-22-020 | Cuffed and Uncuffed Tubes | CUFFED NASAL NORTH-FACING | CUFFED NASAL NORTH-FACING PREFORMED (RAE) ENDOTRACHEAL TUBE - ADULT Purpose: Short term oral intubation where tube is North facing and removed from surgical field Consists of a clear thermo-sensitive PVC tube that is kink resistant, with a North polar preformed curve with a soft flexible pointed tip and two Murphy eyes Must be able to temporarily straightened the tube to allow the passage of a suction catheter Must have a radiopaque line and must be marked with numbered depth markings in 1cm intervals, tube size, length, ID, OD, trade name. The markings on tube must be clear and must not rub off. Must also have marks/reference lines proximal to cuff to aid with depth of intubation and guide proper placement. Solid or circumferential precision band markings accepted Preference will be given to tubes with a second circumferential band (not more than 2cm from each other) to guide proper placement. Must also have a distinct mark at preformed curve to aid positioning Proximal end - secure 15mm OD: circuit connector made from medical grade plastic Distal end - Atraumatic left-facing bevelled tip with 2 Murphy eyes on the opposite side Must have a high-volume-low-pressure polyurethane barrel shaped cuff with pilot cuff. Cuff thickness ± 50 microns. The pilot cuff must have a one-way valve and must indicate the degree of inflation accurately. The cuff must inflate and deflate easily and when deflated must conform closely to tube to minimise the risk of trauma during intubation and extubation and not increase the OD of the tube. The cuff must be resistant to Nitrous Oxide effects All components must be latex-free sterile and individually packed in peel pouch that is easy to open - For SINGLE use only To comply with the latest ISO 5361 Information brochure or the outer package of the item should clearly indicate latex free | Cuffed Nasal North-Facing Preformed (RAE) Endotracheal Tube - Adult ID: 7.0mm
160 | RT13-22-021 | Cuffed and Uncuffed Tubes | CUFFED NASAL NORTH-FACING | CUFFED NASAL NORTH-FACING PREFORMED (RAE) ENDOTRACHEAL TUBE - ADULT Purpose: Short term oral intubation where tube is North facing and removed from surgical field Consists of a clear thermo-sensitive PVC tube that is kink resistant, with a North polar preformed curve with a soft flexible pointed tip and two Murphy eyes Must be able to temporarily straightened the tube to allow the passage of a suction catheter Must have a radiopaque line and must be marked with numbered depth markings in 1cm intervals, tube size, length, ID, OD, trade name. The markings on tube must be clear and must not rub off. Must also have marks/reference lines proximal to cuff to aid with depth of intubation and guide proper placement. Solid or circumferential precision band markings accepted Preference will be given to tubes with a second circumferential band (not more than 2cm from each other) to guide proper placement. Must also have a distinct mark at preformed curve to aid positioning Proximal end - secure 15mm OD: circuit connector made from medical grade plastic Distal end - Atraumatic left-facing bevelled tip with 2 Murphy eyes on the opposite side Must have a high-volume-low-pressure polyurethane barrel shaped cuff with pilot cuff. Cuff thickness ± 50 microns. The pilot cuff must have a one-way valve and must indicate the degree of inflation accurately. The cuff must inflate and deflate easily and when deflated must conform closely to tube to minimise the risk of trauma during intubation and extubation and not increase the OD of the tube. The cuff must be resistant to Nitrous Oxide effects All components must be latex-free sterile and individually packed in peel pouch that is easy to open - For SINGLE use only To comply with the latest ISO 5361 Information brochure or the outer package of the item should clearly indicate latex free | Cuffed Nasal North-Facing Preformed (RAE) Endotracheal Tube - Adult ID: 7.5mm
161 | RT13-22-022 | Cuffed and Uncuffed Tubes | CUFFED NASAL NORTH-FACING | CUFFED NASAL NORTH-FACING PREFORMED (RAE) ENDOTRACHEAL TUBE - ADULT Purpose: Short term oral intubation where tube is North facing and removed from surgical field Consists of a clear thermo-sensitive PVC tube that is kink resistant, with a North polar preformed curve with a soft flexible pointed tip and two Murphy eyes Must be able to temporarily straightened the tube to allow the passage of a suction catheter Must have a radiopaque line and must be marked with numbered depth markings in 1cm intervals, tube size, length, ID, OD, trade name. The markings on tube must be clear and must not rub off. Must also have marks/reference lines proximal to cuff to aid with depth of intubation and guide proper placement. Solid or circumferential precision band markings accepted Preference will be given to tubes with a second circumferential band (not more than 2cm from each other) to guide proper placement. Must also have a distinct mark at preformed curve to aid positioning Proximal end - secure 15mm OD: circuit connector made from medical grade plastic Distal end - Atraumatic left-facing bevelled tip with 2 Murphy eyes on the opposite side Must have a high-volume-low-pressure polyurethane barrel shaped cuff with pilot cuff. Cuff thickness ± 50 microns. The pilot cuff must have a one-way valve and must indicate the degree of inflation accurately. The cuff must inflate and deflate easily and when deflated must conform closely to tube to minimise the risk of trauma during intubation and extubation and not increase the OD of the tube. The cuff must be resistant to Nitrous Oxide effects All components must be latex-free sterile and individually packed in peel pouch that is easy to open - For SINGLE use only To comply with the latest ISO 5361 Information brochure or the outer package of the item should clearly indicate latex free | Cuffed Nasal North-Facing Preformed (RAE) Endotracheal Tube - Adult ID: 8.0mm
162 | RT13-23-001 | Cuffed and Uncuffed Tubes | UNCUFFED ORAL SOUTH-FACING | UNCUFFED ORAL SOUTH-FACING PREFORMED (RAE) ENDOTRACHEAL TUBE WITH 1 OR 2 MURPHY EYE Purpose: Short term oral intubation where tube is South facing and removed from surgical field Consists of a clear thermo-sensitive PVC tube that is kink resistant, with a South polar preformed curve with 1 or 2 Murphy eye Must be able to temporarily straightened the tube to allow the passage of a suction catheter Must have a radiopaque line and must be marked with numbered depth markings in 1cm intervals, tube size, length, ID, OD, trade name. The markings on tube must be clear and must not rub off. Must also have marks/reference lines proximal to cuff to aid with depth of intubation and guide proper placement. Preference will be given to tubes with a second circumferential band (not more than 1cm from each other) to guide proper placement. Must also have a distinct mark at preformed curve to aid positioning Proximal end - secure 15mm OD circuit connector made from medical grade plastic Distal end - Atraumatic left-facing bevelled tip with Murphy eye on the opposite side All components must be latex-free sterile and individually packed in peel pouch that is easy to open - For SINGLE use only To comply with the latest ISO 5361 Information brochure or the outer package of the item should clearly indicate latex free | Uncuffed Oral South-Facing Preformed (RAE) Endotracheal Tube with Murphy eye. ID: 2.0mm
163 | RT13-23-002 | Cuffed and Uncuffed Tubes | UNCUFFED ORAL SOUTH-FACING | UNCUFFED ORAL SOUTH-FACING PREFORMED (RAE) ENDOTRACHEAL TUBE WITH 1 OR 2 MURPHY EYE Purpose: Short term oral intubation where tube is South facing and removed from surgical field Consists of a clear thermo-sensitive PVC tube that is kink resistant, with a South polar preformed curve with 1 or 2 Murphy eye Must be able to temporarily straightened the tube to allow the passage of a suction catheter Must have a radiopaque line and must be marked with numbered depth markings in 1cm intervals, tube size, length, ID, OD, trade name. The markings on tube must be clear and must not rub off. Must also have marks/reference lines proximal to cuff to aid with depth of intubation and guide proper placement. Preference will be given to tubes with a second circumferential band (not more than 1cm from each other) to guide proper placement. Must also have a distinct mark at preformed curve to aid positioning Proximal end - secure 15mm OD circuit connector made from medical grade plastic Distal end - Atraumatic left-facing bevelled tip with Murphy eye on the opposite side All components must be latex-free sterile and individually packed in peel pouch that is easy to open - For SINGLE use only To comply with the latest ISO 5361 Information brochure or the outer package of the item should clearly indicate latex free | Uncuffed Oral South-Facing Preformed (RAE) Endotracheal Tube with Murphy eye. ID: 2.5mm
164 | RT13-23-003 | Cuffed and Uncuffed Tubes | UNCUFFED ORAL SOUTH-FACING | UNCUFFED ORAL SOUTH-FACING PREFORMED (RAE) ENDOTRACHEAL TUBE WITH 1 OR 2 MURPHY EYE Purpose: Short term oral intubation where tube is South facing and removed from surgical field Consists of a clear thermo-sensitive PVC tube that is kink resistant, with a South polar preformed curve with 1 or 2 Murphy eye Must be able to temporarily straightened the tube to allow the passage of a suction catheter Must have a radiopaque line and must be marked with numbered depth markings in 1cm intervals, tube size, length, ID, OD, trade name. The markings on tube must be clear and must not rub off. Must also have marks/reference lines proximal to cuff to aid with depth of intubation and guide proper placement. Preference will be given to tubes with a second circumferential band (not more than 1cm from each other) to guide proper placement. Must also have a distinct mark at preformed curve to aid positioning Proximal end - secure 15mm OD circuit connector made from medical grade plastic Distal end - Atraumatic left-facing bevelled tip with Murphy eye on the opposite side All components must be latex-free sterile and individually packed in peel pouch that is easy to open - For SINGLE use only To comply with the latest ISO 5361 Information brochure or the outer package of the item should clearly indicate latex free | Uncuffed Oral South-Facing Preformed (RAE) Endotracheal Tube with Murphy eye. ID: 3.0mm OD 4.0mm Length ± 105mm
165 | RT13-23-004 | Cuffed and Uncuffed Tubes | UNCUFFED ORAL SOUTH-FACING | UNCUFFED ORAL SOUTH-FACING PREFORMED (RAE) ENDOTRACHEAL TUBE WITH 1 OR 2 MURPHY EYE Purpose: Short term oral intubation where tube is South facing and removed from surgical field Consists of a clear thermo-sensitive PVC tube that is kink resistant, with a South polar preformed curve with 1 or 2 Murphy eye Must be able to temporarily straightened the tube to allow the passage of a suction catheter Must have a radiopaque line and must be marked with numbered depth markings in 1cm intervals, tube size, length, ID, OD, trade name. The markings on tube must be clear and must not rub off. Must also have marks/reference lines proximal to cuff to aid with depth of intubation and guide proper placement. Preference will be given to tubes with a second circumferential band (not more than 1cm from each other) to guide proper placement. Must also have a distinct mark at preformed curve to aid positioning Proximal end - secure 15mm OD circuit connector made from medical grade plastic Distal end - Atraumatic left-facing bevelled tip with Murphy eye on the opposite side All components must be latex-free sterile and individually packed in peel pouch that is easy to open - For SINGLE use only To comply with the latest ISO 5361 Information brochure or the outer package of the item should clearly indicate latex free | Uncuffed Oral South-Facing Preformed (RAE) Endotracheal Tube with Murphy eye. ID: 3.5mm OD 4.7mm Length ± 120mm
166 | RT13-23-005 | Cuffed and Uncuffed Tubes | UNCUFFED ORAL SOUTH-FACING | UNCUFFED ORAL SOUTH-FACING PREFORMED (RAE) ENDOTRACHEAL TUBE WITH 1 OR 2 MURPHY EYE Purpose: Short term oral intubation where tube is South facing and removed from surgical field Consists of a clear thermo-sensitive PVC tube that is kink resistant, with a South polar preformed curve with 1 or 2 Murphy eye Must be able to temporarily straightened the tube to allow the passage of a suction catheter Must have a radiopaque line and must be marked with numbered depth markings in 1cm intervals, tube size, length, ID, OD, trade name. The markings on tube must be clear and must not rub off. Must also have marks/reference lines proximal to cuff to aid with depth of intubation and guide proper placement. Preference will be given to tubes with a second circumferential band (not more than 1cm from each other) to guide proper placement. Must also have a distinct mark at preformed curve to aid positioning Proximal end - secure 15mm OD circuit connector made from medical grade plastic Distal end - Atraumatic left-facing bevelled tip with Murphy eye on the opposite side All components must be latex-free sterile and individually packed in peel pouch that is easy to open - For SINGLE use only To comply with the latest ISO 5361 Information brochure or the outer package of the item should clearly indicate latex free | Uncuffed Oral South-Facing Preformed (RAE) Endotracheal Tube with Murphy eye. ID: 4.0mm OD 5.3mm Length ± 130mm
167 | RT13-23-006 | Cuffed and Uncuffed Tubes | UNCUFFED ORAL SOUTH-FACING | UNCUFFED ORAL SOUTH-FACING PREFORMED (RAE) ENDOTRACHEAL TUBE WITH 1 OR 2 MURPHY EYE Purpose: Short term oral intubation where tube is South facing and removed from surgical field Consists of a clear thermo-sensitive PVC tube that is kink resistant, with a South polar preformed curve with 1 or 2 Murphy eye Must be able to temporarily straightened the tube to allow the passage of a suction catheter Must have a radiopaque line and must be marked with numbered depth markings in 1cm intervals, tube size, length, ID, OD, trade name. The markings on tube must be clear and must not rub off. Must also have marks/reference lines proximal to cuff to aid with depth of intubation and guide proper placement. Preference will be given to tubes with a second circumferential band (not more than 1cm from each other) to guide proper placement. Must also have a distinct mark at preformed curve to aid positioning Proximal end - secure 15mm OD circuit connector made from medical grade plastic Distal end - Atraumatic left-facing bevelled tip with Murphy eye on the opposite side All components must be latex-free sterile and individually packed in peel pouch that is easy to open - For SINGLE use only To comply with the latest ISO 5361 Information brochure or the outer package of the item should clearly indicate latex free | Uncuffed Oral South-Facing Preformed (RAE) Endotracheal Tube with Murphy eye. ID: 4.5mm OD 6.0mm Length ± 145mm
168 | RT13-23-007 | Cuffed and Uncuffed Tubes | UNCUFFED ORAL SOUTH-FACING | UNCUFFED ORAL SOUTH-FACING PREFORMED (RAE) ENDOTRACHEAL TUBE WITH 1 OR 2 MURPHY EYE Purpose: Short term oral intubation where tube is South facing and removed from surgical field Consists of a clear thermo-sensitive PVC tube that is kink resistant, with a South polar preformed curve with 1 or 2 Murphy eye Must be able to temporarily straightened the tube to allow the passage of a suction catheter Must have a radiopaque line and must be marked with numbered depth markings in 1cm intervals, tube size, length, ID, OD, trade name. The markings on tube must be clear and must not rub off. Must also have marks/reference lines proximal to cuff to aid with depth of intubation and guide proper placement. Preference will be given to tubes with a second circumferential band (not more than 1cm from each other) to guide proper placement. Must also have a distinct mark at preformed curve to aid positioning Proximal end - secure 15mm OD circuit connector made from medical grade plastic Distal end - Atraumatic left-facing bevelled tip with Murphy eye on the opposite side All components must be latex-free sterile and individually packed in peel pouch that is easy to open - For SINGLE use only To comply with the latest ISO 5361 Information brochure or the outer package of the item should clearly indicate latex free | Uncuffed Oral South-Facing Preformed (RAE) Endotracheal Tube with Murphy eye. ID: 5.0mm OD 6.7mm Length ± 155mm
169 | RT13-23-008 | Cuffed and Uncuffed Tubes | UNCUFFED ORAL SOUTH-FACING | UNCUFFED ORAL SOUTH-FACING PREFORMED (RAE) ENDOTRACHEAL TUBE WITH 1 OR 2 MURPHY EYE Purpose: Short term oral intubation where tube is South facing and removed from surgical field Consists of a clear thermo-sensitive PVC tube that is kink resistant, with a South polar preformed curve with 1 or 2 Murphy eye Must be able to temporarily straightened the tube to allow the passage of a suction catheter Must have a radiopaque line and must be marked with numbered depth markings in 1cm intervals, tube size, length, ID, OD, trade name. The markings on tube must be clear and must not rub off. Must also have marks/reference lines proximal to cuff to aid with depth of intubation and guide proper placement. Preference will be given to tubes with a second circumferential band (not more than 1cm from each other) to guide proper placement. Must also have a distinct mark at preformed curve to aid positioning Proximal end - secure 15mm OD circuit connector made from medical grade plastic Distal end - Atraumatic left-facing bevelled tip with Murphy eye on the opposite side All components must be latex-free sterile and individually packed in peel pouch that is easy to open - For SINGLE use only To comply with the latest ISO 5361 Information brochure or the outer package of the item should clearly indicate latex free | Uncuffed Oral South-Facing Preformed (RAE) Endotracheal Tube with Murphy eye. ID: 5.5mm OD 7.3mm Length ± 165mm
170 | RT13-23-009 | Cuffed and Uncuffed Tubes | UNCUFFED ORAL SOUTH-FACING | UNCUFFED ORAL SOUTH-FACING PREFORMED (RAE) ENDOTRACHEAL TUBE WITH 1 OR 2 MURPHY EYE Purpose: Short term oral intubation where tube is South facing and removed from surgical field Consists of a clear thermo-sensitive PVC tube that is kink resistant, with a South polar preformed curve with 1 or 2 Murphy eye Must be able to temporarily straightened the tube to allow the passage of a suction catheter Must have a radiopaque line and must be marked with numbered depth markings in 1cm intervals, tube size, length, ID, OD, trade name. The markings on tube must be clear and must not rub off. Must also have marks/reference lines proximal to cuff to aid with depth of intubation and guide proper placement. Preference will be given to tubes with a second circumferential band (not more than 1cm from each other) to guide proper placement. Must also have a distinct mark at preformed curve to aid positioning Proximal end - secure 15mm OD circuit connector made from medical grade plastic Distal end - Atraumatic left-facing bevelled tip with Murphy eye on the opposite side All components must be latex-free sterile and individually packed in peel pouch that is easy to open - For SINGLE use only To comply with the latest ISO 5361 Information brochure or the outer package of the item should clearly indicate latex free | Uncuffed Oral South-Facing Preformed (RAE) Endotracheal Tube with Murphy eye. ID: 6.0mm OD 8.0mm Length ± 180mm
171 | RT13-23-010 | Cuffed and Uncuffed Tubes | CUFFED ORAL SOUTH-FACING | CUFFED ORAL SOUTH-FACING PREFORMED (RAE) ENDOTRACHEAL TUBE WITH 1 OR 2 MURPHY EYE - PAEDIATRIC Purpose: Short term oral intubation where tube is South facing and removed from surgical field Consists of a clear thermo-sensitive PVC tube that is kink resistant, with a South polar preformed curve with one or two Murphy eyes Must be able to temporarily straightened the tube to allow the passage of a suction catheter Must have a radiopaque line and must be marked with numbered depth markings in 1cm intervals, tube size, length, ID, OD, trade name. The markings on tube must be clear and must not rub off. Must also have marks/reference lines proximal to cuff to aid with depth of intubation and guide proper placement. Preference will be given to tubes with a second circumferential band (not more than 1cm from each other) to guide proper placement. Must also have a distinct mark at preformed curve to aid positioning Proximal end - secure 15mm OD circuit connector made from medical grade plastic Distal end - Atraumatic left-facing bevelled tip with Murphy eye on the opposite side Must have a high-volume-low-pressure polyurethane barrel shaped cuff with pilot cuff. Cuff thickness ± 50 microns. The pilot cuff must have a one-way valve and must indicate the degree of inflation accurately. The cuff must inflate and deflate easily and when deflated must conform closely to tube to minimise the risk of trauma during intubation and extubation and not increase the OD of the tube. The cuff must be resistant to Nitrous Oxide effects All components must be latex-free sterile and individually packed in peel pouch that is easy to open - For SINGLE use only To comply with the latest ISO 5361 Size 2.0mm and 2.5mm will be considered individually, with or without murphy eye. Information brochure or the outer package of the item should clearly indicate latex free | Cuffed Oral South-Facing Preformed (RAE) Endotracheal Tube with Murphy eye - Paediatric. ID: 2.0mm
172 | RT13-23-011 | Cuffed and Uncuffed Tubes | CUFFED ORAL SOUTH-FACING | CUFFED ORAL SOUTH-FACING PREFORMED (RAE) ENDOTRACHEAL TUBE WITH 1 OR 2 MURPHY EYE - PAEDIATRIC Purpose: Short term oral intubation where tube is South facing and removed from surgical field Consists of a clear thermo-sensitive PVC tube that is kink resistant, with a South polar preformed curve with one or two Murphy eyes Must be able to temporarily straightened the tube to allow the passage of a suction catheter Must have a radiopaque line and must be marked with numbered depth markings in 1cm intervals, tube size, length, ID, OD, trade name. The markings on tube must be clear and must not rub off. Must also have marks/reference lines proximal to cuff to aid with depth of intubation and guide proper placement. Preference will be given to tubes with a second circumferential band (not more than 1cm from each other) to guide proper placement. Must also have a distinct mark at preformed curve to aid positioning Proximal end - secure 15mm OD circuit connector made from medical grade plastic Distal end - Atraumatic left-facing bevelled tip with Murphy eye on the opposite side Must have a high-volume-low-pressure polyurethane barrel shaped cuff with pilot cuff. Cuff thickness ± 50 microns. The pilot cuff must have a one-way valve and must indicate the degree of inflation accurately. The cuff must inflate and deflate easily and when deflated must conform closely to tube to minimise the risk of trauma during intubation and extubation and not increase the OD of the tube. The cuff must be resistant to Nitrous Oxide effects All components must be latex-free sterile and individually packed in peel pouch that is easy to open - For SINGLE use only To comply with the latest ISO 5361 Size 2.0mm and 2.5mm will be considered individually, with or without murphy eye. Information brochure or the outer package of the item should clearly indicate latex free | Cuffed Oral South-Facing Preformed (RAE) Endotracheal Tube with Murphy eye - Paediatric. ID: 2.5mm
173 | RT13-23-012 | Cuffed and Uncuffed Tubes | CUFFED ORAL SOUTH-FACING | CUFFED ORAL SOUTH-FACING PREFORMED (RAE) ENDOTRACHEAL TUBE WITH 1 OR 2 MURPHY EYE - PAEDIATRIC Purpose: Short term oral intubation where tube is South facing and removed from surgical field Consists of a clear thermo-sensitive PVC tube that is kink resistant, with a South polar preformed curve with one or two Murphy eyes Must be able to temporarily straightened the tube to allow the passage of a suction catheter Must have a radiopaque line and must be marked with numbered depth markings in 1cm intervals, tube size, length, ID, OD, trade name. The markings on tube must be clear and must not rub off. Must also have marks/reference lines proximal to cuff to aid with depth of intubation and guide proper placement. Preference will be given to tubes with a second circumferential band (not more than 1cm from each other) to guide proper placement. Must also have a distinct mark at preformed curve to aid positioning Proximal end - secure 15mm OD circuit connector made from medical grade plastic Distal end - Atraumatic left-facing bevelled tip with Murphy eye on the opposite side Must have a high-volume-low-pressure polyurethane barrel shaped cuff with pilot cuff. Cuff thickness ± 50 microns. The pilot cuff must have a one-way valve and must indicate the degree of inflation accurately. The cuff must inflate and deflate easily and when deflated must conform closely to tube to minimise the risk of trauma during intubation and extubation and not increase the OD of the tube. The cuff must be resistant to Nitrous Oxide effects All components must be latex-free sterile and individually packed in peel pouch that is easy to open - For SINGLE use only To comply with the latest ISO 5361 Size 2.0mm and 2.5mm will be considered individually, with or without murphy eye. Information brochure or the outer package of the item should clearly indicate latex free | Cuffed Oral South-Facing Preformed (RAE) Endotracheal Tube with Murphy eye - Paediatric. ID: 3.0mm OD 4.0mm Length ± 105mm
174 | RT13-23-013 | Cuffed and Uncuffed Tubes | CUFFED ORAL SOUTH-FACING | CUFFED ORAL SOUTH-FACING PREFORMED (RAE) ENDOTRACHEAL TUBE WITH 1 OR 2 MURPHY EYE - PAEDIATRIC Purpose: Short term oral intubation where tube is South facing and removed from surgical field Consists of a clear thermo-sensitive PVC tube that is kink resistant, with a South polar preformed curve with one or two Murphy eyes Must be able to temporarily straightened the tube to allow the passage of a suction catheter Must have a radiopaque line and must be marked with numbered depth markings in 1cm intervals, tube size, length, ID, OD, trade name. The markings on tube must be clear and must not rub off. Must also have marks/reference lines proximal to cuff to aid with depth of intubation and guide proper placement. Preference will be given to tubes with a second circumferential band (not more than 1cm from each other) to guide proper placement. Must also have a distinct mark at preformed curve to aid positioning Proximal end - secure 15mm OD circuit connector made from medical grade plastic Distal end - Atraumatic left-facing bevelled tip with Murphy eye on the opposite side Must have a high-volume-low-pressure polyurethane barrel shaped cuff with pilot cuff. Cuff thickness ± 50 microns. The pilot cuff must have a one-way valve and must indicate the degree of inflation accurately. The cuff must inflate and deflate easily and when deflated must conform closely to tube to minimise the risk of trauma during intubation and extubation and not increase the OD of the tube. The cuff must be resistant to Nitrous Oxide effects All components must be latex-free sterile and individually packed in peel pouch that is easy to open - For SINGLE use only To comply with the latest ISO 5361 Size 2.0mm and 2.5mm will be considered individually, with or without murphy eye. Information brochure or the outer package of the item should clearly indicate latex free | Cuffed Oral South-Facing Preformed (RAE) Endotracheal Tube with Murphy eye - Paediatric. ID: 3.5mm OD 4.7mm Length ± 120mm
175 | RT13-23-014 | Cuffed and Uncuffed Tubes | CUFFED ORAL SOUTH-FACING | CUFFED ORAL SOUTH-FACING PREFORMED (RAE) ENDOTRACHEAL TUBE WITH 1 OR 2 MURPHY EYE - PAEDIATRIC Purpose: Short term oral intubation where tube is South facing and removed from surgical field Consists of a clear thermo-sensitive PVC tube that is kink resistant, with a South polar preformed curve with one or two Murphy eyes Must be able to temporarily straightened the tube to allow the passage of a suction catheter Must have a radiopaque line and must be marked with numbered depth markings in 1cm intervals, tube size, length, ID, OD, trade name. The markings on tube must be clear and must not rub off. Must also have marks/reference lines proximal to cuff to aid with depth of intubation and guide proper placement. Preference will be given to tubes with a second circumferential band (not more than 1cm from each other) to guide proper placement. Must also have a distinct mark at preformed curve to aid positioning Proximal end - secure 15mm OD circuit connector made from medical grade plastic Distal end - Atraumatic left-facing bevelled tip with Murphy eye on the opposite side Must have a high-volume-low-pressure polyurethane barrel shaped cuff with pilot cuff. Cuff thickness ± 50 microns. The pilot cuff must have a one-way valve and must indicate the degree of inflation accurately. The cuff must inflate and deflate easily and when deflated must conform closely to tube to minimise the risk of trauma during intubation and extubation and not increase the OD of the tube. The cuff must be resistant to Nitrous Oxide effects All components must be latex-free sterile and individually packed in peel pouch that is easy to open - For SINGLE use only To comply with the latest ISO 5361 Size 2.0mm and 2.5mm will be considered individually, with or without murphy eye. Information brochure or the outer package of the item should clearly indicate latex free | Cuffed Oral South-Facing Preformed (RAE) Endotracheal Tube with Murphy eye - Paediatric. ID: 4.0mm OD 5.3mm Length ± 130mm
176 | RT13-23-015 | Cuffed and Uncuffed Tubes | CUFFED ORAL SOUTH-FACING | CUFFED ORAL SOUTH-FACING PREFORMED (RAE) ENDOTRACHEAL TUBE WITH 1 OR 2 MURPHY EYE - PAEDIATRIC Purpose: Short term oral intubation where tube is South facing and removed from surgical field Consists of a clear thermo-sensitive PVC tube that is kink resistant, with a South polar preformed curve with one or two Murphy eyes Must be able to temporarily straightened the tube to allow the passage of a suction catheter Must have a radiopaque line and must be marked with numbered depth markings in 1cm intervals, tube size, length, ID, OD, trade name. The markings on tube must be clear and must not rub off. Must also have marks/reference lines proximal to cuff to aid with depth of intubation and guide proper placement. Preference will be given to tubes with a second circumferential band (not more than 1cm from each other) to guide proper placement. Must also have a distinct mark at preformed curve to aid positioning Proximal end - secure 15mm OD circuit connector made from medical grade plastic Distal end - Atraumatic left-facing bevelled tip with Murphy eye on the opposite side Must have a high-volume-low-pressure polyurethane barrel shaped cuff with pilot cuff. Cuff thickness ± 50 microns. The pilot cuff must have a one-way valve and must indicate the degree of inflation accurately. The cuff must inflate and deflate easily and when deflated must conform closely to tube to minimise the risk of trauma during intubation and extubation and not increase the OD of the tube. The cuff must be resistant to Nitrous Oxide effects All components must be latex-free sterile and individually packed in peel pouch that is easy to open - For SINGLE use only To comply with the latest ISO 5361 Size 2.0mm and 2.5mm will be considered individually, with or without murphy eye. Information brochure or the outer package of the item should clearly indicate latex free | Cuffed Oral South-Facing Preformed (RAE) Endotracheal Tube with Murphy eye - Paediatric. ID: 4.5mm OD 6.0mm Length ± 145mm
177 | RT13-23-019 | Cuffed and Uncuffed Tubes | CUFFED ORAL SOUTH-FACING | CUFFED ORAL SOUTH-FACING PREFORMED (RAE) ENDOTRACHEAL TUBE WITH 1 or 2 MURPHY EYE - ADULT Purpose: Short term oral intubation where tube is South facing and removed from surgical field Consists of a clear thermo-sensitive PVC tube that is kink resistant, with a South polar preformed curve with one or two Murphy eyes Must be able to temporarily straightened the tube to allow the passage of a suction catheter Must have a radiopaque line and must be marked with numbered depth markings in 1cm intervals, tube size, length, ID, OD, trade name. The markings on tube must be clear and must not rub off. Must also have marks/reference lines proximal to cuff to aid with depth of intubation and guide proper placement. Preference will be given to tubes with a second circumferential band (not more than 2cm from each other) to guide proper placement. Must also have a distinct mark at preformed curve to aid positioning Proximal end - secure 15mm OD circuit connector made from medical grade plastic Distal end - Atraumatic left-facing bevelled tip with Murphy eye on the opposite side Must have a high-volume-low-pressure polyurethane barrel shaped cuff with pilot cuff. Cuff thickness ± 50 microns. The pilot cuff must have a one-way valve and must indicate the degree of inflation accurately. The cuff must inflate and deflate easily and when deflated must conform closely to tube to minimise the risk of trauma during intubation and extubation and not increase the OD of the tube. The cuff must be resistant to Nitrous Oxide effects All components must be latex-free sterile and individually packed in peel pouch that is easy to open - For SINGLE use only To comply with the latest ISO 5361 Information brochure or the outer package of the item should clearly indicate latex free
178 | RT13-23-020 | Cuffed and Uncuffed Tubes | CUFFED ORAL SOUTH-FACING | CUFFED ORAL SOUTH-FACING PREFORMED (RAE) ENDOTRACHEAL TUBE WITH 1 or 2 MURPHY EYE - ADULT Purpose: Short term oral intubation where tube is South facing and removed from surgical field Consists of a clear thermo-sensitive PVC tube that is kink resistant, with a South polar preformed curve with one or two Murphy eyes Must be able to temporarily straightened the tube to allow the passage of a suction catheter Must have a radiopaque line and must be marked with numbered depth markings in 1cm intervals, tube size, length, ID, OD, trade name. The markings on tube must be clear and must not rub off. Must also have marks/reference lines proximal to cuff to aid with depth of intubation and guide proper placement. Preference will be given to tubes with a second circumferential band (not more than 2cm from each other) to guide proper placement. Must also have a distinct mark at preformed curve to aid positioning Proximal end - secure 15mm OD circuit connector made from medical grade plastic Distal end - Atraumatic left-facing bevelled tip with Murphy eye on the opposite side Must have a high-volume-low-pressure polyurethane barrel shaped cuff with pilot cuff. Cuff thickness ± 50 microns. The pilot cuff must have a one-way valve and must indicate the degree of inflation accurately. The cuff must inflate and deflate easily and when deflated must conform closely to tube to minimise the risk of trauma during intubation and extubation and not increase the OD of the tube. The cuff must be resistant to Nitrous Oxide effects All components must be latex-free sterile and individually packed in peel pouch that is easy to open - For SINGLE use only To comply with the latest ISO 5361 Information brochure or the outer package of the item should clearly indicate latex free | Cuffed Oral South-Facing Preformed (RAE) Endotracheal Tube with Murphy eye - Adult ID:7.0mm OD 9.3mm Length ± 200mm
179 | RT13-23-021 | Cuffed and Uncuffed Tubes | CUFFED ORAL SOUTH-FACING | CUFFED ORAL SOUTH-FACING PREFORMED (RAE) ENDOTRACHEAL TUBE WITH 1 or 2 MURPHY EYE - ADULT Purpose: Short term oral intubation where tube is South facing and removed from surgical field Consists of a clear thermo-sensitive PVC tube that is kink resistant, with a South polar preformed curve with one or two Murphy eyes Must be able to temporarily straightened the tube to allow the passage of a suction catheter Must have a radiopaque line and must be marked with numbered depth markings in 1cm intervals, tube size, length, ID, OD, trade name. The markings on tube must be clear and must not rub off. Must also have marks/reference lines proximal to cuff to aid with depth of intubation and guide proper placement. Preference will be given to tubes with a second circumferential band (not more than 2cm from each other) to guide proper placement. Must also have a distinct mark at preformed curve to aid positioning Proximal end - secure 15mm OD circuit connector made from medical grade plastic Distal end - Atraumatic left-facing bevelled tip with Murphy eye on the opposite side Must have a high-volume-low-pressure polyurethane barrel shaped cuff with pilot cuff. Cuff thickness ± 50 microns. The pilot cuff must have a one-way valve and must indicate the degree of inflation accurately. The cuff must inflate and deflate easily and when deflated must conform closely to tube to minimise the risk of trauma during intubation and extubation and not increase the OD of the tube. The cuff must be resistant to Nitrous Oxide effects All components must be latex-free sterile and individually packed in peel pouch that is easy to open - For SINGLE use only To comply with the latest ISO 5361 Information brochure or the outer package of the item should clearly indicate latex free | Cuffed Oral South-Facing Preformed (RAE) Endotracheal Tube with Murphy eye - Adult ID:7.5mm OD 10mm Length ± 210mm
180 | RT13-23-022 | Cuffed and Uncuffed Tubes | CUFFED ORAL SOUTH-FACING | CUFFED ORAL SOUTH-FACING PREFORMED (RAE) ENDOTRACHEAL TUBE WITH 1 or 2 MURPHY EYE - ADULT Purpose: Short term oral intubation where tube is South facing and removed from surgical field Consists of a clear thermo-sensitive PVC tube that is kink resistant, with a South polar preformed curve with one or two Murphy eyes Must be able to temporarily straightened the tube to allow the passage of a suction catheter Must have a radiopaque line and must be marked with numbered depth markings in 1cm intervals, tube size, length, ID, OD, trade name. The markings on tube must be clear and must not rub off. Must also have marks/reference lines proximal to cuff to aid with depth of intubation and guide proper placement. Preference will be given to tubes with a second circumferential band (not more than 2cm from each other) to guide proper placement. Must also have a distinct mark at preformed curve to aid positioning Proximal end - secure 15mm OD circuit connector made from medical grade plastic Distal end - Atraumatic left-facing bevelled tip with Murphy eye on the opposite side Must have a high-volume-low-pressure polyurethane barrel shaped cuff with pilot cuff. Cuff thickness ± 50 microns. The pilot cuff must have a one-way valve and must indicate the degree of inflation accurately. The cuff must inflate and deflate easily and when deflated must conform closely to tube to minimise the risk of trauma during intubation and extubation and not increase the OD of the tube. The cuff must be resistant to Nitrous Oxide effects All components must be latex-free sterile and individually packed in peel pouch that is easy to open - For SINGLE use only To comply with the latest ISO 5361 Information brochure or the outer package of the item should clearly indicate latex free | Cuffed Oral South-Facing Preformed (RAE) Endotracheal Tube with Murphy eye - Adult ID:8.0mm OD 10.7mm Length ± 220mm
181 | RT13-23-023 | Cuffed and Uncuffed Tubes | CUFFED ORAL SOUTH-FACING | CUFFED ORAL SOUTH-FACING PREFORMED (RAE) ENDOTRACHEAL TUBE WITH 1 or 2 MURPHY EYE - ADULT Purpose: Short term oral intubation where tube is South facing and removed from surgical field Consists of a clear thermo-sensitive PVC tube that is kink resistant, with a South polar preformed curve with one or two Murphy eyes Must be able to temporarily straightened the tube to allow the passage of a suction catheter Must have a radiopaque line and must be marked with numbered depth markings in 1cm intervals, tube size, length, ID, OD, trade name. The markings on tube must be clear and must not rub off. Must also have marks/reference lines proximal to cuff to aid with depth of intubation and guide proper placement. Preference will be given to tubes with a second circumferential band (not more than 2cm from each other) to guide proper placement. Must also have a distinct mark at preformed curve to aid positioning Proximal end - secure 15mm OD circuit connector made from medical grade plastic Distal end - Atraumatic left-facing bevelled tip with Murphy eye on the opposite side Must have a high-volume-low-pressure polyurethane barrel shaped cuff with pilot cuff. Cuff thickness ± 50 microns. The pilot cuff must have a one-way valve and must indicate the degree of inflation accurately. The cuff must inflate and deflate easily and when deflated must conform closely to tube to minimise the risk of trauma during intubation and extubation and not increase the OD of the tube. The cuff must be resistant to Nitrous Oxide effects All components must be latex-free sterile and individually packed in peel pouch that is easy to open - For SINGLE use only To comply with the latest ISO 5361 Information brochure or the outer package of the item should clearly indicate latex free
182 | RT13-23-024 | Cuffed and Uncuffed Tubes | CUFFED ORAL SOUTH-FACING | CUFFED ORAL SOUTH-FACING PREFORMED (RAE) ENDOTRACHEAL TUBE WITH 1 or 2 MURPHY EYE - ADULT Purpose: Short term oral intubation where tube is South facing and removed from surgical field Consists of a clear thermo-sensitive PVC tube that is kink resistant, with a South polar preformed curve with one or two Murphy eyes Must be able to temporarily straightened the tube to allow the passage of a suction catheter Must have a radiopaque line and must be marked with numbered depth markings in 1cm intervals, tube size, length, ID, OD, trade name. The markings on tube must be clear and must not rub off. Must also have marks/reference lines proximal to cuff to aid with depth of intubation and guide proper placement. Preference will be given to tubes with a second circumferential band (not more than 2cm from each other) to guide proper placement. Must also have a distinct mark at preformed curve to aid positioning Proximal end - secure 15mm OD circuit connector made from medical grade plastic Distal end - Atraumatic left-facing bevelled tip with Murphy eye on the opposite side Must have a high-volume-low-pressure polyurethane barrel shaped cuff with pilot cuff. Cuff thickness ± 50 microns. The pilot cuff must have a one-way valve and must indicate the degree of inflation accurately. The cuff must inflate and deflate easily and when deflated must conform closely to tube to minimise the risk of trauma during intubation and extubation and not increase the OD of the tube. The cuff must be resistant to Nitrous Oxide effects All components must be latex-free sterile and individually packed in peel pouch that is easy to open - For SINGLE use only To comply with the latest ISO 5361 Information brochure or the outer package of the item should clearly indicate latex free | Cuffed Oral South-Facing Preformed (RAE) Endotracheal Tube with Murphy eye - Adult. ID: 6.5mm OD 8.7mm Length ± 190mm
183 | RT13-24-001 | ENDOTRACHEAL TUBE | DOUBLE LUMEN ENDOBRONCHIAL TUBE | DOUBLE LUMEN ENDOBRONCHIAL TUBE- LEFT Purpose: For isolation of one lung to prevent contamination, or for ventilation of one lung in thoracic surgery Consists of a double Lumen clear PVC tube that is thermo-sensitive with radiopaque lines and markers at key features. Markings starting at 26cm from distal end marked every 2cm up to and about 32cm Tubes to be flexible enough to allow manipulation through distorted airways, but without luminal collapse. Must have a smooth transition of the inner lumen without narrowing, to allow easy passage of flexible fibre-optic-bronchoscope. Tracheal and Bronchial extensions on the main tube must be colour coded - blue for bronchial, clear for tracheal Stylet: Malleable shape forming - packaged in place down the bronchial lumen Supplied with 2 proximal right-angled connectors with a Bronchoscopy/Aspiration Port The port must have a toggled cap and allow for easy passage of either a flexible fibre-optic-bronchoscope, or a suction catheter – both without a leak. The proximal angled connectors to connect to the bronchial and tracheal tubes and the distal connectors to connect to a standard Y-connector that is compatible with a standard breathing circuit. 2 - 4 Suction catheters: with suction control, to fit the total length of the tube Colourless tracheal high-volume-low-pressure polyvinyl chloride cuff – with colourless pilot balloon, clearly labelled “tracheal” Blue bronchial low pressure Polyurethane cuff - with blue pilot balloon, clearly labelled “bronchial”. Bronchial cuff to be short and cylindrical. One-way valves in pilot balloons. The cuffs must inflate and deflate easily, and the pilot cuffs must indicate the degree of inflation accurately. Cuffs must be resistant to the effects of nitrous oxide. When deflated, cuffs must conform closely to the tube to minimise the risk of trauma during intubation and extubation. All components must be latex-free, sterile and individually packed in peel pouch - SINGLE use only To comply with ISO 16628 Information brochure or the outer package of the item should clearly indicate latex free | Double Lumen Endobronchial tube: Left Size: 26Fr Shaft Length: > mm ID: > (bilaterally) OD: <
184 | RT13-24-002 | ENDOTRACHEAL TUBE | DOUBLE LUMEN ENDOBRONCHIAL TUBE | DOUBLE LUMEN ENDOBRONCHIAL TUBE- LEFT Purpose: For isolation of one lung to prevent contamination, or for ventilation of one lung in thoracic surgery Consists of a double Lumen clear PVC tube that is thermo-sensitive with radiopaque lines and markers at key features. Markings starting at 26cm from distal end marked every 2cm up to and about 32cm Tubes to be flexible enough to allow manipulation through distorted airways, but without luminal collapse. Must have a smooth transition of the inner lumen without narrowing, to allow easy passage of flexible fibre-optic-bronchoscope. Tracheal and Bronchial extensions on the main tube must be colour coded - blue for bronchial, clear for tracheal Stylet: Malleable shape forming - packaged in place down the bronchial lumen Supplied with 2 proximal right-angled connectors with a Bronchoscopy/Aspiration Port The port must have a toggled cap and allow for easy passage of either a flexible fibre-optic-bronchoscope, or a suction catheter – both without a leak. The proximal angled connectors to connect to the bronchial and tracheal tubes and the distal connectors to connect to a standard Y-connector that is compatible with a standard breathing circuit. 2 - 4 Suction catheters: with suction control, to fit the total length of the tube Colourless tracheal high-volume-low-pressure polyvinyl chloride cuff – with colourless pilot balloon, clearly labelled “tracheal” Blue bronchial low pressure Polyurethane cuff - with blue pilot balloon, clearly labelled “bronchial”. Bronchial cuff to be short and cylindrical. One-way valves in pilot balloons. The cuffs must inflate and deflate easily, and the pilot cuffs must indicate the degree of inflation accurately. Cuffs must be resistant to the effects of nitrous oxide. When deflated, cuffs must conform closely to the tube to minimise the risk of trauma during intubation and extubation. All components must be latex-free, sterile and individually packed in peel pouch - SINGLE use only To comply with ISO 16628 Information brochure or the outer package of the item should clearly indicate latex free | Double Lumen Endobronchial tube: Left Size: 28Fr Shaft Length: < 420mm ID: > 3.1mm (bilaterally) OD: < 10mm ± 2mm
185 | RT13-24-003 | ENDOTRACHEAL TUBE | DOUBLE LUMEN ENDOBRONCHIAL TUBE | DOUBLE LUMEN ENDOBRONCHIAL TUBE- LEFT Purpose: For isolation of one lung to prevent contamination, or for ventilation of one lung in thoracic surgery Consists of a double Lumen clear PVC tube that is thermo-sensitive with radiopaque lines and markers at key features. Markings starting at 26cm from distal end marked every 2cm up to and about 32cm Tubes to be flexible enough to allow manipulation through distorted airways, but without luminal collapse. Must have a smooth transition of the inner lumen without narrowing, to allow easy passage of flexible fibre-optic-bronchoscope. Tracheal and Bronchial extensions on the main tube must be colour coded - blue for bronchial, clear for tracheal Stylet: Malleable shape forming - packaged in place down the bronchial lumen Supplied with 2 proximal right-angled connectors with a Bronchoscopy/Aspiration Port The port must have a toggled cap and allow for easy passage of either a flexible fibre-optic-bronchoscope, or a suction catheter – both without a leak. The proximal angled connectors to connect to the bronchial and tracheal tubes and the distal connectors to connect to a standard Y-connector that is compatible with a standard breathing circuit. 2 - 4 Suction catheters: with suction control, to fit the total length of the tube Colourless tracheal high-volume-low-pressure polyvinyl chloride cuff – with colourless pilot balloon, clearly labelled “tracheal” Blue bronchial low pressure Polyurethane cuff - with blue pilot balloon, clearly labelled “bronchial”. Bronchial cuff to be short and cylindrical. One-way valves in pilot balloons. The cuffs must inflate and deflate easily, and the pilot cuffs must indicate the degree of inflation accurately. Cuffs must be resistant to the effects of nitrous oxide. When deflated, cuffs must conform closely to the tube to minimise the risk of trauma during intubation and extubation. All components must be latex-free, sterile and individually packed in peel pouch - SINGLE use only To comply with ISO 16628 Information brochure or the outer package of the item should clearly indicate latex free | Double Lumen Endobronchial tube: Left Size: 32Fr Shaft Length: < 420mm ID: > 3.3mm (bilaterally) OD: < 10.7mm ± 2mm
186 | RT13-24-004 | ENDOTRACHEAL TUBE | DOUBLE LUMEN ENDOBRONCHIAL TUBE | DOUBLE LUMEN ENDOBRONCHIAL TUBE- LEFT Purpose: For isolation of one lung to prevent contamination, or for ventilation of one lung in thoracic surgery Consists of a double Lumen clear PVC tube that is thermo-sensitive with radiopaque lines and markers at key features. Markings starting at 26cm from distal end marked every 2cm up to and about 32cm Tubes to be flexible enough to allow manipulation through distorted airways, but without luminal collapse. Must have a smooth transition of the inner lumen without narrowing, to allow easy passage of flexible fibre-optic-bronchoscope. Tracheal and Bronchial extensions on the main tube must be colour coded - blue for bronchial, clear for tracheal Stylet: Malleable shape forming - packaged in place down the bronchial lumen Supplied with 2 proximal right-angled connectors with a Bronchoscopy/Aspiration Port The port must have a toggled cap and allow for easy passage of either a flexible fibre-optic-bronchoscope, or a suction catheter – both without a leak. The proximal angled connectors to connect to the bronchial and tracheal tubes and the distal connectors to connect to a standard Y-connector that is compatible with a standard breathing circuit. 2 - 4 Suction catheters: with suction control, to fit the total length of the tube Colourless tracheal high-volume-low-pressure polyvinyl chloride cuff – with colourless pilot balloon, clearly labelled “tracheal” Blue bronchial low pressure Polyurethane cuff - with blue pilot balloon, clearly labelled “bronchial”. Bronchial cuff to be short and cylindrical. One-way valves in pilot balloons. The cuffs must inflate and deflate easily, and the pilot cuffs must indicate the degree of inflation accurately. Cuffs must be resistant to the effects of nitrous oxide. When deflated, cuffs must conform closely to the tube to minimise the risk of trauma during intubation and extubation. All components must be latex-free, sterile and individually packed in peel pouch - SINGLE use only To comply with ISO 16628 Information brochure or the outer package of the item should clearly indicate latex free | Double Lumen Endobronchial tube: Left Size: 35Fr Shaft Length: < 420mm ID: > 4.7mm (bilaterally) OD: < 11.7mm ± 2mm
187 | RT13-24-005 | ENDOTRACHEAL TUBE | DOUBLE LUMEN ENDOBRONCHIAL TUBE | DOUBLE LUMEN ENDOBRONCHIAL TUBE- LEFT Purpose: For isolation of one lung to prevent contamination, or for ventilation of one lung in thoracic surgery Consists of a double Lumen clear PVC tube that is thermo-sensitive with radiopaque lines and markers at key features. Markings starting at 26cm from distal end marked every 2cm up to and about 32cm Tubes to be flexible enough to allow manipulation through distorted airways, but without luminal collapse. Must have a smooth transition of the inner lumen without narrowing, to allow easy passage of flexible fibre-optic-bronchoscope. Tracheal and Bronchial extensions on the main tube must be colour coded - blue for bronchial, clear for tracheal Stylet: Malleable shape forming - packaged in place down the bronchial lumen Supplied with 2 proximal right-angled connectors with a Bronchoscopy/Aspiration Port The port must have a toggled cap and allow for easy passage of either a flexible fibre-optic-bronchoscope, or a suction catheter – both without a leak. The proximal angled connectors to connect to the bronchial and tracheal tubes and the distal connectors to connect to a standard Y-connector that is compatible with a standard breathing circuit. 2 - 4 Suction catheters: with suction control, to fit the total length of the tube Colourless tracheal high-volume-low-pressure polyvinyl chloride cuff – with colourless pilot balloon, clearly labelled “tracheal” Blue bronchial low pressure Polyurethane cuff - with blue pilot balloon, clearly labelled “bronchial”. Bronchial cuff to be short and cylindrical. One-way valves in pilot balloons. The cuffs must inflate and deflate easily, and the pilot cuffs must indicate the degree of inflation accurately. Cuffs must be resistant to the effects of nitrous oxide. When deflated, cuffs must conform closely to the tube to minimise the risk of trauma during intubation and extubation. All components must be latex-free, sterile and individually packed in peel pouch - SINGLE use only To comply with ISO 16628 Information brochure or the outer package of the item should clearly indicate latex free | Double Lumen Endobronchial tube: Left Size: 37Fr Shaft Length: < 420mm ID: > 5.0mm (bilaterally) OD: < 12.3mm ± 2mm
188 | RT13-24-006 | ENDOTRACHEAL TUBE | DOUBLE LUMEN ENDOBRONCHIAL TUBE | DOUBLE LUMEN ENDOBRONCHIAL TUBE- LEFT Purpose: For isolation of one lung to prevent contamination, or for ventilation of one lung in thoracic surgery Consists of a double Lumen clear PVC tube that is thermo-sensitive with radiopaque lines and markers at key features. Markings starting at 26cm from distal end marked every 2cm up to and about 32cm Tubes to be flexible enough to allow manipulation through distorted airways, but without luminal collapse. Must have a smooth transition of the inner lumen without narrowing, to allow easy passage of flexible fibre-optic-bronchoscope. Tracheal and Bronchial extensions on the main tube must be colour coded - blue for bronchial, clear for tracheal Stylet: Malleable shape forming - packaged in place down the bronchial lumen Supplied with 2 proximal right-angled connectors with a Bronchoscopy/Aspiration Port The port must have a toggled cap and allow for easy passage of either a flexible fibre-optic-bronchoscope, or a suction catheter – both without a leak. The proximal angled connectors to connect to the bronchial and tracheal tubes and the distal connectors to connect to a standard Y-connector that is compatible with a standard breathing circuit. 2 - 4 Suction catheters: with suction control, to fit the total length of the tube Colourless tracheal high-volume-low-pressure polyvinyl chloride cuff – with colourless pilot balloon, clearly labelled “tracheal” Blue bronchial low pressure Polyurethane cuff - with blue pilot balloon, clearly labelled “bronchial”. Bronchial cuff to be short and cylindrical. One-way valves in pilot balloons. The cuffs must inflate and deflate easily, and the pilot cuffs must indicate the degree of inflation accurately. Cuffs must be resistant to the effects of nitrous oxide. When deflated, cuffs must conform closely to the tube to minimise the risk of trauma during intubation and extubation. All components must be latex-free, sterile and individually packed in peel pouch - SINGLE use only To comply with ISO 16628 Information brochure or the outer package of the item should clearly indicate latex free | Double Lumen Endobronchial tube: Left Size: 39Fr Shaft Length: < 420mm ID: > 5.2mm (bilaterally) OD: < 13mm ± 2mm
189 | RT13-24-007 | ENDOTRACHEAL TUBE | DOUBLE LUMEN ENDOBRONCHIAL TUBE | DOUBLE LUMEN ENDOBRONCHIAL TUBE- LEFT Purpose: For isolation of one lung to prevent contamination, or for ventilation of one lung in thoracic surgery Consists of a double Lumen clear PVC tube that is thermo-sensitive with radiopaque lines and markers at key features. Markings starting at 26cm from distal end marked every 2cm up to and about 32cm Tubes to be flexible enough to allow manipulation through distorted airways, but without luminal collapse. Must have a smooth transition of the inner lumen without narrowing, to allow easy passage of flexible fibre-optic-bronchoscope. Tracheal and Bronchial extensions on the main tube must be colour coded - blue for bronchial, clear for tracheal Stylet: Malleable shape forming - packaged in place down the bronchial lumen Supplied with 2 proximal right-angled connectors with a Bronchoscopy/Aspiration Port The port must have a toggled cap and allow for easy passage of either a flexible fibre-optic-bronchoscope, or a suction catheter – both without a leak. The proximal angled connectors to connect to the bronchial and tracheal tubes and the distal connectors to connect to a standard Y-connector that is compatible with a standard breathing circuit. 2 - 4 Suction catheters: with suction control, to fit the total length of the tube Colourless tracheal high-volume-low-pressure polyvinyl chloride cuff – with colourless pilot balloon, clearly labelled “tracheal” Blue bronchial low pressure Polyurethane cuff - with blue pilot balloon, clearly labelled “bronchial”. Bronchial cuff to be short and cylindrical. One-way valves in pilot balloons. The cuffs must inflate and deflate easily, and the pilot cuffs must indicate the degree of inflation accurately. Cuffs must be resistant to the effects of nitrous oxide. When deflated, cuffs must conform closely to the tube to minimise the risk of trauma during intubation and extubation. All components must be latex-free, sterile and individually packed in peel pouch - SINGLE use only To comply with ISO 16628 Information brochure or the outer package of the item should clearly indicate latex free | Double Lumen Endobronchial tube: Left Size: 41Fr Shaft Length:< 420mm ID: > 5.3mm (bilaterally) OD: < 13.7mm ± 2mm
190 | RT13-24-008 | ENDOTRACHEAL TUBE | DOUBLE LUMEN ENDOBRONCHIAL TUBE | DOUBLE LUMEN ENDOBRONCHIAL TUBE- RIGHT Purpose: For isolation of one lung to prevent contamination, or for ventilation of one lung in thoracic surgery Consists of a double Lumen clear PVC tube that is thermo-sensitive with radiopaque lines and markers at key features. Markings starting at 26cm from distal end marked every 2cm up to and about 32cm Tubes to be flexible enough to allow manipulation through distorted airways, but without luminal collapse. Must have a smooth transition of the inner lumen without narrowing, to allow easy passage of flexible fibre-optic-bronchoscope. Tracheal and Bronchial extensions on the main tube must be colour coded - blue for bronchial, clear for tracheal Stylet: Malleable shape forming - packaged in place down the bronchial lumen Supplied with 2 proximal right-angled connectors with a Bronchoscopy/Aspiration Port The port must have a toggled cap and allow for easy passage of either a flexible fibre-optic-bronchoscope, or a suction catheter – both without a leak. The proximal angled connectors to connect to the bronchial and tracheal tubes and the distal connectors to connect to a standard Y-connector that is compatible with a standard breathing circuit. 2 - 4 Suction catheters: with suction control, to fit the total length of the tube Colourless tracheal high-volume-low-pressure polyvinyl chloride cuff – with colourless pilot balloon, clearly labelled “tracheal” Blue bronchial low pressure Polyurethane cuff - with blue pilot balloon, clearly labelled “bronchial”. Bronchial cuff to be saddle-shaped One-way valves in pilot balloons. The cuffs must inflate and deflate easily, and the pilot cuffs must indicate the degree of inflation accurately. Cuffs must be resistant to the effects of nitrous oxide. When deflated, cuffs must conform closely to the tube to minimise the risk of trauma during intubation and extubation. All components must be latex-free, sterile and individually packed in peel pouch - SINGLE use only To comply with ISO 16628 Information brochure or the outer package of the item should clearly indicate latex free | Double Lumen Endobronchial tube: Right Size: 26Fr Shaft Length: 420mm ID: > mm (bilaterally) OD: <
191 | RT13-24-009 | ENDOTRACHEAL TUBE | DOUBLE LUMEN ENDOBRONCHIAL TUBE | DOUBLE LUMEN ENDOBRONCHIAL TUBE- RIGHT Purpose: For isolation of one lung to prevent contamination, or for ventilation of one lung in thoracic surgery Consists of a double Lumen clear PVC tube that is thermo-sensitive with radiopaque lines and markers at key features. Markings starting at 26cm from distal end marked every 2cm up to and about 32cm Tubes to be flexible enough to allow manipulation through distorted airways, but without luminal collapse. Must have a smooth transition of the inner lumen without narrowing, to allow easy passage of flexible fibre-optic-bronchoscope. Tracheal and Bronchial extensions on the main tube must be colour coded - blue for bronchial, clear for tracheal Stylet: Malleable shape forming - packaged in place down the bronchial lumen Supplied with 2 proximal right-angled connectors with a Bronchoscopy/Aspiration Port The port must have a toggled cap and allow for easy passage of either a flexible fibre-optic-bronchoscope, or a suction catheter – both without a leak. The proximal angled connectors to connect to the bronchial and tracheal tubes and the distal connectors to connect to a standard Y-connector that is compatible with a standard breathing circuit. 2 - 4 Suction catheters: with suction control, to fit the total length of the tube Colourless tracheal high-volume-low-pressure polyvinyl chloride cuff – with colourless pilot balloon, clearly labelled “tracheal” Blue bronchial low pressure Polyurethane cuff - with blue pilot balloon, clearly labelled “bronchial”. Bronchial cuff to be saddle-shaped One-way valves in pilot balloons. The cuffs must inflate and deflate easily, and the pilot cuffs must indicate the degree of inflation accurately. Cuffs must be resistant to the effects of nitrous oxide. When deflated, cuffs must conform closely to the tube to minimise the risk of trauma during intubation and extubation. All components must be latex-free, sterile and individually packed in peel pouch - SINGLE use only To comply with ISO 16628 Information brochure or the outer package of the item should clearly indicate latex free | Double Lumen Endobronchial tube: Right Size: 28Fr Shaft Length: 420mm ID: > 3.2mm (bilaterally) OD: < 9.3mm ± 2mm
192 | RT13-24-010 | ENDOTRACHEAL TUBE | DOUBLE LUMEN ENDOBRONCHIAL TUBE | DOUBLE LUMEN ENDOBRONCHIAL TUBE- RIGHT Purpose: For isolation of one lung to prevent contamination, or for ventilation of one lung in thoracic surgery Consists of a double Lumen clear PVC tube that is thermo-sensitive with radiopaque lines and markers at key features. Markings starting at 26cm from distal end marked every 2cm up to and about 32cm Tubes to be flexible enough to allow manipulation through distorted airways, but without luminal collapse. Must have a smooth transition of the inner lumen without narrowing, to allow easy passage of flexible fibre-optic-bronchoscope. Tracheal and Bronchial extensions on the main tube must be colour coded - blue for bronchial, clear for tracheal Stylet: Malleable shape forming - packaged in place down the bronchial lumen Supplied with 2 proximal right-angled connectors with a Bronchoscopy/Aspiration Port The port must have a toggled cap and allow for easy passage of either a flexible fibre-optic-bronchoscope, or a suction catheter – both without a leak. The proximal angled connectors to connect to the bronchial and tracheal tubes and the distal connectors to connect to a standard Y-connector that is compatible with a standard breathing circuit. 2 - 4 Suction catheters: with suction control, to fit the total length of the tube Colourless tracheal high-volume-low-pressure polyvinyl chloride cuff – with colourless pilot balloon, clearly labelled “tracheal” Blue bronchial low pressure Polyurethane cuff - with blue pilot balloon, clearly labelled “bronchial”. Bronchial cuff to be saddle-shaped One-way valves in pilot balloons. The cuffs must inflate and deflate easily, and the pilot cuffs must indicate the degree of inflation accurately. Cuffs must be resistant to the effects of nitrous oxide. When deflated, cuffs must conform closely to the tube to minimise the risk of trauma during intubation and extubation. All components must be latex-free, sterile and individually packed in peel pouch - SINGLE use only To comply with ISO 16628 Information brochure or the outer package of the item should clearly indicate latex free | Double Lumen Endobronchial tube: Right Size: 32Fr Shaft Length: 420mm ID: > 3.4mm (bilaterally) OD: < 10.7mm ± 2mm
193 | RT13-24-011 | ENDOTRACHEAL TUBE | DOUBLE LUMEN ENDOBRONCHIAL TUBE | DOUBLE LUMEN ENDOBRONCHIAL TUBE- RIGHT Purpose: For isolation of one lung to prevent contamination, or for ventilation of one lung in thoracic surgery Consists of a double Lumen clear PVC tube that is thermo-sensitive with radiopaque lines and markers at key features. Markings starting at 26cm from distal end marked every 2cm up to and about 32cm Tubes to be flexible enough to allow manipulation through distorted airways, but without luminal collapse. Must have a smooth transition of the inner lumen without narrowing, to allow easy passage of flexible fibre-optic-bronchoscope. Tracheal and Bronchial extensions on the main tube must be colour coded - blue for bronchial, clear for tracheal Stylet: Malleable shape forming - packaged in place down the bronchial lumen Supplied with 2 proximal right-angled connectors with a Bronchoscopy/Aspiration Port The port must have a toggled cap and allow for easy passage of either a flexible fibre-optic-bronchoscope, or a suction catheter – both without a leak. The proximal angled connectors to connect to the bronchial and tracheal tubes and the distal connectors to connect to a standard Y-connector that is compatible with a standard breathing circuit. 2 - 4 Suction catheters: with suction control, to fit the total length of the tube Colourless tracheal high-volume-low-pressure polyvinyl chloride cuff – with colourless pilot balloon, clearly labelled “tracheal” Blue bronchial low pressure Polyurethane cuff - with blue pilot balloon, clearly labelled “bronchial”. Bronchial cuff to be saddle-shaped One-way valves in pilot balloons. The cuffs must inflate and deflate easily, and the pilot cuffs must indicate the degree of inflation accurately. Cuffs must be resistant to the effects of nitrous oxide. When deflated, cuffs must conform closely to the tube to minimise the risk of trauma during intubation and extubation. All components must be latex-free, sterile and individually packed in peel pouch - SINGLE use only To comply with ISO 16628 Information brochure or the outer package of the item should clearly indicate latex free | Double Lumen Endobronchial tube: Right Size: 35Fr Shaft Length: 420mm ID: > 4.7mm (bilaterally) OD: < 11.7mm ± 2mm
194 | RT13-24-012 | ENDOTRACHEAL TUBE | DOUBLE LUMEN ENDOBRONCHIAL TUBE | DOUBLE LUMEN ENDOBRONCHIAL TUBE- RIGHT Purpose: For isolation of one lung to prevent contamination, or for ventilation of one lung in thoracic surgery Consists of a double Lumen clear PVC tube that is thermo-sensitive with radiopaque lines and markers at key features. Markings starting at 26cm from distal end marked every 2cm up to and about 32cm Tubes to be flexible enough to allow manipulation through distorted airways, but without luminal collapse. Must have a smooth transition of the inner lumen without narrowing, to allow easy passage of flexible fibre-optic-bronchoscope. Tracheal and Bronchial extensions on the main tube must be colour coded - blue for bronchial, clear for tracheal Stylet: Malleable shape forming - packaged in place down the bronchial lumen Supplied with 2 proximal right-angled connectors with a Bronchoscopy/Aspiration Port The port must have a toggled cap and allow for easy passage of either a flexible fibre-optic-bronchoscope, or a suction catheter – both without a leak. The proximal angled connectors to connect to the bronchial and tracheal tubes and the distal connectors to connect to a standard Y-connector that is compatible with a standard breathing circuit. 2 - 4 Suction catheters: with suction control, to fit the total length of the tube Colourless tracheal high-volume-low-pressure polyvinyl chloride cuff – with colourless pilot balloon, clearly labelled “tracheal” Blue bronchial low pressure Polyurethane cuff - with blue pilot balloon, clearly labelled “bronchial”. Bronchial cuff to be saddle-shaped One-way valves in pilot balloons. The cuffs must inflate and deflate easily, and the pilot cuffs must indicate the degree of inflation accurately. Cuffs must be resistant to the effects of nitrous oxide. When deflated, cuffs must conform closely to the tube to minimise the risk of trauma during intubation and extubation. All components must be latex-free, sterile and individually packed in peel pouch - SINGLE use only To comply with ISO 16628 Information brochure or the outer package of the item should clearly indicate latex free | Double Lumen Endobronchial tube: Right Size: 37Fr Shaft Length: 420mm ID: > 5.0mm (bilaterally) OD: < 12.3mm ± 2mm
195 | RT13-24-013 | ENDOTRACHEAL TUBE | DOUBLE LUMEN ENDOBRONCHIAL TUBE | DOUBLE LUMEN ENDOBRONCHIAL TUBE- RIGHT Purpose: For isolation of one lung to prevent contamination, or for ventilation of one lung in thoracic surgery Consists of a double Lumen clear PVC tube that is thermo-sensitive with radiopaque lines and markers at key features. Markings starting at 26cm from distal end marked every 2cm up to and about 32cm Tubes to be flexible enough to allow manipulation through distorted airways, but without luminal collapse. Must have a smooth transition of the inner lumen without narrowing, to allow easy passage of flexible fibre-optic-bronchoscope. Tracheal and Bronchial extensions on the main tube must be colour coded - blue for bronchial, clear for tracheal Stylet: Malleable shape forming - packaged in place down the bronchial lumen Supplied with 2 proximal right-angled connectors with a Bronchoscopy/Aspiration Port The port must have a toggled cap and allow for easy passage of either a flexible fibre-optic-bronchoscope, or a suction catheter – both without a leak. The proximal angled connectors to connect to the bronchial and tracheal tubes and the distal connectors to connect to a standard Y-connector that is compatible with a standard breathing circuit. 2 - 4 Suction catheters: with suction control, to fit the total length of the tube Colourless tracheal high-volume-low-pressure polyvinyl chloride cuff – with colourless pilot balloon, clearly labelled “tracheal” Blue bronchial low pressure Polyurethane cuff - with blue pilot balloon, clearly labelled “bronchial”. Bronchial cuff to be saddle-shaped One-way valves in pilot balloons. The cuffs must inflate and deflate easily, and the pilot cuffs must indicate the degree of inflation accurately. Cuffs must be resistant to the effects of nitrous oxide. When deflated, cuffs must conform closely to the tube to minimise the risk of trauma during intubation and extubation. All components must be latex-free, sterile and individually packed in peel pouch - SINGLE use only To comply with ISO 16628 Information brochure or the outer package of the item should clearly indicate latex free | Double Lumen Endobronchial tube: Right Size: 39Fr Shaft Length: 420mm ID: > 5.2mm (bilaterally) OD: < 13mm ± 2mm
196 | RT13-24-014 | ENDOTRACHEAL TUBE | DOUBLE LUMEN ENDOBRONCHIAL TUBE | DOUBLE LUMEN ENDOBRONCHIAL TUBE- RIGHT Purpose: For isolation of one lung to prevent contamination, or for ventilation of one lung in thoracic surgery Consists of a double Lumen clear PVC tube that is thermo-sensitive with radiopaque lines and markers at key features. Markings starting at 26cm from distal end marked every 2cm up to and about 32cm Tubes to be flexible enough to allow manipulation through distorted airways, but without luminal collapse. Must have a smooth transition of the inner lumen without narrowing, to allow easy passage of flexible fibre-optic-bronchoscope. Tracheal and Bronchial extensions on the main tube must be colour coded - blue for bronchial, clear for tracheal Stylet: Malleable shape forming - packaged in place down the bronchial lumen Supplied with 2 proximal right-angled connectors with a Bronchoscopy/Aspiration Port The port must have a toggled cap and allow for easy passage of either a flexible fibre-optic-bronchoscope, or a suction catheter – both without a leak. The proximal angled connectors to connect to the bronchial and tracheal tubes and the distal connectors to connect to a standard Y-connector that is compatible with a standard breathing circuit. 2 - 4 Suction catheters: with suction control, to fit the total length of the tube Colourless tracheal high-volume-low-pressure polyvinyl chloride cuff – with colourless pilot balloon, clearly labelled “tracheal” Blue bronchial low pressure Polyurethane cuff - with blue pilot balloon, clearly labelled “bronchial”. Bronchial cuff to be saddle-shaped One-way valves in pilot balloons. The cuffs must inflate and deflate easily, and the pilot cuffs must indicate the degree of inflation accurately. Cuffs must be resistant to the effects of nitrous oxide. When deflated, cuffs must conform closely to the tube to minimise the risk of trauma during intubation and extubation. All components must be latex-free, sterile and individually packed in peel pouch - SINGLE use only To comply with ISO 16628 Information brochure or the outer package of the item should clearly indicate latex free | Double Lumen Endobronchial tube: Right Size: 41Fr Shaft Length: 420mm ID: > 5.3mm (bilaterally) OD: < 13.7mm ± 2mm
197 | RT13-25-001 | ENDOTRACHEAL TUBE | COMBINED ENDOTRACHEAL TUBE AND ENDOBRONCHIAL BLOCKER | COMBINED ENDOTRACHEAL TUBE AND ENDOBRONCHIAL BLOCKER Purpose: For isolation of one lung to prevent contamination, or for ventilation of one lung in thoracic surgery Consists of a single main lumen endotracheal tube for ventilation and a secondary integrated lumen provided to guide a bi-directional, torque controlled bronchus blocker. The tube should function as a conventional endotracheal tube once the bronchus blocker is removed The blocker must allow occlusion of either the right or the left bronchus or for selective lobar blockade The endobronchial blocker tube should be hollow and allow CPAP or high frequency jet ventilation, irrigation, suctioning and/or gas sampling from the occluded lung. The endobronchial blocker should have a malleable shaft with high torque control for manipulation, a soft distal tip with a high volume low pressure cuff, and should extend 8-10cm beyond the tip. All components must be latex-free, sterile and individually packed in peel pouch - SINGLE use only Information brochure or the outer package of the item should clearly indicate latex free | Double Lumen Endobronchial tube: Right Size: 26Fr Shaft Length: 420mm ID: > mm (bilaterally) OD: <
198 | RT13-25-002 | ENDOTRACHEAL TUBE | COMBINED ENDOTRACHEAL TUBE AND ENDOBRONCHIAL BLOCKER | COMBINED ENDOTRACHEAL TUBE AND ENDOBRONCHIAL BLOCKER Purpose: For isolation of one lung to prevent contamination, or for ventilation of one lung in thoracic surgery Consists of a single main lumen endotracheal tube for ventilation and a secondary integrated lumen provided to guide a bi-directional, torque controlled bronchus blocker. The tube should function as a conventional endotracheal tube once the bronchus blocker is removed The blocker must allow occlusion of either the right or the left bronchus or for selective lobar blockade The endobronchial blocker tube should be hollow and allow CPAP or high frequency jet ventilation, irrigation, suctioning and/or gas sampling from the occluded lung. The endobronchial blocker should have a malleable shaft with high torque control for manipulation, a soft distal tip with a high volume low pressure cuff, and should extend 8-10cm beyond the tip. All components must be latex-free, sterile and individually packed in peel pouch - SINGLE use only Information brochure or the outer package of the item should clearly indicate latex free | Double Lumen Endobronchial tube: Right Size: 28Fr Shaft Length: 420mm ID: > 3.2mm (bilaterally) OD: < 9.3mm ± 2mm
199 | RT13-25-003 | ENDOTRACHEAL TUBE | COMBINED ENDOTRACHEAL TUBE AND ENDOBRONCHIAL BLOCKER | COMBINED ENDOTRACHEAL TUBE AND ENDOBRONCHIAL BLOCKER Purpose: For isolation of one lung to prevent contamination, or for ventilation of one lung in thoracic surgery Consists of a single main lumen endotracheal tube for ventilation and a secondary integrated lumen provided to guide a bi-directional, torque controlled bronchus blocker. The tube should function as a conventional endotracheal tube once the bronchus blocker is removed The blocker must allow occlusion of either the right or the left bronchus or for selective lobar blockade The endobronchial blocker tube should be hollow and allow CPAP or high frequency jet ventilation, irrigation, suctioning and/or gas sampling from the occluded lung. The endobronchial blocker should have a malleable shaft with high torque control for manipulation, a soft distal tip with a high volume low pressure cuff, and should extend 8-10cm beyond the tip. All components must be latex-free, sterile and individually packed in peel pouch - SINGLE use only Information brochure or the outer package of the item should clearly indicate latex free | Double Lumen Endobronchial tube: Right Size: 32Fr Shaft Length: 420mm ID: > 3.4mm (bilaterally) OD: < 10.7mm ± 2mm
200 | RT13-25-004 | ENDOTRACHEAL TUBE | COMBINED ENDOTRACHEAL TUBE AND ENDOBRONCHIAL BLOCKER | COMBINED ENDOTRACHEAL TUBE AND ENDOBRONCHIAL BLOCKER Purpose: For isolation of one lung to prevent contamination, or for ventilation of one lung in thoracic surgery Consists of a single main lumen endotracheal tube for ventilation and a secondary integrated lumen provided to guide a bi-directional, torque controlled bronchus blocker. The tube should function as a conventional endotracheal tube once the bronchus blocker is removed The blocker must allow occlusion of either the right or the left bronchus or for selective lobar blockade The endobronchial blocker tube should be hollow and allow CPAP or high frequency jet ventilation, irrigation, suctioning and/or gas sampling from the occluded lung. The endobronchial blocker should have a malleable shaft with high torque control for manipulation, a soft distal tip with a high volume low pressure cuff, and should extend 8-10cm beyond the tip. All components must be latex-free, sterile and individually packed in peel pouch - SINGLE use only Information brochure or the outer package of the item should clearly indicate latex free | Double Lumen Endobronchial tube: Right Size: 35Fr Shaft Length: 420mm ID: > 4.7mm (bilaterally) OD: < 11.7mm ± 2mm
201 | RT13-27-003 | AIRWAY DEVICES | MICROLARYNGEAL TUBE | MICROLARYNGEAL TUBES Purpose: Longer tube for use in narrowed airways or laryngeal surgery Consists of a clear thermo-sensitive extra-long PVC tube that is kink resistant, atraumatic to airways and have a full Magill curvature. Must have a radiopaque line and must be marked with numbered depth markings in 1 cm intervals, tube size, length, ID, OD, trade name. The markings on the tube must be clear and must not rub off. Proximal end - secure 15mm OD circuit connector made from medical grade plastic Distal end - bevelled atraumatic tip with a murphy eye on opposite side to bevel. The high-volume-low-pressure polyurethane cuff with pilot cuff . The pilot cuff must have a one-way valve and must indicate the degree of inflation accurately. The cuff must inflate and deflate easily, and when deflated must conform closely to tube to minimise the risk of trauma during intubation, and extubation and not increase the OD of the tube. Must be appropriately sized in relation to the tube size and appropriately placed with regards to the distance from the distal tip. The cuff must be resistant to Nitrous Oxide effects Cuff thickness ± 50 micron All components must be latex-free sterile and individually packed in peel pouch that is easy to open - For SINGLE use only To comply with the latest ISO 5361 Information brochure or the outer package of the item should clearly indicate latex free | Cuffed MicroLaryngeal Oral/Nasal tracheal tube with Murphy eye Size: ID: 6.0mm OD: 8.2mm Tube length ≥ 320mm
202 | RT13-28-001 | Cuffed and Uncuffed Tubes | UNCUFFED STAINLESS STEEL LASER ENDOTRACHEAL TUBES | UNCUFFED STAINLESS STEEL LASER ENDOTRACHEAL TUBES WITHOUT A MURPHY EYE Purpose: Endotracheal tube for use during laser airway surgery. Consists of an airtight stainless steel endotracheal tube that ends in a permanently attached 15mm connector proximally. The tube must have a Magill curvature, be flexible but kink-resistant to allow it to be positioned outside of the surgical field. The stainless steel design must have a matte finish to defocus the reflected laser beams. The tube must be resistant to C02 and KTP lasers and non-flammable. Company must produce evidence of laser power settings at which ignition does not occur. Distal end must end in a medical grade plastic atraumatic left-facing bevelled tip without a Murphy eye. Must have a precision band mark to aid with depth of intubation and guide proper placement. Solid or circumferential precision band markings accepted Packaging must be marked with tube size, length, ID, OD, trade name. . All components must be sterile and individually packed in peel pouch that is easy to open - For SINGLE use only To comply with the latest ISO 11990 Information brochure or the outer package of the item should be laser compatible | Uncuffed LASER Endotracheal Tubes without a Murphy Eye ID: 3.0mm OD: 5.2mm
203 | RT13-28-002 | Cuffed and Uncuffed Tubes | UNCUFFED STAINLESS STEEL LASER ENDOTRACHEAL TUBES | UNCUFFED STAINLESS STEEL LASER ENDOTRACHEAL TUBES WITHOUT A MURPHY EYE Purpose: Endotracheal tube for use during laser airway surgery. Consists of an airtight stainless steel endotracheal tube that ends in a permanently attached 15mm connector proximally. The tube must have a Magill curvature, be flexible but kink-resistant to allow it to be positioned outside of the surgical field. The stainless steel design must have a matte finish to defocus the reflected laser beams. The tube must be resistant to C02 and KTP lasers and non-flammable. Company must produce evidence of laser power settings at which ignition does not occur. Distal end must end in a medical grade plastic atraumatic left-facing bevelled tip without a Murphy eye. Must have a precision band mark to aid with depth of intubation and guide proper placement. Solid or circumferential precision band markings accepted Packaging must be marked with tube size, length, ID, OD, trade name. . All components must be sterile and individually packed in peel pouch that is easy to open - For SINGLE use only To comply with the latest ISO 11990 Information brochure or the outer package of the item should be laser compatible | Uncuffed LASER Endotracheal Tubes without a Murphy Eye ID: 3.5mm OD: 5.7mm
204 | RT13-28-003 | Cuffed and Uncuffed Tubes | UNCUFFED STAINLESS STEEL LASER ENDOTRACHEAL TUBES | UNCUFFED STAINLESS STEEL LASER ENDOTRACHEAL TUBES WITHOUT A MURPHY EYE Purpose: Endotracheal tube for use during laser airway surgery. Consists of an airtight stainless steel endotracheal tube that ends in a permanently attached 15mm connector proximally. The tube must have a Magill curvature, be flexible but kink-resistant to allow it to be positioned outside of the surgical field. The stainless steel design must have a matte finish to defocus the reflected laser beams. The tube must be resistant to C02 and KTP lasers and non-flammable. Company must produce evidence of laser power settings at which ignition does not occur. Distal end must end in a medical grade plastic atraumatic left-facing bevelled tip without a Murphy eye. Must have a precision band mark to aid with depth of intubation and guide proper placement. Solid or circumferential precision band markings accepted Packaging must be marked with tube size, length, ID, OD, trade name. . All components must be sterile and individually packed in peel pouch that is easy to open - For SINGLE use only To comply with the latest ISO 11990 Information brochure or the outer package of the item should be laser compatible | Uncuffed LASER Endotracheal Tubes without a Murphy Eye ID: 4.0mm OD: 6.1mm
205 | RT13-29-001 | ENDOTRACHEAL TUBE | DOUBLE-CUFFED STAINLESS STEEL LASER ENDOTRACHEAL TUBES | DOUBLE-CUFFED STAINLESS STEEL LASER ENDOTRACHEAL TUBES -FOR VENTILATION DURING CO2 AND KTP LASER Purpose: Endotracheal tube for use during laser airway surgery. Consists of an airtight stainless steel endotracheal tube that ends in a permanently attached 15mm connector proximally. The tube must have a Magill curvature, be flexible but kink-resistant to allow it to be positioned outside of the surgical field. The stainless steel design must have a matte finish to defocus the reflected laser beams. Must have two cuffs that can be filled with saline to seal the airway and maintain ventilation should one cuff be punctured. The pilot balloons for the cuffs, must be colour-coded, have a self sealing valve and must indicate proximal or distal. The tube must be resistant to C02 and KTP lasers and non-flammable. Company must produce evidence of laser power settings at which ignition does not occur. Distal end must end in a medical grade plastic atraumatic left-facing bevelled tip with Murphy eye on the opposite side Must have a precision band mark to aid with depth of intubation and guide proper placement. Solid or circumferential precision band markings accepted Packaging must be marked with tube size, length, ID, OD, trade name. . All components must be sterile and individually packed in peel pouch that is easy to open - For SINGLE use only To comply with the latest ISO 11990 Information brochure or the outer package of the item should be laser compatible | Double-Cuffed Stainless Steel LASER Endotracheal Tubes ID: 4.5mm OD: 7.0mm
206 | RT13-29-002 | ENDOTRACHEAL TUBE | DOUBLE-CUFFED STAINLESS STEEL LASER ENDOTRACHEAL TUBES | DOUBLE-CUFFED STAINLESS STEEL LASER ENDOTRACHEAL TUBES -FOR VENTILATION DURING CO2 AND KTP LASER Purpose: Endotracheal tube for use during laser airway surgery. Consists of an airtight stainless steel endotracheal tube that ends in a permanently attached 15mm connector proximally. The tube must have a Magill curvature, be flexible but kink-resistant to allow it to be positioned outside of the surgical field. The stainless steel design must have a matte finish to defocus the reflected laser beams. Must have two cuffs that can be filled with saline to seal the airway and maintain ventilation should one cuff be punctured. The pilot balloons for the cuffs, must be colour-coded, have a self sealing valve and must indicate proximal or distal. The tube must be resistant to C02 and KTP lasers and non-flammable. Company must produce evidence of laser power settings at which ignition does not occur. Distal end must end in a medical grade plastic atraumatic left-facing bevelled tip with Murphy eye on the opposite side Must have a precision band mark to aid with depth of intubation and guide proper placement. Solid or circumferential precision band markings accepted Packaging must be marked with tube size, length, ID, OD, trade name. . All components must be sterile and individually packed in peel pouch that is easy to open - For SINGLE use only To comply with the latest ISO 11990 Information brochure or the outer package of the item should be laser compatible | Double-Cuffed Stainless Steel LASER Endotracheal Tubes ID: 5.0mm OD: 7.5mm
207 | RT13-29-003 | ENDOTRACHEAL TUBE | DOUBLE-CUFFED STAINLESS STEEL LASER ENDOTRACHEAL TUBES | DOUBLE-CUFFED STAINLESS STEEL LASER ENDOTRACHEAL TUBES -FOR VENTILATION DURING CO2 AND KTP LASER Purpose: Endotracheal tube for use during laser airway surgery. Consists of an airtight stainless steel endotracheal tube that ends in a permanently attached 15mm connector proximally. The tube must have a Magill curvature, be flexible but kink-resistant to allow it to be positioned outside of the surgical field. The stainless steel design must have a matte finish to defocus the reflected laser beams. Must have two cuffs that can be filled with saline to seal the airway and maintain ventilation should one cuff be punctured. The pilot balloons for the cuffs, must be colour-coded, have a self sealing valve and must indicate proximal or distal. The tube must be resistant to C02 and KTP lasers and non-flammable. Company must produce evidence of laser power settings at which ignition does not occur. Distal end must end in a medical grade plastic atraumatic left-facing bevelled tip with Murphy eye on the opposite side Must have a precision band mark to aid with depth of intubation and guide proper placement. Solid or circumferential precision band markings accepted Packaging must be marked with tube size, length, ID, OD, trade name. . All components must be sterile and individually packed in peel pouch that is easy to open - For SINGLE use only To comply with the latest ISO 11990 Information brochure or the outer package of the item should be laser compatible | Double-Cuffed Stainless Steel LASER Endotracheal Tubes ID: 5.5mm OD: 7.9mm
208 | RT13-29-004 | ENDOTRACHEAL TUBE | DOUBLE-CUFFED STAINLESS STEEL LASER ENDOTRACHEAL TUBES | DOUBLE-CUFFED STAINLESS STEEL LASER ENDOTRACHEAL TUBES -FOR VENTILATION DURING CO2 AND KTP LASER Purpose: Endotracheal tube for use during laser airway surgery. Consists of an airtight stainless steel endotracheal tube that ends in a permanently attached 15mm connector proximally. The tube must have a Magill curvature, be flexible but kink-resistant to allow it to be positioned outside of the surgical field. The stainless steel design must have a matte finish to defocus the reflected laser beams. Must have two cuffs that can be filled with saline to seal the airway and maintain ventilation should one cuff be punctured. The pilot balloons for the cuffs, must be colour-coded, have a self sealing valve and must indicate proximal or distal. The tube must be resistant to C02 and KTP lasers and non-flammable. Company must produce evidence of laser power settings at which ignition does not occur. Distal end must end in a medical grade plastic atraumatic left-facing bevelled tip with Murphy eye on the opposite side Must have a precision band mark to aid with depth of intubation and guide proper placement. Solid or circumferential precision band markings accepted Packaging must be marked with tube size, length, ID, OD, trade name. . All components must be sterile and individually packed in peel pouch that is easy to open - For SINGLE use only To comply with the latest ISO 11990 Information brochure or the outer package of the item should be laser compatible | Double-Cuffed Stainless Steel LASER Endotracheal Tubes ID: 6.0mm OD: 8.5mm
209 | RT13-29-005 | ENDOTRACHEAL TUBE | DOUBLE-CUFFED RUBBER LASER ENDOTRACHEAL TUBES | DOUBLE-CUFFED RUBBER LASER ENDOTRACHEAL TUBES -FOR VENTILATION DURING ARGON, ND/YAG AND CO2 TYPE LASER Purpose: Endotracheal tube for use during laser airway surgery. Consists of a soft natural rubber tube with micro-corrugated silver foil that ends in a 15mm connector. The tube must have a Magill curvature, be flexible but kink-resistant to allow it to be positioned outside of the surgical field. The tube must be marked with tube size, length, ID, OD, trade name. . Proximal end - secure 15mm OD circuit connector made from medical grade plastic Distal end - Atraumatic left-facing bevelled tip with Murphy eye on the opposite side Must have two cuffs that must be able to be saline filled to seal the airway and maintain ventilation should one cuff be punctured. The cuffs must be resistant to the diffusion of anesthetic gas. The tube must be resistant to Argon, Nd/YAG and CO2 type laser in the Wave range of 0.488mm to 10.6mm. Company must produce evidence of laser power settings at which ignition does not occur. All components must be sterile and individually packed in peel pouch that is easy to open - For SINGLE use only To comply with the latest ISO 11990 Information brochure or the outer package of the item should be laser compatible | Double-Cuffed Rubber LASER Endotracheal Tubes ID: 4.0mm OD: 8.0mm Length: 400mm
210 | RT13-29-006 | ENDOTRACHEAL TUBE | DOUBLE-CUFFED RUBBER LASER ENDOTRACHEAL TUBES | DOUBLE-CUFFED RUBBER LASER ENDOTRACHEAL TUBES -FOR VENTILATION DURING ARGON, ND/YAG AND CO2 TYPE LASER Purpose: Endotracheal tube for use during laser airway surgery. Consists of a soft natural rubber tube with micro-corrugated silver foil that ends in a 15mm connector. The tube must have a Magill curvature, be flexible but kink-resistant to allow it to be positioned outside of the surgical field. The tube must be marked with tube size, length, ID, OD, trade name. . Proximal end - secure 15mm OD circuit connector made from medical grade plastic Distal end - Atraumatic left-facing bevelled tip with Murphy eye on the opposite side Must have two cuffs that must be able to be saline filled to seal the airway and maintain ventilation should one cuff be punctured. The cuffs must be resistant to the diffusion of anesthetic gas. The tube must be resistant to Argon, Nd/YAG and CO2 type laser in the Wave range of 0.488mm to 10.6mm. Company must produce evidence of laser power settings at which ignition does not occur. All components must be sterile and individually packed in peel pouch that is easy to open - For SINGLE use only To comply with the latest ISO 11990 Information brochure or the outer package of the item should be laser compatible | Double-Cuffed Rubber LASER Endotracheal Tubes ID: 5.0mm OD: 9.0mm Length: 400mm
211 | RT13-29-007 | ENDOTRACHEAL TUBE | DOUBLE-CUFFED RUBBER LASER ENDOTRACHEAL TUBES | DOUBLE-CUFFED RUBBER LASER ENDOTRACHEAL TUBES -FOR VENTILATION DURING ARGON, ND/YAG AND CO2 TYPE LASER Purpose: Endotracheal tube for use during laser airway surgery. Consists of a soft natural rubber tube with micro-corrugated silver foil that ends in a 15mm connector. The tube must have a Magill curvature, be flexible but kink-resistant to allow it to be positioned outside of the surgical field. The tube must be marked with tube size, length, ID, OD, trade name. . Proximal end - secure 15mm OD circuit connector made from medical grade plastic Distal end - Atraumatic left-facing bevelled tip with Murphy eye on the opposite side Must have two cuffs that must be able to be saline filled to seal the airway and maintain ventilation should one cuff be punctured. The cuffs must be resistant to the diffusion of anesthetic gas. The tube must be resistant to Argon, Nd/YAG and CO2 type laser in the Wave range of 0.488mm to 10.6mm. Company must produce evidence of laser power settings at which ignition does not occur. All components must be sterile and individually packed in peel pouch that is easy to open - For SINGLE use only To comply with the latest ISO 11990 Information brochure or the outer package of the item should be laser compatible | Double-Cuffed Rubber LASER Endotracheal Tubes ID: 6.0mm OD: 10.8mm Length: 400mm
212 | RT13-29-008 | ENDOTRACHEAL TUBE | DOUBLE-CUFFED RUBBER LASER ENDOTRACHEAL TUBES | DOUBLE-CUFFED RUBBER LASER ENDOTRACHEAL TUBES -FOR VENTILATION DURING ARGON, ND/YAG AND CO2 TYPE LASER Purpose: Endotracheal tube for use during laser airway surgery. Consists of a soft natural rubber tube with micro-corrugated silver foil that ends in a 15mm connector. The tube must have a Magill curvature, be flexible but kink-resistant to allow it to be positioned outside of the surgical field. The tube must be marked with tube size, length, ID, OD, trade name. . Proximal end - secure 15mm OD circuit connector made from medical grade plastic Distal end - Atraumatic left-facing bevelled tip with Murphy eye on the opposite side Must have two cuffs that must be able to be saline filled to seal the airway and maintain ventilation should one cuff be punctured. The cuffs must be resistant to the diffusion of anesthetic gas. The tube must be resistant to Argon, Nd/YAG and CO2 type laser in the Wave range of 0.488mm to 10.6mm. Company must produce evidence of laser power settings at which ignition does not occur. All components must be sterile and individually packed in peel pouch that is easy to open - For SINGLE use only To comply with the latest ISO 11990 Information brochure or the outer package of the item should be laser compatible | Double-Cuffed Rubber LASER Endotracheal Tubes ID: 7.0mm OD: 12.3mm Length: 400mm
213 | RT13-29-009 | ENDOTRACHEAL TUBE | DOUBLE-CUFFED RUBBER LASER ENDOTRACHEAL TUBES | DOUBLE-CUFFED RUBBER LASER ENDOTRACHEAL TUBES -FOR VENTILATION DURING ARGON, ND/YAG AND CO2 TYPE LASER Purpose: Endotracheal tube for use during laser airway surgery. Consists of a soft natural rubber tube with micro-corrugated silver foil that ends in a 15mm connector. The tube must have a Magill curvature, be flexible but kink-resistant to allow it to be positioned outside of the surgical field. The tube must be marked with tube size, length, ID, OD, trade name. . Proximal end - secure 15mm OD circuit connector made from medical grade plastic Distal end - Atraumatic left-facing bevelled tip with Murphy eye on the opposite side Must have two cuffs that must be able to be saline filled to seal the airway and maintain ventilation should one cuff be punctured. The cuffs must be resistant to the diffusion of anesthetic gas. The tube must be resistant to Argon, Nd/YAG and CO2 type laser in the Wave range of 0.488mm to 10.6mm. Company must produce evidence of laser power settings at which ignition does not occur. All components must be sterile and individually packed in peel pouch that is easy to open - For SINGLE use only To comply with the latest ISO 11990 Information brochure or the outer package of the item should be laser compatible | Double-Cuffed Rubber LASER Endotracheal Tubes ID: 8.0mm OD: 13.4mm Length: 400mm
214 | RT13-30-019 | ENDOTRACHEAL TUBE | DOUBLE-CUFFED RUBBER LASER ENDOTRACHEAL TUBES | DOUBLE-CUFFED RUBBER LASER ENDOTRACHEAL TUBES -FOR VENTILATION DURING ARGON, ND/YAG AND CO2 TYPE LASER Purpose: Endotracheal tube for use during laser airway surgery. Consists of a soft natural rubber tube with micro-corrugated silver foil that ends in a 15mm connector. The tube must have a Magill curvature, be flexible but kink-resistant to allow it to be positioned outside of the surgical field. The tube must be marked with tube size, length, ID, OD, trade name. . Proximal end - secure 15mm OD circuit connector made from medical grade plastic Distal end - Atraumatic left-facing bevelled tip with Murphy eye on the opposite side Must have two cuffs that must be able to be saline filled to seal the airway and maintain ventilation should one cuff be punctured. The cuffs must be resistant to the diffusion of anesthetic gas. The tube must be resistant to Argon, Nd/YAG and CO2 type laser in the Wave range of 0.488mm to 10.6mm. Company must produce evidence of laser power settings at which ignition does not occur. All components must be sterile and individually packed in peel pouch that is easy to open - For SINGLE use only To comply with the latest ISO 11990 Information brochure or the outer package of the item should be laser compatible | Double-Cuffed Stainless Steel LASER Endotracheal Tubes ID: 5.5mm OD: 7.9mm
215 | RT13-30-021 | ENDOTRACHEAL TUBE | DOUBLE-CUFFED RUBBER LASER ENDOTRACHEAL TUBES | DOUBLE-CUFFED RUBBER LASER ENDOTRACHEAL TUBES -FOR VENTILATION DURING ARGON, ND/YAG AND CO2 TYPE LASER Purpose: Endotracheal tube for use during laser airway surgery. Consists of a soft natural rubber tube with micro-corrugated silver foil that ends in a 15mm connector. The tube must have a Magill curvature, be flexible but kink-resistant to allow it to be positioned outside of the surgical field. The tube must be marked with tube size, length, ID, OD, trade name. . Proximal end - secure 15mm OD circuit connector made from medical grade plastic Distal end - Atraumatic left-facing bevelled tip with Murphy eye on the opposite side Must have two cuffs that must be able to be saline filled to seal the airway and maintain ventilation should one cuff be punctured. The cuffs must be resistant to the diffusion of anesthetic gas. The tube must be resistant to Argon, Nd/YAG and CO2 type laser in the Wave range of 0.488mm to 10.6mm. Company must produce evidence of laser power settings at which ignition does not occur. All components must be sterile and individually packed in peel pouch that is easy to open - For SINGLE use only To comply with the latest ISO 11990 Information brochure or the outer package of the item should be laser compatible | Double-Cuffed Stainless Steel LASER Endotracheal Tubes ID: 6.0mm OD: 8.5mm
216 | RT13-30-023 | ENDOTRACHEAL TUBE | DOUBLE-CUFFED RUBBER LASER ENDOTRACHEAL TUBES | DOUBLE-CUFFED RUBBER LASER ENDOTRACHEAL TUBES -FOR VENTILATION DURING ARGON, ND/YAG AND CO2 TYPE LASER Purpose: Endotracheal tube for use during laser airway surgery. Consists of a soft natural rubber tube with micro-corrugated silver foil that ends in a 15mm connector. The tube must have a Magill curvature, be flexible but kink-resistant to allow it to be positioned outside of the surgical field. The tube must be marked with tube size, length, ID, OD, trade name. . Proximal end - secure 15mm OD circuit connector made from medical grade plastic Distal end - Atraumatic left-facing bevelled tip with Murphy eye on the opposite side Must have two cuffs that must be able to be saline filled to seal the airway and maintain ventilation should one cuff be punctured. The cuffs must be resistant to the diffusion of anesthetic gas. The tube must be resistant to Argon, Nd/YAG and CO2 type laser in the Wave range of 0.488mm to 10.6mm. Company must produce evidence of laser power settings at which ignition does not occur. All components must be sterile and individually packed in peel pouch that is easy to open - For SINGLE use only To comply with the latest ISO 11990 Information brochure or the outer package of the item should be laser compatible | Double-Cuffed Rubber LASER Endotracheal Tubes ID: 4.0mm OD: 8.0mm Length: 400mm
217 | RT13-30-032 | AIRWAY DEVICES | 1st GENERATION SUPRAGLOTTIC AIRWAY DEVICE WITH APERTURE BARS | 1st GENERATION SUPRAGLOTTIC AIRWAY DEVICE WITH APERTURE BARS AND REINFORCED TUBING - SINGLE USE Purpose: 1st Generation Supraglottic Airway Device control in a spontaneously breathing or ventilated patient. Consists of a kink resistant wire reinforced flexible airway tube ending in a soft silicone elliptical pharyngeal cuff with a one-way valve and bulb for inflation and deflation of the cuff. The cuff must be suitable for positive pressure ventilation up to 20cm/H20. Must have aperture bars to prevent obstruction and protect the epiglottis. The tip should resist folding over during insertion. The airway tube must have a smooth transition to the posterior aspect of the boot of the cuff and a 15mm connector proximally Must have clear markings that do not rub off, to indicate Trade name, Size, Recommended patient weight There must also be a guiding line on the posterior aspect to assist with orientation Manufactured from latex-free medical grade PVC /Silicone/ Styrene-ethylene-butylene-styrene (SEBS) Must be sterile and individually packed in peel pouch that is easy to open - For SINGLE use only Must comply with CE 0086 Information brochure or the outer package of the item should clearly indicate latex free | 1st Generation Supraglottic Airway Device with Aperture Bars and Reinforced Tubing - Single Use Size: 6 For weight: > 100kg
218 | RT13-31-008 | AIRWAY DEVICES | 2nd GENERATION SUPRAGLOTTIC AIRWAY DEVICE WITH INFLATABLE CUFF- SINGLE USE | 2nd GENERATION SUPRAGLOTTIC AIRWAY DEVICE WITH INFLATABLE CUFF AND A GASTRIC OUTLET - SINGLE USE Purpose: 2nd Generation Supraglottic airway device in a spontaneously breathing or ventilated patient. Consists of an preformed anatomically curved, kink resistant soft airway tube ending in a soft elliptical silicone pharyngeal cuff. The airway tube must allow for easy one hand insertion, must have an integral bite block, have a 15mm connector proximally and a smooth transition to the posterior aspect of the boot distally The cuff must create a high-pressure seal and must be suitable for positive pressure ventilation of at least 30cm/H20 The tip should resist folding over during insertion. Must NOT have aperture bars. The proximal end of the gastric outlet tube must be able to connect directly to a suction source or allow a suction catheter to be passed for gastric deflation as required. This gastric outlet must be transparent. The gastric drain outlet must open at the tip of the elliptical cuff distally. The cuff must have a one-way valve and bulb that indicates the degree of inflation and deflation of pharyngeal cuff. The device must have clear markings that do not rub off to indicate Trade name; Size; Recommended patient weight Guideline on posterior aspect to assist with orientation Manufactured from latex-free medical grade PVC /Silicone/ Styrene-ethylene-butylene-styrene (SEBS) Must be sterile and individually packed in peel pouch that is easy to open - For SINGLE use only Must comply with CE 0086 Size 6 considered individually, Information brochure or the outer package of the item should clearly indicate latex free | 2nd Generation Supraglottic Airway Device with Inflatable Cuff and a Gastric Outlet - Single Use Size: 6 For weight: > 100kg Gastric drain to allow 14Fr catheter
219 | RT13-36-004 | AIRWAY DEVICES | BLIND INSERTION AIRWAY | BLIND INSERTION AIRWAY DEVICES: DOUBLE-CUFFED Purpose: Dual lumen tube designed for difficult or emergency airway intubation and ventilation without the use of a laryngoscope Consists of a cuffed, DOUBLE -cuffed with a soft non-traumatic tip. High-volume proximal cuff for oro and nasopharyngeal occlusion. The tube must be held firmly in place by the proximal cuff and it must prevent the escape of gases through the nose or mouth Low-volume distal cuff for oesophageal occlusion Colour coded pilot balloons and 2 syringes prefilled with the correct amount of air for the inflation and deflation of the proximal and distal cuffs Bilateral ventilation ports between the two cuffs The tube must allow for rapid intubation of difficult airways and should be a viable airway whether it is placed in the trachea or in the oesophagus It should also allow for a suction catheter into the oesophagus when the distal cuff is placed in the oesophagus All components must be latex-free sterile and individually packed in peel pouch that is easy to open - For SINGLE use only Must comply with ISO 11712 Information brochure or the outer package of the item should clearly indicate latex free | Blind Insertion Airway Device: Single-Lumen Double-Cuffed Size: 1 Infant patient's: 5-12 kg
220 | RT13-36-005 | AIRWAY DEVICES | BLIND INSERTION AIRWAY | BLIND INSERTION AIRWAY DEVICES: DOUBLE-CUFFED Purpose: Dual lumen tube designed for difficult or emergency airway intubation and ventilation without the use of a laryngoscope Consists of a cuffed, DOUBLE -cuffed with a soft non-traumatic tip. High-volume proximal cuff for oro and nasopharyngeal occlusion. The tube must be held firmly in place by the proximal cuff and it must prevent the escape of gases through the nose or mouth Low-volume distal cuff for oesophageal occlusion Colour coded pilot balloons and 2 syringes prefilled with the correct amount of air for the inflation and deflation of the proximal and distal cuffs Bilateral ventilation ports between the two cuffs The tube must allow for rapid intubation of difficult airways and should be a viable airway whether it is placed in the trachea or in the oesophagus It should also allow for a suction catheter into the oesophagus when the distal cuff is placed in the oesophagus All components must be latex-free sterile and individually packed in peel pouch that is easy to open - For SINGLE use only Must comply with ISO 11712 Information brochure or the outer package of the item should clearly indicate latex free | Blind Insertion Airway Device: Single-Lumen Double-Cuffed Size: 2 Child patient's: 12-25 kg
221 | RT13-37-001 | Emergency Cricothyroidotomy Kit | Emergency Cricothyroidotomy Kit – with Varess needle | Emergency Cricothyrotomy Kit – with Varess needle Purpose: Used for emergency airway access when conventional endotracheal intubation or ventilation can’t be performed Consists of a cuffed, 6mm Cricothyroidotomy tube with a 15mm connector pre-assembled on a varess needle Varess needle to confirm entry into the trachea and indicate any subsequent contact with the posterior tracheal wall. To include Syringe, Scalpel , Nylon sutures, Catheter mount, HME filter, neck strap. Manufactured from medical grade metal and plastic. All components must be latex-free sterile and individually packed in peel pouch that is easy to open - For SINGLE use only Information brochure or the outer package of the item should clearly indicate latex free | Emergency Cricothyrotomy Kit – with Varess needle Consists of a radiopaque, cuffed, 6mm Cricothyroidotomy tube with a 15mm connector pre-assembled on a varess needle
222 | RT13-37-002 | Emergency Cricothyroidotomy Kit | Emergency Cricothyroidotomy Kit | Emergency Cricothyrotomy Kit –Seldinger dilation technique via the cricothyroid membrane Purpose: Used for emergency airway access when conventional endotracheal intubation or ventilation can’t be performed Consists of a radiopaque, cuffed 5mm Cricothyroidotomy tube, 9cm long with 15mm connector Catheter to sit securely on tapered dilator and must be easy to advance over the dilator Short handled scalpel blade; Syringe with introducer needle; Guide-wire and curved dilator with a tapered end Manufactured from medical grade metal and plastic. All components must be latex-free sterile and individually packed in peel pouch that is easy to open - For SINGLE use only Information brochure or the outer package of the item should clearly indicate latex free | Emergency Cricothyrotomy Kit –Seldinger dilation technique via the cricothyroid membrane Consists of a radiopaque, cuffed 5mm Cricothyroidotomy tube, 9cm long with 15mm connector
223 | RT13-40-008 | Intubation Devices | INTUBATION STYLET | INTUBATION STYLET with metal core Purpose: To pre-form an endotracheal tube to a desired shape to aid intubation Consists of a malleable aluminium core coated and tightly bonded with high density polyurethane/ PVC / medical grade plastic/ polyethylene Must have a soft atraumatic tip Must have a U-bend proximally or other mechanism at proximal end to limit depth Latex-free clinically clean and individually packed in peel pouch that is easy to open - For SINGLE use only Information brochure or the outer package of the item should clearly indicate latex free | Intubation Stylet, Extra-Length Adult OD: 5.0mm Length: 365mm To fit ETT's > 7.5mm
224 | RT13-41-001 | ARTICULATED FLEXIBLE STYLET | ARTICULATED FLEXIBLE STYLET | ARTICULATED FLEXIBLE STYLET - REUSABLE Purpose: To increase the angle of an endotracheal tube during difficult intubations or the use of a D Blade Consists of a stylet with articulating arm, that manoeuvres the tip of an endotracheal tube in a single vertical plane to enable the proper placement of the tube in a difficult airway. Single thumb movement must increase the tube curvature Individually packed and surgical clean in a peel pouch Reusable. | Articulated flexible stylet. Size: 5mm
225 | RT13-41-002 | ARTICULATED FLEXIBLE STYLET | ARTICULATED FLEXIBLE STYLET | ARTICULATED FLEXIBLE STYLET - REUSABLE Purpose: To increase the angle of an endotracheal tube during difficult intubations or the use of a D Blade Consists of a stylet with articulating arm, that manoeuvres the tip of an endotracheal tube in a single vertical plane to enable the proper placement of the tube in a difficult airway. Single thumb movement must increase the tube curvature Individually packed and surgical clean in a peel pouch Reusable. | Articulated flexible stylet. Size: 5.5 - 6.0mm
226 | RT13-41-003 | ARTICULATED FLEXIBLE STYLET | ARTICULATED FLEXIBLE STYLET | ARTICULATED FLEXIBLE STYLET - REUSABLE Purpose: To increase the angle of an endotracheal tube during difficult intubations or the use of a D Blade Consists of a stylet with articulating arm, that manoeuvres the tip of an endotracheal tube in a single vertical plane to enable the proper placement of the tube in a difficult airway. Single thumb movement must increase the tube curvature Individually packed and surgical clean in a peel pouch Reusable. | Articulated flexible stylet. Size: 6.5 - 7.0mm
227 | RT13-41-004 | ARTICULATED FLEXIBLE STYLET | ARTICULATED FLEXIBLE STYLET | ARTICULATED FLEXIBLE STYLET - REUSABLE Purpose: To increase the angle of an endotracheal tube during difficult intubations or the use of a D Blade Consists of a stylet with articulating arm, that manoeuvres the tip of an endotracheal tube in a single vertical plane to enable the proper placement of the tube in a difficult airway. Single thumb movement must increase the tube curvature Individually packed and surgical clean in a peel pouch Reusable. | Articulated flexible stylet. Size: 7.5 - 8.0mm
228 | RT13-42-007 | Intubation Devices | SOLID INTUBATING BOUGIE | SOLID INTUBATING BOUGIE - CURVED-TIP- REUSABLE Purpose: To assist with endotracheal intubation where visualisation of the glottis is impossible or incomplete and for railroad exchange of endotracheal tubes. Consists of a solid tube manufactured from medical grade PVC that has a non-malleable core that is flexible but kink resistant and will aid intubation Must have an atraumatic "hockey stick" 35 - 40 degrees angled tip. The tip must clicks on tracheal rings as the bougie is advanced All components must be latex-free clinically clean and individually packed in peel pouch that is easy to open. Must be autoclavable – at least 40 times Information brochure or the outer package of the item should clearly indicate latex free | Solid Intubating bougie Curved tip Reusable Neonatal OD: 1.7mm Total Length: 500mm To fit ETT's 2.5 – 4.0mm
229 | RT13-42-008 | Intubation Devices | SOLID INTUBATING BOUGIE | SOLID INTUBATING BOUGIE - CURVED-TIP- REUSABLE Purpose: To assist with endotracheal intubation where visualisation of the glottis is impossible or incomplete and for railroad exchange of endotracheal tubes. Consists of a solid tube manufactured from medical grade PVC that has a non-malleable core that is flexible but kink resistant and will aid intubation Must have an atraumatic "hockey stick" 35 - 40 degrees angled tip. The tip must clicks on tracheal rings as the bougie is advanced All components must be latex-free clinically clean and individually packed in peel pouch that is easy to open. Must be autoclavable – at least 40 times Information brochure or the outer package of the item should clearly indicate latex free | Solid Intubating bougie Curved tip Reusable Paediatric OD: 3.3mm Total Length: 600 - 700mm To fit ETT's 4.0 – 6.0mm
230 | RT13-42-010 | AIRWAY DEVICES | AIRWAY EXCHANGE CATHETER | AIRWAY EXCHANGE CATHETER Purpose: For simple, atraumatic endotracheal tube exchange whilst providing oxygenation Consists of a catheter manufactured from radiopaque medical grade plastic Must have a hollow lumen for constant oxygenation of the patient, with distal side ports to ensure adequate airflow Must be supplied with connection and adaptor for attachment of a ventilatory device. Must have a blunt tip for atraumatic intubation, centimetre markings for accurate tube placement. Latex-free, sterile and individually packed in peel pouch that is easy to open - For SINGLE use only Information brochure or the outer package of the item should clearly indicate latex free | Airway Exchange Catheter. Infant 8Fr Length: ± 450mm To fit ETT's Size: 3.0mm
231 | RT13-42-011 | AIRWAY DEVICES | AIRWAY EXCHANGE CATHETER | AIRWAY EXCHANGE CATHETER Purpose: For simple, atraumatic endotracheal tube exchange whilst providing oxygenation Consists of a catheter manufactured from radiopaque medical grade plastic Must have a hollow lumen for constant oxygenation of the patient, with distal side ports to ensure adequate airflow Must be supplied with connection and adaptor for attachment of a ventilatory device. Must have a blunt tip for atraumatic intubation, centimetre markings for accurate tube placement. Latex-free, sterile and individually packed in peel pouch that is easy to open - For SINGLE use only Information brochure or the outer package of the item should clearly indicate latex free | Airway Exchange Catheter. Paediatric 11Fr Length: ± 830mm To fit ETT's Size: 4.0mm
232 | RT13-42-013 | AIRWAY DEVICES | AIRWAY EXCHANGE CATHETER | AIRWAY EXCHANGE CATHETER Purpose: For simple, atraumatic endotracheal tube exchange whilst providing oxygenation Consists of a catheter manufactured from radiopaque medical grade plastic Must have a hollow lumen for constant oxygenation of the patient, with distal side ports to ensure adequate airflow Must be supplied with connection and adaptor for attachment of a ventilatory device. Must have a blunt tip for atraumatic intubation, centimetre markings for accurate tube placement. Latex-free, sterile and individually packed in peel pouch that is easy to open - For SINGLE use only Information brochure or the outer package of the item should clearly indicate latex free | Airway Exchange Catheter. Adult 19Fr Length: ± 830mm To fit ETT's Size: 7.0mm
233 | RT13-42-014 | AIRWAY DEVICES | AIRWAY EXCHANGE CATHETER | AIRWAY EXCHANGE CATHETER FOR FIBRE-OPTIC GUIDED INTUBATION Purpose: For simple, atraumatic fibre-optic guided intubation while providing oxygenation. Consists of a large diameter hollow lumen catheter that fits over a Size 4.2 fibre-optic flexible bronchoscope Manufactured from radiopaque medical grade plastic, with distal side ports to ensure adequate airflow Must be supplied with connection and adaptor for attachment of a ventilatory device. Must have a blunt tip slightly tapered tip to fit snugly on the bronchoscope. Centimetre markings from 5cm in 1cm increments Preference will be given to items with a distal side hole for ventilation Latex-free, sterile and individually packed in peel pouch that is easy to open - For SINGLE use only Information brochure or the outer package of the item should clearly indicate latex free | Airway Exchange Catheter for Fibre-Optic Guided Intubation Size: 19Fr ID: 4.7mm OD: < 5mm Length: 560mm To fit ETT's 7.0mm
234 | RT13-42-015 | AIRWAY DEVICES | Airway Exchange Catheter | Airway Exchange Catheter Kit with Guidewire for staged extubation. Purpose: For staged or difficult extubation of patient's with cardiorespiratory dysfunction and potentially difficult reintubation Consists of an exchange catheter, extra stiff guidewire, Bronchoscopic port, 2 Adaptors 1 x luer lock, 15mm connector Catheter 14Fr ID: 2.3mm OD: < 4mm Catheter Length: 700 - 900mm Guidewire: Diameter: 0.1cm Length: 160cm The wire guide to fit into the working channel of a bronchoscope. The catheter must have distal side ports to allow for airflow. Must have an airway adaptor to allow for ventilation To fit an endotracheal tube of Size 5 or larger Latex-free medical grade metal and plastic, sterile and individually packed in peel pouch that is easy to open - For SINGLE use only Information brochure or the outer package of the item should clearly indicate latex free | Airway Exchange Catheter Kit with Guidewire for staged extubation.Catheter 14Fr ID: 2.3mm OD: < 4mm Catheter Length: 700 - 900mm
235 | RT13-43-001 | Intubation Devices | RETROGRADE INTUBATION KIT | RETROGRADE INTUBATION KIT 11Fr Length: 70cm for ETT's ID: 4mm or larger Purpose: Used to assist in the placement of an endotracheal tube using retrograde techniques during a difficult airway access procedure. Consists of an Introducer needle/ Catheter introducer needle:18G and Guide Wire: Size: 0.038mm Length 110cm Guidewire made from stainless steel with two positioning marks 20cm from each end; atraumatic tips; one end straight and the other a flexible J tip. Radiopaque catheter with distal side ports and a 15mm connector proximally for ventilatory devices. Luer slip syringe to be included. Seldinger technique access to the airway via the cricothyroid membrane. Retrograde (cephalad) advancement of wire towards the oral/ nasal cavities. Antegrade advancement of a hollow catheter with central & peripheral side ports for optimal ventilation. Latex-free, sterile and individually packed in peel pouch that is easy to open - For SINGLE use only Information brochure or the outer package of the item should clearly indicate latex free | RETROGRADE INTUBATION KIT 11Fr Length: 70cm for ETT's ID: 4mm or larger
236 | RT13-47-001 | AIRWAY DEVICES | OROPHARYNGEAL AIRWAY ATOMISER | OROPHARYNGEAL AIRWAY ATOMISER Purpose: An adult oropharyngeal airway with integral atomiser to topicalize airways while facilitating oxygenation. Must have two ports connected to the distal end of the oropharyngeal airway. One port for atomisation - must connect to a 2ml or 5ml syringe and allow for atomization with a particle size of 30-100microns The 2nd port must have a universal connector that connects to standard oxygen tubing. Manufactured from medical grade plastic, and must be latex-free and DEHP-free. Single use, individually packed and surgical clean. For airways requiring a fiberoptic scope Delivers topical anesthetic and oxygenation Information brochure or the outer package of the item should clearly indicate latex free and DEHP-free | OROPHARYNGEAL AIRWAY ATOMISER
237 | RT13-48-002 | ANAESTHETIC CONNECTOR | ANAESTHETIC ELBOW (angled) CONNECTOR | ANAESTHETIC ELBOW (angled) CONNECTORS Consists of a connector made from latex-free, non-pyrogenic, hypoallergenic, DEHP free, transparent, non-conductive medical grade plastic. All components latex-free, clinically clean and individually packed in peel pouch that is easy to open - For SINGLE use only To comply with ISO 5367 Information brochure or the outer package of the item should clearly indicate latex free | Anaesthetic circuit elbow (angled) connector 22mm male/ 15mm female (Patient end) 22mm female (Circuit end)
238 | RT13-48-004 | ANAESTHETIC CONNECTOR | ANAESTHETIC ELBOW (angled) CONNECTOR | ANAESTHETIC ELBOW (angled) CONNECTORS Consists of a connector made from latex-free, non-pyrogenic, hypoallergenic, DEHP free, transparent, non-conductive medical grade plastic. All components latex-free, clinically clean and individually packed in peel pouch that is easy to open - For SINGLE use only To comply with ISO 5367 Information brochure or the outer package of the item should clearly indicate latex free | Anaesthetic circuit elbow (angled) connector, with toggled, capped luer lock sample port for Capnography: 22mm male/ 15mm female (Patient end) 22mm female (Circuit end)
239 | RT13-48-013 | ANAESTHETIC CONNECTOR | ANAESTHETIC T-PIECE CONNECTORS | ANAESTHETIC T-PIECE CONNECTORS Consists of a connector made from latex-free, non-pyrogenic, hypoallergenic, DEHP free, transparent, non-conductive medical grade plastic. All components latex-free, clinically clean and individually packed in peel pouch that is easy to open - For SINGLE use only To comply with ISO 5367 Information brochure or the outer package of the item should clearly indicate latex free | Anaesthetic T-Piece connectors: T - opening (90°) 15mm male; Straight opening: 22mm Male to 15mm Female
240 | RT13-48-014 | ANAESTHETIC CONNECTOR | ANAESTHETIC T-PIECE CONNECTORS | ANAESTHETIC T-PIECE CONNECTORS Consists of a connector made from latex-free, non-pyrogenic, hypoallergenic, DEHP free, transparent, non-conductive medical grade plastic. All components latex-free, clinically clean and individually packed in peel pouch that is easy to open - For SINGLE use only To comply with ISO 5367 Information brochure or the outer package of the item should clearly indicate latex free | Anaesthetic T-Piece connectors: T - opening (90°) 22mm Male/15 Female; Straight opening: 22mm Male to 22mm Male with toggled 7.6mm fibreoptic/suction port on straight part
241 | RT13-48-015 | ANAESTHETIC CONNECTOR | ANAESTHETIC T-PIECE CONNECTORS | ANAESTHETIC T-PIECE CONNECTORS Consists of a connector made from latex-free, non-pyrogenic, hypoallergenic, DEHP free, transparent, non-conductive medical grade plastic. All components latex-free, clinically clean and individually packed in peel pouch that is easy to open - For SINGLE use only To comply with ISO 5367 Information brochure or the outer package of the item should clearly indicate latex free | Anaesthetic T-Piece connectors: T - opening (90°) 22mm Male/15 Female; Straight opening: 22mm Male to 22mm Male with toggled 7.6mm fibreoptic/suction port on straight part
242 | RT13-49-004 | ANAESTHETIC CONNECTOR | CONNECTORS | CONNECTORS and OXYGEN DELIVERY ACCESSORIES Consists of a connector made from latex-free, non-pyrogenic, hypoallergenic, DEHP free, transparent, non-conductive medical grade plastic. All components latex-free, clinically clean and individually packed in peel pouch that is easy to open - For SINGLE use only To comply with ISO 5367 Information brochure or the outer package of the item should clearly indicate latex free | Scavenging connector straight connector 30mm male to 30mm male
243 | RT13-49-005 | ANAESTHETIC CONNECTOR | CONNECTORS | CONNECTORS and OXYGEN DELIVERY ACCESSORIES Consists of a connector made from latex-free, non-pyrogenic, hypoallergenic, DEHP free, transparent, non-conductive medical grade plastic. All components latex-free, clinically clean and individually packed in peel pouch that is easy to open - For SINGLE use only To comply with ISO 5367 Information brochure or the outer package of the item should clearly indicate latex free | Scavenging connector straight connector 30mm male to 30mm female
244 | RT13-49-006 | AIRWAY DEVICES | AEROSOL STRAIGHT CONNECTOR | AEROSOL STRAIGHT CONNECTOR with toggled port to fit nebulising canister: 22mm Male to 22mm Female Consists of a connector made from latex-free, non-pyrogenic, hypoallergenic, DEHP free, transparent, non-conductive medical grade plastic. MDI must be directed towards the patient All components latex-free, clinically clean and individually packed in peel pouch that is easy to open - For SINGLE use only To comply with ISO 5367 Information brochure or the outer package of the item should clearly indicate latex free and DEHP free | AEROSOL STRAIGHT CONNECTOR with toggled port to fit nebulising canister: 22mm Male to 22mm Female
245 | RT13-50-002 | AIRWAY DEVICES | NIPPLE NUT-AND-STEM “CHRISTMAS TREE” OXYGEN ADAPTER | NIPPLE NUT-AND-STEM “CHRISTMAS TREE” OXYGEN ADAPTER Purpose: Connector between standard flow meter DISS connection and oxygen tubing Wing-nut to fit standard DISS connection on oxygen flow meter. Thread to be secure without leaking. Tapered stem to allow easy connection of standard bore oxygen tubing, with deep grooved barbs on the stem to prevent disconnection of tubing Manufactured from latex free medical grade plastic/ medical grade metal To comply with ISO 5367 or equivalent Clinically clean packaging Information brochure or the outer package of the item should clearly indicate latex free | Nipple Nut-and-stem “Christmas tree” oxygen adapter - Metal
246 | RT13-51-001 | CATHETER MOUNT CONNECTOR | CATHETER MOUNT - STRAIGHT CONNECTOR | CATHETER MOUNT - STRAIGHT CONNECTOR Purpose: Connector to join an endotracheal or tracheostomy tube to a breathing circuit Consists of flexible, low-flow smooth-bore tubing with an external coil to prevent kinking.. All components latex-free, sterile and individually packed in peel pouch that is easy to open - For SINGLE use only Information brochure or the outer package of the item should clearly indicate latex free | Catheter mount with Straight Connector- Neonatal Length ± 90mm Connectors: 22mm Male/ 15mm Female (patient end); 15mm Male (machine end) Volume: < 15ml
247 | RT13-51-002 | CATHETER MOUNT CONNECTOR | CATHETER MOUNT - STRAIGHT CONNECTOR | CATHETER MOUNT - STRAIGHT CONNECTOR Purpose: Connector to join an endotracheal or tracheostomy tube to a breathing circuit Consists of flexible, low-flow smooth-bore tubing with an external coil to prevent kinking.. All components latex-free, sterile and individually packed in peel pouch that is easy to open - For SINGLE use only Information brochure or the outer package of the item should clearly indicate latex free | Catheter mount (15mm) with Straight Connector- Paediatric Length ± 90mm Connectors: 22mm Male/ 15mm Female (patient end); 15mm Male (machine end) Volume: < 20ml
248 | RT13-51-004 | CATHETER MOUNT CONNECTOR | CATHETER MOUNT WITH ELBOW CONNECTOR | CATHETER MOUNT WITH ELBOW CONNECTOR, LUER LOCK SAMPLE PORT AND TOGGLED CAP Purpose: Connector to join an endotracheal or tracheostomy tube to a breathing circuit with sample port for Capnography Consists of flexible, low-flow smooth-bore tubing with an external coil to prevent kinks. Elbow connector with a Luer lock sample port and toggled cap All components latex-free, sterile and individually packed in peel pouch that is easy to open - For SINGLE use only Information brochure or the outer package of the item should clearly indicate latex free | Catheter mount with Elbow connector, C02 sampling port and toggled cap - Neonatal Volume: ± 10-16ml Length ± 90mm Connector: 22mm Male/ 15mm Female (patient end); 15mm Male (machine end)
249 | RT13-51-005 | CATHETER MOUNT CONNECTOR | CATHETER MOUNT WITH ELBOW CONNECTOR | CATHETER MOUNT WITH ELBOW CONNECTOR, LUER LOCK SAMPLE PORT AND TOGGLED CAP Purpose: Connector to join an endotracheal or tracheostomy tube to a breathing circuit with sample port for Capnography Consists of flexible, low-flow smooth-bore tubing with an external coil to prevent kinks. Elbow connector with a Luer lock sample port and toggled cap All components latex-free, sterile and individually packed in peel pouch that is easy to open - For SINGLE use only Information brochure or the outer package of the item should clearly indicate latex free | Catheter mount with Elbow connector, C02 sampling port and toggled cap - Paediatric Volume: ± 15-30ml Length ± 90mm Connectors: 22mm Male/ 15mm Female (patient end); 22mm Female (machine end)
250 | RT13-51-006 | CATHETER MOUNT CONNECTOR | CATHETER MOUNT WITH ELBOW CONNECTOR | CATHETER MOUNT WITH ELBOW CONNECTOR, LUER LOCK SAMPLE PORT AND TOGGLED CAP Purpose: Connector to join an endotracheal or tracheostomy tube to a breathing circuit with sample port for Capnography Consists of flexible, low-flow smooth-bore tubing with an external coil to prevent kinks. Elbow connector with a Luer lock sample port and toggled cap All components latex-free, sterile and individually packed in peel pouch that is easy to open - For SINGLE use only Information brochure or the outer package of the item should clearly indicate latex free | Catheter mount with Elbow connector, C02 sampling port and toggled cap - Adult Length ± 120-150mm Connectors: 22mm Male/ 15mm Female (patient end); 22mm Female (machine end)
251 | RT13-51-007 | CATHETER MOUNT CONNECTOR | CATHETER MOUNT WITH DOUBLE SWIVEL ELBOW CONNECTOR | CATHETER MOUNT WITH DOUBLE SWIVEL ELBOW CONNECTOR, BRONCHOSCOPY/ASPIRATION PORT AND TOGGLED CAP Purpose: Connector to join an endotracheal or tracheostomy tube to a breathing circuit with a bronchoscopy/aspiration port. Consists of flexible, low-flow smooth-bore tubing with an external coil to prevent kinks. Double swivel Elbow connector with a Bronchoscopy/Aspiration Port. Must have a diaphragm which hugs fibre-optic bronchoscope. All components latex-free, sterile and individually packed in peel pouch that is easy to open - For SINGLE use only Information brochure or the outer package of the item should clearly indicate latex free | Catheter mount with DOUBLE swivel elbow connector, Bronchoscopy/Aspiration Port and toggled cap - Neonatal Volume: ± 10-16ml Length ± 90mm Connectors: 22mm Male/ 15mm Female (patient end); 15mm Male (machine end)
252 | RT13-51-009 | CATHETER MOUNT CONNECTOR | CATHETER MOUNT WITH DOUBLE SWIVEL ELBOW CONNECTOR | CATHETER MOUNT WITH DOUBLE SWIVEL ELBOW CONNECTOR, BRONCHOSCOPY/ASPIRATION PORT AND TOGGLED CAP Purpose: Connector to join an endotracheal or tracheostomy tube to a breathing circuit with a bronchoscopy/aspiration port. Consists of flexible, low-flow smooth-bore tubing with an external coil to prevent kinks. Double swivel Elbow connector with a Bronchoscopy/Aspiration Port. Must have a diaphragm which hugs fibre-optic bronchoscope. All components latex-free, sterile and individually packed in peel pouch that is easy to open - For SINGLE use only Information brochure or the outer package of the item should clearly indicate latex free | Catheter mount with DOUBLE swivel elbow connector, Bronchoscopy/Aspiration Port and toggled cap - Adult Length: 120-150mm Connectors: 22mm Male/ 15mm Female to patient and 22mm Female to machine
253 | RT13-51-010 | CATHETER MOUNT CONNECTOR | CATHETER MOUNT- EXTENDIBLE | CATHETER MOUNT - Extendible with Elbow Connector, Luer Lock Sample Port and Toggled Cap Purpose: Connector to join an endotracheal or tracheostomy tube to a breathing circuit with sample port for Capnography. Consists of flexible, extendible tubing. The internal diameter must remain constant in every position. Must have an Elbow connector with a Luer lock sample port and toggled cap All components latex-free, sterile and individually packed in peel pouch that is easy to open - For SINGLE use only Information brochure or the outer package of the item should clearly indicate latex free | Catheter mount: Extendible with elbow connector, Luer Lock Sample Port and Toggled Cap - Premature Volume: ± 10-16ml Length ± 45mm - 100mm Connectors: 8.5mm Female (patient end); 15mm Male (machine end)
254 | RT13-51-011 | CATHETER MOUNT CONNECTOR | CATHETER MOUNT- EXTENDIBLE | CATHETER MOUNT - Extendible with Elbow Connector, Luer Lock Sample Port and Toggled Cap Purpose: Connector to join an endotracheal or tracheostomy tube to a breathing circuit with sample port for Capnography. Consists of flexible, extendible tubing. The internal diameter must remain constant in every position. Must have an Elbow connector with a Luer lock sample port and toggled cap All components latex-free, sterile and individually packed in peel pouch that is easy to open - For SINGLE use only Information brochure or the outer package of the item should clearly indicate latex free | Catheter mount: Extendible with connector, Luer Lock Sample Port and Toggled Cap - Neonatal Volume: ± 10-16ml Length ± 45mm - 100mm Connectors: 15mm Female (patient end) ; 15mm Male (machine end)
255 | RT13-51-012 | CATHETER MOUNT CONNECTOR | CATHETER MOUNT- EXTENDIBLE | CATHETER MOUNT - Extendible with Elbow Connector, Luer Lock Sample Port and Toggled Cap Purpose: Connector to join an endotracheal or tracheostomy tube to a breathing circuit with sample port for Capnography. Consists of flexible, extendible tubing. The internal diameter must remain constant in every position. Must have an Elbow connector with a Luer lock sample port and toggled cap All components latex-free, sterile and individually packed in peel pouch that is easy to open - For SINGLE use only Information brochure or the outer package of the item should clearly indicate latex free | Catheter mount: Extendible connector, Luer Lock Sample Port and Toggled Cap - Paediatric Volume: ± 15-30ml Length ± 45mm - 100mm Connectors: 22mm Male/ 15mm Female (patient end) 22mm Female (machine end)
256 | RT13-51-013 | CATHETER MOUNT CONNECTOR | CATHETER MOUNT- EXTENDIBLE | CATHETER MOUNT - Extendible with DOUBLE swivel elbow connector, Bronchoscopy/Aspiration Port and toggled cap Purpose: Connector to join an endotracheal or tracheostomy tube to a breathing circuit with with a bronchoscopy/aspiration port. Consists of flexible, extendible tubing. The internal diameter must remain constant in every position. Double swivel Elbow connector with a Bronchoscopy/Aspiration Port. Must have a diaphragm which hugs fibre-optic bronchoscope All components latex-free, sterile and individually packed in peel pouch that is easy to open - For SINGLE use only Information brochure or the outer package of the item should clearly indicate latex free | Catheter mount: Extendible with DOUBLE swivel elbow connector, Bronchoscopy/Aspiration Port and toggled cap - Premature Volume: ± 10-16ml Length ± 45mm - 100mm Connectors: 8.5mm Female (patient end); 15mm Male (machine end)
257 | RT13-51-014 | CATHETER MOUNT CONNECTOR | CATHETER MOUNT- EXTENDIBLE | CATHETER MOUNT - Extendible with DOUBLE swivel elbow connector, Bronchoscopy/Aspiration Port and toggled cap Purpose: Connector to join an endotracheal or tracheostomy tube to a breathing circuit with with a bronchoscopy/aspiration port. Consists of flexible, extendible tubing. The internal diameter must remain constant in every position. Double swivel Elbow connector with a Bronchoscopy/Aspiration Port. Must have a diaphragm which hugs fibre-optic bronchoscope All components latex-free, sterile and individually packed in peel pouch that is easy to open - For SINGLE use only Information brochure or the outer package of the item should clearly indicate latex free | Catheter mount: Extendible with DOUBLE swivel elbow connector, Bronchoscopy/Aspiration Port and toggled cap - Neonatal Volume: ± 10-16ml Length ± 45mm - 100mm Connectors: 15mm Female (patient end) ; 15mm Male (machine end)
258 | RT13-51-015 | CATHETER MOUNT CONNECTOR | CATHETER MOUNT- EXTENDIBLE | CATHETER MOUNT - Extendible with DOUBLE swivel elbow connector, Bronchoscopy/Aspiration Port and toggled cap Purpose: Connector to join an endotracheal or tracheostomy tube to a breathing circuit with with a bronchoscopy/aspiration port. Consists of flexible, extendible tubing. The internal diameter must remain constant in every position. Double swivel Elbow connector with a Bronchoscopy/Aspiration Port. Must have a diaphragm which hugs fibre-optic bronchoscope All components latex-free, sterile and individually packed in peel pouch that is easy to open - For SINGLE use only Information brochure or the outer package of the item should clearly indicate latex free | Catheter mount: Extendible with DOUBLE swivel elbow connector, Bronchoscopy/Aspiration Port and toggled cap - Paediatric Volume: ± 15-30ml Length ± 45mm - 100mm Connectors: 22mm Male/ 15mm Female (patient end) 22mm Female (machine end)
259 | RT13-52-005 | CLOSED SUCTION CATHETER | CLOSED SUCTION CATHETER PAEDIATRIC 48HR | CLOSED SUCTION CATHETER - PAEDIATRIC ENDOTRACHEAL Y-TYPE 48HR Purpose: Closed system for respiratory suctioning through endotracheal tubes without disconnecting or interrupting ventilation. Consists of a closed suction catheter with an atraumatic tip. The catheter must have markings along the length to indicate depth of suctioning and a radiopaque line on the catheter tip. While connected but not in suction mode the catheter tip must not extend beyond the sheath or impede airflow. The suction catheter must be enclosed within a durable, flexible, plastic sheath for storage to avoid contamination. The transparent sheath must allow visualisation of markings on the suction catheter and must connect to the endotracheal tube via a clear, firmly-fitting, plastic connector or adaptor. The sheath and connector must allow for smooth advancement of catheter. The system must seal the airway and catheter from atmospheric contamination and allow uninterrupted positive pressure ventilation throughout the suction process. An effective seal around the catheter must prevent ventilation gas entering the lumen of the sheath. The system must protect the user from contamination by secretions and suctioned material. Proximal end of suction catheter must have clear chamber, so that suctioned contents can be visualised and must have lockable suction valve with toggled capped suction tubing connector. Paediatric Catheters must have a Y- connector (without locking mechanism) with an appropriate range of diameters for each size catheter. The connector must allow variations in patient positions without causing torque, twisting, or obstruction of the tracheal tube, and must not cause tissue or skin pressure. The system must allow for irrigation via a lateral irrigation port which must be a capped with a one-way valve that connects to a standard luer-slip syringe. The irrigation port must be situated at the patient end of the system for catheters and at the suction end of the catheter for catheters.made from medically safe PVC The suction end of the system must be fitted with capped, lockable thumb valve to control suctioning. It must have a tapered connection to fit a range of suction tubing. The system must be as light in weight as possible. The system is for single patient use for a specified 48 or 72 hours. Must have coloured day sticker labels to facilitate regular changing of the catheter system. Made from medically safe PVC. Sterile and individually packed in easy to open pouch with one transparent side; For SINGLE use only The system must conform to ISO 8836:2019. Information brochure or the outer package of the item should clearly indicate latex free | Closed Suction Catheter - Paediatric Endotracheal Y-Type 48Hr 10Fr OD: 3.3mm For ETT with ID: 5.0mm Y connectors: 4.5mm, 5mm, 5.5mm
260 | RT13-52-013 | CLOSED SUCTION CATHETER | CLOSED SUCTION CATHETER ADULT 48HR | CLOSED SUCTION CATHETER ADULT ENDOTRACHEAL - 48HRS Purpose: Closed system for respiratory suctioning through endotracheal tubes without disconnecting or interrupting ventilation. Consists of a closed suction catheter with an atraumatic tip. The catheter must have markings along the length to indicate depth of suctioning and a radiopaque line on the catheter tip. While connected but not in suction mode the catheter tip must not extend beyond the sheath or impede airflow. The suction catheter must be enclosed within a durable, flexible, plastic sheath for storage to avoid contamination. The transparent sheath must allow visualisation of markings on the suction catheter and must connect to the endotracheal tube via a clear, firmly-fitting, plastic connector or adaptor. The sheath and connector must allow for smooth advancement of catheter. The system must seal the airway and catheter from atmospheric contamination and allow uninterrupted positive pressure ventilation throughout the suction process. An effective seal around the catheter must prevent ventilation gas entering the lumen of the sheath. The system must protect the user from contamination by secretions and suctioned material. Proximal end of suction catheter must have clear chamber, so that suctioned contents can be visualised and must have lockable suction valve with capped suction tubing connector. The suction catheter must connect onto the tube with a DOUBLE swivel connector or a double-swivel Y-connector with standard 15 mm OD taper fittings. This connection must be secure and accidental disconnection must be minimal The whole device must be as lightweight as possible. Must include a light, straight, flexible, clear, concertina catheter-mount (15x22 mm connections) tube for connection to the breathing circuit The connector must allow variations in patient positions without causing torque, twisting, or obstruction of the tracheal tube, and must not cause tissue or skin pressure. The system must allow for irrigation via a lateral irrigation port which must be a capped with a one-way valve that connects to a standard luer-slip syringe. The irrigation port must be situated at the patient end of the system for catheters and at the suction end of the catheter for catheters.made from medically safe PVC The suction end of the system must be fitted with capped, lockable thumb valve to control suctioning. It must have a tapered connection to fit a range of suction tubing. The system is for single patient use for a specified 48 hours Must have coloured day sticker labels to facilitate regular changing of the catheter system. Made from medically safe PVC. Sterile and individually packed in easy to open pouch with one transparent side; For SINGLE use only The system must conform to ISO 8836 Information brochure or the outer package of the item should clearly indicate latex free | Closed Suction Catheter - Adult Endotracheal 48Hr 10Fr OD: 3.3mm Length range: 520 - 600mm
261 | RT13-52-014 | CLOSED SUCTION CATHETER | CLOSED SUCTION CATHETER ADULT 48HR | CLOSED SUCTION CATHETER ADULT ENDOTRACHEAL - 48HRS Purpose: Closed system for respiratory suctioning through endotracheal tubes without disconnecting or interrupting ventilation. Consists of a closed suction catheter with an atraumatic tip. The catheter must have markings along the length to indicate depth of suctioning and a radiopaque line on the catheter tip. While connected but not in suction mode the catheter tip must not extend beyond the sheath or impede airflow. The suction catheter must be enclosed within a durable, flexible, plastic sheath for storage to avoid contamination. The transparent sheath must allow visualisation of markings on the suction catheter and must connect to the endotracheal tube via a clear, firmly-fitting, plastic connector or adaptor. The sheath and connector must allow for smooth advancement of catheter. The system must seal the airway and catheter from atmospheric contamination and allow uninterrupted positive pressure ventilation throughout the suction process. An effective seal around the catheter must prevent ventilation gas entering the lumen of the sheath. The system must protect the user from contamination by secretions and suctioned material. Proximal end of suction catheter must have clear chamber, so that suctioned contents can be visualised and must have lockable suction valve with capped suction tubing connector. The suction catheter must connect onto the tube with a DOUBLE swivel connector or a double-swivel Y-connector with standard 15 mm OD taper fittings. This connection must be secure and accidental disconnection must be minimal The whole device must be as lightweight as possible. Must include a light, straight, flexible, clear, concertina catheter-mount (15x22 mm connections) tube for connection to the breathing circuit The connector must allow variations in patient positions without causing torque, twisting, or obstruction of the tracheal tube, and must not cause tissue or skin pressure. The system must allow for irrigation via a lateral irrigation port which must be a capped with a one-way valve that connects to a standard luer-slip syringe. The irrigation port must be situated at the patient end of the system for catheters and at the suction end of the catheter for catheters.made from medically safe PVC The suction end of the system must be fitted with capped, lockable thumb valve to control suctioning. It must have a tapered connection to fit a range of suction tubing. The system is for single patient use for a specified 48 hours Must have coloured day sticker labels to facilitate regular changing of the catheter system. Made from medically safe PVC. Sterile and individually packed in easy to open pouch with one transparent side; For SINGLE use only The system must conform to ISO 8836 Information brochure or the outer package of the item should clearly indicate latex free | Closed Suction Catheter - Adult Endotracheal 48Hr 12Fr OD: 4.0mm Length range: 520 - 600mm
262 | RT13-52-015 | CLOSED SUCTION CATHETER | CLOSED SUCTION CATHETER ADULT 48HR | CLOSED SUCTION CATHETER ADULT ENDOTRACHEAL - 48HRS Purpose: Closed system for respiratory suctioning through endotracheal tubes without disconnecting or interrupting ventilation. Consists of a closed suction catheter with an atraumatic tip. The catheter must have markings along the length to indicate depth of suctioning and a radiopaque line on the catheter tip. While connected but not in suction mode the catheter tip must not extend beyond the sheath or impede airflow. The suction catheter must be enclosed within a durable, flexible, plastic sheath for storage to avoid contamination. The transparent sheath must allow visualisation of markings on the suction catheter and must connect to the endotracheal tube via a clear, firmly-fitting, plastic connector or adaptor. The sheath and connector must allow for smooth advancement of catheter. The system must seal the airway and catheter from atmospheric contamination and allow uninterrupted positive pressure ventilation throughout the suction process. An effective seal around the catheter must prevent ventilation gas entering the lumen of the sheath. The system must protect the user from contamination by secretions and suctioned material. Proximal end of suction catheter must have clear chamber, so that suctioned contents can be visualised and must have lockable suction valve with capped suction tubing connector. The suction catheter must connect onto the tube with a DOUBLE swivel connector or a double-swivel Y-connector with standard 15 mm OD taper fittings. This connection must be secure and accidental disconnection must be minimal The whole device must be as lightweight as possible. Must include a light, straight, flexible, clear, concertina catheter-mount (15x22 mm connections) tube for connection to the breathing circuit The connector must allow variations in patient positions without causing torque, twisting, or obstruction of the tracheal tube, and must not cause tissue or skin pressure. The system must allow for irrigation via a lateral irrigation port which must be a capped with a one-way valve that connects to a standard luer-slip syringe. The irrigation port must be situated at the patient end of the system for catheters and at the suction end of the catheter for catheters.made from medically safe PVC The suction end of the system must be fitted with capped, lockable thumb valve to control suctioning. It must have a tapered connection to fit a range of suction tubing. The system is for single patient use for a specified 48 hours Must have coloured day sticker labels to facilitate regular changing of the catheter system. Made from medically safe PVC. Sterile and individually packed in easy to open pouch with one transparent side; For SINGLE use only The system must conform to ISO 8836 Information brochure or the outer package of the item should clearly indicate latex free | Closed Suction Catheter - Adult Endotracheal 48Hr 14Fr OD: 4.7mm Length range: 520 - 600mm
263 | RT13-52-016 | CLOSED SUCTION CATHETER | CLOSED SUCTION CATHETER ADULT 48HR | CLOSED SUCTION CATHETER ADULT ENDOTRACHEAL - 48HRS Purpose: Closed system for respiratory suctioning through endotracheal tubes without disconnecting or interrupting ventilation. Consists of a closed suction catheter with an atraumatic tip. The catheter must have markings along the length to indicate depth of suctioning and a radiopaque line on the catheter tip. While connected but not in suction mode the catheter tip must not extend beyond the sheath or impede airflow. The suction catheter must be enclosed within a durable, flexible, plastic sheath for storage to avoid contamination. The transparent sheath must allow visualisation of markings on the suction catheter and must connect to the endotracheal tube via a clear, firmly-fitting, plastic connector or adaptor. The sheath and connector must allow for smooth advancement of catheter. The system must seal the airway and catheter from atmospheric contamination and allow uninterrupted positive pressure ventilation throughout the suction process. An effective seal around the catheter must prevent ventilation gas entering the lumen of the sheath. The system must protect the user from contamination by secretions and suctioned material. Proximal end of suction catheter must have clear chamber, so that suctioned contents can be visualised and must have lockable suction valve with capped suction tubing connector. The suction catheter must connect onto the tube with a DOUBLE swivel connector or a double-swivel Y-connector with standard 15 mm OD taper fittings. This connection must be secure and accidental disconnection must be minimal The whole device must be as lightweight as possible. Must include a light, straight, flexible, clear, concertina catheter-mount (15x22 mm connections) tube for connection to the breathing circuit The connector must allow variations in patient positions without causing torque, twisting, or obstruction of the tracheal tube, and must not cause tissue or skin pressure. The system must allow for irrigation via a lateral irrigation port which must be a capped with a one-way valve that connects to a standard luer-slip syringe. The irrigation port must be situated at the patient end of the system for catheters and at the suction end of the catheter for catheters.made from medically safe PVC The suction end of the system must be fitted with capped, lockable thumb valve to control suctioning. It must have a tapered connection to fit a range of suction tubing. The system is for single patient use for a specified 48 hours Must have coloured day sticker labels to facilitate regular changing of the catheter system. Made from medically safe PVC. Sterile and individually packed in easy to open pouch with one transparent side; For SINGLE use only The system must conform to ISO 8836 Information brochure or the outer package of the item should clearly indicate latex free | Closed Suction Catheter - Adult Endotracheal 48Hr 16Fr OD: 5.3mm Length range: 520 - 600mm
264 | RT13-52-021 | CLOSED SUCTION CATHETER | CLOSED SUCTION CATHETER - PAEDIATRIC TRACHEOSTOMY | CLOSED SUCTION CATHETER - PAEDIATRIC TRACHEOSTOMY DOUBLE SWIVEL ELBOW CONNECTOR 48HR Purpose: Closed system for respiratory suctioning through tracheostomy tubes without disconnecting or interrupting ventilation. Consists of a closed suction catheter with an atraumatic tip. The catheter must have markings along the length to indicate depth of suctioning and a radiopaque line on the catheter tip. While connected but not in suction mode the catheter tip must not extend beyond the sheath or impede airflow. The suction catheter must be enclosed within a durable, flexible, plastic sheath for storage to avoid contamination. The transparent sheath must allow visualisation of markings on the suction catheter and must connect to the tracheal tube via a clear, firmly-fitting, plastic connector or adaptor. The sheath and connector must allow for smooth advancement of catheter. The system must seal the airway and catheter from atmospheric contamination and allow uninterrupted positive pressure ventilation throughout the suction process. An effective seal around the catheter must prevent ventilation gas entering the lumen of the sheath. The system must protect the user from contamination by secretions and suctioned material. Proximal end of suction catheter must have clear chamber, so that suctioned contents can be visualised and must have lockable suction valve with capped suction tubing connector. Paediatric Catheters must have a double swivel elbow connector with an appropriate range of diameters for each size catheter. The connector must allow variations in patient positions without causing torque, twisting, or obstruction of the tracheal tube, and must not cause tissue or skin pressure. The system must allow for irrigation via a lateral irrigation port which must be a capped with a one-way valve that connects to a standard luer-slip syringe. The irrigation port must be situated at the patient end of the system for catheters and at the suction end of the catheter for catheters.made from medically safe PVC The suction end of the system must be fitted with capped, lockable thumb valve to control suctioning. It must have a tapered connection to fit a range of suction tubing. The system must be as light in weight as possible. The system is for single patient use for a specified 48 hours Must have coloured day sticker labels to facilitate regular changing of the catheter system. Made from medically safe PVC. Sterile and individually packed in easy to open pouch with one transparent side; For SINGLE use only The system must conform to ISO 8836 Information brochure or the outer package of the item should clearly indicate latex free | Closed Suction Catheter - Paediatric Tracheostomy Y-Type 48Hr 6Fr OD: 2.0mm For ETT
265 | RT13-52-022 | CLOSED SUCTION CATHETER | CLOSED SUCTION CATHETER - PAEDIATRIC TRACHEOSTOMY | CLOSED SUCTION CATHETER - PAEDIATRIC TRACHEOSTOMY DOUBLE SWIVEL ELBOW CONNECTOR 48HR Purpose: Closed system for respiratory suctioning through tracheostomy tubes without disconnecting or interrupting ventilation. Consists of a closed suction catheter with an atraumatic tip. The catheter must have markings along the length to indicate depth of suctioning and a radiopaque line on the catheter tip. While connected but not in suction mode the catheter tip must not extend beyond the sheath or impede airflow. The suction catheter must be enclosed within a durable, flexible, plastic sheath for storage to avoid contamination. The transparent sheath must allow visualisation of markings on the suction catheter and must connect to the tracheal tube via a clear, firmly-fitting, plastic connector or adaptor. The sheath and connector must allow for smooth advancement of catheter. The system must seal the airway and catheter from atmospheric contamination and allow uninterrupted positive pressure ventilation throughout the suction process. An effective seal around the catheter must prevent ventilation gas entering the lumen of the sheath. The system must protect the user from contamination by secretions and suctioned material. Proximal end of suction catheter must have clear chamber, so that suctioned contents can be visualised and must have lockable suction valve with capped suction tubing connector. Paediatric Catheters must have a double swivel elbow connector with an appropriate range of diameters for each size catheter. The connector must allow variations in patient positions without causing torque, twisting, or obstruction of the tracheal tube, and must not cause tissue or skin pressure. The system must allow for irrigation via a lateral irrigation port which must be a capped with a one-way valve that connects to a standard luer-slip syringe. The irrigation port must be situated at the patient end of the system for catheters and at the suction end of the catheter for catheters.made from medically safe PVC The suction end of the system must be fitted with capped, lockable thumb valve to control suctioning. It must have a tapered connection to fit a range of suction tubing. The system must be as light in weight as possible. The system is for single patient use for a specified 48 hours Must have coloured day sticker labels to facilitate regular changing of the catheter system. Made from medically safe PVC. Sterile and individually packed in easy to open pouch with one transparent side; For SINGLE use only The system must conform to ISO 8836 Information brochure or the outer package of the item should clearly indicate latex free | Closed Suction Catheter - Paediatric Tracheostomy Y-Type 48Hr 7Fr OD: 2.3mm For ETT
266 | RT13-52-023 | CLOSED SUCTION CATHETER | CLOSED SUCTION CATHETER - PAEDIATRIC TRACHEOSTOMY | CLOSED SUCTION CATHETER - PAEDIATRIC TRACHEOSTOMY DOUBLE SWIVEL ELBOW CONNECTOR 48HR Purpose: Closed system for respiratory suctioning through tracheostomy tubes without disconnecting or interrupting ventilation. Consists of a closed suction catheter with an atraumatic tip. The catheter must have markings along the length to indicate depth of suctioning and a radiopaque line on the catheter tip. While connected but not in suction mode the catheter tip must not extend beyond the sheath or impede airflow. The suction catheter must be enclosed within a durable, flexible, plastic sheath for storage to avoid contamination. The transparent sheath must allow visualisation of markings on the suction catheter and must connect to the tracheal tube via a clear, firmly-fitting, plastic connector or adaptor. The sheath and connector must allow for smooth advancement of catheter. The system must seal the airway and catheter from atmospheric contamination and allow uninterrupted positive pressure ventilation throughout the suction process. An effective seal around the catheter must prevent ventilation gas entering the lumen of the sheath. The system must protect the user from contamination by secretions and suctioned material. Proximal end of suction catheter must have clear chamber, so that suctioned contents can be visualised and must have lockable suction valve with capped suction tubing connector. Paediatric Catheters must have a double swivel elbow connector with an appropriate range of diameters for each size catheter. The connector must allow variations in patient positions without causing torque, twisting, or obstruction of the tracheal tube, and must not cause tissue or skin pressure. The system must allow for irrigation via a lateral irrigation port which must be a capped with a one-way valve that connects to a standard luer-slip syringe. The irrigation port must be situated at the patient end of the system for catheters and at the suction end of the catheter for catheters.made from medically safe PVC The suction end of the system must be fitted with capped, lockable thumb valve to control suctioning. It must have a tapered connection to fit a range of suction tubing. The system must be as light in weight as possible. The system is for single patient use for a specified 48 hours Must have coloured day sticker labels to facilitate regular changing of the catheter system. Made from medically safe PVC. Sterile and individually packed in easy to open pouch with one transparent side; For SINGLE use only The system must conform to ISO 8836 Information brochure or the outer package of the item should clearly indicate latex free | Closed Suction Catheter - Paediatric Tracheostomy Y-Type 48Hr 8Fr Suction Catheter OD: 2.6mm For ETT with ID: 4mm Length: 30.5cm Y connectors: 3.5mm, 4mm, 4.5mm
267 | RT13-52-024 | CLOSED SUCTION CATHETER | CLOSED SUCTION CATHETER - PAEDIATRIC TRACHEOSTOMY | CLOSED SUCTION CATHETER - PAEDIATRIC TRACHEOSTOMY DOUBLE SWIVEL ELBOW CONNECTOR 48HR Purpose: Closed system for respiratory suctioning through tracheostomy tubes without disconnecting or interrupting ventilation. Consists of a closed suction catheter with an atraumatic tip. The catheter must have markings along the length to indicate depth of suctioning and a radiopaque line on the catheter tip. While connected but not in suction mode the catheter tip must not extend beyond the sheath or impede airflow. The suction catheter must be enclosed within a durable, flexible, plastic sheath for storage to avoid contamination. The transparent sheath must allow visualisation of markings on the suction catheter and must connect to the tracheal tube via a clear, firmly-fitting, plastic connector or adaptor. The sheath and connector must allow for smooth advancement of catheter. The system must seal the airway and catheter from atmospheric contamination and allow uninterrupted positive pressure ventilation throughout the suction process. An effective seal around the catheter must prevent ventilation gas entering the lumen of the sheath. The system must protect the user from contamination by secretions and suctioned material. Proximal end of suction catheter must have clear chamber, so that suctioned contents can be visualised and must have lockable suction valve with capped suction tubing connector. Paediatric Catheters must have a double swivel elbow connector with an appropriate range of diameters for each size catheter. The connector must allow variations in patient positions without causing torque, twisting, or obstruction of the tracheal tube, and must not cause tissue or skin pressure. The system must allow for irrigation via a lateral irrigation port which must be a capped with a one-way valve that connects to a standard luer-slip syringe. The irrigation port must be situated at the patient end of the system for catheters and at the suction end of the catheter for catheters.made from medically safe PVC The suction end of the system must be fitted with capped, lockable thumb valve to control suctioning. It must have a tapered connection to fit a range of suction tubing. The system must be as light in weight as possible. The system is for single patient use for a specified 48 hours Must have coloured day sticker labels to facilitate regular changing of the catheter system. Made from medically safe PVC. Sterile and individually packed in easy to open pouch with one transparent side; For SINGLE use only The system must conform to ISO 8836 Information brochure or the outer package of the item should clearly indicate latex free | Closed Suction Catheter - Paediatric Tracheostomy Y-Type 48Hr 10Fr OD: 3.3mm For ET
268 | RT13-52-025 | CLOSED SUCTION CATHETER | CLOSED SUCTION CATHETER - PAEDIATRIC TRACHEOSTOMY | CLOSED SUCTION CATHETER - PAEDIATRIC TRACHEOSTOMY DOUBLE SWIVEL ELBOW CONNECTOR 48HR Purpose: Closed system for respiratory suctioning through tracheostomy tubes without disconnecting or interrupting ventilation. Consists of a closed suction catheter with an atraumatic tip. The catheter must have markings along the length to indicate depth of suctioning and a radiopaque line on the catheter tip. While connected but not in suction mode the catheter tip must not extend beyond the sheath or impede airflow. The suction catheter must be enclosed within a durable, flexible, plastic sheath for storage to avoid contamination. The transparent sheath must allow visualisation of markings on the suction catheter and must connect to the tracheal tube via a clear, firmly-fitting, plastic connector or adaptor. The sheath and connector must allow for smooth advancement of catheter. The system must seal the airway and catheter from atmospheric contamination and allow uninterrupted positive pressure ventilation throughout the suction process. An effective seal around the catheter must prevent ventilation gas entering the lumen of the sheath. The system must protect the user from contamination by secretions and suctioned material. Proximal end of suction catheter must have clear chamber, so that suctioned contents can be visualised and must have lockable suction valve with capped suction tubing connector. Paediatric Catheters must have a double swivel elbow connector with an appropriate range of diameters for each size catheter. The connector must allow variations in patient positions without causing torque, twisting, or obstruction of the tracheal tube, and must not cause tissue or skin pressure. The system must allow for irrigation via a lateral irrigation port which must be a capped with a one-way valve that connects to a standard luer-slip syringe. The irrigation port must be situated at the patient end of the system for catheters and at the suction end of the catheter for catheters.made from medically safe PVC The suction end of the system must be fitted with capped, lockable thumb valve to control suctioning. It must have a tapered connection to fit a range of suction tubing. The system must be as light in weight as possible. The system is for single patient use for a specified 48 hours Must have coloured day sticker labels to facilitate regular changing of the catheter system. Made from medically safe PVC. Sterile and individually packed in easy to open pouch with one transparent side; For SINGLE use only The system must conform to ISO 8836 Information brochure or the outer package of the item should clearly indicate latex free | Closed Suction Catheter - Paediatric Tracheostomy Y-Type 48Hr 12Fr OD: 4.0mm For ETT with ID: 6.0mm Length: 40.5cm Y connectors: 5.5mm, 6mm, 6.5mm
269 | RT13-52-031 | CLOSED SUCTION CATHETER | CLOSED SUCTION CATHETER ADULT TRACHEOSTOMY | CLOSED SUCTION CATHETER ADULT TRACHEOSTOMY - 48HRS Purpose: Closed system for respiratory suctioning through tracheostomy tubes without disconnecting or interrupting ventilation. Consists of a closed suction catheter with an atraumatic tip. The catheter must have markings along the length to indicate depth of suctioning and a radiopaque line on the catheter tip. While connected but not in suction mode the catheter tip must not extend beyond the sheath or impede airflow. The suction catheter must be enclosed within a durable, flexible, plastic sheath for storage to avoID:contamination. The transparent sheath must allow visualisation of markings on the suction catheter and must connect to the tracheal tube via a clear, firmly-fitting, plastic connector or adaptor. The sheath and connector must allow for smooth advancement of catheter. The system must seal the airway and catheter from atmospheric contamination and allow uninterrupted positive pressure ventilation throughout the suction process. An effective seal around the catheter must prevent ventilation gas entering the lumen of the sheath. The system must protect the user from contamination by secretions and suctioned material. Proximal end of suction catheter must have clear chamber, so that suctioned contents can be visualised and must have lockable suction valve with capped suction tubing connector. The suction catheter must connect onto the tube with a DOUBLE swivel connector or a double-swivel Y-connector with standard 15 mm OD taper fittings. This connection must be secure and accidental disconnection must be minimal The whole device must be as lightweight as possible. Must include a light, straight, flexible, clear, concertina catheter-mount (15x22 mm connections) tube for connection to the breathing circuit The connector must allow variations in patient positions without causing torque, twisting, or obstruction of the tracheal tube, and must not cause tissue or skin pressure. The system must allow for irrigation via a lateral irrigation port which must be a capped with a one-way valve that connects to a standard luer-slip syringe. The irrigation port must be situated at the patient end of the system for catheters and at the suction end of the catheter for catheters.made from medically safe PVC The suction end of the system must be fitted with capped, lockable thumb valve to control suctioning. It must have a tapered connection to fit a range of suction tubing. The system is for single patient use for a specified 48 hours. Must have coloured day sticker labels to facilitate regular changing of the catheter system. Made from medically safe PVC. Sterile and individually packed in easy to open pouch with one transparent side; For SINGLE use only The system must conform to ISO 8836 Information brochure or the outer package of the item should clearly indicate latex free | Closed Suction Catheter - Adult Tracheostomy 48Hr 10Fr OD: 3.3mm Length range: 300 - 400mm
270 | RT13-52-032 | CLOSED SUCTION CATHETER | CLOSED SUCTION CATHETER ADULT TRACHEOSTOMY | CLOSED SUCTION CATHETER ADULT TRACHEOSTOMY - 48HRS Purpose: Closed system for respiratory suctioning through tracheostomy tubes without disconnecting or interrupting ventilation. Consists of a closed suction catheter with an atraumatic tip. The catheter must have markings along the length to indicate depth of suctioning and a radiopaque line on the catheter tip. While connected but not in suction mode the catheter tip must not extend beyond the sheath or impede airflow. The suction catheter must be enclosed within a durable, flexible, plastic sheath for storage to avoID:contamination. The transparent sheath must allow visualisation of markings on the suction catheter and must connect to the tracheal tube via a clear, firmly-fitting, plastic connector or adaptor. The sheath and connector must allow for smooth advancement of catheter. The system must seal the airway and catheter from atmospheric contamination and allow uninterrupted positive pressure ventilation throughout the suction process. An effective seal around the catheter must prevent ventilation gas entering the lumen of the sheath. The system must protect the user from contamination by secretions and suctioned material. Proximal end of suction catheter must have clear chamber, so that suctioned contents can be visualised and must have lockable suction valve with capped suction tubing connector. The suction catheter must connect onto the tube with a DOUBLE swivel connector or a double-swivel Y-connector with standard 15 mm OD taper fittings. This connection must be secure and accidental disconnection must be minimal The whole device must be as lightweight as possible. Must include a light, straight, flexible, clear, concertina catheter-mount (15x22 mm connections) tube for connection to the breathing circuit The connector must allow variations in patient positions without causing torque, twisting, or obstruction of the tracheal tube, and must not cause tissue or skin pressure. The system must allow for irrigation via a lateral irrigation port which must be a capped with a one-way valve that connects to a standard luer-slip syringe. The irrigation port must be situated at the patient end of the system for catheters and at the suction end of the catheter for catheters.made from medically safe PVC The suction end of the system must be fitted with capped, lockable thumb valve to control suctioning. It must have a tapered connection to fit a range of suction tubing. The system is for single patient use for a specified 48 hours. Must have coloured day sticker labels to facilitate regular changing of the catheter system. Made from medically safe PVC. Sterile and individually packed in easy to open pouch with one transparent side; For SINGLE use only The system must conform to ISO 8836 Information brochure or the outer package of the item should clearly indicate latex free | Closed Suction Catheter - Adult Tracheostomy 48Hr 12Fr OD: 4.0mm Length range: 300 - 400mm
271 | RT13-52-033 | CLOSED SUCTION CATHETER | CLOSED SUCTION CATHETER ADULT TRACHEOSTOMY | CLOSED SUCTION CATHETER ADULT TRACHEOSTOMY - 48HRS Purpose: Closed system for respiratory suctioning through tracheostomy tubes without disconnecting or interrupting ventilation. Consists of a closed suction catheter with an atraumatic tip. The catheter must have markings along the length to indicate depth of suctioning and a radiopaque line on the catheter tip. While connected but not in suction mode the catheter tip must not extend beyond the sheath or impede airflow. The suction catheter must be enclosed within a durable, flexible, plastic sheath for storage to avoID:contamination. The transparent sheath must allow visualisation of markings on the suction catheter and must connect to the tracheal tube via a clear, firmly-fitting, plastic connector or adaptor. The sheath and connector must allow for smooth advancement of catheter. The system must seal the airway and catheter from atmospheric contamination and allow uninterrupted positive pressure ventilation throughout the suction process. An effective seal around the catheter must prevent ventilation gas entering the lumen of the sheath. The system must protect the user from contamination by secretions and suctioned material. Proximal end of suction catheter must have clear chamber, so that suctioned contents can be visualised and must have lockable suction valve with capped suction tubing connector. The suction catheter must connect onto the tube with a DOUBLE swivel connector or a double-swivel Y-connector with standard 15 mm OD taper fittings. This connection must be secure and accidental disconnection must be minimal The whole device must be as lightweight as possible. Must include a light, straight, flexible, clear, concertina catheter-mount (15x22 mm connections) tube for connection to the breathing circuit The connector must allow variations in patient positions without causing torque, twisting, or obstruction of the tracheal tube, and must not cause tissue or skin pressure. The system must allow for irrigation via a lateral irrigation port which must be a capped with a one-way valve that connects to a standard luer-slip syringe. The irrigation port must be situated at the patient end of the system for catheters and at the suction end of the catheter for catheters.made from medically safe PVC The suction end of the system must be fitted with capped, lockable thumb valve to control suctioning. It must have a tapered connection to fit a range of suction tubing. The system is for single patient use for a specified 48 hours. Must have coloured day sticker labels to facilitate regular changing of the catheter system. Made from medically safe PVC. Sterile and individually packed in easy to open pouch with one transparent side; For SINGLE use only The system must conform to ISO 8836 Information brochure or the outer package of the item should clearly indicate latex free | Closed Suction Catheter - Adult Tracheostomy 48Hr 14Fr OD: 4.7mm Length range: 300 - 400mm
272 | RT13-52-034 | CLOSED SUCTION CATHETER | CLOSED SUCTION CATHETER ADULT TRACHEOSTOMY | CLOSED SUCTION CATHETER ADULT TRACHEOSTOMY - 48HRS Purpose: Closed system for respiratory suctioning through tracheostomy tubes without disconnecting or interrupting ventilation. Consists of a closed suction catheter with an atraumatic tip. The catheter must have markings along the length to indicate depth of suctioning and a radiopaque line on the catheter tip. While connected but not in suction mode the catheter tip must not extend beyond the sheath or impede airflow. The suction catheter must be enclosed within a durable, flexible, plastic sheath for storage to avoID:contamination. The transparent sheath must allow visualisation of markings on the suction catheter and must connect to the tracheal tube via a clear, firmly-fitting, plastic connector or adaptor. The sheath and connector must allow for smooth advancement of catheter. The system must seal the airway and catheter from atmospheric contamination and allow uninterrupted positive pressure ventilation throughout the suction process. An effective seal around the catheter must prevent ventilation gas entering the lumen of the sheath. The system must protect the user from contamination by secretions and suctioned material. Proximal end of suction catheter must have clear chamber, so that suctioned contents can be visualised and must have lockable suction valve with capped suction tubing connector. The suction catheter must connect onto the tube with a DOUBLE swivel connector or a double-swivel Y-connector with standard 15 mm OD taper fittings. This connection must be secure and accidental disconnection must be minimal The whole device must be as lightweight as possible. Must include a light, straight, flexible, clear, concertina catheter-mount (15x22 mm connections) tube for connection to the breathing circuit The connector must allow variations in patient positions without causing torque, twisting, or obstruction of the tracheal tube, and must not cause tissue or skin pressure. The system must allow for irrigation via a lateral irrigation port which must be a capped with a one-way valve that connects to a standard luer-slip syringe. The irrigation port must be situated at the patient end of the system for catheters and at the suction end of the catheter for catheters.made from medically safe PVC The suction end of the system must be fitted with capped, lockable thumb valve to control suctioning. It must have a tapered connection to fit a range of suction tubing. The system is for single patient use for a specified 48 hours. Must have coloured day sticker labels to facilitate regular changing of the catheter system. Made from medically safe PVC. Sterile and individually packed in easy to open pouch with one transparent side; For SINGLE use only The system must conform to ISO 8836 Information brochure or the outer package of the item should clearly indicate latex free | Closed Suction Catheter - Adult Tracheostomy 48Hr 16Fr OD: 5.3mm Length range: 300 - 400mm
273 | RT13-54-03 | AIRWAY DEVICES | HEAT AND MOISTURE EXCHANGER | HEAT AND MOISTURE EXCHANGER (HME) - NEONATAL HME REPLACEMENT CONNECTOR Purpose: To warm and moisten air breathed through a endotracheal tube during anaesthesia Consists of a bidirectional hygroscopic bacteriostatic paper filter, with low resistance throughout clinical use Must remain effective when wet. For 24 hourly use The HME to be used as a replacement for size 3.0 and/or 3.5 ETT circuit connectors. The connectors must fit standard breathing circuits and ventilator outlets. Must have a side port allowing for access for suctioning and C02 sampling. The patient end must be tapered to fit securely into the endotracheal tube. Must have a non-conductive smooth housing profile. Latex-free non-pyrogenic, hypoallergenic and DEHP-free transparent medical grade plastic Sterile and individually packed in peel pouch with view paper; For SINGLE use only Must comply with ISO 9360 Information brochure or the outer package of the item should clearly indicate latex free | HEAT AND MOISTURE EXCHANGER (HME) -NEONATAL HME REPLACEMENT CONNECTOR Size 3.0 for VT 10-50ml Dead space: < 3ml; Maximum filter weight: < 5g; Resistance: 1.8 cmH2O at 5l/min Gas flow. Humidification: ± 30mg H2O /litre for VT 20ml
274 | RT13-54-04 | AIRWAY DEVICES | HEAT AND MOISTURE EXCHANGER | HEAT AND MOISTURE EXCHANGER (HME) - NEONATAL HME REPLACEMENT CONNECTOR Purpose: To warm and moisten air breathed through a endotracheal tube during anaesthesia Consists of a bidirectional hygroscopic bacteriostatic paper filter, with low resistance throughout clinical use Must remain effective when wet. For 24 hourly use The HME to be used as a replacement for size 3.0 and/or 3.5 ETT circuit connectors. The connectors must fit standard breathing circuits and ventilator outlets. Must have a side port allowing for access for suctioning and C02 sampling. The patient end must be tapered to fit securely into the endotracheal tube. Must have a non-conductive smooth housing profile. Latex-free non-pyrogenic, hypoallergenic and DEHP-free transparent medical grade plastic Sterile and individually packed in peel pouch with view paper; For SINGLE use only Must comply with ISO 9360 Information brochure or the outer package of the item should clearly indicate latex free | HEAT AND MOISTURE EXCHANGER (HME) - NEONATAL HME REPLACEMENT CONNECTOR Size 3.5 for VT 10-50ml Dead space: < 3ml; Maximum filter weight: < 5g; Resistance: 1.1cmH2O at 5l/min Gas flow. Humidification: ± 30mg H2O /litre for VT 10ml
275 | RT13-54-05 | AIRWAY DEVICES | HEAT AND MOISTURE EXCHANGER | HEAT AND MOISTURE EXCHANGER (HME) -NEONATAL Straight Purpose: To warm and moisten air breathed through a tracheostomy. Bidirectional use filter must be hygroscopic with low resistance throughout clinical use. Foam/ Pleated paper acceptable Must remain effective when wet. For 24 hourly use Must have a non-conductive smooth housing profile. Latex-free non-pyrogenic, hypoallergenic and DEHP-free transparent medical grade plastic Where applicable connections to fit standard breathing circuits and ventilator outlets: Patient ends 15F/22M ISO connections and circuit end to be compatible with commonly used anaesthesia breathing circuits and ventilator outlets Sterile and individually packed in peel pouch with view paper; For SINGLE use only Must comply with ISO 9360 Information brochure or the outer package of the item should clearly indicate latex free | HEAT AND MOISTURE EXCHANGER (HME) -NEONATAL Straight For Tidal Volumes VT 10-50ml Dead space: < 3ml; Maximum filter weight: < 5g; Resistance: < 0.9cmH2O at 10l/min Gas flow. Humidification: ± 30mg H2O /litre.
276 | RT13-54-06 | AIRWAY DEVICES | HEAT AND MOISTURE EXCHANGER | HEAT AND MOISTURE EXCHANGER (HME) -NEONATAL Straight Purpose: To warm and moisten air breathed through a tracheostomy. Bidirectional use filter must be hygroscopic with low resistance throughout clinical use. Foam/ Pleated paper acceptable Must remain effective when wet. For 24 hourly use Must have a non-conductive smooth housing profile. Latex-free non-pyrogenic, hypoallergenic and DEHP-free transparent medical grade plastic Where applicable connections to fit standard breathing circuits and ventilator outlets: Patient ends 15F/22M ISO connections and circuit end to be compatible with commonly used anaesthesia breathing circuits and ventilator outlets Sterile and individually packed in peel pouch with view paper; For SINGLE use only Must comply with ISO 9360 Information brochure or the outer package of the item should clearly indicate latex free | Closed Suction Catheter - Adult Tracheostomy 48Hr 10Fr OD: 3.3mm Length range: 300 - 400mm
277 | RT13-54-07 | AIRWAY DEVICES | HEAT AND MOISTURE EXCHANGER INFANT | HEAT AND MOISTURE EXCHANGER INFANT Purpose: Heat and moisture exchange device to provide passive humidification to patient's airways - for direct connection to tracheostomies Bidirectional use filter must be hydrophobic with low resistance throughout clinical use and must remain effective when wet. For short term use - transport or short procedures Consists of an ergonomic, non-conductive smooth housing profile. Latex-free non-pyrogenic, hypoallergenic and DEHP-free transparent medical grade plastic Connections to fit standard breathing circuits and ventilator outlets: Patient ends 15F/22M ISO connections and circuit end to be compatible with commonly used anaesthesia breathing circuits and ventilator outlets Sterile and individually packed in peel pouch with view paper; For SINGLE use only Information brochure or the outer package of the item should clearly indicate latex free | Tracheostomy Heat and Moisture Exchanger (HME) - Infant straight with toggled Capnography port For Tidal Volumes: > 50ml Dead space: < 11ml; Maximum filter weight: < 12g; Resistance: < 0.6cmH2O at 5 l/min Gas flow. Humidification: > 30 mg H2O /litre.
278 | RT13-57-002 | BREATHING CIRCUITS | CORRUGATED BREATHING CIRCUITS WITH EXTRA LIMB AND BAG | CORRUGATED BREATHING CIRCUITS WITH EXTRA LIMB AND BAG - SINGLE USE Purpose: Used to deliver oxygen and Volatile Gases to patient's and to eliminate C02 during anaesthesia Transparent circuit with low resistance, non-extendable, corrugated tubing that is non-kinking, lightweight and crush resistant. The tubing must be pliable and must not retain a shape. Uninterrupted tubing (not joined). Shallow corrugations to minimise flow resistance and turbulence. Must have an extra limb with an anti-static reservoir bag that gives good feedback regarding patient lung compliance. The reservoir bag must not twist on itself and must have a mechanism to prevent this. Connectors Malleable 22mm female connectors bonded to the ends of inspiratory and expiratory limbs – must fit securely to the machine. Spare 22mm M/M connector. Y-piece at patient's end. Separate elbow-piece with a toggled cap on a luer lock port. Good quality to ensure no leakage or easy disconnections Latex-free, non-pyrogenic, hypoallergenic, non-toxic and non-conductive medical grade plastic No variation in Circuit limb length or Extra limb will be accepted Clinically clean and individually packed, For SINGLE use only To comply with ISO 5367 Information brochure or the outer package of the item should clearly indicate latex free | Corrugated breathing circuit with extra limb and breathing bag - Single Use - Neonatal 15mm Circuit: 1.6m; 15mm Extra limb: 1.8m Reservoir bag: 500ml
279 | RT13-57-004 | BREATHING CIRCUITS | CORRUGATED BREATHING CIRCUITS WITH EXTRA LIMB AND BAG | CORRUGATED BREATHING CIRCUITS WITH EXTRA LIMB AND BAG - SINGLE USE Purpose: Used to deliver oxygen and Volatile Gases to patient's and to eliminate C02 during anaesthesia Transparent circuit with low resistance, non-extendable, corrugated tubing that is non-kinking, lightweight and crush resistant. The tubing must be pliable and must not retain a shape. Uninterrupted tubing (not joined). Shallow corrugations to minimise flow resistance and turbulence. Must have an extra limb with an anti-static reservoir bag that gives good feedback regarding patient lung compliance. The reservoir bag must not twist on itself and must have a mechanism to prevent this. Connectors Malleable 22mm female connectors bonded to the ends of inspiratory and expiratory limbs – must fit securely to the machine. Spare 22mm M/M connector. Y-piece at patient's end. Separate elbow-piece with a toggled cap on a luer lock port. Good quality to ensure no leakage or easy disconnections Latex-free, non-pyrogenic, hypoallergenic, non-toxic and non-conductive medical grade plastic No variation in Circuit limb length or Extra limb will be accepted Clinically clean and individually packed, For SINGLE use only To comply with ISO 5367 Information brochure or the outer package of the item should clearly indicate latex free | Corrugated breathing circuit with extra limb and breathing bag - Single Use - Paediatric > 3 yrs 15mm Circuit: 1.6m; 15mm Extra limb: 1.8m Reservoir bag: 1l
280 | RT13-57-011 | BREATHING CIRCUITS | CORRUGATED BREATHING CIRCUITS WITH WATER TRAPS AND EXTRA LIMB AND BAG-SINGLE USE | SINGLE CORRUGATED BREATHING TUBING - SINGLE USE Purpose: Used as extension or extra limb in breathing circuits Transparent circuit with low resistance, non-extendable, corrugated tubing that is non-kinking, lightweight and crush resistant. The tubing must be pliable and must not retain a shape Shallow corrugations to minimise flow resistance and turbulence. Connectors Malleable 22mm female connectors bonded to the end of the tubing Spare connectors to be provided Latex-free, non-pyrogenic, hypoallergenic, non-toxic and non-conductive medical grade plastic No variation in Circuit limb length or Extra limb will be accepted Clinically clean and individually packed, For SINGLE use only To comply with ISO 5367 Information brochure or the outer package of the item should clearly indicate latex free | Corrugated breathing circuit with water traps and extra limb with breathing bag - Single Use - Paediatric > 3 yrs 15mm Circuit: 1.6m; 15mm Extra limb: 1.2m Reservoir bag: 1l
281 | RT13-57-012 | BREATHING CIRCUITS | CORRUGATED BREATHING CIRCUITS WITH WATER TRAPS AND EXTRA LIMB AND BAG-SINGLE USE | SINGLE CORRUGATED BREATHING TUBING - SINGLE USE Purpose: Used as extension or extra limb in breathing circuits Transparent circuit with low resistance, non-extendable, corrugated tubing that is non-kinking, lightweight and crush resistant. The tubing must be pliable and must not retain a shape Shallow corrugations to minimise flow resistance and turbulence. Connectors Malleable 22mm female connectors bonded to the end of the tubing Spare connectors to be provided Latex-free, non-pyrogenic, hypoallergenic, non-toxic and non-conductive medical grade plastic No variation in Circuit limb length or Extra limb will be accepted Clinically clean and individually packed, For SINGLE use only To comply with ISO 5367 Information brochure or the outer package of the item should clearly indicate latex free | Corrugated breathing circuit with water traps and extra limb with breathing bag - Single use Paediatric > 3 yrs 15mm Circuit: 1.6m; 15mm Extra limb: 1.8m Reservoir bag: 1l
282 | RT13-57-013 | BREATHING CIRCUITS | CORRUGATED BREATHING CIRCUITS WITH WATER TRAPS AND EXTRA LIMB AND BAG-SINGLE USE | SINGLE CORRUGATED BREATHING TUBING - SINGLE USE Purpose: Used as extension or extra limb in breathing circuits Transparent circuit with low resistance, non-extendable, corrugated tubing that is non-kinking, lightweight and crush resistant. The tubing must be pliable and must not retain a shape Shallow corrugations to minimise flow resistance and turbulence. Connectors Malleable 22mm female connectors bonded to the end of the tubing Spare connectors to be provided Latex-free, non-pyrogenic, hypoallergenic, non-toxic and non-conductive medical grade plastic No variation in Circuit limb length or Extra limb will be accepted Clinically clean and individually packed, For SINGLE use only To comply with ISO 5367 Information brochure or the outer package of the item should clearly indicate latex free | Corrugated breathing circuit with water traps and extra limb with breathing bag - Single Use Adult 22mm Circuit: 1.8m; 22mm Extra limb: 1.2m OD: 22mm Reservoir bag: 2l
283 | RT13-57-015 | BREATHING CIRCUITS | CORRUGATED BREATHING CIRCUITS WITH WATER TRAPS AND EXTRA LIMB AND BAG-SINGLE USE | SINGLE CORRUGATED BREATHING TUBING - SINGLE USE Purpose: Used as extension or extra limb in breathing circuits Transparent circuit with low resistance, non-extendable, corrugated tubing that is non-kinking, lightweight and crush resistant. The tubing must be pliable and must not retain a shape Shallow corrugations to minimise flow resistance and turbulence. Connectors Malleable 22mm female connectors bonded to the end of the tubing Spare connectors to be provided Latex-free, non-pyrogenic, hypoallergenic, non-toxic and non-conductive medical grade plastic No variation in Circuit limb length or Extra limb will be accepted Clinically clean and individually packed, For SINGLE use only To comply with ISO 5367 Information brochure or the outer package of the item should clearly indicate latex free | Heat and Moisture Exchanger (HME) - Infant straight For Tidal Volumes: > 50ml Dead space: < 11ml; Maximum filter weight: < 12g; Resistance: < 0.6cmH2O at 5 l/min Gas flow. Humidification: > 30 mg H2O /litre.
284 | RT13-57-016 | BREATHING CIRCUITS | CORRUGATED BREATHING CIRCUITS WITH WATER TRAPS AND EXTRA LIMB AND BAG-SINGLE USE | SINGLE CORRUGATED BREATHING TUBING - SINGLE USE Purpose: Used as extension or extra limb in breathing circuits Transparent circuit with low resistance, non-extendable, corrugated tubing that is non-kinking, lightweight and crush resistant. The tubing must be pliable and must not retain a shape Shallow corrugations to minimise flow resistance and turbulence. Connectors Malleable 22mm female connectors bonded to the end of the tubing Spare connectors to be provided Latex-free, non-pyrogenic, hypoallergenic, non-toxic and non-conductive medical grade plastic No variation in Circuit limb length or Extra limb will be accepted Clinically clean and individually packed, For SINGLE use only To comply with ISO 5367 Information brochure or the outer package of the item should clearly indicate latex free | Corrugated breathing circuit with water traps and extra limb with breathing bag - Single Use Adult Long 22mm Circuit: 2.5m; 22mm Extra limb: 1.2m OD: 22mm Reservoir bag: 2l
285 | RT13-57-017 | BREATHING CIRCUITS | CORRUGATED BREATHING CIRCUITS WITH WATER TRAPS AND EXTRA LIMB AND BAG-SINGLE USE | SINGLE CORRUGATED BREATHING TUBING - SINGLE USE Purpose: Used as extension or extra limb in breathing circuits Transparent circuit with low resistance, non-extendable, corrugated tubing that is non-kinking, lightweight and crush resistant. The tubing must be pliable and must not retain a shape Shallow corrugations to minimise flow resistance and turbulence. Connectors Malleable 22mm female connectors bonded to the end of the tubing Spare connectors to be provided Latex-free, non-pyrogenic, hypoallergenic, non-toxic and non-conductive medical grade plastic No variation in Circuit limb length or Extra limb will be accepted Clinically clean and individually packed, For SINGLE use only To comply with ISO 5367 Information brochure or the outer package of the item should clearly indicate latex free | Corrugated breathing circuit with water traps and extra limb with breathing bag - Single Use Adult Long 22mm Circuit: 2.5m; 22mm Extra limb: 1.5m OD: 22mm Reservoir bag: 2l
286 | RT13-57-018 | BREATHING CIRCUITS | CORRUGATED BREATHING CIRCUITS WITH WATER TRAPS AND EXTRA LIMB AND BAG-SINGLE USE | SINGLE CORRUGATED BREATHING TUBING - SINGLE USE Purpose: Used as extension or extra limb in breathing circuits Transparent circuit with low resistance, non-extendable, corrugated tubing that is non-kinking, lightweight and crush resistant. The tubing must be pliable and must not retain a shape Shallow corrugations to minimise flow resistance and turbulence. Connectors Malleable 22mm female connectors bonded to the end of the tubing Spare connectors to be provided Latex-free, non-pyrogenic, hypoallergenic, non-toxic and non-conductive medical grade plastic No variation in Circuit limb length or Extra limb will be accepted Clinically clean and individually packed, For SINGLE use only To comply with ISO 5367 Information brochure or the outer package of the item should clearly indicate latex free | Corrugated breathing circuit with water traps and extra limb with breathing bag - Single Use Adult Long 22mm Circuit: 2.5m; 22mm Extra limb: 1.8m OD: 22mm Reservoir bag: 2l
287 | RT13-58-001 | BREATHING CIRCUITS | SINGLE CORRUGATED BREATHING TUBING - SINGLE USE | SINGLE CORRUGATED BREATHING TUBING - SINGLE USE Purpose: Used as extension or extra limb in breathing circuits Transparent circuit with low resistance, non-extendable, corrugated tubing that is non-kinking, lightweight and crush resistant. The tubing must be pliable and must not retain a shape Shallow corrugations to minimise flow resistance and turbulence. Connectors Malleable 22mm female connectors bonded to the end of the tubing Spare connectors to be provided Latex-free, non-pyrogenic, hypoallergenic, non-toxic and non-conductive medical grade plastic No variation in Circuit limb length or Extra limb will be accepted Clinically clean and individually packed, For SINGLE use only To comply with ISO 5367 Information brochure or the outer package of the item should clearly indicate latex free | Single Corrugated Breathing Tubing - Single Use Paediatric 15mm Length: 1.2m Connector: 15mm M/F; 15mm M/M; 15mm F/F
288 | RT13-58-002 | BREATHING CIRCUITS | SINGLE CORRUGATED BREATHING TUBING - SINGLE USE | SINGLE CORRUGATED BREATHING TUBING - SINGLE USE Purpose: Used as extension or extra limb in breathing circuits Transparent circuit with low resistance, non-extendable, corrugated tubing that is non-kinking, lightweight and crush resistant. The tubing must be pliable and must not retain a shape Shallow corrugations to minimise flow resistance and turbulence. Connectors Malleable 22mm female connectors bonded to the end of the tubing Spare connectors to be provided Latex-free, non-pyrogenic, hypoallergenic, non-toxic and non-conductive medical grade plastic No variation in Circuit limb length or Extra limb will be accepted Clinically clean and individually packed, For SINGLE use only To comply with ISO 5367 Information brochure or the outer package of the item should clearly indicate latex free | Single Corrugated Breathing Tubing - Single Use Paediatric 15mm Length: 1.8m Connector: 15mm M/F; 15mm M/M; 15mm F/F
289 | RT13-58-003 | BREATHING CIRCUITS | SINGLE CORRUGATED BREATHING TUBING - SINGLE USE | SINGLE CORRUGATED BREATHING TUBING - SINGLE USE Purpose: Used as extension or extra limb in breathing circuits Transparent circuit with low resistance, non-extendable, corrugated tubing that is non-kinking, lightweight and crush resistant. The tubing must be pliable and must not retain a shape Shallow corrugations to minimise flow resistance and turbulence. Connectors Malleable 22mm female connectors bonded to the end of the tubing Spare connectors to be provided Latex-free, non-pyrogenic, hypoallergenic, non-toxic and non-conductive medical grade plastic No variation in Circuit limb length or Extra limb will be accepted Clinically clean and individually packed, For SINGLE use only To comply with ISO 5367 Information brochure or the outer package of the item should clearly indicate latex free | Single Corrugated Breathing Tubing - Single Use Adult 22mm Length: 1.2m Connector: 22mm M/F; 22mm M/M; 22mm F/F
290 | RT13-58-004 | BREATHING CIRCUITS | SINGLE CORRUGATED BREATHING TUBING - SINGLE USE | SINGLE CORRUGATED BREATHING TUBING - SINGLE USE Purpose: Used as extension or extra limb in breathing circuits Transparent circuit with low resistance, non-extendable, corrugated tubing that is non-kinking, lightweight and crush resistant. The tubing must be pliable and must not retain a shape Shallow corrugations to minimise flow resistance and turbulence. Connectors Malleable 22mm female connectors bonded to the end of the tubing Spare connectors to be provided Latex-free, non-pyrogenic, hypoallergenic, non-toxic and non-conductive medical grade plastic No variation in Circuit limb length or Extra limb will be accepted Clinically clean and individually packed, For SINGLE use only To comply with ISO 5367 Information brochure or the outer package of the item should clearly indicate latex free | Single Corrugated Breathing Tubing - Single Use Adult 22mm Length: 1.8m Connector: 22mm M/F; 22mm M/M; 22mm F/F
291 | RT13-59-001 | BREATHING CIRCUITS | MAPLESON F PAEDIATRIC BREATHING CIRCUIT | MAPLESON F PAEDIATRIC BREATHING CIRCUIT - (Modified Ayre’s T piece) for children < 20-25 kg Purpose: Used to deliver oxygen and volatile gases to patient's and to eliminate C02 during anaesthesia Transparent circuit with low resistance, non-extendable, corrugated tubing that is non-kinking, lightweight and crush resistant. The tubing must be pliable and must not retain a shape with shallow corrugations to minimise flow resistance and turbulence. Tubing must be connected to a straight T-Piece with inspiratory port, expiratory port and patient connection port 22M/15F. Inspiratory tubing: 10mm ± 1.8 - 2.1m Bonded 15mm male connector at the end of the inspiratory limb– must fit securely to the machine. . Expiratory tubing 15mm diameter 0.5m with a 0.5l open-ended reservoir bag to allow gas to escape via the tail of the bag The reservoir bag must give good feedback regarding patient lung compliance and must not twist on itself and must have a mechanism to prevent this. Connectors: Elbow-piece with toggled cap on a luer lock port, capped connector at patient’s end. Good quality to ensure no leakage or easy disconnections Manufactured from transparent non-conductive medical grade plastic that is latex-free non-pyrogenic, hypoallergenic and non-toxic. Information brochure or the outer package of the item should clearly indicate latex free | Corrugated breathing circuit with water traps and extra limb with breathing bag - Single Use Adult Long 22mm Circuit: 2.5m; 22mm Extra limb: 1.2m OD: 22mm Reservoir bag: 2l
292 | RT13-60-001 | BREATHING CIRCUITS | SMOOTH BORE BREATHING CIRCUITS WITH EXTRA LIMB AND BREATHING BAG - SINGLE USE | SMOOTH BORE BREATHING CIRCUITS WITH EXTRA LIMB AND BREATHING BAG - SINGLE USE Purpose: Used to deliver oxygen and Volatile Gases to patient's and to eliminate C02 during anaesthesia Transparent circuit with low resistance, non-extendable, smooth-bore tubing that is non-kinking, lightweight and crush resistant. The tubing must be pliable and must not retain a shape Spiral outer casing that supports the inner lumen to maintain tube integrity. Must have an extra limb with an anti-static reservoir bag that gives good feedback regarding patient lung compliance. The reservoir bag must not twist on itself and must have a mechanism to prevent this. Connectors Malleable 22mm female connectors bonded to the ends of inspiratory and expiratory limbs – must fit securely to the machine. Spare 22mm M/M connector. Y-piece at patient's end. Separate elbow-piece with capped luer lock port. Good quality to ensure no leakage or easy disconnections Latex-free, non-pyrogenic, hypoallergenic, non-toxic and non-conductive medical grade plastic No variation in Circuit limb length or Extra limb will be accepted Clinically clean and individually packed, For SINGLE use only To comply with ISO 5367 Information brochure or the outer package of the item should clearly indicate latex free | Smooth bore breathing circuit with extra limb and breathing bag - Single Use Paediatric > 3yrs 15mm Circuit: 1.6m; Extra limb: 1.2m Reservoir bag: 1l
293 | RT13-60-002 | BREATHING CIRCUITS | SMOOTH BORE BREATHING CIRCUITS WITH EXTRA LIMB AND BREATHING BAG - SINGLE USE | SMOOTH BORE BREATHING CIRCUITS WITH EXTRA LIMB AND BREATHING BAG - SINGLE USE Purpose: Used to deliver oxygen and Volatile Gases to patient's and to eliminate C02 during anaesthesia Transparent circuit with low resistance, non-extendable, smooth-bore tubing that is non-kinking, lightweight and crush resistant. The tubing must be pliable and must not retain a shape Spiral outer casing that supports the inner lumen to maintain tube integrity. Must have an extra limb with an anti-static reservoir bag that gives good feedback regarding patient lung compliance. The reservoir bag must not twist on itself and must have a mechanism to prevent this. Connectors Malleable 22mm female connectors bonded to the ends of inspiratory and expiratory limbs – must fit securely to the machine. Spare 22mm M/M connector. Y-piece at patient's end. Separate elbow-piece with capped luer lock port. Good quality to ensure no leakage or easy disconnections Latex-free, non-pyrogenic, hypoallergenic, non-toxic and non-conductive medical grade plastic No variation in Circuit limb length or Extra limb will be accepted Clinically clean and individually packed, For SINGLE use only To comply with ISO 5367 Information brochure or the outer package of the item should clearly indicate latex free | Smooth bore breathing circuit with extra limb and breathing bag - Single Use Paediatric > 3yrs 15mm Circuit: 1.6m; Extra limb: 1.8m Reservoir bag: 1l
294 | RT13-60-003 | BREATHING CIRCUITS | SMOOTH BORE BREATHING CIRCUITS WITH EXTRA LIMB AND BREATHING BAG - SINGLE USE | SMOOTH BORE BREATHING CIRCUITS WITH EXTRA LIMB AND BREATHING BAG - SINGLE USE Purpose: Used to deliver oxygen and Volatile Gases to patient's and to eliminate C02 during anaesthesia Transparent circuit with low resistance, non-extendable, smooth-bore tubing that is non-kinking, lightweight and crush resistant. The tubing must be pliable and must not retain a shape Spiral outer casing that supports the inner lumen to maintain tube integrity. Must have an extra limb with an anti-static reservoir bag that gives good feedback regarding patient lung compliance. The reservoir bag must not twist on itself and must have a mechanism to prevent this. Connectors Malleable 22mm female connectors bonded to the ends of inspiratory and expiratory limbs – must fit securely to the machine. Spare 22mm M/M connector. Y-piece at patient's end. Separate elbow-piece with capped luer lock port. Good quality to ensure no leakage or easy disconnections Latex-free, non-pyrogenic, hypoallergenic, non-toxic and non-conductive medical grade plastic No variation in Circuit limb length or Extra limb will be accepted Clinically clean and individually packed, For SINGLE use only To comply with ISO 5367 Information brochure or the outer package of the item should clearly indicate latex free | Smooth bore breathing circuit with extra limb and breathing bag - Single Use Adult 22mm Circuit: 1.8m; Extra limb: 1.2m OD: Reservoir bag: 2l
295 | RT13-60-004 | BREATHING CIRCUITS | SMOOTH BORE BREATHING CIRCUITS WITH EXTRA LIMB AND BREATHING BAG - SINGLE USE | SMOOTH BORE BREATHING CIRCUITS WITH EXTRA LIMB AND BREATHING BAG - SINGLE USE Purpose: Used to deliver oxygen and Volatile Gases to patient's and to eliminate C02 during anaesthesia Transparent circuit with low resistance, non-extendable, smooth-bore tubing that is non-kinking, lightweight and crush resistant. The tubing must be pliable and must not retain a shape Spiral outer casing that supports the inner lumen to maintain tube integrity. Must have an extra limb with an anti-static reservoir bag that gives good feedback regarding patient lung compliance. The reservoir bag must not twist on itself and must have a mechanism to prevent this. Connectors Malleable 22mm female connectors bonded to the ends of inspiratory and expiratory limbs – must fit securely to the machine. Spare 22mm M/M connector. Y-piece at patient's end. Separate elbow-piece with capped luer lock port. Good quality to ensure no leakage or easy disconnections Latex-free, non-pyrogenic, hypoallergenic, non-toxic and non-conductive medical grade plastic No variation in Circuit limb length or Extra limb will be accepted Clinically clean and individually packed, For SINGLE use only To comply with ISO 5367 Information brochure or the outer package of the item should clearly indicate latex free | Smooth bore breathing circuit with extra limb and breathing bag - Single Use Adult 22mm Circuit: 1.8m; Extra limb: 1.5m OD: Reservoir bag: 2l
296 | RT13-60-005 | BREATHING CIRCUITS | SMOOTH BORE BREATHING CIRCUITS WITH EXTRA LIMB AND BREATHING BAG - SINGLE USE | SMOOTH BORE BREATHING CIRCUITS WITH EXTRA LIMB AND BREATHING BAG - SINGLE USE Purpose: Used to deliver oxygen and Volatile Gases to patient's and to eliminate C02 during anaesthesia Transparent circuit with low resistance, non-extendable, smooth-bore tubing that is non-kinking, lightweight and crush resistant. The tubing must be pliable and must not retain a shape Spiral outer casing that supports the inner lumen to maintain tube integrity. Must have an extra limb with an anti-static reservoir bag that gives good feedback regarding patient lung compliance. The reservoir bag must not twist on itself and must have a mechanism to prevent this. Connectors Malleable 22mm female connectors bonded to the ends of inspiratory and expiratory limbs – must fit securely to the machine. Spare 22mm M/M connector. Y-piece at patient's end. Separate elbow-piece with capped luer lock port. Good quality to ensure no leakage or easy disconnections Latex-free, non-pyrogenic, hypoallergenic, non-toxic and non-conductive medical grade plastic No variation in Circuit limb length or Extra limb will be accepted Clinically clean and individually packed, For SINGLE use only To comply with ISO 5367 Information brochure or the outer package of the item should clearly indicate latex free | Smooth bore breathing circuit with extra limb and breathing bag - Single Use Adult 22mm Circuit: 1.8m; Extra limb: 1.8m OD: Reservoir bag: 2l
297 | RT13-60-006 | BREATHING CIRCUITS | SMOOTH BORE BREATHING CIRCUITS WITH EXTRA LIMB AND BREATHING BAG - SINGLE USE | SMOOTH BORE BREATHING CIRCUITS WITH EXTRA LIMB AND BREATHING BAG - SINGLE USE Purpose: Used to deliver oxygen and Volatile Gases to patient's and to eliminate C02 during anaesthesia Transparent circuit with low resistance, non-extendable, smooth-bore tubing that is non-kinking, lightweight and crush resistant. The tubing must be pliable and must not retain a shape Spiral outer casing that supports the inner lumen to maintain tube integrity. Must have an extra limb with an anti-static reservoir bag that gives good feedback regarding patient lung compliance. The reservoir bag must not twist on itself and must have a mechanism to prevent this. Connectors Malleable 22mm female connectors bonded to the ends of inspiratory and expiratory limbs – must fit securely to the machine. Spare 22mm M/M connector. Y-piece at patient's end. Separate elbow-piece with capped luer lock port. Good quality to ensure no leakage or easy disconnections Latex-free, non-pyrogenic, hypoallergenic, non-toxic and non-conductive medical grade plastic No variation in Circuit limb length or Extra limb will be accepted Clinically clean and individually packed, For SINGLE use only To comply with ISO 5367 Information brochure or the outer package of the item should clearly indicate latex free | Smooth bore breathing circuit with extra limb and breathing bag - Single Use Adult Long 22mm Circuit: 2.6m; Extra limb: 1.2m Reservoir bag: 2l
298 | RT13-60-007 | BREATHING CIRCUITS | SMOOTH BORE BREATHING CIRCUITS WITH EXTRA LIMB AND BREATHING BAG - SINGLE USE | SMOOTH BORE BREATHING CIRCUITS WITH EXTRA LIMB AND BREATHING BAG - SINGLE USE Purpose: Used to deliver oxygen and Volatile Gases to patient's and to eliminate C02 during anaesthesia Transparent circuit with low resistance, non-extendable, smooth-bore tubing that is non-kinking, lightweight and crush resistant. The tubing must be pliable and must not retain a shape Spiral outer casing that supports the inner lumen to maintain tube integrity. Must have an extra limb with an anti-static reservoir bag that gives good feedback regarding patient lung compliance. The reservoir bag must not twist on itself and must have a mechanism to prevent this. Connectors Malleable 22mm female connectors bonded to the ends of inspiratory and expiratory limbs – must fit securely to the machine. Spare 22mm M/M connector. Y-piece at patient's end. Separate elbow-piece with capped luer lock port. Good quality to ensure no leakage or easy disconnections Latex-free, non-pyrogenic, hypoallergenic, non-toxic and non-conductive medical grade plastic No variation in Circuit limb length or Extra limb will be accepted Clinically clean and individually packed, For SINGLE use only To comply with ISO 5367 Information brochure or the outer package of the item should clearly indicate latex free | Smooth bore breathing circuit with extra limb and breathing bag - Single Use Adult Long 22mm Circuit: 2.6m; Extra limb: 1.5m Reservoir bag: 2l
299 | RT13-60-008 | BREATHING CIRCUITS | SMOOTH BORE BREATHING CIRCUITS WITH EXTRA LIMB AND BREATHING BAG - SINGLE USE | SMOOTH BORE BREATHING CIRCUITS WITH EXTRA LIMB AND BREATHING BAG - SINGLE USE Purpose: Used to deliver oxygen and Volatile Gases to patient's and to eliminate C02 during anaesthesia Transparent circuit with low resistance, non-extendable, smooth-bore tubing that is non-kinking, lightweight and crush resistant. The tubing must be pliable and must not retain a shape Spiral outer casing that supports the inner lumen to maintain tube integrity. Must have an extra limb with an anti-static reservoir bag that gives good feedback regarding patient lung compliance. The reservoir bag must not twist on itself and must have a mechanism to prevent this. Connectors Malleable 22mm female connectors bonded to the ends of inspiratory and expiratory limbs – must fit securely to the machine. Spare 22mm M/M connector. Y-piece at patient's end. Separate elbow-piece with capped luer lock port. Good quality to ensure no leakage or easy disconnections Latex-free, non-pyrogenic, hypoallergenic, non-toxic and non-conductive medical grade plastic No variation in Circuit limb length or Extra limb will be accepted Clinically clean and individually packed, For SINGLE use only To comply with ISO 5367 Information brochure or the outer package of the item should clearly indicate latex free | Smooth bore breathing circuit with extra limb and breathing bag - Single Use Adult Long 22mm Circuit: 2.6m; Extra limb: 1.8m Reservoir bag: 2l
300 | RT13-60-009 | BREATHING CIRCUITS | SMOOTH BORE BREATHING CIRCUITS WITH WATER TRAPS AND EXTRA LIMB AND BAG - SINGLE USE | SMOOTH BORE BREATHING CIRCUITS WITH WATER TRAPS AND EXTRA LIMB AND BAG - SINGLE USE Purpose: Used to deliver oxygen and Volatile Gases to patient's and to eliminate C02 during anaesthesia Transparent circuit with low resistance, non-extendable, smooth-bore tubing that is non-kinking, lightweight and crush resistant. The tubing must be pliable and must not retain a shape Spiral outer casing that supports the inner lumen to maintain tube integrity. Lightweight water traps centrally on inspiratory and expiratory limbs Must have an extra limb with an anti-static reservoir bag that gives good feedback regarding patient lung compliance. The reservoir bag must not twist on itself and must have a mechanism to prevent this. Connectors Malleable 22mm female connectors bonded to the ends of inspiratory and expiratory limbs – must fit securely to the machine. Spare 22mm M/M connector. Y-piece at patient's end. Separate angle-piece with luer lock port Good quality to ensure no leakage or easy disconnections Latex-free, non-pyrogenic, hypoallergenic, non-toxic and non-conductive medical grade plastic No variation in Circuit limb length or Extra limb will be accepted Clinically clean and individually packed, For SINGLE use only To comply with ISO 5367 Information brochure or the outer package of the item should clearly indicate latex free | Smooth bore breathing circuit with with water traps, extra limb and breathing bag - Single Use Adult Circuit: ± 1.8m; Extra limb: ± 1.2m OD: 22mm Reservoir bag: 2l
301 | RT13-60-011 | BREATHING CIRCUITS | SMOOTH BORE BREATHING CIRCUITS WITH WATER TRAPS AND EXTRA LIMB AND BAG - SINGLE USE | SMOOTH BORE BREATHING CIRCUITS WITH WATER TRAPS AND EXTRA LIMB AND BAG - SINGLE USE Purpose: Used to deliver oxygen and Volatile Gases to patient's and to eliminate C02 during anaesthesia Transparent circuit with low resistance, non-extendable, smooth-bore tubing that is non-kinking, lightweight and crush resistant. The tubing must be pliable and must not retain a shape Spiral outer casing that supports the inner lumen to maintain tube integrity. Lightweight water traps centrally on inspiratory and expiratory limbs Must have an extra limb with an anti-static reservoir bag that gives good feedback regarding patient lung compliance. The reservoir bag must not twist on itself and must have a mechanism to prevent this. Connectors Malleable 22mm female connectors bonded to the ends of inspiratory and expiratory limbs – must fit securely to the machine. Spare 22mm M/M connector. Y-piece at patient's end. Separate angle-piece with luer lock port Good quality to ensure no leakage or easy disconnections Latex-free, non-pyrogenic, hypoallergenic, non-toxic and non-conductive medical grade plastic No variation in Circuit limb length or Extra limb will be accepted Clinically clean and individually packed, For SINGLE use only To comply with ISO 5367 Information brochure or the outer package of the item should clearly indicate latex free | Smooth bore breathing circuit with with water traps, extra limb and breathing bag - Single Use Adult Circuit: ± 1.8m; Extra limb: ± 1.8m OD: 22mm Reservoir bag: 2l
302 | RT13-60-012 | BREATHING CIRCUITS | SMOOTH BORE BREATHING CIRCUITS WITH WATER TRAPS AND EXTRA LIMB AND BAG - SINGLE USE | SMOOTH BORE BREATHING CIRCUITS WITH WATER TRAPS AND EXTRA LIMB AND BAG - SINGLE USE Purpose: Used to deliver oxygen and Volatile Gases to patient's and to eliminate C02 during anaesthesia Transparent circuit with low resistance, non-extendable, smooth-bore tubing that is non-kinking, lightweight and crush resistant. The tubing must be pliable and must not retain a shape Spiral outer casing that supports the inner lumen to maintain tube integrity. Lightweight water traps centrally on inspiratory and expiratory limbs Must have an extra limb with an anti-static reservoir bag that gives good feedback regarding patient lung compliance. The reservoir bag must not twist on itself and must have a mechanism to prevent this. Connectors Malleable 22mm female connectors bonded to the ends of inspiratory and expiratory limbs – must fit securely to the machine. Spare 22mm M/M connector. Y-piece at patient's end. Separate angle-piece with luer lock port Good quality to ensure no leakage or easy disconnections Latex-free, non-pyrogenic, hypoallergenic, non-toxic and non-conductive medical grade plastic No variation in Circuit limb length or Extra limb will be accepted Clinically clean and individually packed, For SINGLE use only To comply with ISO 5367 Information brochure or the outer package of the item should clearly indicate latex free | Smooth bore breathing circuit with with water traps, extra limb and breathing bag - Single Use Adult Long Circuit: ± 2.6m; Extra limb: ± 1.2m OD: 22mm Reservoir bag: 2l
303 | RT13-60-013 | BREATHING CIRCUITS | SMOOTH BORE BREATHING CIRCUITS WITH WATER TRAPS AND EXTRA LIMB AND BAG - SINGLE USE | SMOOTH BORE BREATHING CIRCUITS WITH WATER TRAPS AND EXTRA LIMB AND BAG - SINGLE USE Purpose: Used to deliver oxygen and Volatile Gases to patient's and to eliminate C02 during anaesthesia Transparent circuit with low resistance, non-extendable, smooth-bore tubing that is non-kinking, lightweight and crush resistant. The tubing must be pliable and must not retain a shape Spiral outer casing that supports the inner lumen to maintain tube integrity. Lightweight water traps centrally on inspiratory and expiratory limbs Must have an extra limb with an anti-static reservoir bag that gives good feedback regarding patient lung compliance. The reservoir bag must not twist on itself and must have a mechanism to prevent this. Connectors Malleable 22mm female connectors bonded to the ends of inspiratory and expiratory limbs – must fit securely to the machine. Spare 22mm M/M connector. Y-piece at patient's end. Separate angle-piece with luer lock port Good quality to ensure no leakage or easy disconnections Latex-free, non-pyrogenic, hypoallergenic, non-toxic and non-conductive medical grade plastic No variation in Circuit limb length or Extra limb will be accepted Clinically clean and individually packed, For SINGLE use only To comply with ISO 5367 Information brochure or the outer package of the item should clearly indicate latex free | Smooth bore breathing circuit with with water traps, extra limb and breathing bag - Single Use Adult Long Circuit: ± 2.6m; Extra limb: ± 1.5m OD: 22mm Reservoir bag: 2l
304 | RT13-60-014 | BREATHING CIRCUITS | SMOOTH BORE BREATHING CIRCUITS WITH WATER TRAPS AND EXTRA LIMB AND BAG - SINGLE USE | SMOOTH BORE BREATHING CIRCUITS WITH WATER TRAPS AND EXTRA LIMB AND BAG - SINGLE USE Purpose: Used to deliver oxygen and Volatile Gases to patient's and to eliminate C02 during anaesthesia Transparent circuit with low resistance, non-extendable, smooth-bore tubing that is non-kinking, lightweight and crush resistant. The tubing must be pliable and must not retain a shape Spiral outer casing that supports the inner lumen to maintain tube integrity. Lightweight water traps centrally on inspiratory and expiratory limbs Must have an extra limb with an anti-static reservoir bag that gives good feedback regarding patient lung compliance. The reservoir bag must not twist on itself and must have a mechanism to prevent this. Connectors Malleable 22mm female connectors bonded to the ends of inspiratory and expiratory limbs – must fit securely to the machine. Spare 22mm M/M connector. Y-piece at patient's end. Separate angle-piece with luer lock port Good quality to ensure no leakage or easy disconnections Latex-free, non-pyrogenic, hypoallergenic, non-toxic and non-conductive medical grade plastic No variation in Circuit limb length or Extra limb will be accepted Clinically clean and individually packed, For SINGLE use only To comply with ISO 5367 Information brochure or the outer package of the item should clearly indicate latex free | Smooth bore breathing circuit with with water traps, extra limb and breathing bag - Single Use Adult Long Circuit: ± 2.6m; Extra limb: ± 1.8m OD: 22mm Reservoir bag: 2l
305 | RT13-60-015 | BREATHING CIRCUITS | SINGLE SMOOTH-BORE BREATHING TUBING - SINGLE USE | SINGLE SMOOTH-BORE BREATHING TUBING - SINGLE USE Purpose: Used as extension or extra limb in breathing circuits Transparent circuit with low resistance, non-extendable, smooth-bore tubing that is non-kinking, lightweight and crush resistant. The tubing must be pliable and must not retain a shape Spiral outer casing that supports the inner lumen to maintain tube integrity. Connectors Malleable 22mm female connectors bonded to the end of the tubing Spare connectors to be provided Latex-free, non-pyrogenic, hypoallergenic, non-toxic and non-conductive medical grade plastic Clinically clean and individually packed, For SINGLE use only To comply with ISO 5367 Information brochure or the outer package of the item should clearly indicate latex free | Single Smooth-bore Breathing Tubing - Single Use Adult 22mm Length: 1.2m Connector: 22mm M/F; 22mm M/M; 22mm F/F
306 | RT13-60-016 | BREATHING CIRCUITS | SINGLE SMOOTH-BORE BREATHING TUBING - SINGLE USE | SINGLE SMOOTH-BORE BREATHING TUBING - SINGLE USE Purpose: Used as extension or extra limb in breathing circuits Transparent circuit with low resistance, non-extendable, smooth-bore tubing that is non-kinking, lightweight and crush resistant. The tubing must be pliable and must not retain a shape Spiral outer casing that supports the inner lumen to maintain tube integrity. Connectors Malleable 22mm female connectors bonded to the end of the tubing Spare connectors to be provided Latex-free, non-pyrogenic, hypoallergenic, non-toxic and non-conductive medical grade plastic Clinically clean and individually packed, For SINGLE use only To comply with ISO 5367 Information brochure or the outer package of the item should clearly indicate latex free | Single Smooth-bore Breathing Tubing - Single Use Adult 22mm Length: 1.8m Connector: 22mm M/F; 22mm M/M; 22mm F/F
307 | RT13-60-017 | BREATHING CIRCUITS | SINGLE SMOOTH-BORE BREATHING TUBING - SINGLE USE | SINGLE SMOOTH-BORE BREATHING TUBING - SINGLE USE Purpose: Used as extension or extra limb in breathing circuits Transparent circuit with low resistance, non-extendable, smooth-bore tubing that is non-kinking, lightweight and crush resistant. The tubing must be pliable and must not retain a shape Spiral outer casing that supports the inner lumen to maintain tube integrity. Connectors Malleable 22mm female connectors bonded to the end of the tubing Spare connectors to be provided Latex-free, non-pyrogenic, hypoallergenic, non-toxic and non-conductive medical grade plastic Clinically clean and individually packed, For SINGLE use only To comply with ISO 5367 Information brochure or the outer package of the item should clearly indicate latex free | Single Smooth-bore Breathing Tubing - Single Use Paediatric 15mm Length: 1.2m Connector: 15mm M/F; 15mm M/M; 15mm F/F
308 | RT13-60-018 | BREATHING CIRCUITS | SINGLE SMOOTH-BORE BREATHING TUBING - SINGLE USE | SINGLE SMOOTH-BORE BREATHING TUBING - SINGLE USE Purpose: Used as extension or extra limb in breathing circuits Transparent circuit with low resistance, non-extendable, smooth-bore tubing that is non-kinking, lightweight and crush resistant. The tubing must be pliable and must not retain a shape Spiral outer casing that supports the inner lumen to maintain tube integrity. Connectors Malleable 22mm female connectors bonded to the end of the tubing Spare connectors to be provided Latex-free, non-pyrogenic, hypoallergenic, non-toxic and non-conductive medical grade plastic Clinically clean and individually packed, For SINGLE use only To comply with ISO 5367 Information brochure or the outer package of the item should clearly indicate latex free | Single Smooth-bore Breathing Tubing - Single Use Paediatric 15mm Length: 1.8m Connector: 15mm M/F; 15mm M/M; 15mm F/F
309 | RT13-61-001 | BREATHING CIRCUITS | CORRUGATED BREATHING CIRCUITS WITH EXTRA LIMB AND BAG - REUSABLE | CORRUGATED BREATHING CIRCUITS WITH EXTRA LIMB AND BAG - REUSABLE Purpose: Used to deliver oxygen and Volatile Gases to patient's and to eliminate C02 during anaesthesia Consists of a circuit with low resistance, non-extendable, corrugated tubing that is non-kinking, clear and lightweight and crush resistant. The tubing must be pliable and must not retain a shape The corrugations must be shallow to minimise flow resistance and turbulence. Must have an extra limb with an anti-static Reservoir bag that gives good feedback regarding patient lung compliance and does not twist on itself. Must have a mechanism to prevent this. Connectors Malleable 22mm female connectors at the end of inspiratory and expiratory limbs – must fit securely to the machine. Spare 22mm M/M connector. Y-piece at patient's end. Good quality to ensure no leakage or easy disconnections No variation in Circuit limb length or Extra limb will be accepted Manufactured from latex- free, non-pyrogenic hypoallergenic and non-toxic Silicone/ Hydro-tubing . Reusable, autoclavable, individually packed and surgical clean. Information brochure or the outer package of the item should clearly indicate latex free | Smooth bore breathing circuit with with water traps, extra limb and breathing bag - Single Use Adult Long Circuit: 2.6m; Extra limb: 1.2m OD: 22mm Reservoir bag: 2l
310 | RT13-61-002 | BREATHING CIRCUITS | CORRUGATED BREATHING CIRCUITS WITH EXTRA LIMB AND BAG - REUSABLE | CORRUGATED BREATHING CIRCUITS WITH EXTRA LIMB AND BAG - REUSABLE Purpose: Used to deliver oxygen and Volatile Gases to patient's and to eliminate C02 during anaesthesia Consists of a circuit with low resistance, non-extendable, corrugated tubing that is non-kinking, clear and lightweight and crush resistant. The tubing must be pliable and must not retain a shape The corrugations must be shallow to minimise flow resistance and turbulence. Must have an extra limb with an anti-static Reservoir bag that gives good feedback regarding patient lung compliance and does not twist on itself. Must have a mechanism to prevent this. Connectors Malleable 22mm female connectors at the end of inspiratory and expiratory limbs – must fit securely to the machine. Spare 22mm M/M connector. Y-piece at patient's end. Good quality to ensure no leakage or easy disconnections No variation in Circuit limb length or Extra limb will be accepted Manufactured from latex- free, non-pyrogenic hypoallergenic and non-toxic Silicone/ Hydro-tubing . Reusable, autoclavable, individually packed and surgical clean. Information brochure or the outer package of the item should clearly indicate latex free | Smooth bore breathing circuit with with water traps, extra limb and breathing bag - Single Use Adult Long Circuit: 2.6m; Extra limb: 1.5m OD: 22mm Reservoir bag: 2l
311 | RT13-61-003 | BREATHING CIRCUITS | CORRUGATED BREATHING CIRCUITS WITH EXTRA LIMB AND BAG - REUSABLE | CORRUGATED BREATHING CIRCUITS WITH EXTRA LIMB AND BAG - REUSABLE Purpose: Used to deliver oxygen and Volatile Gases to patient's and to eliminate C02 during anaesthesia Consists of a circuit with low resistance, non-extendable, corrugated tubing that is non-kinking, clear and lightweight and crush resistant. The tubing must be pliable and must not retain a shape The corrugations must be shallow to minimise flow resistance and turbulence. Must have an extra limb with an anti-static Reservoir bag that gives good feedback regarding patient lung compliance and does not twist on itself. Must have a mechanism to prevent this. Connectors Malleable 22mm female connectors at the end of inspiratory and expiratory limbs – must fit securely to the machine. Spare 22mm M/M connector. Y-piece at patient's end. Good quality to ensure no leakage or easy disconnections No variation in Circuit limb length or Extra limb will be accepted Manufactured from latex- free, non-pyrogenic hypoallergenic and non-toxic Silicone/ Hydro-tubing . Reusable, autoclavable, individually packed and surgical clean. Information brochure or the outer package of the item should clearly indicate latex free | Smooth bore breathing circuit with with water traps, extra limb and breathing bag - Single Use Adult Long Circuit: 2.6m; Extra limb: 1.8m OD: 22mm Reservoir bag: 2l
312 | RT13-61-004 | BREATHING CIRCUITS | CORRUGATED BREATHING CIRCUITS WITH EXTRA LIMB AND BAG - REUSABLE | CORRUGATED BREATHING CIRCUITS WITH EXTRA LIMB AND BAG - REUSABLE Purpose: Used to deliver oxygen and Volatile Gases to patient's and to eliminate C02 during anaesthesia Consists of a circuit with low resistance, non-extendable, corrugated tubing that is non-kinking, clear and lightweight and crush resistant. The tubing must be pliable and must not retain a shape The corrugations must be shallow to minimise flow resistance and turbulence. Must have an extra limb with an anti-static Reservoir bag that gives good feedback regarding patient lung compliance and does not twist on itself. Must have a mechanism to prevent this. Connectors Malleable 22mm female connectors at the end of inspiratory and expiratory limbs – must fit securely to the machine. Spare 22mm M/M connector. Y-piece at patient's end. Good quality to ensure no leakage or easy disconnections No variation in Circuit limb length or Extra limb will be accepted Manufactured from latex- free, non-pyrogenic hypoallergenic and non-toxic Silicone/ Hydro-tubing . Reusable, autoclavable, individually packed and surgical clean. Information brochure or the outer package of the item should clearly indicate latex free | Single Smooth-bore Breathing Tubing - Single Use Adult 22mm Length: 1.2m Connector: 22mm M/F; 22mm M/M; 22mm F/F
313 | RT13-61-005 | BREATHING CIRCUITS | CORRUGATED BREATHING CIRCUITS WITH EXTRA LIMB AND BAG - REUSABLE | CORRUGATED BREATHING CIRCUITS WITH EXTRA LIMB AND BAG - REUSABLE Purpose: Used to deliver oxygen and Volatile Gases to patient's and to eliminate C02 during anaesthesia Consists of a circuit with low resistance, non-extendable, corrugated tubing that is non-kinking, clear and lightweight and crush resistant. The tubing must be pliable and must not retain a shape The corrugations must be shallow to minimise flow resistance and turbulence. Must have an extra limb with an anti-static Reservoir bag that gives good feedback regarding patient lung compliance and does not twist on itself. Must have a mechanism to prevent this. Connectors Malleable 22mm female connectors at the end of inspiratory and expiratory limbs – must fit securely to the machine. Spare 22mm M/M connector. Y-piece at patient's end. Good quality to ensure no leakage or easy disconnections No variation in Circuit limb length or Extra limb will be accepted Manufactured from latex- free, non-pyrogenic hypoallergenic and non-toxic Silicone/ Hydro-tubing . Reusable, autoclavable, individually packed and surgical clean. Information brochure or the outer package of the item should clearly indicate latex free | Single Smooth-bore Breathing Tubing - Single Use Adult 22mm Length: 1.8m Connector: 22mm M/F; 22mm M/M; 22mm F/F
314 | RT13-61-006 | BREATHING CIRCUITS | CORRUGATED BREATHING CIRCUITS WITH EXTRA LIMB AND BAG - REUSABLE | CORRUGATED BREATHING CIRCUITS WITH EXTRA LIMB AND BAG - REUSABLE Purpose: Used to deliver oxygen and Volatile Gases to patient's and to eliminate C02 during anaesthesia Consists of a circuit with low resistance, non-extendable, corrugated tubing that is non-kinking, clear and lightweight and crush resistant. The tubing must be pliable and must not retain a shape The corrugations must be shallow to minimise flow resistance and turbulence. Must have an extra limb with an anti-static Reservoir bag that gives good feedback regarding patient lung compliance and does not twist on itself. Must have a mechanism to prevent this. Connectors Malleable 22mm female connectors at the end of inspiratory and expiratory limbs – must fit securely to the machine. Spare 22mm M/M connector. Y-piece at patient's end. Good quality to ensure no leakage or easy disconnections No variation in Circuit limb length or Extra limb will be accepted Manufactured from latex- free, non-pyrogenic hypoallergenic and non-toxic Silicone/ Hydro-tubing . Reusable, autoclavable, individually packed and surgical clean. Information brochure or the outer package of the item should clearly indicate latex free | Single Smooth-bore Breathing Tubing - Single Use Paediatric 15mm Length: 1.2m Connector: 15mm M/F; 15mm M/M; 15mm F/F
315 | RT13-61-007 | BREATHING CIRCUITS | CORRUGATED BREATHING CIRCUITS WITH EXTRA LIMB AND BAG - REUSABLE | CORRUGATED BREATHING CIRCUITS WITH EXTRA LIMB AND BAG - REUSABLE Purpose: Used to deliver oxygen and Volatile Gases to patient's and to eliminate C02 during anaesthesia Consists of a circuit with low resistance, non-extendable, corrugated tubing that is non-kinking, clear and lightweight and crush resistant. The tubing must be pliable and must not retain a shape The corrugations must be shallow to minimise flow resistance and turbulence. Must have an extra limb with an anti-static Reservoir bag that gives good feedback regarding patient lung compliance and does not twist on itself. Must have a mechanism to prevent this. Connectors Malleable 22mm female connectors at the end of inspiratory and expiratory limbs – must fit securely to the machine. Spare 22mm M/M connector. Y-piece at patient's end. Good quality to ensure no leakage or easy disconnections No variation in Circuit limb length or Extra limb will be accepted Manufactured from latex- free, non-pyrogenic hypoallergenic and non-toxic Silicone/ Hydro-tubing . Reusable, autoclavable, individually packed and surgical clean. Information brochure or the outer package of the item should clearly indicate latex free | Single Smooth-bore Breathing Tubing - Single Use Paediatric 15mm Length: 1.8m Connector: 15mm M/F; 15mm M/M; 15mm F/F
316 | RT13-63-004 | Resuscitation and Re-breathing bags | REBREATHING BAGS REUSABLE | REBREATHING BAGS REUSABLE Purpose: Stores gas between respiratory cycles. Provides information to anaesthetist in manual ventilation regarding pulmonary compliance and airway patency. Consists of alatex-free antistatic, non-conductive reservoir bag with a tapered neck and collar Compliance: 30-60cmH20 when filled to 4 X nominal volume Must hold at least 10 X nominal volume before it will burst Must retain original volume and shape (within 10%) at deflation Must must not adhere to itself and must not twist or kink on itself –must have a mechanism to prevent this. Must clearly indicate that it is Autoclavable on the bag To comply with ISO 5362 Information brochure or the outer package of the item should clearly indicate latex free | Rebreathing reservoir bag 0.5 litre with 15mm F connector - Reusable
317 | RT13-63-005 | Resuscitation and Re-breathing bags | REBREATHING BAGS REUSABLE | REBREATHING BAGS REUSABLE Purpose: Stores gas between respiratory cycles. Provides information to anaesthetist in manual ventilation regarding pulmonary compliance and airway patency. Consists of alatex-free antistatic, non-conductive reservoir bag with a tapered neck and collar Compliance: 30-60cmH20 when filled to 4 X nominal volume Must hold at least 10 X nominal volume before it will burst Must retain original volume and shape (within 10%) at deflation Must must not adhere to itself and must not twist or kink on itself –must have a mechanism to prevent this. Must clearly indicate that it is Autoclavable on the bag To comply with ISO 5362 Information brochure or the outer package of the item should clearly indicate latex free | Rebreathing reservoir bag 1 litre with 15mm F connector - Reusable
318 | RT13-63-006 | Resuscitation and Re-breathing bags | REBREATHING BAGS REUSABLE | REBREATHING BAGS REUSABLE Purpose: Stores gas between respiratory cycles. Provides information to anaesthetist in manual ventilation regarding pulmonary compliance and airway patency. Consists of alatex-free antistatic, non-conductive reservoir bag with a tapered neck and collar Compliance: 30-60cmH20 when filled to 4 X nominal volume Must hold at least 10 X nominal volume before it will burst Must retain original volume and shape (within 10%) at deflation Must must not adhere to itself and must not twist or kink on itself –must have a mechanism to prevent this. Must clearly indicate that it is Autoclavable on the bag To comply with ISO 5362 Information brochure or the outer package of the item should clearly indicate latex free | Rebreathing reservoir bag 2 litre with 22mm F connector- Reusable
319 | RT13-64-001 | BREATHING CIRCUITS | CORRUGATED BREATHING CIRCUITS WITH WATER TRAPS AND EXTRA LIMB AND BAG - SINGLE USE | CORRUGATED BREATHING CIRCUITS WITH WATER TRAPS AND EXTRA LIMB AND BAG - SINGLE USE Purpose: Used to deliver oxygen and Volatile Gases to patient's and to eliminate C02 during anaesthesia Transparent circuit with low resistance, non-extendable, corrugated tubing that is non-kinking, lightweight and crush resistant. The tubing must be pliable and must not retain a shape Shallow corrugations to minimise flow resistance and turbulence. Uninterrupted tubing (not joined). Lightweight water traps centrally on inspiratory and expiratory limb Must have an extra limb with an anti-static reservoir bag that gives good feedback regarding patient lung compliance. The reservoir bag must not twist on itself and must have a mechanism to prevent this. Connectors Malleable 22mm female connectors bonded to the ends of inspiratory and expiratory limbs – must fit securely to the machine. Spare 22mm M/M connector. Y-piece at patient's end. Separate angle-piece with luer lock port Good quality to ensure no leakage or easy disconnections Latex-free, non-pyrogenic, hypoallergenic, non-toxic and non-conductive medical grade plastic No variation in Circuit limb length or Extra limb will be accepted Clinically clean and individually packed, For SINGLE use only To comply with ISO 5367 Information brochure or the outer package of the item should clearly indicate latex free | Corrugated breathing circuit with water traps and extra limb with breathing bag - Single Use - Paediatric > 3 yrs 15mm Circuit: 1.6m; 15mm Extra limb: 1.2m Reservoir bag: 1l
320 | RT13-64-002 | BREATHING CIRCUITS | CORRUGATED BREATHING CIRCUITS WITH WATER TRAPS AND EXTRA LIMB AND BAG - SINGLE USE | CORRUGATED BREATHING CIRCUITS WITH WATER TRAPS AND EXTRA LIMB AND BAG - SINGLE USE Purpose: Used to deliver oxygen and Volatile Gases to patient's and to eliminate C02 during anaesthesia Transparent circuit with low resistance, non-extendable, corrugated tubing that is non-kinking, lightweight and crush resistant. The tubing must be pliable and must not retain a shape Shallow corrugations to minimise flow resistance and turbulence. Uninterrupted tubing (not joined). Lightweight water traps centrally on inspiratory and expiratory limb Must have an extra limb with an anti-static reservoir bag that gives good feedback regarding patient lung compliance. The reservoir bag must not twist on itself and must have a mechanism to prevent this. Connectors Malleable 22mm female connectors bonded to the ends of inspiratory and expiratory limbs – must fit securely to the machine. Spare 22mm M/M connector. Y-piece at patient's end. Separate angle-piece with luer lock port Good quality to ensure no leakage or easy disconnections Latex-free, non-pyrogenic, hypoallergenic, non-toxic and non-conductive medical grade plastic No variation in Circuit limb length or Extra limb will be accepted Clinically clean and individually packed, For SINGLE use only To comply with ISO 5367 Information brochure or the outer package of the item should clearly indicate latex free | Corrugated breathing circuit with water traps and extra limb with breathing bag - Single use Paediatric > 3 yrs 15mm Circuit: 1.6m; 15mm Extra limb: 1.8m Reservoir bag: 1l
321 | RT13-64-003 | BREATHING CIRCUITS | CORRUGATED BREATHING CIRCUITS WITH WATER TRAPS AND EXTRA LIMB AND BAG - SINGLE USE | CORRUGATED BREATHING CIRCUITS WITH WATER TRAPS AND EXTRA LIMB AND BAG - SINGLE USE Purpose: Used to deliver oxygen and Volatile Gases to patient's and to eliminate C02 during anaesthesia Transparent circuit with low resistance, non-extendable, corrugated tubing that is non-kinking, lightweight and crush resistant. The tubing must be pliable and must not retain a shape Shallow corrugations to minimise flow resistance and turbulence. Uninterrupted tubing (not joined). Lightweight water traps centrally on inspiratory and expiratory limb Must have an extra limb with an anti-static reservoir bag that gives good feedback regarding patient lung compliance. The reservoir bag must not twist on itself and must have a mechanism to prevent this. Connectors Malleable 22mm female connectors bonded to the ends of inspiratory and expiratory limbs – must fit securely to the machine. Spare 22mm M/M connector. Y-piece at patient's end. Separate angle-piece with luer lock port Good quality to ensure no leakage or easy disconnections Latex-free, non-pyrogenic, hypoallergenic, non-toxic and non-conductive medical grade plastic No variation in Circuit limb length or Extra limb will be accepted Clinically clean and individually packed, For SINGLE use only To comply with ISO 5367 Information brochure or the outer package of the item should clearly indicate latex free | Corrugated breathing circuit with water traps and extra limb with breathing bag - Single Use Adult 22mm Circuit: 1.8m; 22mm Extra limb: 1.2m OD: 22mm Reservoir bag: 2l
322 | RT13-64-005 | BREATHING CIRCUITS | CORRUGATED BREATHING CIRCUITS WITH WATER TRAPS AND EXTRA LIMB AND BAG - SINGLE USE | CORRUGATED BREATHING CIRCUITS WITH WATER TRAPS AND EXTRA LIMB AND BAG - SINGLE USE Purpose: Used to deliver oxygen and Volatile Gases to patient's and to eliminate C02 during anaesthesia Transparent circuit with low resistance, non-extendable, corrugated tubing that is non-kinking, lightweight and crush resistant. The tubing must be pliable and must not retain a shape Shallow corrugations to minimise flow resistance and turbulence. Uninterrupted tubing (not joined). Lightweight water traps centrally on inspiratory and expiratory limb Must have an extra limb with an anti-static reservoir bag that gives good feedback regarding patient lung compliance. The reservoir bag must not twist on itself and must have a mechanism to prevent this. Connectors Malleable 22mm female connectors bonded to the ends of inspiratory and expiratory limbs – must fit securely to the machine. Spare 22mm M/M connector. Y-piece at patient's end. Separate angle-piece with luer lock port Good quality to ensure no leakage or easy disconnections Latex-free, non-pyrogenic, hypoallergenic, non-toxic and non-conductive medical grade plastic No variation in Circuit limb length or Extra limb will be accepted Clinically clean and individually packed, For SINGLE use only To comply with ISO 5367 Information brochure or the outer package of the item should clearly indicate latex free | Corrugated breathing circuit with water traps and extra limb with breathing bag - Single Use Adult 22mm Circuit: 1.8m; 22mm Extra limb: 1.8m OD: 22mm Reservoir bag: 2l
323 | RT13-64-006 | BREATHING CIRCUITS | CORRUGATED BREATHING CIRCUITS WITH WATER TRAPS AND EXTRA LIMB AND BAG - SINGLE USE | CORRUGATED BREATHING CIRCUITS WITH WATER TRAPS AND EXTRA LIMB AND BAG - SINGLE USE Purpose: Used to deliver oxygen and Volatile Gases to patient's and to eliminate C02 during anaesthesia Transparent circuit with low resistance, non-extendable, corrugated tubing that is non-kinking, lightweight and crush resistant. The tubing must be pliable and must not retain a shape Shallow corrugations to minimise flow resistance and turbulence. Uninterrupted tubing (not joined). Lightweight water traps centrally on inspiratory and expiratory limb Must have an extra limb with an anti-static reservoir bag that gives good feedback regarding patient lung compliance. The reservoir bag must not twist on itself and must have a mechanism to prevent this. Connectors Malleable 22mm female connectors bonded to the ends of inspiratory and expiratory limbs – must fit securely to the machine. Spare 22mm M/M connector. Y-piece at patient's end. Separate angle-piece with luer lock port Good quality to ensure no leakage or easy disconnections Latex-free, non-pyrogenic, hypoallergenic, non-toxic and non-conductive medical grade plastic No variation in Circuit limb length or Extra limb will be accepted Clinically clean and individually packed, For SINGLE use only To comply with ISO 5367 Information brochure or the outer package of the item should clearly indicate latex free | Corrugated breathing circuit with water traps and extra limb with breathing bag - Single Use Adult Long 22mm Circuit: 2.5m; 22mm Extra limb: 1.2m OD: 22mm Reservoir bag: 2l
324 | RT13-64-007 | BREATHING CIRCUITS | CORRUGATED BREATHING CIRCUITS WITH WATER TRAPS AND EXTRA LIMB AND BAG - SINGLE USE | CORRUGATED BREATHING CIRCUITS WITH WATER TRAPS AND EXTRA LIMB AND BAG - SINGLE USE Purpose: Used to deliver oxygen and Volatile Gases to patient's and to eliminate C02 during anaesthesia Transparent circuit with low resistance, non-extendable, corrugated tubing that is non-kinking, lightweight and crush resistant. The tubing must be pliable and must not retain a shape Shallow corrugations to minimise flow resistance and turbulence. Uninterrupted tubing (not joined). Lightweight water traps centrally on inspiratory and expiratory limb Must have an extra limb with an anti-static reservoir bag that gives good feedback regarding patient lung compliance. The reservoir bag must not twist on itself and must have a mechanism to prevent this. Connectors Malleable 22mm female connectors bonded to the ends of inspiratory and expiratory limbs – must fit securely to the machine. Spare 22mm M/M connector. Y-piece at patient's end. Separate angle-piece with luer lock port Good quality to ensure no leakage or easy disconnections Latex-free, non-pyrogenic, hypoallergenic, non-toxic and non-conductive medical grade plastic No variation in Circuit limb length or Extra limb will be accepted Clinically clean and individually packed, For SINGLE use only To comply with ISO 5367 Information brochure or the outer package of the item should clearly indicate latex free | Corrugated breathing circuit with water traps and extra limb with breathing bag - Single Use Adult Long 22mm Circuit: 2.5m; 22mm Extra limb: 1.5m OD: 22mm Reservoir bag: 2l
325 | RT13-64-008 | BREATHING CIRCUITS | CORRUGATED BREATHING CIRCUITS WITH WATER TRAPS AND EXTRA LIMB AND BAG - SINGLE USE | CORRUGATED BREATHING CIRCUITS WITH WATER TRAPS AND EXTRA LIMB AND BAG - SINGLE USE Purpose: Used to deliver oxygen and Volatile Gases to patient's and to eliminate C02 during anaesthesia Transparent circuit with low resistance, non-extendable, corrugated tubing that is non-kinking, lightweight and crush resistant. The tubing must be pliable and must not retain a shape Shallow corrugations to minimise flow resistance and turbulence. Uninterrupted tubing (not joined). Lightweight water traps centrally on inspiratory and expiratory limb Must have an extra limb with an anti-static reservoir bag that gives good feedback regarding patient lung compliance. The reservoir bag must not twist on itself and must have a mechanism to prevent this. Connectors Malleable 22mm female connectors bonded to the ends of inspiratory and expiratory limbs – must fit securely to the machine. Spare 22mm M/M connector. Y-piece at patient's end. Separate angle-piece with luer lock port Good quality to ensure no leakage or easy disconnections Latex-free, non-pyrogenic, hypoallergenic, non-toxic and non-conductive medical grade plastic No variation in Circuit limb length or Extra limb will be accepted Clinically clean and individually packed, For SINGLE use only To comply with ISO 5367 Information brochure or the outer package of the item should clearly indicate latex free | Corrugated breathing circuit with water traps and extra limb with breathing bag - Single Use Adult Long 22mm Circuit: 2.5m; 22mm Extra limb: 1.8m OD: 22mm Reservoir bag: 2l
326 | RT13-65-001 | BREATHING CIRCUITS | SINGLE CORRUGATED BREATHING TUBING - SINGLE USE | SINGLE CORRUGATED BREATHING TUBING - SINGLE USE Purpose: Used as extension or extra limb in breathing circuits Transparent circuit with low resistance, non-extendable, corrugated tubing that is non-kinking, lightweight and crush resistant. The tubing must be pliable and must not retain a shape Shallow corrugations to minimise flow resistance and turbulence. Connectors Malleable 22mm female connectors bonded to the end of the tubing Spare connectors to be provided Latex-free, non-pyrogenic, hypoallergenic, non-toxic and non-conductive medical grade plastic No variation in Circuit limb length or Extra limb will be accepted Clinically clean and individually packed, For SINGLE use only To comply with ISO 5367 Information brochure or the outer package of the item should clearly indicate latex free | Corrugated breathing circuit with water traps and extra limb with breathing bag - Single Use - Paediatric > 3 yrs 15mm Circuit: 1.6m; 15mm Extra limb: 1.2m Reservoir bag: 1l
327 | RT13-65-002 | BREATHING CIRCUITS | SINGLE CORRUGATED BREATHING TUBING - SINGLE USE | SINGLE CORRUGATED BREATHING TUBING - SINGLE USE Purpose: Used as extension or extra limb in breathing circuits Transparent circuit with low resistance, non-extendable, corrugated tubing that is non-kinking, lightweight and crush resistant. The tubing must be pliable and must not retain a shape Shallow corrugations to minimise flow resistance and turbulence. Connectors Malleable 22mm female connectors bonded to the end of the tubing Spare connectors to be provided Latex-free, non-pyrogenic, hypoallergenic, non-toxic and non-conductive medical grade plastic No variation in Circuit limb length or Extra limb will be accepted Clinically clean and individually packed, For SINGLE use only To comply with ISO 5367 Information brochure or the outer package of the item should clearly indicate latex free | Corrugated breathing circuit with water traps and extra limb with breathing bag - Single use Paediatric > 3 yrs 15mm Circuit: 1.6m; 15mm Extra limb: 1.8m Reservoir bag: 1l
328 | RT13-65-003 | BREATHING CIRCUITS | SINGLE CORRUGATED BREATHING TUBING - SINGLE USE | SINGLE CORRUGATED BREATHING TUBING - SINGLE USE Purpose: Used as extension or extra limb in breathing circuits Transparent circuit with low resistance, non-extendable, corrugated tubing that is non-kinking, lightweight and crush resistant. The tubing must be pliable and must not retain a shape Shallow corrugations to minimise flow resistance and turbulence. Connectors Malleable 22mm female connectors bonded to the end of the tubing Spare connectors to be provided Latex-free, non-pyrogenic, hypoallergenic, non-toxic and non-conductive medical grade plastic No variation in Circuit limb length or Extra limb will be accepted Clinically clean and individually packed, For SINGLE use only To comply with ISO 5367 Information brochure or the outer package of the item should clearly indicate latex free | Corrugated breathing circuit with water traps and extra limb with breathing bag - Single Use Adult 22mm Circuit: 1.8m; 22mm Extra limb: 1.2m OD: 22mm Reservoir bag: 2l
329 | RT13-68-015 | BREATHING CIRCUITS | SINGLE SMOOTH-BORE BREATHING TUBING - SINGLE USE | SINGLE SMOOTH-BORE BREATHING TUBING - SINGLE USE Purpose: Used as extension or extra limb in breathing circuits Transparent circuit with low resistance, non-extendable, smooth-bore tubing that is non-kinking, lightweight and crush resistant. The tubing must be pliable and must not retain a shape Spiral outer casing that supports the inner lumen to maintain tube integrity. Connectors Malleable 22mm female connectors bonded to the end of the tubing Spare connectors to be provided Latex-free, non-pyrogenic, hypoallergenic, non-toxic and non-conductive medical grade plastic Clinically clean and individually packed, For SINGLE use only To comply with ISO 5367 Information brochure or the outer package of the item should clearly indicate latex free | Single Smooth-bore Breathing Tubing - Single Use Adult 22mm Length: ± 1.2m Connector: 22mm M/F; 22mm M/M; 22mm F/F
330 | RT13-68-016 | BREATHING CIRCUITS | SINGLE SMOOTH-BORE BREATHING TUBING - SINGLE USE | SINGLE SMOOTH-BORE BREATHING TUBING - SINGLE USE Purpose: Used as extension or extra limb in breathing circuits Transparent circuit with low resistance, non-extendable, smooth-bore tubing that is non-kinking, lightweight and crush resistant. The tubing must be pliable and must not retain a shape Spiral outer casing that supports the inner lumen to maintain tube integrity. Connectors Malleable 22mm female connectors bonded to the end of the tubing Spare connectors to be provided Latex-free, non-pyrogenic, hypoallergenic, non-toxic and non-conductive medical grade plastic Clinically clean and individually packed, For SINGLE use only To comply with ISO 5367 Information brochure or the outer package of the item should clearly indicate latex free | Single Smooth-bore Breathing Tubing - Single Use Adult 22mm Length: ± 1.8m Connector: 22mm M/F; 22mm M/M; 22mm F/F
331 | RT13-68-017 | BREATHING CIRCUITS | SINGLE SMOOTH-BORE BREATHING TUBING - SINGLE USE | SINGLE SMOOTH-BORE BREATHING TUBING - SINGLE USE Purpose: Used as extension or extra limb in breathing circuits Transparent circuit with low resistance, non-extendable, smooth-bore tubing that is non-kinking, lightweight and crush resistant. The tubing must be pliable and must not retain a shape Spiral outer casing that supports the inner lumen to maintain tube integrity. Connectors Malleable 22mm female connectors bonded to the end of the tubing Spare connectors to be provided Latex-free, non-pyrogenic, hypoallergenic, non-toxic and non-conductive medical grade plastic Clinically clean and individually packed, For SINGLE use only To comply with ISO 5367 Information brochure or the outer package of the item should clearly indicate latex free | Single Smooth-bore Breathing Tubing - Single Use Paediatric 15mm Length: ± 1.2m Connector: 15mm M/F; 15mm M/M; 15mm F/F
332 | RT13-68-018 | BREATHING CIRCUITS | SINGLE SMOOTH-BORE BREATHING TUBING - SINGLE USE | SINGLE SMOOTH-BORE BREATHING TUBING - SINGLE USE Purpose: Used as extension or extra limb in breathing circuits Transparent circuit with low resistance, non-extendable, smooth-bore tubing that is non-kinking, lightweight and crush resistant. The tubing must be pliable and must not retain a shape Spiral outer casing that supports the inner lumen to maintain tube integrity. Connectors Malleable 22mm female connectors bonded to the end of the tubing Spare connectors to be provided Latex-free, non-pyrogenic, hypoallergenic, non-toxic and non-conductive medical grade plastic Clinically clean and individually packed, For SINGLE use only To comply with ISO 5367 Information brochure or the outer package of the item should clearly indicate latex free | Single Smooth-bore Breathing Tubing - Single Use Paediatric 15mm Length: ± 1.8m Connector: 15mm M/F; 15mm M/M; 15mm F/F
333 | RT13-69-001 | BREATHING CIRCUITS | CORRUGATED BREATHING CIRCUITS WITH EXTRA LIMB AND BAG - REUSABLE | CORRUGATED BREATHING CIRCUITS WITH EXTRA LIMB AND BAG - REUSABLE Purpose: Used to deliver oxygen and Volatile Gases to patient's and to eliminate C02 during anaesthesia Consists of a circuit with low resistance, non-extendable, corrugated tubing that is non-kinking, clear and lightweight and crush resistant. The tubing must be pliable and must not retain a shape The corrugations must be shallow to minimise flow resistance and turbulence. Must have an extra limb with an anti-static Reservoir bag that gives good feedback regarding patient lung compliance and does not twist on itself. Must have a mechanism to prevent this. Connectors Malleable 22mm female connectors at the end of inspiratory and expiratory limbs – must fit securely to the machine. Spare 22mm M/M connector. Y-piece at patient's end. Good quality to ensure no leakage or easy disconnections No variation in Circuit limb length or Extra limb will be accepted Manufactured from latex- free, non-pyrogenic hypoallergenic and non-toxic Silicone/ Hydro-tubing . Reusable, autoclavable, individually packed and surgical clean. Information brochure or the outer package of the item should clearly indicate latex free | Corrugated breathing circuit with extra limb and breathing bag - Reusable- Paediatric > 3 yrs Circuit: ± 1.6m; Extra limb: ± 1.2m OD: 15mm Reservoir bag: 500ml
334 | RT13-69-002 | BREATHING CIRCUITS | CORRUGATED BREATHING CIRCUITS WITH EXTRA LIMB AND BAG - REUSABLE | CORRUGATED BREATHING CIRCUITS WITH EXTRA LIMB AND BAG - REUSABLE Purpose: Used to deliver oxygen and Volatile Gases to patient's and to eliminate C02 during anaesthesia Consists of a circuit with low resistance, non-extendable, corrugated tubing that is non-kinking, clear and lightweight and crush resistant. The tubing must be pliable and must not retain a shape The corrugations must be shallow to minimise flow resistance and turbulence. Must have an extra limb with an anti-static Reservoir bag that gives good feedback regarding patient lung compliance and does not twist on itself. Must have a mechanism to prevent this. Connectors Malleable 22mm female connectors at the end of inspiratory and expiratory limbs – must fit securely to the machine. Spare 22mm M/M connector. Y-piece at patient's end. Good quality to ensure no leakage or easy disconnections No variation in Circuit limb length or Extra limb will be accepted Manufactured from latex- free, non-pyrogenic hypoallergenic and non-toxic Silicone/ Hydro-tubing . Reusable, autoclavable, individually packed and surgical clean. Information brochure or the outer package of the item should clearly indicate latex free | Corrugated breathing circuit with extra limb and breathing bag - Reusable- Paediatric > 3 yrs Circuit: ± 1.6m; Extra limb: ± 1.8m OD: 15mm Reservoir bag: 500ml
335 | RT13-69-003 | BREATHING CIRCUITS | CORRUGATED BREATHING CIRCUITS WITH EXTRA LIMB AND BAG - REUSABLE | CORRUGATED BREATHING CIRCUITS WITH EXTRA LIMB AND BAG - REUSABLE Purpose: Used to deliver oxygen and Volatile Gases to patient's and to eliminate C02 during anaesthesia Consists of a circuit with low resistance, non-extendable, corrugated tubing that is non-kinking, clear and lightweight and crush resistant. The tubing must be pliable and must not retain a shape The corrugations must be shallow to minimise flow resistance and turbulence. Must have an extra limb with an anti-static Reservoir bag that gives good feedback regarding patient lung compliance and does not twist on itself. Must have a mechanism to prevent this. Connectors Malleable 22mm female connectors at the end of inspiratory and expiratory limbs – must fit securely to the machine. Spare 22mm M/M connector. Y-piece at patient's end. Good quality to ensure no leakage or easy disconnections No variation in Circuit limb length or Extra limb will be accepted Manufactured from latex- free, non-pyrogenic hypoallergenic and non-toxic Silicone/ Hydro-tubing . Reusable, autoclavable, individually packed and surgical clean. Information brochure or the outer package of the item should clearly indicate latex free | Corrugated breathing circuit with extra limb and breathing bag - Reusable- Paediatric > 3 yrs Circuit: ± 1.6m; Extra limb: ± 1.2m OD: 15mm Reservoir bag: 1l
336 | RT13-69-004 | BREATHING CIRCUITS | CORRUGATED BREATHING CIRCUITS WITH EXTRA LIMB AND BAG - REUSABLE | CORRUGATED BREATHING CIRCUITS WITH EXTRA LIMB AND BAG - REUSABLE Purpose: Used to deliver oxygen and Volatile Gases to patient's and to eliminate C02 during anaesthesia Consists of a circuit with low resistance, non-extendable, corrugated tubing that is non-kinking, clear and lightweight and crush resistant. The tubing must be pliable and must not retain a shape The corrugations must be shallow to minimise flow resistance and turbulence. Must have an extra limb with an anti-static Reservoir bag that gives good feedback regarding patient lung compliance and does not twist on itself. Must have a mechanism to prevent this. Connectors Malleable 22mm female connectors at the end of inspiratory and expiratory limbs – must fit securely to the machine. Spare 22mm M/M connector. Y-piece at patient's end. Good quality to ensure no leakage or easy disconnections No variation in Circuit limb length or Extra limb will be accepted Manufactured from latex- free, non-pyrogenic hypoallergenic and non-toxic Silicone/ Hydro-tubing . Reusable, autoclavable, individually packed and surgical clean. Information brochure or the outer package of the item should clearly indicate latex free | Corrugated breathing circuit with extra limb and breathing bag - Reusable- Paediatric > 3 yrs Circuit: ± 1.6m; Extra limb: ± 1.8m OD: 15mm Reservoir bag: 1l
337 | RT13-69-005 | BREATHING CIRCUITS | CORRUGATED BREATHING CIRCUITS WITH EXTRA LIMB AND BAG - REUSABLE | CORRUGATED BREATHING CIRCUITS WITH EXTRA LIMB AND BAG - REUSABLE Purpose: Used to deliver oxygen and Volatile Gases to patient's and to eliminate C02 during anaesthesia Consists of a circuit with low resistance, non-extendable, corrugated tubing that is non-kinking, clear and lightweight and crush resistant. The tubing must be pliable and must not retain a shape The corrugations must be shallow to minimise flow resistance and turbulence. Must have an extra limb with an anti-static Reservoir bag that gives good feedback regarding patient lung compliance and does not twist on itself. Must have a mechanism to prevent this. Connectors Malleable 22mm female connectors at the end of inspiratory and expiratory limbs – must fit securely to the machine. Spare 22mm M/M connector. Y-piece at patient's end. Good quality to ensure no leakage or easy disconnections No variation in Circuit limb length or Extra limb will be accepted Manufactured from latex- free, non-pyrogenic hypoallergenic and non-toxic Silicone/ Hydro-tubing . Reusable, autoclavable, individually packed and surgical clean. Information brochure or the outer package of the item should clearly indicate latex free | Corrugated breathing circuit with extra limb and breathing bag - Reusable - Adult Circuit: ± 1.8m; Extra limb: ± 1.2m OD: 22mm Reservoir bag: 2l
338 | RT13-69-006 | BREATHING CIRCUITS | CORRUGATED BREATHING CIRCUITS WITH EXTRA LIMB AND BAG - REUSABLE | CORRUGATED BREATHING CIRCUITS WITH EXTRA LIMB AND BAG - REUSABLE Purpose: Used to deliver oxygen and Volatile Gases to patient's and to eliminate C02 during anaesthesia Consists of a circuit with low resistance, non-extendable, corrugated tubing that is non-kinking, clear and lightweight and crush resistant. The tubing must be pliable and must not retain a shape The corrugations must be shallow to minimise flow resistance and turbulence. Must have an extra limb with an anti-static Reservoir bag that gives good feedback regarding patient lung compliance and does not twist on itself. Must have a mechanism to prevent this. Connectors Malleable 22mm female connectors at the end of inspiratory and expiratory limbs – must fit securely to the machine. Spare 22mm M/M connector. Y-piece at patient's end. Good quality to ensure no leakage or easy disconnections No variation in Circuit limb length or Extra limb will be accepted Manufactured from latex- free, non-pyrogenic hypoallergenic and non-toxic Silicone/ Hydro-tubing . Reusable, autoclavable, individually packed and surgical clean. Information brochure or the outer package of the item should clearly indicate latex free | Corrugated breathing circuit with extra limb and breathing bag - Reusable - Adult Circuit: ± 1.8m; Extra limb: ± 1.5m OD: 22mm Reservoir bag: 2l
339 | RT13-69-007 | BREATHING CIRCUITS | CORRUGATED BREATHING CIRCUITS WITH EXTRA LIMB AND BAG - REUSABLE | CORRUGATED BREATHING CIRCUITS WITH EXTRA LIMB AND BAG - REUSABLE Purpose: Used to deliver oxygen and Volatile Gases to patient's and to eliminate C02 during anaesthesia Consists of a circuit with low resistance, non-extendable, corrugated tubing that is non-kinking, clear and lightweight and crush resistant. The tubing must be pliable and must not retain a shape The corrugations must be shallow to minimise flow resistance and turbulence. Must have an extra limb with an anti-static Reservoir bag that gives good feedback regarding patient lung compliance and does not twist on itself. Must have a mechanism to prevent this. Connectors Malleable 22mm female connectors at the end of inspiratory and expiratory limbs – must fit securely to the machine. Spare 22mm M/M connector. Y-piece at patient's end. Good quality to ensure no leakage or easy disconnections No variation in Circuit limb length or Extra limb will be accepted Manufactured from latex- free, non-pyrogenic hypoallergenic and non-toxic Silicone/ Hydro-tubing . Reusable, autoclavable, individually packed and surgical clean. Information brochure or the outer package of the item should clearly indicate latex free | Single Smooth-bore Breathing Tubing - Single Use Paediatric 15mm Length: ± 1.8m Connector: 15mm M/F; 15mm M/M; 15mm F/F
340 | RT13-71-001 | Resuscitation and Re-breathing bags | REBREATHING BAGS REUSABLE | REBREATHING BAGS REUSABLE Purpose: Stores gas between respiratory cycles. Provides information to anaesthetist in manual ventilation regarding pulmonary compliance and airway patency. Consists of alatex-free antistatic, non-conductive reservoir bag with a tapered neck and collar Compliance: 30-60cmH20 when filled to 4 X nominal volume Must hold at least 10 X nominal volume before it will burst Must retain original volume and shape (within 10%) at deflation Must must not adhere to itself and must not twist or kink on itself –must have a mechanism to prevent this. Must clearly indicate that it is Autoclavable on the bag To comply with ISO 5362 Information brochure or the outer package of the item should clearly indicate latex free | Rebreathing reservoir bag 0.5 litre with 15mm F connector - Reusable
341 | RT13-71-002 | Resuscitation and Re-breathing bags | REBREATHING BAGS REUSABLE | REBREATHING BAGS REUSABLE Purpose: Stores gas between respiratory cycles. Provides information to anaesthetist in manual ventilation regarding pulmonary compliance and airway patency. Consists of alatex-free antistatic, non-conductive reservoir bag with a tapered neck and collar Compliance: 30-60cmH20 when filled to 4 X nominal volume Must hold at least 10 X nominal volume before it will burst Must retain original volume and shape (within 10%) at deflation Must must not adhere to itself and must not twist or kink on itself –must have a mechanism to prevent this. Must clearly indicate that it is Autoclavable on the bag To comply with ISO 5362 Information brochure or the outer package of the item should clearly indicate latex free | Rebreathing reservoir bag 1 litre with 15mm F connector - Reusable
342 | RT13-71-003 | Resuscitation and Re-breathing bags | REBREATHING BAGS REUSABLE | REBREATHING BAGS REUSABLE Purpose: Stores gas between respiratory cycles. Provides information to anaesthetist in manual ventilation regarding pulmonary compliance and airway patency. Consists of alatex-free antistatic, non-conductive reservoir bag with a tapered neck and collar Compliance: 30-60cmH20 when filled to 4 X nominal volume Must hold at least 10 X nominal volume before it will burst Must retain original volume and shape (within 10%) at deflation Must must not adhere to itself and must not twist or kink on itself –must have a mechanism to prevent this. Must clearly indicate that it is Autoclavable on the bag To comply with ISO 5362 Information brochure or the outer package of the item should clearly indicate latex free | Rebreathing reservoir bag 1 litre with 15mm F connector - Reusable
343 | RT13-72-001 | PERCUTANEOUS TRACHEOSTOMY KITS | PERCUTANEOUS TRACHEOSTOMY KITS - WITH DILATOR | PERCUTANEOUS TRACHEOSTOMY KITS - WITH DILATOR Purpose: Tracheostomy insertion at the bedside Consists of a, radiopaque, cuffed Tracheostomy tube with an inner cannula and a 15mm connector, with hypoallergenic neck tape. The set must have a short-handled scalpel blade; a 14-15G puncture needle with a catheter; a kink-resistant guidewire with a guidewire inserter; a short 14-15FR dilator, an angled tracheostomy dilator on a guiding catheter with a safety stop and a loading dilator. The guiding catheter must have clear marking lines to prevent the dilator from being inserted too deeply The KIT must include either a one-step dilator or a set of serial dilators ( Must be appropriate for the specific Tracheostomy size) Manufactured from latex-free medical grade metal and plastic Individually packed in peel pouch that is easy to open, SINGLE use, Sterile To comply with latest ISO 5366 Information brochure or the outer package of the item should clearly indicate latex free | Percutaneous Tracheostomy Kit Size 7.0
344 | RT13-72-002 | PERCUTANEOUS TRACHEOSTOMY KITS | PERCUTANEOUS TRACHEOSTOMY KITS - WITH DILATOR | PERCUTANEOUS TRACHEOSTOMY KITS - WITH DILATOR Purpose: Tracheostomy insertion at the bedside Consists of a, radiopaque, cuffed Tracheostomy tube with an inner cannula and a 15mm connector, with hypoallergenic neck tape. The set must have a short-handled scalpel blade; a 14-15G puncture needle with a catheter; a kink-resistant guidewire with a guidewire inserter; a short 14-15FR dilator, an angled tracheostomy dilator on a guiding catheter with a safety stop and a loading dilator. The guiding catheter must have clear marking lines to prevent the dilator from being inserted too deeply The KIT must include either a one-step dilator or a set of serial dilators ( Must be appropriate for the specific Tracheostomy size) Manufactured from latex-free medical grade metal and plastic Individually packed in peel pouch that is easy to open, SINGLE use, Sterile To comply with latest ISO 5366 Information brochure or the outer package of the item should clearly indicate latex free | Percutaneous Tracheostomy Kit Size 7.5
345 | RT13-72-003 | PERCUTANEOUS TRACHEOSTOMY KITS | PERCUTANEOUS TRACHEOSTOMY KITS - WITH DILATOR | PERCUTANEOUS TRACHEOSTOMY KITS - WITH DILATOR Purpose: Tracheostomy insertion at the bedside Consists of a, radiopaque, cuffed Tracheostomy tube with an inner cannula and a 15mm connector, with hypoallergenic neck tape. The set must have a short-handled scalpel blade; a 14-15G puncture needle with a catheter; a kink-resistant guidewire with a guidewire inserter; a short 14-15FR dilator, an angled tracheostomy dilator on a guiding catheter with a safety stop and a loading dilator. The guiding catheter must have clear marking lines to prevent the dilator from being inserted too deeply The KIT must include either a one-step dilator or a set of serial dilators ( Must be appropriate for the specific Tracheostomy size) Manufactured from latex-free medical grade metal and plastic Individually packed in peel pouch that is easy to open, SINGLE use, Sterile To comply with latest ISO 5366 Information brochure or the outer package of the item should clearly indicate latex free | Percutaneous Tracheostomy Kit Size 8.0
346 | RT13-72-004 | PERCUTANEOUS TRACHEOSTOMY KITS | PERCUTANEOUS TRACHEOSTOMY KITS - WITH DILATOR | PERCUTANEOUS TRACHEOSTOMY KITS - WITH DILATOR Purpose: Tracheostomy insertion at the bedside Consists of a, radiopaque, cuffed Tracheostomy tube with an inner cannula and a 15mm connector, with hypoallergenic neck tape. The set must have a short-handled scalpel blade; a 14-15G puncture needle with a catheter; a kink-resistant guidewire with a guidewire inserter; a short 14-15FR dilator, an angled tracheostomy dilator on a guiding catheter with a safety stop and a loading dilator. The guiding catheter must have clear marking lines to prevent the dilator from being inserted too deeply The KIT must include either a one-step dilator or a set of serial dilators ( Must be appropriate for the specific Tracheostomy size) Manufactured from latex-free medical grade metal and plastic Individually packed in peel pouch that is easy to open, SINGLE use, Sterile To comply with latest ISO 5366 Information brochure or the outer package of the item should clearly indicate latex free | Percutaneous Tracheostomy Kit Size 8.5
347 | RT13-72-005 | PERCUTANEOUS TRACHEOSTOMY KITS | PERCUTANEOUS TRACHEOSTOMY KITS - WITH DILATING FORCEPS | PERCUTANEOUS TRACHEOSTOMY KITS - WITH DILATING FORCEPS Purpose: Tracheostomy insertion at the bedside Consists of a, radiopaque, cuffed Tracheostomy tube with an inner cannula and a 15mm connector, with hypoallergenic neck tape. The set must have a short-handled scalpel blade; a curved dissecting forceps; a 14-15G puncture needle with a catheter; a kink-resistant guidewire with a guidewire inserter; a short 14-15FR dilator, angled dilating forceps with a specialised notch for the guidewire. The angled dilating forceps must have an ergonomic handle and must pass through the skin and tissue easily. . The forceps must lock over the guidewire catheter and allow for single step dilation. The KIT must include either a one-step dilator or a set of serial dilators ( Must be appropriate for the specific Tracheostomy size) Manufactured from latex-free medical grade metal and plastic Individually packed in peel pouch that is easy to open, SINGLE use, Sterile To comply with latest ISO 5366 Information brochure or the outer package of the item should clearly indicate latex free | Percutaneous Tracheostomy kits with dilating forceps Size 7.0
348 | RT13-72-006 | PERCUTANEOUS TRACHEOSTOMY KITS | PERCUTANEOUS TRACHEOSTOMY KITS - WITH DILATING FORCEPS | PERCUTANEOUS TRACHEOSTOMY KITS - WITH DILATING FORCEPS Purpose: Tracheostomy insertion at the bedside Consists of a, radiopaque, cuffed Tracheostomy tube with an inner cannula and a 15mm connector, with hypoallergenic neck tape. The set must have a short-handled scalpel blade; a curved dissecting forceps; a 14-15G puncture needle with a catheter; a kink-resistant guidewire with a guidewire inserter; a short 14-15FR dilator, angled dilating forceps with a specialised notch for the guidewire. The angled dilating forceps must have an ergonomic handle and must pass through the skin and tissue easily. . The forceps must lock over the guidewire catheter and allow for single step dilation. The KIT must include either a one-step dilator or a set of serial dilators ( Must be appropriate for the specific Tracheostomy size) Manufactured from latex-free medical grade metal and plastic Individually packed in peel pouch that is easy to open, SINGLE use, Sterile To comply with latest ISO 5366 Information brochure or the outer package of the item should clearly indicate latex free | Percutaneous Tracheostomy kits with dilating forceps Size 8.0
349 | RT13-72-007 | PERCUTANEOUS TRACHEOSTOMY KITS | PERCUTANEOUS TRACHEOSTOMY KITS - WITH DILATING FORCEPS | PERCUTANEOUS TRACHEOSTOMY KITS - WITH DILATING FORCEPS Purpose: Tracheostomy insertion at the bedside Consists of a, radiopaque, cuffed Tracheostomy tube with an inner cannula and a 15mm connector, with hypoallergenic neck tape. The set must have a short-handled scalpel blade; a curved dissecting forceps; a 14-15G puncture needle with a catheter; a kink-resistant guidewire with a guidewire inserter; a short 14-15FR dilator, angled dilating forceps with a specialised notch for the guidewire. The angled dilating forceps must have an ergonomic handle and must pass through the skin and tissue easily. . The forceps must lock over the guidewire catheter and allow for single step dilation. The KIT must include either a one-step dilator or a set of serial dilators ( Must be appropriate for the specific Tracheostomy size) Manufactured from latex-free medical grade metal and plastic Individually packed in peel pouch that is easy to open, SINGLE use, Sterile To comply with latest ISO 5366 Information brochure or the outer package of the item should clearly indicate latex free | Percutaneous Tracheostomy kits with dilating forceps Size 9.0