## Pricing Circular
RT13-1-2024 PRICING CIRCULAR
ITEM NUMBER | ITEM CATEGORY | ITEM SUB-CATEGORY | ITEM DESCRIPTION | ITEM SIZE DESCRIPTION | OFFER OPTION | UOM | BIDDER'S NAME | FINAL PRICE | TOTAL POINTS | BRAND NAME | NAME OF ORIGINAL PRODUCT MANUFACTURER | COUNTRY OF ORIGIN
RT13-52-005 | CLOSED SUCTION CATHETER | CLOSED SUCTION CATHETER PAEDIATRIC 48HR | CLOSED SUCTION CATHETER - PAEDIATRIC ENDOTRACHEAL Y-TYPE 48HR Purpose: Closed system for respiratory suctioning through endotracheal tubes without disconnecting or interrupting ventilation. Consists of a closed suction catheter with an atraumatic tip. The catheter must have markings along the length to indicate depth of suctioning and a radiopaque line on the catheter tip. While connected but not in suction mode the catheter tip must not extend beyond the sheath or impede airflow. The suction catheter must be enclosed within a durable, flexible, plastic sheath for storage to avoid contamination. The transparent sheath must allow visualisation of markings on the suction catheter and must connect to the endotracheal tube via a clear, firmly-fitting, plastic connector or adaptor. The sheath and connector must allow for smooth advancement of catheter. The system must seal the airway and catheter from atmospheric contamination and allow uninterrupted positive pressure ventilation throughout the suction process. An effective seal around the catheter must prevent ventilation gas entering the lumen of the sheath. The system must protect the user from contamination by secretions and suctioned material. Proximal end of suction catheter must have clear chamber, so that suctioned contents can be visualised and must have lockable suction valve with toggled capped suction tubing connector. Paediatric Catheters must have a Y- connector (without locking mechanism) with an appropriate range of diameters for each size catheter. The connector must allow variations in patient positions without causing torque, twisting, or obstruction of the tracheal tube, and must not cause tissue or skin pressure. The system must allow for irrigation via a lateral irrigation port which must be a capped with a one-way valve that connects to a standard luer-slip syringe. The irrigation port must be situated at the patient end of the system for catheters and at the suction end of the catheter for catheters.made from medically safe PVC The suction end of the system must be fitted with capped, lockable thumb valve to control suctioning. It must have a tapered connection to fit a range of suction tubing. The system must be as light in weight as possible. The system is for single patient use for a specified 48 or 72 hours. Must have coloured day sticker labels to facilitate regular changing of the catheter system. Made from medically safe PVC. Sterile and individually packed in easy to open pouch with one transparent side; For SINGLE use only The system must conform to ISO 8836:2019. Information brochure or the outer package of the item should clearly indicate latex free | Closed Suction Catheter - Paediatric Endotracheal Y-Type 48Hr 10Fr OD: 3.3mm For ETT with ID: 5.0mm Y connectors: 4.5mm, 5mm, 5.5mm | Offer 1 | Each | Access Medical | 53.06 | 93.7587 | ACCESS | GWS SURGICAL LLP | INDIA
RT13-52-013 | CLOSED SUCTION CATHETER | CLOSED SUCTION CATHETER ADULT 48HR | CLOSED SUCTION CATHETER ADULT ENDOTRACHEAL - 48HRS Purpose: Closed system for respiratory suctioning through endotracheal tubes without disconnecting or interrupting ventilation. Consists of a closed suction catheter with an atraumatic tip. The catheter must have markings along the length to indicate depth of suctioning and a radiopaque line on the catheter tip. While connected but not in suction mode the catheter tip must not extend beyond the sheath or impede airflow. The suction catheter must be enclosed within a durable, flexible, plastic sheath for storage to avoid contamination. The transparent sheath must allow visualisation of markings on the suction catheter and must connect to the endotracheal tube via a clear, firmly-fitting, plastic connector or adaptor. The sheath and connector must allow for smooth advancement of catheter. The system must seal the airway and catheter from atmospheric contamination and allow uninterrupted positive pressure ventilation throughout the suction process. An effective seal around the catheter must prevent ventilation gas entering the lumen of the sheath. The system must protect the user from contamination by secretions and suctioned material. Proximal end of suction catheter must have clear chamber, so that suctioned contents can be visualised and must have lockable suction valve with capped suction tubing connector. The suction catheter must connect onto the tube with a DOUBLE swivel connector or a double-swivel Y-connector with standard 15 mm OD taper fittings. This connection must be secure and accidental disconnection must be minimal The whole device must be as lightweight as possible. Must include a light, straight, flexible, clear, concertina catheter-mount (15x22 mm connections) tube for connection to the breathing circuit The connector must allow variations in patient positions without causing torque, twisting, or obstruction of the tracheal tube, and must not cause tissue or skin pressure. The system must allow for irrigation via a lateral irrigation port which must be a capped with a one-way valve that connects to a standard luer-slip syringe. The irrigation port must be situated at the patient end of the system for catheters and at the suction end of the catheter for catheters.made from medically safe PVC The suction end of the system must be fitted with capped, lockable thumb valve to control suctioning. It must have a tapered connection to fit a range of suction tubing. The system is for single patient use for a specified 48 hours Must have coloured day sticker labels to facilitate regular changing of the catheter system. Made from medically safe PVC. Sterile and individually packed in easy to open pouch with one transparent side; For SINGLE use only The system must conform to ISO 8836 Information brochure or the outer package of the item should clearly indicate latex free | Closed Suction Catheter - Adult Endotracheal 48Hr 10Fr OD: 3.3mm Length range: 520 - 600mm | Offer 1 | Each | Access Medical | 53.06 | 93.7587 | ACCESS | GWS SURGICAL LLP | INDIA
RT13-52-014 | CLOSED SUCTION CATHETER | CLOSED SUCTION CATHETER ADULT 48HR | CLOSED SUCTION CATHETER ADULT ENDOTRACHEAL - 48HRS Purpose: Closed system for respiratory suctioning through endotracheal tubes without disconnecting or interrupting ventilation. Consists of a closed suction catheter with an atraumatic tip. The catheter must have markings along the length to indicate depth of suctioning and a radiopaque line on the catheter tip. While connected but not in suction mode the catheter tip must not extend beyond the sheath or impede airflow. The suction catheter must be enclosed within a durable, flexible, plastic sheath for storage to avoid contamination. The transparent sheath must allow visualisation of markings on the suction catheter and must connect to the endotracheal tube via a clear, firmly-fitting, plastic connector or adaptor. The sheath and connector must allow for smooth advancement of catheter. The system must seal the airway and catheter from atmospheric contamination and allow uninterrupted positive pressure ventilation throughout the suction process. An effective seal around the catheter must prevent ventilation gas entering the lumen of the sheath. The system must protect the user from contamination by secretions and suctioned material. Proximal end of suction catheter must have clear chamber, so that suctioned contents can be visualised and must have lockable suction valve with capped suction tubing connector. The suction catheter must connect onto the tube with a DOUBLE swivel connector or a double-swivel Y-connector with standard 15 mm OD taper fittings. This connection must be secure and accidental disconnection must be minimal The whole device must be as lightweight as possible. Must include a light, straight, flexible, clear, concertina catheter-mount (15x22 mm connections) tube for connection to the breathing circuit The connector must allow variations in patient positions without causing torque, twisting, or obstruction of the tracheal tube, and must not cause tissue or skin pressure. The system must allow for irrigation via a lateral irrigation port which must be a capped with a one-way valve that connects to a standard luer-slip syringe. The irrigation port must be situated at the patient end of the system for catheters and at the suction end of the catheter for catheters.made from medically safe PVC The suction end of the system must be fitted with capped, lockable thumb valve to control suctioning. It must have a tapered connection to fit a range of suction tubing. The system is for single patient use for a specified 48 hours Must have coloured day sticker labels to facilitate regular changing of the catheter system. Made from medically safe PVC. Sterile and individually packed in easy to open pouch with one transparent side; For SINGLE use only The system must conform to ISO 8836 Information brochure or the outer package of the item should clearly indicate latex free | Closed Suction Catheter - Adult Endotracheal 48Hr 12Fr OD: 4.0mm Length range: 520 - 600mm | Offer 1 | Each | Access Medical | 53.06 | 93.7587 | ACCESS | GWS SURGICAL LLP | INDIA
RT13-52-015 | CLOSED SUCTION CATHETER | CLOSED SUCTION CATHETER ADULT 48HR | CLOSED SUCTION CATHETER ADULT ENDOTRACHEAL - 48HRS Purpose: Closed system for respiratory suctioning through endotracheal tubes without disconnecting or interrupting ventilation. Consists of a closed suction catheter with an atraumatic tip. The catheter must have markings along the length to indicate depth of suctioning and a radiopaque line on the catheter tip. While connected but not in suction mode the catheter tip must not extend beyond the sheath or impede airflow. The suction catheter must be enclosed within a durable, flexible, plastic sheath for storage to avoid contamination. The transparent sheath must allow visualisation of markings on the suction catheter and must connect to the endotracheal tube via a clear, firmly-fitting, plastic connector or adaptor. The sheath and connector must allow for smooth advancement of catheter. The system must seal the airway and catheter from atmospheric contamination and allow uninterrupted positive pressure ventilation throughout the suction process. An effective seal around the catheter must prevent ventilation gas entering the lumen of the sheath. The system must protect the user from contamination by secretions and suctioned material. Proximal end of suction catheter must have clear chamber, so that suctioned contents can be visualised and must have lockable suction valve with capped suction tubing connector. The suction catheter must connect onto the tube with a DOUBLE swivel connector or a double-swivel Y-connector with standard 15 mm OD taper fittings. This connection must be secure and accidental disconnection must be minimal The whole device must be as lightweight as possible. Must include a light, straight, flexible, clear, concertina catheter-mount (15x22 mm connections) tube for connection to the breathing circuit The connector must allow variations in patient positions without causing torque, twisting, or obstruction of the tracheal tube, and must not cause tissue or skin pressure. The system must allow for irrigation via a lateral irrigation port which must be a capped with a one-way valve that connects to a standard luer-slip syringe. The irrigation port must be situated at the patient end of the system for catheters and at the suction end of the catheter for catheters.made from medically safe PVC The suction end of the system must be fitted with capped, lockable thumb valve to control suctioning. It must have a tapered connection to fit a range of suction tubing. The system is for single patient use for a specified 48 hours Must have coloured day sticker labels to facilitate regular changing of the catheter system. Made from medically safe PVC. Sterile and individually packed in easy to open pouch with one transparent side; For SINGLE use only The system must conform to ISO 8836 Information brochure or the outer package of the item should clearly indicate latex free | Closed Suction Catheter - Adult Endotracheal 48Hr 14Fr OD: 4.7mm Length range: 520 - 600mm | Offer 1 | Each | Access Medical | 56.78 | 93.7587 | ACCESS | GWS SURGICAL LLP | INDIA
RT13-52-016 | CLOSED SUCTION CATHETER | CLOSED SUCTION CATHETER ADULT 48HR | CLOSED SUCTION CATHETER ADULT ENDOTRACHEAL - 48HRS Purpose: Closed system for respiratory suctioning through endotracheal tubes without disconnecting or interrupting ventilation. Consists of a closed suction catheter with an atraumatic tip. The catheter must have markings along the length to indicate depth of suctioning and a radiopaque line on the catheter tip. While connected but not in suction mode the catheter tip must not extend beyond the sheath or impede airflow. The suction catheter must be enclosed within a durable, flexible, plastic sheath for storage to avoid contamination. The transparent sheath must allow visualisation of markings on the suction catheter and must connect to the endotracheal tube via a clear, firmly-fitting, plastic connector or adaptor. The sheath and connector must allow for smooth advancement of catheter. The system must seal the airway and catheter from atmospheric contamination and allow uninterrupted positive pressure ventilation throughout the suction process. An effective seal around the catheter must prevent ventilation gas entering the lumen of the sheath. The system must protect the user from contamination by secretions and suctioned material. Proximal end of suction catheter must have clear chamber, so that suctioned contents can be visualised and must have lockable suction valve with capped suction tubing connector. The suction catheter must connect onto the tube with a DOUBLE swivel connector or a double-swivel Y-connector with standard 15 mm OD taper fittings. This connection must be secure and accidental disconnection must be minimal The whole device must be as lightweight as possible. Must include a light, straight, flexible, clear, concertina catheter-mount (15x22 mm connections) tube for connection to the breathing circuit The connector must allow variations in patient positions without causing torque, twisting, or obstruction of the tracheal tube, and must not cause tissue or skin pressure. The system must allow for irrigation via a lateral irrigation port which must be a capped with a one-way valve that connects to a standard luer-slip syringe. The irrigation port must be situated at the patient end of the system for catheters and at the suction end of the catheter for catheters.made from medically safe PVC The suction end of the system must be fitted with capped, lockable thumb valve to control suctioning. It must have a tapered connection to fit a range of suction tubing. The system is for single patient use for a specified 48 hours Must have coloured day sticker labels to facilitate regular changing of the catheter system. Made from medically safe PVC. Sterile and individually packed in easy to open pouch with one transparent side; For SINGLE use only The system must conform to ISO 8836 Information brochure or the outer package of the item should clearly indicate latex free | Closed Suction Catheter - Adult Endotracheal 48Hr 16Fr OD: 5.3mm Length range: 520 - 600mm | Offer 1 | Each | Access Medical | 53.06 | 93.7587 | ACCESS | GWS SURGICAL LLP | INDIA
RT13-52-031 | CLOSED SUCTION CATHETER | CLOSED SUCTION CATHETER ADULT TRACHEOSTOMY | CLOSED SUCTION CATHETER ADULT TRACHEOSTOMY - 48HRS Purpose: Closed system for respiratory suctioning through tracheostomy tubes without disconnecting or interrupting ventilation. Consists of a closed suction catheter with an atraumatic tip. The catheter must have markings along the length to indicate depth of suctioning and a radiopaque line on the catheter tip. While connected but not in suction mode the catheter tip must not extend beyond the sheath or impede airflow. The suction catheter must be enclosed within a durable, flexible, plastic sheath for storage to avoID:contamination. The transparent sheath must allow visualisation of markings on the suction catheter and must connect to the tracheal tube via a clear, firmly-fitting, plastic connector or adaptor. The sheath and connector must allow for smooth advancement of catheter. The system must seal the airway and catheter from atmospheric contamination and allow uninterrupted positive pressure ventilation throughout the suction process. An effective seal around the catheter must prevent ventilation gas entering the lumen of the sheath. The system must protect the user from contamination by secretions and suctioned material. Proximal end of suction catheter must have clear chamber, so that suctioned contents can be visualised and must have lockable suction valve with capped suction tubing connector. The suction catheter must connect onto the tube with a DOUBLE swivel connector or a double-swivel Y-connector with standard 15 mm OD taper fittings. This connection must be secure and accidental disconnection must be minimal The whole device must be as lightweight as possible. Must include a light, straight, flexible, clear, concertina catheter-mount (15x22 mm connections) tube for connection to the breathing circuit The connector must allow variations in patient positions without causing torque, twisting, or obstruction of the tracheal tube, and must not cause tissue or skin pressure. The system must allow for irrigation via a lateral irrigation port which must be a capped with a one-way valve that connects to a standard luer-slip syringe. The irrigation port must be situated at the patient end of the system for catheters and at the suction end of the catheter for catheters.made from medically safe PVC The suction end of the system must be fitted with capped, lockable thumb valve to control suctioning. It must have a tapered connection to fit a range of suction tubing. The system is for single patient use for a specified 48 hours. Must have coloured day sticker labels to facilitate regular changing of the catheter system. Made from medically safe PVC. Sterile and individually packed in easy to open pouch with one transparent side; For SINGLE use only The system must conform to ISO 8836 Information brochure or the outer package of the item should clearly indicate latex free | Closed Suction Catheter - Adult Tracheostomy 48Hr 10Fr OD: 3.3mm Length range: 300 - 400mm | Offer 1 | Each | Access Medical | 53.06 | 93.7587 | ACCESS | GWS SURGICAL LLP | INDIA
RT13-52-032 | CLOSED SUCTION CATHETER | CLOSED SUCTION CATHETER ADULT TRACHEOSTOMY | CLOSED SUCTION CATHETER ADULT TRACHEOSTOMY - 48HRS Purpose: Closed system for respiratory suctioning through tracheostomy tubes without disconnecting or interrupting ventilation. Consists of a closed suction catheter with an atraumatic tip. The catheter must have markings along the length to indicate depth of suctioning and a radiopaque line on the catheter tip. While connected but not in suction mode the catheter tip must not extend beyond the sheath or impede airflow. The suction catheter must be enclosed within a durable, flexible, plastic sheath for storage to avoID:contamination. The transparent sheath must allow visualisation of markings on the suction catheter and must connect to the tracheal tube via a clear, firmly-fitting, plastic connector or adaptor. The sheath and connector must allow for smooth advancement of catheter. The system must seal the airway and catheter from atmospheric contamination and allow uninterrupted positive pressure ventilation throughout the suction process. An effective seal around the catheter must prevent ventilation gas entering the lumen of the sheath. The system must protect the user from contamination by secretions and suctioned material. Proximal end of suction catheter must have clear chamber, so that suctioned contents can be visualised and must have lockable suction valve with capped suction tubing connector. The suction catheter must connect onto the tube with a DOUBLE swivel connector or a double-swivel Y-connector with standard 15 mm OD taper fittings. This connection must be secure and accidental disconnection must be minimal The whole device must be as lightweight as possible. Must include a light, straight, flexible, clear, concertina catheter-mount (15x22 mm connections) tube for connection to the breathing circuit The connector must allow variations in patient positions without causing torque, twisting, or obstruction of the tracheal tube, and must not cause tissue or skin pressure. The system must allow for irrigation via a lateral irrigation port which must be a capped with a one-way valve that connects to a standard luer-slip syringe. The irrigation port must be situated at the patient end of the system for catheters and at the suction end of the catheter for catheters.made from medically safe PVC The suction end of the system must be fitted with capped, lockable thumb valve to control suctioning. It must have a tapered connection to fit a range of suction tubing. The system is for single patient use for a specified 48 hours. Must have coloured day sticker labels to facilitate regular changing of the catheter system. Made from medically safe PVC. Sterile and individually packed in easy to open pouch with one transparent side; For SINGLE use only The system must conform to ISO 8836 Information brochure or the outer package of the item should clearly indicate latex free | Closed Suction Catheter - Adult Tracheostomy 48Hr 12Fr OD: 4.0mm Length range: 300 - 400mm | Offer 1 | Each | Access Medical | 53.06 | 93.7587 | ACCESS | GWS SURGICAL LLP | INDIA
RT13-52-033 | CLOSED SUCTION CATHETER | CLOSED SUCTION CATHETER ADULT TRACHEOSTOMY | CLOSED SUCTION CATHETER ADULT TRACHEOSTOMY - 48HRS Purpose: Closed system for respiratory suctioning through tracheostomy tubes without disconnecting or interrupting ventilation. Consists of a closed suction catheter with an atraumatic tip. The catheter must have markings along the length to indicate depth of suctioning and a radiopaque line on the catheter tip. While connected but not in suction mode the catheter tip must not extend beyond the sheath or impede airflow. The suction catheter must be enclosed within a durable, flexible, plastic sheath for storage to avoID:contamination. The transparent sheath must allow visualisation of markings on the suction catheter and must connect to the tracheal tube via a clear, firmly-fitting, plastic connector or adaptor. The sheath and connector must allow for smooth advancement of catheter. The system must seal the airway and catheter from atmospheric contamination and allow uninterrupted positive pressure ventilation throughout the suction process. An effective seal around the catheter must prevent ventilation gas entering the lumen of the sheath. The system must protect the user from contamination by secretions and suctioned material. Proximal end of suction catheter must have clear chamber, so that suctioned contents can be visualised and must have lockable suction valve with capped suction tubing connector. The suction catheter must connect onto the tube with a DOUBLE swivel connector or a double-swivel Y-connector with standard 15 mm OD taper fittings. This connection must be secure and accidental disconnection must be minimal The whole device must be as lightweight as possible. Must include a light, straight, flexible, clear, concertina catheter-mount (15x22 mm connections) tube for connection to the breathing circuit The connector must allow variations in patient positions without causing torque, twisting, or obstruction of the tracheal tube, and must not cause tissue or skin pressure. The system must allow for irrigation via a lateral irrigation port which must be a capped with a one-way valve that connects to a standard luer-slip syringe. The irrigation port must be situated at the patient end of the system for catheters and at the suction end of the catheter for catheters.made from medically safe PVC The suction end of the system must be fitted with capped, lockable thumb valve to control suctioning. It must have a tapered connection to fit a range of suction tubing. The system is for single patient use for a specified 48 hours. Must have coloured day sticker labels to facilitate regular changing of the catheter system. Made from medically safe PVC. Sterile and individually packed in easy to open pouch with one transparent side; For SINGLE use only The system must conform to ISO 8836 Information brochure or the outer package of the item should clearly indicate latex free | Closed Suction Catheter - Adult Tracheostomy 48Hr 14Fr OD: 4.7mm Length range: 300 - 400mm | Offer 1 | Each | Access Medical | 53.06 | 93.7587 | ACCESS | GWS SURGICAL LLP | INDIA
RT13-52-034 | CLOSED SUCTION CATHETER | CLOSED SUCTION CATHETER ADULT TRACHEOSTOMY | CLOSED SUCTION CATHETER ADULT TRACHEOSTOMY - 48HRS Purpose: Closed system for respiratory suctioning through tracheostomy tubes without disconnecting or interrupting ventilation. Consists of a closed suction catheter with an atraumatic tip. The catheter must have markings along the length to indicate depth of suctioning and a radiopaque line on the catheter tip. While connected but not in suction mode the catheter tip must not extend beyond the sheath or impede airflow. The suction catheter must be enclosed within a durable, flexible, plastic sheath for storage to avoID:contamination. The transparent sheath must allow visualisation of markings on the suction catheter and must connect to the tracheal tube via a clear, firmly-fitting, plastic connector or adaptor. The sheath and connector must allow for smooth advancement of catheter. The system must seal the airway and catheter from atmospheric contamination and allow uninterrupted positive pressure ventilation throughout the suction process. An effective seal around the catheter must prevent ventilation gas entering the lumen of the sheath. The system must protect the user from contamination by secretions and suctioned material. Proximal end of suction catheter must have clear chamber, so that suctioned contents can be visualised and must have lockable suction valve with capped suction tubing connector. The suction catheter must connect onto the tube with a DOUBLE swivel connector or a double-swivel Y-connector with standard 15 mm OD taper fittings. This connection must be secure and accidental disconnection must be minimal The whole device must be as lightweight as possible. Must include a light, straight, flexible, clear, concertina catheter-mount (15x22 mm connections) tube for connection to the breathing circuit The connector must allow variations in patient positions without causing torque, twisting, or obstruction of the tracheal tube, and must not cause tissue or skin pressure. The system must allow for irrigation via a lateral irrigation port which must be a capped with a one-way valve that connects to a standard luer-slip syringe. The irrigation port must be situated at the patient end of the system for catheters and at the suction end of the catheter for catheters.made from medically safe PVC The suction end of the system must be fitted with capped, lockable thumb valve to control suctioning. It must have a tapered connection to fit a range of suction tubing. The system is for single patient use for a specified 48 hours. Must have coloured day sticker labels to facilitate regular changing of the catheter system. Made from medically safe PVC. Sterile and individually packed in easy to open pouch with one transparent side; For SINGLE use only The system must conform to ISO 8836 Information brochure or the outer package of the item should clearly indicate latex free | Closed Suction Catheter - Adult Tracheostomy 48Hr 16Fr OD: 5.3mm Length range: 300 - 400mm | Offer 1 | Each | Access Medical | 53.06 | 93.7587 | ACCESS | GWS SURGICAL LLP | INDIA
RT13-57-004 | BREATHING CIRCUITS | CORRUGATED BREATHING CIRCUITS WITH EXTRA LIMB AND BAG | CORRUGATED BREATHING CIRCUITS WITH EXTRA LIMB AND BAG - SINGLE USE Purpose: Used to deliver oxygen and Volatile Gases to patient's and to eliminate C02 during anaesthesia Transparent circuit with low resistance, non-extendable, corrugated tubing that is non-kinking, lightweight and crush resistant. The tubing must be pliable and must not retain a shape. Uninterrupted tubing (not joined). Shallow corrugations to minimise flow resistance and turbulence. Must have an extra limb with an anti-static reservoir bag that gives good feedback regarding patient lung compliance. The reservoir bag must not twist on itself and must have a mechanism to prevent this. Connectors Malleable 22mm female connectors bonded to the ends of inspiratory and expiratory limbs – must fit securely to the machine. Spare 22mm M/M connector. Y-piece at patient's end. Separate elbow-piece with a toggled cap on a luer lock port. Good quality to ensure no leakage or easy disconnections Latex-free, non-pyrogenic, hypoallergenic, non-toxic and non-conductive medical grade plastic No variation in Circuit limb length or Extra limb will be accepted Clinically clean and individually packed, For SINGLE use only To comply with ISO 5367 Information brochure or the outer package of the item should clearly indicate latex free | Corrugated breathing circuit with extra limb and breathing bag - Single Use - Paediatric > 3 yrs 15mm Circuit: 1.6m; 15mm Extra limb: 1.8m Reservoir bag: 1l | Offer 1 | Each | Bestech Medical Suppliers | 139.83 | 46.6562 | Plasti-Med | Plasti-Med | Turkey
RT13-57-017 | BREATHING CIRCUITS | CORRUGATED BREATHING CIRCUITS WITH WATER TRAPS AND EXTRA LIMB AND BAG-SINGLE USE | SINGLE CORRUGATED BREATHING TUBING - SINGLE USE Purpose: Used as extension or extra limb in breathing circuits Transparent circuit with low resistance, non-extendable, corrugated tubing that is non-kinking, lightweight and crush resistant. The tubing must be pliable and must not retain a shape Shallow corrugations to minimise flow resistance and turbulence. Connectors Malleable 22mm female connectors bonded to the end of the tubing Spare connectors to be provided Latex-free, non-pyrogenic, hypoallergenic, non-toxic and non-conductive medical grade plastic No variation in Circuit limb length or Extra limb will be accepted Clinically clean and individually packed, For SINGLE use only To comply with ISO 5367 Information brochure or the outer package of the item should clearly indicate latex free | Corrugated breathing circuit with water traps and extra limb with breathing bag - Single Use Adult Long 22mm Circuit: 2.5m; 22mm Extra limb: 1.5m OD: 22mm Reservoir bag: 2l | Offer 1 | Each | Bestech Medical Suppliers | 19.95 | 94 | Plasti-Med | Plasti-Med | Turkey
RT13-60-003 | BREATHING CIRCUITS | SMOOTH BORE BREATHING CIRCUITS WITH EXTRA LIMB AND BREATHING BAG - SINGLE USE | SMOOTH BORE BREATHING CIRCUITS WITH EXTRA LIMB AND BREATHING BAG - SINGLE USE Purpose: Used to deliver oxygen and Volatile Gases to patient's and to eliminate C02 during anaesthesia Transparent circuit with low resistance, non-extendable, smooth-bore tubing that is non-kinking, lightweight and crush resistant. The tubing must be pliable and must not retain a shape Spiral outer casing that supports the inner lumen to maintain tube integrity. Must have an extra limb with an anti-static reservoir bag that gives good feedback regarding patient lung compliance. The reservoir bag must not twist on itself and must have a mechanism to prevent this. Connectors Malleable 22mm female connectors bonded to the ends of inspiratory and expiratory limbs – must fit securely to the machine. Spare 22mm M/M connector. Y-piece at patient's end. Separate elbow-piece with capped luer lock port. Good quality to ensure no leakage or easy disconnections Latex-free, non-pyrogenic, hypoallergenic, non-toxic and non-conductive medical grade plastic No variation in Circuit limb length or Extra limb will be accepted Clinically clean and individually packed, For SINGLE use only To comply with ISO 5367 Information brochure or the outer package of the item should clearly indicate latex free | Smooth bore breathing circuit with extra limb and breathing bag - Single Use Adult 22mm Circuit: 1.8m; Extra limb: 1.2m OD: Reservoir bag: 2l | Offer 1 | Each | Bestech Medical Suppliers | 237.17 | 23.2153 | Plasti-Med | Plasti-Med | Turkey
RT13-60-005 | BREATHING CIRCUITS | SMOOTH BORE BREATHING CIRCUITS WITH EXTRA LIMB AND BREATHING BAG - SINGLE USE | SMOOTH BORE BREATHING CIRCUITS WITH EXTRA LIMB AND BREATHING BAG - SINGLE USE Purpose: Used to deliver oxygen and Volatile Gases to patient's and to eliminate C02 during anaesthesia Transparent circuit with low resistance, non-extendable, smooth-bore tubing that is non-kinking, lightweight and crush resistant. The tubing must be pliable and must not retain a shape Spiral outer casing that supports the inner lumen to maintain tube integrity. Must have an extra limb with an anti-static reservoir bag that gives good feedback regarding patient lung compliance. The reservoir bag must not twist on itself and must have a mechanism to prevent this. Connectors Malleable 22mm female connectors bonded to the ends of inspiratory and expiratory limbs – must fit securely to the machine. Spare 22mm M/M connector. Y-piece at patient's end. Separate elbow-piece with capped luer lock port. Good quality to ensure no leakage or easy disconnections Latex-free, non-pyrogenic, hypoallergenic, non-toxic and non-conductive medical grade plastic No variation in Circuit limb length or Extra limb will be accepted Clinically clean and individually packed, For SINGLE use only To comply with ISO 5367 Information brochure or the outer package of the item should clearly indicate latex free | Smooth bore breathing circuit with extra limb and breathing bag - Single Use Adult 22mm Circuit: 1.8m; Extra limb: 1.8m OD: Reservoir bag: 2l | Offer 1 | Each | Bestech Medical Suppliers | 237.17 | 34.1634 | Plasti-Med | Plasti-Med | Turkey
RT13-21-010 | ENDOTRACHEAL TUBE | TAPER-CUFFED ORAL/NASAL ENDOTRACHEAL TUBE WITH EVACUATION LUMEN AT THE SUBGLOTTIC SPACE | TAPER-CUFFED ORAL/NASAL ENDOTRACHEAL TUBE WITH EVACUATION LUMEN AT THE SUBGLOTTIC SPACE Purpose: endotracheal tube with a tapered cuff for long stay oral/nasal intubation aimed to reduce micro-aspiration Consists of a clear thermo-sensitive PVC tube that is kink resistant, atraumatic to airways and have a full Magill curvature. Must have a radiopaque line and must be marked with numbered depth markings in 1cm intervals, tube size, length, ID, OD, trade name. The markings on tube must be clear and must not rub off. Must have a precision band mark just above the cuff to aid with depth of intubation and guide proper placement. Preference will be given to tubes with a second circumferential band (not more than 2cm from each other) to guide proper placement. Proximal end - secure 15mm OD circuit connector made from medical grade plastic Distal end - Atraumatic left-facing bevelled tip with Murphy eye on the opposite side Must have a high-volume-low-pressure polyurethane taper-shaped cuff with pilot cuff. Cuff thickness ± 50 microns The integrated evacuation lumen above the tapered cuff should be on the posterior aspect of the tube for the suction of subglottic secretions. The suction lumen must have a toggled capped The cuff must inflate and deflate easily and when deflated must conform closely to tube to minimise the risk of trauma during intubation and extubation. Cuff must be barrel shaped to facilitate optimal pressure dispersion against the tracheal wall The cuff must be resistant to Nitrous Oxide effects The pilot cuff must have a one-way valve and must indicate the degree of inflation accurately. All components must be latex-free sterile and individually packed in peel pouch that is easy to open - For SINGLE use only To comply with the latest ISO 5361:2016 Information brochure or the outer package of the item should clearly indicate latex free | Taper-Cuffed Oral/Nasal endotracheal tube with Evacuation Lumen: ID: 6.0mm OD: 9.0mm | Offer 1 | Each | Bioclin Solutions | 43.24 | 94 | Bonreemed | Bonree Medical Co.Ltd Medical Co.Ltd | China
RT13-40-008 | Intubation Devices | INTUBATION STYLET | INTUBATION STYLET with metal core Purpose: To pre-form an endotracheal tube to a desired shape to aid intubation Consists of a malleable aluminium core coated and tightly bonded with high density polyurethane/ PVC / medical grade plastic/ polyethylene Must have a soft atraumatic tip Must have a U-bend proximally or other mechanism at proximal end to limit depth Latex-free clinically clean and individually packed in peel pouch that is easy to open - For SINGLE use only Information brochure or the outer package of the item should clearly indicate latex free | Intubation Stylet, Extra-Length Adult OD: 5.0mm Length: 365mm To fit ETT's > 7.5mm | Offer 1 | Each | Bioclin Solutions | 10.35 | 94 | Bonreemed | Bonree Medical Co.Ltd Medical Co.Ltd | China
RT13-02-035 | Oxygen Mask | VENTURI OXYGEN MASK - MRI COMPATIBLE | VENTURI OXYGEN MASK - MRI COMPATIBLE 24% - Adult, Colour coded venturi: Blue Purpose: Delivery of oxygen and entrained room air to a spontaneously breathing patient. Suitable for the MRI Suite Modular design correctly sized and comfortable anatomically shaped preformed mask shaped to fit over the patient’s nose and mouth with under the chin cover for a better seal Atraumatic to the nose. 2 exhalation holes in the side of the mask, a secure retaining strap. All components to be MRI compatible. Must have clear kink resistant, crush proof delivery tubing > 2 m Colour coded Venturi device with the Oxygen concentration in % and required flow rate must be moulded or clearly indicated on the device The components must be securely attached to each other and not separate accidentally in normal use. Manufactured from latex-free medical grade plastic - must comply with BS EN 13544-3:2001 standards Individually packed, easy to open without tubing getting tangled, clinically clean - SINGLE use Information brochure or the outer package of the item should clearly indicate latex free | Venturi Oxygen Mask MRI COMPATIBLE 24% - Adult, Colour coded venturi: Blue | Offer 1 | Each | Biocor Hospital Supplies Pty Ltd | 28.75 | -35.5882 | Intersurgical | Intersurgical | Lithuania
RT13-02-036 | Oxygen Mask | VENTURI OXYGEN MASK - MRI COMPATIBLE | VENTURI OXYGEN MASK - MRI COMPATIBLE 28% - Adult, Colour coded venturi: White Purpose: Delivery of oxygen and entrained room air to a spontaneously breathing patient. Suitable for the MRI Suite Modular design correctly sized and comfortable anatomically shaped preformed mask shaped to fit over the patient’s nose and mouth with under the chin cover for a better seal Atraumatic to the nose. 2 exhalation holes in the side of the mask, a secure retaining strap. All components to be MRI compatible. Must have clear kink resistant, crush proof delivery tubing > 2 m Colour coded Venturi device with the Oxygen concentration in % and required flow rate must be moulded or clearly indicated on the device The components must be securely attached to each other and not separate accidentally in normal use. Manufactured from latex-free medical grade plastic - must comply with BS EN 13544-3:2001 standards Individually packed, easy to open without tubing getting tangled, clinically clean - SINGLE use Information brochure or the outer package of the item should clearly indicate latex free | Venturi Oxygen Mask MRI COMPATIBLE 28% - Adult, Colour coded venturi: White | Offer 1 | Each | Biocor Hospital Supplies pty Ltd | 28.75 | 95 | Intersurgical | Intersurgical | Lithuania
RT13-02-037 | Oxygen Mask | VENTURI OXYGEN MASK - MRI COMPATIBLE | VENTURI OXYGEN MASK - MRI COMPATIBLE 40% - Adult, Colour coded venturi: Red Purpose: Delivery of oxygen and entrained room air to a spontaneously breathing patient. Suitable for the MRI Suite Modular design correctly sized and comfortable anatomically shaped preformed mask shaped to fit over the patient’s nose and mouth with under the chin cover for a better seal Atraumatic to the nose. 2 exhalation holes in the side of the mask, a secure retaining strap. All components to be MRI compatible. Must have clear kink resistant, crush proof delivery tubing > 2 m Colour coded Venturi device with the Oxygen concentration in % and required flow rate must be moulded or clearly indicated on the device The components must be securely attached to each other and not separate accidentally in normal use. Manufactured from latex-free medical grade plastic - must comply with BS EN 13544-3:2001 standards Individually packed, easy to open without tubing getting tangled, clinically clean - SINGLE use Information brochure or the outer package of the item should clearly indicate latex free | Venturi Oxygen Mask MRI COMPATIBLE 40% - Adult, Colour coded venturi: Red | Offer 1 | Each | Biocor Hospital Supplies pty Ltd | 28.75 | -90.266 | Intersurgical | Intersurgical | Lithuania
RT13-02-038 | Oxygen Mask | VENTURI OXYGEN MASK - MRI COMPATIBLE | VENTURI OXYGEN MASK - MRI COMPATIBLE 60% - Adult ,Colour coded venturi: Green Purpose: Delivery of oxygen and entrained room air to a spontaneously breathing patient. Suitable for the MRI Suite Modular design correctly sized and comfortable anatomically shaped preformed mask shaped to fit over the patient’s nose and mouth with under the chin cover for a better seal Atraumatic to the nose. 2 exhalation holes in the side of the mask, a secure retaining strap. All components to be MRI compatible. Must have clear kink resistant, crush proof delivery tubing > 2 m Colour coded Venturi device with the Oxygen concentration in % and required flow rate must be moulded or clearly indicated on the device The components must be securely attached to each other and not separate accidentally in normal use. Manufactured from latex-free medical grade plastic - must comply with BS EN 13544-3:2001 standards Individually packed, easy to open without tubing getting tangled, clinically clean - SINGLE use Information brochure or the outer package of the item should clearly indicate latex free | Venturi Oxygen Mask MRI COMPATIBLE 60% - Adult ,Colour coded venturi: Green | Offer 1 | Each | Biocor Hospital Supplies pty Ltd | 28.75 | 95 | Intersurgical | Intersurgical | Lithuania
RT13-11-004 | Resuscitation and Re-breathing bags | resuscitation bag | BAG VALVE MASK RESUSCITATION BAG WITH PEEP VALVE - SINGLE USE ONLY Purpose: Used to provide positive pressure ventilation to a patient who is not breathing or who is breathing inadequately Ventilation bag: Clear, robust silicone. Self-inflating with a fast recoil time, and thin-walled bag to allow lung compliance assessment. Must have an ergonomic grip, Oxygen reservoir bag: PVC One-way Intake / Reservoir valve between ventilation bag and reservoir bag that only allows one way flow into the ventilation bag Oxygen from reservoir bag and room air must enter the ventilation bag if fresh gas flow is inadequate. Secure connector for oxygen tubing. Oxygen tubing: >1 m Patient valves between ventilation bag and mask. Comprising of a lip valve (duck-billed valve), expiratory valve to prevent entrainment of room air, pressure release valve – 35cmH20 (must be printed on valve) Standard connector 15mm ID, 22mm OD. Must have a PEEP valve Masks with comfortable patient fit and user interface. Malleable, non-collapsible with smooth surface or ridges. 1 piece modular face mask Able to create a tight seal between the mask and patient's’ face and must be transparent. There must be a secure fit between mask connector and breathing circuit All connections between components must be easily assembled and disassembled with tight fit and no leak. Must not disassemble during normal use Manufactured from latex-free, clear silicone, PVC, and medical grade plastic Information brochure or the outer package of the item should clearly indicate latex free | Bag valve mask resuscitation bag with PEEP valve. Infant Patient Ventilation bag: Volume: 240ml , Oxygen reservoir bag: Volume: 600ml Peep valve adjustable to ± 1.5-10 cmH2O Must be supplied with 2 round masks – size 00 and 0/1 | Offer 1 | Each | Biocor Hospital Supplies pty Ltd | 282.79 | 95 | Intersurgical | Intersurgical | Lithuania
RT13-11-005 | Resuscitation and Re-breathing bags | resuscitation bag | BAG VALVE MASK RESUSCITATION BAG WITH PEEP VALVE - SINGLE USE ONLY Purpose: Used to provide positive pressure ventilation to a patient who is not breathing or who is breathing inadequately Ventilation bag: Clear, robust silicone. Self-inflating with a fast recoil time, and thin-walled bag to allow lung compliance assessment. Must have an ergonomic grip, Oxygen reservoir bag: PVC One-way Intake / Reservoir valve between ventilation bag and reservoir bag that only allows one way flow into the ventilation bag Oxygen from reservoir bag and room air must enter the ventilation bag if fresh gas flow is inadequate. Secure connector for oxygen tubing. Oxygen tubing: >1 m Patient valves between ventilation bag and mask. Comprising of a lip valve (duck-billed valve), expiratory valve to prevent entrainment of room air, pressure release valve – 35cmH20 (must be printed on valve) Standard connector 15mm ID, 22mm OD. Must have a PEEP valve Masks with comfortable patient fit and user interface. Malleable, non-collapsible with smooth surface or ridges. 1 piece modular face mask Able to create a tight seal between the mask and patient's’ face and must be transparent. There must be a secure fit between mask connector and breathing circuit All connections between components must be easily assembled and disassembled with tight fit and no leak. Must not disassemble during normal use Manufactured from latex-free, clear silicone, PVC, and medical grade plastic Information brochure or the outer package of the item should clearly indicate latex free | Bag valve mask resuscitation bag with PEEP valve. Paediatric Patient Ventilation bag: Volume: 500ml Oxygen reservoir bag: Volume: 1600ml Peep valve adjustable to ± 1.5-10 cmH2O Must be supplied with 2 round masks – size 0/1 and 2 | Offer 1 | Each | Biocor Hospital Supplies pty Ltd | 282.79 | 95 | Intersurgical | Intersurgical | Lithuania
RT13-51-009 | CATHETER MOUNT CONNECTOR | CATHETER MOUNT WITH DOUBLE SWIVEL ELBOW CONNECTOR | CATHETER MOUNT WITH DOUBLE SWIVEL ELBOW CONNECTOR, BRONCHOSCOPY/ASPIRATION PORT AND TOGGLED CAP Purpose: Connector to join an endotracheal or tracheostomy tube to a breathing circuit with a bronchoscopy/aspiration port. Consists of flexible, low-flow smooth-bore tubing with an external coil to prevent kinks. Double swivel Elbow connector with a Bronchoscopy/Aspiration Port. Must have a diaphragm which hugs fibre-optic bronchoscope. All components latex-free, sterile and individually packed in peel pouch that is easy to open - For SINGLE use only Information brochure or the outer package of the item should clearly indicate latex free | Catheter mount with DOUBLE swivel elbow connector, Bronchoscopy/Aspiration Port and toggled cap - Adult Length: 120-150mm Connectors: 22mm Male/ 15mm Female to patient and 22mm Female to machine | Offer 1 | Each | Biocor Hospital Supplies pty Ltd | 27.26 | 45.523 | Intersurgical | Intersurgical | Lithuania
RT13-59-001 | BREATHING CIRCUITS | MAPLESON F PAEDIATRIC BREATHING CIRCUIT | MAPLESON F PAEDIATRIC BREATHING CIRCUIT - (Modified Ayre’s T piece) for children < 20-25 kg Purpose: Used to deliver oxygen and volatile gases to patient's and to eliminate C02 during anaesthesia Transparent circuit with low resistance, non-extendable, corrugated tubing that is non-kinking, lightweight and crush resistant. The tubing must be pliable and must not retain a shape with shallow corrugations to minimise flow resistance and turbulence. Tubing must be connected to a straight T-Piece with inspiratory port, expiratory port and patient connection port 22M/15F. Inspiratory tubing: 10mm ± 1.8 - 2.1m Bonded 15mm male connector at the end of the inspiratory limb– must fit securely to the machine. . Expiratory tubing 15mm diameter 0.5m with a 0.5l open-ended reservoir bag to allow gas to escape via the tail of the bag The reservoir bag must give good feedback regarding patient lung compliance and must not twist on itself and must have a mechanism to prevent this. Connectors: Elbow-piece with toggled cap on a luer lock port, capped connector at patient’s end. Good quality to ensure no leakage or easy disconnections Manufactured from transparent non-conductive medical grade plastic that is latex-free non-pyrogenic, hypoallergenic and non-toxic. Information brochure or the outer package of the item should clearly indicate latex free | Corrugated breathing circuit with water traps and extra limb with breathing bag - Single Use Adult Long 22mm Circuit: 2.5m; 22mm Extra limb: 1.2m OD: 22mm Reservoir bag: 2l | Offer 1 | Each | Biocor Hospital Supplies pty Ltd | 156.35 | 95 | Intersurgical | Intersurgical | Lithuania
RT13-02-028 | Oxygen Mask | Oxygen Mask | VENTURI OXYGEN MASK - MRI COMPATIBLE Purpose: Delivery of oxygen and entrained room air to a spontaneously breathing patient. Suitable for the MRI Suite Modular design correctly sized and comfortable anatomically shaped preformed mask shaped to fit over the patient’s nose and mouth with under the chin cover for a better seal Atraumatic to the nose. 2 exhalation holes in the side of the mask, a secure retaining strap. All components to be MRI compatible. Flexible corrugated tube: Length 10 - 20cm with a 22mm connector Must have clear kink resistant, crush proof delivery tubing > 2 m Colour coded Venturi device with the Oxygen concentration in % and required flow rate must be moulded or clearly indicated on the device The components must be securely attached to each other and not separate accidentally in normal use. Manufactured from latex-free medical grade plastic - must comply with BS EN 13544-3:2001 standards Individually packed, easy to open without tubing getting tangled, clinically clean - SINGLE use Information brochure or the outer package of the item should clearly indicate latex free | Venturi Oxygen Mask MRI COMPATIBLE 28% - Adult ,Colour coded venturi: White | Offer 1 | Each | Diverse Pharmaceuticals | 11.73 | 94 | Prime +Care | NINGBO SHENGYURUI | China
RT13-02-029 | Oxygen Mask | Oxygen Mask | VENTURI OXYGEN MASK - MRI COMPATIBLE Purpose: Delivery of oxygen and entrained room air to a spontaneously breathing patient. Suitable for the MRI Suite Modular design correctly sized and comfortable anatomically shaped preformed mask shaped to fit over the patient’s nose and mouth with under the chin cover for a better seal Atraumatic to the nose. 2 exhalation holes in the side of the mask, a secure retaining strap. All components to be MRI compatible. Flexible corrugated tube: Length 10 - 20cm with a 22mm connector Must have clear kink resistant, crush proof delivery tubing > 2 m Colour coded Venturi device with the Oxygen concentration in % and required flow rate must be moulded or clearly indicated on the device The components must be securely attached to each other and not separate accidentally in normal use. Manufactured from latex-free medical grade plastic - must comply with BS EN 13544-3:2001 standards Individually packed, easy to open without tubing getting tangled, clinically clean - SINGLE use Information brochure or the outer package of the item should clearly indicate latex free | Venturi Oxygen Mask MRI COMPATIBLE 40% - Adult, Colour coded venturi: Red | Offer 1 | Each | Diverse Pharmaceuticals | 11.73 | 94 | Prime +Care | NINGBO SHENGYURUI | China
RT13-02-030 | Oxygen Mask | Oxygen Mask | VENTURI OXYGEN MASK - MRI COMPATIBLE Purpose: Delivery of oxygen and entrained room air to a spontaneously breathing patient. Suitable for the MRI Suite Modular design correctly sized and comfortable anatomically shaped preformed mask shaped to fit over the patient’s nose and mouth with under the chin cover for a better seal Atraumatic to the nose. 2 exhalation holes in the side of the mask, a secure retaining strap. All components to be MRI compatible. Flexible corrugated tube: Length 10 - 20cm with a 22mm connector Must have clear kink resistant, crush proof delivery tubing > 2 m Colour coded Venturi device with the Oxygen concentration in % and required flow rate must be moulded or clearly indicated on the device The components must be securely attached to each other and not separate accidentally in normal use. Manufactured from latex-free medical grade plastic - must comply with BS EN 13544-3:2001 standards Individually packed, easy to open without tubing getting tangled, clinically clean - SINGLE use Information brochure or the outer package of the item should clearly indicate latex free | Venturi Oxygen Mask MRI COMPATIBLE 60% - Adult ,Colour coded venturi: Green | Offer 1 | Each | Diverse Pharmaceuticals | 11.73 | 94 | Prime +Care | NINGBO SHENGYURUI | China
RT13-02-031 | Oxygen Mask | VENTURI OXYGEN MASK - MRI COMPATIBLE | VENTURI OXYGEN MASK - MRI COMPATIBLE - 24% - Paediatric, Colour coded venturi: Blue Purpose: Delivery of oxygen and entrained room air to a spontaneously breathing patient. Suitable for the MRI Suite Modular design correctly sized and comfortable anatomically shaped preformed mask shaped to fit over the patient’s nose and mouth with under the chin cover for a better seal Atraumatic to the nose. 2 exhalation holes in the side of the mask, a secure retaining strap. All components to be MRI compatible. Must have clear kink resistant, crush proof delivery tubing > 2 m Colour coded Venturi device with the Oxygen concentration in % and required flow rate must be moulded or clearly indicated on the device The components must be securely attached to each other and not separate accidentally in normal use. Manufactured from latex-free medical grade plastic - must comply with BS EN 13544-3:2001 standards Individually packed, easy to open without tubing getting tangled, clinically clean - SINGLE use Information brochure or the outer package of the item should clearly indicate latex free | Venturi Oxygen Mask MRI COMPATIBLE 24% - Paediatric, Colour coded venturi: Blue | Offer 1 | Each | Diverse Pharmaceuticals | 11.5 | 94 | Prime +Care | NINGBO SHENGYURUI | China
RT13-02-032 | Oxygen Mask | VENTURI OXYGEN MASK - MRI COMPATIBLE | VENTURI OXYGEN MASK - MRI COMPATIBLE - 28% - Paediatric, Colour coded venturi: White Purpose: Delivery of oxygen and entrained room air to a spontaneously breathing patient. Suitable for the MRI Suite Modular design correctly sized and comfortable anatomically shaped preformed mask shaped to fit over the patient’s nose and mouth with under the chin cover for a better seal Atraumatic to the nose. 2 exhalation holes in the side of the mask, a secure retaining strap. All components to be MRI compatible. Must have clear kink resistant, crush proof delivery tubing > 2 m Colour coded Venturi device with the Oxygen concentration in % and required flow rate must be moulded or clearly indicated on the device The components must be securely attached to each other and not separate accidentally in normal use. Manufactured from latex-free medical grade plastic - must comply with BS EN 13544-3:2001 standards Individually packed, easy to open without tubing getting tangled, clinically clean - SINGLE use Information brochure or the outer package of the item should clearly indicate latex free | Venturi Oxygen Mask MRI COMPATIBLE 28% - Paediatric, Colour coded venturi: White | Offer 1 | Each | Diverse Pharmaceuticals | 11.5 | 94 | Prime +Care | NINGBO SHENGYURUI | China
RT13-02-033 | Oxygen Mask | VENTURI OXYGEN MASK - MRI COMPATIBLE | VENTURI OXYGEN MASK - MRI COMPATIBLE - 40% - Paediatric, Colour coded venturi: Red Purpose: Delivery of oxygen and entrained room air to a spontaneously breathing patient. Suitable for the MRI Suite Modular design correctly sized and comfortable anatomically shaped preformed mask shaped to fit over the patient’s nose and mouth with under the chin cover for a better seal Atraumatic to the nose. 2 exhalation holes in the side of the mask, a secure retaining strap. All components to be MRI compatible. Must have clear kink resistant, crush proof delivery tubing > 2 m Colour coded Venturi device with the Oxygen concentration in % and required flow rate must be moulded or clearly indicated on the device The components must be securely attached to each other and not separate accidentally in normal use. Manufactured from latex-free medical grade plastic - must comply with BS EN 13544-3:2001 standards Individually packed, easy to open without tubing getting tangled, clinically clean - SINGLE use Information brochure or the outer package of the item should clearly indicate latex free | Venturi Oxygen Mask MRI COMPATIBLE 40% - Paediatric, Colour coded venturi: Red | Offer 1 | Each | Diverse Pharmaceuticals | 11.5 | 94 | Prime +Care | NINGBO SHENGYURUI | China
RT13-02-034 | Oxygen Mask | VENTURI OXYGEN MASK - MRI COMPATIBLE | VENTURI OXYGEN MASK - MRI COMPATIBLE Purpose: Delivery of oxygen and entrained room air to a spontaneously breathing patient. Suitable for the MRI Suite Modular design correctly sized and comfortable anatomically shaped preformed mask shaped to fit over the patient’s nose and mouth with under the chin cover for a better seal Atraumatic to the nose. 2 exhalation holes in the side of the mask, a secure retaining strap. All components to be MRI compatible. Must have clear kink resistant, crush proof delivery tubing > 2 m Colour coded Venturi device with the Oxygen concentration in % and required flow rate must be moulded or clearly indicated on the device The components must be securely attached to each other and not separate accidentally in normal use. Manufactured from latex-free medical grade plastic - must comply with BS EN 13544-3:2001 standards Individually packed, easy to open without tubing getting tangled, clinically clean - SINGLE use Information brochure or the outer package of the item should clearly indicate latex free | Venturi Oxygen Mask MRI COMPATIBLE 60% - Paediatric ,Colour coded venturi: Green | Offer 1 | Each | Diverse Pharmaceuticals | 11.5 | 94 | Prime +Care | NINGBO SHENGYURUI | China
RT13-02-035 | Oxygen Mask | VENTURI OXYGEN MASK - MRI COMPATIBLE | VENTURI OXYGEN MASK - MRI COMPATIBLE 24% - Adult, Colour coded venturi: Blue Purpose: Delivery of oxygen and entrained room air to a spontaneously breathing patient. Suitable for the MRI Suite Modular design correctly sized and comfortable anatomically shaped preformed mask shaped to fit over the patient’s nose and mouth with under the chin cover for a better seal Atraumatic to the nose. 2 exhalation holes in the side of the mask, a secure retaining strap. All components to be MRI compatible. Must have clear kink resistant, crush proof delivery tubing > 2 m Colour coded Venturi device with the Oxygen concentration in % and required flow rate must be moulded or clearly indicated on the device The components must be securely attached to each other and not separate accidentally in normal use. Manufactured from latex-free medical grade plastic - must comply with BS EN 13544-3:2001 standards Individually packed, easy to open without tubing getting tangled, clinically clean - SINGLE use Information brochure or the outer package of the item should clearly indicate latex free | Venturi Oxygen Mask MRI COMPATIBLE 24% - Adult, Colour coded venturi: Blue | Offer 1 | Each | Diverse Pharmaceuticals | 11.73 | 94 | Prime +Care | NINGBO SHENGYURUI | China
RT13-13-001 | MOUTHPIECE | MOUTHPIECE | MOUTHPIECE FOR IPPB MACHINE ID: 15mm OD: 22mm Consists of a mouthpiece made of translucent molded-plastic . Must have a comfortable fit and be compatible with IPPB machines. Single use, individually packed and surgical clean | MOUTHPIECE FOR IPPB MACHINE ID: 15mm OD: 22mm | Offer 1 | Each | Diverse Pharmaceuticals | 1.96 | 94 | Prime +Care | ZHEJIANG MEDTEC | China
RT13-11-006 | Resuscitation and Re-breathing bags | resuscitation bag | BAG VALVE MASK RESUSCITATION BAG WITH PEEP VALVE - SINGLE USE ONLY Purpose: Used to provide positive pressure ventilation to a patient who is not breathing or who is breathing inadequately Ventilation bag: Clear, robust silicone. Self-inflating with a fast recoil time, and thin-walled bag to allow lung compliance assessment. Must have an ergonomic grip, Oxygen reservoir bag: PVC One-way Intake / Reservoir valve between ventilation bag and reservoir bag that only allows one way flow into the ventilation bag Oxygen from reservoir bag and room air must enter the ventilation bag if fresh gas flow is inadequate. Secure connector for oxygen tubing. Oxygen tubing: >1 m Patient valves between ventilation bag and mask. Comprising of a lip valve (duck-billed valve), expiratory valve to prevent entrainment of room air, pressure release valve – 35cmH20 (must be printed on valve) Standard connector 15mm ID, 22mm OD. Must have a PEEP valve Masks with comfortable patient fit and user interface. Malleable, non-collapsible with smooth surface or ridges. 1 piece modular face mask Able to create a tight seal between the mask and patient's’ face and must be transparent. There must be a secure fit between mask connector and breathing circuit All connections between components must be easily assembled and disassembled with tight fit and no leak. Must not disassemble during normal use Manufactured from latex-free, clear silicone, PVC, and medical grade plastic Information brochure or the outer package of the item should clearly indicate latex free | Bag valve mask resuscitation bag with PEEP valve.Adult Patient Ventilation bag: Volume: 1000- 1600ml Oxygen reservoir bag: Volume: 2600ml Peep valve adjustable to 5-20 cm/H20 Must be supplied with 2 anatomically correct silicone masks – size 4 and 5 | Offer 1 | Each | Dynalife Healthcare and Medical Supplies | 330.05 | 92 | Bescare | Gammer International | China
RT13-02-037 | Oxygen Mask | VENTURI OXYGEN MASK - MRI COMPATIBLE | VENTURI OXYGEN MASK - MRI COMPATIBLE 40% - Adult, Colour coded venturi: Red Purpose: Delivery of oxygen and entrained room air to a spontaneously breathing patient. Suitable for the MRI Suite Modular design correctly sized and comfortable anatomically shaped preformed mask shaped to fit over the patient’s nose and mouth with under the chin cover for a better seal Atraumatic to the nose. 2 exhalation holes in the side of the mask, a secure retaining strap. All components to be MRI compatible. Must have clear kink resistant, crush proof delivery tubing > 2 m Colour coded Venturi device with the Oxygen concentration in % and required flow rate must be moulded or clearly indicated on the device The components must be securely attached to each other and not separate accidentally in normal use. Manufactured from latex-free medical grade plastic - must comply with BS EN 13544-3:2001 standards Individually packed, easy to open without tubing getting tangled, clinically clean - SINGLE use Information brochure or the outer package of the item should clearly indicate latex free | Venturi Oxygen Mask MRI COMPATIBLE 40% - Adult, Colour coded venturi: Red | Offer 1 | Each | Emergency Hospital Supplies, | 9 | 91 | Bonreemed | Bonree Medical Co., Ltd. | China
RT13-23-002 | Cuffed and Uncuffed Tubes | UNCUFFED ORAL SOUTH-FACING | UNCUFFED ORAL SOUTH-FACING PREFORMED (RAE) ENDOTRACHEAL TUBE WITH 1 OR 2 MURPHY EYE Purpose: Short term oral intubation where tube is South facing and removed from surgical field Consists of a clear thermo-sensitive PVC tube that is kink resistant, with a South polar preformed curve with 1 or 2 Murphy eye Must be able to temporarily straightened the tube to allow the passage of a suction catheter Must have a radiopaque line and must be marked with numbered depth markings in 1cm intervals, tube size, length, ID, OD, trade name. The markings on tube must be clear and must not rub off. Must also have marks/reference lines proximal to cuff to aid with depth of intubation and guide proper placement. Preference will be given to tubes with a second circumferential band (not more than 1cm from each other) to guide proper placement. Must also have a distinct mark at preformed curve to aid positioning Proximal end - secure 15mm OD circuit connector made from medical grade plastic Distal end - Atraumatic left-facing bevelled tip with Murphy eye on the opposite side All components must be latex-free sterile and individually packed in peel pouch that is easy to open - For SINGLE use only To comply with the latest ISO 5361 Information brochure or the outer package of the item should clearly indicate latex free | Uncuffed Oral South-Facing Preformed (RAE) Endotracheal Tube with Murphy eye. ID: 2.5mm | Offer 1 | Each | Emergency Hospital Supplies, | 8.8 | 91 | Bonreemed | Bonree Medical Co., Ltd. | China
RT13-23-003 | Cuffed and Uncuffed Tubes | UNCUFFED ORAL SOUTH-FACING | UNCUFFED ORAL SOUTH-FACING PREFORMED (RAE) ENDOTRACHEAL TUBE WITH 1 OR 2 MURPHY EYE Purpose: Short term oral intubation where tube is South facing and removed from surgical field Consists of a clear thermo-sensitive PVC tube that is kink resistant, with a South polar preformed curve with 1 or 2 Murphy eye Must be able to temporarily straightened the tube to allow the passage of a suction catheter Must have a radiopaque line and must be marked with numbered depth markings in 1cm intervals, tube size, length, ID, OD, trade name. The markings on tube must be clear and must not rub off. Must also have marks/reference lines proximal to cuff to aid with depth of intubation and guide proper placement. Preference will be given to tubes with a second circumferential band (not more than 1cm from each other) to guide proper placement. Must also have a distinct mark at preformed curve to aid positioning Proximal end - secure 15mm OD circuit connector made from medical grade plastic Distal end - Atraumatic left-facing bevelled tip with Murphy eye on the opposite side All components must be latex-free sterile and individually packed in peel pouch that is easy to open - For SINGLE use only To comply with the latest ISO 5361 Information brochure or the outer package of the item should clearly indicate latex free | Uncuffed Oral South-Facing Preformed (RAE) Endotracheal Tube with Murphy eye. ID: 3.0mm OD 4.0mm Length ± 105mm | Offer 1 | Each | Emergency Hospital Supplies, | 8.8 | 91 | Bonreemed | Bonree Medical Co., Ltd. | China
RT13-23-004 | Cuffed and Uncuffed Tubes | UNCUFFED ORAL SOUTH-FACING | UNCUFFED ORAL SOUTH-FACING PREFORMED (RAE) ENDOTRACHEAL TUBE WITH 1 OR 2 MURPHY EYE Purpose: Short term oral intubation where tube is South facing and removed from surgical field Consists of a clear thermo-sensitive PVC tube that is kink resistant, with a South polar preformed curve with 1 or 2 Murphy eye Must be able to temporarily straightened the tube to allow the passage of a suction catheter Must have a radiopaque line and must be marked with numbered depth markings in 1cm intervals, tube size, length, ID, OD, trade name. The markings on tube must be clear and must not rub off. Must also have marks/reference lines proximal to cuff to aid with depth of intubation and guide proper placement. Preference will be given to tubes with a second circumferential band (not more than 1cm from each other) to guide proper placement. Must also have a distinct mark at preformed curve to aid positioning Proximal end - secure 15mm OD circuit connector made from medical grade plastic Distal end - Atraumatic left-facing bevelled tip with Murphy eye on the opposite side All components must be latex-free sterile and individually packed in peel pouch that is easy to open - For SINGLE use only To comply with the latest ISO 5361 Information brochure or the outer package of the item should clearly indicate latex free | Uncuffed Oral South-Facing Preformed (RAE) Endotracheal Tube with Murphy eye. ID: 3.5mm OD 4.7mm Length ± 120mm | Offer 1 | Each | Emergency Hospital Supplies, | 8.8 | 91 | Bonreemed | Bonree Medical Co., Ltd. | China
RT13-23-005 | Cuffed and Uncuffed Tubes | UNCUFFED ORAL SOUTH-FACING | UNCUFFED ORAL SOUTH-FACING PREFORMED (RAE) ENDOTRACHEAL TUBE WITH 1 OR 2 MURPHY EYE Purpose: Short term oral intubation where tube is South facing and removed from surgical field Consists of a clear thermo-sensitive PVC tube that is kink resistant, with a South polar preformed curve with 1 or 2 Murphy eye Must be able to temporarily straightened the tube to allow the passage of a suction catheter Must have a radiopaque line and must be marked with numbered depth markings in 1cm intervals, tube size, length, ID, OD, trade name. The markings on tube must be clear and must not rub off. Must also have marks/reference lines proximal to cuff to aid with depth of intubation and guide proper placement. Preference will be given to tubes with a second circumferential band (not more than 1cm from each other) to guide proper placement. Must also have a distinct mark at preformed curve to aid positioning Proximal end - secure 15mm OD circuit connector made from medical grade plastic Distal end - Atraumatic left-facing bevelled tip with Murphy eye on the opposite side All components must be latex-free sterile and individually packed in peel pouch that is easy to open - For SINGLE use only To comply with the latest ISO 5361 Information brochure or the outer package of the item should clearly indicate latex free | Uncuffed Oral South-Facing Preformed (RAE) Endotracheal Tube with Murphy eye. ID: 4.0mm OD 5.3mm Length ± 130mm | Offer 1 | Each | Emergency Hospital Supplies, | 8.8 | 91 | Bonreemed | Bonree Medical Co., Ltd. | China
RT13-23-006 | Cuffed and Uncuffed Tubes | UNCUFFED ORAL SOUTH-FACING | UNCUFFED ORAL SOUTH-FACING PREFORMED (RAE) ENDOTRACHEAL TUBE WITH 1 OR 2 MURPHY EYE Purpose: Short term oral intubation where tube is South facing and removed from surgical field Consists of a clear thermo-sensitive PVC tube that is kink resistant, with a South polar preformed curve with 1 or 2 Murphy eye Must be able to temporarily straightened the tube to allow the passage of a suction catheter Must have a radiopaque line and must be marked with numbered depth markings in 1cm intervals, tube size, length, ID, OD, trade name. The markings on tube must be clear and must not rub off. Must also have marks/reference lines proximal to cuff to aid with depth of intubation and guide proper placement. Preference will be given to tubes with a second circumferential band (not more than 1cm from each other) to guide proper placement. Must also have a distinct mark at preformed curve to aid positioning Proximal end - secure 15mm OD circuit connector made from medical grade plastic Distal end - Atraumatic left-facing bevelled tip with Murphy eye on the opposite side All components must be latex-free sterile and individually packed in peel pouch that is easy to open - For SINGLE use only To comply with the latest ISO 5361 Information brochure or the outer package of the item should clearly indicate latex free | Uncuffed Oral South-Facing Preformed (RAE) Endotracheal Tube with Murphy eye. ID: 4.5mm OD 6.0mm Length ± 145mm | Offer 1 | Each | Emergency Hospital Supplies, | 8.8 | 91 | Bonreemed | Bonree Medical Co., Ltd. | China
RT13-23-007 | Cuffed and Uncuffed Tubes | UNCUFFED ORAL SOUTH-FACING | UNCUFFED ORAL SOUTH-FACING PREFORMED (RAE) ENDOTRACHEAL TUBE WITH 1 OR 2 MURPHY EYE Purpose: Short term oral intubation where tube is South facing and removed from surgical field Consists of a clear thermo-sensitive PVC tube that is kink resistant, with a South polar preformed curve with 1 or 2 Murphy eye Must be able to temporarily straightened the tube to allow the passage of a suction catheter Must have a radiopaque line and must be marked with numbered depth markings in 1cm intervals, tube size, length, ID, OD, trade name. The markings on tube must be clear and must not rub off. Must also have marks/reference lines proximal to cuff to aid with depth of intubation and guide proper placement. Preference will be given to tubes with a second circumferential band (not more than 1cm from each other) to guide proper placement. Must also have a distinct mark at preformed curve to aid positioning Proximal end - secure 15mm OD circuit connector made from medical grade plastic Distal end - Atraumatic left-facing bevelled tip with Murphy eye on the opposite side All components must be latex-free sterile and individually packed in peel pouch that is easy to open - For SINGLE use only To comply with the latest ISO 5361 Information brochure or the outer package of the item should clearly indicate latex free | Uncuffed Oral South-Facing Preformed (RAE) Endotracheal Tube with Murphy eye. ID: 5.0mm OD 6.7mm Length ± 155mm | Offer 1 | Each | Emergency Hospital Supplies, | 8.8 | 91 | Bonreemed | Bonree Medical Co., Ltd. | China
RT13-23-008 | Cuffed and Uncuffed Tubes | UNCUFFED ORAL SOUTH-FACING | UNCUFFED ORAL SOUTH-FACING PREFORMED (RAE) ENDOTRACHEAL TUBE WITH 1 OR 2 MURPHY EYE Purpose: Short term oral intubation where tube is South facing and removed from surgical field Consists of a clear thermo-sensitive PVC tube that is kink resistant, with a South polar preformed curve with 1 or 2 Murphy eye Must be able to temporarily straightened the tube to allow the passage of a suction catheter Must have a radiopaque line and must be marked with numbered depth markings in 1cm intervals, tube size, length, ID, OD, trade name. The markings on tube must be clear and must not rub off. Must also have marks/reference lines proximal to cuff to aid with depth of intubation and guide proper placement. Preference will be given to tubes with a second circumferential band (not more than 1cm from each other) to guide proper placement. Must also have a distinct mark at preformed curve to aid positioning Proximal end - secure 15mm OD circuit connector made from medical grade plastic Distal end - Atraumatic left-facing bevelled tip with Murphy eye on the opposite side All components must be latex-free sterile and individually packed in peel pouch that is easy to open - For SINGLE use only To comply with the latest ISO 5361 Information brochure or the outer package of the item should clearly indicate latex free | Uncuffed Oral South-Facing Preformed (RAE) Endotracheal Tube with Murphy eye. ID: 5.5mm OD 7.3mm Length ± 165mm | Offer 1 | Each | Emergency Hospital Supplies, | 8.8 | 91 | Bonreemed | Bonree Medical Co., Ltd. | China
RT13-23-020 | Cuffed and Uncuffed Tubes | CUFFED ORAL SOUTH-FACING | CUFFED ORAL SOUTH-FACING PREFORMED (RAE) ENDOTRACHEAL TUBE WITH 1 or 2 MURPHY EYE - ADULT Purpose: Short term oral intubation where tube is South facing and removed from surgical field Consists of a clear thermo-sensitive PVC tube that is kink resistant, with a South polar preformed curve with one or two Murphy eyes Must be able to temporarily straightened the tube to allow the passage of a suction catheter Must have a radiopaque line and must be marked with numbered depth markings in 1cm intervals, tube size, length, ID, OD, trade name. The markings on tube must be clear and must not rub off. Must also have marks/reference lines proximal to cuff to aid with depth of intubation and guide proper placement. Preference will be given to tubes with a second circumferential band (not more than 2cm from each other) to guide proper placement. Must also have a distinct mark at preformed curve to aid positioning Proximal end - secure 15mm OD circuit connector made from medical grade plastic Distal end - Atraumatic left-facing bevelled tip with Murphy eye on the opposite side Must have a high-volume-low-pressure polyurethane barrel shaped cuff with pilot cuff. Cuff thickness ± 50 microns. The pilot cuff must have a one-way valve and must indicate the degree of inflation accurately. The cuff must inflate and deflate easily and when deflated must conform closely to tube to minimise the risk of trauma during intubation and extubation and not increase the OD of the tube. The cuff must be resistant to Nitrous Oxide effects All components must be latex-free sterile and individually packed in peel pouch that is easy to open - For SINGLE use only To comply with the latest ISO 5361 Information brochure or the outer package of the item should clearly indicate latex free | Cuffed Oral South-Facing Preformed (RAE) Endotracheal Tube with Murphy eye - Adult ID:7.0mm OD 9.3mm Length ± 200mm | Offer 1 | Each | Emergency Hospital Supplies, | 11.2 | 91 | Bonreemed | Bonree Medical Co., Ltd. | China
RT13-23-021 | Cuffed and Uncuffed Tubes | CUFFED ORAL SOUTH-FACING | CUFFED ORAL SOUTH-FACING PREFORMED (RAE) ENDOTRACHEAL TUBE WITH 1 or 2 MURPHY EYE - ADULT Purpose: Short term oral intubation where tube is South facing and removed from surgical field Consists of a clear thermo-sensitive PVC tube that is kink resistant, with a South polar preformed curve with one or two Murphy eyes Must be able to temporarily straightened the tube to allow the passage of a suction catheter Must have a radiopaque line and must be marked with numbered depth markings in 1cm intervals, tube size, length, ID, OD, trade name. The markings on tube must be clear and must not rub off. Must also have marks/reference lines proximal to cuff to aid with depth of intubation and guide proper placement. Preference will be given to tubes with a second circumferential band (not more than 2cm from each other) to guide proper placement. Must also have a distinct mark at preformed curve to aid positioning Proximal end - secure 15mm OD circuit connector made from medical grade plastic Distal end - Atraumatic left-facing bevelled tip with Murphy eye on the opposite side Must have a high-volume-low-pressure polyurethane barrel shaped cuff with pilot cuff. Cuff thickness ± 50 microns. The pilot cuff must have a one-way valve and must indicate the degree of inflation accurately. The cuff must inflate and deflate easily and when deflated must conform closely to tube to minimise the risk of trauma during intubation and extubation and not increase the OD of the tube. The cuff must be resistant to Nitrous Oxide effects All components must be latex-free sterile and individually packed in peel pouch that is easy to open - For SINGLE use only To comply with the latest ISO 5361 Information brochure or the outer package of the item should clearly indicate latex free | Cuffed Oral South-Facing Preformed (RAE) Endotracheal Tube with Murphy eye - Adult ID:7.5mm OD 10mm Length ± 210mm | Offer 1 | Each | Emergency Hospital Supplies, | 11.2 | 91 | Bonreemed | Bonree Medical Co., Ltd. | China
RT13-23-022 | Cuffed and Uncuffed Tubes | CUFFED ORAL SOUTH-FACING | CUFFED ORAL SOUTH-FACING PREFORMED (RAE) ENDOTRACHEAL TUBE WITH 1 or 2 MURPHY EYE - ADULT Purpose: Short term oral intubation where tube is South facing and removed from surgical field Consists of a clear thermo-sensitive PVC tube that is kink resistant, with a South polar preformed curve with one or two Murphy eyes Must be able to temporarily straightened the tube to allow the passage of a suction catheter Must have a radiopaque line and must be marked with numbered depth markings in 1cm intervals, tube size, length, ID, OD, trade name. The markings on tube must be clear and must not rub off. Must also have marks/reference lines proximal to cuff to aid with depth of intubation and guide proper placement. Preference will be given to tubes with a second circumferential band (not more than 2cm from each other) to guide proper placement. Must also have a distinct mark at preformed curve to aid positioning Proximal end - secure 15mm OD circuit connector made from medical grade plastic Distal end - Atraumatic left-facing bevelled tip with Murphy eye on the opposite side Must have a high-volume-low-pressure polyurethane barrel shaped cuff with pilot cuff. Cuff thickness ± 50 microns. The pilot cuff must have a one-way valve and must indicate the degree of inflation accurately. The cuff must inflate and deflate easily and when deflated must conform closely to tube to minimise the risk of trauma during intubation and extubation and not increase the OD of the tube. The cuff must be resistant to Nitrous Oxide effects All components must be latex-free sterile and individually packed in peel pouch that is easy to open - For SINGLE use only To comply with the latest ISO 5361 Information brochure or the outer package of the item should clearly indicate latex free | Cuffed Oral South-Facing Preformed (RAE) Endotracheal Tube with Murphy eye - Adult ID:8.0mm OD 10.7mm Length ± 220mm | Offer 1 | Each | Emergency Hospital Supplies, | 11.2 | 91 | Bonreemed | Bonree Medical Co., Ltd. | China
RT13-51-005 | CATHETER MOUNT CONNECTOR | CATHETER MOUNT WITH ELBOW CONNECTOR | CATHETER MOUNT WITH ELBOW CONNECTOR, LUER LOCK SAMPLE PORT AND TOGGLED CAP Purpose: Connector to join an endotracheal or tracheostomy tube to a breathing circuit with sample port for Capnography Consists of flexible, low-flow smooth-bore tubing with an external coil to prevent kinks. Elbow connector with a Luer lock sample port and toggled cap All components latex-free, sterile and individually packed in peel pouch that is easy to open - For SINGLE use only Information brochure or the outer package of the item should clearly indicate latex free | Catheter mount with Elbow connector, C02 sampling port and toggled cap - Paediatric Volume: ± 15-30ml Length ± 90mm Connectors: 22mm Male/ 15mm Female (patient end); 22mm Female (machine end) | Offer 1 | Each | Ithuse Medical | 14 | 95 | Hisern | Zhejiang Haisheng Medical Device | China
RT13-51-012 | CATHETER MOUNT CONNECTOR | CATHETER MOUNT- EXTENDIBLE | CATHETER MOUNT - Extendible with Elbow Connector, Luer Lock Sample Port and Toggled Cap Purpose: Connector to join an endotracheal or tracheostomy tube to a breathing circuit with sample port for Capnography. Consists of flexible, extendible tubing. The internal diameter must remain constant in every position. Must have an Elbow connector with a Luer lock sample port and toggled cap All components latex-free, sterile and individually packed in peel pouch that is easy to open - For SINGLE use only Information brochure or the outer package of the item should clearly indicate latex free | Catheter mount: Extendible connector, Luer Lock Sample Port and Toggled Cap - Paediatric Volume: ± 15-30ml Length ± 45mm - 100mm Connectors: 22mm Male/ 15mm Female (patient end) 22mm Female (machine end) | Offer 1 | Each | Ithuse Medical | 19.5 | 95 | Hisern | Zhejiang Haisheng Medical Device | China
RT13-51-015 | CATHETER MOUNT CONNECTOR | CATHETER MOUNT- EXTENDIBLE | CATHETER MOUNT - Extendible with DOUBLE swivel elbow connector, Bronchoscopy/Aspiration Port and toggled cap Purpose: Connector to join an endotracheal or tracheostomy tube to a breathing circuit with with a bronchoscopy/aspiration port. Consists of flexible, extendible tubing. The internal diameter must remain constant in every position. Double swivel Elbow connector with a Bronchoscopy/Aspiration Port. Must have a diaphragm which hugs fibre-optic bronchoscope All components latex-free, sterile and individually packed in peel pouch that is easy to open - For SINGLE use only Information brochure or the outer package of the item should clearly indicate latex free | Catheter mount: Extendible with DOUBLE swivel elbow connector, Bronchoscopy/Aspiration Port and toggled cap - Paediatric Volume: ± 15-30ml Length ± 45mm - 100mm Connectors: 22mm Male/ 15mm Female (patient end) 22mm Female (machine end) | Offer 1 | Each | Ithuse Medical | 19.5 | 82.6691 | Hisern | Zhejiang Haisheng Medical Device | China
RT13-24-003 | ENDOTRACHEAL TUBE | DOUBLE LUMEN ENDOBRONCHIAL TUBE | DOUBLE LUMEN ENDOBRONCHIAL TUBE- LEFT Purpose: For isolation of one lung to prevent contamination, or for ventilation of one lung in thoracic surgery Consists of a double Lumen clear PVC tube that is thermo-sensitive with radiopaque lines and markers at key features. Markings starting at 26cm from distal end marked every 2cm up to and about 32cm Tubes to be flexible enough to allow manipulation through distorted airways, but without luminal collapse. Must have a smooth transition of the inner lumen without narrowing, to allow easy passage of flexible fibre-optic-bronchoscope. Tracheal and Bronchial extensions on the main tube must be colour coded - blue for bronchial, clear for tracheal Stylet: Malleable shape forming - packaged in place down the bronchial lumen Supplied with 2 proximal right-angled connectors with a Bronchoscopy/Aspiration Port The port must have a toggled cap and allow for easy passage of either a flexible fibre-optic-bronchoscope, or a suction catheter – both without a leak. The proximal angled connectors to connect to the bronchial and tracheal tubes and the distal connectors to connect to a standard Y-connector that is compatible with a standard breathing circuit. 2 - 4 Suction catheters: with suction control, to fit the total length of the tube Colourless tracheal high-volume-low-pressure polyvinyl chloride cuff – with colourless pilot balloon, clearly labelled “tracheal” Blue bronchial low pressure Polyurethane cuff - with blue pilot balloon, clearly labelled “bronchial”. Bronchial cuff to be short and cylindrical. One-way valves in pilot balloons. The cuffs must inflate and deflate easily, and the pilot cuffs must indicate the degree of inflation accurately. Cuffs must be resistant to the effects of nitrous oxide. When deflated, cuffs must conform closely to the tube to minimise the risk of trauma during intubation and extubation. All components must be latex-free, sterile and individually packed in peel pouch - SINGLE use only To comply with ISO 16628 Information brochure or the outer package of the item should clearly indicate latex free | Double Lumen Endobronchial tube: Left Size: 32Fr Shaft Length: < 420mm ID: > 3.3mm (bilaterally) OD: < 10.7mm ± 2mm | Offer 1 | Each | Levingston (Pty) Ltd | 299 | 95 | Bonreemed | Bonree Medical | China
RT13-01-006 | Intubation Devices | Nasal Cannula | Nasal Cannula with integrated C02 sampling for use in an Pre mature Neonates Prong diameter: 1.5mm Prong length: 5.5mm Distance between medial surfaces of the prongs: 4mm NASAL CANNULA WITH INTEGRATED C02 SAMPLING Purpose: For low flow nasal oxygen delivery in a spontaneously breathing patient, with simultaneous C02 sampling Clear or colour tinted non-kinking, non-crush PVC tubing at least 2m long Over-the-ear style with under-chin toggle or sleeve. Nasal cannula tips must be soft, anatomically shaped (curved) and tapered at the ends Any portion of the cannula which comes in contact with facial skin must be soft and uniformly contoured without protruding edges or ridges or flaps Additional C02 sampling line with a male luer connector/ attached to the PVC tubing All components must be latex-free clinically clean and individually packed in peel pouch that is easy to open - For SINGLE use only Information brochure or the outer package of the item should clearly indicate latex free | Nasal Cannula with integrated C02 sampling for use in an Pre mature Neonates Prong diameter: 1.5mm Prong length: 5.5mm Distance between medial surfaces of the prongs: 4mm | Offer 1 | Each | Logan Medical and Surgical | R13.80 | 95 | TruAir | SHAOXING MEDPLY MEDICAL PRODUCTS CO. LTD | China
RT13-01-016 | Intubation Devices | Nasal Cannula | Nasal Cannula with integrated C02 sampling for use in an pre mature neonates - Hypoallergenic Prong diameter: 1.5mm Prong length: 5.5mm Distance between medial surfaces of the prongs: 4mm NASAL CANNULA WITH INTEGRATED C02 SAMPLING - HYPOALLERGENIC Purpose: For low flow nasal oxygen delivery in a spontaneously breathing patient, with simultaneous C02 sampling Clear or colour tinted non-kinking, non-crush PVC tubing at least 2m long Over-the-ear style with under-chin toggle or sleeve. Nasal cannula tips must be soft, anatomically shaped (curved) and tapered at the ends Any portion of the cannula which comes in contact with facial skin must be soft and uniformly contoured without protruding edges or ridges or flaps Additional C02 sampling line with a male luer connector/ attached to the PVC tubing All components must be latex-free, DEHP free and safe for use in hypoallergenic patients Clinically clean and individually packed in peel pouch that is easy to open - For SINGLE use only Information brochure or the outer package of the item should clearly indicate latex free and DEHP free | Nasal Cannula with integrated C02 sampling for use in an pre mature neonates - Hypoallergenic Prong diameter: 1.5mm Prong length: 5.5mm Distance between medial surfaces of the prongs: 4mm | Offer 1 | Each | Logan Medical and Surgical | 13.8 | 95 | TruAir | SHAOXING MEDPLY MEDICAL PRODUCTS CO. LTD | China
RT13-01-018 | Intubation Devices | Nasal Cannula | NASAL CANNULA WITH INTEGRATED C02 SAMPLING - HYPOALLERGENIC Purpose: For low flow nasal oxygen delivery in a spontaneously breathing patient, with simultaneous C02 sampling Clear or colour tinted non-kinking, non-crush PVC tubing at least 2m long Over-the-ear style with under-chin toggle or sleeve. Nasal cannula tips must be soft, anatomically shaped (curved) and tapered at the ends Any portion of the cannula which comes in contact with facial skin must be soft and uniformly contoured without protruding edges or ridges or flaps Additional C02 sampling line with a male luer connector/ attached to the PVC tubing All components must be latex-free, DEHP free and safe for use in hypoallergenic patients Clinically clean and individually packed in peel pouch that is easy to open - For SINGLE use only Information brochure or the outer package of the item should clearly indicate latex free and DEHP free | Nasal Cannula with integrated C02 sampling for use in Infants - Hypoallergenic Prong diameter: 4mm Prong length: 10mm Distance between medial surfaces of the prongs: 7mm | Offer 1 | Each | Logan Medical and Surgical | 13.8 | 95 | TruAir | SHAOXING MEDPLY MEDICAL PRODUCTS CO. LTD | China
RT13-37-002 | Emergency Cricothyroidotomy Kit | Emergency Cricothyroidotomy Kit | Emergency Cricothyrotomy Kit –Seldinger dilation technique via the cricothyroid membrane Purpose: Used for emergency airway access when conventional endotracheal intubation or ventilation can’t be performed Consists of a radiopaque, cuffed 5mm Cricothyroidotomy tube, 9cm long with 15mm connector Catheter to sit securely on tapered dilator and must be easy to advance over the dilator Short handled scalpel blade; Syringe with introducer needle; Guide-wire and curved dilator with a tapered end Manufactured from medical grade metal and plastic. All components must be latex-free sterile and individually packed in peel pouch that is easy to open - For SINGLE use only Information brochure or the outer package of the item should clearly indicate latex free | Emergency Cricothyrotomy Kit –Seldinger dilation technique via the cricothyroid membrane Consists of a radiopaque, cuffed 5mm Cricothyroidotomy tube, 9cm long with 15mm connector | Offer 1 | Each | Marcus Medical PTY Ltd. | 5323.8071 | 92.55 | Cook Medical (C-TCCSB-500 Melker Emergency Cricothyrotomy Set (Seldinger)) | Cook Medical | USA
RT13-42-010 | AIRWAY DEVICES | AIRWAY EXCHANGE CATHETER | AIRWAY EXCHANGE CATHETER Purpose: For simple, atraumatic endotracheal tube exchange whilst providing oxygenation Consists of a catheter manufactured from radiopaque medical grade plastic Must have a hollow lumen for constant oxygenation of the patient, with distal side ports to ensure adequate airflow Must be supplied with connection and adaptor for attachment of a ventilatory device. Must have a blunt tip for atraumatic intubation, centimetre markings for accurate tube placement. Latex-free, sterile and individually packed in peel pouch that is easy to open - For SINGLE use only Information brochure or the outer package of the item should clearly indicate latex free | Airway Exchange Catheter. Infant 8Fr Length: ± 450mm To fit ETT's Size: 3.0mm | Offer 1 | Each | Marcus Medical PTY Ltd. | 1678.5651 | 92.55 | Cook Medical (C-CAE-8.0-45 Cook Arway Exchange Catheter 8Fr) | Cook Medical | USA
RT13-42-011 | AIRWAY DEVICES | AIRWAY EXCHANGE CATHETER | AIRWAY EXCHANGE CATHETER Purpose: For simple, atraumatic endotracheal tube exchange whilst providing oxygenation Consists of a catheter manufactured from radiopaque medical grade plastic Must have a hollow lumen for constant oxygenation of the patient, with distal side ports to ensure adequate airflow Must be supplied with connection and adaptor for attachment of a ventilatory device. Must have a blunt tip for atraumatic intubation, centimetre markings for accurate tube placement. Latex-free, sterile and individually packed in peel pouch that is easy to open - For SINGLE use only Information brochure or the outer package of the item should clearly indicate latex free | Airway Exchange Catheter. Paediatric 11Fr Length: ± 830mm To fit ETT's Size: 4.0mm | Offer 1 | Each | Marcus Medical PTY Ltd. | 1678.5651 | 92.55 | Cook Medical (C-CAE-11.0-83 Cook Airway Exchange Catheter 11Fr) | Cook Medical | USA
RT13-42-013 | AIRWAY DEVICES | AIRWAY EXCHANGE CATHETER | AIRWAY EXCHANGE CATHETER Purpose: For simple, atraumatic endotracheal tube exchange whilst providing oxygenation Consists of a catheter manufactured from radiopaque medical grade plastic Must have a hollow lumen for constant oxygenation of the patient, with distal side ports to ensure adequate airflow Must be supplied with connection and adaptor for attachment of a ventilatory device. Must have a blunt tip for atraumatic intubation, centimetre markings for accurate tube placement. Latex-free, sterile and individually packed in peel pouch that is easy to open - For SINGLE use only Information brochure or the outer package of the item should clearly indicate latex free | Airway Exchange Catheter. Adult 19Fr Length: ± 830mm To fit ETT's Size: 7.0mm | Offer 1 | Each | Marcus Medical PTY Ltd. | 1678.5651 | 20.7633 | Cook Medical (C-CAE-19.0-83 Cook Airway Exchange Catheter 19Fr) | Cook Medical | USA
RT13-42-014 | AIRWAY DEVICES | AIRWAY EXCHANGE CATHETER | AIRWAY EXCHANGE CATHETER FOR FIBRE-OPTIC GUIDED INTUBATION Purpose: For simple, atraumatic fibre-optic guided intubation while providing oxygenation. Consists of a large diameter hollow lumen catheter that fits over a Size 4.2 fibre-optic flexible bronchoscope Manufactured from radiopaque medical grade plastic, with distal side ports to ensure adequate airflow Must be supplied with connection and adaptor for attachment of a ventilatory device. Must have a blunt tip slightly tapered tip to fit snugly on the bronchoscope. Centimetre markings from 5cm in 1cm increments Preference will be given to items with a distal side hole for ventilation Latex-free, sterile and individually packed in peel pouch that is easy to open - For SINGLE use only Information brochure or the outer package of the item should clearly indicate latex free | Airway Exchange Catheter for Fibre-Optic Guided Intubation Size: 19Fr ID: 4.7mm OD: < 5mm Length: 560mm To fit ETT's 7.0mm | Offer 1 | Each | Marcus Medical PTY Ltd. | 1678.5651 | 92.55 | Cook Medical (C-CAE-19.0-83-AIC Aintree Intubation Catheter 19Fr) | Cook Medical | USA
RT13-42-015 | AIRWAY DEVICES | Airway Exchange Catheter | Airway Exchange Catheter Kit with Guidewire for staged extubation. Purpose: For staged or difficult extubation of patient's with cardiorespiratory dysfunction and potentially difficult reintubation Consists of an exchange catheter, extra stiff guidewire, Bronchoscopic port, 2 Adaptors 1 x luer lock, 15mm connector Catheter 14Fr ID: 2.3mm OD: < 4mm Catheter Length: 700 - 900mm Guidewire: Diameter: 0.1cm Length: 160cm The wire guide to fit into the working channel of a bronchoscope. The catheter must have distal side ports to allow for airflow. Must have an airway adaptor to allow for ventilation To fit an endotracheal tube of Size 5 or larger Latex-free medical grade metal and plastic, sterile and individually packed in peel pouch that is easy to open - For SINGLE use only Information brochure or the outer package of the item should clearly indicate latex free | Airway Exchange Catheter Kit with Guidewire for staged extubation.Catheter 14Fr ID: 2.3mm OD: < 4mm Catheter Length: 700 - 900mm | Offer 1 | Each | Marcus Medical PTY Ltd. | 4321.6787 | 92.55 | Cook Medical (C-CAE-14.0-83-EXTU Cook Staged Extubation Set 14Fr) | Cook Medical | USA
RT13-43-001 | Intubation Devices | RETROGRADE INTUBATION KIT | RETROGRADE INTUBATION KIT 11Fr Length: 70cm for ETT's ID: 4mm or larger Purpose: Used to assist in the placement of an endotracheal tube using retrograde techniques during a difficult airway access procedure. Consists of an Introducer needle/ Catheter introducer needle:18G and Guide Wire: Size: 0.038mm Length 110cm Guidewire made from stainless steel with two positioning marks 20cm from each end; atraumatic tips; one end straight and the other a flexible J tip. Radiopaque catheter with distal side ports and a 15mm connector proximally for ventilatory devices. Luer slip syringe to be included. Seldinger technique access to the airway via the cricothyroid membrane. Retrograde (cephalad) advancement of wire towards the oral/ nasal cavities. Antegrade advancement of a hollow catheter with central & peripheral side ports for optimal ventilation. Latex-free, sterile and individually packed in peel pouch that is easy to open - For SINGLE use only Information brochure or the outer package of the item should clearly indicate latex free | RETROGRADE INTUBATION KIT 11Fr Length: 70cm for ETT's ID: 4mm or larger | Offer 1 | Each | Marcus Medical PTY Ltd. | 3319.5503 | 92.55 | Cook Medical (C-RETRO-11.0-70-38J-110 Cook Retrograde Intubation Set 11Fr) | Cook Medical | USA
RT13-72-002 | PERCUTANEOUS TRACHEOSTOMY KITS | PERCUTANEOUS TRACHEOSTOMY KITS - WITH DILATOR | PERCUTANEOUS TRACHEOSTOMY KITS - WITH DILATOR Purpose: Tracheostomy insertion at the bedside Consists of a, radiopaque, cuffed Tracheostomy tube with an inner cannula and a 15mm connector, with hypoallergenic neck tape. The set must have a short-handled scalpel blade; a 14-15G puncture needle with a catheter; a kink-resistant guidewire with a guidewire inserter; a short 14-15FR dilator, an angled tracheostomy dilator on a guiding catheter with a safety stop and a loading dilator. The guiding catheter must have clear marking lines to prevent the dilator from being inserted too deeply The KIT must include either a one-step dilator or a set of serial dilators ( Must be appropriate for the specific Tracheostomy size) Manufactured from latex-free medical grade metal and plastic Individually packed in peel pouch that is easy to open, SINGLE use, Sterile To comply with latest ISO 5366 Information brochure or the outer package of the item should clearly indicate latex free | Percutaneous Tracheostomy Kit Size 7.5 | Offer 1 | Each | Marcus Medical PTY Ltd. | 7265.4309 | 92.55 | Cook Medical (C-PTIS-100-HC-G-EU-FLEX7.5 Cook Blue Rhino G2 Multi Percutaneous Tracheostomy Set 7.5) | Cook Medical | USA
RT13-72-004 | PERCUTANEOUS TRACHEOSTOMY KITS | PERCUTANEOUS TRACHEOSTOMY KITS - WITH DILATOR | PERCUTANEOUS TRACHEOSTOMY KITS - WITH DILATOR Purpose: Tracheostomy insertion at the bedside Consists of a, radiopaque, cuffed Tracheostomy tube with an inner cannula and a 15mm connector, with hypoallergenic neck tape. The set must have a short-handled scalpel blade; a 14-15G puncture needle with a catheter; a kink-resistant guidewire with a guidewire inserter; a short 14-15FR dilator, an angled tracheostomy dilator on a guiding catheter with a safety stop and a loading dilator. The guiding catheter must have clear marking lines to prevent the dilator from being inserted too deeply The KIT must include either a one-step dilator or a set of serial dilators ( Must be appropriate for the specific Tracheostomy size) Manufactured from latex-free medical grade metal and plastic Individually packed in peel pouch that is easy to open, SINGLE use, Sterile To comply with latest ISO 5366 Information brochure or the outer package of the item should clearly indicate latex free | Percutaneous Tracheostomy Kit Size 8.5 | Offer 1 | Each | Marcus Medical PTY Ltd. | 7265.4309 | 92.55 | Cook Medical (C-PTIS-100-HC-G-EU-FLEX8.5 Cook Blue Rhino G2 Multi Percutaneous Tracheostomy Set 8.5) | Cook Medical | USA
RT13-51-009 | CATHETER MOUNT CONNECTOR | CATHETER MOUNT WITH DOUBLE SWIVEL ELBOW CONNECTOR | CATHETER MOUNT WITH DOUBLE SWIVEL ELBOW CONNECTOR, BRONCHOSCOPY/ASPIRATION PORT AND TOGGLED CAP Purpose: Connector to join an endotracheal or tracheostomy tube to a breathing circuit with a bronchoscopy/aspiration port. Consists of flexible, low-flow smooth-bore tubing with an external coil to prevent kinks. Double swivel Elbow connector with a Bronchoscopy/Aspiration Port. Must have a diaphragm which hugs fibre-optic bronchoscope. All components latex-free, sterile and individually packed in peel pouch that is easy to open - For SINGLE use only Information brochure or the outer package of the item should clearly indicate latex free | Catheter mount with DOUBLE swivel elbow connector, Bronchoscopy/Aspiration Port and toggled cap - Adult Length: 120-150mm Connectors: 22mm Male/ 15mm Female to patient and 22mm Female to machine | Offer 1 | Each | Mbuso Medical Supplies | 16.71 | 94 | CREATE BIOTECH | CREATE BIOTECH CO,.LTD | TAIWAN
RT13-51-015 | CATHETER MOUNT CONNECTOR | CATHETER MOUNT- EXTENDIBLE | CATHETER MOUNT - Extendible with DOUBLE swivel elbow connector, Bronchoscopy/Aspiration Port and toggled cap Purpose: Connector to join an endotracheal or tracheostomy tube to a breathing circuit with with a bronchoscopy/aspiration port. Consists of flexible, extendible tubing. The internal diameter must remain constant in every position. Double swivel Elbow connector with a Bronchoscopy/Aspiration Port. Must have a diaphragm which hugs fibre-optic bronchoscope All components latex-free, sterile and individually packed in peel pouch that is easy to open - For SINGLE use only Information brochure or the outer package of the item should clearly indicate latex free | Catheter mount: Extendible with DOUBLE swivel elbow connector, Bronchoscopy/Aspiration Port and toggled cap - Paediatric Volume: ± 15-30ml Length ± 45mm - 100mm Connectors: 22mm Male/ 15mm Female (patient end) 22mm Female (machine end) | Offer 1 | Each | Mbuso Medical Supplies | 16.71 | 94 | CREATE BIOTECH | CREATE BIOTECH CO,.LTD | TAIWAN
RT13-40-008 | Intubation Devices | INTUBATION STYLET | INTUBATION STYLET with metal core Purpose: To pre-form an endotracheal tube to a desired shape to aid intubation Consists of a malleable aluminium core coated and tightly bonded with high density polyurethane/ PVC / medical grade plastic/ polyethylene Must have a soft atraumatic tip Must have a U-bend proximally or other mechanism at proximal end to limit depth Latex-free clinically clean and individually packed in peel pouch that is easy to open - For SINGLE use only Information brochure or the outer package of the item should clearly indicate latex free | Intubation Stylet, Extra-Length Adult OD: 5.0mm Length: 365mm To fit ETT's > 7.5mm | Offer 1 | Each | Mr First Aid | 13.14 | 66.1143 | Trident | Wuxi Emsrun | China
RT13-52-014 | CLOSED SUCTION CATHETER | CLOSED SUCTION CATHETER ADULT 48HR | CLOSED SUCTION CATHETER ADULT ENDOTRACHEAL - 48HRS Purpose: Closed system for respiratory suctioning through endotracheal tubes without disconnecting or interrupting ventilation. Consists of a closed suction catheter with an atraumatic tip. The catheter must have markings along the length to indicate depth of suctioning and a radiopaque line on the catheter tip. While connected but not in suction mode the catheter tip must not extend beyond the sheath or impede airflow. The suction catheter must be enclosed within a durable, flexible, plastic sheath for storage to avoid contamination. The transparent sheath must allow visualisation of markings on the suction catheter and must connect to the endotracheal tube via a clear, firmly-fitting, plastic connector or adaptor. The sheath and connector must allow for smooth advancement of catheter. The system must seal the airway and catheter from atmospheric contamination and allow uninterrupted positive pressure ventilation throughout the suction process. An effective seal around the catheter must prevent ventilation gas entering the lumen of the sheath. The system must protect the user from contamination by secretions and suctioned material. Proximal end of suction catheter must have clear chamber, so that suctioned contents can be visualised and must have lockable suction valve with capped suction tubing connector. The suction catheter must connect onto the tube with a DOUBLE swivel connector or a double-swivel Y-connector with standard 15 mm OD taper fittings. This connection must be secure and accidental disconnection must be minimal The whole device must be as lightweight as possible. Must include a light, straight, flexible, clear, concertina catheter-mount (15x22 mm connections) tube for connection to the breathing circuit The connector must allow variations in patient positions without causing torque, twisting, or obstruction of the tracheal tube, and must not cause tissue or skin pressure. The system must allow for irrigation via a lateral irrigation port which must be a capped with a one-way valve that connects to a standard luer-slip syringe. The irrigation port must be situated at the patient end of the system for catheters and at the suction end of the catheter for catheters.made from medically safe PVC The suction end of the system must be fitted with capped, lockable thumb valve to control suctioning. It must have a tapered connection to fit a range of suction tubing. The system is for single patient use for a specified 48 hours Must have coloured day sticker labels to facilitate regular changing of the catheter system. Made from medically safe PVC. Sterile and individually packed in easy to open pouch with one transparent side; For SINGLE use only The system must conform to ISO 8836 Information brochure or the outer package of the item should clearly indicate latex free | Closed Suction Catheter - Adult Endotracheal 48Hr 12Fr OD: 4.0mm Length range: 520 - 600mm | Offer 1 | Each | Quality Medical Supplies (PTY)Ltd | 85 | 27.8447 | Bonree | Bonree Medical Co. Ltd | China
RT13-52-015 | CLOSED SUCTION CATHETER | CLOSED SUCTION CATHETER ADULT 48HR | CLOSED SUCTION CATHETER ADULT ENDOTRACHEAL - 48HRS Purpose: Closed system for respiratory suctioning through endotracheal tubes without disconnecting or interrupting ventilation. Consists of a closed suction catheter with an atraumatic tip. The catheter must have markings along the length to indicate depth of suctioning and a radiopaque line on the catheter tip. While connected but not in suction mode the catheter tip must not extend beyond the sheath or impede airflow. The suction catheter must be enclosed within a durable, flexible, plastic sheath for storage to avoid contamination. The transparent sheath must allow visualisation of markings on the suction catheter and must connect to the endotracheal tube via a clear, firmly-fitting, plastic connector or adaptor. The sheath and connector must allow for smooth advancement of catheter. The system must seal the airway and catheter from atmospheric contamination and allow uninterrupted positive pressure ventilation throughout the suction process. An effective seal around the catheter must prevent ventilation gas entering the lumen of the sheath. The system must protect the user from contamination by secretions and suctioned material. Proximal end of suction catheter must have clear chamber, so that suctioned contents can be visualised and must have lockable suction valve with capped suction tubing connector. The suction catheter must connect onto the tube with a DOUBLE swivel connector or a double-swivel Y-connector with standard 15 mm OD taper fittings. This connection must be secure and accidental disconnection must be minimal The whole device must be as lightweight as possible. Must include a light, straight, flexible, clear, concertina catheter-mount (15x22 mm connections) tube for connection to the breathing circuit The connector must allow variations in patient positions without causing torque, twisting, or obstruction of the tracheal tube, and must not cause tissue or skin pressure. The system must allow for irrigation via a lateral irrigation port which must be a capped with a one-way valve that connects to a standard luer-slip syringe. The irrigation port must be situated at the patient end of the system for catheters and at the suction end of the catheter for catheters.made from medically safe PVC The suction end of the system must be fitted with capped, lockable thumb valve to control suctioning. It must have a tapered connection to fit a range of suction tubing. The system is for single patient use for a specified 48 hours Must have coloured day sticker labels to facilitate regular changing of the catheter system. Made from medically safe PVC. Sterile and individually packed in easy to open pouch with one transparent side; For SINGLE use only The system must conform to ISO 8836 Information brochure or the outer package of the item should clearly indicate latex free | Closed Suction Catheter - Adult Endotracheal 48Hr 14Fr OD: 4.7mm Length range: 520 - 600mm | Offer 1 | Each | Quality Medical Supplies (PTY)Ltd | 85 | 37.8954 | Bonree | Bonree Medical Co. Ltd | China
RT13-52-016 | CLOSED SUCTION CATHETER | CLOSED SUCTION CATHETER ADULT 48HR | CLOSED SUCTION CATHETER ADULT ENDOTRACHEAL - 48HRS Purpose: Closed system for respiratory suctioning through endotracheal tubes without disconnecting or interrupting ventilation. Consists of a closed suction catheter with an atraumatic tip. The catheter must have markings along the length to indicate depth of suctioning and a radiopaque line on the catheter tip. While connected but not in suction mode the catheter tip must not extend beyond the sheath or impede airflow. The suction catheter must be enclosed within a durable, flexible, plastic sheath for storage to avoid contamination. The transparent sheath must allow visualisation of markings on the suction catheter and must connect to the endotracheal tube via a clear, firmly-fitting, plastic connector or adaptor. The sheath and connector must allow for smooth advancement of catheter. The system must seal the airway and catheter from atmospheric contamination and allow uninterrupted positive pressure ventilation throughout the suction process. An effective seal around the catheter must prevent ventilation gas entering the lumen of the sheath. The system must protect the user from contamination by secretions and suctioned material. Proximal end of suction catheter must have clear chamber, so that suctioned contents can be visualised and must have lockable suction valve with capped suction tubing connector. The suction catheter must connect onto the tube with a DOUBLE swivel connector or a double-swivel Y-connector with standard 15 mm OD taper fittings. This connection must be secure and accidental disconnection must be minimal The whole device must be as lightweight as possible. Must include a light, straight, flexible, clear, concertina catheter-mount (15x22 mm connections) tube for connection to the breathing circuit The connector must allow variations in patient positions without causing torque, twisting, or obstruction of the tracheal tube, and must not cause tissue or skin pressure. The system must allow for irrigation via a lateral irrigation port which must be a capped with a one-way valve that connects to a standard luer-slip syringe. The irrigation port must be situated at the patient end of the system for catheters and at the suction end of the catheter for catheters.made from medically safe PVC The suction end of the system must be fitted with capped, lockable thumb valve to control suctioning. It must have a tapered connection to fit a range of suction tubing. The system is for single patient use for a specified 48 hours Must have coloured day sticker labels to facilitate regular changing of the catheter system. Made from medically safe PVC. Sterile and individually packed in easy to open pouch with one transparent side; For SINGLE use only The system must conform to ISO 8836 Information brochure or the outer package of the item should clearly indicate latex free | Closed Suction Catheter - Adult Endotracheal 48Hr 16Fr OD: 5.3mm Length range: 520 - 600mm | Offer 1 | Each | Quality Medical Supplies (PTY)Ltd | 85 | 27.8447 | Bonree | Bonree Medical Co. Ltd | China
RT13-52-032 | CLOSED SUCTION CATHETER | CLOSED SUCTION CATHETER ADULT TRACHEOSTOMY | CLOSED SUCTION CATHETER ADULT TRACHEOSTOMY - 48HRS Purpose: Closed system for respiratory suctioning through tracheostomy tubes without disconnecting or interrupting ventilation. Consists of a closed suction catheter with an atraumatic tip. The catheter must have markings along the length to indicate depth of suctioning and a radiopaque line on the catheter tip. While connected but not in suction mode the catheter tip must not extend beyond the sheath or impede airflow. The suction catheter must be enclosed within a durable, flexible, plastic sheath for storage to avoID:contamination. The transparent sheath must allow visualisation of markings on the suction catheter and must connect to the tracheal tube via a clear, firmly-fitting, plastic connector or adaptor. The sheath and connector must allow for smooth advancement of catheter. The system must seal the airway and catheter from atmospheric contamination and allow uninterrupted positive pressure ventilation throughout the suction process. An effective seal around the catheter must prevent ventilation gas entering the lumen of the sheath. The system must protect the user from contamination by secretions and suctioned material. Proximal end of suction catheter must have clear chamber, so that suctioned contents can be visualised and must have lockable suction valve with capped suction tubing connector. The suction catheter must connect onto the tube with a DOUBLE swivel connector or a double-swivel Y-connector with standard 15 mm OD taper fittings. This connection must be secure and accidental disconnection must be minimal The whole device must be as lightweight as possible. Must include a light, straight, flexible, clear, concertina catheter-mount (15x22 mm connections) tube for connection to the breathing circuit The connector must allow variations in patient positions without causing torque, twisting, or obstruction of the tracheal tube, and must not cause tissue or skin pressure. The system must allow for irrigation via a lateral irrigation port which must be a capped with a one-way valve that connects to a standard luer-slip syringe. The irrigation port must be situated at the patient end of the system for catheters and at the suction end of the catheter for catheters.made from medically safe PVC The suction end of the system must be fitted with capped, lockable thumb valve to control suctioning. It must have a tapered connection to fit a range of suction tubing. The system is for single patient use for a specified 48 hours. Must have coloured day sticker labels to facilitate regular changing of the catheter system. Made from medically safe PVC. Sterile and individually packed in easy to open pouch with one transparent side; For SINGLE use only The system must conform to ISO 8836 Information brochure or the outer package of the item should clearly indicate latex free | Closed Suction Catheter - Adult Tracheostomy 48Hr 12Fr OD: 4.0mm Length range: 300 - 400mm | Offer 1 | Each | Quality Medical Supplies (PTY)Ltd | 85 | 27.8447 | Bonree | Bonree Medical Co. Ltd | China
RT13-52-033 | CLOSED SUCTION CATHETER | CLOSED SUCTION CATHETER ADULT TRACHEOSTOMY | CLOSED SUCTION CATHETER ADULT TRACHEOSTOMY - 48HRS Purpose: Closed system for respiratory suctioning through tracheostomy tubes without disconnecting or interrupting ventilation. Consists of a closed suction catheter with an atraumatic tip. The catheter must have markings along the length to indicate depth of suctioning and a radiopaque line on the catheter tip. While connected but not in suction mode the catheter tip must not extend beyond the sheath or impede airflow. The suction catheter must be enclosed within a durable, flexible, plastic sheath for storage to avoID:contamination. The transparent sheath must allow visualisation of markings on the suction catheter and must connect to the tracheal tube via a clear, firmly-fitting, plastic connector or adaptor. The sheath and connector must allow for smooth advancement of catheter. The system must seal the airway and catheter from atmospheric contamination and allow uninterrupted positive pressure ventilation throughout the suction process. An effective seal around the catheter must prevent ventilation gas entering the lumen of the sheath. The system must protect the user from contamination by secretions and suctioned material. Proximal end of suction catheter must have clear chamber, so that suctioned contents can be visualised and must have lockable suction valve with capped suction tubing connector. The suction catheter must connect onto the tube with a DOUBLE swivel connector or a double-swivel Y-connector with standard 15 mm OD taper fittings. This connection must be secure and accidental disconnection must be minimal The whole device must be as lightweight as possible. Must include a light, straight, flexible, clear, concertina catheter-mount (15x22 mm connections) tube for connection to the breathing circuit The connector must allow variations in patient positions without causing torque, twisting, or obstruction of the tracheal tube, and must not cause tissue or skin pressure. The system must allow for irrigation via a lateral irrigation port which must be a capped with a one-way valve that connects to a standard luer-slip syringe. The irrigation port must be situated at the patient end of the system for catheters and at the suction end of the catheter for catheters.made from medically safe PVC The suction end of the system must be fitted with capped, lockable thumb valve to control suctioning. It must have a tapered connection to fit a range of suction tubing. The system is for single patient use for a specified 48 hours. Must have coloured day sticker labels to facilitate regular changing of the catheter system. Made from medically safe PVC. Sterile and individually packed in easy to open pouch with one transparent side; For SINGLE use only The system must conform to ISO 8836 Information brochure or the outer package of the item should clearly indicate latex free | Closed Suction Catheter - Adult Tracheostomy 48Hr 14Fr OD: 4.7mm Length range: 300 - 400mm | Offer 1 | Each | Quality Medical Supplies (PTY)Ltd | 85 | 27.8447 | Bonree | Bonree Medical Co. Ltd | China
RT13-52-034 | CLOSED SUCTION CATHETER | CLOSED SUCTION CATHETER ADULT TRACHEOSTOMY | CLOSED SUCTION CATHETER ADULT TRACHEOSTOMY - 48HRS Purpose: Closed system for respiratory suctioning through tracheostomy tubes without disconnecting or interrupting ventilation. Consists of a closed suction catheter with an atraumatic tip. The catheter must have markings along the length to indicate depth of suctioning and a radiopaque line on the catheter tip. While connected but not in suction mode the catheter tip must not extend beyond the sheath or impede airflow. The suction catheter must be enclosed within a durable, flexible, plastic sheath for storage to avoID:contamination. The transparent sheath must allow visualisation of markings on the suction catheter and must connect to the tracheal tube via a clear, firmly-fitting, plastic connector or adaptor. The sheath and connector must allow for smooth advancement of catheter. The system must seal the airway and catheter from atmospheric contamination and allow uninterrupted positive pressure ventilation throughout the suction process. An effective seal around the catheter must prevent ventilation gas entering the lumen of the sheath. The system must protect the user from contamination by secretions and suctioned material. Proximal end of suction catheter must have clear chamber, so that suctioned contents can be visualised and must have lockable suction valve with capped suction tubing connector. The suction catheter must connect onto the tube with a DOUBLE swivel connector or a double-swivel Y-connector with standard 15 mm OD taper fittings. This connection must be secure and accidental disconnection must be minimal The whole device must be as lightweight as possible. Must include a light, straight, flexible, clear, concertina catheter-mount (15x22 mm connections) tube for connection to the breathing circuit The connector must allow variations in patient positions without causing torque, twisting, or obstruction of the tracheal tube, and must not cause tissue or skin pressure. The system must allow for irrigation via a lateral irrigation port which must be a capped with a one-way valve that connects to a standard luer-slip syringe. The irrigation port must be situated at the patient end of the system for catheters and at the suction end of the catheter for catheters.made from medically safe PVC The suction end of the system must be fitted with capped, lockable thumb valve to control suctioning. It must have a tapered connection to fit a range of suction tubing. The system is for single patient use for a specified 48 hours. Must have coloured day sticker labels to facilitate regular changing of the catheter system. Made from medically safe PVC. Sterile and individually packed in easy to open pouch with one transparent side; For SINGLE use only The system must conform to ISO 8836 Information brochure or the outer package of the item should clearly indicate latex free | Closed Suction Catheter - Adult Tracheostomy 48Hr 16Fr OD: 5.3mm Length range: 300 - 400mm | Offer 1 | Each | Quality Medical Supplies (PTY)Ltd | 85 | 27.8447 | Bonree | Bonree Medical Co. Ltd | China
RT13-06-009 | Oral Fibre-Optic Guide | Oral Fibre-Optic Guide | Oral Fibre-optic Guide - Ovassapian type Size 2 ORAL FIBRE OPTIC GUIDES - Ovassapian type Adult Purpose: Modified oral airways specifically designed for oral fibre-optic intubation. Anatomically shaped EVA or PVC airway with an appropriate atraumatic contour The edges or corners that come into contact with patient tissue, must be smooth One-piece design with a strong bite-block to protect the fibre-optic should the patient bite down Must split longitudinally to facilitate the placing of an endotracheal tube over the fibre optic scope All components must be latex-free, clinically clean and individually packed in peel pouch that is easy to open - For SINGLE use only Information brochure or the outer package of the item should clearly indicate latex free | Oral Fibre-optic Guide - Ovassapian type Size 2 | Offer 1 | Each | SSEM Mthembu Medical | 173.9835 | 92.22 | VBM | VBM | Germany
RT13-06-011 | Oral Fibre-Optic Guide | Oral Fibre-Optic Guide | Oral Fibre-optic Guide - Ovassapian type Size 4 ORAL FIBRE OPTIC GUIDES - Ovassapian type Adult Purpose: Modified oral airways specifically designed for oral fibre-optic intubation. Anatomically shaped EVA or PVC airway with an appropriate atraumatic contour The edges or corners that come into contact with patient tissue, must be smooth One-piece design with a strong bite-block to protect the fibre-optic should the patient bite down Must split longitudinally to facilitate the placing of an endotracheal tube over the fibre optic scope All components must be latex-free, clinically clean and individually packed in peel pouch that is easy to open - For SINGLE use only Information brochure or the outer package of the item should clearly indicate latex free | Oral Fibre-optic Guide - Ovassapian type Size 4 | Offer 1 | Each | SSEM Mthembu Medical | 173.9835 | 92.22 | VBM | VBM | Germany
RT13-36-004 | AIRWAY DEVICES | BLIND INSERTION AIRWAY | BLIND INSERTION AIRWAY DEVICES: DOUBLE-CUFFED Purpose: Dual lumen tube designed for difficult or emergency airway intubation and ventilation without the use of a laryngoscope Consists of a cuffed, DOUBLE -cuffed with a soft non-traumatic tip. High-volume proximal cuff for oro and nasopharyngeal occlusion. The tube must be held firmly in place by the proximal cuff and it must prevent the escape of gases through the nose or mouth Low-volume distal cuff for oesophageal occlusion Colour coded pilot balloons and 2 syringes prefilled with the correct amount of air for the inflation and deflation of the proximal and distal cuffs Bilateral ventilation ports between the two cuffs The tube must allow for rapid intubation of difficult airways and should be a viable airway whether it is placed in the trachea or in the oesophagus It should also allow for a suction catheter into the oesophagus when the distal cuff is placed in the oesophagus All components must be latex-free sterile and individually packed in peel pouch that is easy to open - For SINGLE use only Must comply with ISO 11712 Information brochure or the outer package of the item should clearly indicate latex free | Blind Insertion Airway Device: Single-Lumen Double-Cuffed Size: 1 Infant patient's: 5-12 kg | Offer 1 | Each | SSEM Mthembu Medical | 476.422 | 92.22 | VBM | VBM | Germany
RT13-36-005 | AIRWAY DEVICES | BLIND INSERTION AIRWAY | BLIND INSERTION AIRWAY DEVICES: DOUBLE-CUFFED Purpose: Dual lumen tube designed for difficult or emergency airway intubation and ventilation without the use of a laryngoscope Consists of a cuffed, DOUBLE -cuffed with a soft non-traumatic tip. High-volume proximal cuff for oro and nasopharyngeal occlusion. The tube must be held firmly in place by the proximal cuff and it must prevent the escape of gases through the nose or mouth Low-volume distal cuff for oesophageal occlusion Colour coded pilot balloons and 2 syringes prefilled with the correct amount of air for the inflation and deflation of the proximal and distal cuffs Bilateral ventilation ports between the two cuffs The tube must allow for rapid intubation of difficult airways and should be a viable airway whether it is placed in the trachea or in the oesophagus It should also allow for a suction catheter into the oesophagus when the distal cuff is placed in the oesophagus All components must be latex-free sterile and individually packed in peel pouch that is easy to open - For SINGLE use only Must comply with ISO 11712 Information brochure or the outer package of the item should clearly indicate latex free | Blind Insertion Airway Device: Single-Lumen Double-Cuffed Size: 2 Child patient's: 12-25 kg | Offer 1 | Each | SSEM Mthembu Medical | 476.422 | 92.22 | VBM | VBM | Germany
RT13-02-028 | Oxygen Mask | Oxygen Mask | VENTURI OXYGEN MASK - MRI COMPATIBLE Purpose: Delivery of oxygen and entrained room air to a spontaneously breathing patient. Suitable for the MRI Suite Modular design correctly sized and comfortable anatomically shaped preformed mask shaped to fit over the patient’s nose and mouth with under the chin cover for a better seal Atraumatic to the nose. 2 exhalation holes in the side of the mask, a secure retaining strap. All components to be MRI compatible. Flexible corrugated tube: Length 10 - 20cm with a 22mm connector Must have clear kink resistant, crush proof delivery tubing > 2 m Colour coded Venturi device with the Oxygen concentration in % and required flow rate must be moulded or clearly indicated on the device The components must be securely attached to each other and not separate accidentally in normal use. Manufactured from latex-free medical grade plastic - must comply with BS EN 13544-3:2001 standards Individually packed, easy to open without tubing getting tangled, clinically clean - SINGLE use Information brochure or the outer package of the item should clearly indicate latex free | Venturi Oxygen Mask MRI COMPATIBLE 28% - Adult ,Colour coded venturi: White | Offer 1 | Each | STAT Tiakeni Medical | 28.75 | -39.4482 | Flexicare 032-10-041CT | Flexicare Limited | China
RT13-02-029 | Oxygen Mask | Oxygen Mask | VENTURI OXYGEN MASK - MRI COMPATIBLE Purpose: Delivery of oxygen and entrained room air to a spontaneously breathing patient. Suitable for the MRI Suite Modular design correctly sized and comfortable anatomically shaped preformed mask shaped to fit over the patient’s nose and mouth with under the chin cover for a better seal Atraumatic to the nose. 2 exhalation holes in the side of the mask, a secure retaining strap. All components to be MRI compatible. Flexible corrugated tube: Length 10 - 20cm with a 22mm connector Must have clear kink resistant, crush proof delivery tubing > 2 m Colour coded Venturi device with the Oxygen concentration in % and required flow rate must be moulded or clearly indicated on the device The components must be securely attached to each other and not separate accidentally in normal use. Manufactured from latex-free medical grade plastic - must comply with BS EN 13544-3:2001 standards Individually packed, easy to open without tubing getting tangled, clinically clean - SINGLE use Information brochure or the outer package of the item should clearly indicate latex free | Venturi Oxygen Mask MRI COMPATIBLE 40% - Adult, Colour coded venturi: Red | Offer 1 | Each | STAT Tiakeni Medical | 28.75 | -39.4482 | Flexicare 032-10-043CT | Flexicare Limited | China
RT13-02-030 | Oxygen Mask | Oxygen Mask | VENTURI OXYGEN MASK - MRI COMPATIBLE Purpose: Delivery of oxygen and entrained room air to a spontaneously breathing patient. Suitable for the MRI Suite Modular design correctly sized and comfortable anatomically shaped preformed mask shaped to fit over the patient’s nose and mouth with under the chin cover for a better seal Atraumatic to the nose. 2 exhalation holes in the side of the mask, a secure retaining strap. All components to be MRI compatible. Flexible corrugated tube: Length 10 - 20cm with a 22mm connector Must have clear kink resistant, crush proof delivery tubing > 2 m Colour coded Venturi device with the Oxygen concentration in % and required flow rate must be moulded or clearly indicated on the device The components must be securely attached to each other and not separate accidentally in normal use. Manufactured from latex-free medical grade plastic - must comply with BS EN 13544-3:2001 standards Individually packed, easy to open without tubing getting tangled, clinically clean - SINGLE use Information brochure or the outer package of the item should clearly indicate latex free | Venturi Oxygen Mask MRI COMPATIBLE 60% - Adult ,Colour coded venturi: Green | Offer 1 | Each | STAT Tiakeni Medical | 28.75 | -39.4482 | Flexicare 032-10-044CT | Flexicare Limited | China
RT13-02-035 | Oxygen Mask | VENTURI OXYGEN MASK - MRI COMPATIBLE | VENTURI OXYGEN MASK - MRI COMPATIBLE 24% - Adult, Colour coded venturi: Blue Purpose: Delivery of oxygen and entrained room air to a spontaneously breathing patient. Suitable for the MRI Suite Modular design correctly sized and comfortable anatomically shaped preformed mask shaped to fit over the patient’s nose and mouth with under the chin cover for a better seal Atraumatic to the nose. 2 exhalation holes in the side of the mask, a secure retaining strap. All components to be MRI compatible. Must have clear kink resistant, crush proof delivery tubing > 2 m Colour coded Venturi device with the Oxygen concentration in % and required flow rate must be moulded or clearly indicated on the device The components must be securely attached to each other and not separate accidentally in normal use. Manufactured from latex-free medical grade plastic - must comply with BS EN 13544-3:2001 standards Individually packed, easy to open without tubing getting tangled, clinically clean - SINGLE use Information brochure or the outer package of the item should clearly indicate latex free | Venturi Oxygen Mask MRI COMPATIBLE 24% - Adult, Colour coded venturi: Blue | Offer 1 | Each | STAT Tiakeni Medical | 28.75 | -39.4482 | Flexicare 032-10-040CT | Flexicare Limited | China
RT13-02-036 | Oxygen Mask | VENTURI OXYGEN MASK - MRI COMPATIBLE | VENTURI OXYGEN MASK - MRI COMPATIBLE 28% - Adult, Colour coded venturi: White Purpose: Delivery of oxygen and entrained room air to a spontaneously breathing patient. Suitable for the MRI Suite Modular design correctly sized and comfortable anatomically shaped preformed mask shaped to fit over the patient’s nose and mouth with under the chin cover for a better seal Atraumatic to the nose. 2 exhalation holes in the side of the mask, a secure retaining strap. All components to be MRI compatible. Must have clear kink resistant, crush proof delivery tubing > 2 m Colour coded Venturi device with the Oxygen concentration in % and required flow rate must be moulded or clearly indicated on the device The components must be securely attached to each other and not separate accidentally in normal use. Manufactured from latex-free medical grade plastic - must comply with BS EN 13544-3:2001 standards Individually packed, easy to open without tubing getting tangled, clinically clean - SINGLE use Information brochure or the outer package of the item should clearly indicate latex free | Venturi Oxygen Mask MRI COMPATIBLE 28% - Adult, Colour coded venturi: White | Offer 1 | Each | STAT Tiakeni Medical | 28.75 | 91.14 | Flexicare 032-10-041CT | Flexicare Limited | China
RT13-02-037 | Oxygen Mask | VENTURI OXYGEN MASK - MRI COMPATIBLE | VENTURI OXYGEN MASK - MRI COMPATIBLE 40% - Adult, Colour coded venturi: Red Purpose: Delivery of oxygen and entrained room air to a spontaneously breathing patient. Suitable for the MRI Suite Modular design correctly sized and comfortable anatomically shaped preformed mask shaped to fit over the patient’s nose and mouth with under the chin cover for a better seal Atraumatic to the nose. 2 exhalation holes in the side of the mask, a secure retaining strap. All components to be MRI compatible. Must have clear kink resistant, crush proof delivery tubing > 2 m Colour coded Venturi device with the Oxygen concentration in % and required flow rate must be moulded or clearly indicated on the device The components must be securely attached to each other and not separate accidentally in normal use. Manufactured from latex-free medical grade plastic - must comply with BS EN 13544-3:2001 standards Individually packed, easy to open without tubing getting tangled, clinically clean - SINGLE use Information brochure or the outer package of the item should clearly indicate latex free | Venturi Oxygen Mask MRI COMPATIBLE 40% - Adult, Colour coded venturi: Red | Offer 1 | Each | STAT Tiakeni Medical | 28.75 | -94.126 | Flexicare 032-10-043CT | Flexicare Limited | China
RT13-02-038 | Oxygen Mask | VENTURI OXYGEN MASK - MRI COMPATIBLE | VENTURI OXYGEN MASK - MRI COMPATIBLE 60% - Adult ,Colour coded venturi: Green Purpose: Delivery of oxygen and entrained room air to a spontaneously breathing patient. Suitable for the MRI Suite Modular design correctly sized and comfortable anatomically shaped preformed mask shaped to fit over the patient’s nose and mouth with under the chin cover for a better seal Atraumatic to the nose. 2 exhalation holes in the side of the mask, a secure retaining strap. All components to be MRI compatible. Must have clear kink resistant, crush proof delivery tubing > 2 m Colour coded Venturi device with the Oxygen concentration in % and required flow rate must be moulded or clearly indicated on the device The components must be securely attached to each other and not separate accidentally in normal use. Manufactured from latex-free medical grade plastic - must comply with BS EN 13544-3:2001 standards Individually packed, easy to open without tubing getting tangled, clinically clean - SINGLE use Information brochure or the outer package of the item should clearly indicate latex free | Venturi Oxygen Mask MRI COMPATIBLE 60% - Adult ,Colour coded venturi: Green | Offer 1 | Each | STAT Tiakeni Medical | 28.75 | 91.14 | Flexicare 032-10-044CT | Flexicare Limited | China
RT13-42-013 | AIRWAY DEVICES | AIRWAY EXCHANGE CATHETER | AIRWAY EXCHANGE CATHETER Purpose: For simple, atraumatic endotracheal tube exchange whilst providing oxygenation Consists of a catheter manufactured from radiopaque medical grade plastic Must have a hollow lumen for constant oxygenation of the patient, with distal side ports to ensure adequate airflow Must be supplied with connection and adaptor for attachment of a ventilatory device. Must have a blunt tip for atraumatic intubation, centimetre markings for accurate tube placement. Latex-free, sterile and individually packed in peel pouch that is easy to open - For SINGLE use only Information brochure or the outer package of the item should clearly indicate latex free | Airway Exchange Catheter. Adult 19Fr Length: ± 830mm To fit ETT's Size: 7.0mm | Offer 1 503100-000060 | Each | Teleflex Medical | 850.99 | 90 | RUSCH | Teleflex | Germany
RT13-40-008 | Intubation Devices | INTUBATION STYLET | INTUBATION STYLET with metal core Purpose: To pre-form an endotracheal tube to a desired shape to aid intubation Consists of a malleable aluminium core coated and tightly bonded with high density polyurethane/ PVC / medical grade plastic/ polyethylene Must have a soft atraumatic tip Must have a U-bend proximally or other mechanism at proximal end to limit depth Latex-free clinically clean and individually packed in peel pouch that is easy to open - For SINGLE use only Information brochure or the outer package of the item should clearly indicate latex free | Intubation Stylet, Extra-Length Adult OD: 5.0mm Length: 365mm To fit ETT's > 7.5mm | Offer 1 | Each | Phoshi Investments | 17 | 27.8261 | STERIMED | STERIMED | INDIA
RT13-52-005 | CLOSED SUCTION CATHETER | CLOSED SUCTION CATHETER PAEDIATRIC 48HR | CLOSED SUCTION CATHETER - PAEDIATRIC ENDOTRACHEAL Y-TYPE 48HR Purpose: Closed system for respiratory suctioning through endotracheal tubes without disconnecting or interrupting ventilation. Consists of a closed suction catheter with an atraumatic tip. The catheter must have markings along the length to indicate depth of suctioning and a radiopaque line on the catheter tip. While connected but not in suction mode the catheter tip must not extend beyond the sheath or impede airflow. The suction catheter must be enclosed within a durable, flexible, plastic sheath for storage to avoid contamination. The transparent sheath must allow visualisation of markings on the suction catheter and must connect to the endotracheal tube via a clear, firmly-fitting, plastic connector or adaptor. The sheath and connector must allow for smooth advancement of catheter. The system must seal the airway and catheter from atmospheric contamination and allow uninterrupted positive pressure ventilation throughout the suction process. An effective seal around the catheter must prevent ventilation gas entering the lumen of the sheath. The system must protect the user from contamination by secretions and suctioned material. Proximal end of suction catheter must have clear chamber, so that suctioned contents can be visualised and must have lockable suction valve with toggled capped suction tubing connector. Paediatric Catheters must have a Y- connector (without locking mechanism) with an appropriate range of diameters for each size catheter. The connector must allow variations in patient positions without causing torque, twisting, or obstruction of the tracheal tube, and must not cause tissue or skin pressure. The system must allow for irrigation via a lateral irrigation port which must be a capped with a one-way valve that connects to a standard luer-slip syringe. The irrigation port must be situated at the patient end of the system for catheters and at the suction end of the catheter for catheters.made from medically safe PVC The suction end of the system must be fitted with capped, lockable thumb valve to control suctioning. It must have a tapered connection to fit a range of suction tubing. The system must be as light in weight as possible. The system is for single patient use for a specified 48 or 72 hours. Must have coloured day sticker labels to facilitate regular changing of the catheter system. Made from medically safe PVC. Sterile and individually packed in easy to open pouch with one transparent side; For SINGLE use only The system must conform to ISO 8836:2019. Information brochure or the outer package of the item should clearly indicate latex free | Closed Suction Catheter - Paediatric Endotracheal Y-Type 48Hr 10Fr OD: 3.3mm For ETT with ID: 5.0mm Y connectors: 4.5mm, 5mm, 5.5mm | Offer 1 | Each | Promed Technologies | 53 | 94 | PROCARE | Hangzhou Demo Medical Technology Co., Ltd. | China
RT13-52-013 | CLOSED SUCTION CATHETER | CLOSED SUCTION CATHETER ADULT 48HR | CLOSED SUCTION CATHETER ADULT ENDOTRACHEAL - 48HRS Purpose: Closed system for respiratory suctioning through endotracheal tubes without disconnecting or interrupting ventilation. Consists of a closed suction catheter with an atraumatic tip. The catheter must have markings along the length to indicate depth of suctioning and a radiopaque line on the catheter tip. While connected but not in suction mode the catheter tip must not extend beyond the sheath or impede airflow. The suction catheter must be enclosed within a durable, flexible, plastic sheath for storage to avoid contamination. The transparent sheath must allow visualisation of markings on the suction catheter and must connect to the endotracheal tube via a clear, firmly-fitting, plastic connector or adaptor. The sheath and connector must allow for smooth advancement of catheter. The system must seal the airway and catheter from atmospheric contamination and allow uninterrupted positive pressure ventilation throughout the suction process. An effective seal around the catheter must prevent ventilation gas entering the lumen of the sheath. The system must protect the user from contamination by secretions and suctioned material. Proximal end of suction catheter must have clear chamber, so that suctioned contents can be visualised and must have lockable suction valve with capped suction tubing connector. The suction catheter must connect onto the tube with a DOUBLE swivel connector or a double-swivel Y-connector with standard 15 mm OD taper fittings. This connection must be secure and accidental disconnection must be minimal The whole device must be as lightweight as possible. Must include a light, straight, flexible, clear, concertina catheter-mount (15x22 mm connections) tube for connection to the breathing circuit The connector must allow variations in patient positions without causing torque, twisting, or obstruction of the tracheal tube, and must not cause tissue or skin pressure. The system must allow for irrigation via a lateral irrigation port which must be a capped with a one-way valve that connects to a standard luer-slip syringe. The irrigation port must be situated at the patient end of the system for catheters and at the suction end of the catheter for catheters.made from medically safe PVC The suction end of the system must be fitted with capped, lockable thumb valve to control suctioning. It must have a tapered connection to fit a range of suction tubing. The system is for single patient use for a specified 48 hours Must have coloured day sticker labels to facilitate regular changing of the catheter system. Made from medically safe PVC. Sterile and individually packed in easy to open pouch with one transparent side; For SINGLE use only The system must conform to ISO 8836 Information brochure or the outer package of the item should clearly indicate latex free | Closed Suction Catheter - Adult Endotracheal 48Hr 10Fr OD: 3.3mm Length range: 520 - 600mm | Offer 1 | Each | Promed Technologies | 53 | 94 | PROCARE | Hangzhou Demo Medical Technology Co., Ltd. | China
RT13-52-014 | CLOSED SUCTION CATHETER | CLOSED SUCTION CATHETER ADULT 48HR | CLOSED SUCTION CATHETER ADULT ENDOTRACHEAL - 48HRS Purpose: Closed system for respiratory suctioning through endotracheal tubes without disconnecting or interrupting ventilation. Consists of a closed suction catheter with an atraumatic tip. The catheter must have markings along the length to indicate depth of suctioning and a radiopaque line on the catheter tip. While connected but not in suction mode the catheter tip must not extend beyond the sheath or impede airflow. The suction catheter must be enclosed within a durable, flexible, plastic sheath for storage to avoid contamination. The transparent sheath must allow visualisation of markings on the suction catheter and must connect to the endotracheal tube via a clear, firmly-fitting, plastic connector or adaptor. The sheath and connector must allow for smooth advancement of catheter. The system must seal the airway and catheter from atmospheric contamination and allow uninterrupted positive pressure ventilation throughout the suction process. An effective seal around the catheter must prevent ventilation gas entering the lumen of the sheath. The system must protect the user from contamination by secretions and suctioned material. Proximal end of suction catheter must have clear chamber, so that suctioned contents can be visualised and must have lockable suction valve with capped suction tubing connector. The suction catheter must connect onto the tube with a DOUBLE swivel connector or a double-swivel Y-connector with standard 15 mm OD taper fittings. This connection must be secure and accidental disconnection must be minimal The whole device must be as lightweight as possible. Must include a light, straight, flexible, clear, concertina catheter-mount (15x22 mm connections) tube for connection to the breathing circuit The connector must allow variations in patient positions without causing torque, twisting, or obstruction of the tracheal tube, and must not cause tissue or skin pressure. The system must allow for irrigation via a lateral irrigation port which must be a capped with a one-way valve that connects to a standard luer-slip syringe. The irrigation port must be situated at the patient end of the system for catheters and at the suction end of the catheter for catheters.made from medically safe PVC The suction end of the system must be fitted with capped, lockable thumb valve to control suctioning. It must have a tapered connection to fit a range of suction tubing. The system is for single patient use for a specified 48 hours Must have coloured day sticker labels to facilitate regular changing of the catheter system. Made from medically safe PVC. Sterile and individually packed in easy to open pouch with one transparent side; For SINGLE use only The system must conform to ISO 8836 Information brochure or the outer package of the item should clearly indicate latex free | Closed Suction Catheter - Adult Endotracheal 48Hr 12Fr OD: 4.0mm Length range: 520 - 600mm | Offer 1 | Each | Promed Technologies | 53 | 94 | PROCARE | Hangzhou Demo Medical Technology Co., Ltd. | China
RT13-52-015 | CLOSED SUCTION CATHETER | CLOSED SUCTION CATHETER ADULT 48HR | CLOSED SUCTION CATHETER ADULT ENDOTRACHEAL - 48HRS Purpose: Closed system for respiratory suctioning through endotracheal tubes without disconnecting or interrupting ventilation. Consists of a closed suction catheter with an atraumatic tip. The catheter must have markings along the length to indicate depth of suctioning and a radiopaque line on the catheter tip. While connected but not in suction mode the catheter tip must not extend beyond the sheath or impede airflow. The suction catheter must be enclosed within a durable, flexible, plastic sheath for storage to avoid contamination. The transparent sheath must allow visualisation of markings on the suction catheter and must connect to the endotracheal tube via a clear, firmly-fitting, plastic connector or adaptor. The sheath and connector must allow for smooth advancement of catheter. The system must seal the airway and catheter from atmospheric contamination and allow uninterrupted positive pressure ventilation throughout the suction process. An effective seal around the catheter must prevent ventilation gas entering the lumen of the sheath. The system must protect the user from contamination by secretions and suctioned material. Proximal end of suction catheter must have clear chamber, so that suctioned contents can be visualised and must have lockable suction valve with capped suction tubing connector. The suction catheter must connect onto the tube with a DOUBLE swivel connector or a double-swivel Y-connector with standard 15 mm OD taper fittings. This connection must be secure and accidental disconnection must be minimal The whole device must be as lightweight as possible. Must include a light, straight, flexible, clear, concertina catheter-mount (15x22 mm connections) tube for connection to the breathing circuit The connector must allow variations in patient positions without causing torque, twisting, or obstruction of the tracheal tube, and must not cause tissue or skin pressure. The system must allow for irrigation via a lateral irrigation port which must be a capped with a one-way valve that connects to a standard luer-slip syringe. The irrigation port must be situated at the patient end of the system for catheters and at the suction end of the catheter for catheters.made from medically safe PVC The suction end of the system must be fitted with capped, lockable thumb valve to control suctioning. It must have a tapered connection to fit a range of suction tubing. The system is for single patient use for a specified 48 hours Must have coloured day sticker labels to facilitate regular changing of the catheter system. Made from medically safe PVC. Sterile and individually packed in easy to open pouch with one transparent side; For SINGLE use only The system must conform to ISO 8836 Information brochure or the outer package of the item should clearly indicate latex free | Closed Suction Catheter - Adult Endotracheal 48Hr 14Fr OD: 4.7mm Length range: 520 - 600mm | Offer 1 | Each | Promed Technologies | 56.5 | 94 | PROCARE | Hangzhou Demo Medical Technology Co., Ltd. | China
RT13-52-016 | CLOSED SUCTION CATHETER | CLOSED SUCTION CATHETER ADULT 48HR | CLOSED SUCTION CATHETER ADULT ENDOTRACHEAL - 48HRS Purpose: Closed system for respiratory suctioning through endotracheal tubes without disconnecting or interrupting ventilation. Consists of a closed suction catheter with an atraumatic tip. The catheter must have markings along the length to indicate depth of suctioning and a radiopaque line on the catheter tip. While connected but not in suction mode the catheter tip must not extend beyond the sheath or impede airflow. The suction catheter must be enclosed within a durable, flexible, plastic sheath for storage to avoid contamination. The transparent sheath must allow visualisation of markings on the suction catheter and must connect to the endotracheal tube via a clear, firmly-fitting, plastic connector or adaptor. The sheath and connector must allow for smooth advancement of catheter. The system must seal the airway and catheter from atmospheric contamination and allow uninterrupted positive pressure ventilation throughout the suction process. An effective seal around the catheter must prevent ventilation gas entering the lumen of the sheath. The system must protect the user from contamination by secretions and suctioned material. Proximal end of suction catheter must have clear chamber, so that suctioned contents can be visualised and must have lockable suction valve with capped suction tubing connector. The suction catheter must connect onto the tube with a DOUBLE swivel connector or a double-swivel Y-connector with standard 15 mm OD taper fittings. This connection must be secure and accidental disconnection must be minimal The whole device must be as lightweight as possible. Must include a light, straight, flexible, clear, concertina catheter-mount (15x22 mm connections) tube for connection to the breathing circuit The connector must allow variations in patient positions without causing torque, twisting, or obstruction of the tracheal tube, and must not cause tissue or skin pressure. The system must allow for irrigation via a lateral irrigation port which must be a capped with a one-way valve that connects to a standard luer-slip syringe. The irrigation port must be situated at the patient end of the system for catheters and at the suction end of the catheter for catheters.made from medically safe PVC The suction end of the system must be fitted with capped, lockable thumb valve to control suctioning. It must have a tapered connection to fit a range of suction tubing. The system is for single patient use for a specified 48 hours Must have coloured day sticker labels to facilitate regular changing of the catheter system. Made from medically safe PVC. Sterile and individually packed in easy to open pouch with one transparent side; For SINGLE use only The system must conform to ISO 8836 Information brochure or the outer package of the item should clearly indicate latex free | Closed Suction Catheter - Adult Endotracheal 48Hr 16Fr OD: 5.3mm Length range: 520 - 600mm | Offer 1 | Each | Promed Technologies | 53 | 94 | PROCARE | Hangzhou Demo Medical Technology Co., Ltd. | China
RT13-52-031 | CLOSED SUCTION CATHETER | CLOSED SUCTION CATHETER ADULT TRACHEOSTOMY | CLOSED SUCTION CATHETER ADULT TRACHEOSTOMY - 48HRS Purpose: Closed system for respiratory suctioning through tracheostomy tubes without disconnecting or interrupting ventilation. Consists of a closed suction catheter with an atraumatic tip. The catheter must have markings along the length to indicate depth of suctioning and a radiopaque line on the catheter tip. While connected but not in suction mode the catheter tip must not extend beyond the sheath or impede airflow. The suction catheter must be enclosed within a durable, flexible, plastic sheath for storage to avoID:contamination. The transparent sheath must allow visualisation of markings on the suction catheter and must connect to the tracheal tube via a clear, firmly-fitting, plastic connector or adaptor. The sheath and connector must allow for smooth advancement of catheter. The system must seal the airway and catheter from atmospheric contamination and allow uninterrupted positive pressure ventilation throughout the suction process. An effective seal around the catheter must prevent ventilation gas entering the lumen of the sheath. The system must protect the user from contamination by secretions and suctioned material. Proximal end of suction catheter must have clear chamber, so that suctioned contents can be visualised and must have lockable suction valve with capped suction tubing connector. The suction catheter must connect onto the tube with a DOUBLE swivel connector or a double-swivel Y-connector with standard 15 mm OD taper fittings. This connection must be secure and accidental disconnection must be minimal The whole device must be as lightweight as possible. Must include a light, straight, flexible, clear, concertina catheter-mount (15x22 mm connections) tube for connection to the breathing circuit The connector must allow variations in patient positions without causing torque, twisting, or obstruction of the tracheal tube, and must not cause tissue or skin pressure. The system must allow for irrigation via a lateral irrigation port which must be a capped with a one-way valve that connects to a standard luer-slip syringe. The irrigation port must be situated at the patient end of the system for catheters and at the suction end of the catheter for catheters.made from medically safe PVC The suction end of the system must be fitted with capped, lockable thumb valve to control suctioning. It must have a tapered connection to fit a range of suction tubing. The system is for single patient use for a specified 48 hours. Must have coloured day sticker labels to facilitate regular changing of the catheter system. Made from medically safe PVC. Sterile and individually packed in easy to open pouch with one transparent side; For SINGLE use only The system must conform to ISO 8836 Information brochure or the outer package of the item should clearly indicate latex free | Closed Suction Catheter - Adult Tracheostomy 48Hr 10Fr OD: 3.3mm Length range: 300 - 400mm | Offer 1 | Each | Promed Technologies | 53 | 94 | PROCARE | Hangzhou Demo Medical Technology Co., Ltd. | China
RT13-52-032 | CLOSED SUCTION CATHETER | CLOSED SUCTION CATHETER ADULT TRACHEOSTOMY | CLOSED SUCTION CATHETER ADULT TRACHEOSTOMY - 48HRS Purpose: Closed system for respiratory suctioning through tracheostomy tubes without disconnecting or interrupting ventilation. Consists of a closed suction catheter with an atraumatic tip. The catheter must have markings along the length to indicate depth of suctioning and a radiopaque line on the catheter tip. While connected but not in suction mode the catheter tip must not extend beyond the sheath or impede airflow. The suction catheter must be enclosed within a durable, flexible, plastic sheath for storage to avoID:contamination. The transparent sheath must allow visualisation of markings on the suction catheter and must connect to the tracheal tube via a clear, firmly-fitting, plastic connector or adaptor. The sheath and connector must allow for smooth advancement of catheter. The system must seal the airway and catheter from atmospheric contamination and allow uninterrupted positive pressure ventilation throughout the suction process. An effective seal around the catheter must prevent ventilation gas entering the lumen of the sheath. The system must protect the user from contamination by secretions and suctioned material. Proximal end of suction catheter must have clear chamber, so that suctioned contents can be visualised and must have lockable suction valve with capped suction tubing connector. The suction catheter must connect onto the tube with a DOUBLE swivel connector or a double-swivel Y-connector with standard 15 mm OD taper fittings. This connection must be secure and accidental disconnection must be minimal The whole device must be as lightweight as possible. Must include a light, straight, flexible, clear, concertina catheter-mount (15x22 mm connections) tube for connection to the breathing circuit The connector must allow variations in patient positions without causing torque, twisting, or obstruction of the tracheal tube, and must not cause tissue or skin pressure. The system must allow for irrigation via a lateral irrigation port which must be a capped with a one-way valve that connects to a standard luer-slip syringe. The irrigation port must be situated at the patient end of the system for catheters and at the suction end of the catheter for catheters.made from medically safe PVC The suction end of the system must be fitted with capped, lockable thumb valve to control suctioning. It must have a tapered connection to fit a range of suction tubing. The system is for single patient use for a specified 48 hours. Must have coloured day sticker labels to facilitate regular changing of the catheter system. Made from medically safe PVC. Sterile and individually packed in easy to open pouch with one transparent side; For SINGLE use only The system must conform to ISO 8836 Information brochure or the outer package of the item should clearly indicate latex free | Closed Suction Catheter - Adult Tracheostomy 48Hr 12Fr OD: 4.0mm Length range: 300 - 400mm | Offer 1 | Each | Promed Technologies | 53 | 94 | PROCARE | Hangzhou Demo Medical Technology Co., Ltd. | China
RT13-52-033 | CLOSED SUCTION CATHETER | CLOSED SUCTION CATHETER ADULT TRACHEOSTOMY | CLOSED SUCTION CATHETER ADULT TRACHEOSTOMY - 48HRS Purpose: Closed system for respiratory suctioning through tracheostomy tubes without disconnecting or interrupting ventilation. Consists of a closed suction catheter with an atraumatic tip. The catheter must have markings along the length to indicate depth of suctioning and a radiopaque line on the catheter tip. While connected but not in suction mode the catheter tip must not extend beyond the sheath or impede airflow. The suction catheter must be enclosed within a durable, flexible, plastic sheath for storage to avoID:contamination. The transparent sheath must allow visualisation of markings on the suction catheter and must connect to the tracheal tube via a clear, firmly-fitting, plastic connector or adaptor. The sheath and connector must allow for smooth advancement of catheter. The system must seal the airway and catheter from atmospheric contamination and allow uninterrupted positive pressure ventilation throughout the suction process. An effective seal around the catheter must prevent ventilation gas entering the lumen of the sheath. The system must protect the user from contamination by secretions and suctioned material. Proximal end of suction catheter must have clear chamber, so that suctioned contents can be visualised and must have lockable suction valve with capped suction tubing connector. The suction catheter must connect onto the tube with a DOUBLE swivel connector or a double-swivel Y-connector with standard 15 mm OD taper fittings. This connection must be secure and accidental disconnection must be minimal The whole device must be as lightweight as possible. Must include a light, straight, flexible, clear, concertina catheter-mount (15x22 mm connections) tube for connection to the breathing circuit The connector must allow variations in patient positions without causing torque, twisting, or obstruction of the tracheal tube, and must not cause tissue or skin pressure. The system must allow for irrigation via a lateral irrigation port which must be a capped with a one-way valve that connects to a standard luer-slip syringe. The irrigation port must be situated at the patient end of the system for catheters and at the suction end of the catheter for catheters.made from medically safe PVC The suction end of the system must be fitted with capped, lockable thumb valve to control suctioning. It must have a tapered connection to fit a range of suction tubing. The system is for single patient use for a specified 48 hours. Must have coloured day sticker labels to facilitate regular changing of the catheter system. Made from medically safe PVC. Sterile and individually packed in easy to open pouch with one transparent side; For SINGLE use only The system must conform to ISO 8836 Information brochure or the outer package of the item should clearly indicate latex free | Closed Suction Catheter - Adult Tracheostomy 48Hr 14Fr OD: 4.7mm Length range: 300 - 400mm | Offer 1 | Each | Promed Technologies | 53 | 94 | PROCARE | Hangzhou Demo Medical Technology Co., Ltd. | China
RT13-52-034 | CLOSED SUCTION CATHETER | CLOSED SUCTION CATHETER ADULT TRACHEOSTOMY | CLOSED SUCTION CATHETER ADULT TRACHEOSTOMY - 48HRS Purpose: Closed system for respiratory suctioning through tracheostomy tubes without disconnecting or interrupting ventilation. Consists of a closed suction catheter with an atraumatic tip. The catheter must have markings along the length to indicate depth of suctioning and a radiopaque line on the catheter tip. While connected but not in suction mode the catheter tip must not extend beyond the sheath or impede airflow. The suction catheter must be enclosed within a durable, flexible, plastic sheath for storage to avoID:contamination. The transparent sheath must allow visualisation of markings on the suction catheter and must connect to the tracheal tube via a clear, firmly-fitting, plastic connector or adaptor. The sheath and connector must allow for smooth advancement of catheter. The system must seal the airway and catheter from atmospheric contamination and allow uninterrupted positive pressure ventilation throughout the suction process. An effective seal around the catheter must prevent ventilation gas entering the lumen of the sheath. The system must protect the user from contamination by secretions and suctioned material. Proximal end of suction catheter must have clear chamber, so that suctioned contents can be visualised and must have lockable suction valve with capped suction tubing connector. The suction catheter must connect onto the tube with a DOUBLE swivel connector or a double-swivel Y-connector with standard 15 mm OD taper fittings. This connection must be secure and accidental disconnection must be minimal The whole device must be as lightweight as possible. Must include a light, straight, flexible, clear, concertina catheter-mount (15x22 mm connections) tube for connection to the breathing circuit The connector must allow variations in patient positions without causing torque, twisting, or obstruction of the tracheal tube, and must not cause tissue or skin pressure. The system must allow for irrigation via a lateral irrigation port which must be a capped with a one-way valve that connects to a standard luer-slip syringe. The irrigation port must be situated at the patient end of the system for catheters and at the suction end of the catheter for catheters.made from medically safe PVC The suction end of the system must be fitted with capped, lockable thumb valve to control suctioning. It must have a tapered connection to fit a range of suction tubing. The system is for single patient use for a specified 48 hours. Must have coloured day sticker labels to facilitate regular changing of the catheter system. Made from medically safe PVC. Sterile and individually packed in easy to open pouch with one transparent side; For SINGLE use only The system must conform to ISO 8836 Information brochure or the outer package of the item should clearly indicate latex free | Closed Suction Catheter - Adult Tracheostomy 48Hr 16Fr OD: 5.3mm Length range: 300 - 400mm | Offer 1 | Each | Promed Technologies | 53 | 94 | PROCARE | Hangzhou Demo Medical Technology Co., Ltd. | China
RT13-57-002 | BREATHING CIRCUITS | CORRUGATED BREATHING CIRCUITS WITH EXTRA LIMB AND BAG | CORRUGATED BREATHING CIRCUITS WITH EXTRA LIMB AND BAG - SINGLE USE Purpose: Used to deliver oxygen and Volatile Gases to patient's and to eliminate C02 during anaesthesia Transparent circuit with low resistance, non-extendable, corrugated tubing that is non-kinking, lightweight and crush resistant. The tubing must be pliable and must not retain a shape. Uninterrupted tubing (not joined). Shallow corrugations to minimise flow resistance and turbulence. Must have an extra limb with an anti-static reservoir bag that gives good feedback regarding patient lung compliance. The reservoir bag must not twist on itself and must have a mechanism to prevent this. Connectors Malleable 22mm female connectors bonded to the ends of inspiratory and expiratory limbs – must fit securely to the machine. Spare 22mm M/M connector. Y-piece at patient's end. Separate elbow-piece with a toggled cap on a luer lock port. Good quality to ensure no leakage or easy disconnections Latex-free, non-pyrogenic, hypoallergenic, non-toxic and non-conductive medical grade plastic No variation in Circuit limb length or Extra limb will be accepted Clinically clean and individually packed, For SINGLE use only To comply with ISO 5367 Information brochure or the outer package of the item should clearly indicate latex free | Corrugated breathing circuit with extra limb and breathing bag - Single Use - Neonatal 15mm Circuit: 1.6m; 15mm Extra limb: 1.8m Reservoir bag: 500ml | Offer 1 | Each | Surgicare Medical | 94 | 94.25 | SURGICARE | Ningbo Haukun Medical Equipment | China
RT13-57-004 | BREATHING CIRCUITS | CORRUGATED BREATHING CIRCUITS WITH EXTRA LIMB AND BAG | CORRUGATED BREATHING CIRCUITS WITH EXTRA LIMB AND BAG - SINGLE USE Purpose: Used to deliver oxygen and Volatile Gases to patient's and to eliminate C02 during anaesthesia Transparent circuit with low resistance, non-extendable, corrugated tubing that is non-kinking, lightweight and crush resistant. The tubing must be pliable and must not retain a shape. Uninterrupted tubing (not joined). Shallow corrugations to minimise flow resistance and turbulence. Must have an extra limb with an anti-static reservoir bag that gives good feedback regarding patient lung compliance. The reservoir bag must not twist on itself and must have a mechanism to prevent this. Connectors Malleable 22mm female connectors bonded to the ends of inspiratory and expiratory limbs – must fit securely to the machine. Spare 22mm M/M connector. Y-piece at patient's end. Separate elbow-piece with a toggled cap on a luer lock port. Good quality to ensure no leakage or easy disconnections Latex-free, non-pyrogenic, hypoallergenic, non-toxic and non-conductive medical grade plastic No variation in Circuit limb length or Extra limb will be accepted Clinically clean and individually packed, For SINGLE use only To comply with ISO 5367 Information brochure or the outer package of the item should clearly indicate latex free | Corrugated breathing circuit with extra limb and breathing bag - Single Use - Paediatric > 3 yrs 15mm Circuit: 1.6m; 15mm Extra limb: 1.8m Reservoir bag: 1l | Offer 1 | Each | Surgicare Medical | 94 | 94.25 | SURGICARE | Ningbo Haukun Medical Equipment | China
RT13-57-012 | BREATHING CIRCUITS | CORRUGATED BREATHING CIRCUITS WITH WATER TRAPS AND EXTRA LIMB AND BAG-SINGLE USE | SINGLE CORRUGATED BREATHING TUBING - SINGLE USE Purpose: Used as extension or extra limb in breathing circuits Transparent circuit with low resistance, non-extendable, corrugated tubing that is non-kinking, lightweight and crush resistant. The tubing must be pliable and must not retain a shape Shallow corrugations to minimise flow resistance and turbulence. Connectors Malleable 22mm female connectors bonded to the end of the tubing Spare connectors to be provided Latex-free, non-pyrogenic, hypoallergenic, non-toxic and non-conductive medical grade plastic No variation in Circuit limb length or Extra limb will be accepted Clinically clean and individually packed, For SINGLE use only To comply with ISO 5367 Information brochure or the outer package of the item should clearly indicate latex free | Corrugated breathing circuit with water traps and extra limb with breathing bag - Single use Paediatric > 3 yrs 15mm Circuit: 1.6m; 15mm Extra limb: 1.8m Reservoir bag: 1l | Offer 1 | Each | Surgicare Medical | 110.08 | 94.25 | SURGICARE | Ningbo Haukun Medical Equipment | China
RT13-57-013 | BREATHING CIRCUITS | CORRUGATED BREATHING CIRCUITS WITH WATER TRAPS AND EXTRA LIMB AND BAG-SINGLE USE | SINGLE CORRUGATED BREATHING TUBING - SINGLE USE Purpose: Used as extension or extra limb in breathing circuits Transparent circuit with low resistance, non-extendable, corrugated tubing that is non-kinking, lightweight and crush resistant. The tubing must be pliable and must not retain a shape Shallow corrugations to minimise flow resistance and turbulence. Connectors Malleable 22mm female connectors bonded to the end of the tubing Spare connectors to be provided Latex-free, non-pyrogenic, hypoallergenic, non-toxic and non-conductive medical grade plastic No variation in Circuit limb length or Extra limb will be accepted Clinically clean and individually packed, For SINGLE use only To comply with ISO 5367 Information brochure or the outer package of the item should clearly indicate latex free | Corrugated breathing circuit with water traps and extra limb with breathing bag - Single Use Adult 22mm Circuit: 1.8m; 22mm Extra limb: 1.2m OD: 22mm Reservoir bag: 2l | Offer 1 | Each | Surgicare Medical | 110.08 | 94.25 | SURGICARE | Ningbo Haukun Medical Equipment | China
RT13-57-016 | BREATHING CIRCUITS | CORRUGATED BREATHING CIRCUITS WITH WATER TRAPS AND EXTRA LIMB AND BAG-SINGLE USE | SINGLE CORRUGATED BREATHING TUBING - SINGLE USE Purpose: Used as extension or extra limb in breathing circuits Transparent circuit with low resistance, non-extendable, corrugated tubing that is non-kinking, lightweight and crush resistant. The tubing must be pliable and must not retain a shape Shallow corrugations to minimise flow resistance and turbulence. Connectors Malleable 22mm female connectors bonded to the end of the tubing Spare connectors to be provided Latex-free, non-pyrogenic, hypoallergenic, non-toxic and non-conductive medical grade plastic No variation in Circuit limb length or Extra limb will be accepted Clinically clean and individually packed, For SINGLE use only To comply with ISO 5367 Information brochure or the outer package of the item should clearly indicate latex free | Corrugated breathing circuit with water traps and extra limb with breathing bag - Single Use Adult Long 22mm Circuit: 2.5m; 22mm Extra limb: 1.2m OD: 22mm Reservoir bag: 2l | Offer 1 | Each | Surgicare Medical | 124.86 | 94.25 | SURGICARE | Ningbo Haukun Medical Equipment | China
RT13-57-017 | BREATHING CIRCUITS | CORRUGATED BREATHING CIRCUITS WITH WATER TRAPS AND EXTRA LIMB AND BAG-SINGLE USE | SINGLE CORRUGATED BREATHING TUBING - SINGLE USE Purpose: Used as extension or extra limb in breathing circuits Transparent circuit with low resistance, non-extendable, corrugated tubing that is non-kinking, lightweight and crush resistant. The tubing must be pliable and must not retain a shape Shallow corrugations to minimise flow resistance and turbulence. Connectors Malleable 22mm female connectors bonded to the end of the tubing Spare connectors to be provided Latex-free, non-pyrogenic, hypoallergenic, non-toxic and non-conductive medical grade plastic No variation in Circuit limb length or Extra limb will be accepted Clinically clean and individually packed, For SINGLE use only To comply with ISO 5367 Information brochure or the outer package of the item should clearly indicate latex free | Corrugated breathing circuit with water traps and extra limb with breathing bag - Single Use Adult Long 22mm Circuit: 2.5m; 22mm Extra limb: 1.5m OD: 22mm Reservoir bag: 2l | Offer 1 | Each | Surgicare Medical | 124.86 | -383.5395 | SURGICARE | Ningbo Haukun Medical Equipment | China
RT13-57-018 | BREATHING CIRCUITS | CORRUGATED BREATHING CIRCUITS WITH WATER TRAPS AND EXTRA LIMB AND BAG-SINGLE USE | SINGLE CORRUGATED BREATHING TUBING - SINGLE USE Purpose: Used as extension or extra limb in breathing circuits Transparent circuit with low resistance, non-extendable, corrugated tubing that is non-kinking, lightweight and crush resistant. The tubing must be pliable and must not retain a shape Shallow corrugations to minimise flow resistance and turbulence. Connectors Malleable 22mm female connectors bonded to the end of the tubing Spare connectors to be provided Latex-free, non-pyrogenic, hypoallergenic, non-toxic and non-conductive medical grade plastic No variation in Circuit limb length or Extra limb will be accepted Clinically clean and individually packed, For SINGLE use only To comply with ISO 5367 Information brochure or the outer package of the item should clearly indicate latex free | Corrugated breathing circuit with water traps and extra limb with breathing bag - Single Use Adult Long 22mm Circuit: 2.5m; 22mm Extra limb: 1.8m OD: 22mm Reservoir bag: 2l | Offer 1 | Each | Surgicare Medical | 124.86 | 94.25 | SURGICARE | Ningbo Haukun Medical Equipment | China
RT13-60-001 | BREATHING CIRCUITS | SMOOTH BORE BREATHING CIRCUITS WITH EXTRA LIMB AND BREATHING BAG - SINGLE USE | SMOOTH BORE BREATHING CIRCUITS WITH EXTRA LIMB AND BREATHING BAG - SINGLE USE Purpose: Used to deliver oxygen and Volatile Gases to patient's and to eliminate C02 during anaesthesia Transparent circuit with low resistance, non-extendable, smooth-bore tubing that is non-kinking, lightweight and crush resistant. The tubing must be pliable and must not retain a shape Spiral outer casing that supports the inner lumen to maintain tube integrity. Must have an extra limb with an anti-static reservoir bag that gives good feedback regarding patient lung compliance. The reservoir bag must not twist on itself and must have a mechanism to prevent this. Connectors Malleable 22mm female connectors bonded to the ends of inspiratory and expiratory limbs – must fit securely to the machine. Spare 22mm M/M connector. Y-piece at patient's end. Separate elbow-piece with capped luer lock port. Good quality to ensure no leakage or easy disconnections Latex-free, non-pyrogenic, hypoallergenic, non-toxic and non-conductive medical grade plastic No variation in Circuit limb length or Extra limb will be accepted Clinically clean and individually packed, For SINGLE use only To comply with ISO 5367 Information brochure or the outer package of the item should clearly indicate latex free | Smooth bore breathing circuit with extra limb and breathing bag - Single Use Paediatric > 3yrs 15mm Circuit: 1.6m; Extra limb: 1.2m Reservoir bag: 1l | Offer 1 | Each | Surgicare Medical | 135.86 | 94.25 | SURGICARE | Ningbo Haukun Medical Equipment | China
RT13-60-002 | BREATHING CIRCUITS | SMOOTH BORE BREATHING CIRCUITS WITH EXTRA LIMB AND BREATHING BAG - SINGLE USE | SMOOTH BORE BREATHING CIRCUITS WITH EXTRA LIMB AND BREATHING BAG - SINGLE USE Purpose: Used to deliver oxygen and Volatile Gases to patient's and to eliminate C02 during anaesthesia Transparent circuit with low resistance, non-extendable, smooth-bore tubing that is non-kinking, lightweight and crush resistant. The tubing must be pliable and must not retain a shape Spiral outer casing that supports the inner lumen to maintain tube integrity. Must have an extra limb with an anti-static reservoir bag that gives good feedback regarding patient lung compliance. The reservoir bag must not twist on itself and must have a mechanism to prevent this. Connectors Malleable 22mm female connectors bonded to the ends of inspiratory and expiratory limbs – must fit securely to the machine. Spare 22mm M/M connector. Y-piece at patient's end. Separate elbow-piece with capped luer lock port. Good quality to ensure no leakage or easy disconnections Latex-free, non-pyrogenic, hypoallergenic, non-toxic and non-conductive medical grade plastic No variation in Circuit limb length or Extra limb will be accepted Clinically clean and individually packed, For SINGLE use only To comply with ISO 5367 Information brochure or the outer package of the item should clearly indicate latex free | Smooth bore breathing circuit with extra limb and breathing bag - Single Use Paediatric > 3yrs 15mm Circuit: 1.6m; Extra limb: 1.8m Reservoir bag: 1l | Offer 1 | Each | Surgicare Medical | 136.2 | 94.25 | SURGICARE | Ningbo Haukun Medical Equipment | China
RT13-60-003 | BREATHING CIRCUITS | SMOOTH BORE BREATHING CIRCUITS WITH EXTRA LIMB AND BREATHING BAG - SINGLE USE | SMOOTH BORE BREATHING CIRCUITS WITH EXTRA LIMB AND BREATHING BAG - SINGLE USE Purpose: Used to deliver oxygen and Volatile Gases to patient's and to eliminate C02 during anaesthesia Transparent circuit with low resistance, non-extendable, smooth-bore tubing that is non-kinking, lightweight and crush resistant. The tubing must be pliable and must not retain a shape Spiral outer casing that supports the inner lumen to maintain tube integrity. Must have an extra limb with an anti-static reservoir bag that gives good feedback regarding patient lung compliance. The reservoir bag must not twist on itself and must have a mechanism to prevent this. Connectors Malleable 22mm female connectors bonded to the ends of inspiratory and expiratory limbs – must fit securely to the machine. Spare 22mm M/M connector. Y-piece at patient's end. Separate elbow-piece with capped luer lock port. Good quality to ensure no leakage or easy disconnections Latex-free, non-pyrogenic, hypoallergenic, non-toxic and non-conductive medical grade plastic No variation in Circuit limb length or Extra limb will be accepted Clinically clean and individually packed, For SINGLE use only To comply with ISO 5367 Information brochure or the outer package of the item should clearly indicate latex free | Smooth bore breathing circuit with extra limb and breathing bag - Single Use Adult 22mm Circuit: 1.8m; Extra limb: 1.2m OD: Reservoir bag: 2l | Offer 1 | Each | Surgicare Medical | 136.2 | 94.25 | SURGICARE | Ningbo Haukun Medical Equipment | China
RT13-60-004 | BREATHING CIRCUITS | SMOOTH BORE BREATHING CIRCUITS WITH EXTRA LIMB AND BREATHING BAG - SINGLE USE | SMOOTH BORE BREATHING CIRCUITS WITH EXTRA LIMB AND BREATHING BAG - SINGLE USE Purpose: Used to deliver oxygen and Volatile Gases to patient's and to eliminate C02 during anaesthesia Transparent circuit with low resistance, non-extendable, smooth-bore tubing that is non-kinking, lightweight and crush resistant. The tubing must be pliable and must not retain a shape Spiral outer casing that supports the inner lumen to maintain tube integrity. Must have an extra limb with an anti-static reservoir bag that gives good feedback regarding patient lung compliance. The reservoir bag must not twist on itself and must have a mechanism to prevent this. Connectors Malleable 22mm female connectors bonded to the ends of inspiratory and expiratory limbs – must fit securely to the machine. Spare 22mm M/M connector. Y-piece at patient's end. Separate elbow-piece with capped luer lock port. Good quality to ensure no leakage or easy disconnections Latex-free, non-pyrogenic, hypoallergenic, non-toxic and non-conductive medical grade plastic No variation in Circuit limb length or Extra limb will be accepted Clinically clean and individually packed, For SINGLE use only To comply with ISO 5367 Information brochure or the outer package of the item should clearly indicate latex free | Smooth bore breathing circuit with extra limb and breathing bag - Single Use Adult 22mm Circuit: 1.8m; Extra limb: 1.5m OD: Reservoir bag: 2l | Offer 1 | Each | Surgicare Medical | 135.84 | 94.25 | SURGICARE | Ningbo Haukun Medical Equipment | China
RT13-60-005 | BREATHING CIRCUITS | SMOOTH BORE BREATHING CIRCUITS WITH EXTRA LIMB AND BREATHING BAG - SINGLE USE | SMOOTH BORE BREATHING CIRCUITS WITH EXTRA LIMB AND BREATHING BAG - SINGLE USE Purpose: Used to deliver oxygen and Volatile Gases to patient's and to eliminate C02 during anaesthesia Transparent circuit with low resistance, non-extendable, smooth-bore tubing that is non-kinking, lightweight and crush resistant. The tubing must be pliable and must not retain a shape Spiral outer casing that supports the inner lumen to maintain tube integrity. Must have an extra limb with an anti-static reservoir bag that gives good feedback regarding patient lung compliance. The reservoir bag must not twist on itself and must have a mechanism to prevent this. Connectors Malleable 22mm female connectors bonded to the ends of inspiratory and expiratory limbs – must fit securely to the machine. Spare 22mm M/M connector. Y-piece at patient's end. Separate elbow-piece with capped luer lock port. Good quality to ensure no leakage or easy disconnections Latex-free, non-pyrogenic, hypoallergenic, non-toxic and non-conductive medical grade plastic No variation in Circuit limb length or Extra limb will be accepted Clinically clean and individually packed, For SINGLE use only To comply with ISO 5367 Information brochure or the outer package of the item should clearly indicate latex free | Smooth bore breathing circuit with extra limb and breathing bag - Single Use Adult 22mm Circuit: 1.8m; Extra limb: 1.8m OD: Reservoir bag: 2l | Offer 1 | Each | Surgicare Medical | 145.84 | 94.25 | SURGICARE | Ningbo Haukun Medical Equipment | China
RT13-60-006 | BREATHING CIRCUITS | SMOOTH BORE BREATHING CIRCUITS WITH EXTRA LIMB AND BREATHING BAG - SINGLE USE | SMOOTH BORE BREATHING CIRCUITS WITH EXTRA LIMB AND BREATHING BAG - SINGLE USE Purpose: Used to deliver oxygen and Volatile Gases to patient's and to eliminate C02 during anaesthesia Transparent circuit with low resistance, non-extendable, smooth-bore tubing that is non-kinking, lightweight and crush resistant. The tubing must be pliable and must not retain a shape Spiral outer casing that supports the inner lumen to maintain tube integrity. Must have an extra limb with an anti-static reservoir bag that gives good feedback regarding patient lung compliance. The reservoir bag must not twist on itself and must have a mechanism to prevent this. Connectors Malleable 22mm female connectors bonded to the ends of inspiratory and expiratory limbs – must fit securely to the machine. Spare 22mm M/M connector. Y-piece at patient's end. Separate elbow-piece with capped luer lock port. Good quality to ensure no leakage or easy disconnections Latex-free, non-pyrogenic, hypoallergenic, non-toxic and non-conductive medical grade plastic No variation in Circuit limb length or Extra limb will be accepted Clinically clean and individually packed, For SINGLE use only To comply with ISO 5367 Information brochure or the outer package of the item should clearly indicate latex free | Smooth bore breathing circuit with extra limb and breathing bag - Single Use Adult Long 22mm Circuit: 2.6m; Extra limb: 1.2m Reservoir bag: 2l | Offer 1 | Each | Surgicare Medical | 173.86 | 94.25 | SURGICARE | Ningbo Haukun Medical Equipment | China
RT13-60-007 | BREATHING CIRCUITS | SMOOTH BORE BREATHING CIRCUITS WITH EXTRA LIMB AND BREATHING BAG - SINGLE USE | SMOOTH BORE BREATHING CIRCUITS WITH EXTRA LIMB AND BREATHING BAG - SINGLE USE Purpose: Used to deliver oxygen and Volatile Gases to patient's and to eliminate C02 during anaesthesia Transparent circuit with low resistance, non-extendable, smooth-bore tubing that is non-kinking, lightweight and crush resistant. The tubing must be pliable and must not retain a shape Spiral outer casing that supports the inner lumen to maintain tube integrity. Must have an extra limb with an anti-static reservoir bag that gives good feedback regarding patient lung compliance. The reservoir bag must not twist on itself and must have a mechanism to prevent this. Connectors Malleable 22mm female connectors bonded to the ends of inspiratory and expiratory limbs – must fit securely to the machine. Spare 22mm M/M connector. Y-piece at patient's end. Separate elbow-piece with capped luer lock port. Good quality to ensure no leakage or easy disconnections Latex-free, non-pyrogenic, hypoallergenic, non-toxic and non-conductive medical grade plastic No variation in Circuit limb length or Extra limb will be accepted Clinically clean and individually packed, For SINGLE use only To comply with ISO 5367 Information brochure or the outer package of the item should clearly indicate latex free | Smooth bore breathing circuit with extra limb and breathing bag - Single Use Adult Long 22mm Circuit: 2.6m; Extra limb: 1.5m Reservoir bag: 2l | Offer 1 | Each | Surgicare Medical | 173.86 | 94.25 | SURGICARE | Ningbo Haukun Medical Equipment | China
RT13-60-008 | BREATHING CIRCUITS | SMOOTH BORE BREATHING CIRCUITS WITH EXTRA LIMB AND BREATHING BAG - SINGLE USE | SMOOTH BORE BREATHING CIRCUITS WITH EXTRA LIMB AND BREATHING BAG - SINGLE USE Purpose: Used to deliver oxygen and Volatile Gases to patient's and to eliminate C02 during anaesthesia Transparent circuit with low resistance, non-extendable, smooth-bore tubing that is non-kinking, lightweight and crush resistant. The tubing must be pliable and must not retain a shape Spiral outer casing that supports the inner lumen to maintain tube integrity. Must have an extra limb with an anti-static reservoir bag that gives good feedback regarding patient lung compliance. The reservoir bag must not twist on itself and must have a mechanism to prevent this. Connectors Malleable 22mm female connectors bonded to the ends of inspiratory and expiratory limbs – must fit securely to the machine. Spare 22mm M/M connector. Y-piece at patient's end. Separate elbow-piece with capped luer lock port. Good quality to ensure no leakage or easy disconnections Latex-free, non-pyrogenic, hypoallergenic, non-toxic and non-conductive medical grade plastic No variation in Circuit limb length or Extra limb will be accepted Clinically clean and individually packed, For SINGLE use only To comply with ISO 5367 Information brochure or the outer package of the item should clearly indicate latex free | Smooth bore breathing circuit with extra limb and breathing bag - Single Use Adult Long 22mm Circuit: 2.6m; Extra limb: 1.8m Reservoir bag: 2l | Offer 1 | Each | Surgicare Medical | 174.26 | 94.25 | SURGICARE | Ningbo Haukun Medical Equipment | China
RT13-60-009 | BREATHING CIRCUITS | SMOOTH BORE BREATHING CIRCUITS WITH WATER TRAPS AND EXTRA LIMB AND BAG - SINGLE USE | SMOOTH BORE BREATHING CIRCUITS WITH WATER TRAPS AND EXTRA LIMB AND BAG - SINGLE USE Purpose: Used to deliver oxygen and Volatile Gases to patient's and to eliminate C02 during anaesthesia Transparent circuit with low resistance, non-extendable, smooth-bore tubing that is non-kinking, lightweight and crush resistant. The tubing must be pliable and must not retain a shape Spiral outer casing that supports the inner lumen to maintain tube integrity. Lightweight water traps centrally on inspiratory and expiratory limbs Must have an extra limb with an anti-static reservoir bag that gives good feedback regarding patient lung compliance. The reservoir bag must not twist on itself and must have a mechanism to prevent this. Connectors Malleable 22mm female connectors bonded to the ends of inspiratory and expiratory limbs – must fit securely to the machine. Spare 22mm M/M connector. Y-piece at patient's end. Separate angle-piece with luer lock port Good quality to ensure no leakage or easy disconnections Latex-free, non-pyrogenic, hypoallergenic, non-toxic and non-conductive medical grade plastic No variation in Circuit limb length or Extra limb will be accepted Clinically clean and individually packed, For SINGLE use only To comply with ISO 5367 Information brochure or the outer package of the item should clearly indicate latex free | Smooth bore breathing circuit with with water traps, extra limb and breathing bag - Single Use Adult Circuit: ± 1.8m; Extra limb: ± 1.2m OD: 22mm Reservoir bag: 2l | Offer 1 | Each | Surgicare Medical | 174.26 | 94.25 | SURGICARE | Ningbo Haukun Medical Equipment | China
RT13-60-011 | BREATHING CIRCUITS | SMOOTH BORE BREATHING CIRCUITS WITH WATER TRAPS AND EXTRA LIMB AND BAG - SINGLE USE | SMOOTH BORE BREATHING CIRCUITS WITH WATER TRAPS AND EXTRA LIMB AND BAG - SINGLE USE Purpose: Used to deliver oxygen and Volatile Gases to patient's and to eliminate C02 during anaesthesia Transparent circuit with low resistance, non-extendable, smooth-bore tubing that is non-kinking, lightweight and crush resistant. The tubing must be pliable and must not retain a shape Spiral outer casing that supports the inner lumen to maintain tube integrity. Lightweight water traps centrally on inspiratory and expiratory limbs Must have an extra limb with an anti-static reservoir bag that gives good feedback regarding patient lung compliance. The reservoir bag must not twist on itself and must have a mechanism to prevent this. Connectors Malleable 22mm female connectors bonded to the ends of inspiratory and expiratory limbs – must fit securely to the machine. Spare 22mm M/M connector. Y-piece at patient's end. Separate angle-piece with luer lock port Good quality to ensure no leakage or easy disconnections Latex-free, non-pyrogenic, hypoallergenic, non-toxic and non-conductive medical grade plastic No variation in Circuit limb length or Extra limb will be accepted Clinically clean and individually packed, For SINGLE use only To comply with ISO 5367 Information brochure or the outer package of the item should clearly indicate latex free | Smooth bore breathing circuit with with water traps, extra limb and breathing bag - Single Use Adult Circuit: ± 1.8m; Extra limb: ± 1.8m OD: 22mm Reservoir bag: 2l | Offer 1 | Each | Surgicare Medical | 174.26 | 94.25 | SURGICARE | Ningbo Haukun Medical Equipment | China
RT13-60-012 | BREATHING CIRCUITS | SMOOTH BORE BREATHING CIRCUITS WITH WATER TRAPS AND EXTRA LIMB AND BAG - SINGLE USE | SMOOTH BORE BREATHING CIRCUITS WITH WATER TRAPS AND EXTRA LIMB AND BAG - SINGLE USE Purpose: Used to deliver oxygen and Volatile Gases to patient's and to eliminate C02 during anaesthesia Transparent circuit with low resistance, non-extendable, smooth-bore tubing that is non-kinking, lightweight and crush resistant. The tubing must be pliable and must not retain a shape Spiral outer casing that supports the inner lumen to maintain tube integrity. Lightweight water traps centrally on inspiratory and expiratory limbs Must have an extra limb with an anti-static reservoir bag that gives good feedback regarding patient lung compliance. The reservoir bag must not twist on itself and must have a mechanism to prevent this. Connectors Malleable 22mm female connectors bonded to the ends of inspiratory and expiratory limbs – must fit securely to the machine. Spare 22mm M/M connector. Y-piece at patient's end. Separate angle-piece with luer lock port Good quality to ensure no leakage or easy disconnections Latex-free, non-pyrogenic, hypoallergenic, non-toxic and non-conductive medical grade plastic No variation in Circuit limb length or Extra limb will be accepted Clinically clean and individually packed, For SINGLE use only To comply with ISO 5367 Information brochure or the outer package of the item should clearly indicate latex free | Smooth bore breathing circuit with with water traps, extra limb and breathing bag - Single Use Adult Long Circuit: ± 2.6m; Extra limb: ± 1.2m OD: 22mm Reservoir bag: 2l | Offer 1 | Each | Surgicare Medical | 208.28 | 94.25 | SURGICARE | Ningbo Haukun Medical Equipment | China
RT13-60-013 | BREATHING CIRCUITS | SMOOTH BORE BREATHING CIRCUITS WITH WATER TRAPS AND EXTRA LIMB AND BAG - SINGLE USE | SMOOTH BORE BREATHING CIRCUITS WITH WATER TRAPS AND EXTRA LIMB AND BAG - SINGLE USE Purpose: Used to deliver oxygen and Volatile Gases to patient's and to eliminate C02 during anaesthesia Transparent circuit with low resistance, non-extendable, smooth-bore tubing that is non-kinking, lightweight and crush resistant. The tubing must be pliable and must not retain a shape Spiral outer casing that supports the inner lumen to maintain tube integrity. Lightweight water traps centrally on inspiratory and expiratory limbs Must have an extra limb with an anti-static reservoir bag that gives good feedback regarding patient lung compliance. The reservoir bag must not twist on itself and must have a mechanism to prevent this. Connectors Malleable 22mm female connectors bonded to the ends of inspiratory and expiratory limbs – must fit securely to the machine. Spare 22mm M/M connector. Y-piece at patient's end. Separate angle-piece with luer lock port Good quality to ensure no leakage or easy disconnections Latex-free, non-pyrogenic, hypoallergenic, non-toxic and non-conductive medical grade plastic No variation in Circuit limb length or Extra limb will be accepted Clinically clean and individually packed, For SINGLE use only To comply with ISO 5367 Information brochure or the outer package of the item should clearly indicate latex free | Smooth bore breathing circuit with with water traps, extra limb and breathing bag - Single Use Adult Long Circuit: ± 2.6m; Extra limb: ± 1.5m OD: 22mm Reservoir bag: 2l | Offer 1 | Each | Surgicare Medical | 208.28 | 94.25 | SURGICARE | Ningbo Haukun Medical Equipment | China
RT13-60-014 | BREATHING CIRCUITS | SMOOTH BORE BREATHING CIRCUITS WITH WATER TRAPS AND EXTRA LIMB AND BAG - SINGLE USE | SMOOTH BORE BREATHING CIRCUITS WITH WATER TRAPS AND EXTRA LIMB AND BAG - SINGLE USE Purpose: Used to deliver oxygen and Volatile Gases to patient's and to eliminate C02 during anaesthesia Transparent circuit with low resistance, non-extendable, smooth-bore tubing that is non-kinking, lightweight and crush resistant. The tubing must be pliable and must not retain a shape Spiral outer casing that supports the inner lumen to maintain tube integrity. Lightweight water traps centrally on inspiratory and expiratory limbs Must have an extra limb with an anti-static reservoir bag that gives good feedback regarding patient lung compliance. The reservoir bag must not twist on itself and must have a mechanism to prevent this. Connectors Malleable 22mm female connectors bonded to the ends of inspiratory and expiratory limbs – must fit securely to the machine. Spare 22mm M/M connector. Y-piece at patient's end. Separate angle-piece with luer lock port Good quality to ensure no leakage or easy disconnections Latex-free, non-pyrogenic, hypoallergenic, non-toxic and non-conductive medical grade plastic No variation in Circuit limb length or Extra limb will be accepted Clinically clean and individually packed, For SINGLE use only To comply with ISO 5367 Information brochure or the outer package of the item should clearly indicate latex free | Smooth bore breathing circuit with with water traps, extra limb and breathing bag - Single Use Adult Long Circuit: ± 2.6m; Extra limb: ± 1.8m OD: 22mm Reservoir bag: 2l | Offer 1 | Each | Surgicare Medical | 208.28 | 94.25 | SURGICARE | Ningbo Haukun Medical Equipment | China
RT13-64-005 | BREATHING CIRCUITS | CORRUGATED BREATHING CIRCUITS WITH WATER TRAPS AND EXTRA LIMB AND BAG - SINGLE USE | CORRUGATED BREATHING CIRCUITS WITH WATER TRAPS AND EXTRA LIMB AND BAG - SINGLE USE Purpose: Used to deliver oxygen and Volatile Gases to patient's and to eliminate C02 during anaesthesia Transparent circuit with low resistance, non-extendable, corrugated tubing that is non-kinking, lightweight and crush resistant. The tubing must be pliable and must not retain a shape Shallow corrugations to minimise flow resistance and turbulence. Uninterrupted tubing (not joined). Lightweight water traps centrally on inspiratory and expiratory limb Must have an extra limb with an anti-static reservoir bag that gives good feedback regarding patient lung compliance. The reservoir bag must not twist on itself and must have a mechanism to prevent this. Connectors Malleable 22mm female connectors bonded to the ends of inspiratory and expiratory limbs – must fit securely to the machine. Spare 22mm M/M connector. Y-piece at patient's end. Separate angle-piece with luer lock port Good quality to ensure no leakage or easy disconnections Latex-free, non-pyrogenic, hypoallergenic, non-toxic and non-conductive medical grade plastic No variation in Circuit limb length or Extra limb will be accepted Clinically clean and individually packed, For SINGLE use only To comply with ISO 5367 Information brochure or the outer package of the item should clearly indicate latex free | Corrugated breathing circuit with water traps and extra limb with breathing bag - Single Use Adult 22mm Circuit: 1.8m; 22mm Extra limb: 1.8m OD: 22mm Reservoir bag: 2l | Offer 1 | Each | Surgicare Medical | 208.29 | 94.25 | SURGICARE | Ningbo Haukun Medical Equipment | China
RT13-57-002 | BREATHING CIRCUITS | CORRUGATED BREATHING CIRCUITS WITH EXTRA LIMB AND BAG | CORRUGATED BREATHING CIRCUITS WITH EXTRA LIMB AND BAG - SINGLE USE Purpose: Used to deliver oxygen and Volatile Gases to patient's and to eliminate C02 during anaesthesia Transparent circuit with low resistance, non-extendable, corrugated tubing that is non-kinking, lightweight and crush resistant. The tubing must be pliable and must not retain a shape. Uninterrupted tubing (not joined). Shallow corrugations to minimise flow resistance and turbulence. Must have an extra limb with an anti-static reservoir bag that gives good feedback regarding patient lung compliance. The reservoir bag must not twist on itself and must have a mechanism to prevent this. Connectors Malleable 22mm female connectors bonded to the ends of inspiratory and expiratory limbs – must fit securely to the machine. Spare 22mm M/M connector. Y-piece at patient's end. Separate elbow-piece with a toggled cap on a luer lock port. Good quality to ensure no leakage or easy disconnections Latex-free, non-pyrogenic, hypoallergenic, non-toxic and non-conductive medical grade plastic No variation in Circuit limb length or Extra limb will be accepted Clinically clean and individually packed, For SINGLE use only To comply with ISO 5367 Information brochure or the outer package of the item should clearly indicate latex free | Corrugated breathing circuit with extra limb and breathing bag - Single Use - Neonatal 15mm Circuit: 1.6m; 15mm Extra limb: 1.8m Reservoir bag: 500ml | Offer 1 | Each | Lechoba Medical Technologies | 178.63 | -5.5092 | NON-CHANGE | NON-CHANGE ENTERPRISE CO., LTD | TAIWAN
RT13-58-002 | BREATHING CIRCUITS | SINGLE CORRUGATED BREATHING TUBING - SINGLE USE | SINGLE CORRUGATED BREATHING TUBING - SINGLE USE Purpose: Used as extension or extra limb in breathing circuits Transparent circuit with low resistance, non-extendable, corrugated tubing that is non-kinking, lightweight and crush resistant. The tubing must be pliable and must not retain a shape Shallow corrugations to minimise flow resistance and turbulence. Connectors Malleable 22mm female connectors bonded to the end of the tubing Spare connectors to be provided Latex-free, non-pyrogenic, hypoallergenic, non-toxic and non-conductive medical grade plastic No variation in Circuit limb length or Extra limb will be accepted Clinically clean and individually packed, For SINGLE use only To comply with ISO 5367 Information brochure or the outer package of the item should clearly indicate latex free | Single Corrugated Breathing Tubing - Single Use Paediatric 15mm Length: 1.8m Connector: 15mm M/F; 15mm M/M; 15mm F/F | Offer 1 | Each | Lechoba Medical Technologies | 34.28 | 95 | NON-CHANGE | NON-CHANGE ENTERPRISE CO., LTD | TAIWAN
RT13-60-002 | BREATHING CIRCUITS | SMOOTH BORE BREATHING CIRCUITS WITH EXTRA LIMB AND BREATHING BAG - SINGLE USE | SMOOTH BORE BREATHING CIRCUITS WITH EXTRA LIMB AND BREATHING BAG - SINGLE USE Purpose: Used to deliver oxygen and Volatile Gases to patient's and to eliminate C02 during anaesthesia Transparent circuit with low resistance, non-extendable, smooth-bore tubing that is non-kinking, lightweight and crush resistant. The tubing must be pliable and must not retain a shape Spiral outer casing that supports the inner lumen to maintain tube integrity. Must have an extra limb with an anti-static reservoir bag that gives good feedback regarding patient lung compliance. The reservoir bag must not twist on itself and must have a mechanism to prevent this. Connectors Malleable 22mm female connectors bonded to the ends of inspiratory and expiratory limbs – must fit securely to the machine. Spare 22mm M/M connector. Y-piece at patient's end. Separate elbow-piece with capped luer lock port. Good quality to ensure no leakage or easy disconnections Latex-free, non-pyrogenic, hypoallergenic, non-toxic and non-conductive medical grade plastic No variation in Circuit limb length or Extra limb will be accepted Clinically clean and individually packed, For SINGLE use only To comply with ISO 5367 Information brochure or the outer package of the item should clearly indicate latex free | Smooth bore breathing circuit with extra limb and breathing bag - Single Use Paediatric > 3yrs 15mm Circuit: 1.6m; Extra limb: 1.8m Reservoir bag: 1l | Offer 1 | Each | Lechoba Medical Technologies | 209.77 | 36.6769 | NON-CHANGE | NON-CHANGE ENTERPRISE CO., LTD | TAIWAN
RT13-60-003 | BREATHING CIRCUITS | SMOOTH BORE BREATHING CIRCUITS WITH EXTRA LIMB AND BREATHING BAG - SINGLE USE | SMOOTH BORE BREATHING CIRCUITS WITH EXTRA LIMB AND BREATHING BAG - SINGLE USE Purpose: Used to deliver oxygen and Volatile Gases to patient's and to eliminate C02 during anaesthesia Transparent circuit with low resistance, non-extendable, smooth-bore tubing that is non-kinking, lightweight and crush resistant. The tubing must be pliable and must not retain a shape Spiral outer casing that supports the inner lumen to maintain tube integrity. Must have an extra limb with an anti-static reservoir bag that gives good feedback regarding patient lung compliance. The reservoir bag must not twist on itself and must have a mechanism to prevent this. Connectors Malleable 22mm female connectors bonded to the ends of inspiratory and expiratory limbs – must fit securely to the machine. Spare 22mm M/M connector. Y-piece at patient's end. Separate elbow-piece with capped luer lock port. Good quality to ensure no leakage or easy disconnections Latex-free, non-pyrogenic, hypoallergenic, non-toxic and non-conductive medical grade plastic No variation in Circuit limb length or Extra limb will be accepted Clinically clean and individually packed, For SINGLE use only To comply with ISO 5367 Information brochure or the outer package of the item should clearly indicate latex free | Smooth bore breathing circuit with extra limb and breathing bag - Single Use Adult 22mm Circuit: 1.8m; Extra limb: 1.2m OD: Reservoir bag: 2l | Offer 1 | Each | Lechoba Medical Technologies | 208.66 | 40.8856 | NON-CHANGE | NON-CHANGE ENTERPRISE CO., LTD | TAIWAN
RT13-60-004 | BREATHING CIRCUITS | SMOOTH BORE BREATHING CIRCUITS WITH EXTRA LIMB AND BREATHING BAG - SINGLE USE | SMOOTH BORE BREATHING CIRCUITS WITH EXTRA LIMB AND BREATHING BAG - SINGLE USE Purpose: Used to deliver oxygen and Volatile Gases to patient's and to eliminate C02 during anaesthesia Transparent circuit with low resistance, non-extendable, smooth-bore tubing that is non-kinking, lightweight and crush resistant. The tubing must be pliable and must not retain a shape Spiral outer casing that supports the inner lumen to maintain tube integrity. Must have an extra limb with an anti-static reservoir bag that gives good feedback regarding patient lung compliance. The reservoir bag must not twist on itself and must have a mechanism to prevent this. Connectors Malleable 22mm female connectors bonded to the ends of inspiratory and expiratory limbs – must fit securely to the machine. Spare 22mm M/M connector. Y-piece at patient's end. Separate elbow-piece with capped luer lock port. Good quality to ensure no leakage or easy disconnections Latex-free, non-pyrogenic, hypoallergenic, non-toxic and non-conductive medical grade plastic No variation in Circuit limb length or Extra limb will be accepted Clinically clean and individually packed, For SINGLE use only To comply with ISO 5367 Information brochure or the outer package of the item should clearly indicate latex free | Smooth bore breathing circuit with extra limb and breathing bag - Single Use Adult 22mm Circuit: 1.8m; Extra limb: 1.5m OD: Reservoir bag: 2l | Offer 1 | Each | Lechoba Medical Technologies | 209.8 | 35.8549 | NON-CHANGE | NON-CHANGE ENTERPRISE CO., LTD | TAIWAN
RT13-60-005 | BREATHING CIRCUITS | SMOOTH BORE BREATHING CIRCUITS WITH EXTRA LIMB AND BREATHING BAG - SINGLE USE | SMOOTH BORE BREATHING CIRCUITS WITH EXTRA LIMB AND BREATHING BAG - SINGLE USE Purpose: Used to deliver oxygen and Volatile Gases to patient's and to eliminate C02 during anaesthesia Transparent circuit with low resistance, non-extendable, smooth-bore tubing that is non-kinking, lightweight and crush resistant. The tubing must be pliable and must not retain a shape Spiral outer casing that supports the inner lumen to maintain tube integrity. Must have an extra limb with an anti-static reservoir bag that gives good feedback regarding patient lung compliance. The reservoir bag must not twist on itself and must have a mechanism to prevent this. Connectors Malleable 22mm female connectors bonded to the ends of inspiratory and expiratory limbs – must fit securely to the machine. Spare 22mm M/M connector. Y-piece at patient's end. Separate elbow-piece with capped luer lock port. Good quality to ensure no leakage or easy disconnections Latex-free, non-pyrogenic, hypoallergenic, non-toxic and non-conductive medical grade plastic No variation in Circuit limb length or Extra limb will be accepted Clinically clean and individually packed, For SINGLE use only To comply with ISO 5367 Information brochure or the outer package of the item should clearly indicate latex free | Smooth bore breathing circuit with extra limb and breathing bag - Single Use Adult 22mm Circuit: 1.8m; Extra limb: 1.8m OD: Reservoir bag: 2l | Offer 1 | Each | Lechoba Medical Technologies | 219.83 | 40.4337 | NON-CHANGE | NON-CHANGE ENTERPRISE CO., LTD | TAIWAN
RT13-60-009 | BREATHING CIRCUITS | SMOOTH BORE BREATHING CIRCUITS WITH WATER TRAPS AND EXTRA LIMB AND BAG - SINGLE USE | SMOOTH BORE BREATHING CIRCUITS WITH WATER TRAPS AND EXTRA LIMB AND BAG - SINGLE USE Purpose: Used to deliver oxygen and Volatile Gases to patient's and to eliminate C02 during anaesthesia Transparent circuit with low resistance, non-extendable, smooth-bore tubing that is non-kinking, lightweight and crush resistant. The tubing must be pliable and must not retain a shape Spiral outer casing that supports the inner lumen to maintain tube integrity. Lightweight water traps centrally on inspiratory and expiratory limbs Must have an extra limb with an anti-static reservoir bag that gives good feedback regarding patient lung compliance. The reservoir bag must not twist on itself and must have a mechanism to prevent this. Connectors Malleable 22mm female connectors bonded to the ends of inspiratory and expiratory limbs – must fit securely to the machine. Spare 22mm M/M connector. Y-piece at patient's end. Separate angle-piece with luer lock port Good quality to ensure no leakage or easy disconnections Latex-free, non-pyrogenic, hypoallergenic, non-toxic and non-conductive medical grade plastic No variation in Circuit limb length or Extra limb will be accepted Clinically clean and individually packed, For SINGLE use only To comply with ISO 5367 Information brochure or the outer package of the item should clearly indicate latex free | Smooth bore breathing circuit with with water traps, extra limb and breathing bag - Single Use Adult Circuit: ± 1.8m; Extra limb: ± 1.2m OD: 22mm Reservoir bag: 2l | Offer 1 | Each | Lechoba Medical Technologies | 259.31 | 40.0612 | NON-CHANGE | NON-CHANGE ENTERPRISE CO., LTD | TAIWAN
RT13-60-011 | BREATHING CIRCUITS | SMOOTH BORE BREATHING CIRCUITS WITH WATER TRAPS AND EXTRA LIMB AND BAG - SINGLE USE | SMOOTH BORE BREATHING CIRCUITS WITH WATER TRAPS AND EXTRA LIMB AND BAG - SINGLE USE Purpose: Used to deliver oxygen and Volatile Gases to patient's and to eliminate C02 during anaesthesia Transparent circuit with low resistance, non-extendable, smooth-bore tubing that is non-kinking, lightweight and crush resistant. The tubing must be pliable and must not retain a shape Spiral outer casing that supports the inner lumen to maintain tube integrity. Lightweight water traps centrally on inspiratory and expiratory limbs Must have an extra limb with an anti-static reservoir bag that gives good feedback regarding patient lung compliance. The reservoir bag must not twist on itself and must have a mechanism to prevent this. Connectors Malleable 22mm female connectors bonded to the ends of inspiratory and expiratory limbs – must fit securely to the machine. Spare 22mm M/M connector. Y-piece at patient's end. Separate angle-piece with luer lock port Good quality to ensure no leakage or easy disconnections Latex-free, non-pyrogenic, hypoallergenic, non-toxic and non-conductive medical grade plastic No variation in Circuit limb length or Extra limb will be accepted Clinically clean and individually packed, For SINGLE use only To comply with ISO 5367 Information brochure or the outer package of the item should clearly indicate latex free | Smooth bore breathing circuit with with water traps, extra limb and breathing bag - Single Use Adult Circuit: ± 1.8m; Extra limb: ± 1.8m OD: 22mm Reservoir bag: 2l | Offer 1 | Each | Lechoba Medical Technologies | 271.11 | 34.4458 | NON-CHANGE | NON-CHANGE ENTERPRISE CO., LTD | TAIWAN