## List
ANNEXURE A
No | Description of the Item | Equipment to be Submitted | Comments
1 | Anaesthesia Delivery System: Advanced without physiological monitor
2 | Anaesthesia Delivery System: Advanced with integrated physiological monitor
3 | Anaesthesia Delivery System: Standard without physiological monitor
4 | Anaesthesia Delivery System: Standard with integrated physiological monitor
5 | Apnea Monitor (Impedance sensor)
6 | Apnea Monitor (Pressure sensor)
7 | Apnea Monitor (Airflow/CO2 sensor)
8 | Billirubin meter Xenon
9 | Billirubin meter LED
10 | Cardiotocograph (twins)
11 | Defibrillator, with SpO2 and pacing
12 | Electrocardiograph (12 channel) with trolley.
13 | Electrosurgery unit, advanced
14 | Electrosurgical unit basic
15 | Electrosurgery unit, general surgery (350 W)
16 | Examination light, LED, ceiling mount
17 | Examination light, LED, mobile, with battery
18 | Examination light, LED, mobile, without battery
19 | Examination light, LED, wall mount
20 | Doppler Fetal Heart Rate Monitor
21 | ICU Crib
22 | Incubator ICU
23 | Incubator (Standard)
24 | Infant radiant warmer
25 | Infant radiant warmer (Wall-mount)
26 | Infusion Pump
27 | Monitor basic multiparameter
28 | Monitor ICU
29 | Central Station
30 | NIBP + SpO2 monitor, desktop
31 | NIBP + SpO2 monitor, on trolley
32 | NIBP Monitor, desktop
33 | NIBP Monitor on mobile stand
34 | Oxygen Flowmeter, single
35 | Oxygen Flowmeter, double
36 | Plaster of Paris Cutter, electric, with vacuum
37 | Phototherapy Light (LED)
38 | Phototherapy Blanket LED
39 | Pulse Oximeter, desktop
40 | Pulse Oximeter on trolley
41 | Resuscitator, infant, gas powered
42 | Resuscitator, infant, gas powered
43 | Suction Unit, mobile, battery operated
44 | Suction Unit, Mobile
45 | Suction unit, wall-mount
46 | Thermometer, electronic
47 | Ventilator CPAP
48 | Ventilator: Neonatal/Paediatric
49 | Ventilator Paediatric/Adult

## Anaesthetic 1
| ANAESTHESIA DELIVERY SYSTEM: Advanced without physiological monitor | Complies Y/N | Comment
1 | An advanced anaesthesia delivery system comprising a completely integrated electronic anaesthesia workstation.
2 | The unit must include electronicaly integrated oxygen, nitrous oxide and air delivery and monitoring.
3 | Ventilation and anaesthesia vapour monitoring must be fully integrated within the unit.
4 | The anaesthesia unit must be able to communicate/interface with a physiological monitor
5 | The anaesthesia system must comprise:
5.1 | Anaesthesia system suitable for neonatal (down to 500g), paediatric and adult anaesthesia.
5.2 | Ventilator suitable for ventilating neonatal (down to 500g), paediatric, adult and morbid/obese patients
5.3 | Multigas analyser:
5.3.1 | Oxygen
5.3.2 | Nitrous oxide
5.3.3 | Carbon dioxide
5.4 | Anaesthetic vapour monitoring:
5.4.1 | Isoflurane
5.4.2 | Sevoflurane
5.4.3 | Desflurane
5.4.4 | Halothane
5.4.5 | Enflurane
6 | Must comply with ISO 13485
7 | Must comply with SANS/IEC 60601-2-13
8 | Complete and automated self test requiring very limited user intervention.
9 | The self test must provide the measurements and values for automatic leakage and compliance.
10 | Automatic calibration of sensors for anaesthetic gases and vapours must form part of the self test and require very limited user intervention.
11 | It must be possible to override the self test facility. The system must display the override status
12 | The unit must provide conventional, low-flow (down to 200 ml/min) anaesthesia.
13 | The unit must return analysed gas back to the circle system
14 | A safety oxygen function is required for the delivery of oxygen through the anaesthesia unit in the event of an electronic failure of the gas delivery system.
15 | The unit must be supplied with a minimum of 1 storage drawer
16 | The unit must be supplied with a minimum of 3 additional 220 V power sockets. These sockets must be excluded from the battery backup circuit in case of a power failure
17 | The unit must have a working area or a retractable shelf. Width: not less than 300 mm. Depth: not less than 200 mm
18 | The unit must have a light to enable the anaesthetist to work in a darkened theatre.
19 | Trolley
19.1 | The unit must be mounted on a strong , robust and compact trolley manufactured from corrosion free material.
19.2 | The trolley must have no sharp corners and be easy to clean.
19.3 | The trolley must have at least four antistatic castors of diameter 100 mm or greater.
19.4 | Trolley must have two lockable castors or a central locking brake.
19.5 | The unit must accommodate an oxygen cylinder. Cylinder size according to user requirement. Size to be compatible with current state vendor of oxygen.
19.6 | Trolley must have accessible handles for ease of movement
19.7 | The trolley, with all components mounted, must be stable without toppling over when placed on a 30° slope.
19.8 | The unit must be supplied with a castor guard system, preventing the entaglement and trapping of cables within the wheels.
19.9 | The trolley must have a mounting system suitable for a physiologial monitor.
20 | Pipeline and Cylinder Gas Supplies
20.1 | The anaesthesia system must be designed to use external piped medical gasses (O2, N2O and medical air) and vacuum from medical service outlet sockets.
20.2 | Connection hoses for above must be provided
20.3 | Hoses must be fitted with Heyer type non-interchangeable probes for the existing medical gas outlets as specified by the purchaser.
20.4 | The medical air hose connection or CO2 circle system must include a water trap
20.5 | A safety pressure relief valve of 60 - 80 cm H2O must be provided to protect the system from excessive pressure .
20.6 | The unit must function safely and correctly within the pipeline supply pressure range of: 270 - 600 kPa
20.7 | The system must preferentially utilise the piped oxygen supply over the cylinder supply, if pipeline pressure is > 270 kPa
20.8 | Connection hoses must be:
20.8.1 | Flexible and at least 6 metres long
20.8.2 | Antistatic
20.8.3 | Colour coded as follows:
20.8.3.1 | O2: white
20.8.3.2 | N2O: blue
20.8.3.3 | Vacuum: yellow
20.8.3.4 | Medical air: black
20.8.3.5 | Scavenge low vacuum: orange or yellow
20.8.4 | Where relevant, fitted with outlet probes complying with SANS 1409:2014
20.8.5 | Permanently bonded to the outlet probe and the machine connector
20.9 | Hoses for gases above atmospheric pressure, must designed for pressures of at least 1 000 kPa.
20.10 | Vacuum hoses may not collapse or kink at 90 kPa below atmospheric pressure, when subjected to 180° bend of inside radius equal to the outside diameter of the hose.
20.11 | All connection hoses must be fed to a central manifold
20.12 | All high pressure gas and vacuum pipe connections (other than hoses) and all pipes connected to pressure regulators and pressure gauges must be stainless steel (18/8 grade) pipes that are oil-free and colour coded as for hoses (see above) over their entire lengths
20.13 | An in-line non-return valve must be fitted to each gas pipeline on the high pressure (inlet) line
20.14 | An in-line non-return valve must be fitted to each gas cylinder yoke
20.15 | A replaceable gas filter, to block particles greater than 65 µm, must be fitted on the upstream side of the non-return valve
20.16 | Each gas pipeline inlet must have a pressure gauge or electronic display of pressure. The pressure range (415 kPa ± 75 kPa) must be clearly indicated in colour
20.17 | Must be fitted with a gas cylinder yoke to accommodate one oxygen gas cylinder with a pin index system
20.18 | Each gas cylinder yoke must have a pressure gauge or electronic display of pressure
20.19 | The pressure gauges or electronic displays of pressure for both pipeline & cylinder supplies, calibrated in kPa for each line, must be labelled and visible from the front of the unit.
20.20 | O2 and N2O regulators must be provided complying with CKS 605:1987.
20.21 | A pressure regulator complying with CKS 605 must be supplied for each gas inlet and must reduce the gas pressure to 415 kPa without freezing up, creeping, chattering or whistling
20.22 | Outlet sockets and probes must comply with SANS 1409:2014.
20.23 | Sockets with hoses must be colour-coded as per SANS 1409:2014:
20.23.1 | Oxygen: White background with black letters
20.23.2 | Nitrous oxide: Blue background with white letters
20.23.3 | Medical air: One half black with white letters, the other half white with black letters
20.23.4 | Vacuum: Yellow background with black letters
20.24 | Continuous display, visible from the front of the machine, of both pipeline and cylinder pressures even during power failure
20.25 | Pipeline and cylinder pressures must be differentially displayed, and clearly indicated
20.26 | Safety pressure relief valves must protect the entire pneumatic system from excessive input pressures from pipeline and cylinder supplies.
20.27 | When a cylinder is detached from the pin-indexed yoke no gas leak must occur at the yoke.
20.28 | Airway pressure must be monitored and displayed
20.29 | Safety relief valve for all gases must be supplied and set at 600kPa
21 | Scavenging
21.1 | The unit must have a scavenging system with a hose at least 5 metres long with the suitable connector as per the institutions scavenging installation.
21.2 | Air-break device to prevent negative pressure on the breathing system
21.3 | If an active sytem is supplied a pressure indicator must be present
21.4 | The scavenging outlet must be a 30mm ISO connector.
22 | Patient Gas Delivery System
22.1 | Patient gas delivery must be electronically controlled
22.2 | Must display the delivered patient gas flow using an ISO colour standard for each gas.
22.3 | The unit must be able to deliver a total patient gas flow of 200 ml/min to 18 l/min (Minimum). The unit must have the ability to deliver at least 10 l/min of oxygen. These must be incremental adjustments
22.4 | Gas supply must conform to:
22.4.1 | O2 minimum flow range of 200 ml/min to 10 l/min. Accuracy ± 5%
22.4.2 | medical air minimum flow range 0 to 10 l/min. Accuracy ± 5%
22..4.3 | nitrous oxide minimum flow range 0 to 10 l/min. Accuracy ± 5%
22.5 | Flowmeter sequence to be according to standards approved by the South Africal Society of Anaesthesiologists - Oxygen on the right
22.6 | The delivery system must prevent simultaneous delivery of two anaesthetic vapours.
22.7 | A separate common gas outlet to accept a standard 22 mm outside diameter and 15 mm internal diameter.
23 | Hypoxic guard
23.1 | The unit must include a safety mechanism that ensures a minimum concentration of 25% oxygen in any oxygen/nitrous oxide gas mixture across all flow rates.
23.2 | The unit must include a safety mechanism that ensures a minimum concentration of 21% oxygen in any oxygen/air gas mixture across all flow rates.
23.3 | The unit must include a high-flow oxygen flush mechanism with a flow rate of at least 35 l/min
23.4 | Oxygen flush must be non-lockable, recessed and self closing
23.5 | This system must bypass the flowmeters and vaporisers
24 | Breathing System
24.1 | Total circuit volume, including absorber, must not exceed 2,7 litres
24.2 | With APL valve 5 - 75 cm H2O
24.3 | With ventilator to manual switch. This switch when directed to the ventilator must isolate the APL valve from the circle
24.4 | Breathing system must be fully autoclavable to 134°C
24.5 | There must be a mechanism to remove water from the breathing system
24.6 | Breathing system must be latex free
24.7 | Bag-to-ventilator switch automatically begins in the ventilator position
24.8 | Reusable corrugated breathing circuit adequate for both adult and paediatric use to be supplied
24.9 | Bidders to supply the breathing circuit as required by the user
25 | Integrated Ventilator
25.1 | Ventilator suitable for ventilating adults, paediatric and neonates (> 500g) patients without changing ventilator components (e.g. bellows)
25.2 | The ventilator must be electronically controlled
25.3 | The ventilation technology offered may be turbine, volume reflector, piston or ascending bellows
25.4 | The minimum basic modes to be supplied:
25.4.1 | Manual/Spontaneous Ventilation
25.4.2 | Volume Control Ventilation
25.4.3 | Pressure Control Ventilation
25.4.4 | Pressure Support with Mandatory Apnoea back-up ventilation
25.4.5 | Volume Controlled Ventilation + SIMV + Pressure Support
25.4.6 | Pressure Control Ventilation + SIMV + Presssure Support
25.4.7 | CPAP
25.5 | Additional modes of ventilation may be offered and quoted for separately
25.6 | The ventilator must be integrated with the workstation and be user friendly
25.7 | Frequency: minimum range of 5 – 60 bpm.
25.8 | Inspiratory time: minimum range of 0.2 – 10.0 seconds.
25.9 | Minimum Inspiration/Expiration Ratio: 4:1 – 1:4.
25.10 | Adjustable inspiratory pause/plateau time (volume control ventilation): 0-30% of inspiratory time
25.11 | Peak Inspiratory Pressure: minimum range of: PEEP + 5 to 70 cm H2O.
25.12 | Positive End Expiratory Pressure (PEEP): Integrated, electronically controlled. Minimum range: 0 – 30 cm H2O.
25.13 | Tidal Volume: 20ml – 1500ml per breath. Minimum accuracy: 15%
25.14 | Peak Inspiratory Flow: At least 120 lpm
25.15 | The ventilator must be capable of flow or pressure triggering.
25.16 | Adjustable Inspiratory Rise time/Ramp with a range of 0.1 - 2 sec
25.17 | Compliance compensation/Tidal volume compensation/Compressible volume compensation or equivalent.
25.18 | Patient gas flow compensation must be automatic so that alterations in patient gas flow does not affect delivered tidal volumes
15.19 | It must be possible to set CPAP in manual/spontaneous mode within a range of: 1 to 10 cm H2O
25.20 | In event of power or technical failure, manual ventilation must be possible.
26 | Display:
26.1 | All displays must be in acceptable English
26.2 | Colour display must include numerical and graphic display of parameters measured
26.3 | Graphic display to feature at least 3 waveforms: - pressure, flow and volume.
26.4 | Must also display flow-volume and pressure-volume loops with a reference loop
27 | Alarms : The ventilator must have the following user adjustable alarms
27.1 | Apnoea alarm
27.2 | High airway pressure alarm
27.3 | Continuous pressure alarm
27.4 | High and low expiratory tidal volume alarm
27.5 | Minute volume, high/low alarm
27.6 | Inspired oxygen (Fi02), high/low alarm
27.7 | PEEP, low alarm
27.8 | Leakage / Check breathing circuit alarm or equivalent
27.9 | Alarm profile must be user definable or default to patient appropriate settings
27.10 | The ventilator must have a mute alarm function, active only for a limited time.
17.11 | It must not be possible to cancel the visual alarm until the problem has been rectified.
27.12 | All components of the ventilator and breathing system in contact with the patient’s airway must be latex free.
27.13 | All components of the ventilator and breathing system in contact with the patient’s airway must be steam autoclavable up to 134ᴼ C
28 | Integrated Gas Monitoring
28.1 | Oxygen Monitoring
28.1.1 | The system must display both inspired and end-tidal oxygen concentration.
28.1.2 | The oxygen analyser must have user adjustable low and high alarms – visual and audible. The low alarm must automatically default to no less than 21%.
28.1.3 | The system must be accurate (+/- 2% in 21-100%) range, and have a rapid response time (≤ 500ms).
28.1.4 | The unit must sound a continuous alarm if the oxygen pressure falls below 190 kPa, then all nitrous oxide must be automatically cut from the patient circuit.
28.1.5 | Bidder to give details of the oxygen analyser
28.1.6 | The lifespan of the oxygen analyser must be guaranteed for at least 10 years.
28.1.7 | A system using replaceable oxygen fuel cells will not be accepted
28.2 | Anaesthetic vapour, carbon dioxide and nitrous oxide monitoring:
28.2.1 | Anaesthetic vapour, N2O monitoring and capnography must be integrated into the anaesthetic workstation.
28.2.2 | The monitoring must be of the side-stream type.
28.2.3 | Sampling: 100 - 200 ml/min
28.2.4 | Sampled gases must be internally returned to the breathing system.
28.2.5 | CO2 must be presented as a numerical display of inspired and end-tidal concentration and a capnography waveform.The displayed units must be user-selectable: kPa or mmHg.
28.2.6 | The monitoring system must have automatic identification of anaesthetic vapours.
28.2.7 | The system must have a graphical display indicating the type of anaesthetic vapour/s currently in use and the concentration thereof. The colour coding must be as per international colour coding for anaesthetic vapours.
28.2.8 | The anaestheric vapour must be displayed as vol % and must include inspiratory and expiratory concentrations.
28.2.9 | The monitoring system must be able to simultaneously detect, identify and display mixed anaesthetic vapours.
28.2.10 | High and low anaesthetic vapour alarms must be present and these must be user-definable.
28.2.11 | There must be automatic calculation of minimum alveolar concentration.
28.2.12 | The system must be able to calculate anaesthetic gas utilization in ml per case
28.2.13 | The unit must have an indication of the level of vapour in the vaporiser
28.2.14 | Integrated vapour analysis must deliver an audible "low vapour " alarm and a critical alarm for an empty tank
29 | Circle System
29.1 | The circle absorber system must be integrated as part of the ventilation system.
29.2 | The circle absorber system must have an effective system to eliminate excess water created by the chemical reactions either by heating or water traps.
29.3 | The circle absorber system must be autoclavable at 134ºC.
29.4 | The circle absorber system must be supplied with a carbon dioxide absorber canister.
29.5 | The absorber canister may be reusable/refillable
29.6 | It must be easy to remove and replace the absorber canister while the workstation is in operation.
29.7 | Misalignment of the canister must be readily apparent to the user.
29.8 | The system must have a circle bypass facility while changing/refilling the absorber canister.
29.9 | The breathing circuit must contain an integrated adjustable pressure limiting (APL) valve with a minimum range of 0-70 cm H2O.
29.10 | The APL valve must have a tactile or electronic indicator when a value of between 20-30 cm H2O is exceeded.
29.11 | The circle absorber system must be latex free.
29.12 | The bag-to-ventilator switching mechanism must be clearly labeled and user-friendly. Mechanical or electronic
30 | General
30.1 | All components must bear the CE mark or equivalent.
30.2 | The following safety standards must apply: ISO 9001/ ISO 46001; Annex 11 of Directive 93/42/EEC, ISO 13485, I.E.C 60601-2-13
30.3 | The unit must operate from both the 220 V, 50 Hz mains supply and internal rechargeable batteries
30.4 | The batteries must be charged while the unit is operating from the mains supply.
30.5 | The batteries must automatically take over in the event of a mains power failure.
30.6 | The batteries must have the capacity for full operation of the anaesthetic machine and ventilator for at least 90 minutes with batteries fully charged, under typical operating conditions.
30.7 | Graphic battery power indicator
30.8 | Recharge time < 8 hrs
30.9 | Mains cable at least three metres long with SABS approved 15A three prong plug.
30.10 | The unit must have a rail for attaching a suction regulator.
31 | Data management: The workstation must be capable of downloading data to a USB device or memory card.
32 | The unit must be handed over in full operating order
33 | The following accessories must be supplied with the unit:
33.1 | Re-breathing bag, reusable, rubber
33.2 | Reservoir bag
33.3 | Pressure limiting valve, adjustable
33.4 | Oxygen fuel cell analyser
33.5 | Adult & Paediatric volume sensor re-usable
33.6 | Water trap for multigas monitor, re-usable (10 pcs.)
34 | A starter-pack of consumables must be supplied with the unit comprising at least the following:
34.1 | Anaesthesia circuit (Adult, single/double limb design) Disposable. 1.8m (Complete with Y-piece, Elbow, Catheter mount, Rebreathing bag)
34.2 | Anaesthesia circuit (Paediatric, single/double limb design) Disposable. 1.8m (Complete with Y-piece, Elbow, Catheter mount, Rebreathing bag)
34.3 | Anaesthesia circuit (neonatal, single/double limb design) Disposable. 1.8m (Complete with Y-piece, Elbow, Catheter mount, Rebreathing bag)
34.4 | Re-breathing circuit complete
34.5 | Non-re-breathing circuit complete
34.6 | Flow sensor (pack of 2)
34.7 | Spirometry line
34.8 | Sample line (pack of 10)
34.9 | Face masks: sizes 00, 0, 1, 2, 3, 4, 5.
34.10 | Bacterial/viral filter
34.11 | Y connector
34.12 | Manual ventilation hose
34.13 | Fresh gas tubing
34.14 | Seal for back-up cylinder connection on the unit
34.15 | Scavenger tubing
34.16 | Inspiratory and expiratory unidirectional valves
34.17 | Side stream CO2 kit (adult/paed)
34.18 | Disposable Adult Spirometry Set complete with spirometry tube, sample line and flow sensor, 3 m, box of 20 pieces
34.19 | Disposable Paediatric Spirometry Set complete with spirometry tube, sample line and flow sensor, 3 m, box of 20 pieces
34.20 | Gas sample tube with luer lock connection, 3m, disposable (10 pcs)
34.21 | Suction Tubing Set inc connector
34.22 | Resuscitator Adult, 1500mL, autoclavable, silicone
34.23 | Resuscitator Paediatric, 550mL, autoclavable, silicone
34.24 | Tubing set Paediatric for circle systems
35 | List ALL available consumables with prices incl VAT
36 | List ALL available accessories with prices incl VAT
37 | OPTIONAL ITEMS: Any accessories and modules may be offered as optional items
37.1 | Prices incl VAT must be given for all optional items
37.2 | Brochures with full technical specifications must be supplied for each optional item offered
37.3 | Vaporisers
37.3.1 | Isofluorane
37.3.2 | Sevofluorane
37.3.3 | Desfluorane
37.3.4 | Halothane
37.3.5 | Enflurane
37.3.6 | Vaporisers must be tilt, pressure and temperature compensated
37.3.7 | Mechanical, electronic or injection technology vaporisers are acceptable
37.3.8 | Filling mechanism:
37.3.9 | Vaporiser filling mechanism must be compatible with the current state approved vendor of inhalational volatile vapours.
37.3.10 | Each vaporiser must be supplied with 2 compatible filling adapters as per user requirement
37.3.11 | Minimum accuracy of the supplied volatile vapour delivery system must be at least 20% of set value
37.3.12 | Visible level indicator
37.4 | Auxilliary Delivery System
37.4.1 | An external auxilliary flowmeter suitable for delivery of oxygen via a face mask, supplied with a 6mm tapered connection. Flow range 0-10 l/min and above
37.5 | Trending:
37.5.1 | Trend information and event recording for:
37.5.2 | peak pressure
37.5.3 | plateau pressure
37.5.4 | PEEP
37.5.5 | expiratory tidal volume
37.5.6 | respiratory rate
37.5.7 | This information must be provided in graphical and tabular formats.
37.5.8 | The trend and event information must cover at least the last 24 hours of operation.
37.5.9 | Resolution must be adjustable, with a minimum resolution of 1 minute intervals.
37.5.10 | The adjustment of the resolution must not affect the display of parameters which were sampled at different times to the resolution setting in such a way as to display a null reading, when one was in fact obtained.
37.6 | Software
37.6.1 | Software for upgrades may be offered
37.7 | CO2 absorber canister disposable
37.8 | Suction Regulator
37.8.1 | A suction regulator capable of -50 kPa (minimum)
37.8.2 | The suction regulator must have a built-in diaphragm with a 0 – 100 kPa vacuum gauge and an “on-off” switch.
37.8.3 | It must aspirate 1 litre water in 45 seconds at a pressure of at least - 50 kPa
37.8.4 | A collection jar calibrated to at least 1 500 ml must be included.
37.8.5 | The collection jar must accommodate a disposable liner, as per user requirement
37.9 | Neuromuscular transmission monitoring
37.9.1 | The monitor sensor must use piezoelectric technology
37.9.2 | The sensor must be easily fixed to the thumb and index finger.
37.9.3 | The sensor must be available in adult and paediatric sizes.
37.9.4 | The monitor must have the following stimulation modes:
37.9.5 | Train-of-four stimulation (TOF): four stimuli given at a frequency of 0.5 Hz and which can be repeated every 10 s to at least every 30 minutes. The TOF must be displayed graphically and the TOF count and TOF ratio must be displayed.
37.9.6 | Single twitch of a 50 Hz stimulus at a frequency ranging from about 0.1 Hz to 1 Hz
37.9.7 | Tetanic 50Hz for 5 s followed after 3 s by a posttetanic count at 1 Hz. The posttetanic count must be displayed.
37.9.8 | List any other monitoring options, synchronization capabilities, networking or data management capabilities available. State cost of these additional options
38 | MAINTENANCE CONTRACT
38.1 | Bidder must offer a five year full maintenance contract effective after the two year guarantee expires.
38.2 | The cost of the maintenance contract must not be included in the price of the bid
38.3 | State the cost per year and the total cost
38.4 | Provide details of the contract in a separate document

## Anaesthetic 2
| ANAESTHESIA DELIVERY SYSTEM: Advanced with integrated physiological monitor | Complies Y/N | Comment
1 | An advanced anaesthesia delivery system comprising a completely integrated electronic anaesthesia workstation is required.
2 | The unit must include electronicaly integrated oxygen, nitrous oxide and air delivery and monitoring.
3 | Ventilation and anaesthesia vapour monitoring must be fully integrated within the unit.
4 | The patient physiological monitoring must communicate/interface with the anaesthesia unit.
5 | The anaesthesia system must comprise:
5.1 | Anaesthesia system suitable for neonatal (down to 500g), paediatric and adult anaesthesia.
5.2 | Ventilator suitable for ventilating neonatal (down to 500g), paediatric, adult and morbid/obese patients
5.3 | Physiological monitor
5.4 | Multigas analyser:
5.4.1 | Oxygen
5.4.2 | Nitrous oxide
5.4.3 | Carbon dioxide
5.5 | Anaesthetic vapour monitoring:
5.5.1 | Isoflurane
5.5.2 | Sevoflurane
5.5.3 | Desflurane
5.5.4 | Halothane
5.5.5 | Enflurane
6 | Must comply with ISO 13485
7 | Must comply with SANS/IEC 60601-2-13
8 | Complete and automated self test requiring very limited user intervention.
9 | The self test must provide the measurements and values for automatic leakage and compliance.
10 | Automatic calibration of sensors for anaesthetic gases and vapours must form part of the self test and require very limited user intervention.
11 | It must be possible to override the self test facility. The system must display the override status
12 | The unit must provide conventional, low-flow (down to 200 ml/min) anaesthesia.
13 | The unit must return analysed gas back to the circle system
14 | A safety oxygen function is required for the delivery of oxygen through the anaesthesia unit in the event of an electronic failure of the gas delivery system.
15 | The patient monitor must be fixed to the anaesthesia device with a suitable mounting system or integrated into the device.
16 | The unit must be supplied with a minimum of 1 storage drawer
17 | The unit must be supplied with a minimum of 2 additional 220 V power sockets. These sockets must be excluded from the battery backup circuit in case of a power failure
18 | The unit must have a working area or a retractable shelf. Width: not less than 300 mm. Depth: not less than 200 mm
18 | The unit must have a light to enable the anaesthetist to work in a darkened theatre.
20 | Trolley
20.1 | The unit must be mounted on a strong , robust and compact trolley manufactured from corrosion free material.
20.2 | The trolley must have no sharp corners and be easy to clean.
20.3 | The trolley must have at least four antistatic castors of diameter 100 mm or greater.
20.4 | Trolley must have two lockable castors or a central locking brake.
20.5 | The unit must accommodate an oxygen cylinder. Cylinder size according to user requirement. Size to be compatible with current state vendor of oxygen.
20.6 | Trolley must have accessible handles for ease of movement
20.7 | The trolley, with all components mounted, must be stable without toppling over when placed on a 30° slope.
20.8 | The unit must be supplied with a castor guard system, preventing the entaglement and trapping of cables within the wheels.
21 | Pipeline and Cylinder Gas Supplies
21.1 | The anaesthesia system must be designed to use external piped medical gasses (O2, N2O and medical air) and vacuum from medical service outlet sockets.
21.2 | Connection hoses for above must be provided
21.3 | Hoses must be fitted with Heyer type non-interchangeable probes for the existing medical gas outlets as specified by the purchaser.
21.4 | The medical air hose connection or CO2 circle system must include a water trap
21.5 | A safety pressure relief valve of 60 - 80 cm H2O must be provided to protect the system from excessive pressure .
21.6 | The unit must function safely and correctly within the pipeline supply pressure range of: 270 - 600 kPa
21.7 | The system must preferentially utilise the piped oxygen supply over the cylinder supply, if pipeline pressure is > 270 kPa
21.8 | Connection hoses must be:
21.8.1 | Flexible and at least 6 metres long
21.8.2 | Antistatic
21.8.3 | Colour coded as follows:
21.8.3.1 | O2: white
21.8.3.2 | N2O: blue
21.8.3.3 | Vacuum: yellow
21.8.3.4 | Medical air: black/white
21.8.3.5 | Scavenge low vacuum: orange or yellow
21.8.4 | Where relevant, fitted with outlet probes complying with SANS 1409:2014
21.8.5 | Permanently bonded to the outlet probe and the machine connector
21.9 | Hoses for gases above atmospheric pressure, must designed for pressures of at least 1 000 kPa.
21.10 | Vacuum hoses may not collapse or kink at 90 kPa below atmospheric pressure, when subjected to 180° bend of inside radius equal to the outside diameter of the hose.
21.11 | All connection hoses must be fed to a central manifold
21.12 | All high pressure gas and vacuum pipe connections (other than hoses) and all pipes connected to pressure regulators and pressure gauges must be stainless steel (18/8 grade) pipes that are oil-free and colour coded as for hoses (see above) over their entire lengths
21.13 | An in-line non-return valve must be fitted to each gas pipeline on the high pressure (inlet) line
21.14 | An in-line non-return valve must be fitted to each gas cylinder yoke
21.15 | A replaceable gas filter, to block particles greater than 65 µm, must be fitted on the upstream side of the non-return valve
21.16 | Each gas pipeline inlet must have a pressure gauge or electronic display of pressure. The pressure range (415 kPa ± 75 kPa) must be clearly indicated in colour
21.17 | Must be fitted with a gas cylinder yoke to accommodate one oxygen gas cylinder with a pin index system
21.18 | Each gas cylinder yoke must have a pressure gauge or electronic display of pressure
21.19 | The pressure gauges or electronic displays of pressure for both pipeline & cylinder supplies, calibrated in kPa for each line, must be labelled and visible from the front of the unit.
21.20 | O2 and N2O regulators must be provided complying with CKS 605:1987.
21.21 | A pressure regulator complying with CKS 605 must be supplied for each gas inlet and must reduce the gas pressure to 415 kPa without freezing up, creeping, chattering or whistling
21.22 | Outlet sockets and probes must comply with SANS 1409:2014.
21.23 | Sockets with hoses must be colour-coded as per SANS 1409:2014:
21.23.1 | Oxygen: White background with black letters
21.23.2 | Nitrous oxide: Blue background with white letters
21.23.3 | Medical air: One half black with white letters, the other half white with black letters
21.23.4 | Vacuum: Yellow background with black letters
21.24 | Continuous display, visible from the front of the machine, of both pipeline and cylinder pressures even during power failure
21.25 | Pipeline and cylinder pressures must be differentially displayed, and clearly indicated
21.26 | Safety pressure relief valves must protect the entire pneumatic system from excessive input pressures from pipeline and cylinder supplies.
21.27 | When a cylinder is detached from the pin-indexed yoke no gas leak must occur at the yoke.
21.28 | Airway pressure must be monitored and displayed
21.29 | Safety relief valve for all gases must be supplied and set at 600kPa
22 | Scavenging
22.1 | The unit must have a scavenging system with a hose at least 5 metres long with the suitable connector as per the institutions scavenging installation.
22.2 | Air-break device to prevent negative pressure on the breathing system
22.3 | If an active sytem is supplied a pressure indicator must be present
22.4 | The scavenging outlet must be a 30mm ISO connector.
23 | Patient Gas Delivery System
23.1 | Patient gas delivery must be electronically controlled
23.2 | and display the delivered patient gas flow using an ISO colour standard for each gas.
23.3 | The unit must be able to deliver a total patient gas flow of 200 ml/min to 18 l/min (Minimum). The unit must have the ability to deliver at least 10 l/min of oxygen. These must be incremental adjustments
23.4 | Gas supply must conform to:
23.4.1 | O2 minimum flow range of 200 ml/min to 10 l/min. Accuracy ± 5%
23.4.2 | medical air minimum flow range 0 to 10 l/min. Accuracy ± 5%
23.4.3 | nitrous oxide minimum flow range 0 to 10 l/min. Accuracy ± 5%
23.5 | Flowmeter sequence to be according to standards approved by the South Africal Society of Anaesthesiologists - Oxygen on the right
23.6 | The delivery system must prevent simultaneous delivery of two anaesthetic vapours.
23.7 | A separate common gas outlet to accept a standard 22 mm outside diameter and 15 mm internal diameter.
24 | Hypoxic guard
24.1 | The unit must include a safety mechanism that ensures a minimum concentration of 25% oxygen in any oxygen/nitrous oxide gas mixture across all flow rates.
24.2 | The unit must include a safety mechanism that ensures a minimum concentration of 21% oxygen in any oxygen/air gas mixture across all flow rates.
24.3 | The unit must include a high-flow oxygen flush mechanism with a flow rate of at least 35 l/min
24.4 | Oxygen flush must be non-lockable, recessed and self closing
24.5 | This system must bypass the flowmeters and vaporisers
25 | Breathing System
25.1 | Total circuit volume, including absorber, must not exceed 2,7 litres
25.2 | With APL valve 5 - 75 cm H2O
25.3 | With ventilator to manual switch. This switch when directed to the ventilator must isolate the APL valve from the circle
25.4 | There must be a mechanism to remove water from the breathing system
25.5 | Breathing system must be fully autoclavable to 134°C
25.6 | Breathing system must be latex free
25.7 | Bag-to-ventilator switch automatically begins in the ventilator position
25.8 | Reusable corrugated breathing circuit adequate for both adult and paediatric use to be supplied
25.9 | Bidders to supply the breathing circuit as required by the user
26 | Integrated Ventilator
26.1 | Ventilator suitable for ventilating adults, paediatric and neonates (> 500g) patients without changing ventilator components (e.g. bellows)
26.2 | The ventilator must be electronically controlled
26.3 | The ventilation technology offered may be turbine, volume reflector, piston or ascending bellows
26.4 | The minimum basic modes to be supplied:
26.4.1 | Manual/Spontaneous Ventilation
26.4.2 | Volume Control Ventilation
26.4.3 | Pressure Control Ventilation
26.4.4 | Pressure Support with Mandatory Apnoea back-up ventilation
26.4.5 | Volume Controlled Ventilation + SIMV + Pressure Support
26.4.6 | Pressure Control Ventilation + SIMV + Presssure Support
26.4.7 | CPAP
26.5 | Additional modes of ventilation may be offered and quoted for separately
26.6 | The ventilator must be integrated with the workstation and user friendly
26.7 | Frequency: minimum range of 5 – 60 bpm.
26.8 | Inspiratory time: minimum range of 0.2 – 10.0 seconds.
26.9 | Minimum Inspiration/Expiration Ratio: 4:1 – 1:4.
26.10 | Adjustable inspiratory pause/plateau time (volume control ventilation): 0-30% of inspiratory time
26.11 | Peak Inspiratory Pressure: minimum range of: PEEP + 5 to 70 cm H2O.
26.12 | Positive End Expiratory Pressure (PEEP): Integrated, electronically controlled. Minimum range: 0 – 30 cm H2O.
26.13 | Tidal Volume: 20ml – 1500ml per breath. Minimum accuracy: 15%
26.14 | Peak Inspiratory Flow: At least 120 lpm
26.15 | The ventilator must be capable of flow or pressure triggering.
26.16 | Adjustable Inspiratory Rise time/Ramp with a range of 0.1 - 2 sec
26.17 | Compliance compensation/Tidal volume compensation/Compressible volume compensation or equivalent.
26.18 | Patient gas flow compensation must be automatic so that alterations in patient gas flow does not affect delivered tidal volumes
26.19 | It must be possible to set CPAP in manual/spontaneous mode within a range of: 1 to 10 cm H2O
26.20 | In event of power or technical failure, manual ventilation must be possible.
27 | Display:
27.1 | All displays must be in acceptable English
27.1 | Colour display must include numerical and graphic display of parameters measured
27.3 | Graphic display to feature at least 3 waveforms: - pressure, flow and volume.
27.4 | Must also display flow-volume and pressure-volume loops with a reference loop
28 | Alarms : The ventilator must have the following user adjustable alarms:
28.1 | Apnoea alarm
28.2 | High airway pressure alarm
28.3 | Continuous pressure alarm
28.4 | High and low expiratory tidal volume alarm
28.5 | Minute volume, high/low alarm
28.6 | Inspired oxygen (Fi02), high/low alarm
28.7 | PEEP, low alarm
28.8 | Leakage / Check breathing circuit alarm or equivalent
28.9 | Alarm profile must be user definable or default to patient appropriate settings
28.10 | The ventilator must have a mute alarm function, active only for a limited time.
28.11 | It must not be possible to cancel the visual alarm until the problem has been rectified.
28.12 | All components of the ventilator and breathing system in contact with the patient’s airway must be latex free.
28.13 | All components of the ventilator and breathing system in contact with the patient’s airway must be steam autoclavable up to 134ᴼ C
29 | Integrated Gas Monitoring
29.1 | Oxygen Monitoring
29.1.1 | The system must display both inspired and end-tidal oxygen concentration.
29.1.2 | The oxygen analyser must have user adjustable low and high alarms – visual and audible. The low alarm must automatically default to no less than 21%.
29.1.3 | The system must be accurate (+/- 2% in 21-100%) range, and have a rapid response time (≤ 500ms).
29.1.4 | The unit must sound a continuous alarm if the oxygen pressure falls below 190 kPa, then all nitrous oxide must be automatically cut from the patient circuit.
29.1.5 | Bidder to give details of the oxygen analyser
29.1.6 | The lifespan of the oxygen analyser must be guaranteed for at least 10 years.
29.1.7 | A system using replaceable oxygen fuel cells will not be accepted
29.2 | Anaesthetic vapour, carbon dioxide and nitrous oxide monitoring:
29.2.1 | Anaesthetic vapour, N2O monitoring and capnography must be integrated into the anaesthetic workstation.
29.2.2 | The monitoring must be of the side-stream type.
29.2.3 | Sampling: 100 - 200 ml/min
29.2.4 | Sampled gases must be internally returned to the breathing system.
29.2.5 | There must be no water build-up in the measuring line.
29.2.6 | CO2 must be presented as a numerical display of inspired and end-tidal concentration and a capnography waveform.The displayed units must be user-selectable: kPa or mmHg.
29.2.7 | The monitoring system must have automatic identification of anaesthetic vapours.
29.2.8 | The system must have a graphical display indicating the type of anaesthetic vapour/s currently in use and the concentration thereof. The colour coding must be as per international colour coding for anaesthetic vapours.
29.2.9 | The anaestheric vapour must be displayed as vol % and must include inspiratory and expiratory concentrations.
29.2.10 | The monitoring system must be able to simultaneously detect, identify and display mixed anaesthetic vapours.
29.2.11 | High and low anaesthetic vapour alarms must be present and these must be user-definable.
29.2.12 | There must be automatic calculation of minimum alveolar concentration.
29.2.13 | The system must be able to calculate anaesthetic gas utilization in ml per case
29.2.14 | The unit must have an indication of the level of vapour in the vaporiser
29.2.15 | Integrated vapour analysis must deliver an audible "low vapour " alarm and a critical alarm for an empty tank
30 | Circle System
30.1 | The circle absorber system must be integrated as part of the ventilation system.
30.2 | The circle absorber system must have an effective system to eliminate excess water created by the chemical reactions either by heating or water traps.
30.3 | The circle absorber system must be autoclavable at 134ºC.
30.4 | The circle absorber system must be supplied with a non-disposable CO2 absorber canister
30.5 | The absorber canister may be reusable/refillable
30.6 | It must be easy to remove and replace the absorber canister while the workstation is in operation.
30.7 | Misalignment of the canister must be readily apparent to the user.
30.8 | The system must have a circle bypass facility while changing/refilling the absorber canister.
30.9 | The breathing circuit must contain an integrated adjustable pressure limiting (APL) valve with a minimum range of 0-70 cm H2O.
30.10 | The APL valve must have a tactile or electronic indicator when a value of between 20-30 cm H2O is exceeded.
30.11 | The circle absorber system must be latex free.
30.12 | The bag-to-ventilator switching mechanism must be clearly labeled and user-friendly. Mechanical or electronic
31 | Patient Physiological Monitoring
31.1 | The patient monitor must be modular and suitable for use for neonatal, paediatric and adult patients.
31.2 | The following paramaters must be monitored:
31.2.1 | - ECG: 3-lead and 5-lead. Must be capable of displaying 2 channels simultaneously.
31.2.2 | - ST-segment and arrhythmia analysis.
31.2.3 | - NIBP
31.2.4 | - Oxygen saturation (SpO2)
31.2.5 | - Plethysmography/respiratory rate
31.2.6 | - Temperature x 2 (minimum)
31.2.7 | - Invasive pressure x 4 (minimum)
31.3 | The patient monitor must be fully integrated with the workstation.
31.4 | Display must be no less than 400 mm diagonally.
31.5 | Minimum 6 waveforms must be able to be displayed simultaneously.
31.6 | Full colour display.
31.7 | The monitor must be capable of displaying trends for at least 24 hours in tabular and graphical format.
31.8 | The trend resolution must be user-adjustable and have a minimum resolution of 1 min.
31.9 | The front panel must be easy to clean and disinfect.
31.10 | The display must be flicker free and easy to read.
31.11 | Brightness control must be possible.
31.12 | In addition to the waveform display section, the monitor must have alphanumeric display of parameter identification, measured values, alarm and status messages.
31.13 | Audible and visual alarms must be available for all measured parameters.
31.14 | Default alarm limits if the limits are not set by the user
31.15 | Alarm indication: Visual and audible with mute
31.16 | Adjustable audio alarm volume
31.17 | Low battery: Alarm at least 10 minutes before battery power is exhausted
31.18 | It must be possible to mute the audible alarm for a short period, but no override of the visual alarm must be possible.
31.19 | It must be possible to suspend the alarms of certain parameters during clinical procedures, but the alarm must reactivate automatically after a suitable time delay. Any silencing or disabling of alarms must be clearly indicated on the monitor.
31.20 | Input impedance of the system > 10 MΩ.
31.21 | Maximum allowable leakage current to the patient must be less than 10 µA.
31.22 | The monitor must be protected against defibrillator and electrosurgical interference.
31.23 | The monitor must be of low power consumption, and must not require fan or other forced ventilation.
31.24 | All cables and sockets must be colour coded for easy recognition, and must conform to international standards
32 | ECG
32.1 | A full set (adult and paediatric) of 3/5 lead ECG cables must be included in the offer.
32.2 | The ECG amplifier must employ a differential input with high noise suppression.
32.3 | Sweep speed: at least 25 to 50 mm/sec.
32.4 | Freeze facility for ECG.
32.5 | The input impedance must be > 10 MΩ.
32.6 | The heart rate measuring range: 20 to 300 bpm, with an accuracy < 5 bpm.
32.7 | An 50 Hz filter must be provided.
32.8 | Alarm limits must be adjustable in the following range:
32.9 | Low: 40 to 150
32.10 | High: 70 to 250.
33 | Non Invasive Blood Pressure
33.1 | Must be suitable for adults, paeds and neonates
33.2 | The unit must use the oscillometric method of measurement.
33.3 | Modes: Manual and automatic
33.4 | Reusable neonate, paediatric, adult and obese cuffs must be supplied
33.5 | The cuffs must be robust, interchangeable and easy to clean.
33.6 | Must display systolic, diastolic and mean pressures for each measurement
33.7 | Adult: Systolic: 30 - 270 mm Hg, Diastolic: 20 - 240 mm Hg, MAP: 20 - 250 mm Hg
33.8 | Paeds: Systolic: 30 - 160 mm Hg, Diastolic: 20 - 120, MAP: 20 - 120 mm Hg
33.9 | Values must be displayed numerically
33.10 | Auto zero when unit is switched on
33.11 | The interval between measurements must be selectable from 1 – 120 minutes.
33.12 | Adjustable high and low alarm limits for systolic, diastolic, mean arterial pressure and heart rate.
33.13 | Over pressure protection: Automatic deflation at 300 mm Hg adult, 150 mm Hg neonate
33.14 | Average measurement time: approximately 30 s
33.15 | If measurement exceeds two minutes the cuff must deflate, an alarm must sound and likely faults indicated.
33.16 | The air hose must be at least two metres long.
34 | Oxygen Saturation.
34.1 | Continuous monitoring of oxygen saturation levels by means of the photometric principle.
34.2 | Must be suitable for adults, paeds and neonates
34.3 | The unit must continuously display the SaO2 % waveform
34.4 | Oxygen saturation (SaO2 %): Range of measurement: 0 - 100%
34.5 | The accuracy must be better than 2,5 % (70 - 100%) and 5% (50 - 70%).
34.6 | Pulse rate (beats/min): Range of measurement: 20 - 300 bpm
34.7 | The pulse rate accuracy must be better than ± 2 %.
34.8 | An audible bleep should sound for each heartbeat, the pitch of which should vary with the oxygen saturation.
34.9 | Automatic rejection of motion artefacts
35 | SpO2 Sensor:
35.1 | Sensor: Dual LED, Infra-red/visible red, which does not heat the point of application to the patient.
35.2 | The following reusable sensors must be provided:
35.3 | Finger sensor for adults and paeds.
35.4 | Earlobe sensor for adults and paeds.
35.5 | Finger wrap-around probe for neonates
35.6 | The sensor lead must be approximately 3 metres long.
36 | Temperature
36.1 | Temperature range must be approximately 15º to 45º C.
36.2 | Two reusable temperature probes must be included
37 | General
37.1 | All components must bear the CE mark or equivalent.
37.2 | The following safety standards must apply: ISO 9001/ ISO 46001; Annex 11 of Directive 93/42/EEC, ISO 13485, I.E.C 60601-2-13
37.3 | The unit must operate from both the 220 V, 50 Hz mains supply and internal rechargeable batteries
37.4 | The batteries must be charged while the unit is operating from the mains supply.
37.5 | The batteries must automatically take over in the event of a mains power failure.
37.6 | The batteries must have the capacity for full operation of the anaesthetic machine, ventilator and monitor for at least 90 minutes with batteries fully charged, under typical operating conditions.
37.7 | Graphic battery power indicator
37.8 | Recharge time < 8 hrs
37.9 | Mains cable at least three metres long with SABS approved 15A three prong plug.
37.10 | The unit must have a rail for attaching a suction regulator.
38 | Data management:
38.1 | The workstation must be capable of downloading data to a USB device or memory card.
38.2 | The monitor must be able to communicate with a central station
38.3 | A printer must be incorporated in the unit
39 | The unit must be handed over in full operating order
40 | The following accessories must be supplied with the unit:
40.1 | Re-breathing bag, reusable, rubber
40.2 | Reservoir bag
40.3 | Pressure limiting valve, adjustable
40.4 | Oxygen fuel cell analyser
40.5 | Adult & Paediatric volume sensor re-usable
40.6 | Water trap for multigas monitor, re-usable (10 pcs.)
40.7 | A full set paediatric 3/5 lead ECG cables
40.8 | A full set adult 3/5 lead ECG cables
40.9 | Reusable paediatric cuffs
40.10 | Reusable adult cuffs
40.11 | Reusable neonate cuffs
40.12 | Reusable obese cuffs
40.13 | Air hose
40.14 | Finger sensor for paeds
40.15 | Finger sensor for adults
40.16 | Earlobe sensor for paeds
40.17 | Earlobe sensor for adults
40.18 | Finger wrap-around probe for neonates
40.19 | Sensor lead
40.20 | Two reusable temperature probes
41 | A starter-pack of consumables must be supplied with the unit comprising at least the following:
41.1 | Anaesthesia circuit (Adult, single/double limb design) Disposable. 1.8m (Complete with Y-piece, Elbow, Catheter mount, Rebreathing bag)
41.2 | Anaesthesia circuit (Paediatric, single/double limb design) Disposable. 1.8m (Complete with Y-piece, Elbow, Catheter mount, Rebreathing bag)
41.3 | Anaesthesia circuit (neonatal, single/double limb design) Disposable. 1.8m (Complete with Y-piece, Elbow, Catheter mount, Rebreathing bag)
41.4 | Re-breathing circuit complete
41.5 | Non-re-breathing circuit complete
41.6 | Flow sensor (pack of 2)
41.7 | Spirometry line
41.8 | Sample line (pack of 10)
41.9 | Face masks: sizes 00, 0, 1, 2, 3, 4, 5.
41.10 | Bacterial/viral filter
41.11 | Y connector
41.12 | Manual ventilation hose
41.13 | Fresh gas tubing
41.14 | Seal for back-up cylinder connection on the unit
41.15 | Scavenger tubing
41.16 | Inspiratory and expiratory unidirectional valves
41.17 | Side stream CO2 kit (adult/paed)
41.18 | Disposable Adult Spirometry Set complete with spirometry tube, sample line and flow sensor, 3 m, box of 20 pieces
41.19 | Disposable Paediatric Spirometry Set complete with spirometry tube, sample line and flow sensor, 3 m, box of 20 pieces
41.20 | Gas sample tube with luer lock connection, 3m, disposable (10 pcs)
41.21 | Suction Tubing Set inc connector
41.22 | Resuscitator Adult, 1500mL, autoclavable, silicone
41.23 | Resuscitator Paediatric, 550mL, autoclavable, silicone
41.24 | Tubing set Paediatric for circle systems
42 | List ALL available consumables with prices incl VAT
43 | List ALL available accessories with prices incl VAT
44 | OPTIONAL ITEMS: Any accessories and modules may be offered as optional items
44.1 | Prices incl VAT must be given for all optional items
44.2 | Brochures with full technical specifications must be supplied for each optional item offered
44.3 | Vaporisers
44.3.1 | Isofluorane
44.3.2 | Sevofluorane
44.3.3 | Desfluorane
44.3.4 | Halothane
44.3.5 | Enflurane
44.3.6 | Vaporisers must be tilt, pressure and temperature compensated
44.3.7 | Mechanical, electronic or injection technology vaporisers are acceptable
44.3.8 | Filling mechanism:
44.3.9 | Vaporiser filling mechanism must be compatible with the current state approved vendor of inhalational volatile vapours.
44.3.10 | Each vaporiser must be supplied with 2 compatible filling adapters as per user requirement
44.3.11 | Minimum accuracy of the supplied volatile vapour delivery system must be at least 20% of set value
44.3.12 | Visible level indicator
44.4 | AIMS (Anaesthesia information management system)
44.5 | Auxilliary Delivery System
44.5.1 | An external auxilliary flowmeter suitable for delivery of oxygen via a face mask, supplied with a 6mm tapered connection. Flow range 0-10 l/min and above
44.6 | Trending:
44.6.1 | Trend information and event recording for:
44.6.2 | peak pressure
44.6.3 | plateau pressure
44.6.4 | PEEP
44.6.5 | expiratory tidal volume
44.6.6 | respiratory rate
44.6.7 | This must be provided in graphical and tabular formats.
44.6.8 | The trend and event information must cover at least the last 24 hours of operation.
44.6.9 | Resolution must be adjustable, with a minimum resolution of 1 minute intervals.
44.6.10 | The adjustment of the resolution must not affect the display of parameters which were sampled at different times to the resolution setting in such a way as to display a null reading, when one was in fact obtained.
44.7 | Software
44.7.1 | Software for upgrades may be offered
44.8 | CO2 absorber canister disposable
44.9 | Suction Regulator
44.9.1 | A suction regulator capable of -50 kPa (minimum)
44.9.2 | The suction regulator must have a built-in diaphragm with a 0 – 100 kPa vacuum gauge and an “on-off” switch.
44.9.3 | It must aspirate 1 litre water in 45 seconds at a pressure of at least - 50 kPa
44.9.4 | A collection jar calibrated to at least 1 500 ml must be included.
44.9.5 | The collection jar must accommodate a disposable liner, as per user requirement
44.10 | Neuromuscular transmission monitoring
44.10.1 | The monitor sensor must use piezoelectric technology
44.10.2 | The sensor must be easily fixed to the thumb and index finger.
44.10.3 | The sensor must be available in adult and paediatric sizes.
44.10.4 | The monitor must have the following stimulation modes:
44.10.5 | Train-of-four stimulation (TOF): four stimuli given at a frequency of 0.5 Hz and which can be repeated every 10 s to at least every 30 minutes. The TOF must be displayed graphically and the TOF count and TOF ratio must be displayed.
44.10.6 | Single twitch of a 50 Hz stimulus at a frequency ranging from about 0.1 Hz to 1 Hz
44.10.7 | Tetanic 50Hz for 5 s followed after 3 s by a posttetanic count at 1 Hz. The posttetanic count must be displayed.
44.11 | Metabolic monitoring
44.12 | Depth of anaesthesia monitoring system
44.12.1 | Processed EEG monitoring system including bipolar quantitative EEG, Bispectral index, and spectral entropy or equivalent.
44.12.2 | The monitor must display burst suppression ratio.
44.13 | Invasive cardiac output monitoring
44.14 | Non-invasive cardiac output monitoring
44.15 | Cerebral oximetry
45 | MAINTENANCE CONTRACT
45.1 | Bidder must offer a five year full maintenance contract effective after the two year guarantee expires.
45.2 | The cost of the maintenance contract must not be included in the price of the bid
45.3 | State the cost per year and the total cost
45.4 | Provide details of the contract in a separate document

## Anaesthetic Std 1
| ANAESTHESIA DELIVERY SYSTEM: Standard without physiological monitor | Complies Y/N | Comment
1 | Anaesthesia delivery system suitable for paediatric and adult anaesthesia
2 | Ventilator suitable for ventilating paediatric to adult patients
3 | Multigas analyser:
3.1 | Oxygen
3.2 | Nitrous oxide
3.3 | Carbon dioxide
4 | Anaesthetic agent monitoring:
4.1 | Isofluorane
4.2 | Sevofluorane
4.3 | Desfluorane
4.4 | Halothane
4.5 | Enflurane
5 | Must comply with ISO 13485
6 | Must comply with SANS/IEC 60601-2-13
7 | Self test requiring limited user intervention.
8 | The self test must provide the measurements and values for automatic leak detection and compliance.
9 | Automatic calibration of sensors for gases and anaesthetic vapours must form part of the self test and require limited user intervention.
10 | The unit must provide conventional, low-flow (down to 200 ml/min) anaesthesia.
11 | It must be possible to override the self test facility. The system must display the override status
12 | The unit must return analysed gas back to the circle system
13 | The unit must be supplied with a minimum of 1 storage drawer
14 | The unit must be supplied with a minimum of 3 additional 220 V power sockets. These sockets must be excluded from the battery backup circuit in case of a power failure
15 | The unit must have a working area or a retractable shelf. Width: not less than 300 mm. Depth: not less than 200 mm
16 | The unit must have a light to enable the anaesthetist to work in a darkened theatre.
17 | Trolley
17.1 | The unit must be mounted on a strong , robust and compact trolley manufactured from corrosion free material.
17.2 | The trolley must have no sharp corners and be easy to clean.
17.3 | The trolley must have at least four antistatic castors of diameter 100 mm or greater.
17.4 | Trolley must have two lockable castors or a central locking brake.
17.5 | The unit must accommodate an oxygen cylinder. Cylinder size according to user requirement. Size to be compatible with current state vendor of oxygen.
17.6 | Trolley must have accessible handles for ease of movement
17.7 | The trolley, with all components mounted, must be stable without toppling over when placed on a 30° slope.
17.8 | The unit must be supplied with a castor guard system, preventing the entaglement and trapping of cables within the wheels.
17.9 | The trolley must have a mounting system suitable for a physiologial monitor.
18 | Pipeline and Cylinder Gas Supplies
18.1 | The anaesthesia system must be designed to use external piped medical gasses (O2, N2O and medical air) and vacuum from medical service outlet sockets.
18.2 | Connection hoses for above must be provided
18.3 | Hoses must be fitted with Heyer type non-interchangeable probes for the existing medical gas outlets as specified by the purchaser.
18.4 | The medical air hose connection or CO2 circle system must include a water trap
18.5 | A safety pressure relief valve of 60 - 80 cm H2O must be provided to protect the system from excessive pressure .
18.6 | The unit must function safely and correctly within the pipeline supply pressure range of: 270 - 600 kPa
18.7 | The system must preferentially utilise the piped oxygen supply over the cylinder supply, if pipeline pressure is > 270 kPa
18.8 | Connection hoses must be:
18.8.1 | Flexible and at least 6 metres long
18.8.2 | Antistatic
18.8.3 | Colour coded as follows:
18.8.3.1 | O2: white
18.8.3.2 | N2O: blue
18.8.3.3 | Vacuum: yellow
18.8.3.4 | Medical air: black
18.8.3.5 | Scavenge low vacuum: orange or yellow
18.8.4 | Where relevant, fitted with outlet probes complying with SANS 1409:2014
18.8.5 | Permanantly bonded to the outlet probe and the machine connector
18.9 | Hoses for gases above atmospheric pressure, must designed for pressures of at least 1 000 kPa.
18.10 | Vacuum hoses may not collapse or kink at 90 kPa below atmospheric pressure, when subjected to 180° bend of inside radius equal to the outside diameter of the hose.
18.11 | All connection hoses must be fed to a central manifold
18.12 | All high pressure gas and vacuum pipe connections (other than hoses) and all pipes connected to pressure regulators and pressure gauges must be stainless steel (18/8 grade) pipes that are oil-free and colour coded as for hoses (see above) over their entire lengths
18.13 | An in-line non-return valve must be fitted to each gas pipeline on the high pressure (inlet) line
18.14 | An in-line non-return valve must be fitted to each gas cylinder yoke
18.15 | A replaceable gas filter, to block particles greater than 65 µm, must be fitted on the upstream side of the non-return valve
18.16 | Each gas pipeline inlet must have a pressure gauge or electronic display of pressure. The pressure range (415 kPa ± 75 kPa) must be clearly indicated in colour
18.17 | Must be fitted with a gas cylinder yoke to accommodate one oxygen gas cylinder with a pin index system
18.18 | Each gas cylinder yoke must have a pressure gauge or electronic display of pressure
18.19 | The pressure gauges or electronic displays of pressure for both pipeline & cylinder supplies, calibrated in kPa for each line, must be labelled and visible from the front of the unit.
18.20 | O2 and N2O regulators must be provided complying with CKS 605:1987.
18.21 | A pressure regulator complying with CKS 605 must be supplied for each gas inlet and must reduce the gas pressure to 415 kPa without freezing up, creeping, chattering or whistling
18.22 | Outlet sockets and probes must comply with SANS 1409:2014.
18.23 | Sockets with hoses must be colour-coded as per SANS 1409:2014:
18.24 | Oxygen: White background with black letters
18.25 | Nitrous oxide: Blue background with white letters
18.26 | Medical air: One half black with white letters, the other half white with black letters
18.27 | Vacuum: Yellow background with black letters
18.28 | Continuous display, visible from the front of the machine, of both pipeline and cylinder pressures even during power failure
18.29 | Pipeline and cylinder pressures must be differentially displayed, and clearly indicated
18.30 | Safety pressure relief valves must protect the entire pneumatic system from excessive input pressures from pipeline and cylinder supplies.
18.31 | When a cylinder is detached from the pin-indexed yoke no gas leak must occur at the yoke.
18.32 | Airway pressure must be monitored and displayed
18.33 | Safety relief valve for all gases must be supplied and set at 600kPa
19 | Scavenging
19.1 | The unit must have a scavenging system with a hose at least 5 metres long with the suitable connector as per the institutions scavenging installation.
19.2 | Air-break device to prevent negative pressure on the breathing system
19.3 | If an active sytem is supplied a pressure indicator must be present
19.4 | The scavenging outlet must be a 30mm ISO connector.
20 | Patient Gas Delivery System (Flowmeters/Rotameters)
20.1 | The unit must have cascading flowmeters (mechanical or electronic) that show flow rates for high and low flow anaesthesia in the following ranges:
20.2 | Oxygen low flow: 200 - 1 000 ml/min
20.3 | At higher flow rates the gas must cascade into a second flowmeter (1 – 10 l/min)
20.4 | Nitrous oxide low flow: 200 - 1 000 ml/min
20.5 | At higher flow rates the gas must cascade into a second flowmeter (1 – 10 l/min)
20.6 | Air tube 200ml/min – 10 l/min
20.7 | Flowmeters, if not electronic, must be diameter indexed to prevent incorrect replacement during maintenance.
20.8 | Gas supply must conform to:
20.8.1 | O2 minimum flow range of 200 ml/min to 10 l/min. Accuracy ± 5%
20.8.2 | medical air minimum flow range 0 to 10 l/min. Accuracy ± 5%
20.8.3 | nitrous oxide minimum flow range 0 to 10 l/min. Accuracy ± 5%
20.9 | Flowmeters must be accurate to at least 5%
20.10 | Flowmeter valve control
20.11 | Flowmeter control knobs
20.12 | Flowmeter sequence to be according to standards approved by the South Africal Society of Anaesthesiologists - Oxygen on the right
20.13 | The delivery system must prevent simultaneous delivery of two anaesthetic vapours.
20.14 | A separate common gas outlet to accept a standard 22 mm outside diameter and 15 mm internal diameter.
21 | Hypoxic guard
21.1 | The unit must include a safety mechanism that ensures a minimum concentration of 25% oxygen in any oxygen/nitrous oxide gas mixture across all flow rates.
21.2 | The unit must include a safety mechanism that ensures a minimum concentration of 21% oxygen in any oxygen/air gas mixture across all flow rates.
21.3 | The unit must include a high-flow oxygen flush mechanism with a flow rate of at least 35 l/min
21.4 | Oxygen flush must be non-lockable, recessed and self closing
21.5 | This system must bypass the flowmeters and vaporisers
22 | Breathing System
22.1 | Total circuit volume, including absorber, must not exceed 2,7 litres
22.2 | With APL valve 5 - 75 cm H2O
22.3 | With ventilator to manual switch. This switch when directed to the ventilator must isolate the APL valve from the circle
22.4 | Breathing system must be fully autoclavable to 134°C
22.5 | There must be a mechanism to remove water from the breathing system
22.6 | Breathing system must be latex free
22.7 | Bag-to-ventilator switch automatically begins in the ventilator position
22.8 | Reusable corrugated breathing circuit adequate for both adult and paediatric use to be supplied
22.9 | Bidders to supply the breathing circuit as required by the user
22.10 | Bellows must be visible
23 | Integrated Ventilator
23.1 | Ventilator suitable for ventilating adults and paediatric patients without changing ventilator components (e.g. bellows)
23.2 | The ventilator must be electronically controlled
23.3 | The ventilation technology offered may be turbine, volume reflector, piston or ascending bellows
23.4 | The minimum basic modes to be supplied:
23.4.1 | Manual/Spontaneous Ventilation
23.4.2 | Volume Control Ventilation
23.4.3 | Pressure Control Ventilation
23.5 | Additional modes of ventilation may be offered and quoted for separately
23.6 | The ventilator must be integrated with the workstation and user friendly
23.7 | Frequency: minimum range of 5 – 60 bpm.
23.8 | Inspiratory time: minimum range of 0.2 – 10.0 seconds.
23.9 | Minimum Inspiration/Expiration Ratio: 4:1 – 1:4.
23.10 | Adjustable inspiratory pause/plateau time (volume control ventilation): 0-30% of inspiratory time
23.11 | Peak Inspiratory Pressure: minimum range of: PEEP + 5 to 70 cm H2O.
23.12 | Positive End Expiratory Pressure (PEEP): Integrated, electronically controlled. Minimum range: 0 – 30 cm H2O.
23.13 | Tidal Volume: 20ml – 1500ml per breath. Minimum accuracy: 15%
23.14 | Peak Inspiratory Flow: At least 120 lpm
23.15 | The ventilator must be capable of flow or pressure triggering.
23.16 | Adjustable Inspiratory Rise time/Ramp with a range of 0.1 - 2 sec
23.17 | Compliance compensation/Tidal volume compensation/Compressible volume compensation or equivalent.
23.18 | Patient gas flow compensation must be automatic so that alterations in patient gas flow do not affect delivered tidal volumes
23.19 | In event of power or technical failure, manual ventilation must be possible.
24 | Display:
24.1 | All displays must be in acceptable English
24.2 | Colour display must include numerical and graphic display of parameters measured
24.3 | Graphic display to feature at least 3 waveforms: - pressure, flow and volume.
24.4 | Must also display flow-volume and pressure-volume loops with a reference loop
25 | Alarms : The ventilator must have the following user adjustable alarms
25.1 | Apnoea alarm
25.2 | High airway pressure alarm
25.3 | Continuous pressure alarm
25.4 | High and low expiratory tidal volume alarm
25.5 | Minute volume, high/low alarm
25.6 | Inspired oxygen (Fi02), high/low alarm
25.7 | PEEP, low alarm
25.8 | Leakage / Check breathing circuit alarm or equivalent
25.9 | Alarm profile must be user definable or default to patient appropriate settings
25.10 | The ventilator must have a mute alarm function, active only for a limited time.
26 | It must not be possible to cancel the visual alarm until the problem has been rectified.
27 | All components of the ventilator and breathing system in contact with the patient’s airway must be latex free.
28 | All components of the ventilator and breathing system in contact with the patient’s airway must be steam autoclavable up to 134ᴼ C
29 | Integrated Gas Monitoring
29.1 | Oxygen Monitoring
29.1.1 | The system must display both inspired and end-tidal oxygen concentration.
29.1.2 | The oxygen analyser must have user adjustable low and high alarms – visual and audible. The low alarm must automatically default to no less than 21%.
29.1.3 | The system must be accurate (+/- 2% in 21-100%) range, and have a rapid response time (≤ 500ms).
29.1.4 | The unit must sound a continuous alarm if the oxygen pressure falls below 190 kPa, then all nitrous oxide must be automatically cut from the patient circuit.
29.2 | Circle System
29.2.1 | The circle absorber system must be integrated as part of the ventilation system.
29.2.2 | The circle absorber system must have an effective system to eliminate excess water created by the chemical reactions either by heating or water traps.
29.2.3 | The circle absorber system must be autoclavable at 134ºC.
29.2.4 | The circle absorber system must be supplied with a carbon dioxide absorber canister.
29.2.5 | The absorber canister may be reusable/refillable
29.2.6 | It must be easy to remove and replace the absorber canister while the workstation is in operation.
29.2.7 | Misalignment of the canister must be readily apparent to the user.
29.2.8 | The system must have a circle bypass facility while changing/refilling the absorber canister.
29.2.9 | The breathing circuit must contain an integrated adjustable pressure limiting (APL) valve with a minimum range of 0-70 cm H2O.
29.2.10 | The APL valve must have a tactile or electronic indicator when a value of between 20-30 cm H2O is exceeded.
29.2.11 | The circle absorber system must be latex free.
29.2.12 | The bag-to-ventilator switching mechanism must be clearly labeled and user-friendly. Mechanical or electronic
30 | Patient circuit pressure release valve
30.1 | Adjustable up to 70 cm water
30.2 | Audible alarm when the valve opens
30.3 | Manual override, non-lockable, to keep the release valve closed when the set pressure is exceeded.
31 | Flowmeter pressure release valve
31.1 | Situated between the flowmeters and the vaporisers
31.2 | Must open at 105 kPa +- 5 kPa
33 | General
33.1 | All components must bear the CE mark or equivalent.
33.2 | The following safety standards must apply: ISO 9001/ ISO 46001; Annex 11 of Directive 93/42/EEC, ISO 13485, I.E.C 60601-2-13
33.3 | The unit must operate from both the 220 V, 50 Hz mains supply and internal rechargeable batteries
33.4 | The batteries must be charged while the unit is operating from the mains supply.
33.5 | The batteries must automatically take over in the event of a mains power failure.
33.6 | The batteries must have the capacity for full operation of the anaesthetic machine and ventilator for at least 90 minutes with batteries fully charged, under typical operating conditions.
33.7 | Graphic battery power indicator
33.8 | Recharge time < 8 hrs
33.9 | Mains cable at least three metres long with SABS approved 15A three prong plug.
33.10 | The unit must have a rail for attaching a suction regulator.
33.11 | Mains cable at least three metres long with SABS approved 15A three prong plug.
33.12 | The unit must be handed over in full operating order
34 | The following accessories must be supplied with the unit:
34.1 | Re-breathing bag, reusable, rubber
34.2 | Reservoir bag
34.3 | Pressure limiting valve, adjustable
34.4 | Oxygen fuel cell analyser
34.5 | State the cost and frequency of replacement for the flow sensors.
34.6 | State the cost and frequency of replacement for the fuel cell.
35 | A starter-pack of consumables must be supplied with the unit comprising at least the following:
35.1 | Anaesthesia circuit (Adult, single/double limb design) Disposable. 1.8m (Complete with Y-piece, Elbow, Catheter mount, Rebreathing bag)
35.2 | Anaesthesia circuit (Paediatric, single/double limb design) Disposable. 1.8m (Complete with Y-piece, Elbow, Catheter mount, Rebreathing bag)
35.3 | Anaesthesia circuit (neonatal, single/double limb design) Disposable. 1.8m (Complete with Y-piece, Elbow, Catheter mount, Rebreathing bag)
35.4 | Re-breathing circuit complete
35.5 | Non-re-breathing circuit complete
35.6 | Flow sensor (pack of 2)
35.7 | Spirometry line
35.8 | Sample line (pack of 10)
35.9 | Face masks: sizes 00, 0, 1, 2, 3, 4, 5.
35.10 | Bacterial/viral filter
35.11 | Y connector
35.12 | Manual ventilation hose
35.13 | Fresh gas tubing
35.14 | Seal for back-up cylinder connection on the unit
35.15 | Scavenger tubing
35.16 | Inspiratory and expiratory unidirectional valves
35.17 | Side stream CO2 kit (adult/paed)
36 | List ALL available consumables with prices incl VAT
37 | List ALL available accessories with prices incl VAT
38 | OPTIONAL ITEMS
38.1 | Prices incl VAT must be given for all optional items
38.2 | Brochures with full technical specifications must be supplied for each optional item offered
38.3 | Vaporisers
38.3.1 | Isofluorane
38.3.2 | Sevofluorane
38.3.3 | Desfluorane
38.3.4 | Halothane
38.3.5 | Enflurane
38.3.6 | Vaporisers must be tilt, pressure and temperature compensated
38.3.7 | Mechanical, electronic or injection technology vaporisers are acceptable
38.3.8 | Filling mechanism:
38.3.9 | Vaporiser filling mechanism must be compatible with the current state approved vendor of inhalational volatile vapours.
38.3.10 | Each vaporiser must be supplied with 2 compatible filling adapters as per user requirement
38.3.11 | Minimum accuracy of the supplied volatile vapour delivery system must be at least 20% of set value
38.3.12 | Visible level indicator
38.4 | Suction Regulator
38.4.1 | A suction regulator capable of -50 kPa (minimum)
38.4.2 | The unit must have a rail for attaching the suction regulator.
38.4.3 | The suction regulator must have a built-in diaphragm with a 0 – 100 kPa vacuum gauge and an “on-off” switch.
38.4.4 | It must aspirate 1 litre water in 45 seconds at a pressure of at least - 50 kPa
38.4.5 | A collection jar calibrated to at least 1 500 ml must be supplied.
38.4.6 | The collection jar must accommodate a disposable liner, as per user requirement
38.5 | Anaesthetic vapour, carbon dioxide and nitrous oxide monitoring:
38.5.1 | Anaesthetic vapour, N2O monitoring and capnography must be integrated into the anaesthetic workstation.
38.5.2 | The monitoring must be of the side-stream type.
38.5.3 | Sampling: 100 - 200 ml/min
38.5.4 | Sampled gases must be internally returned to the breathing system.
38.5.5 | CO2 must be presented as a numerical display of inspired and end-tidal concentration and a capnography waveform.The displayed units must be user-selectable: kPa or mmHg.
38.5.6 | The monitoring system must have automatic identification of anaesthetic vapours.
38.5.7 | The system must have a graphical display indicating the type of anaesthetic vapour/s currently in use and the concentration thereof. The colour coding must be as per international colour coding for anaesthetic vapours.
38.5.8 | The anaestheric vapour must be displayed as vol % and must include inspiratory and expiratory concentrations.
38.5.9 | The monitoring system must be able to simultaneously detect, identify and display mixed anaesthetic vapours.
38.5.10 | High and low anaesthetic vapour alarms must be present and these must be user-definable.
38.5.11 | There must be automatic calculation of minimum alveolar concentration.
38.5.12 | The system must be able to calculate anaesthetic gas utilization in ml per case
38.5.13 | The unit must have an indication of the level of vapour in the vaporiser
38.5.14 | Integrated vapour analysis must deliver an audible "low vapour " alarm and a critical alarm for an empty tank
38.6 | Ventilator modes
38.6.1 | Pressure Support with Mandatory Apnoea back-up ventilation
38.6.2 | Volume Controlled Ventilation + SIMV + Pressure Support
38.6.3 | Pressure Control Ventilation + SIMV + Presssure Support
38.6.4 | CPAP
39 | MAINTENANCE CONTRACT
39.1 | Bidder must offer a five year full maintenance contract effective after the two year guarantee expires.
39.2 | The cost of the maintenance contract must not be included in the price of the bid
39.3 | State the cost per year and the total cost
39.4 | Provide details of the contract in a separate document

## Anaesthetic Std 2
| ANAESTHESIA DELIVERY SYSTEM: Standard with integrated physiological monitor | Complies Y/N | Comment
1 | Anaesthesia delivery system suitable for paediatric and adult anaesthesia
2 | Ventilator suitable for ventilating paediatric to adult patients
3 | Multigas analyser:
3.1 | Oxygen
3.2 | Nitrous oxide
3.3 | Carbon dioxide
4 | Anaesthetic agent monitoring:
4.1 | Isofluorane
4.2 | Sevofluorane
4.3 | Desfluorane
4.4 | Halothane
4.5 | Enflurane
5 | Must comply with ISO 13485
6 | Must comply with SANS/IEC 60601-2-13
7 | Self test requiring limited user intervention.
8 | The self test must provide the measurements and values for automatic leak detection and compliance.
9 | Automatic calibration of sensors for gases and anaesthetic vapours must form part of the self test and require very limited user intervention.
10 | The unit must provide conventional, low-flow (down to 200 ml/min) anaesthesia.
11 | It must be possible to override the self test facility. The system must display the override status
12 | The unit must return analysed gas back to the circle system
13 | The unit must be supplied with a minimum of 1 storage drawer
14 | The unit must be supplied with a minimum of 3 additional 220 V power sockets. These sockets must be excluded from the battery backup circuit in case of a power failure
15 | The unit must have a working area or a retractable shelf. Width: not less than 300 mm. Depth: not less than 200 mm
16 | The unit must have a light to enable the anaesthetist to work in a darkened theatre.
17 | Trolley
17.1 | The unit must be mounted on a strong , robust and compact trolley manufactured from corrosion free material.
17.2 | The trolley must have no sharp corners and be easy to clean.
17.3 | The trolley must have at least four antistatic castors of diameter 100 mm or greater.
17.4 | Trolley must have two lockable castors or a central locking brake.
17.5 | The unit must accommodate an oxygen cylinder. Cylinder size according to user requirement. Size to be compatible with current state vendor of oxygen.
17.6 | Trolley must have accessible handles for ease of movement
17.7 | The trolley, with all components mounted, must be stable without toppling over when placed on a 30° slope.
17.8 | The unit must be supplied with a castor guard system, preventing the entaglement and trapping of cables within the wheels.
17.9 | The trolley must have a mounting system suitable for a physiologial monitor.
18 | Pipeline and Cylinder Gas Supplies
18.1 | The anaesthesia system must be designed to use external piped medical gasses (O2, N2O and medical air) and vacuum from medical service outlet sockets.
18.2 | Connection hoses for above must be provided
18.3 | Hoses must be fitted with Heyer type non-interchangeable probes for the existing medical gas outlets as specified by the purchaser.
18.4 | The medical air hose connection or CO2 circle system must include a water trap
18.5 | A safety pressure relief valve of 60 - 80 cm H2O must be provided to protect the system from excessive pressure .
18.6 | The unit must function safely and correctly within the pipeline supply pressure range of: 270 - 600 kPa
18.7 | The system must preferentially utilise the piped oxygen supply over the cylinder supply, if pipeline pressure is > 270 kPa
18.8 | Connection hoses must be:
18.8.1 | Flexible and at least 6 metres long
18.8.2 | Antistatic
18.8.3 | Colour coded as follows:
18.8.3.1 | O2: white
18.8.3.2 | N2O: blue
18.8.3.3 | Vacuum: yellow
18.8.3.4 | Medical air: black
18.8.3.5 | Scavenge low vacuum: orange or yellow
18.8.4 | Where relevant, fitted with outlet probes complying with SANS 1409:2014
18.8.5 | Permanantly bonded to the outlet probe and the machine connector
18.9 | Hoses for gases above atmospheric pressure, must designed for pressures of at least 1 000 kPa.
18.10 | Vacuum hoses may not collapse or kink at 90 kPa below atmospheric pressure, when subjected to 180° bend of inside radius equal to the outside diameter of the hose.
18.11 | All connection hoses must be fed to a central manifold
18.12 | All high pressure gas and vacuum pipe connections (other than hoses) and all pipes connected to pressure regulators and pressure gauges must be stainless steel (18/8 grade) pipes that are oil-free and colour coded as for hoses (see above) over their entire lengths
18.13 | An in-line non-return valve must be fitted to each gas pipeline on the high pressure (inlet) line
18.14 | An in-line non-return valve must be fitted to each gas cylinder yoke
18.15 | A replaceable gas filter, to block particles greater than 65 µm, must be fitted on the upstream side of the non-return valve
18.16 | Each gas pipeline inlet must have a pressure gauge or electronic display of pressure. The pressure range (415 kPa ± 75 kPa) must be clearly indicated in colour
18.17 | Must be fitted with a gas cylinder yoke to accommodate one oxygen gas cylinder with a pin index system
18.18 | Each gas cylinder yoke must have a pressure gauge or electronic display of pressure
18.19 | The pressure gauges or electronic displays of pressure for both pipeline & cylinder supplies, calibrated in kPa for each line, must be labelled and visible from the front of the unit.
18.20 | O2 and N2O regulators must be provided complying with CKS 605:1987.
18.21 | A pressure regulator complying with CKS 605 must be supplied for each gas inlet and must reduce the gas pressure to 415 kPa without freezing up, creeping, chattering or whistling
18.22 | Outlet sockets and probes must comply with SANS 1409:2014.
18.23 | Sockets with hoses must be colour-coded as per SANS 1409:2014:
18.23.1 | Oxygen: White background with black letters
18.23.2 | Nitrous oxide: Blue background with white letters
18.23.3 | Medical air: One half black with white letters, the other half white with black letters
18.23.4 | Vacuum: Yellow background with black letters
18.24 | Continuous display, visible from the front of the machine, of both pipeline and cylinder pressures even during power failure
18.25 | Pipeline and cylinder pressures must be differentially displayed, and clearly indicated
18.26 | Safety pressure relief valves must protect the entire pneumatic system from excessive input pressures from pipeline and cylinder supplies.
18.27 | When a cylinder is detached from the pin-indexed yoke no gas leak must occur at the yoke.
18.28 | Airway pressure must be monitored and displayed
18.29 | Safety relief valve for all gases must be supplied and set at 600kPa
19 | Scavenging
19.1 | The unit must have a scavenging system with a hose at least 5 metres long with the suitable connector as per the institutions scavenging installation.
19.2 | Air-break device to prevent negative pressure on the breathing system
19.3 | If an active sytem is supplied a pressure indicator must be present
19.4 | The scavenging outlet must be a 30mm ISO connector.
20 | Patient Gas Delivery System (Flowmeters/Rotameters)
20.1 | The unit must have cascading flowmeters (mechanical or electronic) that show flow rates for high and low flow anaesthesia in the following ranges:
20.2 | Oxygen low flow: 200 - 1 000 ml/min
20.3 | At higher flow rates the gas must cascade into a second flowmeter (1 – 10 l/min)
20.4 | Nitrous oxide low flow: 200 - 1 000 ml/min
20.5 | At higher flow rates the gas must cascade into a second flowmeter (1 – 10 l/min)
20.6 | Air tube 200ml/min – 10 l/min
20.7 | Flowmeters, if not electronic, must be diameter indexed to prevent incorrect replacement during maintenance.
20.8 | Gas supply must conform to:
20.8.1 | O2 minimum flow range of 200 ml/min to 10 l/min. Accuracy ± 5%
20.8.2 | medical air minimum flow range 0 to 10 l/min. Accuracy ± 5%
20.8.3 | nitrous oxide minimum flow range 0 to 10 l/min. Accuracy ± 5%
20.9 | Flowmeters must be accurate to at least 5%
20.10 | Flowmeter valve control
20.11 | Flowmeter control knobs
20.12 | Flowmeter sequence to be according to standards approved by the South Africal Society of Anaesthesiologists - Oxygen on the right
20.13 | The delivery system must prevent simultaneous delivery of two anaesthetic vapours.
20.14 | A separate common gas outlet to accept a standard 22 mm outside diameter and 15 mm internal diameter.
21 | Hypoxic guard
21.1 | The unit must include a safety mechanism that ensures a minimum concentration of 25% oxygen in any oxygen/nitrous oxide gas mixture across all flow rates.
21.2 | The unit must include a safety mechanism that ensures a minimum concentration of 21% oxygen in any oxygen/air gas mixture across all flow rates.
21.3 | The unit must include a high-flow oxygen flush mechanism with a flow rate of at least 35 l/min
21.4 | Oxygen flush must be non-lockable, recessed and self closing
21.5 | This system must bypass the flowmeters and vaporisers
22 | Breathing System
22.1 | Total circuit volume, including absorber, must not exceed 2,7 litres
22.2 | With APL valve 5 - 75 cm H2O
22.3 | With ventilator to manual switch. This switch when directed to the ventilator must isolate the APL valve from the circle
22.4 | Breathing system must be fully autoclavable to 134°C
22.5 | There must be a mechanism to remove water from the breathing system
22.6 | Breathing system must be latex free
22.7 | Bag-to-ventilator switch automatically begins in the ventilator position
22.8 | Reusable corrugated breathing circuit adequate for both adult and paediatric use to be supplied
22.9 | Bidders to supply the breathing circuit as required by the user
22.10 | Bellows must be visible
23 | Integrated Ventilator
23.1 | Ventilator suitable for ventilating adults and paediatric patients without changing ventilator components (e.g. bellows)
23.2 | The ventilator must be electronically controlled
23.3 | The ventilation technology offered may be turbine, volume reflector, piston or ascending bellows
23.4 | The minimum basic modes to be supplied:
23.4.1 | Manual/Spontaneous Ventilation
23.4.2 | Volume Control Ventilation
23.4.3 | Pressure Control Ventilation
23.5 | Additional modes of ventilation may be offered and quoted for separately
23.6 | The ventilator must be integrated with the workstation and user friendly
23.7 | Frequency: minimum range of 5 – 60 bpm.
23.8 | Inspiratory time: minimum range of 0.2 – 10.0 seconds.
23.9 | Minimum Inspiration/Expiration Ratio: 4:1 – 1:4.
23.10 | Adjustable inspiratory pause/plateau time (volume control ventilation): 0-30% of inspiratory time
23.11 | Peak Inspiratory Pressure: minimum range of: PEEP + 5 to 70 cm H2O.
23.12 | Positive End Expiratory Pressure (PEEP): Integrated, electronically controlled. Minimum range: 0 – 30 cm H2O.
23.13 | Tidal Volume: 20ml – 1500ml per breath. Minimum accuracy: 15%
23.14 | Peak Inspiratory Flow: At least 120 lpm
23.15 | The ventilator must be capable of flow or pressure triggering.
23.16 | Adjustable Inspiratory Rise time/Ramp with a range of 0.1 - 2 sec
23.17 | Compliance compensation/Tidal volume compensation/Compressible volume compensation or equivalent.
23.18 | Patient gas flow compensation must be automatic so that alterations in patient gas flow do not affect delivered tidal volumes
23.19 | In event of power or technical failure, manual ventilation must be possible.
24 | Display:
24.1 | All displays must be in acceptable English
24.2 | Colour display must include numerical and graphic display of parameters measured
24.3 | Graphic display to feature at least 3 waveforms: - pressure, flow and volume.
24.4 | Must also display flow-volume and pressure-volume loops with a reference loop
25 | Alarms : The ventilator must have the following user adjustable alarms
25.1 | Apnoea alarm
25.2 | High airway pressure alarm
25.3 | Continuous pressure alarm
25.4 | High and low expiratory tidal volume alarm
25.5 | Minute volume, high/low alarm
25.6 | Inspired oxygen (Fi02), high/low alarm
25.7 | PEEP, low alarm
25.8 | Leakage / Check breathing circuit alarm or equivalent
25.9 | Alarm profile must be user definable or default to patient appropriate settings
25.10 | The ventilator must have a mute alarm function, active only for a limited time.
25.11 | It must not be possible to cancel the visual alarm until the problem has been rectified.
26 | All components of the ventilator and breathing system in contact with the patient’s airway must be latex free.
27 | All components of the ventilator and breathing system in contact with the patient’s airway must be steam autoclavable up to 134ᴼ C
28 | Integrated Gas Monitoring
28.1 | Oxygen Monitoring
28.1.1 | The system must display both inspired and end-tidal oxygen concentration.
28.1.2 | The oxygen analyser must have user adjustable low and high alarms – visual and audible. The low alarm must automatically default to no less than 21%.
28.1.3 | The system must be accurate (+/- 2% in 21-100%) range, and have a rapid response time (≤ 500ms).
28.1.4 | The unit must sound a continuous alarm if the oxygen pressure falls below 190 kPa, then all nitrous oxide must be automatically cut from the patient circuit.
29 | Patient Physiological Monitoring
29.1 | The patient monitor must be modular and suitable for use for paediatric and adult patients.
29.2 | Must interface with the gas monitoring system if required
29.3 | The following paramaters must be monitored:
29.3.1 | - ECG: 3-lead and 5-lead. Must be capable of displaying 2 channels simultaneously.
29.3.2 | - ST-segment and arrhythmia analysis.
29.3.3 | - NIBP
29.3.4 | - Oxygen saturation (SpO2)
29.3.5 | - Plethysmography/respiratory rate
29.3.6 | - Temperature x 2 (minimum)
29.3.7 | - Invasive pressure x 2 (minimum)
29.4 | The patient monitor must be fully integrated with the workstation.
29.5 | Display must be no less than 400 mm diagonally.
29.6 | Minimum 6 waveforms must be able to be displayed simultaneously.
29.7 | Full colour display.
29.8 | The monitor must be capable of displaying trends for at least 24 hours in tabular and graphical format.
29.9 | The trend resolution must be user-adjustable and have a minimum resolution of 1 min.
29.10 | The front panel must be easy to clean and disinfect.
29.11 | The display must be flicker free and easy to read.
29.12 | Brightness control must be possible.
29.13 | In addition to the waveform display section, the monitor must have alphanumeric display of parameter identification, measured values, alarm and status messages.
29.14 | Audible and visual alarms must be available for all measured parameters.
29.15 | Default alarm limits if the limits are not set by the user
29.16 | Alarm indication: Visual and audible with mute
29.17 | Adjustable audio alarm volume
29.18 | Low battery: Alarm at least 10 minutes before battery power is exhausted
29.19 | It must be possible to mute the audible alarm for a short period, but no override of the visual alarm must be possible.
29.20 | It must be possible to suspend the alarms of certain parameters during clinical procedures, but the alarm must reactivate automatically after a suitable time delay. Any silencing or disabling of alarms must be clearly indicated on the monitor.
29.21 | Input impedance of the system > 10 MΩ.
29.22 | Maximum allowable leakage current to the patient must be less than 10 µA.
29.23 | The monitor must be protected against defibrillator and electrosurgical interference.
29.24 | The monitor must be of low power consumption, and must not require fan or other forced ventilation.
29.25 | All cables and sockets must be colour coded for easy recognition, and must conform to international standards
30 | ECG
30.1 | A full set of 3/5 lead ECG cables must be included in the offer.
30.2 | The ECG amplifier must employ a differential input with high noise suppression.
30.3 | Sweep speed: at least 25 to 50 mm/sec.
30.4 | Freeze facility for ECG.
30.5 | The input impedance must be > 10 MΩ.
30.6 | The heart rate measuring range: 20 to 300 bpm, with an accuracy < 5 bpm.
30.7 | An 50 Hz filter must be provided.
30.8 | Alarm limits must be adjustable in the following range:
30.9 | Low: 40 to 150
30.10 | High: 70 to 250.
31 | Non Invasive Blood Pressure
31.1 | Must be suitable for adults, paeds and neonates
31.2 | The unit must use the oscillometric method of measurement.
31.3 | Modes: Manual and automatic
31.4 | Reusable neonate, paediatric, adult and obese cuffs must be supplied
31.5 | The cuffs must be robust, interchangeable and easy to clean.
31.6 | Must display systolic, diastolic and mean pressures for each measurement
31.7 | Adult: Systolic: 30 - 270 mm Hg, Diastolic: 20 - 240 mm Hg, MAP: 20 - 250 mm Hg
31.8 | Paeds: Systolic: 30 - 160 mm Hg, Diastolic: 20 - 120, MAP: 20 - 120 mm Hg
31.9 | Values must be displayed numerically
31.10 | Auto zero when unit is switched on
31.11 | The interval between measurements must be selectable from 1 – 120 minutes.
31.12 | Adjustable high and low alarm limits for systolic, diastolic, mean arterial pressure and heart rate.
31.13 | Over pressure protection: Automatic deflation at 300 mm Hg adult, 150 mm Hg neonate
31.14 | Average measurement time: approximately 30 s
31.15 | If measurement exceeds two minutes the cuff must deflate, an alarm must sound and likely faults indicated.
31.16 | The air hose must be at least two metres long.
32 | Oxygen Saturation.
32.1 | Continuous monitoring of oxygen saturation levels by means of the photometric principle.
32.2 | Must be suitable for adults, paeds and neonates
32.3 | The unit must continuously display the SaO2 % waveform
32.4 | Oxygen saturation (SaO2 %): Range of measurement: 0 - 100%
32.5 | The accuracy must be better than 2,5 % (70 - 100%) and 5% (50 - 70%).
32.6 | Pulse rate (beats/min): Range of measurement: 20 - 300 bpm
32.7 | The pulse rate accuracy must be better than ± 2 %.
32.8 | An audible bleep should sound for each heartbeat, the pitch of which should vary with the oxygen saturation.
32.9 | Automatic rejection of motion artefacts
33 | SpO2 Sensor:
33.1 | Sensor: Dual LED, Infra-red/visible red, which does not heat the point of application to the patient.
33.2 | The following reusable sensors must be provided:
33.3 | Finger sensor for adults and paeds.
33.4 | Earlobe sensor for adults and paeds.
33.5 | Finger wrap-around probe for neonates
33.6 | The sensor lead must be approximately 3 metres long.
34 | Temperature
34.1 | Temperature range must be approximately 15º to 45º C.
34.2 | Two reusable temperature probes must be included
35 | Circle System
35.1 | The circle absorber system must be integrated as part of the ventilation system.
35.2 | The circle absorber system must have an effective system to eliminate excess water created by the chemical reactions either by heating or water traps.
35.3 | The circle absorber system must be autoclavable at 134ºC.
35.4 | The circle absorber system must be supplied with a carbon dioxide absorber canister.
35.5 | The absorber canister may be reusable/refillable
35.6 | It must be easy to remove and replace the absorber canister while the workstation is in operation.
35.7 | Misalignment of the canister must be readily apparent to the user.
35.8 | The system must have a circle bypass facility while changing/refilling the absorber canister.
35.9 | The breathing circuit must contain an integrated adjustable pressure limiting (APL) valve with a minimum range of 0-70 cm H2O.
35.10 | The APL valve must have a tactile or electronic indicator when a value of between 20-30 cm H2O is exceeded.
35.11 | The circle absorber system must be latex free.
35.12 | The bag-to-ventilator switching mechanism must be clearly labeled and user-friendly. Mechanical or electronic
36 | Patient circuit pressure release valve
36.1 | Adjustable up to 70 cm water
36.2 | Audible alarm when the valve opens
36.3 | Manual override, non-lockable, to keep the release valve closed when the set pressure is exceeded.
37 | Flowmeter pressure release valve
37.1 | Situated between the flowmeters and the vaporisers
37.2 | Must open at 105 kPa +- 5 kPa
38 | General
38.1 | All components must bear the CE mark or equivalent.
38.2 | The following safety standards must apply: ISO 9001/ ISO 46001; Annex 11 of Directive 93/42/EEC, ISO 13485, I.E.C 60601-2-13
38.3 | The unit must operate from both the 220 V, 50 Hz mains supply and internal rechargeable batteries
38.4 | The batteries must be charged while the unit is operating from the mains supply.
38.5 | The batteries must automatically take over in the event of a mains power failure.
38.6 | The batteries must have the capacity for full operation of the anaesthetic machine and ventilator for at least 90 minutes with batteries fully charged, under typical operating conditions.
38.7 | Graphic battery power indicator
38.8 | Recharge time < 8 hrs
38.9 | Mains cable at least three metres long with SABS approved 15A three prong plug.
38.10 | The unit must have a rail for attaching a suction regulator.
39 | Data management:
39.1 | The workstation must be capable of downloading data to a USB device or memory card.
39.2 | The monitor must be able to communicate with a central station
39.3 | A printer must be incorporated in the unit
40 | The unit must be handed over in full operating order
41 | The following accessories must be supplied with the unit:
41.1 | Re-breathing bag, reusable, rubber
41.2 | Reservoir bag
41.3 | Pressure limiting valve, adjustable
41.4 | Oxygen fuel cell analyser
41.5 | Adult & Paediatric volume sensor re-usable
41.6 | Water trap for multigas monitor, re-usable (10 pcs.)
41.7 | A full set paediatric 3/5 lead ECG cables
41.8 | A full set adult 3/5 lead ECG cables
41.9 | Reusable paediatric cuffs
41.10 | Reusable adult cuffs
41.11 | Reusable neonate cuffs
41.12 | Reusable obese cuffs
41.13 | Air hose
41.14 | Finger sensor for paeds
41.15 | Finger sensor for adults
41.16 | Earlobe sensor for paeds
41.17 | Earlobe sensor for adults
41.18 | Finger wrap-around probe for neonates
41.19 | Sensor lead
41.20 | Two reusable temperature probes
41.21 | State the cost and frequency of replacement for the flow sensors.
41.22 | State the cost and frequency of replacement for the fuel cell.
42 | A starter-pack of consumables must be supplied with the unit comprising at least the following:
42.1 | Anaesthesia circuit (Adult, single/double limb design) Disposable. 1.8m (Complete with Y-piece, Elbow, Catheter mount, Rebreathing bag)
42.2 | Anaesthesia circuit (Paediatric, single/double limb design) Disposable. 1.8m (Complete with Y-piece, Elbow, Catheter mount, Rebreathing bag)
42.3 | Anaesthesia circuit (neonatal, single/double limb design) Disposable. 1.8m (Complete with Y-piece, Elbow, Catheter mount, Rebreathing bag)
42.4 | Re-breathing circuit complete
42.5 | Non-re-breathing circuit complete
42.6 | Flow sensor (pack of 2)
42.7 | Spirometry line
42.8 | Sample line (pack of 10)
42.9 | Face masks: sizes 00, 0, 1, 2, 3, 4, 5.
42.10 | Bacterial/viral filter
42.11 | Y connector
42.12 | Manual ventilation hose
42.13 | Fresh gas tubing
42.14 | Seal for back-up cylinder connection on the unit
42.15 | Scavenger tubing
42.16 | Inspiratory and expiratory unidirectional valves
42.17 | Side stream CO2 kit (adult/paed)
42.18 | Disposable Adult Spirometry Set complete with spirometry tube, sample line and flow sensor, 3 m, box of 20 pieces
42.19 | Disposable Paediatric Spirometry Set complete with spirometry tube, sample line and flow sensor, 3 m, box of 20 pieces
42.20 | Gas sample tube with luer lock connection, 3m, disposable (10 pcs)
42.21 | Suction Tubing Set inc connector
42.22 | Resuscitator Adult, 1500mL, autoclavable, silicone
42.23 | Resuscitator Paediatric, 550mL, autoclavable, silicone
42.24 | Tubing set Paediatric for circle systems
43 | List ALL available consumables with prices incl VAT
44 | List ALL available accessories with prices incl VAT
45 | OPTIONAL ITEMS
45.1 | Prices incl VAT must be given for all optional items
45.2 | Brochures with full technical specifications must be supplied for each optional item offered
45.3 | Vaporisers
45.3.1 | Isofluorane
45.3.2 | Sevofluorane
45.3.3 | Desfluorane
45.3.4 | Halothane
45.3.5 | Enflurane
45.3.6 | Vaporisers must be tilt, pressure and temperature compensated
45.3.7 | Mechanical, electronic or injection technology vaporisers are acceptable
45.3.8 | Filling mechanism:
45.3.9 | Vaporiser filling mechanism must be compatible with the current state approved vendor of inhalational volatile vapours.
45.3.10 | Each vaporiser must be supplied with 2 compatible filling adapters as per user requirement
45.3.11 | Minimum accuracy of the supplied volatile vapour delivery system must be at least 20% of set value
45.3.12 | Visible level indicator
45.4 | Suction Regulator
45.4.1 | A suction regulator capable of -50 kPa (minimum)
45.4.2 | The unit must have a rail for attaching the suction regulator.
45.4.3 | The suction regulator must have a built-in diaphragm with a 0 – 100 kPa vacuum gauge and an “on-off” switch.
45.4.4 | It must aspirate 1 litre water in 45 seconds at a pressure of at least - 50 kPa
45.4.5 | A collection jar calibrated to at least 1 500 ml must be supplied.
45.4.6 | The collection jar must accommodate a disposable liner, as per user requirement
45.5 | Anaesthetic vapour, carbon dioxide and nitrous oxide monitoring:
45.5.1 | Anaesthetic vapour, N2O monitoring and capnography must be integrated into the anaesthetic workstation.
45.5.2 | The monitoring must be of the side-stream type.
45.5.3 | Sampling: 100 - 200 ml/min
45.5.4 | Sampled gases must be internally returned to the breathing system.
45.5.5 | CO2 must be presented as a numerical display of inspired and end-tidal concentration and a capnography waveform.The displayed units must be user-selectable: kPa or mmHg.
45.5.6 | The monitoring system must have automatic identification of anaesthetic vapours.
45.5.7 | The system must have a graphical display indicating the type of anaesthetic vapour/s currently in use and the concentration thereof. The colour coding must be as per international colour coding for anaesthetic vapours.
45.5.8 | The anaestheric vapour must be displayed as vol % and must include inspiratory and expiratory concentrations.
45.5.9 | The monitoring system must be able to simultaneously detect, identify and display mixed anaesthetic vapours.
45.5.10 | High and low anaesthetic vapour alarms must be present and these must be user-definable.
45.5.11 | There must be automatic calculation of minimum alveolar concentration.
45.5.12 | The system must be able to calculate anaesthetic gas utilization in ml per case
45.5.13 | The unit must have an indication of the level of vapour in the vaporiser
45.5.14 | Integrated vapour analysis must deliver an audible "low vapour " alarm and a critical alarm for an empty tank
45.6 | Ventilator modes
45.6.1 | Pressure Support with Mandatory Apnoea back-up ventilation
45.6.2 | Volume Controlled Ventilation + SIMV + Pressure Support
45.6.3 | Pressure Control Ventilation + SIMV + Presssure Support
45.6.4 | CPAP
45.7 | MAINTENANCE CONTRACT
45.7.1 | Bidder must offer a five year full maintenance contract effective after the two year guarantee expires.
45.7.2 | The cost of the maintenance contract must not be included in the price of the bid
45.7.3 | State the cost per year and the total cost
45.7.4 | Provide details of the contract in a separate document

## Apnea monitor 1
| APNEA MONITOR (Impedance sensor) | Complies Y/N | Comment
1 | The unit must be a monitoring device, for continuous respiratory surveillance in neonates and infants.
2 | The monitor must trigger both audible and visible alarms after a preselected delay of 10, 15 or 20 sec following the onset of apnea.
3 | The audible alarm must be audible at a distance of at least 10 metres
4 | The monitor must indicate every breathing motion sensed.
5 | Tenderer must describe the working principle of the breathing sensor/transducer.
6 | Digital display for breath rate of the infant. The display must be clearly visible under all lighting conditions
7 | The unit must operate from 220V, 50 Hz single phase supply.
8 | Mains cable at least three metres long with SABS approved 15A three prong plug.
9 | The unit must operate from an internally mounted rechargeable battery in case of mains failure.
10 | Battery charger to be incorporated
11 | Battery to be supplied
12 | Low battery: audible and visual alarms
13 | The sensor/transducer must be safely isolated from the 220V, 50 Hz supply.
14 | The unit must be well sealed from liquids
15 | The unit must be handed over in full operating order
16 | A starter-pack of consumables must be supplied with the unit.
17 | List ALL available consumables with prices incl VAT
18 | List ALL available accessories with prices incl VAT

## Apnea monitor 2
| APNEA MONITOR (Pressure sensor) | Complies Y/N | Comment
1 | The unit must be a monitoring device, for continuous respiratory surveillance in neonates and infants.
2 | The monitor must trigger both audible and visible alarms after a preselected delay of 10, 15 or 20 sec following the onset of apnea.
3 | The audible alarm must be audible at a distance of at least 10 metres
4 | The monitor must indicate every breathing motion sensed.
5 | Tenderer must describe the working principle of the breathing sensor/transducer.
6 | Digital display for breath rate of the infant. The display must be clearly visible under all lighting conditions
7 | The unit must operate from 220V, 50 Hz single phase supply.
8 | Mains cable at least three metres long with SABS approved 15A three prong plug.
9 | The unit must operate from an internally mounted rechargeable battery in case of mains failure.
10 | Battery charger to be incorporated
11 | Battery to be supplied
12 | Low battery: audible and visual alarms
13 | The sensor/transducer must be safely isolated from the 220V, 50 Hz supply.
14 | The unit must be well sealed from liquids
15 | The unit must be handed over in full operating order
16 | A starter-pack of consumables must be supplied with the unit.
17 | List ALL available consumables with prices incl VAT
18 | List ALL available accessories with prices incl VAT

## Apnea monitor 3
| APNEA MONITOR (Airflow/CO2 sensor) | Complies Y/N | Comment
1 | The unit must be a monitoring device, for continuous respiratory surveillance in neonates and infants.
2 | The monitor must trigger both audible and visible alarms after a preselected delay of 10, 15 or 20 sec following the onset of apnea.
3 | The audible alarm must be audible at a distance of at least 10 metres
4 | The monitor must indicate every breathing motion sensed.
5 | Tenderer must describe the working principle of the breathing sensor/transducer.
6 | Digital display for breath rate of the infant. The display must be clearly visible under all lighting conditions
7 | The unit must operate from 220V, 50 Hz single phase supply.
8 | Mains cable at least three metres long with SABS approved 15A three prong plug.
9 | The unit must operate from an internally mounted rechargeable battery in case of mains failure.
10 | Battery charger to be incorporated
11 | Battery to be supplied
12 | Low battery: audible and visual alarms
13 | The sensor/transducer must be safely isolated from the 220V, 50 Hz supply.
14 | The unit must be well sealed from liquids
15 | The unit must be handed over in full operating order
16 | A starter-pack of consumables must be supplied with the unit.
17 | List ALL available consumables with prices incl VAT
18 | List ALL available accessories with prices incl VAT

## Bilirubin meter Xenon
| Billirubin meter Xenon | Complies Y/N | Comment
1 | Non-invasive transcutaneous bilirubin meter to assess the risk of neonatal hyperbilirubinemia
2 | Effective regardless of skin colour
3 | Suitable for neonates of gestational age > 24 weeks
4 | Suitable for neonates of post-natal age 0 - 20 days
5 | The unit must be handheld
6 | The unit must be easy to operate
7 | Measurement by placing the unit on the skin of the newborn and pressing it lightly
8 | Measurements must be done by means of a white light measured at specific wavelengths between 300 and 900 nm
9 | The measuring method must be based on the determination of the yellowness of the subcutaneous tissue by using two optical paths to measure optical density difference at two wavelengths
10 | The light source: Xenon
11 | The unit must take the measurement, process and display the information without need of any peripheral devices.
12 | Must not be sensitive to user-applied pressure or angle of device
13 | Must be able to display result of a single measurement or the average of a number of measurements
14 | Must not be sensitive to motion artifacts (or movements) of the baby
15 | Can be used with or without consumable tips
16 | Must not require routine calibration
17 | The light source life: > 150 000 measurements
18 | The unit must be capable of being used before, during and post phototherapy.
19 | The unit must incorporate a LCD display to display results, operating parameters, errors, date and time.
20 | The unit must report results in mg/dl or µmol/L
21 | The total serum billirubin range must be from 0 to 20 mg/dl or 0 to 340 µmol/L.
22 | Repeatability must be ±1.5mg/dl or 25 µmol/L.
23 | Accuracy during and post phototherapy must be ±1.75mg/dl or 30µmol/L.
24 | The following error signals must be displayed on the unit.
24.1 | Unacceptable calibration
24.2 | Unacceptable patient measurement
24.3 | Too much movement during measurement
24.4 | Measurement is out of clinically proven range
24.5 | Light bulb faulty
24.6 | System error
24.7 | Low battery
25 | The unit must incorporate a rechargeable battery, and must operate only from the battery power.
26 | Battery to be supplied
27 | The battery must allow at least 100 measurements with a full charge.
28 | Expected battery life: at least 1 year
29 | A battery charger, operating from 220 Volt, 50Hz mains power must be supplied
30 | The battery charger must have a battery overcharge protection circuit.
31 | The unit must indicate that the battery is being charged when the unit is plugged into mains power.
32 | Must indicate either the charge or the number of tests left in the battery
33 | A protective carry case must be supplied
34 | An operators manual must be supplied
35 | Physical dimensions: approximately 50mm(W ) X 150mm(H) X 35mm(D)
36 | Weight: approximately 150 grams
37 | The unit must be handed over in full operating order
38 | A starter-pack of consumables must be supplied with the unit.
39 | List ALL available consumables with prices incl VAT
40 | List ALL available accessories with prices incl VAT

## Bilirubin meter LED
| Billirubin meter LED | Complies Y/N | Comment
1 | Non-invasive transcutaneous bilirubin meter to assess the risk of neonatal hyperbilirubinemia
2 | Effective regardless of skin colour
3 | Suitable for neonates of gestational age > 24 weeks
4 | Suitable for neonates of post-natal age 0 - 20 days
5 | The unit must be handheld
6 | The unit must be easy to operate
7 | Measurement by placing the unit on the skin of the newborn and pressing it lightly
8 | Measurements must be done by means of a white light measured at specific wavelengths between 300 and 900 nm
9 | The measuring method must be based on the determination of the yellowness of the subcutaneous tissue by using two optical paths to measure optical density difference at two wavelengths
10 | The light source: LED
11 | The unit must take the measurement, process and display the information without need of any peripheral devices.
12 | Must not be sensitive to user-applied pressure or angle of device
13 | Must be able to display result of a single measurement or the average of a number of measurements
14 | Must not be sensitive to motion artifacts (or movements) of the baby
15 | Can be used with or without consumable tips
16 | Must not require routine calibration
17 | The light source life: > 150 000 measurements
18 | The unit must be capable of being used before, during and post phototherapy.
19 | The unit must incorporate a LCD display to display results, operating parameters, errors, date and time.
20 | The unit must report results in mg/dl or µmol/L
21 | The total serum billirubin range must be from 0 to 20 mg/dl or 0 to 340 µmol/L.
22 | Repeatability must be ±1.5mg/dl or 25 µmol/L.
23 | Accuracy during and post phototherapy must be ±1.75mg/dl or 30µmol/L.
24 | The following error signals must be displayed on the unit.
24.1 | Unacceptable calibration
24.2 | Unacceptable patient measurement
24.3 | Too much movement during measurement
24.4 | Measurement is out of clinically proven range
24.5 | Light bulb faulty
24.6 | System error
24.7 | Low battery
25 | The unit must incorporate a rechargeable battery, and must operate only from the battery power.
26 | Battery to be supplied
27 | The battery must allow at least 100 measurements with a full charge.
28 | Expected battery life: at least 1 year
29 | A battery charger, operating from 220 Volt, 50Hz mains power must be supplied
30 | The battery charger must have a battery overcharge protection circuit.
31 | The unit must indicate that the battery is being charged when the unit is plugged into mains power.
32 | Must indicate either the charge or the number of tests left in the battery
33 | A protective carry case must be supplied
34 | An operators manual must be supplied
35 | Physical dimensions: approximately 50mm(W ) X 150mm(H) X 35mm(D)
36 | Weight: approximately 150 grams
37 | The unit must be handed over in full operating order
38 | A starter-pack of consumables must be supplied with the unit.
39 | List ALL available consumables with prices incl VAT
40 | List ALL available accessories with prices incl VAT

## Cardiotocograph twins
| Cardiotocograph (twins) | Complies Y/N | Comment
1 | Required for external monitoring of fetal heart rate (FHR), and external monitoring of uterine activity of pregnant patients
2 | Twins monitoring is required
3 | Must perform automatic self-check when switched on
4 | There must a visible Fetal Heart Beat indicator which will flash with each detected heart beat
5 | The unit must provide an audio signal, with variable volume, for every heart beat detected
6 | The fetal heart rate must be measured by means of wide beam pulsed Doppler ultrasound, with auto correlation processing.
7 | Must be able to detect and display multiple FHR while discriminating between different foetuses
8 | The oscillator frequency < 1,2 MHz.
9 | Ultrasound intensity must be less than 10 mW/cm2
10 | The transducer must be manufactured of high-impact plastic that is robust and watertight.
11 | The belts and buckles supplied for positioning of the transducer must allow easy angling for optimum positioning.
12 | Adjustable baseline position
13 | A test button must be provided that produces a series of pulses for calibration which must be recorded and displayed.
14 | Display: > 16 cm (corner to corner).
15 | Displays must be clearly visible under all lighting conditions.
16 | Must allow Maternal Heart Rate (MHR) and Fetal Heart Rate(FHR) coincidence detection?
17 | The following parameters must be displayed:
17.1 | Fetal heart rate approximately 50 to 200 beats per minute
17.2 | Uterine activity: approximately 0 to 100 relative units.
18 | Fetal heart rate must update once per second when FHR is greater than 60 and once each beat when FHR is less than 60.
19 | The monitor must have the following controls:
19.1 | Volume Up/Down
19.2 | TOCO/IUP Zero
19.3 | Printer On/Off
19.4 | Printer fast forward
20 | Must have audible and visual alarms for low/high FHR
21 | Electronic storage data capacity.
22 | RECORDER
22.1 | Integrated dual channel thermal recorder
22.2 | Dual heart rate recording for twins must be available.
22.3 | All dot print test
22.4 | The strip chart paper must be easy to load
23 | Fetal heart rate:
23.1 | Range: approximately 50 to 210 beats per minute.
23.2 | Vertical chart: at least 7 cm.
23.3 | Vertical scale sensitivity: approximately 20 bpm per cm.
23.4 | Resolution: 1 bpm
24 | Uterine activity:
24.1 | Range: 0 – 100 units
24.2 | Vertical chart: at least 4 cm.
24.3 | Vertical scale sensitivity: approximately 25 relative units per cm
24.4 | Resolution: 1 relative unit
25 | Paper chart speed must be selectable for 1, 2 and 3 cm per minute.
26 | The time and date must be printed by the recorder at least every 10 minutes.
27 | Recording time must be at least 24 hours at 1 cm per minute.
28 | The unit must have a facility for remote event marking
29 | TRANSDUCERS
29.1 | Ultrasound and toco transducers must be supplied
29.2 | All transducers must be of very robust design. Bidders must guarantee shock resistance on the transducers e.g. will withstand a drop from one metre high onto a hard surface
29.3 | Bidders must state whether the ultrasound transducer is repairable and which parts are replaceable or whether it is a sealed unit that must be replaced when faulty.
29.4 | The ultrasound transducer cable must be a minimum of 2 m long
29.5 | Connecting plug must be keyed and colour coded to avoid it being plugged or forced into the wrong receptacle on the unit.
29.6 | Transducer holders must be provided
30 | General
30.1 | The unit must operate off 220V, 50Hz power supply
30.2 | All metal parts must be earthed.
30.3 | Mains cable at least three metres long with SABS approved 15A three prong plug.
30.4 | Must be fitted with rechargeable batteries
30.5 | The unit must be fitted with an internal battery charger
30.6 | Indicator for battery status
30.7 | Must store trends for at least 8 hours
30.8 | A cart/trolley must be supplied
30.9 | The cart must have four castor wheels, two of them with brakes
30.10 | The cart must have a basket/holder for accessories
30.11 | The unit must be handed over in full operating order
30.12 | A starter-pack of consumables must be supplied with the unit.
30.13 | List ALL available consumables with prices incl VAT
30.14 | List ALL available accessories with prices incl VAT

## Defibrillator
| Defibrillator, with SpO2 and pacing | Complies Y/N | Comment
1 | The defibrillator offered must provide ECG monitoring, a printer and SpO2 and external pacing
2 | The defibrillator must conform to IEC 60601-2-4: Particular Requirements fot the Safety of Cardiac Defibtillators and Cardiac Defibrillator monitors
3 | The defibrillator, monitor, printer, pacing unit, SpO2 unit, battery and battery-charger must be integrated into a single unit.
4 | The defibrillator must deliver a bi-phasic impulse
5 | The defibrillator must have manual and automatic (AED) modes
6 | Automatic mode must measure patient impedance and deliver an appropriate impulse automatically generated, between 0 and 250 J
7 | Audible/voice and on-screen messages for use during AED mode
8 | Manual mode energy selection: 5 - 250 J
9 | Selection of manual energy must be simple with a clear unambiguous indication of the energy selected
10 | The defibrillator must have the following selection of defibrillation:
10.1 | a. Non-synchronous defibrillation
10.2 | b. Synchronous defibrillation
11 | The defibrillator must indicate if the synchronous mode has been selected.
12 | The defibrillator must default to non-synchronous mode at initial switch on.
13 | Synchronous defibrillation must be possible via ECG patient cable and via paddles
14 | The synch marker must be clearly indicated on the ECG trace.
15 | The defibrillator must operate from a 220 V, 50 Hz. supply
16 | The defibrillator must contain a rechargeable battery which charges automatically when the defibrillator is plugged into the 220 V, 50 Hz supply.
17 | The battery must charge whether the defibrillator is switched on not
18 | The defibrillator must have a visual indicator to indicate that the mains supply is switched on.
19 | The internal rechargeable battery must be able to deliver at least 50 defibrillations at the maximum setting and provide a minimum of 2 hours ECG monitoring in the event of a mains power failure
20 | Battery status indicator with audible and visual alarm for low battery
21 | Audible tone while charging
22 | Audible tone, different from above, to indicate full charge
23 | Visual indication of fully charged state
24 | Defibrillating charge time to maximum setting must not exceed 10 seconds.
25 | In the event of a completely depleted battery, the defibrillator must be fully operable on the 220 V, 50 Hz supply.
26 | The defibrillator must do an automatic self-test when switched on.
27 | Override of self test must be possible
28 | Operation of the defibrillator must be user-friendly with clear indications for:
28.1 | Switch on;
28.2 | Select energy;
28.3 | Charge;
28.4 | Discharge.
29 | The defibrillator must be portable and weigh less than 7 kg with all accessories attached
30 | The unit must have a carry handle
31 | The defibrillator must incorporate an internal 50 W test load to allow a check for satisfactory functioning.
32 | The defibrillator must incorporate an emergency dump circuit that will automatically provide a safe current path for unused energy when the paddles are charged but not discharged within sixty seconds.
33 | The defibrillator must be able to do both internal and external defibrillation.
34 | Energy selection must accommodate patients from neonates to adults for both internal and external defibrillation.
35 | The defibrillator must automatically sense when internal paddles are in use and restrict the maximum selectable energy to 50 Joules.
36 | Apex paddle to incorporate a push button to charge the defibrillator.
37 | Energy discharge via defibrillator paddles to patient or test load must be activated by pressing buttons on both apex and sternum paddles simultaneously.
38 | Energy discharge via AED pads or internal electrodes by means of a button on the panel
39 | ECG monitoring must be possible through either the ECG leads or the paddles and unit must display the ECG lead selected (I, II or III)/(aVl, aVR or aVF) or indicate whether the paddles are the source of the ECG
40 | Size of ECG: 0,5; 1; 2 and 4 cm/mV
41 | Sweep rate: 25 mm/s
42 | 50 Hz filter
43 | Heart rate: 30 - 300 bpm
44 | Analysis time: minimum 15 s
45 | Artefact detection
46 | Pacemaker detection
47 | The following alarms must be displayed:
47.1 | low battery,
47.2 | no paper (printer)
47.3 | ECG lead off/fault.
48 | Device must alarm audibly and visually on violation of adjustable setting of LOW and HIGH heart rate
49 | Alarm limit range: adjustable from 30 to 300 bpm (approximately) and 4 seconds of asystole.
50 | The audible alarm should have an mute function to temporarily silence the audible alarm.
51 | Audible alarm-silencing period: 60 s (max)
52 | The screen must be visible under all lighting conditions.
53 | Display size: Approximately 115 x 85 mm.
54 | Display resolution: Approximately 320 x 240 pixels
55 | Display to indicate the following:
55.1 | Alarm messages
55.2 | Heart rate
55.3 | Systole blinker
55.4 | Alarm limits
55.5 | ECG, with freeze capability
56 | A high-resolution, approximately 200 dpi, printer must be incorporated in the unit
57 | Two channel printer for ECG and SpO2 waveform
58 | Printer sensitivity range 1 mV = 5 mm to 1 mV = 25 mm
59 | Muscle artifact filter
60 | Print-out of the following:
60.1 | ECG trace
60.2 | SpO2 wavefrom
60.3 | Date
60.4 | Time
60.5 | Energy selected
60.6 | Energy delivered
61 | Printer paper: 50 mm
62 | Paper speed: 25 to 50 mm/s .
63 | Manual and automatic printer modes.
64 | Automatic mode must print ECG trace before and after alarm/defibrillation conditions
65 | Memory/storage: At least 5 hours of ECG recording with all significant events
66 | External pacemaker
66.1 | Modes:
66.1.1 | Fixed
66.1.2 | Demand
66.2 | 30 - 180 paces/min (Fixed, demand)
66.3 | Impulse current: 10 - 180 mA, in 1 mA steps
66.4 | Impulse shape: Square: 20 ms
66.5 | Refractory period (Demand): 240 ms (HR > 100 bpm), 340 ms (HR < 100 bpm)
66.6 | QRS marker in Demand mode
66.7 | There must be a continuous display of the pacing mode, the selected current and the selected pacing rate.
67 | SpO2
67.1 | Alarm for max and min SpO2
67.2 | Saturation range: 1 - 100%
67.3 | Pulse rate range: 25 - 240 ppm, +- 5 ppm
68 | The unit must be handed over in full operating order
69 | A starter-pack of consumables must be supplied with the unit.
70 | The following accessories must be supplied with the unit
70.1 | Reusable external adult/paediatric defibrillator paddles
70.2 | 3-Lead ECG patient cable (x1)
70.3 | ECG electrodes (x25)
70.4 | AED pads (x2)
70.5 | Cable for AED pads (x1)
70.6 | Pacing electrodes
70.7 | Pacing cable
70.8 | Electrode gel (x1: 250ml)
70.9 | Printer paper (x 10 rolls)
70.10 | SpO2 transducer
70.11 | SpO2 cable
70.12 | User manual
71 | List ALL available consumables with prices incl VAT
72 | List ALL available accessories with prices incl VAT
73 | Optional Items, quote prices incl VAT
73.1 | A trolley must be offered as an optional item.
73.1 | Internal paddles must be offered as optional items

## ECG with trolley
| ELECTROCARDIOGRAPH (12 CHANNEL) with trolley. | Complies Y/N | Comment
1 | Twelve channel portable electrocardiograph
2 | The unit must contain a recorder/printer.
3 | The unit must have a colour TFT LCD display of approximately 115 x 85 mm
4 | High resolution display: at least 800 x 600
5 | Must conform to the following standards:
5.1 | I EN 60601-1:1990 +A1: 1993 + A2:1995 (General requirements for safety)
5.2 | I EN 60601-1-2:2007 (Electromagnetic compatibility)
5.3 | I EN 60601-2-25:1995 + A1:1999 (Particular requirements for safety of ECGs)
5.4 | I EN 60601-2-51:2003 ( Particular requirements for safety)
6 | The unit must do a self-test at power-on
7 | The following must be displayed:
7.1 | Heart rate
7.2 | Patient ID
7.3 | 12 lead waveforms
7.4 | Lead identification
7.5 | ECG gain
7.6 | Paper speed
7.7 | Filters
7.8 | Electrode connection indicator
7.9 | Messages and prompts
7.10 | Clock with date and time
7.11 | Battery power status
8 | R-R analysis for 30 to 300 sec
9 | Pacemaker detection
10 | The display screen must be capable of displaying two lines of text and a minimum of three ECG waveforms.
11 | The unit must be supplied with a 10 lead ECG cable and be able to record 3, 6 and 12 lead configurations.
12 | 12 lead configuration:
12.1 | Lead I, Lead II, Lead III
12.2 | AVR, AVL, AVF
12.3 | V1, V2, V3, V4, V5, V6
13 | Manual and automatic modes
14 | The unit must be capable of ECG interpretation for paediatric to adult.
15 | Full alphanumeric keyboard or touchscreen
16 | The unit must have facility for entering patient data and text
17 | The unit must offer a minimum of six pre-set ECG programs.
18 | Unit must be fitted with a selector switch for QRS sound.
19 | Event markers
20 | Automatic indication of disconnected leads/electrodes
21 | Automatic check for incorrect lead placement
22 | Full screen print preview of ECG waveforms prior to printing
23 | Heart rate recognition: peak-to peak detection. 30 - 300 bpm, accuracy 1 bpm
24 | Electrode connection quality check.
25 | ECG amplifier specs
25.1 | Simultaneous acquisition of 12 leads
25.2 | Frequency response: 0.05 - 150 Hz
25.3 | Sensitivity: Automatic or selectable: 2.5 - 20 mm/mV
25.4 | Automatic amplitude adjustment in the range of 2.5 to 40 mm/mV.
25.5 | Sensitivity threshold: 20 µV
25.6 | Input impedance: > 50 MΩ
25.7 | Patient leakage current: < 10 µA
25.8 | Calibration pulse: 1 mV
25.9 | Filters:
25.9.1 | 50 Hz
25.9.2 | Low pass: 25, 35, 75, 100, 150 Hz
25.10 | CMRR: > 100 dB
25.11 | Sampling frequency: 1 kHz
25.12 | Artifact rejection
25.13 | Baseline stabiliser
26 | Printer
26.1 | Automatic and manual printout mode.
26.2 | In addition to the required ECG waveforms the printer must be able to print:
26.2.1 | device settings
26.2.2 | heart rate
26.2.3 | date and time
26.2.4 | lead identification
26.2.5 | calibration pulse
26.2.6 | patient data
26.3 | The ECG printout must be available on A4 size, Z fold paper.
26.4 | Printer quality: 8 dots/mm (vertical), 20 dots/mm (horiz, 25 mm/s)
26.5 | Paper speed must be selectable as follows: 5, 25 and 50 mm/s.
27 | Internal memory capable of storing at least 200 patient files.
28 | Facility for transmitting this stored ECG to an external device if and when required.
29 | External storage bmo SD card (or other) for 2 000 patient files
30 | Internal protection against damage from 360 J defibrillator discharge.
31 | The unit must have an internal battery with integrated charger
32 | Battery capacity: at least 2 hours of continuous recording
33 | The unit must be fully operable from mains power regardless of the internal battery condition.
34 | A sturdy trolley to house the ECG machine and accessories must be supplied.
34.1 | Trolley must have at least one drawer
34.2 | Trolley must have at least four castors, two with brakes
35 | Mains cable at least three metres long with SABS approved 15A three prong plug .
36 | The unit must be handed over in full operating order
37 | A starter-pack of consumables must be supplied with the unit.
38 | The following accessories must be supplied with the unit
38.1 | ECG cable with adaptor clips
38.2 | Operator's manual
39 | List ALL available consumables with prices incl VAT
40 | List ALL available accessories with prices incl VAT

## Electrosurgery advanced
| Electrosurgery unit, advanced | Complies Y/N | Comment
1 | Must be suitable for the full range of electrosurgical requirements:
1.1 | General surgery
1.2 | Cardiology
1.3 | ENT
1.4 | Gastroenterology
1.5 | Gynaecology
1.6 | Laparoscopy
1.7 | Neurosurgery
1.8 | Thoracic surgery
1.9 | Urology (TURP, TUVP)
2 | User must be able to enter and save preference settings
3 | Power output: at least 400 W with output isolated from ground
4 | Automatic power adjustment for variation in tissue impedance
5 | Automatic recognition of the electrode used and adjustment of the output parameters accordingly
6 | The unit must employ solid state technology for the generation of the various outputs required
7 | The unit must do a self-test at switch-on
8 | The unit must have the following modes of operations:
8.1 | Monopolar mode
8.2 | Bipolar mode
8.3 | Suitable for two operators with active outputs simultaneously
8.4 | Two footswitches active simultaneously
8.5 | One footswitch and one handswitch active simultaneously
8.6 | Two handswitches active simultaneously
8.7 | Monopolar and bipolar active simultaneously
9 | The unit must be able to perform the following functions in monopolar mode:
9.1 | Pure cut
9.2 | Coagulation
9.3 | Blend
9.4 | Spray coagulation
10 | The pure cut function must have variable power from up to 400 W with a 500 Ohm load
11 | The output power selected must be digitally displayed
12 | The coagulation function must have the following output:
12.1 | Variable power from 0 to 120 W delivered into a 500 Ohm load in the coagulation pin point mode
12.2 | Variable power from 0 to 100 W delivered into a 500 Ohm load in the coagulation spray mode
13 | The blend function must have a variable control of the degree of haemostasis.
14 | The Bipolar mode must deliver variable power from 0 to 50 W into a 125 Ohm load in both the Bipolar Cut and Bipolar Coagulation modes.
15 | Frequency: approximately 500 kHz.
16 | Memory for at least 10 programs with their waveforms and power levels
17 | Visual and audible activation indicators
18 | Audible and visual warnings must be given when:
18.1 | Cut function is activated;
18.2 | Coagulation function is activated;
18.3 | Neutral electrode (patient plate) disconnected from the unit
18.4 | The unit has been connected with a faulty neutral electrode or cable;
19 | When an alarm associated within the neutral electrode occurs, the power output of the unit must be reduced to zero.
20 | Reusable neutral electrode to be supplied
21 | The unit must monitor patient contact with the neutral electrode
22 | The unit must fail-safe should the patient make insufficient contact with the neutral electrode.
23 | The output power to the active electrode must be activated by either foot pedal control or a finger-operated switch on the handpiece of the active electrode
24 | There must be no possibility of activating both hand and foot operated accessories simultaneously.
25 | The foot switch must have a barrier between the two pedals
26 | The foot pedal cable must be detachable from both the electrosurgical unit and the foot switch
27 | The foot pedal must be splash proof.
28 | The unit’s power on switch must be enclosed to prevent accidental switch On or Off;
29 | The equipment must function off 220 V, 50 Hz single phase mains supply
30 | Mains cable must be at least 3 metres long and have a SABS approved 15A three pin plug
31 | The unit must comply with IEC 60601-1
32 | The unit must comply to IEC 60601-2-2 Particular Requirements for the Safety of High Frequency Surgical Equipment
33 | The unit must comply with IEC 60601-1-2 Electromagnetic Compatibility
34 | The unit must carry an approved CE or similar mark.
35 | The unit must be supplied with a trolley on antistatic castors, with brakes and a drawer to store footswitch and accessories.
36 | The unit must be handed over in full operating order
37 | A starter-pack of consumables must be supplied with the unit.
38 | The following accessories must be supplied with the unit:
38.1 | Foot switch, double, with cable (1x)
38.2 | Patient electrode, adult, stainless steel or conductive rubber, reusable
38.3 | Patient electrode cable
38.4 | Patient electrode, split, suitable for monitoring patient contact, reusable, sterlizable
38.5 | Cable for split patient electrode
38.6 | Electrode handles with cables, reusable, sterilizable (2x)
38.7 | Blade/knife electrodes, straight, reusable (2x)
38.8 | Needle electrodes (thick) (2x)
38.9 | Needle electrode (thin) (1x)
38.10 | Ball electrodes, straight, 4 mm diameter ((2x)
38.11 | Ball electrode, straight, 2,5 mm diameter (1x)
38.12 | Bipolar electrodes reusable, sterilizable (2x)
39 | Optional Items, quote prices incl VAT
39.1 | Disposable single and split neutral electrode pads
39.2 | Argon electorsurgery
39.3 | Vessel sealing
39.4 | Smoke evacuation
40 | List ALL available consumables with prices incl VAT
41 | List ALL available accessories with prices incl VAT

## Electrosurgery basic
| Electrosurgical unit basic | Complies Y/N | Comment
1 | Electrosurgical cutting and coagulation unit, 100W, mobile
2 | Electrosurgical unit, for general surgery
3 | Microprocessor controlled generator with mono-polar and bipolar output
4 | The unit must employ solid state technology for the generation of the various outputs required
5 | The unit must do a self-test at switch-on
6 | The unit must be able to perform the following functions in monopolar mode:
6.1 | Pure cut
6.2 | Coagulation
6.3 | Blend
6.4 | Spray coagulation
7 | The output power selected must be digitally displayed
8 | Output frequency: approx 400 kHz
9 | Max output power, mono-polar: up to 100 W (cutting),
10 | Bipolar output power: 50 W (coagulation)
11 | Double function foot switch (mono and bipolar)
12 | Hand switch mode when button-activated probes are connected
13 | Visual and audible activation indicators
14 | Audible and visual warnings must be given when:
14.1 | Cut function is activated;
14.2 | Coagulation function is activated;
14.3 | Neutral electrode (patient plate) disconnected from the unit
14.4 | The unit has been connected with a faulty neutral electrode or cable;
15 | When an alarm associated within the neutral electrode occurs, the power output of the unit must be reduced to zero.
16 | Reusable neutral electrode to be supplied
17 | The unit must monitor patient contact with the neutral electrode
18 | The unit must fail-safe should the patient make insufficient contact with the neutral electrode.
19 | Return circuit sensing monitor that deactivates generator in case patient plate or cable fails
20 | Display: output power, system errors and electrode failure
21 | Power requirements: 220 V, 50 Hz
22 | Mounted on a sturdy trolley with 4 antistatic swivel castors, 2 with brakes
23 | Trolley must have a drawer and storage for foot pedal and cable
24 | The unit must comply with IEC 60601-1
25 | The unit must comply to IEC 60601-2-2 Particular Requirements for the Safety of High Frequency Surgical Equipment
26 | The unit must comply with IEC 60601-1-2 Electromagnetic Compatibility
27 | Mains cable at least three metres long with SABS approved 15A three prong plug.
28 | The unit must be handed over in full operating order
29 | A starter-pack of consumables must be supplied with the unit.
30 | The following accessories must be supplied with the unit:
30.1 | Foot switch, double, with cable (1x)
30.2 | Patient electrode, adult, stainless steel or conductive rubber, reusable
30.3 | Patient electrode cable
30.4 | Patient electrode, split, suitable for monitoring patient contact, reusable, sterlizable
30.5 | Cable for split patient electrode
30.6 | Electrode handle, mono-polar, reusable, foot switch controlled, (2x) with connecting cable
30.7 | Electrode handle, mono-polar, reusable, finger switch controlled, (2x) with connecting cable
30.8 | Set different mono-polar reusable electrodes (needle, blade, ball and loop) (1x)
30.9 | Bipolar forceps, reusable, foot switch controlled, with connecting cable (1x)
40 | Optional Items, quote prices incl VAT
40.1 | Disposable single neutral electrode pad
40.2 | Disposable split neutral electrode pad
41 | List ALL available consumables with prices incl VAT
42 | List ALL available accessories with prices incl VAT

## Electrosurgery general
| Electrosurgery unit, general surgery (350 W) | Complies Y/N | Comment
1 | The unit must employ solid state technology for the generation of the various outputs required for general surgery.
2 | The unit must do a self-test at switch-on
3 | The unit must have the following modes of operations:
3.1 | Monopolar mode
3.2 | Bipolar mode
4 | The unit must be able to perform the following functions in monopolar mode:
4.1 | Pure cut
4.2 | Coagulation
4.3 | Blend
4.4 | Spray coagulation
5 | The pure cut function must have variable power from 0 to 350 W when delivered into a 500 Ohm load
6 | The output power selected must be digitally displayed
7 | Power output must remain constant over a wide range of tissue impedance.
8 | Output to be electrically isolated from ground.
9 | The coagulation function must have the following output:
9.1 | Variable power from 0 to 120 W delivered into a 500 Ohm load in the coagulation pin point mode
9.2 | Variable power from 0 to 100 W delivered into a 500 Ohm load in the coagulation spray mode
10 | The blend function must have a variable control of the degree of haemostasis.
11 | The Bipolar mode must deliver variable power from 0 to 50 W into a 125 Ohm load in both the Bipolar Cut and Bipolar Coagulation modes.
12 | Frequency: approximately 500 kHz.
13 | Visual and audible activation indicators
14 | Audible and visual warnings must be given when:
14.1 | Cut function is activated;
14.2 | Coagulation function is activated;
14.3 | Neutral electrode (patient plate) disconnected from the unit
14.4 | The unit has been connected with a faulty neutral electrode or cable;
15 | When an alarm associated within the neutral electrode occurs, the power output of the unit must be reduced to zero.
16 | The unit must monitor patient contact with the neutral electrode
17 | The unit must fail-safe should the patient make insufficient contact with the neutral electrode.
18 | The output power to the active electrode must be activated by either foot pedal control or a finger-operated switch on the handpiece of the active electrode
19 | There must be no possibility of activating both hand and foot operated accessories simultaneously.
20 | The foot switch must have a barrier between the two pedals
21 | The foot pedal cable must be detachable from both the electrosurgical unit and the foot switch
22 | The foot pedal must be splash proof.
23 | The unit’s power on switch must be enclosed to prevent accidental switch On or Off;
24 | The equipment must function off 220 V, 50 Hz single phase mains supply
25 | Mains cable must be at least 3 metres long and have a SABS approved 15A three pin plug
26 | The unit must comply with IEC 60601-1
27 | The unit must comply to IEC 60601-2-2 Particular Requirements for the Safety of High Frequency Surgical Equipment
28 | The unit must comply with IEC 60601-1-2 Electromagnetic Compatibility
29 | The unit must carry an approved CE or similar mark.
30 | The unit must be supplied with a trolley with brakes on antistatic castors with a drawer to store footswitch and accessories.
31 | The unit must be handed over in full operating order
32 | A starter-pack of consumables must be supplied with the unit.
33 | The following accessories must be supplied with the unit:
33.1 | Foot switch, double, with cable (1x)
33.2 | Patient electrode, adult, stainless steel or conductive rubber, reusable
33.3 | Patient electrode cable
33.4 | Patient electrode, split, suitable for monitoring patient contact, reusable, sterlizable
33.5 | Cable for split patient electrode
33.6 | Electrode handle, mono-polar, reusable, foot switch controlled, (2x) with connecting cable
33.7 | Electrode handle, mono-polar, reusable, finger switch controlled, (2x) with connecting cable
33.8 | Set different mono-polar reusable electrodes (needle, blade, ball and loop) (1x)
33.9 | Bipolar forceps, reusable, foot switch controlled, with connecting cable (1x)
33.10 | The bidder must quote on any other accessories required for the unit to be fully operational
34 | Optional Items, quote prices incl VAT
34.1 | Disposable single neutral electrode pad
34.2 | Disposable split neutral electrode pad
34.3 | Vessel sealing
34.4 | Memory for at least 10 programs with their waveforms and power levels
34.5 | Smoke evacuation
35 | List ALL available consumables with prices incl VAT
36 | List ALL available accessories with prices incl VAT

## Exam light 1
| Examination light, LED, ceiling mount | Complies Y/N | Comment
1 | The examination light must be a single head LED with suitable reflector.
2 | The lamp must have sufficient movement to ensure effective lighting over the bed with a minimum length of 1,5 m
3 | The extension must be spring balanced with smooth and easy manoeuvring with one hand
4 | No movement (drift) must take place after positioning the lamp.
5 | Must have friction joints to allow for mobility and positioning.
6 | Must have a removable, autoclavable handle for positioning the light
7 | Light intensity of the lamp must not be less than 30 000 lux measured at 1,0 metre.
8 | The luminaire must be cool to touch, even during multiple examinations.
9 | The light must be separately switched either on the column or on the luminaire.
10 | Colour temp, not less than 4 000K
11 | Colour rendering index: > 90
12 | Lamp life: > 30 000 continuous hours
13 | Spot size (maximum):
13.1 | Working distance of 50 cm = 15 cm
13.2 | Working distance of 100 cm = 25 cm
14 | Must operate from 230V, 50 Hz mains power
15 | Must comply with IEC 60601.1
16 | Smooth easily cleaned surface
17 | Bidder must be responsible for installation/mounting and electrical connection - end user consulation and detailed quotation required per work order
18 | Power supply must be sealed and "hidden"
19 | Bidders to state the replacement cost per LED or range of LEDs.
20 | The unit must be handed over in full operating order
21 | List ALL available consumables with prices incl VAT
22 | List ALL available accessories with prices incl VAT

## Exam light 2
| Examination light, LED, mobile, with battery | Complies Y/N | Comment
1 | The examination light must be a single head LED with suitable reflector.
2 | The lamp must be mounted on a mobile stand
3 | The lamp must have sufficient movement to ensure effective lighting over the bed with a minimum length of 1,5 m
4 | The extension must be spring balanced with smooth and easy manoeuvring with one hand
5 | No movement (drift) must take place after positioning the lamp.
6 | Must have sturdy upright pole at least 1,5 m high
7 | Must have friction joints to allow for mobility and positioning.
8 | Must have a removable, autoclavable handle for positioning the light
9 | Power supply and battery must be hidden in the base of the unit, or neatly installed out of sight.
10 | Light intensity of the lamp must not be less than 30 000 lux measured at 1,0 metre.
11 | The luminaire must be cool to touch, even during multiple examinations.
12 | The light must be separately switched either on the column or on the luminaire.
13 | Colour temp, approximately 4 000K
14 | Colour rendering index: > 90
15 | Lamp life: > 30 000 continuous hours
16 | Spot size (maximum):
16.1 | Working distance of 50 cm = 15 cm
16.2 | Working distance of 100 cm = 25 cm
17 | Base diameter: 35 cm
18 | Must operate from 230V, 50 Hz mains power
19 | Must contain a battery charger
20 | Battery must be capable for at least 2 hours of operation
21 | Mains cable at least three metres, with 15A three pin plug, SABS approved
22 | Mobile stand: lockable castors + hook for mains cable
23 | Durable castors
24 | Must comply with IEC 60601.1
25 | Smooth easily cleaned surface
26 | Bidders to state the replacement cost per LED or range of LEDs.
27 | The unit must be handed over in full operating order
28 | List ALL available consumables with prices incl VAT
29 | List ALL available accessories with prices incl VAT

## Exam light 3
| Examination light, LED, mobile, without battery | Complies Y/N | Comment
1 | The examination light must be a single head LED with suitable reflector.
2 | The lamp must be mounted on a mobile stand.
3 | The lamp must have sufficient movement to ensure effective lighting over the bed with a minimum length of 1,5 m
4 | The extension must be spring balanced with smooth and easy manoeuvring with one hand
5 | No movement (drift) must take place after positioning the lamp.
6 | Must have sturdy upright pole at least 1,5 m high
7 | Must have friction joints to allow for mobility and positioning.
8 | Must have a removable, autoclavable handle for positioning the light
9 | Power supply must be hidden in the base of the unit, or neatly installed out of sight.
10 | Light intensity of the lamp must not be less than 30 000 lux measured at 1,0 metre.
11 | The luminaire must be cool to touch, even during multiple examinations.
12 | The light must be separately switched either on the column or on the luminaire.
13 | Colour temp, approximately 4 000K
14 | Colour rendering index: > 90
15 | Lamp life: > 30 000 continuous hours
16 | Spot size (maximum):
16.1 | Working distance of 50 cm = 15 cm
16.2 | Working distance of 100 cm = 25 cm
17 | Base diameter: 35 cm
18 | Must operate from 230V, 50 Hz mains power
19 | Mains cable at least three metres, with 15A three pin plug, SABS approved
20 | Mobile stand: lockable castors + hook for mains cable
21 | Durable castors
22 | Must comply with IEC 60601.1
23 | Smooth easily cleaned surface
24 | Bidders to state the replacement cost per LED or range of LEDs.
25 | The unit must be handed over in full operating order
26 | List ALL available consumables with prices incl VAT
27 | List ALL available accessories with prices incl VAT

## Exam light 4
| Examination light, LED, wall mount | Complies Y/N | Comment
1 | The examination light must be a single head LED with suitable reflector.
2 | The lamp must have sufficient movement to ensure effective lighting over the bed with a minimum length of 1,5 m
3 | The extension must be spring balanced with smooth and easy manoeuvring with one hand
4 | Must be able to push completely out of the way against the wall for storage.
5 | Restriction to prevent bashing against the wall
6 | No movement (drift) must take place after positioning the lamp.
7 | Must have friction joints to allow for mobility and positioning.
8 | Must have a removable, autoclavable handle for positioning the light
9 | Light intensity of the lamp must not be less than 30 000 lux measured at 1,0 metre.
10 | The luminaire must be cool to touch, even during multiple examinations.
11 | Colour temp, approximately 4 000K
12 | Colour rendering index: > 90
13 | Lamp life: > 30 000 continuous hours
14 | Spot size (maximum):
14.1 | Working distance of 50 cm = 15 cm
14.2 | Working distance of 100 cm = 25 cm
15 | Must operate from 230V, 50 Hz mains power
16 | Power supply to be built in and mounted in/on the wall
17 | Must comply with IEC 60601.1
18 | Smooth easily cleaned surface
19 | Bidder must be responsible for installation/mounting and electrical connection - end user consulation and detailed quotation required per work order
20 | Bidders to state the replacement cost per LED or range of LEDs.
21 | The unit must be handed over in full operating order
22 | List ALL available consumables with prices incl VAT
23 | List ALL available accessories with prices incl VAT

## Fetal Doppler
| Doppler Fetal Heart Rate Monitor | Complies Y/N | Comment
1 | The unit must be a portable Doppler for measuring fetal heart rate.
2 | The unit must be able to accommodate two types of obstetric probes – 2 MHz and 3 MHz
3 | The 2 MHz probe must be supplied with the unit
4 | The probes must be robust and well constructed and the crystals of the transducer must be well protected against accidental droppage.
5 | Fetal heart rate measuring range: approximately 60 - 220 beats per minute.
6 | Resolution: 1 bpm
7 | Accuracy +- 2 bpm
8 | The following must be displayed:
8.1 | Heart rate
8.1 | Heartbeat synchronous indicator
9 | Low battery indicator
10 | Battery charging indicator
11 | The power output of the obstetric probes must not exceed 10 mW/cm².
12 | The unit must operate with rechargable batteries
13 | Battery and charger must be supplied.
14 | Battery charger must operate from a 220 V, 50 Hz power supply
15 | Batteries must have capacity for more than 250 examinations of 1 minute each
16 | Mains cable at least three metres long with SABS approved 15A three prong plug.
17 | The handheld unit must be lightweight, approximately 2 kg
18 | The unit must be equipped with a high quality speaker
19 | The unit must accommodate an external headset. Headset may be offered as an option
20 | If the headset is used for the audio output, the internal loudspeaker of the unit must be disabled.
21 | The volume of the audio output must be adjustable.
22 | The unit must have a clear LED or LCD display showing fetal heart rate
23 | Low battery indication
24 | The unit must automatically shut down after one minute of inactivity
25 | The unit must have a storage holder for the Doppler probes.
26 | The unit must be handed over in full operating order
27 | A starter-pack of consumables must be supplied with the unit.
28 | The following accessories must be supplied with the unit:
28.1 | 2 MHz probe
28.2 | Carry bag
29 | List ALL available consumables with prices incl VAT
30 | List ALL available accessories with prices incl VAT
31 | Optional Items, quote prices incl VAT
31.1 | A 3 MHz probe may be offered as an option
31.2 | Other obstetric probes may be offered as option
31.3 | Headset

## ICU Crib
| ICU Crib | Complies Y/N | Comment
1 | The unit must offer long-term intensive care of a neonate
2 | The unit must conform to (IEC 60601-2-21) “Particular Requirements for the Safety of Infant Radiant Warmers”. The relevant certification from the manufacturer must be supplied
3 | The ICU crib must function off 220 V, 50 Hz single phase power supply.
4 | Platform and mattress:
4.1 | Adjustable tilt to Trendellenburg and reverse Trendellenburg positions
4.2 | Locking mechanism to keep the mattress platform locked in all positions.
4.3 | Must accommodate X-ray film cassettes beneath the mattress.
4.4 | The mattress must be waterproof and must be X-ray transparent.
4.5 | Mattress size: approximately 500 x 650 mm.
4.6 | Mattress thickness: at least 25 mm
4.7 | Mattress must support patient weight up to 4,5 kg.
4.8 | All four sides of the platform must be fitted with transparent guards/panels.
4.9 | The front and two side panels must be hinged.
4.10 | Facility for holding patient cables, tubes, drains etc to prevent then being trapped or dislodged even when sides are opened
5 | Heater
5.1 | Radiant heater housed in an overhead canopy, with internal reflector, mounted on the rear mainframe.
5.2 | Indicator to show heater is switched ON
5.3 | The overhead canopy must swivel left and right and must lock in the desired position.
6 | Controls
6.1 | Microprocessor control with automatic self-test when switched on.
6.2 | The heat controller must provide the following modes for temperature control:
6.3 | Pre-heat
6.4 | Manual mode
6.5 | Skin temperature control
6.6 | Manual heat control: adjustable heater power from 0 - 100%.
6.7 | Skin temperature control: servo-control using a skin temperature probe
6.8 | A reusable skin temperature probe must be supplied
6.9 | Skin temperature probe cable length: 1m
6.10 | The skin temperature set point must be selectable in the range of 34 °C to 38 °C, in 0,1 °C intervals.
6.11 | The skin temperature measurement range: 30 °C to 40 °C.
6.12 | Skin temperature measurement accuracy: 0,5 °C
6.13 | If a malfunction occurs affecting safety or performance the heater must be automatically turned off
7 | Displays:
7.1 | Measured temperature: digital display
7.2 | Heater power
7.3 | Indicator of the selected mode (manual/control)
8 | Alarms
8.1 | Audible and visual alarms when the 220 V power is lost during operation.
8.2 | Audible and visual alarms if the skin temperature deviates by ± 1 °C from the set temperature.
8.3 | Audible and visual alarms if the skin temperature exceeds 39°C
8.4 | Audible and visual alarms if skin temperature probe is disconnected or fails during control mode.
8.5 | Should the skin temperature probe become disconnected or fail, power to the heater circuit must be shut off, the unit must automatically switch to manual mode and the temperature must default to 34°C.
8.6 | Alarm test facility
8.7 | Temporary mute of audible alarms.
8.8 | Apgar timer, with audible alarm, with digital display: 1, 5 and 10 minutes elapsed time.
8.9 | Apgar timer with on/off switch and reset button
9 | Suction
9.1 | Suction regulator adjustable: 0 to 20 kPa
9.2 | Collection bottle (850 ml minimum) with overflow stop valve
9.3 | Adequate tubing must be supplied
9.4 | The suction regulator must be able to operate from the hospital’s piped vacuum system.
9.5 | Suction unit must be mounted onto an accessory rail utilising a snap-lock
9.6 | Collection bottle must accommodate disposable liners
9.7 | Bacterial filter
10 | Other
10.1 | The overhead canopy must be provided with an examination lamp, with on/off switch.
10.2 | The ICU crib must be supplied with an oxygen cylinder holder.
10.3 | The ICU crib must have a drawer
10.4 | The minimum acceptable height of the heater canopy must be 1.85 metres above floor
10.5 | The minimum acceptable height of the mattress platform must 900 mm above floor.
10.6 | The mains cable of the unit must be at least three metres long and be fitted with an SABS approved 15A three pin plug
10.7 | Instrument tray, approximately 200 x 300 mm, 2 kg capacity
10.8 | Swivel mounted monitor shelf, approximately 300 x 400 mm, 10 kg capacity
10.9 | IV pole
10.10 | Accessory rail must be provided on the rear mainframe.
10.11 | The unit must be mobile, but stable during movement
11 | Resuscitator
11.1 | The crib must be equipped with an infant resuscitator integrated into the crib
11.2 | Unit must be gas powered via the wall oxygen supply
11.3 | Oxygen flow regulator: 5 - 15 l/min
11.4 | Adjustable Peak Inspiratory Pressure (PIP): 2 - 75 cm H2O
11.5 | Adjustable Positive End Expiratory Pressure (PEEP): 1 - 25 cm H2O
11.6 | Resuscitator must provide controlled pressure settings by means of a pressure limiting device
11.7 | Pressure gauge: -10 to 80 cm H2O
11.8 | Maximum pressure relief: 5 - 70cmH²O @ 8lpm.
12 | The unit must be handed over in full operating order
13 | A starter-pack of consumables must be supplied with the unit.
14 | The following accessories must be supplied with the unit:
14.1 | Temperature probe
14.2 | Disposable hydrophobic bacterial filters, box of 10
14.3 | IV pole
14.4 | Shelf for monitor
14.5 | Instrument tray
14.6 | Battery
15 | List ALL available consumables with prices incl VAT
16 | List ALL available accessories with prices incl VAT
17 | Optional items, quotes prices incl VAT
17.1 | Air/oxygen blender
17.2 | A variety of patient tubes, face masks, T-pieces and oxygen supply tubes may be offered as options.
17.3 | Built-in scale
17.4 | LED Phototherapy
17.4.1 | The phototherapy light must be either built into the overhead heater canopy or it may be a separate dedicated module that mounts onto the side of the overhead heater canopy.
17.4.2 | The phototherapy light must work off the 220 Volt, 50 Hz single-phase supply
17.4.3 | The phototherapy light must produce UV light with a wavelength between 450 and 470nm
17.4.4 | The phototherapy light must have an intensity of >30uW/cm²/µm
17.4.5 | Must comply with SANS/IEC 60601-2-50:2010
17.5 | SpO2
17.5.1 | Adjustable alarm
17.5.2 | Range: 30 - 100% SpO2
17.5.3 | Pulse rate: 25 - 240 bpm
17.6 | Disposable covers for skin temperature probe (box of 10)

## Incubator ICU
| Incubator ICU | Complies Y/N | Comment
1 | Must comply with SANS/IEC 60601-2-19 (Particular requirements for safety of baby incubator).
2 | The control module must provide an even distribution of heat and humidity within the hood
3 | The unit must perform an automatic self test when switched on.
4 | The incubator frame and chassis must be robust
5 | Chassis must not detach from the frame
6 | Must be fitted with a robust transparent double hood. Approximate dimensions (LxWxH): 900 x 450 x 430 mm
7 | The front of the hood must have a full width transparent access panel, to allow the mattress tray to be extended outside the hood enclosure for examination and treatment of the infant without raising the complete hood.
8 | When the mattress tray is extended outside the hood enclosure, it must remain attached to the incubator and support the weight of the baby.
9 | The hood must be hinged to allow easy access for bedding the infant.
10 | The hood must have a safety latch to prevent accidental opening
11 | The hood must remain in a stable position when opened
12 | The hood must have at least four arm openings with robust latches that are elbow operated. Two of these openings must be on the front and two at the rear of the hood
13 | The hood must also have two iris openings, on the left and right hand sides
14 | The hood must have 2 small self-sealing ports, for IV tubes etc.
15 | The hood must seal well onto the chassis frame to avoid any form of heat loss.
16 | It must be possible to easily manipulate the bed into the Trendellenberg and anti-Trendellenberg positions.
17 | The unit must have an IV pole and a monitor stand.
18 | The monitor stand must support a mass of at least 10 kg
19 | The incubator must be easy to clean, disinfect, and/or sterilize
20 | Switches and controls must be protected against accidental setting changes
21 | TEMPERATURE CONTROL:
21.1 | The control unit must be microprocessor-based.
21.2 | The control module must have servo and manual heat control modes
21.3 | The unit must not take more than ± 30 minutes from power-on to reach the desired temperature setting.
21.4 | User selectable air temperature and skin temperatures.
21.5 | Air temperature must be selectable from ± 25°C to ± 37°C
21.6 | Air temperature must be selectable to 39°C with override confirmation
21.7 | Skin temperature must be selectable from ± 34°C to ± 37°C.
21.8 | Skin temperature must be selectable to 39°C with override confirmation
21.9 | Temperature variation less than +/- 0.2°C
21.10 | A temperature sensor must be available in the hood for temperature control and monitoring.
21.11 | In addition to the primary sensor for controlling the temperature, the unit must have a secondary fail-safe sensor which will act as a safety back-up and activate audible and visible alarms and shut off power to the heater if the temperature exceeds 40oC.
21.12 | The supply to the heater may only be restored after a manual reset and if the incubator’s secondary temperature monitoring device senses that the internal temperature has decreased to a safe level.
21.13 | At switch-on the incubator must automatically heat to 34°C +/- 1°C
21.14 | When the hood’s front access panel is opened, it will have a minimal effect on the set temperature within the incubator.
21.15 | Reusable skin temperature probe must be supplied
22 | HUMIDITY CONTROL:
22.1 | Servo humidity control
22.2 | The humidity must be evenly distributed throughout the hood and must automatically regulate from approximately 40% to 95% relative humidity.
22.3 | The incubator must be provided with an easy to fill water reservoir, which will provide humidity. The water reservoir must be built internally to the unit.
22.4 | The water level in the reservoir must be easy to check
22.5 | Alarm for low water level
22.6 | The water reservoir must be removable and autoclavable
22.6 | Water reservoir volume: approximately 1 000 ml
23 | DISPLAY
23.1 | Display for the following:
23.2 | Skin temperature.
23.3 | Air temperature.
23.4 | Set temperature
23.5 | Relative humidity
23.6 | Details of all alarms
23.7 | Display must be large and clear enough to be viewed from three metres under all lighting conditions.
24 | ALARMS
24.1 | Default alarms must be enabled when the incubator is turned on
24.2 | Audible and visual alarms with heater power cut-off in the event of:
24.2.1 | over temperature
24.2.2 | failure of air circulation
24.3 | Audible and visual alarms in the event of:
24.3.1 | skin temperature differs from set temperature by - 1°C
24.3.2 | air temperature differs from set temperature by - 1°C
24.3.3 | power failure
24.3.4 | heater failure
24.3.5 | skin or air probe disconnection or failure.
24.4 | The audible alarm must have a mute option. Visual alarm may not be reset until the cause of the alarm has been rectified
24.5 | Test facility for audible and visual alarms.
24.6 | Low water alarm
24.7 | The unit must be supplied with an oxygen regulator and dial-in flowmeter
25 | GENERAL
25.1 | Exhaled air must continuously be removed to prevent carbon dioxide build-up inside the hood.
25.2 | Fresh air drawn to the hood must be filtered by an easily replaceable bacterial filter
25.3 | The mattress must be approximately 630 x 350 x 30 mm and must be very durable, comfortable and have a waterproof cover.
25.4 | The sound level within the baby compartment must not exceed 50 dB(A) except if an audible alarm is sounding, then the sound level within the baby compartment must not exceed 80dB(A).
25.5 | The air velocity above the mattress must not exceed 0.35 m/s.
25.6 | The incubator must operate off 220V, 50Hz single phase supply
25.7 | Mains cable at least three metres long with SABS approved 15A three prong plug.
25.8 | All metal parts of the incubator must be well earthed.
25.9 | The incubator’s control module must be well designed to guard against entry of liquids
25.10 | All parts on the incubator and trolley must be rust resistant
25.11 | X-Ray tray incorporated without need to open incubator
25.12 | Rail for accessories on head side
25.13 | Switches and controls must be protected against accidental setting changes
25.14 | Labels and markings must be durable enough to withstand routine cleaning and normal wear
25.15 | Appropriate warning legends must be provided on the unit.
25.16 | The unit must be simple to use and maintain.
25.17 | The unit must be designed for easy access to serviceable parts
25.18 | The incubator must have no sharp edges.
25.19 | The unit must be mounted on a robust trolley
25.20 | Trolley must have 100 mm diameter (minimum) antistatic swivel castors.
25.21 | At least two of the castors must have brakes.
25.22 | Handle (for pushing)
25.23 | Holder for a standard 1 kg oxygen cylinder must be mounted to the trolley.
25.24 | Switches, knobs, and other controls must be designed for conditions of heavy use.
25.25 | Storage space (drawers)
26 | The unit must be handed over in full operating order
27 | A starter-pack of consumables must be supplied with the unit.
28 | The following accessories must be supplied with the unit:
28.1 | Bacterial filter (X 4)
28.2 | Iris port sleeves (one box of 10)
28.3 | Oval port sleeves (one box of 10)
28.4 | Skin temperature probe (reusable)
28.5 | Covers for temperature probe
28.6 | Internal battery for alarms
29 | List ALL available consumables with prices incl VAT
30 | List ALL available accessories with prices incl VAT
31 | Optional Items, quote prices incl VAT
31.1 | Scale
31.2 | Observation lamp.

## Incubator standard
| Incubator (Standard) | Complies Y/N | Comment
1 | Must comply with SANS/IEC 60601-2-19 (Particular requirements for safety of baby incubator).
2 | The control module must provide an even distribution of heat and humidity within the hood
3 | The unit must perform an automatic self test when switched on.
4 | The incubator frame and chassis must be robust
5 | Chassis must not detach from the frame
6 | Must be fitted with a robust transparent double hood. Approximate dimensions (LxWxH): 900 x 450 x 430 mm
7 | The front of the hood must have a full width transparent access panel, to allow the mattress tray to be extended outside the hood enclosure for examination and treatment of the infant without raising the complete hood.
8 | When the mattress tray is extended outside the hood enclosure, it must remain attached to the incubator and support the weight of the baby.
9 | The hood must be hinged to allow easy access for bedding the infant.
10 | The hood must have a safety latch to prevent accidental opening
11 | The hood must remain in a stable position when opened
12 | The hood must have at least four arm openings with robust latches that are elbow operated. Two of these openings must be on the front and two at the rear of the hood
13 | The hood must also have two iris openings, on the left and right hand sides
14 | The hood must have at least 4 small self-sealing ports, for IV tubes etc.
15 | The hood must seal well onto the chassis frame to avoid heat loss.
16 | Easy manipulation of the bed into Trendellenberg and anti-Trendellenberg positions.
17 | The unit must have an IV pole
18 | The incubator must be easy to clean, disinfect, and/or sterilize
19 | Switches and controls must be protected against accidental setting changes
20 | TEMPERATURE CONTROL:
20.1 | The control unit must be microprocessor-based.
20.2 | The control module must have servo and manual heat control modes
20.3 | The unit must not take more than 30 minutes from power-on to reach the desired temperature setting.
20.4 | User selectable air temperature and skin temperatures.
20.5 | Air temperature must be selectable from ± 25°C to ± 37°C
20.6 | Air temperature must be selectable to 39°C with override confirmation
20.7 | Skin temperature must be selectable from ± 34°C to ± 37°C.
20.8 | Skin temperature must be selectable to 39°C with override confirmation
20.9 | Temperature variation less than +/- 0.2°C
20.10 | A temperature sensor must be available in the hood for temperature control and monitoring.
20.11 | In addition to the primary sensor for controlling the temperature, the unit must have a secondary fail-safe sensor which will act as a safety back-up and activate audible and visible alarms and shut off power to the heater if the temperature exceeds 40oC.
20.12 | The supply to the heater may only be restored after a manual reset and if the incubator’s secondary temperature monitoring device senses that the internal temperature has decreased to a safe level.
20.13 | At switch-on the incubator must automatically heat to 34°C +/- 1°C
20.14 | When the hood’s front access panel is opened, it must have a minimal effect on the set temperature within the incubator.
20.15 | Reusable skin temperature probe must be supplied
21 | HUMIDITY :
21.1 | The humidity must be evenly distributed throughout the hood
21.2 | The incubator must have an easy to fill, integrated water reservoir, to provide humidity.
21.3 | Non-spill/splash design of water reservoir
21.4 | The water level in the reservoir must be easy to check
21.5 | Alarm for low humidity
21.6 | The water reservoir must be removable and autoclavable
21.7 | Water reservoir volume: sufficient for approximately 10 hours humidification at 50%
22 | DISPLAY
22.1 | LCD or LED display for the following:
22.2 | Skin temperature.
22.3 | Air temperature.
22.4 | Set temperature
22.5 | Details of all alarms
22.6 | Display must be large and clear enough to be viewed from three metres under all lighting conditions.
23 | ALARMS
23.1 | Default alarms must be enabled when the incubator is turned on
23.2 | Audible and visual alarms with heater power cut-off in the event of:
23.2.1 | Over temperature
23.2.2 | Failure of air circulation
23.3 | Audible and visual alarms in the event of:
23.3.1 | skin temperature differs from set temperature by - 1°C
23.3.2 | air temperature differs from set temperature by - 1°C
23.3.3 | low humidity
23.3.4 | Power failure
23.3.5 | Heater failure
23.3.6 | Skin or air probe disconnection or failure.
23.4 | The audible alarm must have a mute option. Visual alarm may not be reset until the cause of the alarm has been rectified
23.5 | Visual alarm may not be reset until the cause of the alarm has been rectified
23.6 | Test facility for audible and visual alarms.
23.7 | Low water alarm
24 | GENERAL
24.1 | The air velocity above the mattress must not exceed 0.35 m/s.
24.2 | Exhaled air must continuously be removed to prevent carbon dioxide build-up inside the hood.
24.3 | Fresh air drawn to the hood must be filtered by an easily replaceable bacterial filter
24.4 | The mattress must be approximately 630 x 350 x 30 mm and must be very durable, comfortable and have a waterproof cover.
24.5 | The sound level within the baby compartment must not exceed 45 dB(A) except if an audible alarm is sounding, then the sound level within the baby compartment must not exceed 80 dB(A).
24.6 | The incubator must operate from a 220V, 50Hz single phase supply
24.7 | Mains cable at least three metres long with SABS approved 15A three prong plug.
24.8 | All metal parts of the incubator must be well earthed.
24.9 | The incubator’s control module must be well designed to guard against entry of liquids
24.10 | All parts on the incubator and trolley must be rust resistant
24.11 | Labels and markings must be durable enough to withstand routine cleaning and normal wear
24.12 | The unit must be simple to use and maintain.
24.13 | The unit must be mounted on a robust trolley
24.14 | Trolley must have 100 mm diameter (minimum) antistatic swivel castors.
24.15 | At least two of the castors must have brakes.
24.16 | Handles, for pushing, on trolley
24.17 | The unit must be designed for easy access to serviceable parts
24.18 | The incubator must be easy to clean, disinfect, and/or sterilize
24.19 | The incubator must have no sharp edges.
24.20 | Switches, knobs, and other controls must be designed for conditions of heavy use.
24.21 | Storage drawer
25 | The unit must be handed over in full operating order
26 | A starter-pack of consumables must be supplied with the unit.
27 | The following accessories must be supplied with the unit:
27.1 | Bacterial filter (x 4)
27.2 | Skin temperature probe (reusable)
28 | List ALL available consumables with prices incl VAT
29 | List ALL available accessories with prices incl VAT
30 | Optional Items, quote prices incl VAT
30.1 | Blanket cover
30.2 | X-ray tray
30.3 | IV pole
30.4 | Accessory rail
30.5 | Temperature probe covers

## Infant radiant warmer 1
| Infant radiant warmer | Complies Y/N | Comment
1 | The unit must offer long-term care of a neonate
2 | The unit must conform to SANS/IEC 60601-2-21 (Particular Requirements for the Safety of Infant Radiant Warmers). The relevant certification from the manufacturer must be supplied
3 | The infant radiant warmer must function off 220 V, 50 Hz single phase power supply.
4 | Platform and mattress:
4.1 | Adjustable tilt to Trendellenburg and reverse Trendellenburg positions
4.2 | Locking mechanism to keep the mattress platform locked in all positions.
4.3 | The mattress must be waterproof and must be X-ray transparent.
4.4 | Mattress size: approximately 500 x 650 mm.
4.5 | Mattress thickness: at least 25 mm
4.6 | Mattress must support patient weight up to 4,5 kg.
4.7 | All four sides of the platform must be fitted with transparent guards/panels.
4.8 | The front and two side panels must be hinged.
5 | Heater
5.1 | Radiant heater housed in an overhead canopy, with internal reflector, mounted on the rear mainframe.
5.2 | Indicator to show heater is switched ON
5.3 | The overhead canopy must swivel left and right and must lock in the desired position.
6 | Controls
6.1 | Heat control: adjustable heater power from 0 - 100%.
6.2 | If a malfunction occurs affecting safety or performance the heater must be automatically turned off
6.3 | Skin temperature probe to monitor infant temperature.
6.4 | Heat output must be controlled manually or through baby mode (feedback mode) for thermoregulation.
7 | Displays:
7.1 | Heater power
8 | Alarms
8.1 | Visual and audio alarms must be present for safety
8.2 | Apgar timer, with audible alarm, with digital display: 1, 5 and 10 minutes elapsed time.
8.3 | Apgar timer with on/off switch and reset button
9 | Other
9.1 | The overhead canopy must be provided with an examination lamp, with on/off switch.
9.2 | The infant radiant warmer must have a drawer
9.3 | The minimum acceptable height of the heater canopy must be 1.85 metres above floor
9.4 | Adjustable mattress platform height with a minimum of 900 mm
9.5 | The mains cable of the unit must be at least three metres long and be fitted with an SABS approved 15A three pin plug
9.6 | The unit must be mobile, but stable during movement
9.7 | 4 swivel castors of minimum diameter 100 mm
9.8 | Two of the castors must be lockable.
10 | The unit must be handed over in full operating order
11 | A starter-pack of consumables must be supplied with the unit.
12 | The following accessories must be supplied with the unit:
12.1 | Reusable Temperature probe
12.2 | Mattress
13 | List ALL available consumables with prices incl VAT
14 | List ALL available accessories with prices incl VAT
15 | Optional Items, quote prices incl VAT
15.1 | LED Phototherapy
15.1.1 | The phototherapy light must be either built into the overhead heater canopy or it must be a separate dedicated module that mounts onto the side of the overhead heater canopy.
15.1.2 | The phototherapy light must work off the 220 Volt, 50 Hz single-phase supply
15.1.3 | The phototherapy light must produce UV light with a wavelength between 450 and 470nm
15.1.4 | The phototherapy light must have an intensity of >30uW/cm²/µm
15.2 | IV pole
15.3 | Disposable temperature probe
15.4 | Accessory rail
15.5 | Monitor stand
15.6 | Oxygen cylinder holder

## Infant radiant warmer 2
| Infant radiant warmer (Wall-mount) | Complies Y/N | Comment
1 | The unit must offer long-term care of a neonate
2 | The unit must conform to SANS/IEC 60601-2-21 (Particular Requirements for the Safety of Infant Radiant Warmers). The relevant certification from the manufacturer must be supplied
3 | The infant radiant warmer must function off 220 V, 50 Hz single phase power supply.
4 | Platform and mattress:
4.1 | Adjustable tilt to Trendellenburg and reverse Trendellenburg positions
4.2 | Locking mechanism to keep the mattress platform locked in all positions.
4.3 | The mattress must be waterproof and must be X-ray transparent.
4.4 | Mattress size: approximately 500 x 650 mm.
4.5 | Mattress thickness: at least 25 mm
4.6 | Mattress must support patient weight up to 4,5 kg.
4.7 | All four sides of the platform must be fitted with transparent guards/panels.
4.8 | The front and two side panels must be hinged.
5 | Heater
5.1 | Radiant heater housed in an overhead canopy, with internal reflector, mounted on the rear mainframe.
5.2 | Indicator to show heater is switched ON
5.3 | The overhead canopy must swivel left and right and must lock in the desired position.
6 | Controls
6.1 | Heat control: adjustable heater power from 0 - 100%.
6.2 | If a malfunction occurs affecting safety or performance the heater must be automatically turned off
6.3 | Skin temperature probe to monitor infant temperature.
6.4 | Heat output must be controlled manually or through baby mode (feedback mode) for thermoregulation.
7 | Displays:
7.1 | Heater power
8 | Alarms
8.1 | Visual and audio alarms must be present for safety
8.2 | Apgar timer, with audible alarm, with digital display: 1, 5 and 10 minutes elapsed time.
8.3 | Apgar timer with on/off switch and reset button
9 | Other
9.1 | The overhead canopy must be provided with an examination lamp, with on/off switch.
9.2 | The infant radiant warmer must have a drawer
9.3 | The minimum acceptable height of the heater canopy must be 1.85 metres above floor
9.4 | Adjustable mattress platform height with a minimum of 900 mm
9.5 | The mains cable of the unit must be at least three metres long and be fitted with an SABS approved 15A three pin plug
9.6 | Supplier to be responsible for installation
10 | The unit must be handed over in full operating order
11 | A starter-pack of consumables must be supplied with the unit.
12 | The following accessories must be supplied with the unit:
12.1 | Reusable Temperature probe
12.2 | Mattress
13 | List ALL available consumables with prices incl VAT
14 | List ALL available accessories with prices incl VAT
15 | Optional Items, quote prices incl VAT
15.1 | LED Phototherapy
15.1.1 | The phototherapy light must be either built into the overhead heater canopy or it must be a separate dedicated module that mounts onto the side of the overhead heater canopy.
15.1.2 | The phototherapy light must work off the 220 Volt, 50 Hz single-phase supply
15.1.3 | The phototherapy light must produce UV light with a wavelength between 450 and 470nm
15.1.4 | The phototherapy light must have an intensity of >30uW/cm²/µm
15.2 | IV pole
15.3 | Disposable temperature probe
15.4 | Accessory rail
15.5 | Monitor stand
15.6 | Oxygen cylinder holder

## Infusion pump
| INFUSION PUMP. | Complies Y/N | Comment
1 | The unit must use a linear peristaltic drive mechanism. Other drive mechanisms may be offered details to be provided.
2 | The unit must provide user selectable infusion rate from 1 to 999 ml/h, selectable in 1 ml/h increments.
3 | The unit must provide user selectable infusion volume from 1 to 9 999 ml, selectable in 1 ml increments.
4 | The unit must deliver the selected volume with an accuracy better than 5% throughout the whole range of operation/infusion.
5 | Bolus rate, user selectable: 10 - 999 ml/hr
6 | Bolus volume, user selectable: 1 - 1000 ml
7 | The unit must also be capable of carrying out blood administration.
8 | It must not be possible to change the rate while the infusion is in progress.
9 | LCD display
10 | The following must be clearly displayed on the front panel under all lighting conditions:
10.1 | Pump running
10.2 | AC mains supply operation
10.3 | AC power failure/battery operation
10.4 | Infusion rate selected
10.5 | Volume selected
10.6 | Volume infused
10.7 | Alarm condition and alarm/error messages
11 | When an infusion is completed an audible tone must sound.
12 | On completion of an infusion there must be a Keep Vein Open (KVO) rate of 1 - 5 ml/h.
13 | The unit must provide the following alarms:
13.1 | Air in line detection
13.2 | Occlusion during infusion (downstream)
13.3 | Occlusion pressure alarm user selectable: 200 - 500 mmHg
13.4 | Low battery (alarm must sound when there is still approximately 10 minutes battery operation available).
13.5 | Open door
13.6 | Infusion set removed
13.7 | Wrong infusion set detection
13.8 | Completion of selected volume to be infused
13.9 | Zero infusion rate
13.10 | Zero volume selected for infusion
13.11 | Internal malfunction
14 | All alarm conditions must be accompanied by an audible tone.
15 | Free flow protection
16 | State clearly if a dedicated infusion set is required to achieve volumetric accuracy.
17 | The unit must have a memory, which stores alarm messages that can be recalled by service technicians when carrying out preventative maintenance, repairs or servicing.
18 | The unit must have a service mode, to be accessed by service technicians to check important parameters of the unit, without having to first strip the unit for checking/ testing.
19 | The casing of the unit must be impact resistant.
20 | The unit must be provided with a mechanism for attaching to either a mobile drip stand or a gabler rail – NB whichever is requested by the hospital/institution.
21 | The 220 V, 50 Hz AC supply to the unit must be fused in the live or both live and neutral.
22 | Mains cable at least three metres long with SABS approved 15A three prong plug.
23 | The unit must be fitted with an internal rechargeable battery.
24 | In the event of a 220 V mains failure the battery must automatically provide continuity of operation.
25 | The battery must have capacity to operate continuously for a minimum of 3 hours at an infusion rate of 125 ml/h.
26 | The battery charger for the internal rechargeable battery must be an integral part of the infusion pump.
27 | The infusion pump must be provided with circuitry to ensure that the internal rechargeable battery will be protected against over-charging and over-discharging.
28 | The infusion pump must be fully operable from the mains supply regardless of the condition of the battery
29 | The battery must be automatically charged when the unit is connected to the mains supply.
30 | Purge rate (1 - 999 ml/h) Purge volume (1 - 99 ml) (Sci Gr)
31 | Unit can be mounted on standard bed/wall rail or mobile pole/stand
32 | Suitable for all intravenous infusions of fluids
33 | Display/alarm for low/high flow, occlusion, open door, end-of-infusion and built-in battery status
34 | Audio visual alarm with silencing feature
35 | Stand-by mode when not in use
36 | The unit must be handed over in full operating order
37 | A starter-pack of consumables must be supplied with the unit.
38 | The following accessories must be supplied with the unit:
38.1 | Start-up set of 10 giving sets
38.2 | Mounting bracket for fixation to standard bed/wall rail or mobile pole/stand
39 | List ALL available consumables with prices incl VAT
40 | List ALL available accessories with prices incl VAT
41 | Optional Items, quote prices incl VAT
41.1 | Dosing error software
41.2 | Capablity for labeling drug names for display on the pump
41.3 | Docking/stackable station

## Monitor basic
| MONITOR, BASIC MULTI-PARAMETER | Complies Y/N | Comment
1 | Preconfigured, state of the art technology, suitable for neonates to adult
2 | Microprocessor controlled with automatic self-test and calibration
3 | The following parameters must be measured and displayed :
3.1 | ECG waveform
3.2 | Oxygen saturation, waveform + digital
3.3 | Heart rate, digital
3.4 | Respiration, waveform + digital
3.5 | NIBP, digital
3.6 | Temperature, digital
4 | The monitor must operate off both 220 V, 50 Hz single phase AC power supply and an internal rechargeable battery.
5 | If mains power fails the monitor must operate off battery power for at least 2 hours.
6 | Visible indication to show whether the monitor is operating off the mains or battery power.
7 | All control keys must be splashproof and resistant to the entry of fluids into the unit.
8 | The front panel must be easy to clean and disinfect.
9 | The monitor must have a carry handle.
10 | High resolution LCD colour display, 20 - 30 cm diagonal
10.1 | Digits: large: 10 mm, small: 5 mm
10.2 | Sweep speeds : 25 and 50 mm/s
10.3 | It must be possible to freeze the display. This must not affect the QRS signal.
11 | Trending for 24 hours must be possible on the monitor
12 | Defibrillator-proof up to 360 Joules
13 | Protected against HF interference
14 | Patient isolated from all potentially harmful circuits
15 | Must comply with IEC 60601-1
16 | ECG
16.1 | Must comply with IEC 60601-2-27
16.2 | 3 lead ECG with arrhythmia detection
16.3 | Adult ECG lead must be supplied
16.4 | The amplitude of the ECG trace must be adjustable from 1 mV = 5 mm to 1 mV = 20 mm.
16.5 | Digital display of heart rate : 30 - 300 bpm
16.6 | Accuracy : 3 bpm
16.7 | Resolution : 1 bpm
16.8 | Differential input with CMRR > 80 dB
16.9 | Input impedance > 10 MOhm
16.10 | Respiration monitoring through ECG electrodes
16.11 | Pacemaker detection/rejection : 2 to 500 mV; 0,1 to 2 ms
16.12 | Synch pulse for defibrillator
17 | OXYGEN SATURATION
17.1 | Must comply with ISO 80601-2-61
17.2 | The unit should indicate whether the signal from the pulse oximeter sensor is adequate for reliable measurements. If the signal strength or quality is inadequate, this must be clearly displayed.
17.3 | An audible bleep must sound for each heart beat. The pitch must vary with the oxygen saturation.
17.4 | Measurement range : 0 - 100%
17.5 | Accuracy:
17.5.1 | 80 - 100% : 1,5%
17.5.2 | 60 - 80% : 3,5%
17.6 | Continuous waveform display
17.7 | Digital display of SpO2%
17.8 | SpO2 finger sensor to be supplied with the monitor
17.9 | The sensor lead must be at least 3 metres long.
17.10 | Automatic calibration for skin colour
17.11 | Pulse rate: 25 - 250 bpm
18 | NON-INVASIVE BLOOD PRESSURE
18.1 | Must comply with IEC 60601-2-30
18.2 | Oscillometric measurement method
18.3 | Automatic, manual and stat modes
18.4 | Measure and display :
18.4.1 | - systolic pressure
18.4.2 | - diastolic pressure
18.4.3 | - mean pressure
18.5 | The interval between measurements must be selectable from 1 to 60 minutes.
18.6 | Ranges (approximate) :
18.6.1 | Systolic : 30 - 270 mmHg. Resolution : 1 mmHg
18.6.2 | Diastolic : 10 - 150 mmHg. Resolution : 1 mmHg
18.6.3 | Mean : 20 - 250 mmHg. Resolution : 1 mmHg
18.7 | Cuff pressure : 0 - 280 mmHg
18.8 | Cuff inflation time < 10 s
18.9 | Automatic deflation if measurement exceeds 3 min or if mains power fails
18.10 | Adult reusable cuffs must be supplied with the monitor.
19 | TEMPERATURE
19.1 | Range : 10 - 45 ˚C
19.2 | Digital display
19.3 | Surface probe to be supplied
19.4 | Accuracy: 0.3 ˚C
19.5 | Resolution : 0.1 ˚C
20 | ALARMS
20.1 | Visible and audio alarms
20.2 | The monitor must automatically set default alarms when switched on.
20.3 | The alarms must be programmable
20.4 | Alarm limits must be displayed with the parameter digit
20.5 | Heart rate alarm : 30 - 300 bpm
20.6 | Asystole alarm
20.7 | Arrhythmia alarms
20.8 | Apnea alarm
20.9 | Alarms for faulty/not-connected electrode/probe
20.10 | User selectable alarm limits for systolic, diastolic and mean pressures
10.11 | Alarm for incorrect NIBPcuff connection
10.12 | Adjustable SpO2 high and low alarms:
20.12.1 | High : 70 - 100%
20.12.2 | Low : 50 - 95%
20.13 | Adjustable high and low temperature alarms
20.14 | Violation of limits must result in both audible and visible alarms and indicate which parameter is affected.
20.15 | The sound level of the audible alarm must be adjustable.
20.16 | It must be possible to temporarily silence the audible alarm but not remove the visible alarm.
20.17 | Duration of the silenced period: 2 minutes.
20.18 | Low battery alarm
21 | GENERAL
21.1 | Must operate from 220V 50 Hz mains power
21.2 | Mains cable at least three metres long with SABS approved 15A three prong plug.
21.3 | All cables and sockets must be colour coded, labelled or otherwise prevented from being incorrectly coupled
22 | The unit must be handed over in full operating order
23 | A starter-pack of consumables must be supplied with the unit.
24 | The following accessories must be supplied with the unit:
24.1 | 3-lead ECG cable
24.2 | Temperature probe with cable
24.3 | SpO2 finger probe with cable
24.4 | NIBP cuff adult
24.5 | NIBP hose
25 | List ALL available consumables with prices incl VAT
26 | List ALL available accessories with prices incl VAT
27 | Optional Items, quote prices incl VAT
27.1 | Wall bracket, height adjustable on which the monitor can be mounted.
27.2 | Roll stand

## Monitor ICU
| MONITOR, ICU | Complies Y/N | Comment
1 | The monitor must be compatible with the Central Station offered in item xxx.
2 | Modular, state of the art technology, suitable for neonates to adult
3 | Microprocessor controlled with automatic self-test and calibration
4 | The following parameters must be measured and displayed :
4.1 | ECG waveform
4.2 | Oxygen saturation, waveform + digital
4.3 | Heart rate, digital
4.4 | Respiration, waveform + digital
4.5 | Invasive pressure, waveform + digital (systolic, diastolic, mean)
4.6 | NIBP, digital (systolic, diastolic, mean)
4.7 | Temperature, digital
5 | The monitor must operate off both 220 V, 50 Hz single phase AC power supply and an internal rechargeable battery.
6 | If mains power fails the monitor must operate off battery power for at least 2 hours.
7 | Visible indication to show whether the monitor is operating from the mains or battery power.
8 | All control keys must be splashproof and resistant to the entry of fluids into the unit.
9 | The front panel must be easy to clean and disinfect.
10 | High resolution LCD colour display, 35 cm diagonal minimum
11 | At least 6 waveforms must be displayed
11.1 | Digits: large: 10 mm, small: 5 mm
11.2 | Sweep speeds : 25 and 50 mm/s
11.3 | It must be possible to freeze the display. This must not affect the QRS signal.
12 | Trending for 72 hours must be possible on the monitor,
13 | Unit to provide at least 72 hours of full disclosure wave data for review
14 | Defibrillator-proof up to 360 Joules
15 | Protected against HF interference
16 | Patient isolated from all potentially harmful circuits
17 | Must comply with IEC 60601-1
18 | ECG
18.1 | Must comply with IEC 60601-2-27
18.2 | ECG 12 lead with arrhythmia detection
18.3 | Adult ECG lead must be supplied
18.4 | The amplitude of the ECG trace must be adjustable from 1 mV = 5 mm to 1 mV = 20 mm.
18.5 | Digital display of heart rate : 30 - 300 bpm
18.6 | Accuracy : 3 bpm
18.7 | Resolution : 1 bpm
18.8 | Differential input with CMRR > 80 dB
18.9 | Input impedance > 10 MOhm
18.10 | Respiration monitoring through ECG electrodes
18.11 | Pacemaker detection/rejection : 2 to 500 mV; 0,1 to 2 ms
18.12 | Synch pulse for defibrillator
19 | OXYGEN SATURATION
19.1 | Must comply with ISO 80601-2-61
19.2 | The unit should indicate whether the signal from the pulse oximeter sensor is adequate for reliable measurements. If the signal strength or quality is inadequate, this must be clearly displayed.
19.3 | An audible bleep must sound for each heart beat. The pitch must vary with the oxygen saturation.
19.4 | Measurement range : 0 - 100%
19.5 | Accuracy:
19.5.1 | 80 - 100% : 1,5%
19.5.2 | 60 - 80% : 3,5%
19.6 | Continuous waveform display
19.7 | Digital display of SpO2%
19.8 | The sensor lead must be at least 3 metres long.
19.9 | Automatic calibration for skin colour
19.10 | Pulse rate: 25 - 250 bpm
20 | NON-INVASIVE BLOOD PRESSURE
20.1 | Must comply with IEC 60601-2-30
20.2 | Oscillometric measurement method
20.3 | Automatic, manual and stat modes
20.4 | Measure and display :
20.4.1 | - systolic pressure
20.4.2 | - diastolic pressure
20.4.3 | - mean pressure
20.5 | The interval between measurements must be selectable from 1 to 60 minutes.
20.6 | Ranges (approximate) :
20.7 | Systolic : 30 - 270 mmHg. Resolution : 1 mmHg
20.8 | Diastolic : 10 - 150 mmHg. Resolution : 1 mmHg
20.9 | Mean : 20 - 250 mmHg. Resolution : 1 mmHg
20.10 | Cuff pressure : 0 - 280 mmHg
20.11 | Cuff inflation time < 10 s
20.12 | Automatic deflation if measurement exceeds 3 min or if mains power fails
21 | INVASIVE BLOOD PRESSURE
21.1 | Must comply with IEC 60601-2-34
21.2 | Measure and display :
21.2.2 | - systolic pressure
21.2.2 | - diastolic pressure
21.2.3 | - mean pressure
21.3 | Systolic : 0 - 300 mmHg. Resolution : 1 mmHg
21.4 | Diastolic : 0 - 250 mmHg. Resolution : 1 mmHg
21.5 | Mean : 0 - 250 mmHg. Resolution : 1 mmHg
22 | TEMPERATURE
22.1 | Range : 10 - 45 ˚C
22.2 | Digital display
22.3 | Surface probe to be supplied
22.4 | Accuracy: 0.3 ˚C
22.5 | Resolution : 0.1 ˚C
23 | ALARMS
32.1 | Visible and audio alarms
23.2 | The monitor must automatically set default alarms when switched on.
23.3 | The alarms must be programmable
23.4 | Alarm limits must be displayed with the parameter digit
23.5 | Heart rate alarm : 30 - 300 bpm
23.6 | Asystole alarm
23.7 | Arrhythmia alarms
23.8 | Apnea alarm
23.9 | Alarms for faulty/not-connected electrode/probe
23.10 | User selectable alarm limits for systolic, diastolic and mean pressures
23.11 | Alarm for incorrect NIBPcuff connection
23.12 | Adjustable SpO2 high and low alarms:
23.12.1 | High : 70 - 100%
23.12.2 | Low : 50 - 95%
23.13 | Adjustable high and low temperature alarms
23.14 | Violation of limits must result in both audible and visible alarms and indicate which parameter is affected.
23.15 | The sound level of the audible alarm must be adjustable.
23.16 | It must be possible to temporarily silence the audible alarm but not remove the visible alarm.
23.17 | Duration of the silenced period: 2 minutes.
23.18 | Low battery alarm
24 | GENERAL
24.1 | Mains cable at least three metres long with SABS approved 15A three prong plug.
24.2 | All cables and sockets must be colour coded, labelled or otherwise prevented from being incorrectly coupled
25 | The unit must be handed over in full operating order
26 | A starter-pack of consumables must be supplied with the unit.
27 | The following accessories must be supplied with the unit:
27.1 | Adult NIBP cuff with hose
27.2 | Adult SpO2 finger probe with cable
27.3 | Adult reusable temperature sensor
27.4 | ECG cable, 12 lead
28 | List ALL available consumables with prices incl VAT
29 | List ALL available accessories with prices incl VAT
30 | Optional Items, quote prices incl VAT
30.1 | Wall bracket, height adjustable on which the monitor can be mounted.
30.2 | EEG
30.3 | Intra-cranial pressure
30.4 | Cardiac output
30.5 | CO2/EtCO2
30.6 | Printer
30.7 | Apnea monitor for neonatal use
31 | Any additional features may be listed below, state whether they are standard or optional
32 | Technical specifications of any modules not listed in the above specifications must be supplied

## Central station
| Central Station | Complies Y/N | Comment
1 | The central station must have a dual display.
2 | The unit must be compatible with the ICU monitor offered in item xxx.
3 | The unit must be network configurable and wireless.
4 | Must have a viewing screen as well as a review screen for measurements and both screens sizes must be at least 480 mm colour LCD
5 | The second screen must be able to do patient management and reviews without interruption of the live patient view on the first screen
6 | Touchscreen
7 | The unit must be able to remotely manage patient bedside monitors, including viewing active or historic data.
8 | Password protection must be available to protect the critical menus and certain settings.
9 | All waveforms and alphanumeric displays must be available at the central station depending on the parameters of the bedside monitor
10 | Bi-directional comunication between monitors and central station
11 | The central station must be able to download all the patient vitals after the patient has been disconnected and reconnected
12 | Display at least 20 patients
13 | Full arrhythmia and ST segment analysis with 12 lead ECG interpretation
14 | Screen configurations must be able to be saved as favourites for easy access
15 | Up to to 2,000 arrhythmia events per patient monitoring session must be available
16 | Must be able to review full disclosure on the central station for the duration of the disclosure time.
17 | The central station must be able to do the following: Ambulatory ECG review function must be available so the clinician can request full disclosure data from the central station for retrospective analyses
18 | The central station must be able to be the single point of data collection for bedside monitors and a telemetry system at the same time
19 | The central station must allow access to patient demographic information; information arrhythmias, ST-limit violations, and ECG samples in event review; parameter data in graphic and numeric trends; full disclosure data in strip view, page view with zoom, or in calipers.
20 | Printer
20.1 | Paper size must be standard A4 off the shelf paper
20.2 | With the capability of full disclosure waveform printing, multi-wave printing, tabular trend printing and trend graph printing.
20.3 | Printing must be possible at both bedside monitor and central station
21 | Data Storage and Review
21.1 | Trends of at least 120 hours must be available
21.2 | Trends must be displayed graphically and in tabular format
21.3 | Full disclosure of at least 144 hours to be offered
21.4 | Waveform storage must be available for at least 144 hours post discharge
21.5 | Events must be able to be sorted by time and type
21.6 | All patient events, graphical and numerical data and alarms must be available for viewing on a first in first out basis.
22 | Alarms
22.1 | All alarms must be able to be muted using password function at the central station
22.2 | 3-level alarm must be available audible & visual.
22.3 | The sound levels and colour to differentiate alarm type
22.4 | High and low limits of alarm must be offered
22.5 | Monitor alarm limits must be adjustable from the central station. This setting must be enabled/disabled via a password protected menu.
23 | General
23.1 | Network installation and cabling will be done by the bidder and included in the tender price.
23.2 | Height adjustable brackets. Monitor to be adjustable to a lower level than the bottom of the wall mounted baseplate to accommodate shorter staff members.
23.3 | The bidder shall provide fixtures and wall fittings for the entire system including network cable.
23.4 | Bidder must be responsible for installation/mounting and electrical connection - end user consulation and detailed quotation required per work order
24 | The unit must be handed over in full operating order
25 | A starter-pack of consumables must be supplied with the unit.
26 | List ALL available consumables with prices incl VAT
27 | List ALL available accessories with prices incl VAT
Optional Items
Option for monitoring 30 patients
Option for monitoring 40 patients
The unit to be supplied with the UPS (uninterrupted power supply)
Keyboard
Mouse

## NIBP and SpO2 1
| NIBP + SpO2 monitor, desktop | Complies Y/N | Comment
1 | The unit must measure both NIBP and SpO2
2 | Self test procedure at switch-on
3 | The unit must have trending for at least 24 hours
4 | NIBP
4.1 | The unit must measure blood pressure non-invasively.
4.2 | Must be suitable for adults, paeds and neonates
4.3 | The unit must use the oscillometric method of measurement.
4.4 | Modes: Manual, automatic and STAT
4.5 | Pressure measurements:
4.6 | Adult: Systolic: 30 - 270 mm Hg, Diastolic: 20 - 240 mm Hg, MAP: 20 - 250 mm Hg
4.7 | Paeds: Systolic: 30 - 160 mm Hg, Diastolic: 20 - 120, MAP: 20 - 120 mm Hg
4.8 | Neonate: Systolic: 30 - 130 mm Hg, Diastolic: 20 - 100 mm Hg, MAP: 20 - 120 mm Hg
4.9 | Auto zero when unit is switched on
4.10 | The interval between measurements must be selectable from 1 – 120 minutes.
4.11 | Over pressure protection: Automatic deflation at 300 mm Hg adult, 150 mm Hg neonate
4.12 | Average measurement time: approximately 30 s
4.13 | If measurement exceeds two minutes the cuff must deflate, an alarm must sound and likely faults indicated.
4.14 | The air hose must be at least two metres long.
5 | PULSE OXIMETER
5.1 | Suitable for adult, paediatric and neonatal use.
5.2 | The unit must continuously measure and display SpO2 %
5.3 | Oxygen saturation (SpO2 %): Range of measurement: 0 - 100%
5.4 | The accuracy must be better than 2,5 % (70 - 100%) and 5% (50 - 70%).
5.5 | Pulse rate (beats/min): Range of measurement: 20 - 300 bpm
5.6 | The pulse rate accuracy must be better than ± 2 %.
5.7 | An audible bleep should sound for each heartbeat, the pitch of which should vary with the oxygen saturation.
5.8 | Automatic rejection of motion artefacts
6 | SpO2 SENSOR:
6.1 | Sensor: Dual LED, Infra-red/visible red, which does not heat the point of application to the patient.
6.2 | A reusable finger sensor for adults, with cable and plug must be provided:
6.3 | The sensor lead must be approximately 3 metres long.
7 | LCD Display:
7.1 | The display must offer clear viewing under all lighting conditions.
7.2 | The following must be displayed:
7.3 | Systolic, diastolic and mean arterial pressures (numeric)
7.4 | Oxygen saturation (SpO2 %)
7.5 | Pulse rate (beats/min)
7.6 | Alarm limits
7.7 | SpO2 waveform
7.8 | Trend information
7.9 | Battery status
7.10 | Indication of mains or battery operation
7.11 | The unit must indicate whether the signal to the oximeter sensor is adequate for reliable measurements. If the signal strength or quality is inadequate, this must be clearly displayed.
8 | ALARMS:
8.1 | All alarms must have default settings
8.2 | The alarm parameters must be user-adjustable:
8.3 | Adjustable high and low alarm limits for systolic, diastolic, mean arterial pressure and heart rate.
8.4 | SpO2 LOW: 50 % - 99 %
8.5 | Pulse rate LOW: 40 – 150 b.p.m.
8.6 | Pulse rate HIGH: 70 – 250 b.p.m.
8.7 | SpO2 probe not attached to patient or the pulse oximeter
8.8 | Violation of limits must result in both audible and visual alarms and indication of the parameters affected.
8.9 | Adjustable audio alarm volume
8.10 | Audio alarm mute: 1 minute.
8.11 | Low battery: Alarm at least 10 minutes before battery power is exhausted
9 | GENERAL
9.1 | The unit must operate from 220 V, 50 Hz mains supply
9.2 | Mains cable at least three metres long with SABS approved 15A three prong plug.
9.3 | The unit must include an internal rechargeable battery
9.4 | The unit must be able to function from the mains supply regardless of the battery condition.
9.5 | Battery must have the capacity for the unit to operate fully for at least five hours
9.6 | The battery must be charged while the unit is operating from the mains supply.
9.7 | The battery must automatically take over in the event of a mains power failure.
9.8 | The power supply must be an integral part of the unit.
9.9 | The case of the unit must be robust and impact resistant.
9.10 | The design must be such that spilled liquids do not gain access to the internal working of the unit.
10 | The unit must be handed over in full operating order
11 | A starter-pack of consumables must be supplied with the unit.
12 | The following accessories must be supplied with the unit:
12.1 | Adult NIBP cuff
12.2 | Air hose
12.3 | SpO2 Finger sensor for adults with cable.
12.4 | Battery
13 | List ALL available consumables with prices incl VAT
14 | List ALL available accessories with prices incl VAT

## NIBP and SpO2 2
| NIBP + SpO2 monitor, on trolley | Complies Y/N | Comment
1 | The unit must measure both NIBP and SpO2
2 | Self test procedure at switch-on
3 | The unit must have trending for at least 24 hours
4 | NIBP
4.1 | The unit must measure blood pressure non-invasively.
4.2 | Must be suitable for adults, paeds and neonates
4.3 | The unit must use the oscillometric method of measurement.
4.4 | Modes: Manual, automatic and STAT
4.5 | Pressure measurements:
4.6 | Adult: Systolic: 30 - 270 mm Hg, Diastolic: 20 - 240 mm Hg, MAP: 20 - 250 mm Hg
4.7 | Paeds: Systolic: 30 - 160 mm Hg, Diastolic: 20 - 120, MAP: 20 - 120 mm Hg
4.8 | Neonate: Systolic: 30 - 130 mm Hg, Diastolic: 20 - 100 mm Hg, MAP: 20 - 120 mm Hg
4.9 | Auto zero when unit is switched on
4.10 | The interval between measurements must be selectable from 1 – 120 minutes.
4.11 | Over pressure protection: Automatic deflation at 300 mm Hg adult, 150 mm Hg neonate
4.12 | Average measurement time: approximately 30 s
4.13 | If measurement exceeds two minutes the cuff must deflate, an alarm must sound and likely faults indicated.
4.14 | The air hose must be at least two metres long.
5 | PULSE OXIMETER
5.1 | Suitable for adult, paediatric and neonatal use.
5.2 | The unit must continuously measure and display SpO2 %
5.3 | Oxygen saturation (SpO2 %): Range of measurement: 0 - 100%
5.4 | The accuracy must be better than 2,5 % (70 - 100%) and 5% (50 - 70%).
5.5 | Pulse rate (beats/min): Range of measurement: 20 - 300 bpm
5.6 | The pulse rate accuracy must be better than ± 2 %.
5.7 | An audible bleep should sound for each heartbeat, the pitch of which should vary with the oxygen saturation.
5.8 | Automatic rejection of motion artefacts
6 | SpO2 SENSOR:
6.1 | Sensor: Dual LED, Infra-red/visible red, which does not heat the point of application to the patient.
6.2 | A reusable finger sensor for adults, with cable and plug must be provided:
6.3 | The sensor lead must be approximately 3 metres long.
7 | LCD Display:
7.1 | The display must offer clear viewing under all lighting conditions.
7.2 | The following must be displayed:
7.3 | Systolic, diastolic and mean arterial pressures (numeric)
7.4 | Oxygen saturation (SpO2 %)
7.5 | Pulse rate (beats/min)
7.6 | Alarm limits
7.7 | SpO2 waveform
7.8 | Trend information
7.9 | Battery status
7.10 | Indication of mains or battery operation
7.11 | The unit must indicate whether the signal to the oximeter sensor is adequate for reliable measurements. If the signal strength or quality is inadequate, this must be clearly displayed.
8 | ALARMS:
8.1 | All alarms must have default settings
8.2 | The alarm parameters must be user-adjustable:
8.3 | Adjustable high and low alarm limits for systolic, diastolic, mean arterial pressure and heart rate.
8.4 | SpO2 LOW: 50 % - 99 %
8.5 | Pulse rate LOW: 40 – 150 b.p.m.
8.6 | Pulse rate HIGH: 70 – 250 b.p.m.
8.7 | SpO2 probe not attached to patient or the pulse oximeter
8.8 | Violation of limits must result in both audible and visual alarms and indication of the parameters affected.
8.9 | Adjustable audio alarm volume
8.10 | Audio alarm mute: 1 minute.
8.11 | Low battery: Alarm at least 10 minutes before battery power is exhausted
9 | GENERAL
9.1 | The unit must operate from 220 V, 50 Hz mains supply
9.2 | Mains cable at least three metres long with SABS approved 15A three prong plug.
9.3 | The unit must include an internal rechargeable battery
9.4 | The unit must be able to function from the mains supply regardless of the battery condition.
9.5 | Battery must have the capacity for the unit to operate fully for at least five hours
9.6 | The battery must be charged while the unit is operating from the mains supply.
9.7 | The battery must automatically take over in the event of a mains power failure.
9.8 | The power supply must be an integral part of the unit.
9.9 | The length of the mains cable must be a minimum of three metres.
9.10 | The case of the unit must be robust and impact resistant.
9.11 | The design must be such that spilled liquids do not gain access to the internal working of the unit.
10 | The monitor must be mounted on a trolley with basket for accessories
11 | The unit must be handed over in full operating order
12 | A starter-pack of consumables must be supplied with the unit.
13 | The following accessories must be supplied with the unit
13.1 | Adult NIBP cuff
13.2 | Air hose
13.3 | SpO2 Finger sensor for adults with cable.
13.4 | Battery
14 | List ALL available consumables with prices incl VAT
15 | List ALL available accessories with prices incl VAT

## NIBP desktop
| NIBP Monitor, desktop | Complies Y/N | Comment
1 | The unit must measure blood pressure non-invasively.
2 | Must be suitable for adults, paeds and neonates
3 | The unit must use the oscillometric method of measurement.
4 | Modes: Manual, automatic and stat
5 | Pressure measurements:
5.1 | Adult: Systolic: 30 - 270 mm Hg, Diastolic: 20 - 240 mm Hg, MAP: 20 - 250 mm Hg
5.2 | Paeds: Systolic: 30 - 160 mm Hg, Diastolic: 20 - 120, MAP: 20 - 120 mm Hg
5.3 | Neonate: Systolic: 30 - 130 mm Hg, Diastolic: 20 - 100 mm Hg, MAP: 20 - 120 mm Hg
6 | Heart rate: 30 - 240 /min
7 | Values must be displayed numerically
8 | Auto zero when unit is switched on
9 | The interval between measurements must be selectable from 1 – 120 minutes.
10 | Adjustable high and low alarm limits for systolic, diastolic, mean arterial pressure and heart rate.
11 | Default alarm limits if the limits are not set by the user
12 | Alarm indication: Visual and audible with mute
13 | Adjustable audio alarm volume
14 | Audio alarm mute: 1 minute.
15 | Low battery: Alarm at least 10 minutes before battery power is exhausted
16 | Over pressure protection: Automatic deflation at 300 mm Hg adult, 150 mm Hg neonate
17 | Average measurement time: approximately 30 s
18 | If measurement exceeds two minutes the cuff must deflate, an alarm must sound and likely faults indicated.
19 | The air hose must be at least two metres long.
20 | The unit must operate from both the 220 V, 50 Hz mains supply
21 | Mains cable at least three metres long with SABS approved 15A three prong plug.
22 | Rechargeable battery incorporated in the unit
23 | The battery must be charged while the unit is operating from the mains supply.
24 | Battery charger to be incorporated in the unit
25 | The battery must automatically take over in the event of a mains power failure.
26 | The length of the mains cable must be a minimum of three metres.
27 | The battery must operate for at least five hours with measurements every five minutes
28 | The unit must be impervious to liquids
29 | The case of the unit must be robust and impact resistant.
30 | The unit must be handed over in full operating order
31 | A starter-pack of consumables must be supplied with the unit.
32 | The following accessories must be supplied with the unit
32.1 | Adult NIBP cuff
32.2 | Air hose
32.3 | Battery
33 | List ALL available consumables with prices incl VAT
34 | List ALL available accessories with prices incl VAT

## NIBP mobile
| NIBP Monitor on mobile stand | Complies Y/N | Comment
1 | The unit must measure blood pressure non-invasively.
2 | Must be suitable for adults, paeds and neonates
3 | The unit must use the oscillometric method of measurement.
4 | Modes: Manual and automatic
5 | Pressure measurements:
5.1 | Adult: Systolic: 30 - 270 mm Hg, Diastolic: 20 - 240 mm Hg, MAP: 20 - 250 mm Hg
5.2 | Paeds: Systolic: 30 - 160 mm Hg, Diastolic: 20 - 120, MAP: 20 - 120 mm Hg
5.3 | Neonate: Systolic: 30 - 130 mm Hg, Diastolic: 20 - 100 mm Hg, MAP: 20 - 120 mm Hg
6 | Heart rate: 30 - 240 /min
7 | Values must be displayed numerically
8 | Auto zero when unit is switched on
9 | The interval between measurements must be selectable from 1 – 120 minutes.
10 | Adjustable high and low alarm limits for systolic, diastolic, mean arterial pressure and heart rate.
11 | Default alarm limits if the limits are not set by the user
12 | Alarm indication: Visual and audible with mute
13 | Adjustable audio alarm volume
14 | Audio alarm mute: 1 minute.
15 | Low battery: Alarm at least 10 minutes before battery power is exhausted
16 | Over pressure protection: Automatic deflation at 300 mm Hg adult, 150 mm Hg neonate
17 | Average measurement time: approximately 30 s
18 | If measurement exceeds two minutes the cuff must deflate, an alarm must sound and likely faults indicated.
19 | The air hose must be at least two metres long.
20 | The unit must operate from both the 220 V, 50 Hz mains supply
21 | Mains cable at least three metres long with SABS approved 15A three prong plug.
22 | Rechargeable battery incorporated in the unit
23 | The battery must be charged while the unit is operating from the mains supply.
24 | Battery charger to be incorporated in the unit
25 | The battery must automatically take over in the event of a mains power failure.
26 | The length of the mains cable must be a minimum of three metres.
27 | The battery must operate for at least five hours with measurements every five minutes
28 | Data storage: 100 patients
29 | The unit must be impervious to liquids
30 | The case of the unit must be robust and impact resistant.
31 | NIBP must be mounted on a moble stand/trolley with basket for accessories
32 | The unit must be handed over in full operating order
33 | A starter-pack of consumables must be supplied with the unit.
34 | The following accessories must be supplied with the unit
34.1 | Adult NIBP cuff
34.2 | Air hose
34.3 | Battery
35 | List ALL available consumables with prices incl VAT
36 | List ALL available accessories with prices incl VAT

## O2 flowmeter single
| Oxygen Flowmeter, single | Complies Y/N | Comment
1 | Flowmeter must be of the pressure compensated type and fit directly into the Oxygen outlet.
2 | The flowmeter must have a stainless steel float in a calibrated flow tube
3 | Must have a flow accuracy within 10%.
4 | The range of 0 – 15 l/m of the flowmeter must be permanent and easy to read.
5 | It must have one Hayer or Drager connection, 1 flow meter with 1outlet
6 | The supplier must be able to supply the unit with a Hayer or Drager connecter. End-user to specify
7 | Must have a chrome-plated brass body and adjustment knob
8 | The flowmeter must operate accurately on a 4 bar pipeline source.
9 | The flowmeter hood must be made of crack–resistant polycarbonate.
10 | Each flowmeter must be supplied with a metal oxygen outlet ‘tail’ fitting.
11 | The flow meter must be easy to service and maintain
12 | 5 year guarantee
13 | The unit must be handed over in full operating order
14 | A starter-pack of consumables must be supplied with the unit.
15 | List ALL available consumables with prices incl VAT
16 | List ALL available accessories with prices incl VAT
17 | Optional extras
17.1 | Gabler snap lock including oxygen hose and Hayer or Drager connection
17.2 | Re-usable humidifier manufactured from a durable polycarbonate with a metal inlet connector. Operating from a pressure: ±400kpa (4 Bar) with a upper & lower level marking. The unit must have a metal safety pressure release valve.

## O2 flowmeter double
| Oxygen Flowmeter, double | Complies Y/N | Comment
1 | Flowmeter must be of the pressure compensated type and fit directly into the Oxygen outlet.
2 | The flowmeter must have a stainless steel float in a calibrated flow tube
3 | Must have a flow accuracy within 10%.
4 | The range of 0 – 15 l/m of the flowmeter must be permanent and easy to read.
5 | It must have one Hayer or Drager connection, 2 flow meters with 2 outlets
6 | The supplier must be able to supply the unit with a Hayer or Drager connecter. End-user to specify
7 | Must have a chrome-plated brass body and adjustment knob
8 | The flowmeter must operate accurately on a 4 bar pipeline source.
9 | The flowmeter hood must be made of crack–resistant polycarbonate.
10 | Each flowmeter must be supplied with a metal oxygen outlet ‘tail’ fitting.
11 | The flow meter must be easy to service and maintain
12 | 5 year guarantee
13 | The unit must be handed over in full operating order
14 | A starter-pack of consumables must be supplied with the unit.
15 | List ALL available consumables with prices incl VAT
16 | List ALL available accessories with prices incl VAT
17 | Optional extras
17.1 | Gabler snap lock including oxygen hose and Hayer or Drager connection
17.2 | Re usable humidifier manufactured from a durable polycarbonate with a metal inlet connector. Operating from a pressure: ±400kpa (4 Bar) with a upper & lower level marking. The unit must have a metal safety pressure release valve.

## Plaster of paris cutter
| PLASTER OF PARIS CUTTER, electric, with Vacuum | Complies Y/N | Comment
1 | Electric plaster cutter to remove all types/materials of casts.
2 | Vacuum unit to collect cast dust
3 | Small cast spreader (stainless steel)
4 | Large cast spreader (stainless steel)
5 | The system must have a variable speed control
6 | Must be mounted on a stable mobile stand
7 | The unit must have an impact resistant housing.
8 | Noise: < 70 dB
9 | The unit must have a vacuum hose at least 2m long.
10 | The hose mount must allow for easy movement of the vacuum hose.
11 | The vacuum canister and filter must have a quick and easy method of emptying and cleaning
12 | Reusable dust bag, capacity of at least 7.0 litres.
13 | The vacuum must turn on automatically when the cast cutter is activated.
14 | Spare parts must be available for at least ten years after production of the unit stops.
15 | CUTTER:
15.1 | · Weight: ± 1.5 kg
15.2 | · Speed: Variable speed up to 24, 000 cycles/min
16 | The unit must be double insulated
17 | Power: 220V 50Hz
18 | Power cord at least 3 m long with SABS approved 15A three prong plug top.
19 | The unit must have an ON/OFF switch
20 | On switching off the saw, the extraction system must purge the system before switching off.
21 | Must have an overload protection system.
22 | The filtration system must retain > 99% of the dust particles.
23 | The system must be TV and radio suppressed.
24 | The saw must be able to take a range of blades.
25 | The unit must be handed over in full operating order
26 | A starter-pack of consumables must be supplied with the unit.
27 | The following accessories must be supplied with the unit:
27.1 | 1 x 50mm blade
27.2 | 1 x 65mm blade
27.3 | Tools to replace the blades.
27.4 | Re-usable vacuum filters
27.5 | Re-usable vacuum bags
27.6 | Re-usable liners
28 | List ALL available consumables with prices incl VAT
29 | List ALL available accessories with prices incl VAT

## Phototherapy LED
| Phototherapy Light (LED) | Complies Y/N | Comment
1 | The phototherapy light system must be suitable for the treatment of neonatal hyperbilirubinemia.
2 | Microprocessor control with LCD display of set treatment time, therapy time, lamp usage time, intensity level, infant skin temp
3 | The light system must be suitable for babies in a bassinet, incubator, open incubator or radiant warmer.
4 | The light source must be LEDs, emiting light in the range of 450 – 475nm
5 | Target light to ensure correct placement of lamp over patient
6 | The unit must have at least two intensity settings, high and low.
7 | The high setting must give an output of 30-45 µW/cm²/nm at a distance of 300 mm from the light enclosure.
8 | The low setting must give an output of 12-15 µW/cm²/nm at a distance of 300 mm from the light enclosure.
9 | The LED light source must ensure a uniform illumination pattern at the skin surface
10 | The light source housing must tilt at least to 30o in either direction to provide effective phototherapy
11 | The lamp effective surface area must be at least 400 x 200 mm at 300 mm from the light source
12 | The unit must be height adjustable : approximately 1200 - 1600 mm from light source to floor
13 | The light enclosure must be protected from the ingress of fluids.
14 | The light unit must be mounted on a roll stand.
15 | The roll stand must have lockable castors.
16 | The base of the roll stand must not tip, even when the light enclosure is tilted.
17 | Must have a re-settable treatment timer that will switch off the light when the set time has expired
18 | The unit must have a time recorder that operates only when the phototherapy light is switched on. It must be possible to reset the time recorder to zero when the light source is replaced.
19 | It must not be possible to accidently reset the time recorder
20 | LED array replacement must be a simple operation
21 | It must be possible to replace only a section of the LED array
22 | The unit must operate off the 220 Volt, 50Hz power supply.
23 | Mains cable at least three metres long with SABS approved 15A three prong plug.
24 | Warranty period for LED light source of no less than 20000 hrs or 3yrs
25 | Must comply with SANS/IEC 60601-2-50:2010
26 | The unit must be handed over in full operating order
27 | A starter-pack of consumables must be supplied with the unit.
28 | List ALL available consumables with prices incl VAT
29 | List ALL available accessories with prices incl VAT
30 | OPTIONAL
30.1 | Monitoring of baby temperature
30.2 | LED examination light

## Phototherapy Blanket LED
| Phototherapy Blanket LED | Complies Y/N | Comment
1 | The phototherapy light system must be suitable for the treatment of neonatal hyperbilirubinemia.
2 | Microprocessor control with LCD display of set treatment time, therapy time, lamp usage time, intensity level, infant skin temp
3 | The light system must be suitable for babies in a bassinet, incubator, open incubator or radiant warmer.
4 | The light source must be LEDs, emiting light in the range of 450 – 475nm
5 | The unit must have at least two intensity settings, high and low.
6 | Intensity: Maximum 30-35uW/cm2/nm and Minimum 12-15uW/cm2/nm
7 | Effective treatment area: Small pad of 250-350cm2
8 | The unit must not emit light in the Ultraviolet or Infrared radiation range
9 | The unit must consist of a soft, flexible, oval design that conforms to the shape of the baby to deliver phototherapy via an LED light source
10 | The unit must be able to shut off automatically in the event of the blanket overheating
11 | Must have a re-settable treatment timer that will switch off the light when the set time has expired
12 | The unit must have a time recorder that operates only when the phototherapy light is switched on. It must be possible to reset the time recorder to zero when the light source is replaced.
13 | It must not be possible to accidently reset the time recorder
14 | LED array replacement must be a simple operation
15 | It must be possible to replace only a section of the LED array
16 | The unit must operate off the 220 Volt, 50Hz power supply.
17 | Mains cable at least three metres long with SABS approved 15A three prong plug.
18 | Warranty period for LED light source of no less than 20000 hrs or 3yrs
19 | Must comply with SANS/IEC 60601-2-50:2010
20 | The unit must be handed over in full operating order
21 | A starter-pack of consumables must be supplied with the unit.
22 | List ALL available consumables with prices incl VAT
23 | List ALL available accessories with prices incl VAT
24 | OPTIONAL
24.1 | Monitoring of baby temperature
24.2 | Effective treatment area: Large pad of 450-550cm2

## Pulse Oximeter Desk
| PULSE OXIMETER, desktop | Complies Y/N | Comment
1 | Suitable for adult, paediatric and neonatal use.
2 | The unit must continuously measure and display SpO2 %
3 | Oxygen saturation (SpO2 %): Range of measurement: 0 - 100%
4 | The accuracy must be better than 2,5 % (70 - 100%) and 5% (50 - 70%).
5 | Pulse rate (beats/min): Range of measurement: 20 - 300 bpm
6 | The pulse rate accuracy must be better than ± 2 %.
7 | An audible bleep should sound for each heartbeat, the pitch of which should vary with the oxygen saturation.
8 | Automatic rejection of motion artefacts
9 | The unit must have trending for at least 72 hours
10 | Self test procedure at switch-on
11 | The display must offer clear viewing under all lighting conditions.
12 | The following must be displayed:
12.1 | Oxygen saturation (SpO2 %)
12.2 | Pulse rate (beats/min)
12.3 | Alarm limits
12.4 | Pulse waveform
12.5 | Trend information
12.6 | Battery status
12.7 | Indication of mains or battery operation
13 | The unit must indicate whether the signal to the oximeter sensor is adequate for reliable measurements. If the signal strength or quality is inadequate, this must be clearly displayed.
14 | SENSOR:
14.1 | Sensor: Dual LED, Infra-red/visible red, which does not heat the point of application to the patient.
14.2 | The sensor lead must be approximately 3 metres long.
15 | ALARMS:
15.1 | All alarms must have default settings
15.2 | The alarm parameters must be user-adjustable:
15.3 | SpO2 LOW: 50 % - 99 %
15.4 | Pulse rate LOW: 40 – 150 b.p.m.
15.5 | Pulse rate HIGH: 70 – 250 b.p.m.
15.6 | Probe not attached to patient or the pulse oximeter
15.7 | Violation of these limits must result in both audible and visual alarms and indication of the parameters affected.
15.8 | Adjustable audio alarm volume
15.9 | Audio alarm mute: 1 minute. The unit must resume normal alarm functions after the minute has expired.
15.10 | Low battery: Alarm at least 10 minutes before battery power is exhausted
16 | The unit must operate from 220 V, 50 Hz mains supply
17 | Mains cable at least three metres long with SABS approved 15A three prong plug.
18 | The unit must include an internal rechargeable battery
19 | Battery charger to be incorporated in the unit
20 | The unit must be able to function from the mains supply regardless of the battery condition.
21 | Battery must have the capacity for the unit to operate fully for at least five hours
22 | The battery must be charged while the unit is operating from the mains supply.
23 | The battery must automatically take over in the event of a mains power failure.
24 | Automatic switch-off after a period of non-use
25 | The case of the unit must be robust and impact resistant.
26 | The design must be such that spilled liquids do not gain access to the internal working of the unit.
27 | The unit must be handed over in full operating order
28 | A starter-pack of consumables must be supplied with the unit.
29 | The following accessories must be supplied with the unit:
29.1 | SpO2 Finger sensor for adults with cable.
29.2 | Battery
30 | List ALL available consumables with prices incl VAT
31 | List ALL available accessories with prices incl VAT

## Pulse Oximeter trolley
| PULSE OXIMETER on trolley | Complies Y/N | Comment
1 | Suitable for adult, paediatric and neonatal use.
2 | The unit must continuously measure and display SpO2 %
3 | Oxygen saturation (SpO2 %): Range of measurement: 0 - 100%
4 | The accuracy must be better than 2,5 % (70 - 100%) and 5% (50 - 70%).
5 | Pulse rate (beats/min): Range of measurement: 20 - 300 bpm
6 | The pulse rate accuracy must be better than ± 2 %.
7 | An audible bleep should sound for each heartbeat, the pitch of which should vary with the oxygen saturation.
8 | Automatic rejection of motion artefacts
9 | The unit must have trending for at least 72 hours
10 | Self test procedure at switch-on
11 | The display must offer clear viewing under all lighting conditions.
12 | The following must be displayed:
12.1 | Oxygen saturation (SpO2 %)
12.2 | Pulse rate (beats/min)
12.3 | Alarm limits
12.4 | Pulse waveform
12.5 | Trend information
12.6 | Battery status
12.7 | Indication of mains or battery operation
13 | The unit must indicate whether the signal to the oximeter sensor is adequate for reliable measurements. If the signal strength or quality is inadequate, this must be clearly displayed.
14 | SENSOR:
14.1 | Sensor: Dual LED, Infra-red/visible red, which does not heat the point of application to the patient.
14.2 | The sensor lead must be approximately 3 metres long.
15 | ALARMS:
15.1 | All alarms must have default settings
15.2 | The alarm parameters must be user-adjustable:
15.3 | SpO2 LOW: 50 % - 99 %
15.4 | Pulse rate LOW: 40 – 150 b.p.m.
15.5 | Pulse rate HIGH: 70 – 250 b.p.m.
15.6 | Probe not attached to patient or the pulse oximeter
15.7 | Violation of these limits must result in both audible and visual alarms and indication of the parameters affected.
15.8 | Adjustable audio alarm volume
15.9 | Audio alarm mute: 1 minute. The unit must resume normal alarm functions after the minute has expired.
15.10 | Low battery: Alarm at least 10 minutes before battery power is exhausted
16 | The unit must operate from 220 V, 50 Hz mains supply
17 | Mains cable at least three metres long with SABS approved 15A three prong plug.
18 | The unit must include an internal rechargeable battery
19 | Battery charger to be incorporated in the unit
20 | The unit must be able to function from the mains supply regardless of the battery condition.
21 | Battery must have the capacity for the unit to operate fully for at least five hours
22 | The battery must be charged while the unit is operating from the mains supply.
23 | The battery must automatically take over in the event of a mains power failure.
24 | Automatic switch-off after a period of non-use
25 | The case of the unit must be robust and impact resistant.
26 | The design must be such that spilled liquids do not gain access to the internal working of the unit.
27 | The unit must be mounted on a trolley/roll stand.
28 | The trolley must be stable and no easily knocked over
29 | The following accessories must be supplied with the unit:
29.1 | SpO2 Finger sensor for adults with cable.
29.2 | Battery
30 | The unit must be handed over in full operating order
31 | A starter-pack of consumables must be supplied with the unit.
32 | List ALL available consumables with prices incl VAT
33 | List ALL available accessories with prices incl VAT

## Resuscitator gas powered
| Resuscitator, infant, gas powered | Complies Y/N | Comment
1 | Must provide 100% oxygen for infants up to 10 kg
2 | Unit must be gas powered via the wall oxygen supply.
3 | Oxygen flow regulator: 5 - 15 l/min
4 | Adjustable Peak Inspiratory Pressure (PIP): 2 - 75 cm H2O
5 | Adjustable Positive End Expiratory Pressure (PEEP): 1 - 10 cm H2O
6 | Resuscitator must provide controlled pressure settings by means of a pressure limiting device
7
8 | Pressure gauge: -10 to 80 cm H2O
9 | Maximum pressure relief: 5 - 70cmH²O @ 8lpm.
10 | The unit must be handed over in full operating order
11 | A starter-pack of consumables must be supplied with the unit.
12 | List ALL available consumables with prices incl VAT
13 | List ALL available accessories with prices incl VAT
14 | Optional Items, quote prices incl VAT
14.1 | Air/O2 blender

## Suction mobile 1
| Suction Unit, mobile, battery operated | Complies Y/N | Comment
1 | Unit must be mounted on mobile stand with at least 4 x castor wheels.
2 | Must contain a rechargable battery with built-in charger
3 | Two of the castors must be lockable
4 | Operating pressure: 0 to -80 kPa
5 | The gauge must be marked marked in kPa and mmHg
6 | The gauge must be fully removable for maintenance.
7 | The suction capacity must be at least 25 litres per minute.
8 | On-off switch must be illuminated and have a splash-proof cover
9 | Must have suction control for accurate regulation of the vacuum
10 | Two receiver jars, each of 2 litre capacity, graduated in 100 ml steps
11 | Receiver jars must be high impact crack resistant polycarbonate
12 | Receiver jars and lids must be removable and autoclavable
13 | Receiver jar lid must have overflow protection
14 | Receiver jar must operate with and without liners.
15 | Vacuum pump regulator must be fitted with a hydrophobic bacterial filter to prevent aspirated fluid from entering the suction pump.
16 | Working noise less than 50 dB
17 | The suction inlets and connectors must be stainless steel.
18 | The pump must be oil free and maintenance free
19 | Vacuum regulator should operate on high vacuum 7 to 67 kPa and low vacuum 0 to 19 kPa.
20 | The vacuum regulator must be fitted with a probe of the S.A.B.S./Heyer type
21 | The mobile stand must have a handle for pushing
22 | The unit must operate from a 220 V AC supply.
23 | Must have a mains cable at least 3 m long fitted with a SABS approved three pin 15A plug
24 | Battery status indicator
25 | Battery capacity: at least 1 hour operation
26 | Battery must be sealed, maintenance free
27 | The unit must meet the appropriate ISO standards
28 | The unit must be CE approved
29 | The suction unit must be easy to service and maintain.
30 | The unit must be handed over in full operating order
31 | A starter-pack of consumables must be supplied with the unit.
32 | The following accessories must be supplied with the unit:
32.1 | Silicone tubes set 8 x 14 mm
32.2 | Tube connectors
32.3 | Probe connector
32.4 | Aspiration probe.
33 | List ALL available consumables with prices incl VAT
34 | List ALL available accessories with prices incl VAT

## Suction mobile 2
| Suction Unit, Mobile | Complies Y/N | Comment
1 | Unit must be mounted on mobile stand with at least 4 x castor wheels.
2 | Two of the castors must be lockable
3 | Operating pressure: 0 to -80 kPa
4 | The gauge must be marked marked in kPa and mmHg
5 | The gauge must be fully removable for maintenance.
6 | The suction capacity must be at least 25 litres per minute.
7 | On-off switch must be illuminated and have a splash-proof cover
8 | Must have suction control for accurate regulation of the vacuum
9 | Two receiver jars, each of 2 litre capacity, graduated in 100 ml steps
10 | Receiver jars must be high impact crack resistant polycarbonate
11 | Receiver jars and lids must be removable and autoclavable
12 | Receiver jar lid must have overflow protection
13 | Receiver jar must operate with and without liners.
14 | Vacuum pump regulator must be fitted with a hydrophobic bacterial filter to prevent aspirated fluid from entering the suction pump.
15 | Working noise less than 50 dB
16 | The suction inlets and connectors must be stainless steel.
17 | The unit must be CE approved
18 | The suction unit must be easy to service and maintain.
19 | The pump must be oil free and maintenance free
20 | Vacuum regulator should operate on high vacuum 7 to 67 kPa and low vacuum 0 to 19 kPa.
21 | The vacuum regulator must be fitted with a probe of the S.A.B.S./Heyer type
22 | The mobile stand must have a handle for pushing
23 | The unit must operate from a 220 V AC supply.
24 | Must have a mains cable at least 3 m long fitted with a SABS approved three pin 15A plug
25 | Battery status indicator
26 | Battery capacity: at least 1 hour operation
27 | Battery must be sealed, maintenance free
28 | The unit must meet the appropriate ISO standards
29 | The unit must be handed over in full operating order
30 | A starter-pack of consumables must be supplied with the unit.
31 | The following accessories must be supplied with the unit:
31.1 | Silicone tubes set 8 x 14 mm
31.2 | Tube connectors
31.3 | Probe connector
31.4 | Aspiration probe.
32 | List ALL available consumables with prices incl VAT
33 | List ALL available accessories with prices incl VAT

## Suction wallmount
| Suction unit, wall-mount | Complies Y/N | Comment
1 | Must be a complete single suction unit with an accessory rail mounting clip.
2 | The unit must have a Heyer connection
3 | The unit must be of robust design.
4 | Two receiver jars, each of 2 litre capacity, graduated in 100 ml steps
5 | Receiver jars must be high impact crack resistant polycarbonate
6 | The receiving jar with lid must be fully autoclavable.
7 | Receiver jar lid must have overflow protection
8 | The unit offered must be able to be used with and without liners.
9 | The suction inlets and connectors must be stainless steel or chrome plated brass.
10 | Suction power must be controllable by means of a vacuum control
11 | The unit must have a durable on/off switch
12 | Operating pressure: 0 to -80 kPa
13 | The gauge must be marked marked in kPa and mmHg
14 | The gauge must be fully removable for quick maintenance.
15 | The pressure hose must be colour coded yellow and labeled "vacuum".
16 | Length of the pressure hose: 1.5 m
17 | The suction unit must be easy to service and maintain.
18 | The suction capacity must be at least 25 litres per minute.
19 | The unit must be CE approved
20 | Water trap/safety trap
21 | The unit must meet the appropriate ISO standards
22 | The unit must be handed over in full operating order
23 | A starter-pack of consumables must be supplied with the unit.
24 | The following accessories must be supplied with the unit:
24.1 | Silicone tubes set 8 x 14 mm
24.2 | Tube connectors
24.3 | Probe connector
24.4 | Aspiration probe.
24.5 | Particle filter
24.6 | Suction tube with canula, reusable
24.7 | Vacuum probe
24.8 | Hose vacuum probe
25 | List ALL available consumables with prices incl VAT
26 | List ALL available accessories with prices incl VAT

## Thermometer electronic
| Thermometer, electronic | Complies Y/N | Comment
1 | Must be able to measure temporal artery temperature at least 3 cm away from forehead
2 | Suitable for all patients
3 | Must measure temperature in °C
4 | Measurement range: 32 - 42°C
5 | Must be able to take body temperature regardless of room temperature
6 | Accuracy: < 0,2°C
7 | Must provide reliable and stable results
8 | Display: LCD with backlight
9 | Water resistant
10 | Self-test when switched on
11 | Low battery display
12 | Automatic shut off after 2 min of non-use
13 | No. of measurements with one set of batteries > 40 000
14 | Hand-held, convenient and easy to use
15 | The unit must be handed over in full operating order
16 | A starter-pack of consumables must be supplied with the unit.
17 | List ALL available consumables with prices incl VAT
18 | List ALL available accessories with prices incl VAT

## Ventilator CPAP
| VENTILATOR CPAP | Complies Y/N | Comment
1 | An electronically controlled & monitored CPAP flow system to safely and effectively deliver fresh gas via a nasal interface or mask to the airways of neonate and paediatric patients.
2 | Electrically powered and microprocessor driven, monitored and controlled.
3 | The CPAP system must consist of the following:
3.1 | Blender
3.2 | O2 analyser
3.3 | Flowmeter
3.4 | Electronic pressure gauge
3.4 | Alarm system
4 | Manual breath inspiration/hold feature with ability to ensure sustained lung inflation which is used as a recruitment manoeuvre on infants with atelectasis
5 | Manual time for a manual breath adjustable from 2 to 10 sec for this manoeuvre
6 | Manual breath delivered in any mode adjustable 5 - 15 cm H2O
7 | Oxygen flush to ensure good oxygenation during suctioning etc 23 - 100%
8 | Must have leak compensation to allow for various cannulae and prongs to be used
9 | Able to deliver Demand flow in nCPAP thereby reducing the work of breathing and not have pre-set limited flow in nCPAP
10 | CPAP pressure range: 0 to 35 cm H2O
11 | nCPAP 2 - 13 cm H2O
12 | O2 flush time adjustable 30 - 120 sec
13 | The unit must be able to supply two independent pressure levels via 2 flow meters.
14 | Correctly coded high pressure gas supply hoses for oxygen and air of three metre length must be supplied with the unit. Appropriate type oxygen and air probes to be fitted. Probes to be in accordance with SABS 1409.
15 | Supply air/oxygen mixture to the patient at a flow rate between 0 – 15 lpm (95% accuracy)
16 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 and 601-1-2 for medical equipment, attached certification
17 | System must comply to ISO 13458 standards, attach proof of compliance
18 | The control unit must automatically go through a self diagnosis and checking process.
19 | The control unit must have direct input connections for oxygen and medical air
20 | The source for oxygen either from a central wall outlet (gas bank) or an oxygen cylinder.
21 | The source for medical air either from a central wall outlet (gas bank) or a medical air cylinder.
22 | The input supply pressure for oxygen and medical air: 200 – 500 kPa.
23 | The unit must automatically switch between the piped air and oxygen supply should one of the supplies fail. (Alarm condition should be triggered in any event)
24 | Must provide compensation for leaks
25 | The unit must be user friendly with all circuit connections, front panel controls and displays ergonomically grouped by function and clearly marked.
26 | Oxygen concentration setting variable between 21 % and 100 %.
27 | The unit must be able to supply two independent pressure levels via 2 flow meters.
28 | Correctly coded high pressure gas supply hoses for oxygen and air of three metre length must be supplied with the unit. Appropriate type oxygen and air probes to be fitted. Probes to be in accordance with SABS 1409.
29 | Supply air/oxygen mixture to the patient at a flow rate between 0 – 15 lpm (95% accuracy)
30 | Have a manometer with a range of 0 – 12 cm H2O
31 | Breath rate: 1 - 120 bpm
32 | Timed inspiration: 0.1 to 3.0 sec.
33 | Apnoea backup ventilation: 1 - 80 bpm.
34 | Apnoea interval: 10 - 30 seconds.
35 | Maximum safety pressure set to ± 3 cm H20 above measured airway pressure.
36 | Unit must have safety features to prevent volutrauma, barotrauma and over stretching of the lungs in infants.
37 | The unit must have a proximal pressure monitoring line.
38 | Must have a method to prevent it becoming occluded by water vapour.
39 | Must monitor the following parameters:
39.1 | nCPAP
39.2 | MAP
39.3 | PIP
39.4 | PEEP
39.5 | Fi02 (21% to 100%)
40 | Graphics Display
40.1 | Measured values to be displayed:
40.1.1 | CPAP
40.1.2 | PIP
40.1.3 | MAP
40.2 | The graphics display system must have freeze capabilities.
41 | The offered unit must display trending of at least 72 hours on at least 20 parameters of pressure-, flow- and tidal volumes.
42 | Alarms:
42.1 | Patient disconnection
42.2 | High airway pressure
42.3 | Low airway pressure.
42.4 | High oxygen alarm
42.5 | Low oxygen alarm
42.6 | Low nCPAP alarm: -2 to -20 cm H2O
42.7 | High nCPAP alarm: +2 to + 20 cm H2O
42.8 | Oxygen supply failure
42.9 | Air supply failure
42.10 | Apnoea: off; 10 - 20 sec
42.11 | Electrical supply failure
42.12 | Low battery
43 | A safety mechanical pressure release (overpressure) valve must be present
44 | The above alarm parameters must auto set themselves within 2 minutes of power up of drive gas flows and oxygen concentrations or by means of a single key setting.
45 | Violation of any of the set alarm limits must result in both audible and visual alarms.
46 | It must be possible to silence the alarm for a maximum of 120 seconds.
47 | The unit must be mounted on a mobile stand/trolley, with 4 to 5 castors (lockable) with a facility to connect a humidifier.
48 | The unit must have a rail and circuit hanger on both sides of the trolley.
49 | The unit must operate from a 220 V, 50Hz supply.
50 | Mains cable at least three metres long with SABS approved 15A three prong plug.
51 | The offered unit shall have automatic switch over to a built in battery unit with a capacity to ensure full functionality of the unit for not less than 2 hour.
52 | The unit must be fitted with an internal rechargable battery.
53 | Battery must be capable of at least 2 hours operation
54 | Battery charger to be incorporated in the unit
55 | Must display battery status and give an alarm when battery is low
56 | Must comply with EC: 93/42/EEC class iib
57 | The unit must be handed over in full operating order
58 | A starter-pack of consumables must be supplied with the unit comprising at least the following:
58.1 | 5 sets of silicone nasal prongs (each set must include 3 different neonatal sizes)
58.2 | 5 sets of CPAP bonnets (each set to include 3 different neonatal sizes)
58.3 | 5 neonatal circuits
59 | The following accessories must be supplied with the unit:
59.1 | Humidifier
59.2 | Temperature probes for humidifier
59.3 | Circuit holding arm
60 | List ALL available consumables with prices incl VAT
61 | List ALL available accessories with prices incl VAT
62 | MAINTENANCE CONTRACT
62.1 | Bidder must offer a five year full maintenance contract effective after the two year guarantee expires.
62.2 | The cost of the maintenance contract must not be included in the price of the bid
62.3 | State the cost per year and the total cost
62.4 | Provide details of the contract in a separate document

## Ventilator neonatal
| Ventilator: Neonatal/Paediatric | Complies Y/N | Comment
1 | Ventilator with the latest technology to ventilate neonates and paediatric patients (500g to 20kg) in the Neonatal (premature and full term newborns) and Paediatric Intensive Care Units
2 | The ventilator must be pneumatically powered
3 | The ventilator must be microprocessor controlled
4 | The ventilator must be time cycled
5 | The ventilator must be pressure limited
6 | Multi-mode state of the art touch screen or equivalent graphical interface
7 | Combined Pressure-limited ventilation modes
8 | Volume–controlled ventilation modes
9 | Trigger-Ventilation modes
10 | The ventilator must provide the following ventilation modes, both invasively and non-invasively for Neonatal and Paediatric ventilation:
10.1 | Assist /Control in Volume and Pressure modes
10.2 | Time Cycled mode;
10.3 | Synchronized Intermittent Mandatory Ventilation - in volume and pressure mode;
10.4 | Continuous Positive Airway Pressure with Pressure supported ventilation:
10.5 | Spontaneous;
10.6 | Manual Breath;
10.7 | Inspiration Hold;
10.8 | Expiration Hold;
11 | Back up / Apnea ventilation in all modes.
12 | The ventilator must provide the following breath types for Neonatal and Paediatric modes:
12.1 | Pressure Limited Volume Control;
12.2 | Pressure Control with Volume target;
12.3 | Pressure Ventilation with alternating baselines (CPAP/PEEP);
12.4 | Timed Oxygenation (100%).
13 | Respiratory rate per minute: 1 to 150 b/min.
14 | Inspiratory Flow: 0,4 to 100 L/min
15 | Inspired pressure: 0 to at least 60 cm H2O.
16 | Patient triggering: 0,2 to 5 LPM. (Adjustable sensitivity.)
17 | Ventilation frequency: 1 to at least 150 BPM.
18 | Tidal Volume:
18.1 | Paediatric: 20ml – 1 000ml;
18.2 | Neonates: 5 - 500 ml
19 | Inspiratory Time: 0.1 to 3 seconds;
20 | Pressure Support: 0 to 80 cm H2O.
21 | Pressure Limit: 0 to 80cmH2O
22 | Pause adjustable from 0 to 2 seconds.
23 | PEEP/CPAP: 0 to 50cmH2O.
24 | Flow Trigger: 0,1 to 5 L/min
25 | Safety features to prevent volutrama, barotraumas and over stretching of the lungs in infants.
26 | Proximal trigger option for neonate/infant.
27 | Pressure Trigger: 0.5 to 10 mbar below PEEP/CPAP.
28 | Inspiratory Hold: Up to 5 seconds.
29 | Expiratory Hold: Up to 20 seconds.
30 | Pressure High: 0 to 70 mbar.
31 | Pressure Low: 0 to 45 mbar.
32 | Oxygen Concentration from 21% to 100%.
33 | Bias flow.
34 | Leak Compensation: Up to 8 L/min.
35 | Alarms:
35.1 | Both audible and visual alarms.
35.2 | Patient disconnect
35.3 | Oxygen concentration variation
35.4 | Low inspiratory pressure.
35.5 | High breath rates 5 to 250 BPM.
35.6 | Incompatible settings.
35.7 | Low PEEP.
35.8 | Over and under patient pressure
35.9 | Over and under patient volume
35.10 | High continuous pressure
35.11 | Inverse I:E ratio
35.12 | Air/Oxygen % deviation
35.13 | Power failure
35.14 | Air and Oxygen supply failure
35.15 | Air and Oxygen low pressure.
35.16 | APNEA.
35.17 | Alarm mute. (Resets after 60 seconds.)
35.18 | Low battery.
36 | The unit must have automatic switch over to a built in battery with a capacity to ensure full functionality of the unit for not less than 2 hours.
37 | The unit must have automatic switch over that will switch between the piped air and oxygen supply should one of the supplies fail.
38 | The unit must have a circuit compliance compensation 0.0 to 7.5 ml/cmH2O and must be automatically measured.
39 | Proximal flow and pressure sensing technology for neonates.
40 | Ability to accurately deliver, monitor and blend helium/oxygen from 21 to 100% O2.
41 | Graphics Display: The following parameters and measurements must be displayed:
41.1 | Inspired tidal volume in ml
41.2 | Expired tidal volume in ml
41.3 | Peak inspiratory pressure.
41.4 | Proximal airway pressure.
41.5 | Inspired tidal volume per kg in ml/kg
41.6 | Expired tidal volume per kg in ml/kg
41.7 | Spontaneous tidal volume in ml
41.8 | Total breath rate in bpm
41.9 | Spontaneous breath rate in bpm
41.10 | Minute volume in litres
41.11 | Spontaneous minute volume in litres
41.12 | Peak airway pressure in cmH2O
41.13 | Mean airways pressure in cmH2O
41.14 | Plateau pressure in cmH2O
41.15 | I:E ratio
41.16 | Set I:E ratio
41.17 | Inspired time in seconds
41.18 | Expired time in seconds
41.19 | Static airways compliance in ml/cmH2O
41.20 | Static airways compliance per kg in ml/cmH2O
41.21 | Dynamic airways compliance in ml/cmH2O
41.22 | Dynamic airways compliance per kg in ml/cmH2O
41.23 | Peak inspired flow rate in l/m
41.24 | Peak expired flow rate in l/m
41.25 | Inspired O2 in %
41.26 | Tube leak in %
42 | At least 5 selectable waveforms must be displayed at any time on the graphic display screen.
43 | The graphics display system must have freeze capabilities.
44 | Trending of at least 24 hours on at least 8 parameters of pressure-, flow- and tidal volumes.
45 | The unit must be able to calculate and display pressure- and flow/volume loops.
46 | The ventilator must be mounted on a stable trolley with at least two of the castors fitted with brakes that conform to SANS 601.
47 | The unit must have circuit holding arm
48 | The ventilator trolley must be suitable for mounting of the dual servo humidifier
49 | The unit must be supplied a pole or universal clamp unit to mount the exiting humidifiers in use onto the unit.
50 | A 220 V power outlet point must be available on the frame.
51 | Mains cable at least three metres long with SABS approved 15A three prong plug.
52 | The ventilator must come complete with colour-coded air and O2 hoses 3-metre length. (O2 and air Heyer type keys, SABS 1409, must also be supplied.)
53 | The unit must have a rail and circuit hanger of a suitable strength on both sides of the ventilator trolley.
54 | The unit must have an integrated blender with O2 concentration 21 to 100 % ± 3 %.
55 | The unit must be handed over in full operating order
56 | A starter-pack of consumables must be supplied with the unit comprising at least the following:
56.1 | Bacteria filter (pk of 10)
56.2 | Neonatal patient circuit (Pk of 10)
56.3 | Paeds patient circuits (Pk of 10)
56.4 | Test lung for neonates
57 | List ALL available consumables with prices incl VAT
58 | List ALL available accessories with prices incl VAT
59 | Optional Items, quote prices incl VAT
59.1 | Humidifier
60 | MAINTENANCE CONTRACT
60.1 | Bidder must offer a five year full maintenance contract effective after the two year guarantee expires.
60.2 | The cost of the maintenance contract must not be included in the price of the bid
60.3 | State the cost per year and the total cost
60.4 | Provide details of the contract in a separate document

## Ventilator paed adult
| VENTILATOR PAEDIATRIC/ADULT | Complies Y/N | Comment
1 | The unit must be able to ventilate and monitor paediatric and adult patients
2 | The ventilator must be a time-cycled, volume- and pressure-controlled unit.
3 | The ventilator must have an integrated, electronically controlled air compressor or turbine
4 | The ventilator must have an oxygen port for use with an oxygen cylinder during transport
5 | Suitable preset pin indexed regulators must be supplied that can handle flows up to 100 L/pm
6 | The unit must have automatic switch-over unit that will switch between the piped air and turbine/oxygen supply should one of the supplies fail.
7 | Self diagnosis with each start-up and regular integrity testing of all system parameters
8 | Ventilator must have "Standby" mode which suspends alarms, breath cycling and trends.
9 | "Standby" mode must retain most current ventilator settings and alarms, resume normal ventilation with one touch
10 | The following breath types must be selectable:
10.1 | Volume
10.2 | Pressure
10.3 | Pressure regulated volume control
10.4 | Pressure support ventilation
10.5 | Spontaneous
11 | The ventilator must have at the least the following modes of ventilation:
11.1 | Invasive and non-invasive ventilation must be possible
11.2 | Volume ventilation with:
11.2.1 | Assist/control
11.2.2 | SIMV (Synchronous Intermittent Mandatory Ventilation)
11.2.3 | PSV (Pressure Support Ventilation)
11.2.4 | Option to disable demand flow
11.2.5 | Option of square or decelerating flow waveform
11.3 | Pressure ventilation (IPPV) with:
11.3.1 | Assist/control
11.3.2 | SIMV (Synchronous Intermittent Mandatory Ventilation)
11.3.3 | PSV (Pressure Support Ventilation)
11.3.4 | CPAP
11.3.5 | NPPV (Non-invasive positive pressure ventilation) with leak compensation and apnea back-up
11.4 | PRVC (Pressure Regulated Volume Control)
11.5 | APRV (Airway Pressure Release Ventilation) /BiPAP ( Biphasic , Duo Level, BiLevel)
11.6 | Spontaneous ventilation
11.7 | Apnea monitoring must be active in A/C, SIMV, CPAP and APRV
11.7.1 | Apnea mode: user settable time duration before ventilation onset
12 | BREATH CONTROL
12.1 | Breath rate: 1 - 150 bpm
12.2 | Tidal volume: 50 - 3000 ml
12.3 | Peak flow: 1 - 120 litres/min
12.4 | Volume limit: 5 ml - 2,5 litres
12.5 | PEEP: 0 - 50 cm H2O
12.6 | Inspiratory flow trigger: 0.5 - 15 litres/min
12.7 | Pressure trigger: 0,5 - 10 mbar below PEEP/CPAP
12.8 | Bias flow: 0.4 - 5 litres/min
12.9 | Expiratory hold: 0 - 5 sec
12.10 | Inspiratory hold: 0 - 5 sec
12.11 | Inspiratory pause: 0 - 2 sec
12.12 | Inspiratory time: 0,2 - 5 sec
12.13 | Manual breath - add single control breath
12.14 | Inspiratory waveform: square and decelerating waves
12.15 | Inspiratory pressure control level: 1 - 100 cm H2O
13 | In ARPV mode:
13.1 | Pressure high: 0 - 60 cm H2O
13.2 | Time high: 0,2 - 30 sec
13.3 | Pressure low: 0 - 40 cm H2O
13.4 | Time low: 0,3 - 30 sec
14 | Leakage compensation in all non-invasive ventilation modes must be possible.
15 | The unit must have a circuit compliance compensation 0 to 7.5 ml/cmH2O and must be automatically measured.
16 | I:E ratio: 99:1 to 1:99
17 | PEEP/CPAP must be possible from 0 to approximately 50 cm H2O
18 | OXYGEN CONTROL
18.1 | FiO2: 21 – 100%
18.2 | Increase % oxygen: single key activation to deliver up to100% O2 for 2 min. FiO2 to be adjustable by user
18.3 | Ventilator must have an integrated nebulizer port capable of delivering a set FiO2, synchronized with the inspiratory phase. Must be timed and volume compensated with a peak flow ≥ 15 L/min
19 | MONITOR
19.1 | Volume monitoring:
19.1.1 | Monitored values to include Tidal and Minute Volumes divided by patent's weight (ml/kg)
19.1.2 | Tidal volume: 0 - 4 litres
19.1.3 | Minute volume: 0 - 99 litres
19.1.4 | Mandatory expired tidal volume: 0 - 4 litres
19.1.5 | Spontaneous expired tidal volume
19.1.6 | Peak Inspiratory Flow Rate (PIFR): 0 to 300 litres/min
19.1.7 | Peak Expiratory Flow Rate (PEFR): 0 to 300 litres/min
19.2 | Pressure monitoring:
19.2.1 | Peak inspiratory pressure: 0 - 120 cm H2O
19.2.2 | Mean airway pressure: 0 - 120 cm H2O
19.2.3 | Plateau pressure: 0 - 120 cm H2O
19.2.4 | PEEP: 0 - 50 cm H2O
19.2.5 | Oxygen inlet pressure: 40 - 95 psig
19.2.6 | Delivered O2: 21 - 100 %
19.3 | Time cycle monitoring:
19.3.1 | Breath rate: 0 - 150 bpm
19.3.2 | Spontaneous breath rate: 0 - 150 bpm
19.3.3 | Inspiratory time: 0 - 99 sec
19.3.4 | Exhalation time: 0 - 99 sec
19.3.5 | I:E ratio: 1:99 - 99:1
19.3.6 | Rapid Mustow breathing index 0 - 500 bpm/litre
20 | The ventilator must have a colour LCD screen with at least three simultaneous real-time waveforms
21 | The following parameters must be available as time-base curves:
21.1 | Pressure
21.2 | Flow
21.3 | Tidal volume
22 | Display of the following patient parameters must be user selectable and at least three selected parameters must be displayed as digital values
22.1 | Spontaneous minute volume
22.2 | Inspiratory flow
22.3 | Peak airway pressure
22.4 | Mean airway pressure
22.5 | Plateau pressure
22.6 | PEEP
22.7 | Expired tidal volume
22.8 | Expired minute volume
22.9 | Inspiratory time
22.10 | Breath rate
22.11 | Spontaneous rate
22.12 | Internal oxygen inlet pressure (pipeline inlet pressure)
22.13 | Battery charge status
22.14 | I:E ratio
22.15 | Set I:E ratio
22.16 | Static airways compliance in ml/cmH2O
22.17 | Static airways compliance per kg in ml/cmH2O
22.18 | Dynamic airways compliance in ml/cmH2O
22.19 | Dynamic airways compliance per kg in ml/cmH2O
22.20 | Peak inspired flow rate in l/m
22.21 | Peak expired flow rate in l/m
22.22 | Inspired O2 in %
22.23 | Tube leak in %
23 | The ventilator must measure and display inspiratory/expiratory loops.
24 | Waveforms and loops must be colour coded to denote inspiratory and expiratory breath phases
25 | The graphics display system must have a freeze capability.
26 | Loops: Flow/Volume Loop, Pressure/Volume Loop
27 | Reference loops must be able to be stored and displayed in the background in different colours
28 | All primary controls must be visible and changeable from all screens without entering a menu
29 | Monitored values can be arranged to match sequence of hospital's ventilator check sheet
30 | SETUP
30.1 | Patient details
30.2 | Patient size: Adult / Paediatric
30.3 | Patient weight
31 | TRENDS
31.1 | Trends: at least 8 parameters for at least 24 hours
31.2 | Trends must be configurable to display at least 4 values as histograms with ability to review with cursor/time stamp
32 | Automatically recorded events:
32.1 | New patient
32.2 | Settings Change
32.3 | Increase 02
32.4 | Power On
32.5 | Insp Hold
32.6 | Power Off
32.7 | Exp Hold
32.8 | Power Recovery
32.9 | Nebulizer
32.10 | Enter Standby
32.11 | Manual Breath
32.12 | Exit Standby
33 | ALARMS: The ventilator must have the following audible and visual alarms:
33.1 | Low/high airway pressure
33.2 | Low/high minute volume
33.3 | Low/high tidal volume
33.4 | Low/high breath rate:1 to 200 bpm
33.5 | Low/high FiO2
33.6 | Loss of PEEP
33.7 | Low/high PEEP
33.8 | Low/high PIP
33.9 | Leakage alarm
33.10 | Apnea: 10 - 60 sec
33.11 | Patient disconnect
33.12 | Oxygen concentration variation (± 5%)
33.13 | Air/oxygen % deviation
33.14 | Power failure
33.15 | Air and oxygen supply failure
33.16 | Air and oxygen low pressure
33.17 | Battery low (When approximately 15 min operation left).
33.18 | Apnea
34 | Alarm limits must be user-adjustable
35 | Default alarm limits if user does not set alarm limits
36 | Alarm tone differentiates high, medium and low priority alarms
37 | Adjustable alarm volume
38 | Temporary mute for audible alarm: 0 - 2 min
39 | Alarm Reset. Deactivates visual indicators for resolved alarm conditions
40 | Machine Alarms:
40.1 | Safety Valve Open
40.2 | Ventilator Inoperative
40.4 | Fan Failure
40.5 | Loss of Air -Air Supply pressure < 18.0 psig
40.6 | Loss of O2 – O2 Supply pressure <18.0 psig (when % O2 > 21)
40.7 | Loss of Heliox -Heliox Supply pressure
41 | GAS SUPPLY
41.1 | Oxygen pipeline supply: Pressure: 200 to 600 kPa) Minimum flow: 80 L/min at 1.5 bar
41.2 | Air pipeline supply: Pressure: 200 to 600 kPa) Minimum flow: 80 L/min at 1.5 bar
41.3 | Compressor air: 0.2 to 0.7 bar
41.4 | Medical air and oxygen hoses, 3m long, colour coded, mounted with Heyer oxygen and medical air probes. The Heyer probes must be compatible with the wall points used in the instutution.
42 | POWER SUPPLY
42.1 | The unit must be powered by 220V, 50 Hz mains supply
42.2 | Internal rechargeable battery to operate the unit for approximately five hours
42.3 | Automatic switch to battery in case of power failure,
42.4 | Automatic battery recharge when connected to mains
42.5 | Power cord at least 3 m long with SABS approved 15A three prong plug
43 | GENERAL
43.1 | The ventilator must be mounted on a trolley with at least 4 antistatic swivel castors of at least 125mm diameter of which 2 must be lockable.
43.2 | Trolley must have a push-handle
43.3 | The trolley must have an oxygen cylinder holder
43.4 | The trolley must have a rail on both sides to fit the circuit arm
43.5 | Removeable arm with circuit hanger to fit the side rails
43.6 | Expiratory inlet to include integral water trap, valved to allow emptying without interruption of normal patient
43.7 | Air filter capacity at inlet: 99 % (for > 0.5 um)
43.8 | State lifetime and price of Oxygen sensor
43.9 | The unit must be handed over in full operating order
44 | The following accessories must be supplied with the unit:
44.1 | Temperature probes for humidifier
44.2 | Two reusable patient breathing circuits
44.3 | Built-in nebulizer
44.4 | Reusable, sterilizable patient masks
44.5 | Reusable, sterilizable patient connectors
44.6 | Reusable flow sensor.
44.7 | Internal O2-Air mixer.
44.8 | FiO2 analyzer, internal.
44.9 | Test lung and test circuit
45 | A starter-pack of consumables must be supplied with the unit comprising at least the following:
45.1 | Bacteria filter (pk of 10)
45.2 | Adult disposable patient circuit (Pk of 10)
45.3 | Paeds disposable patient circuits (Pk of 10)
45.4 | Air filters (pk of 10)
45.5 | Flow sensor (pk of 5)
45.6 | List ALL available consumables with prices incl VAT
45.7 | List ALL available accessories with prices incl VAT
46 | Optional Items
46.1 | MANEUVRES
46.1.1 | Ventilator must provide separate maneuvres for MIP/NIF P100, AutoPEEP airway, Slow Flow Loop Inflection Points
46.1.2 | Ventilator must return to normal ventilation at end of maneuvre without any changes required to primary controls
46.1.3 | Ventilator must have "Sensitivity" setting for maneuvre only, separate from primary "Sensitivity" setting. Allowing Patient to abort maneuvre.
46.1.4 | Inflection point maneuvre must calculate upper and lower inflection points and display on Pressure/Volume loop
46.1.5 | Stored inflection point loops must be able to be superimposed on current loop for comparison
46.2 | Humidifier
46.2.1 | The humidifier must be dual servo controlled with an alarm system
46.2.2 | All accessories for the humidifier must be included
46.2.3 | Humidifier must have a hose heating element, dual temperature probe and electrical adapter.
46.3 | Heliox
46.3.1 | Heliox gas mixtures in all modes with automatic corrections for all volumes and FiO2
46.3.2 | Heliox Supply: Pressure range: 20 to 80 psig (1.4 to 5.5 bar)
46.3.3 | Ventilator to deliver either Air/Oxygen mixture or Heliox (internally blended to 80:20 oxygen: helium ratio)
46.3.4 | High pressure medical air connection must be interchangeable by the clinician for Helium with appropriate CGA gas fitting
46.4 | End Tidal CO2 curve
46.5 | Printer
47 | MAINTENANCE CONTRACT
47.1 | Bidder must offer a five year full maintenance contract effective after the two year guarantee expires.
47.2 | The cost of the maintenance contract must not be included in the price of the bid
47.3 | State the cost per year and the total cost
47.4 | Provide details of the contract in a separate document