## Bid Title
ANNEXURE E
TECHNICAL SPECIFICATION
TRANSVERSAL CONTRACT RT2-2021
THE SUPPLY, DELIVERY, INSTALLATION, COMMISSIONING AND MAINTENANCE OF MEDICAL EQUIPMENT TO THE STATE FOR THE PERIOD 1 DECEMBER 2021 TO 30 NOVEMBER 2024
NATIONAL TREASURY
TRANSVERSAL CONTRACTING

## Caterogy-1 Anaesthesia
ANAESTHESIA EQUIPMENT
# | Item number | Item Description
1 | RT2-01-01 | Anaesthetic Machine, Basic
2 | RT2-01-02 | Anaesthetic Machine, Basic with Monitor
3 | RT2-01-03 | Anaesthetic Machine, Standard
4 | RT2-01-04 | Anaesthetic Machine, Standard with Monitor
5 | RT2-01-05 | Anaesthetic Machine, Advanced
6 | RT2-01-06 | Anaesthetic Machine, Advanced with Monitor
7 | RT2-01-07 | Monitor, Anaesthetic
8 | RT2-01-08 | Anaesthetic Machine, MRI Compatible
9 | RT2-01-09 | Anaesthetic Machine, MRI Compatible with Monitor
10 | RT2-01-10 | Monitor, MRI Compatible
11 | RT2-01-11 | Video Laryngoscope
12 | RT2-01-12 | Video Laryngoscope with Flexible Bronchoscope
13 | RT2-01-13 | Video Disposable Flexible Bronchoscope
15 | RT2-01-15 | Syringe Driver, TCI
16 | RT2-01-16 | Syringe Driver, PCA
17 | RT2-01-17 | Blood-Fluid Warmer
18 | RT2-01-18 | Patient Warmer, Forced-Air
20 | RT2-01-20 | Nerve Stimulator
23 | RT2-01-23 | Ultrasound Machine, Point-of-care

## Anaesthetic Machine, Basic
RT2-01-01 | ANAESTHESIA DELIVERY SYSTEM: BASIC WITHOUT PHYSIOLOGICAL MONITOR
1 | Description
1.1
2 | Product Offered (Anaesthetic Machine)
2.1 | Make:
2.2 | Model:
2.3 | Initial Manufacturing year of Model:
2.4 | Radiation Control License number:
3 | Technical Specification
3.1 | The complete anaesthesia workstation shall have (but not be limited to) the following aspects, features and components, as described in the specifications:
3.2 | Anaesthesia delivery system including 2x vaporisers
3.3 | Ventilator suitable for ventilating paediatric to adult patients
3.4 | The design of the unit shall comply with all relevant International Standards for Anaesthesia systems. Offered units shall carry the marks for these standards. Bidders to indicate to which standards the offered machine complies
3.5 | The unit shall be easily moved around without the possibility of toppling over
3.6 | The frame of the unit shall be manufactured from corrosion free material, yet be strong and durable
3.7 | The unit shall be provided with at least four die-cast aluminium, or better material, antistatic castors. The castors shall be equipped with a braking system
3.8 | The unit shall be supplied with a swivel monitoring shelf to mount the physiological monitor. The monitor (described later in the specifications) shall be mounted on the frame of the machine with a dedicated monitor mounting system
3.9 | The unit shall have at least two drawers for secure storage of airway management equipment and other critical requirements
3.10 | The unit shall be supplied with at least three (3) extra 220 Volt 15A power sockets for additional monitoring/syringe pumps
3.11 | The unit shall have a working area, as well as a pull-out or flip-up writing table for writing notes
3.12 | The unit shall preferably be equipped with a working light to enable the anaesthetist to work in a darkened theatre
3.13 | The unit shall incorporate a suction bottle mounting rail (Gabler) on the side of the machine, with a mounted suction control
3.14 | All critical controls shall be within reach of an anaesthetist sitting in front of the machine
3.15 | A anaesthetic machine master switch is essential for start-up of both pneumatic and electrical supplies. Both must switch on simultaneously with a single action
3.16 | The unit shall have a routine electronic checklist for diagnostic testing of flow delivery and preferably circuit integrity. It shall be possible to bypass this in the event of urgent machine requirement
3.17 | Pipeline gas inlets shall be non-interchangeable. Hoses shall be colour-coded, and shall be connectable to the medical service outlets in ……… Hospital. It shall be possible to connect the machine to the outlets for Nitrous Oxide, Oxygen, Vacuum and Medical Air (Standard colours according National Standards)
3.18 | At least one cylinder regulator shall be provided for Oxygen. This shall be pin-indexed as per ISO standard. The trolley shall have an attachment to support the cylinder
3.19 | Gas regulators shall preferably have metal diaphragms. Bidders to state the output pressure and burst pressure settings of the regulators
3.20 | A safety pressure relief valve shall protect the entire pneumatic system from excessive input pressure from pipeline and cylinder supply. Bidders to state the pressure relief setting on this system
3.21 | Back flow of gases into pipeline or cylinders shall be prevented by a non-return valve system. With a cylinder detached from the pin-indexed yoke and the pipeline hoses connected, no gas leak shall occur at the yoke. Preferential flow shall be facilitated from the pipeline inlets. Bidders to describe how this is achieved
3.22 | Gauges must be supplied for each individual gas. Pipeline and cylinder gauges shall be differentially positioned, and clearly indicated for purpose
3.23 | The anaesthetic machine or ventilator shall have a built-in manometer gauge, calibrated between the range of -20 to +100cm water. A digital readout is acceptable
3.24 | All internal pneumatic circuitry shall be contained within the enclosed work surface, to avoid pulling and accidental disconnection of gas pipes. All gas tubing shall be diameter indexed with non-interchangeable connections
3.25 | The unit shall be equipped to administer low flow anaesthesia. For this reason, the unit shall be equipped with rotameters that can display(show) the flow rates for high and low flow anaesthesia, typically by using cascading rotameters in the following range:
3.25.1 | Oxygen low flow tube 200ml/minute – 1 litre/min At higher flow rates the gas must cascade into a second flow tube (1 – 10 l/min)
3.25.2 | Nitrous Oxide low flow tube 200ml/minute – 1 litre/min At higher flow rates the gas must cascade into a second flow tube (1 – 10 l/min)
3.25.3 | Air tube 200ml – 10litres/min Cascading tubes as above are acceptable, but not essential
3.26 | Flowmeter tubes shall preferably be diameter indexed to prevent transportation during service and maintenance. Secondary regulators shall be fitted between incoming pipeline gas supply and the flowmeter control valves to ensure stability of the bobbin at low end of scale (This applies if flowmeters are offered, this point does not apply if electronic flowmeters are offered.)
3.27 | Machines supplied with electronic flowmeters shall be provided with a manual flow meter for oxygen delivery independent of electronic control
3.28 | Flowmeter sequence to be in accordance with South African Society of Anaesthesiologists (SASA) Practice Guidelines – (Oxygen on the right)
3.29 | The unit shall include a safety mechanism that ensures a minimum concentration of 25% oxygen in a oxygen/nitrous oxide gas mixture across all flow rates above 1 litre. It is desired that at flow rates of below 1 litre, the Oxygen flow rate cannot be decreased below 250ml/min
3.30 | The Anaesthetic machine shall be equipped with an oxygen monitor – Even if the monitor has an oxygen monitoring capability, it shall also be built into the machine
3.31 | The oxygen monitor shall have a low and high alarm – visual and audible. The low alarm shall automatically default to no less than 18%.
3.32 | It shall be possible to calibrate the low oxygen control to 21%.
3.33 | The unit shall include a high-flow Oxygen flush mechanism. This system shall bypass the rotameters and vaporisers, and shall be non-lockable. The activation mechanism shall be a recessed, self closing push button, and shall provide an Oxygen flush in excess of 30 litres/min
3.34 | The unit shall incorporate a continuous alarm should the Oxygen pressure fall below 193 KPa. Should pressure further fall below 138 KPa, then all Nitrous Oxide shall be automatically cut out from the patient circuit
3.35 | The unit shall have a vaporiser manifold which is of the lock-on type, allowing the vaporiser to be removed from the machine, and replaced with an alternative. No tools shall be required to perform this action
3.36 | The unit shall have 2 slots/stations for at least 2 vaporisers, to be mounted side by side
3.37 | The machine/vaporiser shall automatically prevent more than one vaporiser from being turned on simultaneously. This must be achieved by an interlock mechanism, and not by a separate vaporiser selector switch
3.38 | Two vaporisers shall be supplied with each machine
3.38.1 | 1 X Isoflurane
3.38.2 | 1 X Sevoflurane
3.39 | Preference will be given to vaporisers that has a long service interval period. Bidder to fully describe the service requirements, frequency and cost per service per vaporiser
3.40 | Each vaporiser shall be colour-coded, and provided with a safety device to avoid accidental filling with incorrect product according to International Standards
3.41 | The unit shall have an open-end gas scavenging reservoir fitted. Active scavenging from the reservoir shall be approximately 30 litres/minute. Scavenging outlet shall be compatible with the scavenging system being installed at the users's Hospital
3.42 | The unit shall be equipped with a carbon dioxide circle absorber with APL valve with antipollution hood and ventilator to manual switch. This switch when directed to the ventilator shall isolate the APL valve from the circle. All patient connectors shall conform to ISO standards (male with outer diameter of 22mm and inner diameter of 15mm)
3.43 | Two non-disposable canisters shall be included in the offer. The complete system shall be autoclaveable to 134º.
3.44 | It is essential that the soda lime can be isolated while the circle is in use
3.45 | The absorber system shall be either integrated into the frame or be positioned under the shelf for transport purposes
3.46 | All materials in contact with the exhaled patient gases shall be autoclavable, (excluding flow sensors and Oxygen cell)
3.47 | A water trap to accommodate excess water generated is essential. This shall preferably be integrated into the soda lime canister
3.48 | The frame shall be equipped with a power back-up system to enable full operation of the anaesthetic machine and ventilator for at least 60 minutes with batteries fully charged, under typical operating conditions
3.49 | VENTILATOR
3.49.1 | The ventilator shall allow a “closed-loop” monitoring configuration to ensure that the selected VT or pressure is achieved
3.49.2 | Any change to fresh gas flow shall be compensated for
3.49.3 | Any flow sensors used shall be robust and easily removed for cleaning
3.49.4 | Flow sensors shall be remote from patient, or integrated within the system
3.49.5 | Bidders to clearly state the system used, as well as cost and frequency of replacement required for the flow sensors
3.49.6 | The ventilator shall be able to ventilate patients in Volume Controlled and Pressure Controlled mode
3.49.7 | It shall be possible to ventilate patients with a Tidal Volume from 30ml to 1 500ml
3.49.8 | The following (approximate) ventilator parameter ranges shall be applicable:
3.49.8.1 | Respiration Rate: 4 to 80 breaths/min
3.49.8.2 | Minute Volume: 0 to 60 l/min
3.49.8.3 | Tidal Volume: 30 to 1 500ml. (State different values for Volume and Pressure Controlled modes)
3.49.8.4 | Be able to provide PEEP in the range: 0 to 25cm Water
3.49.9 | Bidders to state all monitoring capabilities offered as part of the ventilator
3.49.10 | The following alarms shall form part of the capabilities of the offered ventilator:
3.49.10.1 | Tidal Volume
3.49.10.2 | Minute volume
3.49.10.3 | Inspired Oxygen (Fi02)
3.49.10.4 | Apnoea
3.49.10.5 | Low airway Pressure
3.49.10.6 | High Airway pressure
3.49.10.7 | Bidders to clearly indicate that these alarms are available, as well as the ranges and settings possible
3.49.11 | The ventilator shall have a mute alarm function, but it shall not be possible to cancel the alarm until the problem has been rectified. Mute function shall only be for a limited time period
3.49.12 | The ventilator shall have power back-up to allow full function of the ventilator and anaesthetic machine for not less than 60 minutes
4 | A starter-pack of consumables must be supplied with the unit comprising at least the following:
4.1 | Anaesthesia circuit (Adult, single/double limb design) Disposable. 1.8m (Complete with Y-piece, Elbow, Catheter mount, Rebreathing bag)
4.2 | Anaesthesia circuit (Paediatric, single/double limb design) Disposable. 1.8m (Complete with Y-piece, Elbow, Catheter mount, Rebreathing bag)
4.3 | Anaesthesia circuit (neonatal, single/double limb design) Disposable. 1.8m (Complete with Y-piece, Elbow, Catheter mount, Rebreathing bag)
4.4 | Re-breathing circuit complete
4.5 | Non-re-breathing circuit complete
4.6 | Flow sensor (pack of 2)
4.7 | Spirometry line
4.8 | Sample line (pack of 10)
4.9 | Face masks: sizes 00, 0, 1, 2, 3, 4, 5.
4.10 | Bacterial/viral filter
4.11 | Y connector
4.12 | Manual ventilation hose
4.13 | Fresh gas tubing
4.14 | Seal for back-up cylinder connection on the unit
4.15 | Scavenger tubing (as per institution's scavenging outlet)
4.16 | Inspiratory and expiratory unidirectional valves
4.17 | Side stream CO2 kit (adult/paed)
4.18 | Disposable Adult Spirometry Set complete with spirometry tube, sample line and flow sensor, 3 m, box of 20 pieces
4.19 | Disposable Paediatric Spirometry Set complete with spirometry tube, sample line and flow sensor, 3 m, box of 20 pieces
4.20 | Gas sample tube with luer lock connection, 3m, disposable (10 pcs)
4.21 | Suction Tubing Set inc connector
4.22 | Resuscitator Adult, 1500mL, autoclavable, silicone
4.23 | Resuscitator Paediatric, 550mL, autoclavable, silicone
4.24 | Tubing set Paediatric for circle systems
5 | Conformity Compliance (Please attach certificate)
5.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
5.2 | The unit must comply with electrical safety and performance standard such as IEC 60601-2-13 for Anaesthetic systems, attached certification
5.2.1 | Please provide unique ref number of the IEC ceriticate:
5.3 | OEM must comply and certified on ISO 9001 quality standards, attach proof of compliance
5.3.1 | Please provide unique ref number of the ISO 9001 ceriticate:
5.4 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
5.4.1 | Please provide unique ref number of the ISO 13485 ceriticate:
5.5 | Model quoted for must be EC certified. Attach a copy of certification
5.5.1 | Please provide unique ref number of the EC ceriticate:
5.6 | Hazardous Substance Act: If this type of equipment / apparatus appears on the schedule of Hazardous substances issued by the Directorate: Health Technology of the Department of Health, a license in terms of the Act on Hazardous Substances (Act 15/1973) must be submitted with the bid / quotation documents. The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.7 | All electrical/electronic medical equipment must be licensed by Radiation Control, where applicable a copy of the license must be submitted The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.7.1 | Radiation Control License number:
5.8 | The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade (rubber to with 2 screws) type and it must be a minimum length of three (3) meters. NB the mains cable of the unit tendered for must be SABS colour coded.
5.8 | The equipment quoted for must be protected against electromagnetic interference
6 | COMPREHENSIVE MAINTENANCE
6.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) should be included in the bid price as per the SCC document
6.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries, Flow sensors and Oxygen sensor), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
6.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
6.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries, Flow sensors and Oxygen sensor), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
6.3 | The bidder to provide the OEM recommended service interval of the unit:
6.4 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
7 | General
7.1 | Has the product on offer or any of its components been recalled by the Regulatory Authority or OEM initiated recall in the last 5 years? If yes, please provide details of the type of recall and how was it addressed
7.2 | The unit on offer must be the latest technology.
ANAESTHESIA DELIVERY SYSTEM: BASIC WITHOUT PHYSIOLOGICAL MONITOR, REPLACEMENT ACCESSORIES
RT2-01-01a | Anaesthesia circuit (Adult, single/double limb design) Disposable. 1.8m (Complete with Y-piece, Elbow, Catheter mount, Rebreathing bag)
RT2-01-01b | Anaesthesia circuit (Paediatric, single/double limb design) Disposable. 1.8m (Complete with Y-piece, Elbow, Catheter mount, Rebreathing bag)
RT2-01-01c | Anaesthesia circuit (neonatal, single/double limb design) Disposable. 1.8m (Complete with Y-piece, Elbow, Catheter mount, Rebreathing bag)
RT2-01-01d | Re-breathing circuit complete
RT2-01-01e | Non-re-breathing circuit complete
RT2-01-01f | Flow sensor (pack of 2)
RT2-01-01g | Spirometry line
RT2-01-01h | Sample line (pack of 10)
RT2-01-01i | Face masks: sizes (00,0,1,2,3,4, 5)- End-user to specify size
RT2-01-01j | Bacterial/viral filter
RT2-01-01k | Y connector
RT2-01-01l | Manual ventilation hose
RT2-01-01m | Fresh gas tubing
RT2-01-01n | Seal for back-up cylinder connection on the unit
RT2-01-01o | Scavenger tubing (as per institution's scavenging outlet)
RT2-01-01p | Inspiratory and expiratory unidirectional valves
RT2-01-01q | Side stream CO2 kit (adult/paed)
RT2-01-01r | Disposable Adult Spirometry Set complete with spirometry tube, sample line and flow sensor, 3 m, box of 20 pieces
RT2-01-01s | Disposable Paediatric Spirometry Set complete with spirometry tube, sample line and flow sensor, 3 m, box of 20 pieces
RT2-01-01t | Gas sample tube with luer lock connection, 3m, disposable (10 pcs)
RT2-01-01u | Suction Tubing Set inc connector
RT2-01-01v | Resuscitator Adult, 1500mL, autoclavable, silicone
RT2-01-01w | Resuscitator Paediatric, 550mL, autoclavable, silicone
RT2-01-01x | Tubing set Paediatric for circle systems
RT2-01-01y | sevoflurane vaporizer
RT2-01-01z | Isoflurane vaporizer
RT2-01-01zz | Desflurane vaporizer
Bidder to list any other Consumables or Accessories (Not listed above) for the model on offer below

## Anaesthetic, Basic + monitor
RT2-01-02 | ANAESTHESIA DELIVERY SYSTEM: Basic with integrated physiological monitor
1 | Product Offered (Anaesthetic Machine)
1.1 | Make:
1.2 | Model:
1.3 | Initial Manufacturing year of Model:
1.4 | Radiation Control License number:
2 | Product Offered (Monitor)
2.1 | Make:
2.2 | Model:
2.3 | Initial Manufacturing year of Model:
2.4 | Radiation Control License number:
3 | Technical Specification
3.1 | The complete anaesthesia workstation shall have (but not be limited to) the following aspects, features and components, as described in the specifications:
3.2 | Anaesthesia delivery system including vaporisers
3.3 | Physiological monitor with multigas analyser and anaesthetic agent monitoring – This can be either built into the machine, or be a function on the physiological monitor
3.4 | Ventilator suitable for ventilating paediatric to adult patients
3.5 | The design of the unit shall comply with all relevant International Standards for Anaesthesia systems. Offered units shall carry the marks for these standards. Bidders to indicate to which standards the offered machine complies
3.6 | The unit shall be easily moved around without the possibility of toppling over
3.7 | The frame of the unit shall be manufactured from corrosion free material, yet be strong and durable
3.8 | The unit shall be provided with at least four die-cast aluminium, or better material, antistatic castors. The castors shall be equipped with a braking system
3.9 | The unit shall be supplied with a swivel monitoring shelf to mount the physiological monitor. The monitor (described later in the specifications) shall be mounted on the frame of the machine with a dedicated monitor mounting system
3.10 | The unit shall have at least two drawers for secure storage of airway management equipment and other critical requirements
3.11 | The unit shall be supplied with at least three (3) extra 220 Volt 15A power sockets for additional monitoring/syringe pumps
3.12 | The unit shall have a working area, as well as a pull-out or flip-up writing table for writing notes
3.13 | The unit shall preferably be equipped with a working light to enable the anaesthetist to work in a darkened theatre
3.14 | The unit shall incorporate a suction bottle mounting rail (Gabler) on the side of the machine, with a mounted suction control
3.15 | All critical controls shall be within reach of an anaesthetist sitting in front of the machine
3.16 | The anaesthetic machine master switch is essential for start-up of both pneumatic and electrical supplies. Both must switch on simultaneously with a single action
3.17 | The unit shall have a routine electronic checklist for diagnostic testing of flow delivery and preferably circuit integrity. It shall be possible to bypass this in the event of urgent machine requirement
3.18 | Pipeline gas inlets shall be non-interchangeable. Hoses shall be colour-coded, and shall be connectable to the medical service outlets in ……… Hospital. It shall be possible to connect the machine to the outlets for Nitrous Oxide, Oxygen, Vacuum and Medical Air
3.19 | At least one cylinder regulator shall be provided for Oxygen. This shall be pin-indexed as per ISO standard. The trolley shall have an attachment to support the cylinder
3.20 | Gas regulators shall preferably have metal diaphragms. Bidders to state the output pressure and burst pressure settings of the regulators
3.21 | A safety pressure relief valve shall protect the entire pneumatic system from excessive input pressure from pipeline and cylinder supply. Bidders to state the pressure relief setting on this system
3.22 | Back flow of gases into pipeline or cylinders shall be prevented by a non-return valve system. With a cylinder detached from the pin-indexed yoke and the pipeline hoses connected, no gas leak shall occur at the yoke. Preferential flow shall be facilitated from the pipeline inlets. Bidders to describe how this is achieved
3.23 | Gauges must be supplied for each individual gas. Pipeline and cylinder gauges shall be differentially positioned, and clearly indicated for purpose
3.24 | The anaesthetic machine or ventilator shall have a built-in manometer gauge, calibrated between the range of -20 to +100cm water. A digital readout is acceptable
3.25 | All internal pneumatic circuitry shall be contained within the enclosed work surface, to avoid pulling and accidental disconnection of gas pipes. All gas tubing shall be diameter indexed with non-interchangeable connections
3.26 | The unit shall be equipped to administer low flow anaesthesia. For this reason, the unit shall be equipped with rotameters that can display(show) the flow rates for high and low flow anaesthesia, typically by using cascading rotameters in the following range:
3.26.1 | Oxygen low flow tube 200ml/minute – 1 litre/min At higher flow rates the gas must cascade into a second flow tube (1 – 10 l/min)
3.26.2 | Nitrous Oxide low flow tube 200ml/minute – 1 litre/min At higher flow rates the gas must cascade into a second flow tube (1 – 10 l/min)
3.26.3 | Air tube 200ml – 10litres/min Cascading tubes as above are acceptable, but not essential
3.27 | Flowmeter tubes shall preferably be diameter indexed to prevent transportation during service and maintenance. Secondary regulators shall be fitted between incoming pipeline gas supply and the flowmeter control valves to ensure stability of the bobbin at low end of scale (This applies if flowmeters are offered, this point does not apply if electronic flowmeters are offered.)
3.28 | Machines supplied with electronic flowmeters shall be provided with a manual flow meter for oxygen delivery independent of electronic control
3.29 | Flowmeter sequence to be in accordance with South African Society of Anaesthesiologists (SASA) Practice Guidelines – (Oxygen on the right)
3.30 | The unit shall include a safety mechanism that ensures a minimum concentration of 25% oxygen in a oxygen/nitrous oxide gas mixture across all flow rates above 1 litre. It is desired that at flow rates of below 1 litre, the Oxygen flow rate cannot be decreased below 250ml/min
3.31 | The Anaesthetic machine shall be equipped with an oxygen monitor – Even if the monitor has an oxygen monitoring capability, it shall also be built into the machine
3.32 | The oxygen monitor shall have a low and high alarm – visual and audible. The low alarm shall automatically default to no less than 18%.
3.33 | It shall be possible to calibrate the low oxygen control to 21%.
3.34 | The unit shall include a high-flow Oxygen flush mechanism. This system shall bypass the rotameters and vaporisers, and shall be non-lockable. The activation mechanism shall be a recessed, self closing push button, and shall provide an Oxygen flush in excess of 30 litres/min
3.35 | The unit shall incorporate a continuous alarm should the Oxygen pressure fall below 193 KPa. Should pressure further fall below 138 KPa, then all Nitrous Oxide shall be automatically cut out from the patient circuit
3.36 | The unit shall have a vaporiser manifold which is of the lock-on type, allowing the vaporiser to be removed from the machine, and replaced with an alternative. No tools shall be required to perform this action
3.37 | The unit shall have 2 slots/stations for at least 2 vaporisers, to be mounted side by side
3.38 | The machine/vaporiser shall automatically prevent more than one vaporiser from being turned on simultaneously. This must be achieved by an interlock mechanism, and not by a separate vaporiser selector switch
3.39 | Two vaporisers shall be supplied with each machine
3.39.1 | 1 X Isoflurane
3.39.2 | 1 X Sevoflurane
3.40 | Preference will be given to vaporisers that has a long service interval period. Bidder to fully describe the service requirements, frequency and cost per service per vaporiser
3.41 | Each vaporiser shall be colour-coded, and provided with a safety device to avoid accidental filling with incorrect product according to International Standards
3.42 | The unit shall have an open-end gas scavenging reservoir fitted. Active scavenging from the reservoir shall be approximately 30 litres/minute. Scavenging outlet shall be compatible with the scavenging system being installed at the users's Hospital
3.43 | The unit shall be equipped with a carbon dioxide circle absorber with APL valve with antipollution hood and ventilator to manual switch. This switch when directed to the ventilator shall isolate the APL valve from the circle. All patient connectors shall conform to ISO standards (male with outer diameter of 22mm and inner diameter of 15mm)
3.44 | Two non-disposable canisters shall be included in the offer. The complete system shall be autoclaveable to 134º.
3.45 | It is essential that the soda lime can be isolated while the circle is in use
3.46 | The absorber system shall be either integrated into the frame or be positioned under the shelf for transport purposes
3.47 | All materials in contact with the exhaled patient gases shall be autoclavable, (excluding flow sensors and Oxygen cell)
3.48 | A water trap to accommodate excess water generated is essential. This shall preferably be integrated into the soda lime canister
3.49 | The frame shall be equipped with a power back-up system to enable full operation of the anaesthetic machine and ventilator for at least 60 minutes with batteries fully charged, under typical operating conditions
3.50 | VENTILATOR
3.50.1 | The ventilator shall allow a “closed-loop” monitoring configuration to ensure that the selected VT or pressure is achieved
3.50.2 | Any change to fresh gas flow shall be compensated for
3.50.3 | Any flow sensors used shall be robust and easily removed for cleaning
3.50.4 | Flow sensors shall be remote from patient, or integrated within the system
3.50.5 | Bidders to clearly state the system used, as well as cost and frequency of replacement required for the flow sensors
3.50.6 | The ventilator shall be able to ventilate patients in Volume Controlled and Pressure Controlled mode
3.50.7 | It shall be possible to ventilate patients with a Tidal Volume from 30ml to 1 500ml
3.50.8 | The following (approximate) ventilator parameter ranges shall be applicable:
3.50.8.1 | Respiration Rate: 4 to 80 breaths/min
3.50.8.2 | Minute Volume: 0 to 60 l/min
3.50.8.3 | Tidal Volume: 30 to 1 500ml. (State different values for Volume and Pressure Controlled modes)
3.50.8.4 | Be able to provide PEEP in the range: 0 to 25cm Water
3.50.9 | Bidders to state all monitoring capabilities offered as part of the ventilator
3.50.10 | The following alarms shall form part of the capabilities of the offered ventilator:
3.50.10.1 | Tidal Volume
3.50.10.2 | Minute volume
3.50.10.3 | Inspired Oxygen (Fi02)
3.50.10.4 | Apnea
3.50.10.5 | Low airway Pressure
3.50.10.6 | High Airway pressure
3.50.10.7 | Bidders to clearly indicate that these alarms are available, as well as the ranges and settings possible
3.50.11 | The ventilator shall have a mute alarm function, but it shall not be possible to cancel the alarm until the problem has been rectified. Mute function shall only be for a limited time period
3.50.12 | The ventilator shall have power back-up to allow full function of the ventilator and anaesthetic machine for not less than 60 minutes
3.51 | MONITOR
3.51.1 | The anaesthetic machine shall be equipped with a monitor specifically designed for use in operating theatres
3.51.2 | The monitor functions can be incorporated in the design of the anaesthetic machine, or can be a separate monitor mounted on the anaesthetic machine on a suitable swivel arm
3.51.3 | The monitor shall be able to monitor and display the following parameters:
3.51.3.1 | ECG
3.51.3.2 | Respiration
3.51.3.3 | Pulse Oximetry (Oxygen Saturation)
3.51.3.4 | Non Invasive Blood Pressure
3.51.3.5 | Side stream capnography
3.51.3.6 | Automatic Agent Monitoring
3.51.3.7 | 2 channels of Invasive Blood Pressure
3.51.3.8 | Temperature
3.51.4 | The monitor shall be of modern design based on ergonomic principles
3.51.5 | The front panel shall be designed to facilitate easy cleaning and disinfecting
3.51.6 | Preference will be given to units incorporating touch screen as well as “navigation point” controls for easy information input and onscreen navigation
3.51.7 | The display of the monitor shall have a display area of not less than 26 cm in size (diagonally measured: from corner to corner)
3.51.8 | The display shall be flicker free and easy to read. Bidders to state if brightness control is possible
3.51.9 | Preference will be given to units with a colour display, and different parameters displayed in different colours
3.51.10 | In addition to the waveform display section, the monitors shall have an area dedicated for alphanumeric display of parameter identification and measured values, alarm and status messages
3.51.11 | Bidders to state if the screen can be configured by the user to display parameters in different set-ups
3.51.12 | Bidders to state if trending is possible on the monitor, and state trending options
3.51.13 | Audible and visual alarms shall be available for all measured parameters. State the method employed to avoid “nuisance alarms”
3.51.14 | Bidders to state if alarms are graded into levels of urgency, and explain the method and grading
3.51.15 | The monitor shall provide clinically reasonable default alarms when switched on. These alarms shall be programmable, and shall be capable of being overridden manually
3.51.16 | It shall be possible to override the audible alarm for a short period, but no override of the visual alarm shall be possible
3.51.17 | It shall be possible to suspend the alarms of certain parameters during clinical procedures, but the alarm shall reactivate automatically after a suitable time delay. Any silencing or disabling of alarms shall be clearly indicated on the monitor
3.51.18 | Input impedance of the system shall be of such a nature as to protect the patient at all times
3.51.19 | Maximum allowable leakage current to the patient shall be less than 10 µA
3.51.20 | The monitor shall be protected against defibrillator and electro surgical interference
3.51.21 | The monitor shall be of low power consumption, and shall preferably not require fan or other forced ventilation
3.51.22 | All cables and sockets shall be colour coded for easy recognition, and shall conform to AAMI quality standards or equivalents
3.51.23 | The monitor shall be equipped with a battery to keep the monitor functional for at least 15 minutes after power failure
3.51.24 | ECG/RESPIRATION
3.51.24.1 | A full set of 3 lead ECG cables shall be included in the offer
3.51.24.2 | The ECG amplifier shall employ a differential input with high noise suppression
3.51.24.3 | It shall be possible to change the sweep speed of the ECG between 25 and 50mm/sec
3.51.24.4 | Bidders to state if a freeze facility for ECG is available
3.51.24.5 | The input impedance shall be higher than 10 MegOhms
3.51.24.6 | The heart rate measuring range shall be between 0 and 250 beats per minute, with an accuracy of not more than 5 bpm
3.51.24.7 | An AC filter shall be provided
3.51.24.8 | Alarm limits shall be adjustable in the following range
3.51.24.8.1 | Low – 40 to 150
3.51.24.8.2 | High – 70 to 250
3.51.24.9 | The violation of alarm limits shall result in an audible and visual alarm. It shall be possible to temporarily silence the audible alarm. State the time duration of the temporary silence
3.51.24.10 | The respiration rate shall be measured by means of impedance via the ECG leads
3.51.24.11 | Measurement range for respiration shall be approximately 4 to 120 breaths/min
3.51.25 | NON INVASIVE BLOOD PRESSURE
3.51.25.1 | The monitor shall make use of the oscillometric or other reliable measuring method
3.51.25.2 | Systolic, Diastolic and Mean pressure measurements shall be displayed
3.51.25.3 | It shall be possible to program the sampling interval for automatic NIBP measurements
3.51.25.4 | The shortest period of sampling shall not be longer than one minute
3.51.25.5 | The cuff pressure shall be limited to a comfortable level for both paediatric and adult patients
3.51.25.6 | The NIBP cuffs shall be robust, interchangeable and easy to clean
3.51.25.7 | A full set of reusable cuffs shall be supplied with each monitor. This shall include the range from neonatal, paediatric to obese and thigh cuffs. Bidders to state which cuffs are included in the offer
3.51.25.8 | Two sets of NIBP hoses shall be included in the offer per monitor– One set for adult measurement, and one set for paediatric measurement
3.51.25.9 | It shall be possible to set high and low alarm limits for Systolic, Diastolic and Mean Arterial Pressure. Default settings to be stated
3.51.26 | OXYGEN SATURATION
3.51.26.1 | Continuous monitoring of oxygen saturation levels by means of the photometric principle shall be possible
3.51.26.2 | The measuring range shall be from low to 100%. Bidders to state the accuracy levels
3.51.26.3 | It must be the latest technology with clinical research support based on accuracy and real-time monitoring. Please submit the published peer reviewed/researched paper as supporting documents.
3.51.26.4 | A finger probe for adults shall be included in the offer, and a suitable probe for paediatric use. These shall be per monitor. Bidders to state the replacement price for each probe, and state other types of probes available
3.51.26.5 | The monitor shall provide an audible tone indicating fall or change in oxygen saturation
3.51.26.6 | The monitor shall have a continuous oxygen saturation trace, and units displaying oxygen saturation with only a bar graph will not be acceptable
3.51.27 | CAPNOGRAPHY / ANAESTHETIC AGENT MONITORING
3.51.27.1 | The capnography shall be an integral part of the monitor, and no loose attachments shall be acceptable
3.51.27.2 | The capnography measurement shall be in kPa and mmHg
3.51.27.3 | Bidders to state the method employed to get rid of water build-up in the measuring line
3.51.27.4 | Bidders to state the cost of the measuring line. A set of at least 10 lines shall be included per monitor
3.51.27.5 | The monitor shall also include automatic agent identification of at least the following anaesthetic agents
3.51.27.5.1 | Halothane
3.51.27.5.2 | Isoflurane
3.51.27.5.3 | Sevoflurane
3.51.27.6 | The unit shall automatically alert when an anesthetic agent mixture is detected
3.51.27.7 | Inspiration and expiration values of all gases shall be displayed
3.51.27.8 | The unit shall have individual alarm limits for each parameter
3.51.27.9 | The unit shall have automatic zero calibration
3.51.28 | INVASIVE BLOOD PRESSURE
3.51.28.1 | The monitor shall have the ability to monitor 2 channels of Invasive Pressure
3.51.28.2 | Measurement range shall be approximately -40 to 320mm Hg
3.51.29 | TEMPERATURE
3.51.29.1 | The monitor shall be able to measure at least one temperature
3.51.29.2 | Temperature range shall be approximately 10º to 45º Celsius
3.51.29.3 | Two temperature probes shall be included per monitor – one oral and one rectal probe
4 | The following Accessories and Consumables must be supplied with the unit:
4.1 | Re-breathing bag, reusable, rubber
4.2 | Reservoir bag
4.3 | Pressure limiting valve, adjustable
4.4 | Oxygen fuel cell analyser with fuel cell
4.5 | Adult & Paediatric volume sensor re-usable
4.6 | Water trap for multigas monitor, re-usable (10 pcs.)
4.7 | A full set paediatric 3/5 lead ECG cables
4.8 | A full set adult 3/5 lead ECG cables
4.9 | Reusable paediatric cuffs
4.10 | Reusable adult cuffs
4.11 | Reusable neonate cuffs
4.12 | Reusable obese cuffs
4.13 | Air hoses
4.14 | Finger sensor for paeds
4.15 | Finger sensor for adults
4.16 | Earlobe sensor for paeds
4.17 | Earlobe sensor for adults
4.18 | Finger wrap-around probe for neonates
4.19 | Sensor lead
4.20 | Two reusable temperature probes
4.21 | State the cost and frequency of replacement for the flow sensors.
4.22 | State the cost and frequency of replacement for the fuel cell.
4.23 | Vaporizer key-fillers (Sevo & Iso)
4.24 | Sevoflurane vaporizer
4.25 | Isoflurane vaporizer
4.26 | Anaesthesia circuit (Adult, single/double limb design) Disposable. 1.8m (Complete with Y-piece, Elbow, Catheter mount, Rebreathing bag)
4.27 | Anaesthesia circuit (Paediatric, single/double limb design) Disposable. 1.8m (Complete with Y-piece, Elbow, Catheter mount, Rebreathing bag)
4.28 | Anaesthesia circuit (neonatal, single/double limb design) Disposable. 1.8m (Complete with Y-piece, Elbow, Catheter mount, Rebreathing bag)
4.29 | Re-breathing circuit complete
4.30 | Non-re-breathing circuit complete
4.31 | Flow sensor (pack of 2)
4.32 | Spirometry line
4.33 | Sample line (pack of 10)
4.34 | Face masks: sizes 00, 0, 1, 2, 3, 4, 5.
4.35 | Bacterial/viral filter
4.36 | Y connector
4.37 | Manual ventilation hose
4.38 | Fresh gas tubing
4.39 | Seal for back-up cylinder connection on the unit
4.40 | Scavenger tubing (as per institution's scavenging outlet)
4.41 | Inspiratory and expiratory unidirectional valves
4.42 | Side stream CO2 kit (adult/paed)
4.43 | Disposable Adult Spirometry Set complete with spirometry tube, sample line and flow sensor, 3 m, box of 20 pieces
4.44 | Disposable Paediatric Spirometry Set complete with spirometry tube, sample line and flow sensor, 3 m, box of 20 pieces
4.45 | Gas sample tube with luer lock connection, 3m, disposable (10 pcs)
4.46 | Suction Tubing Set inc connector
4.47 | Resuscitator Adult, 1500mL, autoclavable, silicone
4.48 | Resuscitator Paediatric, 550mL, autoclavable, silicone
4.49 | Tubing set Paediatric for circle systems
5 | Conformity Compliance (Please attach certificate)
5.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 and 601-1-2 for medical equipment, attached certification
5.2 | The unit must comply with electrical safety and performance standard such as IEC 60601-2-13 for Anaesthetic systems, attached certification
5.2.1 | Please provide unique ref number of the IEC ceriticate:
5.3 | OEM must comply and certified on ISO 9001 quality standards, attach proof of compliance
5.3.1 | Please provide unique ref number of the ISO 9001 ceriticate:
5.4 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
5.4.1 | Please provide unique ref number of the ISO 13485 ceriticate:
5.5 | Model quoted for must be CE certified. Attach a copy of certification
5.5.1 | Please provide unique ref number of the CE ceriticate:
5.6 | Hazardous Substance Act: If this type of equipment / apparatus appears on the schedule of Hazardous substances issued by the Directorate: Health Technology of the Department of Health, a license in terms of the Act on Hazardous Substances (Act 15/1973) must be submitted with the bid / quotation documents. The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.7 | All electrical/electronic medical equipment must be licensed by Radiation Control, where applicable a copy of the license must be submitted The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.7.1 | Radiation Control License number:
5.8 | The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade (rubber to with 2 screws) type and it must be a minimum length of three (3) meters. NB the mains cable of the unit tendered for must be SABS colour coded.
5.9 | The equipment quoted for must be protected against electromagnetic interference, IEC 60601-1-2
5.10 | The unit must comply with electrical safety and performance standard such as IEC 60601-2-49 for Patient Monitoring, attached certification
5.10.1 | Please provide unique ref number of the IEC ceriticate:
6 | COMPREHENSIVE MAINTENANCE
6.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) must be included in the bid price
6.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries, Flow sensors and Oxygen sensor), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
6.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
6.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries, Flow sensors and Oxygen sensor), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
6.3 | The bidder to provide the OEM recommended service interval of the unit:
6.4 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
7 | General
7.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
7.2 | The unit on offer must be the latest technology.
ANAESTHESIA DELIVERY SYSTEM: Basic with integrated physiological monitor, Replacement Accessories
RT2-01-02a | A full set paediatric 3/5 lead ECG cable
RT2-01-02b | A full set adult 3/5 lead ECG cable
RT2-01-02c | Reusable paediatric cuff
RT2-01-02d | Reusable adult cuff
RT2-01-02e | Reusable neonate cuff
RT2-01-02f | Reusable obese cuff
RT2-01-02g | Air hose
RT2-01-02h | Finger sensor for paeds
RT2-01-02i | Finger sensor for adults
RT2-01-02j | Earlobe sensor for paeds
RT2-01-02k | Earlobe sensor for adults
RT2-01-02l | Finger wrap-around probe for neonates
RT2-01-02m | Sensor lead
RT2-01-02n | Reusable temperature probe
Bidder to list any other Consumables or Accessories (Not listed above) for the model on offer below

## Anaesthetic, Standard
RT2-01-03 | ANAESTHESIA DELIVERY SYSTEM: Standard without physiological monitor
1 | Description
2 | Product Offered (Anaesthetic Machine)
2.1 | Make:
2.2 | Model:
2.3 | Initial Manufacturing year of Model:
2.4 | Radiation Control License number:
3 | Technical Specification
3.1 | Anaesthesia delivery system suitable for neonatal, paediatric and adult anaesthesia
3.2 | Ventilator suitable for ventilating paediatric to adult patients
3.3 | Multigas analyser:
3.3.1 | Oxygen
3.3.2 | Nitrous oxide
3.3.3 | Carbon dioxide
3.4 | Anaesthetic agent monitoring:
3.4.1 | Isofluorane
3.4.2 | Sevofluorane
3.4.3 | Desfluorane
3.5 | The monitor must be able to simultaneously display the inspiratory and expiratory O2, CO2, N2O and Anaesthetic agent concentrations.
3.6 | The monitor must be able to identify the Anaesthetic agent automatically. (During a change between agents it is very desirable that the monitor must be able to measure the agents simultaneously with an alert message. i.e.: wash in and wash out values of both agents)
3.7 | Ketone, methane & Alcohol concentrations should not influence the measurements of the anaesthetic gasses analyzed
3.8 | Gas analyzer Mac point should be age corrected. The monitor should be able to display measured MAC values
3.9 | Self test requiring limited user intervention.
3.10 | The self test must provide the measurements and values for automatic leak detection and compliance.
3.11 | Automatic calibration of sensors for gases and anaesthetic vapours must form part of the self test and require limited user intervention.
3.12 | The unit must provide conventional, low-flow (down to 200 ml/min) anaesthesia.
3.13 | It must be possible to override the self test facility. The system must display the override status
3.14 | The unit must return analysed gas back to the circle system
3.15 | The unit must be supplied with a minimum of 2 storage drawer
3.16 | The unit must be supplied with a minimum of 4 additional 220 V power sockets. These sockets must be excluded from the battery backup circuit in case of a power failure
3.17 | The unit must have a working area or a retractable shelf. Width: not less than 300 mm. Depth: not less than 200 mm
3.18 | The unit must have a light to enable the anaesthetist to work in a darkened theatre.
3.19 | Connectivity with automated record system should be possible (Optional)
3.20 | Trolley
3.20.1 | The unit must be mounted on a strong , robust and compact trolley manufactured from corrosion free material.
3.20.2 | The trolley must have no sharp corners and be easy to clean.
3.20.3 | The trolley must have at least four antistatic castors of diameter 100 mm or greater.
3.20.4 | Trolley must have two lockable castors or a central locking brake.
3.20.5 | The unit must accommodate an oxygen cylinder. Cylinder size according to user requirement. Size to be compatible with current state vendor of oxygen.
3.20.6 | Trolley must have accessible handles for ease of movement
3.20.7 | The trolley, with all components mounted, must be stable without toppling over when placed on a 30° slope.
3.20.8 | The unit must be supplied with a castor guard system, preventing the entanglement and trapping of cables within the wheels.
3.20.9 | The trolley must have a mounting system suitable for a physiologial monitor.
3.20.10 | The unit should be ergonomically designed to ensure maximum ease and safety when used.
3.20.11 | All elements of the system shall be software driven allowing easy future upgrades. All parts of the aesthesia machine: gas mixer, vaporizer and ventilator shall preferably be software driven
3.21 | Pipeline and Cylinder Gas Supplies
3.21.1 | The anaesthesia system must be designed to use external piped medical gasses (O2, N2O and medical air) and vacuum from medical service outlet sockets.
3.21.2 | Connection hoses for above must be provided
3.21.3 | Hoses must be fitted with Heyer type non-interchangeable probes for the existing medical gas outlets as specified by the purchaser.
3.21.4 | The medical air hose connection or CO2 circle system must include a water trap
3.21.5 | A safety pressure relief valve of 60 - 80 cm H2O must be provided to protect the system from excessive pressure .
3.21.6 | The unit must function safely and correctly within the pipeline supply pressure range of: 270 - 600 kPa
3.21.7 | The system must preferentially utilise the piped oxygen supply over the cylinder supply, if pipeline pressure is > 270 kPa
3.21.8 | Connection hoses must be:
3.21.8.1 | Flexible and at least 6 metres long
3.21.8.2 | Antistatic
3.21.9 | Colour coded as follows:
3.21.9.1 | O2: white
3.21.9.2 | N2O: blue
3.21.9.3 | Vacuum: yellow
3.21.9.4 | Medical air: black
3.21.9.5 | Scavenge low vacuum: orange or yellow
3.21.9 | Where relevant, fitted with outlet probes complying with SANS 1409:2014
3.21.10 | Permanantly bonded to the outlet probe and the machine connector
3.21.11 | Hoses for gases above atmospheric pressure, must designed for pressures of at least 1 000 kPa.
3.21.12 | Vacuum hoses may not collapse or kink at 90 kPa below atmospheric pressure, when subjected to 180° bend of inside radius equal to the outside diameter of the hose.
3.21.13 | All connection hoses must be fed to a central manifold
3.21.14 | All high pressure gas and vacuum pipe connections (other than hoses) and all pipes connected to pressure regulators and pressure gauges must be stainless steel (18/8 grade) pipes that are oil-free and colour coded as for hoses (see above) over their entire lengths
3.21.15 | An in-line non-return valve must be fitted to each gas pipeline on the high pressure (inlet) line
3.21.16 | An in-line non-return valve must be fitted to each gas cylinder yoke
3.21.17 | A replaceable gas filter, to block particles greater than 65 µm, must be fitted on the upstream side of the non-return valve
3.21.18 | Each gas pipeline inlet must have a pressure gauge or electronic display of pressure. The pressure range (415 kPa ± 75 kPa) must be clearly indicated in colour
3.21.19 | Must be fitted with a gas cylinder yoke to accommodate one oxygen gas cylinder with a pin index system
3.21.20 | Each gas cylinder yoke must have a pressure gauge or electronic display of pressure
3.21.21 | The pressure gauges or electronic displays of pressure for both pipeline & cylinder supplies, calibrated in kPa for each line, must be labelled and visible from the front of the unit.
3.21.22 | O2 and N2O regulators must be provided complying with CKS 605:1987.
3.21.23 | A pressure regulator complying with CKS 605 must be supplied for each gas inlet and must reduce the gas pressure to 415 kPa without freezing up, creeping, chattering or whistling
3.21.24 | Outlet sockets and probes must comply with SANS 1409:2014.
3.21.25 | Sockets with hoses must be colour-coded as per SANS 1409:2014:
3.21.26 | Oxygen: White background with black letters
3.21.27 | Nitrous oxide: Blue background with white letters
3.21.28 | Medical air: One half black with white letters, the other half white with black letters
3.21.29 | Vacuum: Yellow background with black letters
3.21.30 | Continuous display, visible from the front of the machine, of both pipeline and cylinder pressures even during power failure
3.21.31 | Pipeline and cylinder pressures must be differentially displayed, and clearly indicated
3.21.32 | Safety pressure relief valves must protect the entire pneumatic system from excessive input pressures from pipeline and cylinder supplies.
3.21.33 | When a cylinder is detached from the pin-indexed yoke no gas leak must occur at the yoke.
3.21.34 | Airway pressure must be monitored and displayed
3.21.35 | Safety relief valve for all gases must be supplied and set at 600kPa
3.22 | Scavenging
3.22.1 | The unit must have a scavenging system with a hose at least 5 metres long with the suitable connector as per the institutions scavenging installation.
3.22.2 | Air-break device to prevent negative pressure on the breathing system
3.22.3 | If an active sytem is supplied a pressure indicator must be present
3.22.4 | The scavenging outlet must be a 30mm ISO connector.
3.23 | Patient Gas Delivery System (Flowmeters/Rotameters)
3.23.1 | The unit must have cascading flowmeters (mechanical or electronic) that show flow rates for high and low flow anaesthesia in the following ranges:
3.23.1.1 | Oxygen low flow: 200 - 1 000 ml/min
3.23.1.2 | At higher flow rates the gas must cascade into a second flowmeter (1 – 10 l/min)
3.23.1.3 | Nitrous oxide low flow: 200 - 1 000 ml/min
3.23.1.4 | At higher flow rates the gas must cascade into a second flowmeter (1 – 10 l/min)
3.23.1.5 | Air tube 200ml/min – 10 l/min
3.23.1.6 | Flowmeters, if not electronic, must be diameter indexed to prevent incorrect replacement during maintenance.
3.23.2 | Gas supply must conform to:
3.23.2.1 | O2 minimum flow range of 200 ml/min to 10 l/min. Accuracy ± 5%
3.23.2.2 | medical air minimum flow range 0 to 10 l/min. Accuracy ± 5%
3.23.2.3 | nitrous oxide minimum flow range 0 to 10 l/min. Accuracy ± 5%
3.23.3 | Flowmeters must be accurate to at least 5%
3.23.4 | Flowmeter valve control
3.23.5 | Flowmeter control knobs
3.23.6 | Flowmeter sequence to be according to standards approved by the South Africal Society of Anaesthesiologists - Oxygen on the right
3.23.7 | The delivery system must prevent simultaneous delivery of two anaesthetic vapours.
3.23.8 | A separate common gas outlet to accept a standard 22 mm outside diameter and 15 mm internal diameter.
3.23.9 | A separate oxygen gas outlet for connection to and oxygen mask/Ambu bag tubing
3.24 | Hypoxic guard
3.24.1 | The unit must include a safety mechanism that ensures a minimum concentration of 25% oxygen in any oxygen/nitrous oxide gas mixture across all flow rates.
3.24.2 | The unit must include a safety mechanism that ensures a minimum concentration of 21% oxygen in any oxygen/air gas mixture across all flow rates.
3.24.3 | The unit must include a high-flow oxygen flush mechanism with a flow rate of at least 35 l/min
3.24.4 | Oxygen flush must be non-lockable, recessed and self closing
3.24.5 | This system must bypass the flowmeters and vaporisers
3.25 | Breathing System
3.25.1 | Total circuit volume, including absorber, must not exceed 2,7 litres
3.25.2 | With APL valve 5 - 75 cm H2O
3.25.3 | With ventilator to manual switch. This switch when directed to the ventilator must isolate the APL valve from the circle
3.25.4 | Breathing system must be fully autoclavable to 134°C
3.25.5 | There must be a mechanism to remove water from the breathing system
3.25.6 | Breathing system must be latex free
3.25.7 | Reusable corrugated breathing circuit adequate for both adult and paediatric use to be supplied
3.25.8 | Bidders to supply the breathing circuit as required by the user
3.25.9 | If unit is using bellows, they must be rising and visible
3.25.10 | The gas machine must have 15/22 fresh gas outlet and thus accept all commonly used breathing systems. It shall be possible to use Bain and Jackson Rees systems to allow easy spontaneous breathing and fast response times in paediatric anaesthesia
3.25.11 | The absorber canister should be exchangeable during machine use without loss of pressure and circuit gas
3.25.12 | The absorber canister should be exchangeable during machine use without loss of pressure and circuit gas
3.26 | Integrated Ventilator
3.26.1 | Ventilator suitable for ventilating adults, paediatric and neonatal patients without changing ventilator components (e.g. bellows)
3.26.2 | The ventilator must be electronically controlled
3.26.3 | The ventilation technology offered may be turbine, volume reflector, piston or ascending bellows
3.26.4 | The minimum basic modes to be supplied:
3.26.4.1 | Manual/Spontaneous Ventilation
3.26.4.2 | Volume Control Ventilation
3.26.4.3 | Pressure Control Ventilation
3.26.4.4 | SIMV
3.26.5 | Additional modes of ventilation may be offered and quoted for separately
3.26.6 | The ventilator must be integrated with the workstation and user friendly
3.26.7 | Frequency: minimum range of 5 – 60 bpm.
3.26.8 | Inspiratory time: minimum range of 0.2 – 10.0 seconds.
3.26.9 | Minimum Inspiration/Expiration Ratio: 4:1 – 1:4.
3.26.10 | Adjustable inspiratory pause/plateau time (volume control ventilation): 0-30% of inspiratory time
3.26.11 | Peak Inspiratory Pressure: minimum range of: PEEP + 5 to 70 cm H2O.
3.26.12 | Positive End Expiratory Pressure (PEEP): Integrated, electronically controlled. Minimum range: 0 – 30 cm H2O.
3.26.13 | Tidal Volume: 20ml – 1500ml per breath. Minimum accuracy: 15%
3.26.14 | Peak Inspiratory Flow: At least 120 lpm
3.26.15 | The ventilator must be capable of flow or pressure triggering.
3.26.16 | Adjustable Inspiratory Rise time/Ramp with a range of 0.1 - 2 sec
3.26.17 | Compliance compensation/Tidal volume compensation/Compressible volume compensation or equivalent.
3.26.18 | Patient gas flow compensation must be automatic so that alterations in patient gas flow do not affect delivered tidal volumes
3.26.19 | In event of power or technical failure, manual ventilation must be possible.
3.26.20 | The ventilator must be compatible with a circle system and open system
3.27 | Display:
3.27.1 | All displays must be in acceptable English
3.27.2 | Colour display must include numerical and graphic display of parameters measured and set ventilator settings
3.27.3 | Graphic display to feature at least 3 waveforms: - pressure, flow and volume.
3.27.4 | Must also display flow-volume and pressure-volume loops with a reference loop
3.27.5 | The flow sensor has to be service free
3.27.6 | Alarms : The ventilator must have the following user adjustable alarms
3.27.6.1 | Apnoea alarm
3.27.6.2 | High airway pressure alarm
3.27.6.3 | Continuous pressure alarm
3.27.6.4 | High and low expiratory tidal volume alarm
3.27.6.5 | Minute volume, high/low alarm
3.27.6.6 | Inspired oxygen (Fi02), high/low alarm
3.27.6.7 | PEEP, low alarm
3.27.6.8 | Leakage / Check breathing circuit alarm or equivalent
3.27.6.9 | Alarm profile must be user definable or default to patient appropriate settings
3.27.6.10 | The ventilator must have a mute alarm function, active only for a limited time(no more than 2 minutes).
3.27.6.11 | It must not be possible to cancel the visual alarm until the problem has been rectified.
3.27.7 | All components of the ventilator and breathing system in contact with the patient’s airway must be latex free.
3.27.8 | All components of the ventilator and breathing system in contact with the patient’s airway must be steam autoclavable up to 134ᴼ C
3.28 | Integrated Gas Monitoring
3.28.1 | Oxygen Monitoring
3.28.1.1 | The system must display both inspired and end-tidal oxygen concentration.
3.28.1.2 | The oxygen analyser must have user adjustable low and high alarms – visual and audible. The low alarm must automatically default to no less than 21%.
3.28.1.3 | The system must be accurate (+/- 2% in 21-100%) range, and have a rapid response time (≤ 500ms).
3.28.1.4 | The unit must sound a continuous alarm if the oxygen pressure falls below 190 kPa, then all nitrous oxide must be automatically cut from the patient circuit.
3.28.1.5 | Analyser preferrably of paramagnetic type. If fuel cell offered, provide full information on cost of disposables and frequency of replacement
3.28.2 | Circle System
3.28.2.1 | The circle absorber system must be integrated as part of the ventilation system.
3.28.2.2 | The circle absorber system must have an effective system to eliminate excess water created by the chemical reactions either by heating or water traps.
3.28.2.3 | The circle absorber system must be autoclavable at 134ºC.
3.28.2.4 | The circle absorber system must be supplied with a carbon dioxide absorber canister.
3.28.2.5 | The absorber canister should be reusable/refillable. Disposable cannisters should be quoted separately
3.28.2.6 | It must be easy to remove and replace the absorber canister while the workstation is in operation.
3.28.2.7 | Misalignment of the canister must be readily apparent to the user.
3.28.2.8 | The system must have a circle bypass facility while changing/refilling the absorber canister.
3.28.2.9 | The breathing circuit must contain an integrated adjustable pressure limiting (APL) valve with a minimum range of 0-70 cm H2O.
3.28.2.10 | The APL valve must have a tactile or electronic indicator when a value of between 20-30 cm H2O is exceeded.
3.28.2.11 | The circle absorber system must be latex free.
3.28.2.12 | The bag-to-ventilator switching mechanism must be clearly labeled and user-friendly. Mechanical or electronic
3.29 | Patient circuit pressure release valve
3.29.1 | Adjustable up to 70 cm water
3.29.2 | Audible alarm when the valve opens
3.29.3 | Manual override, non-lockable, to keep the release valve closed when the set pressure is exceeded.
3.30 | Flowmeter pressure release valve
3.30.1 | Situated between the flowmeters and the vaporisers
3.30.2 | Must open at 105 kPa +- 5 kPa
3.31 | Other System requirements
3.31.1 | The unit must operate from both the 220 V, 50 Hz mains supply and internal rechargeable batteries
3.31.2 | The batteries must be charged while the unit is operating from the mains supply.
3.31.3 | The batteries must automatically take over in the event of a mains power failure.
3.31.4 | The batteries must have the capacity for full operation of the anaesthetic machine and ventilator for at least 90 minutes with batteries fully charged, under typical operating conditions.
3.31.5 | Graphic battery power indicator
3.31.6 | Recharge time < 8 hrs
3.31.7 | Mains cable at least four (4) metres long with SABS approved 15A three prong plug.
3.31.8 | The unit must have a rail for attaching a suction regulator.
3.31.9 | The unit must be handed over in full operating order
4 | The following Accessories and Consumables must be supplied with the unit:
4.1 | Re-breathing bag, reusable, rubber
4.2 | Reservoir bag
4.3 | Pressure limiting valve, adjustable
4.4 | Oxygen fuel cell analyser
4.5 | Flow sensors.
4.6 | Fuel cell.
4.7 | Vaporizer key-fillers (Sevo, & Iso)
4.8 | The unit must be supplied with two vaporizers (one must sevoflurane, the other must be of the choice of the buying institution)
4.9 | Anaesthesia circuit (Adult, single/double limb design) Disposable. 1.8m (Complete with Y-piece, Elbow, Catheter mount, Rebreathing bag)
4.10 | Anaesthesia circuit (Paediatric, single/double limb design) Disposable. 1.8m (Complete with Y-piece, Elbow, Catheter mount, Rebreathing bag)
4.11 | Anaesthesia circuit (neonatal, single/double limb design) Disposable. 1.8m (Complete with Y-piece, Elbow, Catheter mount, Rebreathing bag)
4.12 | Re-breathing circuit complete
4.13 | Non-re-breathing circuit complete
4.14 | Maintenance-free Flow sensor (pack of 2)
4.15 | Spirometry line
4.16 | Sample line (pack of 10)
4.17 | Face masks: sizes 00, 0, 1, 2, 3, 4, 5.
4.18 | Bacterial/viral filter
4.19 | Y connector
4.20 | Manual ventilation hose
4.21 | Fresh gas tubing
4.22 | Seal for back-up cylinder connection on the unit
4.23 | Scavenger tubing
4.24 | Inspiratory and expiratory unidirectional valves
4.25 | Side stream CO2 kit (adult/paed)
4.26 | Disposable Adult Spirometry Set complete with spirometry tube, sample line and flow sensor, 3 m, box of 20 pieces
4.27 | Disposable Paediatric Spirometry Set complete with spirometry tube, sample line and flow sensor, 3 m, box of 20 pieces
4.28 | Gas sample tube with luer lock connection, 3m, disposable (10 pcs)
4.29 | Suction Tubing Set inc connector
4.30 | Resuscitator Adult, 1500mL, autoclavable, silicone
4.31 | Resuscitator Paediatric, 550mL, autoclavable, silicone
4.32 | Tubing set Paediatric for circle systems
5 | OPTIONAL ITEMS
5.1 | Prices incl VAT must be given for all optional items
5.2 | Brochures with full technical specifications must be supplied for each optional item offered
6 | Conformity Compliance (Please attach certificate)
6.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
6.2 | The unit must comply with electrical safety and performance standard such as IEC 60601-2-13 for Anaesthetic systems, attached certification
6.2.1 | Please provide unique ref number of the IEC ceriticate:
6.3 | OEM must comply and certified on ISO 9001 quality standards, attach proof of compliance
6.3.1 | Please provide unique ref number of the ISO 9001 ceriticate:
6.4 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
6.4.1 | Please provide unique ref number of the ISO 13485 ceriticate:
6.5 | Model quoted for must be EC certified. Attach a copy of certification
6.5.1 | Please provide unique ref number of the EC ceriticate:
6.6 | Hazardous Substance Act: If this type of equipment / apparatus appears on the schedule of Hazardous substances issued by the Directorate: Health Technology of the Department of Health, a license in terms of the Act on Hazardous Substances (Act 15/1973) must be submitted with the bid / quotation documents. The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
6.7 | All electrical/electronic medical equipment must be licensed by Radiation Control, where applicable a copy of the license must be submitted The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
6.7.1 | Radiation Control License number:
6.8 | The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade (rubber to with 2 screws) type and it must be a minimum length of three (3) meters. NB the mains cable of the unit tendered for must be SABS colour coded.
6.9 | The equipment quoted for must be protected against electromagnetic interference IEC 60601-1-2
7 | COMPREHENSIVE MAINTENANCE
7.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) must be included in the bid price
7.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries, Flow sensors and Oxygen sensor), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
7.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
7.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries, Flow sensors and Oxygen sensor), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
7.3 | The bidder to provide the OEM recommended service interval of the unit:
7,4 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
8 | General
8.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
8.2 | The unit on offer must be the latest technology.
ANAESTHESIA DELIVERY SYSTEM: STANDARD WITHOUT PHYSIOLOGICAL MONITOR, REPLACEMENT ACCESSORIES
RT2-01-03a | Anaesthesia circuit (Adult, single/double limb design) Disposable. 1.8m (Complete with Y-piece, Elbow, Catheter mount, Rebreathing bag)
RT2-01-03b | Anaesthesia circuit (Paediatric, single/double limb design) Disposable. 1.8m (Complete with Y-piece, Elbow, Catheter mount, Rebreathing bag)
RT2-01-03c | Anaesthesia circuit (neonatal, single/double limb design) Disposable. 1.8m (Complete with Y-piece, Elbow, Catheter mount, Rebreathing bag)
RT2-01-03d | Re-breathing circuit complete
RT2-01-03e | Non-re-breathing circuit complete
RT2-01-03f | Flow sensor (pack of 2)
RT2-01-03g | Spirometry line
RT2-01-03h | Sample line (pack of 10)
RT2-01-03i | Face masks: sizes (00,0,1,2,3,4, 5)- End-user to specify size
RT2-01-03j | Bacterial/viral filter
RT2-01-03k | Y connector
RT2-01-03l | Manual ventilation hose
RT2-01-03m | Fresh gas tubing
RT2-01-03n | Seal for back-up cylinder connection on the unit
RT2-01-03o | Scavenger tubing (as per institution's scavenging outlet)
RT2-01-03p | Inspiratory and expiratory unidirectional valves
RT2-01-03q | Side stream CO2 kit (adult/paed)
RT2-01-03r | Disposable Adult Spirometry Set complete with spirometry tube, sample line and flow sensor, 3 m, box of 20 pieces
RT2-01-03s | Disposable Paediatric Spirometry Set complete with spirometry tube, sample line and flow sensor, 3 m, box of 20 pieces
RT2-01-03t | Gas sample tube with luer lock connection, 3m, disposable (10 pcs)
RT2-01-03u | Suction Tubing Set inc connector
RT2-01-03v | Resuscitator Adult, 1500mL, autoclavable, silicone
RT2-01-03w | Resuscitator Paediatric, 550mL, autoclavable, silicone
RT2-01-03x | Tubing set Paediatric for circle systems
RT2-01-03y | sevoflurane vaporizer
RT2-01-03z | Isoflurane vaporizer
RT2-01-03zz | Desflurane vaporizer
Bidder to list any other Consumables or Accessories (Not listed above) for the model on offer below

## Anaesthetic, Standard + Monitor
RT2-01-04 | ANAESTHESIA DELIVERY SYSTEM: Standard with integrated physiological monitor
1 | Description
2a | Product Offered (Anaesthetic Machine)
2a.1 | Make:
2a.2 | Model:
2a.3 | Initial Manufacturing year of Model:
2a.4 | Radiation Control License number:
2b | Product Offered (Monitor)
2b.1 | Make:
2b.2 | Model:
2b.3 | Initial Manufacturing year of Model:
2b.4 | Radiation Control License number:
3 | Technical Specification
3.1 | Anaesthesia delivery system suitable for neonatal, paediatric and adult anaesthesia
3.2 | Ventilator suitable for ventilating neonatal to adult patients
3.3 | Multigas analyser:
3.3.1 | Oxygen
3.3.2 | Nitrous oxide
3.3.3 | Carbon dioxide
3.4 | Anaesthetic agent monitoring:
3.4.1 | Isofluorane
3.4.2 | Sevofluorane
3.4.3 | Desfluorane
3.5 | The monitor must be able to simultaneously display the inspiratory and expiratory O2, CO2, N2O and Anaesthetic agent concentrations.
3.6 | The monitor must be able to identify the Anaesthetic agent automatically. (During a change between agents it is very desirable that the monitor must be able to measure the agents simultaneously with an alert message. i.e.: wash in and wash out values of both agents)
3.7 | Ketone, methane & Alcohol concentrations should not influence the measurements of the anaesthetic gasses analyzed
3.8 | Gas analyzer Mac point should be age corrected. The monitor should be able to display measured MAC values
3.9 | Self test requiring limited user intervention.
3.10 | The self test must provide the measurements and values for automatic leak detection and compliance.
3.11 | Automatic calibration of sensors for gases and anaesthetic vapours must form part of the self test and require very limited user intervention.
3.12 | The unit must provide conventional, low-flow (down to 200 ml/min) anaesthesia.
3.13 | It must be possible to override the self test facility. The system must display the override status
3.14 | The unit must return analysed gas back to the circle system
3.15 | The unit must be supplied with a minimum of 2 storage drawers
3.16 | The unit must be supplied with a minimum of 4 additional 220 V power sockets. These sockets must be excluded from the battery backup circuit in case of a power failure
3.17 | The unit must have a working area or a retractable shelf. Width: not less than 300 mm. Depth: not less than 200 mm
3.18 | The unit must have a light to enable the anaesthetist to work in a darkened theatre.
3.20 | Trolley
3.20.1 | The unit must be mounted on a strong , robust and compact trolley manufactured from corrosion free material.
3.20.2 | The trolley must have no sharp corners and be easy to clean.
3.20.3 | The trolley must have at least four antistatic castors of diameter 100 mm or greater.
3.20.4 | Trolley must have two lockable castors or a central locking brake.
3.20.5 | The unit must accommodate an oxygen cylinder. Cylinder size according to user requirement. Size to be compatible with current state vendor of oxygen.
3.20.6 | Trolley must have accessible handles for ease of movement
3.20.7 | The trolley, with all components mounted, must be stable without toppling over when placed on a 30° slope.
3.20.8 | The unit must be supplied with a castor guard system, preventing the entaglement and trapping of cables within the wheels.
3.20.9 | The trolley must have a mounting system suitable for a physiologial monitor.
3.20.10 | The unit should be ergonomically designed to ensure maximum ease and safety when used.
3.20.11 | All elements of the system shall be software driven allowing easy future upgrades. All parts of the aesthesia machine: gas mixer, vaporizer and ventilator shall preferably be software driven
3.21 | Pipeline and Cylinder Gas Supplies
3.21.1 | The anaesthesia system must be designed to use external piped medical gasses (O2, N2O and medical air) and vacuum from medical service outlet sockets.
3.21.2 | Connection hoses for above must be provided
3.21.3 | Hoses must be fitted with Heyer type non-interchangeable probes for the existing medical gas outlets as specified by the purchaser.
3.21.4 | The medical air hose connection or CO2 circle system must include a water trap
3.21.5 | A safety pressure relief valve of 60 - 80 cm H2O must be provided to protect the system from excessive pressure .
3.21.6 | The unit must function safely and correctly within the pipeline supply pressure range of: 270 - 600 kPa
3.21.7 | The system must preferentially utilise the piped oxygen supply over the cylinder supply, if pipeline pressure is > 270 kPa
3.21.8 | Connection hoses must be:
3.21.8.1 | Flexible and at least 6 metres long
3.21.8.2 | Antistatic
3.21.9 | Colour coded as follows:
3.21.9.1 | O2: white
3.21.9.2 | N2O: blue
3.21.9.3 | Vacuum: yellow
3.21.9.4 | Medical air: black
3.21.9.5 | Scavenge low vacuum: orange or yellow
3.21.9 | Where relevant, fitted with outlet probes complying with SANS 1409:2014
3.21.10 | Permanantly bonded to the outlet probe and the machine connector
3.21.11 | Hoses for gases above atmospheric pressure, must designed for pressures of at least 1 000 kPa.
3.21.12 | Vacuum hoses may not collapse or kink at 90 kPa below atmospheric pressure, when subjected to 180° bend of inside radius equal to the outside diameter of the hose.
3.21.13 | All connection hoses must be fed to a central manifold
3.21.14 | All high pressure gas and vacuum pipe connections (other than hoses) and all pipes connected to pressure regulators and pressure gauges must be stainless steel (18/8 grade) pipes that are oil-free and colour coded as for hoses (see above) over their entire lengths
3.21.15 | An in-line non-return valve must be fitted to each gas pipeline on the high pressure (inlet) line
3.21.16 | An in-line non-return valve must be fitted to each gas cylinder yoke
3.21.17 | A replaceable gas filter, to block particles greater than 65 µm, must be fitted on the upstream side of the non-return valve
3.21.18 | Each gas pipeline inlet must have a pressure gauge or electronic display of pressure. The pressure range (415 kPa ± 75 kPa) must be clearly indicated in colour
3.21.19 | Must be fitted with a gas cylinder yoke to accommodate one oxygen gas cylinder with a pin index system
3.21.20 | Each gas cylinder yoke must have a pressure gauge or electronic display of pressure
3.21.21 | The pressure gauges or electronic displays of pressure for both pipeline & cylinder supplies, calibrated in kPa for each line, must be labelled and visible from the front of the unit.
3.21.22 | O2 and N2O regulators must be provided complying with CKS 605:1987.
3.21.23 | A pressure regulator complying with CKS 605 must be supplied for each gas inlet and must reduce the gas pressure to 415 kPa without freezing up, creeping, chattering or whistling
3.21.24 | Outlet sockets and probes must comply with SANS 1409:2014.
3.21.25 | Sockets with hoses must be colour-coded as per SANS 1409:2014:
3.21.26 | Oxygen: White background with black letters
3.21.27 | Nitrous oxide: Blue background with white letters
3.21.28 | Medical air: One half black with white letters, the other half white with black letters
3.21.29 | Vacuum: Yellow background with black letters
3.21.30 | Continuous display, visible from the front of the machine, of both pipeline and cylinder pressures even during power failure
3.21.31 | Pipeline and cylinder pressures must be differentially displayed, and clearly indicated
3.21.32 | Safety pressure relief valves must protect the entire pneumatic system from excessive input pressures from pipeline and cylinder supplies.
3.21.33 | When a cylinder is detached from the pin-indexed yoke no gas leak must occur at the yoke.
3.21.34 | Airway pressure must be monitored and displayed
3.21.35 | Safety relief valve for all gases must be supplied and set at 600kPa
3.22 | Scavenging
3.22.1 | The unit must have a scavenging system with a hose at least 5 metres long with the suitable connector as per the institutions scavenging installation.
3.22.2 | Air-break device to prevent negative pressure on the breathing system
3.22.3 | If an active sytem is supplied a pressure indicator must be present
3.22.4 | The scavenging outlet must be a 30mm ISO connector.
3.23 | Patient Gas Delivery System (Flowmeters/Rotameters)
3.23.1 | The unit must have cascading flowmeters (mechanical or electronic) that show flow rates for high and low flow anaesthesia in the following ranges:
3.23.1.1 | Oxygen low flow: 200 - 1 000 ml/min
3.23.1.2 | At higher flow rates the gas must cascade into a second flowmeter (1 – 10 l/min)
3.23.1.3 | Nitrous oxide low flow: 200 - 1 000 ml/min
3.23.1.4 | At higher flow rates the gas must cascade into a second flowmeter (1 – 10 l/min)
3.23.1.5 | Air tube 200ml/min – 10 l/min
3.23.1.6 | Flowmeters, if not electronic, must be diameter indexed to prevent incorrect replacement during maintenance.
3.23.2 | Gas supply must conform to:
3.23.2.1 | O2 minimum flow range of 200 ml/min to 10 l/min. Accuracy ± 5%
3.23.2.2 | medical air minimum flow range 0 to 10 l/min. Accuracy ± 5%
3.23.2.3 | nitrous oxide minimum flow range 0 to 10 l/min. Accuracy ± 5%
3.23.3 | Flowmeters must be accurate to at least 5%
3.23.4 | Flowmeter valve control
3.23.5 | Flowmeter control knobs
3.23.6 | Flowmeter sequence to be according to standards approved by the South Africal Society of Anaesthesiologists - Oxygen on the right
3.23.7 | The delivery system must prevent simultaneous delivery of two anaesthetic vapours.
3.23.8 | A separate common gas outlet to accept a standard 22 mm outside diameter and 15 mm internal diameter.
3.23.9 | A separate oxygen gas outlet for connection to and oxygen mask/Ambu bag tubing
3.24 | Hypoxic guard
3.24.1 | The unit must include a safety mechanism that ensures a minimum concentration of 25% oxygen in any oxygen/nitrous oxide gas mixture across all flow rates.
3.24.2 | The unit must include a safety mechanism that ensures a minimum concentration of 21% oxygen in any oxygen/air gas mixture across all flow rates.
3.24.3 | The unit must include a high-flow oxygen flush mechanism with a flow rate of at least 35 l/min
3.24.4 | Oxygen flush must be non-lockable, recessed and self closing
3.24.5 | This system must bypass the flowmeters and vaporisers
3.25 | Breathing System
3.25.1 | Total circuit volume, including absorber, must not exceed 2,7 litres
3.25.2 | With APL valve 5 - 75 cm H2O
3.25.3 | With ventilator to manual switch. This switch when directed to the ventilator must isolate the APL valve from the circle
3.25.4 | Breathing system must be fully autoclavable to 134°C
3.25.5 | There must be a mechanism to remove water from the breathing system
3.25.6 | Breathing system must be latex free
3.25.7 | Reusable corrugated breathing circuit adequate for both adult and paediatric use to be supplied
3.25.8 | Bidders to supply the breathing circuit as required by the user
3.25.9 | If unit is using bellows, they must be rising and visible
3.25.10 | The gas machine must have 15/22 fresh gas outlet and thus accept all commonly used breathing systems. It shall be possible to use Bain and Jackson Rees systems to allow easy spontaneous breathing and fast response times in paediatric anaesthesia
3.25.11 | The absorber canister should be exchangeable during machine use without loss of pressure and circuit gas
3.25.12 | The absorber canister should be exchangeable during machine use without loss of pressure and circuit gas
3.26 | Integrated Ventilator
3.26.1 | Ventilator suitable for ventilating adults, paediatric and neonatal patients without changing ventilator components (e.g. bellows)
3.26.2 | The ventilator must be electronically controlled
3.26.3 | The ventilation technology offered may be turbine, volume reflector, piston or ascending bellows
3.26.4 | The minimum basic modes to be supplied:
3.26.4.1 | Manual/Spontaneous Ventilation
3.26.4.2 | Volume Control Ventilation
3.26.4.3 | Pressure Control Ventilation
3.26.4.4 | SIMV
3.26.5 | Additional modes of ventilation may be offered and quoted for separately
3.26.6 | The ventilator must be integrated with the workstation and user friendly
3.26.7 | Frequency: minimum range of 5 – 60 bpm.
3.26.8 | Inspiratory time: minimum range of 0.2 – 10.0 seconds.
3.26.9 | Minimum Inspiration/Expiration Ratio: 4:1 – 1:4.
3.26.10 | Adjustable inspiratory pause/plateau time (volume control ventilation): 0-30% of inspiratory time
3.26.11 | Peak Inspiratory Pressure: minimum range of: PEEP + 5 to 70 cm H2O.
3.26.12 | Positive End Expiratory Pressure (PEEP): Integrated, electronically controlled. Minimum range: 0 – 30 cm H2O.
3.26.13 | Tidal Volume: 20ml – 1500ml per breath. Minimum accuracy: 15%
3.26.14 | Peak Inspiratory Flow: At least 120 lpm
3.26.15 | The ventilator must be capable of flow or pressure triggering.
3.26.16 | Adjustable Inspiratory Rise time/Ramp with a range of 0.1 - 2 sec
3.26.17 | Compliance compensation/Tidal volume compensation/Compressible volume compensation or equivalent.
3.26.18 | Patient gas flow compensation must be automatic so that alterations in patient gas flow do not affect delivered tidal volumes
3.26.19 | In event of power or technical failure, manual ventilation must be possible.
3.27 | Display:
3.27.1 | All displays must be in acceptable English
3.27.2 | Colour display must include numerical and graphic display of parameters measured
3.27.3 | Graphic display to feature at least 3 waveforms: - pressure, flow and volume.
3.27.4 | Must also display flow-volume and pressure-volume loops with a reference loop
3.27.5 | Alarms : The ventilator must have the following user adjustable alarms
3.27.6 | Apnoea alarm
3.27.6.1 | High airway pressure alarm
3.27.6.2 | Continuous pressure alarm
3.27.6.3 | High and low expiratory tidal volume alarm
3.27.6.4 | Minute volume, high/low alarm
3.27.6.5 | Inspired oxygen (Fi02), high/low alarm
3.27.6.6 | PEEP, low alarm
3.27.6.7 | Leakage / Check breathing circuit alarm or equivalent
3.27.6.8 | Alarm profile must be user definable or default to patient appropriate settings
3.27.6.9 | The ventilator must have a mute alarm function, active only for a limited time (no more than 2 minutes).
3.27.6.10 | It must not be possible to cancel the visual alarm until the problem has been rectified.
3.27.6.11 | All components of the ventilator and breathing system in contact with the patient’s airway must be latex free.
3.27.7 | All components of the ventilator and breathing system in contact with the patient’s airway must be steam autoclavable up to 134ᴼ C
3.28 | Integrated Gas Monitoring
3.28.1 | Oxygen Monitoring
3.28.1.1 | The system must display both inspired and end-tidal oxygen concentration.
3.28.1.2 | The oxygen analyser must have user adjustable low and high alarms – visual and audible. The low alarm must automatically default to no less than 21%.
3.28.1.3 | The system must be accurate (+/- 2% in 21-100%) range, and have a rapid response time (≤ 500ms).
3.28.1.4 | The unit must sound a continuous alarm if the oxygen pressure falls below 190 kPa, then all nitrous oxide must be automatically cut from the patient circuit.
3.28.1.5 | Analyser preferrably of paramagnetic type. If fuel cell offered, provide full information on cost of disposables and frequency of replacement
4 | Patient Physiological Monitoring
4.1 | The patient monitor must be modular and suitable for use for neonatal, paediatric and adult patients.
4.2 | Must interface with the gas monitoring system if required
4.3 | The following paramaters must be monitored:
4.3.1 | - ECG: 3-lead and 5-lead. Must be capable of displaying 2 channels simultaneously.
4.3.2 | - ST-segment and arrhythmia analysis.
4.3.3 | - NIBP
4.3.4 | - Oxygen saturation (SpO2)
4.3.5 | - Plethysmography/respiratory rate
4.3.6 | - Temperature x 2 (minimum)
4.3.7 | - Invasive pressure x 2 (minimum)
4.4 | The patient monitor must be fully integrated with the workstation.
4.5 | Display must be no less than 400 mm diagonally.
4.6 | Minimum 6 waveforms must be able to be displayed simultaneously.
4.7 | Full colour display.
4.8 | The monitor must be capable of displaying trends for at least 24 hours in tabular and graphical format.
4.9 | The trend resolution must be user-adjustable and have a minimum resolution of 1 min.
4.10 | The front panel must be easy to clean and disinfect.
4.11 | The display must be flicker free and easy to read.
4.12 | Brightness control must be possible.
4.13 | In addition to the waveform display section, the monitor must have alphanumeric display of parameter identification, measured values, alarm and status messages.
4.13.1 | Audible and visual alarms must be available for all measured parameters.
4.13.2 | Default alarm limits if the limits are not set by the user
4.13.3 | Alarm indication: Visual and audible with mute
4.13.4 | Adjustable audio alarm volume
4.13.5 | Low battery: Alarm at least 10 minutes before battery power is exhausted
4.13.6 | It must be possible to mute the audible alarm for a short period (no longer than 2 minutes), but no override of the visual alarm must be possible.
4.13.7 | It must be possible to suspend the alarms of certain parameters during clinical procedures, but the alarm must reactivate automatically after a suitable time delay. Any silencing or disabling of alarms must be clearly indicated on the monitor.
4.14 | Input impedance of the system > 10 MΩ.
4.15 | Maximum allowable leakage current to the patient must be less than 10 µA.
4.16 | The monitor must be protected against defibrillator and electrosurgical interference.
4.17 | The monitor must be of low power consumption, and must not require fan or other forced ventilation.
4.18 | All cables and sockets must be colour coded for easy recognition, and must conform to international standards
4.19 | ECG
4.19.1 | A full set of 3/5 lead ECG cables must be included in the offer.
4.19.2 | The ECG amplifier must employ a differential input with high noise suppression.
4.19.3 | Sweep speed: at least 25 to 50 mm/sec.
4.19.4 | Freeze facility for ECG.
4.19.5 | The input impedance must be > 10 MΩ.
4.19.6 | The heart rate measuring range: 20 to 300 bpm, with an accuracy < 5 bpm.
4.19.7 | An 50 Hz filter must be provided.
4.19.8 | Alarm limits must be adjustable in the following range:
4.19.9 | Low: 40 to 150
4.19.10 | High: 70 to 250.
4.20 | Non Invasive Blood Pressure
4.20.1 | Must be suitable for adults, paeds and neonates
4.20.2 | The unit must use the oscillometric method of measurement.
4.20.3 | Modes: Manual and automatic
4.20.4 | Reusable neonate, paediatric, adult and obese cuffs must be supplied
4.20.5 | The cuffs must be robust, interchangeable and easy to clean.
4.20.6 | Must display systolic, diastolic and mean pressures for each measurement
4.20.7 | Adult: Systolic: 30 - 270 mm Hg, Diastolic: 20 - 240 mm Hg, MAP: 20 - 250 mm Hg
4.20.8 | Paeds: Systolic: 30 - 160 mm Hg, Diastolic: 20 - 120, MAP: 20 - 120 mm Hg
4.20.9 | Values must be displayed numerically
4.20.10 | Auto zero when unit is switched on
4.20.11 | The interval between measurements must be selectable from 1 – 120 minutes.
4.20.12 | Adjustable high and low alarm limits for systolic, diastolic, mean arterial pressure and heart rate.
4.20.13 | Over pressure protection: Automatic deflation at 300 mm Hg adult, 150 mm Hg neonate
4.20.14 | Average measurement time: approximately 30 s
4.20.15 | If measurement exceeds two minutes the cuff must deflate, an alarm must sound and likely faults indicated.
4.20.16 | The air hose must be at least two metres long.
4.21 | Oxygen Saturation.
4.21.1 | Continuous monitoring of oxygen saturation levels by means of the photometric principle.
4.21.2 | Must be suitable for adults, paeds and neonates
4.21.3 | The unit must continuously display the SaO2 % waveform
4.21.4 | Oxygen saturation (SaO2 %): Range of measurement: 0 - 100%
4.21.5 | The accuracy must be better than 2,5 % (70 - 100%) and 5% (50 - 70%).
4.21.6 | Pulse rate (beats/min): Range of measurement: 20 - 300 bpm
4.21.7 | The pulse rate accuracy must be better than ± 2 %.
4.21.8 | An audible bleep should sound for each heartbeat, the pitch of which should vary with the oxygen saturation.
4.21.9 | Automatic rejection of motion artefacts
4.22 | SpO2 Sensor:
4.22.1 | It must be the latest technology with clinical research support based on accuracy and real-time monitoring. Please submit the published peer reviewed/researched paper as supporting documents.
4.22.2 | The following reusable sensors must be provided:
4.22.2.1 | Finger sensor for adults and paeds.
4.22.2.2 | Earlobe sensor for adults and paeds.
4.22.2.3 | Finger wrap-around probe for neonates
4.22.3 | The sensor lead must be approximately 3 metres long.
4.23 | Temperature
4.23.1 | Temperature range must be approximately 15º to 45º C.
4.23.2 | Two reusable temperature probes must be included
4.24 | Circle System
4.24.1 | The circle absorber system must be integrated as part of the ventilation system.
4.24.2 | The circle absorber system must have an effective system to eliminate excess water created by the chemical reactions either by heating or water traps.
4.24.3 | The circle absorber system must be autoclavable at 134ºC.
4.24.4 | The circle absorber system must be supplied with a carbon dioxide absorber canister.
4.24.5 | The absorber canister should be reusable/refillable. Disposable cannisters should be quoted separately
4.24.6 | It must be easy to remove and replace the absorber canister while the workstation is in operation.
4.24.7 | Misalignment of the canister must be readily apparent to the user.
4.24.8 | The system must have a circle bypass facility while changing/refilling the absorber canister.
4.24.9 | The breathing circuit must contain an integrated adjustable pressure limiting (APL) valve with a minimum range of 0-70 cm H2O.
4.24.10 | The APL valve must have a tactile or electronic indicator when a value of between 20-30 cm H2O is exceeded.
4.24.11 | The circle absorber system must be latex free.
4.24.12 | The bag-to-ventilator switching mechanism must be clearly labeled and user-friendly. Mechanical or electronic
4.25 | Patient circuit pressure release valve
4.25.1 | Adjustable up to 70 cm water
4.25.2 | Audible alarm when the valve opens
4.25.3 | Manual override, non-lockable, to keep the release valve closed when the set pressure is exceeded.
4.26 | Flowmeter pressure release valve
4.26.1 | Situated between the flowmeters and the vaporisers
4.26.2 | Must open at 105 kPa +- 5 kPa
4.27 | Other System Requirements
4.27.1 | The unit must operate from both the 220 V, 50 Hz mains supply and internal rechargeable batteries
4.27.2 | The batteries must be charged while the unit is operating from the mains supply.
4.27.3 | The batteries must automatically take over in the event of a mains power failure.
4.27.4 | The batteries must have the capacity for full operation of the anaesthetic machine and ventilator for at least 90 minutes with batteries fully charged, under typical operating conditions.
4.27.5 | Graphic battery power indicator
4.27.6 | Recharge time < 8 hrs
4.27.7 | Mains cable at least four (4) metres long with SABS approved 15A three prong plug.
4.27.8 | The unit must have a rail for attaching a suction regulator.
4.28 | Data management:
4.28.1 | The workstation must be capable of downloading data to a USB device or memory card.
4.28.2 | The monitor must be able to communicate with a central station
4.28.3 | A printer must be incorporated in the unit
4.29 | The unit must be handed over in full operating order
5 | The following Accessories and Consumablesmust be supplied with the unit:
5.1 | Re-breathing bag, reusable, rubber
5.2 | Reservoir bag
5.3 | Pressure limiting valve, adjustable
5.4 | Oxygen fuel cell analyser with fuel cell
5.5 | Adult & Paediatric volume sensor re-usable
5.6 | Water trap for multigas monitor, re-usable (10 pcs.)
5.7 | A full set paediatric 3/5 lead ECG cables
5.8 | A full set adult 3/5 lead ECG cables
5.9 | Reusable paediatric cuffs
5.10 | Reusable adult cuffs
5.11 | Reusable neonate cuffs
5.12 | Reusable obese cuffs
5.13 | Air hoses
5.14 | Finger sensor for paeds
5.15 | Finger sensor for adults
5.16 | Earlobe sensor for paeds
5.17 | Earlobe sensor for adults
5.18 | Finger wrap-around probe for neonates
5.19 | Sensor lead
5.20 | Two reusable temperature probes
5.21 | State the cost and frequency of replacement for the flow sensors.
5.22 | State the cost and frequency of replacement for the fuel cell.
5.23 | Vaporizer key-fillers (Sevo & Iso)
5.24 | The unit must be supplied with two vaporizers (1x sevoflurane and 1x Iso)
5.25 | Anaesthesia circuit (Adult, single/double limb design) Disposable. 1.8m (Complete with Y-piece, Elbow, Catheter mount, Rebreathing bag)
5.26 | Anaesthesia circuit (Paediatric, single/double limb design) Disposable. 1.8m (Complete with Y-piece, Elbow, Catheter mount, Rebreathing bag)
5.27 | Anaesthesia circuit (neonatal, single/double limb design) Disposable. 1.8m (Complete with Y-piece, Elbow, Catheter mount, Rebreathing bag)
5.28 | Re-breathing circuit complete
5.29 | Non-re-breathing circuit complete
5.30 | Flow sensor (pack of 2)
5.31 | Spirometry line
5.32 | Sample line (pack of 10)
5.33 | Face masks: sizes 00, 0, 1, 2, 3, 4, 5.
5.34 | Bacterial/viral filter
5.35 | Y connector
5.36 | Manual ventilation hose
5.37 | Fresh gas tubing
5.38 | Seal for back-up cylinder connection on the unit
5.39 | Scavenger tubing (as per institution's scavenging outlet)
5.40 | Inspiratory and expiratory unidirectional valves
5.41 | Side stream CO2 kit (adult/paed)
5.42 | Disposable Adult Spirometry Set complete with spirometry tube, sample line and flow sensor, 3 m, box of 20 pieces
5.43 | Disposable Paediatric Spirometry Set complete with spirometry tube, sample line and flow sensor, 3 m, box of 20 pieces
5.44 | Gas sample tube with luer lock connection, 3m, disposable (10 pcs)
5.45 | Suction Tubing Set inc connector
5.46 | Resuscitator Adult, 1500mL, autoclavable, silicone
5.47 | Resuscitator Paediatric, 550mL, autoclavable, silicone
5.48 | Tubing set Paediatric for circle systems
6 | OPTIONAL ITEMS
6.1 | Prices incl VAT must be given for all optional items
6.2 | Brochures with full technical specifications must be supplied for each optional item offered
7 | Conformity Compliance (Please attach certificate)
7.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
7.2 | The unit must comply with electrical safety and performance standard such as IEC 60601-2-13 for Anaesthetic systems, attached certification
7.2.1 | Please provide unique ref number of the IEC ceriticate:
7.3 | OEM must comply and certified on ISO 9001 quality standards, attach proof of compliance
7.3.1 | Please provide unique ref number of the ISO 9001 ceriticate:
7.4 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
7.4.1 | Please provide unique ref number of the ISO 13485 ceriticate:
7.5 | Model quoted for must be EC certified. Attach a copy of certification
7.5.1 | Please provide unique ref number of the EC ceriticate:
7.6 | Hazardous Substance Act: If this type of equipment / apparatus appears on the schedule of Hazardous substances issued by the Directorate: Health Technology of the Department of Health, a license in terms of the Act on Hazardous Substances (Act 15/1973) must be submitted with the bid / quotation documents. The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
7.7 | All electrical/electronic medical equipment must be licensed by Radiation Control, where applicable a copy of the license must be submitted The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
7.7.1 | Radiation Control License number:
7.8 | The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade (rubber to with 2 screws) type and it must be a minimum length of three (3) meters. NB the mains cable of the unit tendered for must be SABS colour coded.
7.9 | The equipment quoted for must be protected against electromagnetic interference IEC 60601-1-2
7.10 | The unit must comply with electrical safety and performance standard such as IEC 60601-2-49 for Patient Monitoring, attached certification
7.10.1 | Please provide unique ref number of the IEC ceriticate:
8 | COMPREHENSIVE MAINTENANCE
8.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) must be included in the bid price
8.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries, Flow sensors and Oxygen sensor), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
8.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
8.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries, Flow sensors and Oxygen sensor), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
8.3 | The bidder to provide the OEM recommended service interval of the unit:
8.4 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
9 | General
9.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
9.2 | The unit on offer must be the latest technology.
ANAESTHESIA DELIVERY SYSTEM: STANDARD with integrated physiological monitor, Replacement Accessories
RT2-01-04a | A full set paediatric 3/5 lead ECG cable
RT2-01-04b | A full set adult 3/5 lead ECG cable
RT2-01-04c | Reusable paediatric cuff
RT2-01-04d | Reusable adult cuff
RT2-01-04e | Reusable neonate cuff
RT2-01-04f | Reusable obese cuff
RT2-01-04g | Air hose
RT2-01-04h | Finger sensor for paeds
RT2-01-04i | Finger sensor for adults
RT2-01-04j | Earlobe sensor for paeds
RT2-01-04k | Earlobe sensor for adults
RT2-01-04l | Finger wrap-around probe for neonates
RT2-01-04m | Sensor lead
RT2-01-04n | Reusable temperature probe
Ventilator modes- UPGRADES
RT2-01-04o | Pressure Support with Mandatory Apnoea back-up ventilation
RT2-01-04p | Volume Controlled Ventilation + SIMV + Pressure Support
RT2-01-04q | Pressure Control Ventilation + SIMV + Presssure Support
RT2-01-04r | CPAP
Bidder to list any other Consumables or Accessories (Not listed above) for the model on offer below

## Anaesthetic, Advanced
RT2-01-05 | ANAESTHESIA DELIVERY SYSTEM: Advanced without physiological monitor
2 | Product Offered (Anaesthetic Machine)
2.1 | Make:
2.2 | Model:
2.3 | Initial Manufacturing year of Model:
2.4 | Radiation Control License number:
3 | Technical Specification
3.1 | An advanced anaesthesia delivery system comprising a completely integrated electronic anaesthesia workstation.
3.2 | The unit must include electronicaly integrated oxygen, nitrous oxide and air delivery and monitoring.
3.3 | Ventilation and anaesthesia vapour monitoring must be fully integrated within the unit.
3.4 | The anaesthesia unit must be able to communicate/interface with a physiological monitor
3.5 | The anaesthesia system must comprise:
3.5.1 | Anaesthesia system suitable for neonatal (down to 500g), paediatric and adult anaesthesia.
3.5.2 | Ventilator suitable for ventilating neonatal (down to 500g), paediatric, adult and morbid/obese patients
3.6 | Multigas analyser:
3.6.1 | Oxygen
3.6.2 | Nitrous oxide
3.6.3 | Carbon dioxide
3.7 | Anaesthetic vapour monitoring:
3.7.1 | Isoflurane
3.7.2 | Sevoflurane
3.7.3 | Desflurane
3.8 | Complete and automated self test requiring very limited user intervention.
3.9 | The self test must provide the measurements and values for automatic leakage and compliance.
3.10 | Automatic calibration of sensors for anaesthetic gases and vapours must form part of the self test and require very limited user intervention.
3.11 | It must be possible to override the self test facility. The system must display the override status
3.12 | The unit must provide conventional, low-flow (down to 200 ml/min) anaesthesia.
3.13 | The unit must return analysed gas back to the circle system
3.14 | A safety oxygen function is required for the delivery of oxygen through the anaesthesia unit in the event of an electronic failure of the gas delivery system.
3.15 | The unit must be supplied with a minimum of 2 storage drawer
3.16 | The unit must be supplied with a minimum of 4 additional 220 V power sockets. These sockets must be excluded from the battery backup circuit in case of a power failure
3.17 | The unit must have a working area or a retractable shelf. Width: not less than 300 mm. Depth: not less than 200 mm
3.18 | The unit must have a light to enable the anaesthetist to work in a darkened theatre.
3.20 | Trolley
3.20.1 | The unit must be mounted on a strong , robust and compact trolley manufactured from corrosion free material.
3.20.2 | The trolley must have no sharp corners and be easy to clean.
3.20.3 | The trolley must have at least four antistatic castors of diameter 100 mm or greater.
3.20.4 | Trolley must have two lockable castors or a central locking brake.
3.20.5 | The unit must accommodate an oxygen cylinder. Cylinder size according to user requirement. Size to be compatible with current state vendor of oxygen.
3.20.6 | Trolley must have accessible handles for ease of movement
3.20.7 | The trolley, with all components mounted, must be stable without toppling over when placed on a 30° slope.
3.20.8 | The unit must be supplied with a castor guard system, preventing the entaglement and trapping of cables within the wheels.
3.20.9 | The trolley must have a mounting system suitable for a physiologial monitor.
3.21 | Pipeline and Cylinder Gas Supplies
3.21.1 | The anaesthesia system must be designed to use external piped medical gasses (O2, N2O and medical air) and vacuum from medical service outlet sockets.
3.21.2 | Connection hoses for above must be provided
3.21.3 | Hoses must be fitted with Heyer type non-interchangeable probes for the existing medical gas outlets as specified by the purchaser.
3.21.4 | The medical air hose connection or CO2 circle system must include a water trap
3.21.5 | A safety pressure relief valve of 60 - 80 cm H2O must be provided to protect the system from excessive pressure .
3.21.6 | The unit must function safely and correctly within the pipeline supply pressure range of: 270 - 600 kPa
3.21.7 | The system must preferentially utilise the piped oxygen supply over the cylinder supply, if pipeline pressure is > 270 kPa
3.21.8 | Connection hoses must be:
3.21.8.1 | Flexible and at least 6 metres long
3.21.8.2 | Antistatic
3.21.9 | Colour coded as follows:
3.21.9.1 | O2: white
3.21.9.2 | N2O: blue
3.21.9.3 | Vacuum: yellow
3.21.9.4 | Medical air: black
3.21.9.5 | Scavenge low vacuum: orange or yellow
3.21.10 | Where relevant, fitted with outlet probes complying with SANS 1409:2014
3.21.11 | Permanently bonded to the outlet probe and the machine connector
3.21.12 | Hoses for gases above atmospheric pressure, must designed for pressures of at least 1 000 kPa.
3.21.13 | Vacuum hoses may not collapse or kink at 90 kPa below atmospheric pressure, when subjected to 180° bend of inside radius equal to the outside diameter of the hose.
3.21.14 | All connection hoses must be fed to a central manifold
3.21.15 | All high pressure gas and vacuum pipe connections (other than hoses) and all pipes connected to pressure regulators and pressure gauges must be stainless steel (18/8 grade) pipes that are oil-free and colour coded as for hoses (see above) over their entire lengths
3.21.16 | An in-line non-return valve must be fitted to each gas pipeline on the high pressure (inlet) line
3.21.17 | An in-line non-return valve must be fitted to each gas cylinder yoke
3.21.18 | A replaceable gas filter, to block particles greater than 65 µm, must be fitted on the upstream side of the non-return valve
3.21.19 | Each gas pipeline inlet must have a pressure gauge or electronic display of pressure. The pressure range (415 kPa ± 75 kPa) must be clearly indicated in colour
3.21.20 | Must be fitted with a gas cylinder yoke to accommodate one oxygen gas cylinder with a pin index system
3.21.21 | Each gas cylinder yoke must have a pressure gauge or electronic display of pressure
3.21.22 | The pressure gauges or electronic displays of pressure for both pipeline & cylinder supplies, calibrated in kPa for each line, must be labelled and visible from the front of the unit.
3.21.23 | O2 and N2O regulators must be provided complying with CKS 605:1987.
3.21.24 | A pressure regulator complying with CKS 605 must be supplied for each gas inlet and must reduce the gas pressure to 415 kPa without freezing up, creeping, chattering or whistling
3.21.25 | Outlet sockets and probes must comply with SANS 1409:2014.
3.21.26 | Sockets with hoses must be colour-coded as per SANS 1409:2014:
3.21.26.1 | Oxygen: White background with black letters
3.21.26.2 | Nitrous oxide: Blue background with white letters
3.21.26.3 | Medical air: One half black with white letters, the other half white with black letters
3.21.26.4 | Vacuum: Yellow background with black letters
3.21.27 | Continuous display, visible from the front of the machine, of both pipeline and cylinder pressures even during power failure
3.21.28 | Pipeline and cylinder pressures must be differentially displayed, and clearly indicated
3.21.29 | Safety pressure relief valves must protect the entire pneumatic system from excessive input pressures from pipeline and cylinder supplies.
3.21.30 | When a cylinder is detached from the pin-indexed yoke no gas leak must occur at the yoke.
3.21.31 | Airway pressure must be monitored and displayed
3.21.32 | Safety relief valve for all gases must be supplied and set at 600kPa
3.22 | Scavenging
3.22.1 | The unit must have a scavenging system with a hose at least 5 metres long with the suitable connector as per the institutions scavenging installation.
3.22.2 | Air-break device to prevent negative pressure on the breathing system
3.22.3 | If an active sytem is supplied a pressure indicator must be present
3.22.4 | The scavenging outlet must be a 30mm ISO connector.
3.23 | Patient Gas Delivery System
3.23.1 | Patient gas delivery must be electronically controlled
3.23.2 | Must display the delivered patient gas flow using an ISO colour standard for each gas.
3.23.3 | The unit must be able to deliver a total patient gas flow of 200 ml/min to 18 l/min (Minimum). The unit must have the ability to deliver at least 10 l/min of oxygen. These must be incremental adjustments
3.23.4 | Gas supply must conform to:
3.23.5 | O2 minimum flow range of 200 ml/min to 10 l/min. Accuracy ± 5%
3.23.6 | medical air minimum flow range 0 to 10 l/min. Accuracy ± 5%
3.23.7 | nitrous oxide minimum flow range 0 to 10 l/min. Accuracy ± 5%
3.23.8 | Flowmeter sequence to be according to standards approved by the South African Society of Anaesthesiologists - Oxygen on the right
3.23.9 | The delivery system must prevent simultaneous delivery of two anaesthetic vapours.
3.23.10 | A separate common gas outlet to accept a standard 22 mm outside diameter and 15 mm internal diameter.
3.24 | Hypoxic guard
3.24.1 | The unit must include a safety mechanism that ensures a minimum concentration of 25% oxygen in any oxygen/nitrous oxide gas mixture across all flow rates.
3.24.2 | The unit must include a safety mechanism that ensures a minimum concentration of 21% oxygen in any oxygen/air gas mixture across all flow rates.
3.24.3 | The unit must include a high-flow oxygen flush mechanism with a flow rate of at least 35 l/min
3.24.4 | Oxygen flush must be non-lockable, recessed and self closing
3.24.5 | This system must bypass the flowmeters and vaporisers
3.25 | Breathing System
3.25.1 | Total circuit volume, including absorber, must not exceed 2,7 litres
3.25.2 | With APL valve 5 - 75 cm H2O
3.25.3 | With ventilator to manual switch. This switch when directed to the ventilator must isolate the APL valve from the circle
3.25.4 | Breathing system must be fully autoclavable to 134°C
3.25.5 | There must be a mechanism to remove water from the breathing system
3.25.6 | Breathing system must be latex free
3.25.7 | Bag-to-ventilator switch automatically begins in the ventilator position
3.25.8 | Reusable corrugated breathing circuit adequate for both adult and paediatric use to be supplied
3.25.9 | Bidders to supply the breathing circuit as required by the user
3.26 | Integrated Ventilator
3.26.1 | Ventilator suitable for ventilating adults, paediatric and neonates (> 500g) patients without changing ventilator components (e.g. bellows)
3.26.2 | The ventilator must be electronically controlled
3.26.3 | The ventilation technology offered may be turbine, volume reflector, piston or ascending bellows
3.26.4 | The minimum basic modes to be supplied:
3.26.4.1 | Manual/Spontaneous Ventilation
3.26.4.2 | Volume Control Ventilation
3.26.4.3 | Pressure Control Ventilation
3.26.4.4 | Pressure Support with Mandatory Apnoea back-up ventilation
3.26.4.5 | Volume Controlled Ventilation + SIMV + Pressure Support
3.26.4.6 | Pressure Control Ventilation + SIMV + Presssure Support
3.26.4.7 | CPAP
3.26.5 | Additional modes of ventilation may be offered and quoted for separately
3.26.6 | The ventilator must be integrated with the workstation and be user friendly
3.26.7 | Frequency: minimum range of 5 – 60 bpm.
3.26.8 | Inspiratory time: minimum range of 0.2 – 10.0 seconds.
3.26.9 | Minimum Inspiration/Expiration Ratio: 4:1 – 1:4.
3.26.10 | Adjustable inspiratory pause/plateau time (volume control ventilation): 0-30% of inspiratory time
3.26.11 | Peak Inspiratory Pressure: minimum range of: PEEP + 5 to 70 cm H2O.
3.26.12 | Positive End Expiratory Pressure (PEEP): Integrated, electronically controlled. Minimum range: 0 – 30 cm H2O.
3.26.13 | Tidal Volume: 20ml – 1500ml per breath. Minimum accuracy: 15%
3.26.14 | Peak Inspiratory Flow: At least 90 lpm
3.26.15 | The ventilator must be capable of flow or pressure triggering.
3.26.16 | Adjustable Inspiratory Rise time/Ramp with a range of 0.1 - 2 sec
3.26.17 | Compliance compensation/Tidal volume compensation/Compressible volume compensation or equivalent.
3.26.18 | Patient gas flow compensation must be automatic so that alterations in patient gas flow does not affect delivered tidal volumes
3.26.19 | It must be possible to set CPAP in manual/spontaneous mode within a range of: 1 to 10 cm H2O
3.26.20 | In event of power or technical failure, manual ventilation must be possible.
3.27 | Display:
3.27.1 | All displays must be in acceptable English
3.27.2 | Colour display must include numerical and graphic display of parameters measured
3.27.3 | Graphic display to feature at least 3 waveforms: - pressure, flow and volume.
3.27.4 | Must also display flow-volume and pressure-volume loops with a reference loop
3.27.5 | Alarms : The ventilator must have the following user adjustable alarms
3.27.5.1 | Apnoea alarm
3.27.5.1 | High airway pressure alarm
3.27.5.1 | Continuous pressure alarm
3.27.5.1 | High and low expiratory tidal volume alarm
3.27.5.1 | Minute volume, high/low alarm
3.27.5.1 | Inspired oxygen (Fi02), high/low alarm
3.27.5.1 | PEEP, low alarm
3.27.5.1 | Leakage / Check breathing circuit alarm or equivalent
3.27.5.1 | Alarm profile must be user definable or default to patient appropriate settings
3.27.5.1 | The ventilator must have a mute alarm function, active only for a limited time.
3.27.5.1 | It must not be possible to cancel the visual alarm until the problem has been rectified.
3.27.6 | All components of the ventilator and breathing system in contact with the patient’s airway must be latex free.
3.27.7 | All components of the ventilator and breathing system in contact with the patient’s airway must be steam autoclavable up to 134ᴼ C
3.28 | Integrated Gas Monitoring
3.28.1 | Oxygen Monitoring
3.28.1.1 | The system must display both inspired and end-tidal oxygen concentration.
3.28.1.2 | The oxygen analyser must have user adjustable low and high alarms – visual and audible. The low alarm must automatically default to no less than 21%.
3.28.1.3 | The system must be accurate (+/- 2% in 21-100%) range, and have a rapid response time (≤ 500ms).
3.28.1.4 | The unit must sound a continuous alarm if the oxygen pressure falls below 190 kPa, then all nitrous oxide must be automatically cut from the patient circuit.
3.28.1.5 | Bidder to give details of the oxygen analyser
3.28.1.6 | The lifespan of the oxygen analyser must be guaranteed for at least 10 years.
3.28.1.7 | A system using replaceable oxygen fuel cells will not be accepted
3.28.2 | Anaesthetic vapour, carbon dioxide and nitrous oxide monitoring:
3.28.2.1 | Anaesthetic vapour, N2O monitoring and capnography must be integrated into the anaesthetic workstation.
3.28.2.2 | The monitoring must be of the side-stream type.
3.28.2.3 | Sampling: 100 - 200 ml/min
3.28.2.4 | Sampled gases must be internally returned to the breathing system.
3.28.2.5 | CO2 must be presented as a numerical display of inspired and end-tidal concentration and a capnography waveform.The displayed units must be user-selectable: kPa or mmHg.
3.28.2.6 | The monitoring system must have automatic identification of anaesthetic vapours.
3.28.2.7 | The system must have a graphical display indicating the type of anaesthetic vapour/s currently in use and the concentration thereof. The colour coding must be as per international colour coding for anaesthetic vapours.
3.28.2.8 | The anaestheric vapour must be displayed as vol % and must include inspiratory and expiratory concentrations.
3.28.2.9 | The monitoring system must be able to simultaneously detect, identify and display mixed anaesthetic vapours.
3.28.2.10 | High and low anaesthetic vapour alarms must be present and these must be user-definable.
3.28.2.11 | There must be automatic calculation of minimum alveolar concentration.
3.28.2.12 | The system must be able to calculate anaesthetic gas utilization in ml per case
3.28.2.13 | The unit must have an indication of the level of vapour in the vaporiser
3.28.2.14 | Integrated vapour analysis must deliver an audible "low vapour " alarm and a critical alarm for an empty tank
3.29 | Circle System
3.29.1 | The circle absorber system must be integrated as part of the ventilation system.
3.29.2 | The circle absorber system must have an effective system to eliminate excess water created by the chemical reactions either by heating or water traps.
3.29.3 | The circle absorber system must be autoclavable at 134ºC.
3.29.4 | The circle absorber system must be supplied with a carbon dioxide absorber canister.
3.29.5 | The absorber canister may be reusable/refillable
3.29.6 | It must be easy to remove and replace the absorber canister while the workstation is in operation.
3.29.7 | Misalignment of the canister must be readily apparent to the user.
3.29.8 | The system must have a circle bypass facility while changing/refilling the absorber canister.
3.29.9 | The breathing circuit must contain an integrated adjustable pressure limiting (APL) valve with a minimum range of 0-70 cm H2O.
3.29.10 | The APL valve must have a tactile or electronic indicator when a value of between 20-30 cm H2O is exceeded.
3.29.11 | The circle absorber system must be latex free.
3.29.12 | The bag-to-ventilator switching mechanism must be clearly labeled and user-friendly. Mechanical or electronic
3.30 | Other system requirements
3.30.1 | The unit must operate from both the 220 V, 50 Hz mains supply and internal rechargeable batteries
3.30.2 | The batteries must be charged while the unit is operating from the mains supply.
3.30.3 | The batteries must automatically take over in the event of a mains power failure.
3.30.4 | The batteries must have the capacity for full operation of the anaesthetic machine and ventilator for at least 90 minutes with batteries fully charged, under typical operating conditions.
3.30.5 | Graphic battery power indicator
3.30.6 | Recharge time < 8 hrs
3.30.7 | Mains cable at least four metres long with SABS approved 15A three prong plug.
3.30.8 | The unit must have a rail for attaching a suction regulator.
3.31 | Data management: The workstation must be capable of downloading data to a USB device or memory card.
3.31.1 | The unit must be handed over in full operating order
4 | The following Accessories and Consumables must be supplied with the unit:
4.1 | Re-breathing bag, reusable, rubber
4.2 | Reservoir bag
4.3 | Pressure limiting valve, adjustable
4.4 | Adult & Paediatric volume sensor re-usable
4.5 | Water trap for multigas monitor, re-usable (10 pcs.)
4.6 | Vaporizer key-fillers (Sevo, & other choice below)
4.7 | The unit must be supplied with two vaporizers (one must sevoflurane, the other must be of the choice of the buying institution)
4.8 | Anaesthesia circuit (Adult, single/double limb design) Disposable. 1.8m (Complete with Y-piece, Elbow, Catheter mount, Rebreathing bag)
4.9 | Anaesthesia circuit (Paediatric, single/double limb design) Disposable. 1.8m (Complete with Y-piece, Elbow, Catheter mount, Rebreathing bag)
4.10 | Anaesthesia circuit (neonatal, single/double limb design) Disposable. 1.8m (Complete with Y-piece, Elbow, Catheter mount, Rebreathing bag)
4.11 | Re-breathing circuit complete
4.12 | Non-re-breathing circuit complete
4.13 | Flow sensor (pack of 2)
4.14 | Spirometry line
4.15 | Sample line (pack of 10)
4.16 | Face masks: sizes 00, 0, 1, 2, 3, 4, 5.
4.17 | Bacterial/viral filter
4.18 | Y connector
4.19 | Manual ventilation hose
4.20 | Fresh gas tubing
4.21 | Seal for back-up cylinder connection on the unit
4.22 | Scavenger tubing
4.23 | Inspiratory and expiratory unidirectional valves
4.24 | Side stream CO2 kit (adult/paed)
4.25 | Disposable Adult Spirometry Set complete with spirometry tube, sample line and flow sensor, 3 m, box of 20 pieces
4.26 | Disposable Paediatric Spirometry Set complete with spirometry tube, sample line and flow sensor, 3 m, box of 20 pieces
4.27 | Gas sample tube with luer lock connection, 3m, disposable (10 pcs)
4.28 | Suction Tubing Set inc connector
4.29 | Resuscitator Adult, 1500mL, autoclavable, silicone
4.30 | Resuscitator Paediatric, 550mL, autoclavable, silicone
4.31 | Tubing set Paediatric for circle systems
5 | OPTIONAL ITEMS: Any accessories and modules may be offered as optional items
5.1 | Prices incl VAT must be given for all optional items
5.2 | Brochures with full technical specifications must be supplied for each optional item offered
6 | Conformity Compliance (Please attach certificate)
6.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
6.2 | The unit must comply with electrical safety and performance standard such as IEC 60601-2-13 for Anaesthetic systems, attached certification
6.2.1 | Please provide unique ref number of the IEC ceriticate:
6.3 | OEM must comply and certified on ISO 9001 quality standards, attach proof of compliance
6.3.1 | Please provide unique ref number of the ISO 9001 ceriticate:
6.4 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
6.4.1 | Please provide unique ref number of the ISO 13485 ceriticate:
6.5 | Model quoted for must be EC certified. Attach a copy of certification
6.5.1 | Please provide unique ref number of the EC ceriticate:
6.6 | Hazardous Substance Act: If this type of equipment / apparatus appears on the schedule of Hazardous substances issued by the Directorate: Health Technology of the Department of Health, a license in terms of the Act on Hazardous Substances (Act 15/1973) must be submitted with the bid / quotation documents. The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
6.7 | All electrical/electronic medical equipment must be licensed by Radiation Control, where applicable a copy of the license must be submitted The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
6.7.1 | Radiation Control License number:
6.8 | The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade (rubber to with 2 screws) type and it must be a minimum length of three (3) meters. NB the mains cable of the unit tendered for must be SABS colour coded.
6.9 | The equipment quoted for must be protected against electromagnetic interference IEC 60601-1-2
7 | COMPREHENSIVE MAINTENANCE
7.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) must be included in the bid price
7.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries, Flow sensors and Oxygen sensor), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
7.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
7.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries, Flow sensors and Oxygen sensor), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
7.3 | The bidder to provide the OEM recommended service interval of the unit:
7.4 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
8 | General
8.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
8.2 | The unit on offer must be the latest technology.
ANAESTHESIA DELIVERY SYSTEM: ADVANCED WITHOUT PHYSIOLOGICAL MONITOR, REPLACEMENT ACCESSORIES
RT2-01-05a | Anaesthesia circuit (Adult, single/double limb design) Disposable. 1.8m (Complete with Y-piece, Elbow, Catheter mount, Rebreathing bag)
RT2-01-05b | Anaesthesia circuit (Paediatric, single/double limb design) Disposable. 1.8m (Complete with Y-piece, Elbow, Catheter mount, Rebreathing bag)
RT2-01-05c | Anaesthesia circuit (neonatal, single/double limb design) Disposable. 1.8m (Complete with Y-piece, Elbow, Catheter mount, Rebreathing bag)
RT2-01-05d | Re-breathing circuit complete
RT2-01-05e | Non-re-breathing circuit complete
RT2-01-05f | Flow sensor (pack of 2)
RT2-01-05g | Spirometry line
RT2-01-05h | Sample line (pack of 10)
RT2-01-05i | Face masks: sizes (00,0,1,2,3,4, 5)- End-user to specify size
RT2-01-05j | Bacterial/viral filter
RT2-01-05k | Y connector
RT2-01-05l | Manual ventilation hose
RT2-01-05m | Fresh gas tubing
RT2-01-05n | Seal for back-up cylinder connection on the unit
RT2-01-05o | Scavenger tubing (as per institution's scavenging outlet)
RT2-01-05p | Inspiratory and expiratory unidirectional valves
RT2-01-05q | Side stream CO2 kit (adult/paed)
RT2-01-05r | Disposable Adult Spirometry Set complete with spirometry tube, sample line and flow sensor, 3 m, box of 20 pieces
RT2-01-05s | Disposable Paediatric Spirometry Set complete with spirometry tube, sample line and flow sensor, 3 m, box of 20 pieces
RT2-01-05t | Gas sample tube with luer lock connection, 3m, disposable (10 pcs)
RT2-01-05u | Suction Tubing Set inc connector
RT2-01-05v | Resuscitator Adult, 1500mL, autoclavable, silicone
RT2-01-05w | Resuscitator Paediatric, 550mL, autoclavable, silicone
RT2-01-05x | Tubing set Paediatric for circle systems
RT2-01-05y | sevoflurane vaporizer
RT2-01-05z | Isoflurane vaporizer
RT2-01-05zz | Desflurane vaporizer
Bidder to list any other Consumables or Accessories (Not listed above) for the model on offer below

## Anaesthetic, Advanced + Monitor
RT2-01-06 | ANAESTHESIA DELIVERY SYSTEM: Advanced with integrated physiological monitor
1 | Description
1a | Product Offered (Anaesthetic Machine)
1a.1 | Make:
1a.2 | Model:
1a.3 | Initial Manufacturing year of Model:
1a.4 | Radiation Control License number:
1b | Product Offered (Monitor)
1b.1 | Make:
1b.2 | Model:
1b.3 | Initial Manufacturing year of Model:
1b.4 | Radiation Control License number:
3 | Technical Specification
3.1 | An advanced anaesthesia delivery system comprising a completely integrated electronic anaesthesia workstation is required.
3.2 | The unit must include electronicaly integrated oxygen, nitrous oxide and air delivery and monitoring.
3.3 | Ventilation and anaesthesia vapour monitoring must be fully integrated within the unit.
3.4 | The patient physiological monitoring must communicate/interface with the anaesthesia unit.
3.5 | The anaesthesia system must comprise:
3.5.1 | Anaesthesia system suitable for neonatal (down to 500g), paediatric and adult anaesthesia.
3.5.2 | Ventilator suitable for ventilating neonatal (down to 500g), paediatric, adult and morbid/obese patients
3.6 | Physiological monitor
3.7 | Multigas analyser:
3.7.1 | Oxygen
3.7.2 | Nitrous oxide
3.7.3 | Carbon dioxide
3.8 | Anaesthetic vapour monitoring:
3.8.1 | Isoflurane
3.8.2 | Sevoflurane
3.8.3 | Desflurane
3.9 | Complete and automated self test requiring very limited user intervention.
3.10 | The self test must provide the measurements and values for automatic leakage and compliance.
3.11 | Automatic calibration of sensors for anaesthetic gases and vapours must form part of the self test and require very limited user intervention.
3.12 | It must be possible to override the self test facility. The system must display the override status
3.13 | The unit must provide conventional, low-flow (down to 200 ml/min) anaesthesia.
3.14 | The unit must return analysed gas back to the circle system
3.15 | A safety oxygen function is required for the delivery of oxygen through the anaesthesia unit in the event of an electronic failure of the gas delivery system.
3.16 | The patient monitor must be fixed to the anaesthesia device with a suitable mounting system or integrated into the device.
3.17 | The unit must be supplied with a minimum of 2 storage drawer
3.18 | The unit must be supplied with a minimum of 4 additional 220 V power sockets. These sockets must be excluded from the battery backup circuit in case of a power failure
3.19 | The unit must have a working area or a retractable shelf. Width: not less than 300 mm. Depth: not less than 200 mm
3.20 | The unit must have a light to enable the anaesthetist to work in a darkened theatre.
3.21 | Trolley
3.21.1 | The unit must be mounted on a strong , robust and compact trolley manufactured from corrosion free material.
3.21.2 | The trolley must have no sharp corners and be easy to clean.
3.21.3 | The trolley must have at least four antistatic castors of diameter 100 mm or greater.
3.21.4 | Trolley must have two lockable castors or a central locking brake.
3.21.5 | The unit must accommodate an oxygen cylinder. Cylinder size according to user requirement. Size to be compatible with current state vendor of oxygen.
3.21.6 | Trolley must have accessible handles for ease of movement
3.21.7 | The trolley, with all components mounted, must be stable without toppling over when placed on a 30° slope.
3.21.8 | The unit must be supplied with a castor guard system, preventing the entaglement and trapping of cables within the wheels.
3.22 | Pipeline and Cylinder Gas Supplies
3.22.1 | The anaesthesia system must be designed to use external piped medical gasses (O2, N2O and medical air) and vacuum from medical service outlet sockets.
3.22.2 | Connection hoses for above must be provided
3.22.3 | Hoses must be fitted with Heyer type non-interchangeable probes for the existing medical gas outlets as specified by the purchaser.
3.22.4 | The medical air hose connection or CO2 circle system must include a water trap
3.22.5 | A safety pressure relief valve of 60 - 80 cm H2O must be provided to protect the system from excessive pressure .
3.22.6 | The unit must function safely and correctly within the pipeline supply pressure range of: 270 - 600 kPa
3.22.7 | The system must preferentially utilise the piped oxygen supply over the cylinder supply, if pipeline pressure is > 270 kPa
3.22.8 | Connection hoses must be:
3.22.8.1 | Flexible and at least 6 metres long
3.22.8.2 | Antistatic
3.22.9 | Colour coded as follows:
3.22.9.1 | O2: white
3.22.9.2 | N2O: blue
3.22.9.3 | Vacuum: yellow
3.22.9.4 | Medical air: black/white
3.22.9.5 | Scavenge low vacuum: orange or yellow
3.22.10 | Where relevant, fitted with outlet probes complying with SANS 1409:2014
3.22.11 | Permanently bonded to the outlet probe and the machine connector
3.22.12 | Hoses for gases above atmospheric pressure, must designed for pressures of at least 1 000 kPa.
3.22.13 | Vacuum hoses may not collapse or kink at 90 kPa below atmospheric pressure, when subjected to 180° bend of inside radius equal to the outside diameter of the hose.
3.22.14 | All connection hoses must be fed to a central manifold
3.22.15 | All high pressure gas and vacuum pipe connections (other than hoses) and all pipes connected to pressure regulators and pressure gauges must be stainless steel (18/8 grade) pipes that are oil-free and colour coded as for hoses (see above) over their entire lengths
3.22.16 | An in-line non-return valve must be fitted to each gas pipeline on the high pressure (inlet) line
3.22.17 | An in-line non-return valve must be fitted to each gas cylinder yoke
3.22.18 | A replaceable gas filter, to block particles greater than 65 µm, must be fitted on the upstream side of the non-return valve
3.22.19 | Each gas pipeline inlet must have a pressure gauge or electronic display of pressure. The pressure range (415 kPa ± 75 kPa) must be clearly indicated in colour
3.22.20 | Must be fitted with a gas cylinder yoke to accommodate one oxygen gas cylinder with a pin index system
3.22.21 | Each gas cylinder yoke must have a pressure gauge or electronic display of pressure
3.22.22 | The pressure gauges or electronic displays of pressure for both pipeline & cylinder supplies, calibrated in kPa for each line, must be labelled and visible from the front of the unit.
3.22.23 | O2 and N2O regulators must be provided complying with CKS 605:1987.
3.22.24 | A pressure regulator complying with CKS 605 must be supplied for each gas inlet and must reduce the gas pressure to 415 kPa without freezing up, creeping, chattering or whistling
3.22.25 | Outlet sockets and probes must comply with SANS 1409:2014.
3.22.26 | Sockets with hoses must be colour-coded as per SANS 1409:2014:
3.22.26.1 | Oxygen: White background with black letters
3.22.26.2 | Nitrous oxide: Blue background with white letters
3.22.26.3 | Medical air: One half black with white letters, the other half white with black letters
3.22.26.4 | Vacuum: Yellow background with black letters
3.22.27 | Continuous display, visible from the front of the machine, of both pipeline and cylinder pressures even during power failure
3.22.28 | Pipeline and cylinder pressures must be differentially displayed, and clearly indicated
3.22.29 | Safety pressure relief valves must protect the entire pneumatic system from excessive input pressures from pipeline and cylinder supplies.
3.22.30 | When a cylinder is detached from the pin-indexed yoke no gas leak must occur at the yoke.
3.22.31 | Airway pressure must be monitored and displayed
3.22.32 | Safety relief valve for all gases must be supplied and set at 600kPa
3.23 | Scavenging
3.23.1 | The unit must have a scavenging system with a hose at least 5 metres long with the suitable connector as per the institutions scavenging installation.
3.23.2 | Air-break device to prevent negative pressure on the breathing system
3.23.3 | If an active sytem is supplied a pressure indicator must be present
3.23.4 | The scavenging outlet must be a 30mm ISO connector.
3.24 | Patient Gas Delivery System
3.24.1 | Patient gas delivery must be electronically controlled
3.24.2 | and display the delivered patient gas flow using an ISO colour standard for each gas.
3.24.3 | The unit must be able to deliver a total patient gas flow of 200 ml/min to 18 l/min (Minimum). The unit must have the ability to deliver at least 10 l/min of oxygen. These must be incremental adjustments
3.24.4 | Gas supply must conform to:
3.24.4.1 | O2 minimum flow range of 200 ml/min to 10 l/min. Accuracy ± 5%
3.24.4.2 | medical air minimum flow range 0 to 10 l/min. Accuracy ± 5%
3.24.4.3 | nitrous oxide minimum flow range 0 to 10 l/min. Accuracy ± 5%
3.24.5 | Flowmeter sequence to be according to standards approved by the South Africal Society of Anaesthesiologists - Oxygen on the right
3.24.6 | The delivery system must prevent simultaneous delivery of two anaesthetic vapours.
3.24.7 | A separate common gas outlet to accept a standard 22 mm outside diameter and 15 mm internal diameter.
3.25 | Hypoxic guard
3.25.1 | The unit must include a safety mechanism that ensures a minimum concentration of 25% oxygen in any oxygen/nitrous oxide gas mixture across all flow rates.
3.25.2 | The unit must include a safety mechanism that ensures a minimum concentration of 21% oxygen in any oxygen/air gas mixture across all flow rates.
3.25.3 | The unit must include a high-flow oxygen flush mechanism with a flow rate of at least 35 l/min
3.25.4 | Oxygen flush must be non-lockable, recessed and self closing
3.25.5 | This system must bypass the flowmeters and vaporisers
3.26 | Breathing System
3.26.1 | Total circuit volume, including absorber, must not exceed 2,7 litres
3.26.2 | With APL valve 5 - 75 cm H2O
3.26.3 | With ventilator to manual switch. This switch when directed to the ventilator must isolate the APL valve from the circle
3.26.4 | There must be a mechanism to remove water from the breathing system
3.26.5 | Breathing system must be fully autoclavable to 134°C
3.26.6 | Breathing system must be latex free
3.26.7 | Bag-to-ventilator switch automatically begins in the ventilator position
3.26.8 | Reusable corrugated breathing circuit adequate for both adult and paediatric use to be supplied
3.26.9 | Bidders to supply the breathing circuit as required by the user
3.27 | Integrated Ventilator
3.27.1 | Ventilator suitable for ventilating adults, paediatric and neonates (> 500g) patients without changing ventilator components (e.g. bellows)
3.27.2 | The ventilator must be electronically controlled
3.27.3 | The ventilation technology offered may be turbine, volume reflector, piston or ascending bellows
3.27.4 | The minimum basic modes to be supplied:
3.27.4.1 | Manual/Spontaneous Ventilation
3.27.4.2 | Volume Control Ventilation
3.27.4.3 | Pressure Control Ventilation
3.27.4.4 | Pressure Support with Mandatory Apnoea back-up ventilation
3.27.4.5 | Volume Controlled Ventilation + SIMV + Pressure Support
3.27.4.6 | Pressure Control Ventilation + SIMV + Presssure Support
3.27.4.7 | CPAP
3.27.5 | Additional modes of ventilation may be offered and quoted for separately
3.27.6 | The ventilator must be integrated with the workstation and user friendly
3.27.7 | Frequency: minimum range of 5 – 60 bpm.
3.27.8 | Inspiratory time: minimum range of 0.2 – 10.0 seconds.
3.27.9 | Minimum Inspiration/Expiration Ratio: 4:1 – 1:4.
3.27.10 | Adjustable inspiratory pause/plateau time (volume control ventilation): 0-30% of inspiratory time
3.27.11 | Peak Inspiratory Pressure: minimum range of: PEEP + 5 to 70 cm H2O.
3.27.12 | Positive End Expiratory Pressure (PEEP): Integrated, electronically controlled. Minimum range: 0 – 30 cm H2O.
3.27.13 | Tidal Volume: 20ml – 1500ml per breath. Minimum accuracy: 15%
3.27.14 | Peak Inspiratory Flow: At least 90 lpm
3.27.15 | The ventilator must be capable of flow or pressure triggering.
3.27.16 | Adjustable Inspiratory Rise time/Ramp with a range of 0.1 - 2 sec
3.27.17 | Compliance compensation/Tidal volume compensation/Compressible volume compensation or equivalent.
3.27.18 | Patient gas flow compensation must be automatic so that alterations in patient gas flow does not affect delivered tidal volumes
3.27.19 | It must be possible to set CPAP in manual/spontaneous mode within a range of: 1 to 10 cm H2O
3.27.20 | In event of power or technical failure, manual ventilation must be possible.
3.28 | Display:
3.28.1 | All displays must be in acceptable English
3.28.2 | Colour display must include numerical and graphic display of parameters measured
3.28.3 | Graphic display to feature at least 3 waveforms: - pressure, flow and volume.
3.28.4 | Must also display flow-volume and pressure-volume loops with a reference loop
3.28.5 | Alarms : The ventilator must have the following user adjustable alarms:
3.28.5.1 | Apnoea alarm
3.28.5.2 | High airway pressure alarm
3.28.5.3 | Continuous pressure alarm
3.28.5.4 | High and low expiratory tidal volume alarm
3.28.5.5 | Minute volume, high/low alarm
3.28.5.6 | Inspired oxygen (Fi02), high/low alarm
3.28.5.7 | PEEP, low alarm
3.28.5.8 | Leakage / Check breathing circuit alarm or equivalent
3.28.5.9 | Alarm profile must be user definable or default to patient appropriate settings
3.28.5.10 | The ventilator must have a mute alarm function, active only for a limited time. No more than 2 minutes
3.28.5.11 | It must not be possible to cancel the visual alarm until the problem has been rectified.
3.28.6 | All components of the ventilator and breathing system in contact with the patient’s airway must be latex free.
3.28.7 | All components of the ventilator and breathing system in contact with the patient’s airway must be steam autoclavable up to 134ᴼ C
3.29 | Integrated Gas Monitoring
3.29.1 | Oxygen Monitoring
3.29.1.1 | The system must display both inspired and end-tidal oxygen concentration.
3.29.1.2 | The oxygen analyser must have user adjustable low and high alarms – visual and audible. The low alarm must automatically default to no less than 21%.
3.29.1.3 | The system must be accurate (+/- 2% in 21-100%) range, and have a rapid response time (≤ 500ms).
3.29.1.4 | The unit must sound a continuous alarm if the oxygen pressure falls below 190 kPa, then all nitrous oxide must be automatically cut from the patient circuit.
3.29.1.5 | Bidder to give details of the oxygen analyser
3.29.1.6 | The lifespan of the oxygen analyser must be guaranteed for at least 10 years.
3.29.1.7 | A system using replaceable oxygen fuel cells will not be accepted
3.29.2 | Anaesthetic vapour, carbon dioxide and nitrous oxide monitoring:
3.29.2.1 | Anaesthetic vapour, N2O monitoring and capnography must be integrated into the anaesthetic workstation.
3.29.2.2 | The monitoring must be of the side-stream type.
3.29.2.3 | Sampling: 100 - 200 ml/min
3.29.2.4 | Sampled gases must be internally returned to the breathing system.
3.29.2.5 | There must be no water build-up in the measuring line.
3.29.2.6 | CO2 must be presented as a numerical display of inspired and end-tidal concentration and a capnography waveform.The displayed units must be user-selectable: kPa or mmHg.
3.29.2.7 | The monitoring system must have automatic identification of anaesthetic vapours.
3.29.2.8 | The system must have a graphical display indicating the type of anaesthetic vapour/s currently in use and the concentration thereof. The colour coding must be as per international colour coding for anaesthetic vapours.
3.29.2.9 | The anaestheric vapour must be displayed as vol % and must include inspiratory and expiratory concentrations.
3.29.2.10 | The monitoring system must be able to simultaneously detect, identify and display mixed anaesthetic vapours.
3.29.2.11 | High and low anaesthetic vapour alarms must be present and these must be user-definable.
3.29.2.12 | There must be automatic calculation of minimum alveolar concentration.
3.29.2.13 | The system must be able to calculate anaesthetic gas utilization in ml per case
3.29.2.14 | The unit must have an indication of the level of vapour in the vaporiser
3.29.2.15 | Integrated vapour analysis must deliver an audible "low vapour " alarm and a critical alarm for an empty tank
3.30 | Circle System
3.30.1 | The circle absorber system must be integrated as part of the ventilation system.
3.30.2 | The circle absorber system must have an effective system to eliminate excess water created by the chemical reactions either by heating or water traps.
3.30.3 | The circle absorber system must be autoclavable at 134ºC.
3.30.4 | The circle absorber system must be supplied with a non-disposable CO2 absorber canister
3.30.5 | The absorber canister may be reusable/refillable
3.30.6 | It must be easy to remove and replace the absorber canister while the workstation is in operation.
3.30.7 | Misalignment of the canister must be readily apparent to the user.
3.30.8 | The system must have a circle bypass facility while changing/refilling the absorber canister.
3.30.9 | The breathing circuit must contain an integrated adjustable pressure limiting (APL) valve with a minimum range of 0-70 cm H2O.
3.30.10 | The APL valve must have a tactile or electronic indicator when a value of between 20-30 cm H2O is exceeded.
3.30.11 | The circle absorber system must be latex free.
3.30.12 | The bag-to-ventilator switching mechanism must be clearly labeled and user-friendly. Mechanical or electronic
3.31 | Patient Physiological Monitoring
3.31.1 | The patient monitor must be modular and suitable for use for neonatal, paediatric and adult patients.
3.31.2 | The following paramaters must be monitored:
3.31.2.1 | - ECG: 3-lead and 5-lead. Must be capable of displaying 2 channels simultaneously.
3.31.2.2 | - ST-segment and arrhythmia analysis.
3.31.2.3 | - NIBP
3.31.2.4 | - Oxygen saturation (SpO2)
3.31.2.5 | - Plethysmography/respiratory rate
3.31.2.6 | - Temperature x 2 (minimum)
3.31.2.7 | - Invasive pressure x 4 (minimum)
3.31.3 | The patient monitor must be fully integrated with the workstation.
3.31.4 | Display must be no less than 400 mm diagonally.
3.31.5 | Minimum 6 waveforms must be able to be displayed simultaneously.
3.31.6 | Full colour display.
3.31.7 | The monitor must be capable of displaying trends for at least 24 hours in tabular and graphical format.
3.31.8 | The trend resolution must be user-adjustable and have a minimum resolution of 1 min.
3.31.9 | The front panel must be easy to clean and disinfect.
3.31.10 | The display must be flicker free and easy to read.
3.31.11 | Brightness control must be possible.
3.31.12 | In addition to the waveform display section, the monitor must have alphanumeric display of parameter identification, measured values, alarm and status messages.
3.31.13 | Audible and visual alarms must be available for all measured parameters.
3.31.14 | Default alarm limits if the limits are not set by the user
3.31.15 | Alarm indication: Visual and audible with mute
3.31.16 | Adjustable audio alarm volume
3.31.17 | Low battery: Alarm at least 10 minutes before battery power is exhausted
3.31.18 | It must be possible to mute the audible alarm for a short period (no longer than 2 minutes), but no override of the visual alarm must be possible.
3.31.19 | It must be possible to suspend the alarms of certain parameters during clinical procedures, but the alarm must reactivate automatically after a suitable time delay. Any silencing or disabling of alarms must be clearly indicated on the monitor.
3.31.20 | Input impedance of the system > 10 MΩ.
3.31.21 | Maximum allowable leakage current to the patient must be less than 10 µA.
3.31.22 | The monitor must be protected against defibrillator and electrosurgical interference.
3.31.23 | The monitor must be of low power consumption, and must not require fan or other forced ventilation.
3.31.24 | All cables and sockets must be colour coded for easy recognition, and must conform to international standards
3.32 | ECG
3.32.1 | A full set (adult and paediatric) of 3/5 lead ECG cables must be included in the offer.
3.32.2 | The ECG amplifier must employ a differential input with high noise suppression.
3.32.3 | Sweep speed: at least 25 to 50 mm/sec.
3.32.4 | Freeze facility for ECG.
3.32.5 | The input impedance must be > 10 MΩ.
3.32.6 | The heart rate measuring range: 20 to 300 bpm, with an accuracy < 5 bpm.
3.32.7 | An 50 Hz filter must be provided.
3.32.8 | Alarm limits must be adjustable in the following range:
3.32.9 | Low: 40 to 150
3.32.10 | High: 70 to 250.
3.33 | Non Invasive Blood Pressure
3.33.1 | Must be suitable for adults, paeds and neonates
3.33.2 | The unit must use the oscillometric method of measurement.
3.33.3 | Modes: Manual and automatic
3.33.4 | Reusable neonate, paediatric, adult and obese cuffs must be supplied
3.33.5 | The cuffs must be robust, interchangeable and easy to clean.
3.33.6 | Must display systolic, diastolic and mean pressures for each measurement
3.33.7 | Adult: Systolic: 30 - 270 mm Hg, Diastolic: 20 - 240 mm Hg, MAP: 20 - 250 mm Hg
3.33.8 | Paeds: Systolic: 30 - 160 mm Hg, Diastolic: 20 - 120, MAP: 20 - 120 mm Hg
3.33.9 | Values must be displayed numerically
3.33.10 | Auto zero when unit is switched on
3.33.11 | The interval between measurements must be selectable from 1 – 120 minutes.
3.33.12 | Adjustable high and low alarm limits for systolic, diastolic, mean arterial pressure and heart rate.
3.33.13 | Over pressure protection: Automatic deflation at 300 mm Hg adult, 150 mm Hg neonate
3.33.14 | Average measurement time: approximately 30 s
3.33.15 | If measurement exceeds two minutes the cuff must deflate, an alarm must sound and likely faults indicated.
3.33.16 | The air hose must be at least two metres long.
3.34 | Oxygen Saturation.
3.34.1 | Continuous monitoring of oxygen saturation levels by means of the photometric principle.
3.34.2 | Must be suitable for adults, paeds and neonates
3.34.3 | The unit must continuously display the SaO2 % waveform
3.34.4 | Oxygen saturation (SaO2 %): Range of measurement: 0 - 100%
3.34.5 | The accuracy must be better than 2,5 % (70 - 100%) and 5% (50 - 70%).
3.34.6 | Pulse rate (beats/min): Range of measurement: 20 - 300 bpm
3.34.7 | The pulse rate accuracy must be better than ± 2 %.
3.34.8 | An audible bleep should sound for each heartbeat, the pitch of which should vary with the oxygen saturation.
3.34.9 | Automatic rejection of motion artefacts
3.35 | SpO2 Sensor:
3.35.1 | It must be the latest technology with clinical research support based on accuracy and real-time monitoring. Please submit the published peer reviewed/researched paper as supporting documents.
3.35.2 | The following reusable sensors must be provided:
3.35.3 | Finger sensor for adults and paeds.
3.35.4 | Earlobe sensor for adults and paeds.
3.35.5 | Finger wrap-around probe for neonates
3.35.6 | The sensor lead must be approximately 3 metres long.
3.36 | Temperature
3.36.1 | Temperature range must be approximately 15º to 45º C.
3.36.2 | Two reusable temperature probes must be included
3.37 | Other system requirements
3.37.3 | The unit must operate from both the 220 V, 50 Hz mains supply and internal rechargeable batteries
3.37.4 | The batteries must be charged while the unit is operating from the mains supply.
3.37.5 | The batteries must automatically take over in the event of a mains power failure.
3.37.6 | The batteries must have the capacity for full operation of the anaesthetic machine, ventilator and monitor for at least 90 minutes with batteries fully charged, under typical operating conditions.
3.37.7 | Graphic battery power indicator
3.37.8 | Recharge time < 8 hrs
3.37.9 | Mains cable at least four metres long with SABS approved 15A three prong plug.
3.37.10 | The unit must have a rail for attaching a suction regulator.
3.38 | Data management:
3.38.1 | The workstation must be capable of downloading data to a USB device or memory card.
3.38.2 | The monitor must be able to communicate with a central station
3.38.3 | A printer must be incorporated in the unit
3.39 | The unit must be handed over in full operating order
4 | The following Accessories and Consumables must be supplied with the unit:
4.1 | Re-breathing bag, reusable, rubber
4.2 | Reservoir bag
4.3 | Pressure limiting valve, adjustable
4.4 | Adult & Paediatric volume sensor re-usable
4.5 | Water trap for multigas monitor, re-usable (10 pcs.)
4.6 | A full set paediatric 3/5 lead ECG cables
4.7 | A full set adult 3/5 lead ECG cables
4.8 | Reusable paediatric cuffs
4.9 | Reusable adult cuffs
4.10 | Reusable neonate cuffs
4.11 | Reusable obese cuffs
4.12 | Air hose
4.13 | Finger sensor for paeds
4.14 | Finger sensor for adults
4.15 | Earlobe sensor for paeds
4.16 | Earlobe sensor for adults
4.17 | Finger wrap-around probe for neonates
4.18 | Sensor lead
4.19 | Two reusable temperature probes
4.20 | Vaporizer key-fillers (Sevo, & Iso)
4.21 | The unit must be supplied with two vaporizers (1x sevoflurane, & 1x Iso)
4.22 | Anaesthesia circuit (Adult, single/double limb design) Disposable. 1.8m (Complete with Y-piece, Elbow, Catheter mount, Rebreathing bag)
4.23 | Anaesthesia circuit (Paediatric, single/double limb design) Disposable. 1.8m (Complete with Y-piece, Elbow, Catheter mount, Rebreathing bag)
4.24 | Anaesthesia circuit (neonatal, single/double limb design) Disposable. 1.8m (Complete with Y-piece, Elbow, Catheter mount, Rebreathing bag)
4.25 | Re-breathing circuit complete
4.26 | Non-re-breathing circuit complete
4.27 | Flow sensor (pack of 2)
4.28 | Spirometry line
4.29 | Sample line (pack of 10)
4.30 | Face masks: sizes 00, 0, 1, 2, 3, 4, 5.
4.31 | Bacterial/viral filter
4.32 | Y connector
4.33 | Manual ventilation hose
4.34 | Fresh gas tubing
4.35 | Seal for back-up cylinder connection on the unit
4.36 | Scavenger tubing (as per institution's scavenging outlet)
4.37 | Inspiratory and expiratory unidirectional valves
4.38 | Side stream CO2 kit (adult/paed)
4.39 | Disposable Adult Spirometry Set complete with spirometry tube, sample line and flow sensor, 3 m, box of 20 pieces
4.40 | Disposable Paediatric Spirometry Set complete with spirometry tube, sample line and flow sensor, 3 m, box of 20 pieces
4.41 | Gas sample tube with luer lock connection, 3m, disposable (10 pcs)
4.42 | Suction Tubing Set inc connector
4.43 | Resuscitator Adult, 1500mL, autoclavable, silicone
4.44 | Resuscitator Paediatric, 550mL, autoclavable, silicone
4.45 | Tubing set Paediatric for circle systems
6 | Conformity Compliance (Please attach certificate)
6.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
6.2 | The unit must comply with electrical safety and performance standard such as IEC 60601-2-13 for Anaesthetic systems, attached certification
6.2.1 | Please provide unique ref number of the IEC ceriticate:
6.3 | OEM must comply and certified on ISO 9001 quality standards, attach proof of compliance
6.3.1 | Please provide unique ref number of the ISO 9001 ceriticate:
6.4 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
6.4.1 | Please provide unique ref number of the ISO 13485 ceriticate:
6.5 | Model quoted for must be EC certified. Attach a copy of certification
6.5.1 | Please provide unique ref number of the EC ceriticate:
6.6 | Hazardous Substance Act: If this type of equipment / apparatus appears on the schedule of Hazardous substances issued by the Directorate: Health Technology of the Department of Health, a license in terms of the Act on Hazardous Substances (Act 15/1973) must be submitted with the bid / quotation documents. The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
6.7 | All electrical/electronic medical equipment must be licensed by Radiation Control, where applicable a copy of the license must be submitted The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
6.7.1 | Radiation Control License number:
6.8 | The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade (rubber to with 2 screws) type and it must be a minimum length of three (3) meters. NB the mains cable of the unit tendered for must be SABS colour coded.
6.9 | The equipment quoted for must be protected against electromagnetic interference IEC 60601-1-2
6.10 | The unit must comply with electrical safety and performance standard such as IEC 60601-2-49 for Patient Monitoring, attached certification
6.10.1 | Please provide unique ref number of the IEC ceriticate:
7 | COMPREHENSIVE MAINTENANCE
7.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) must be included in the bid price
7.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries, Flow sensors and Oxygen sensor), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
7.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
7.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries, Flow sensors and Oxygen sensor), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
7.3 | The bidder to provide the OEM recommended service interval of the unit:
7.4 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
8 | General
8.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
8.2 | The unit on offer must be the latest technology.
ANAESTHESIA DELIVERY SYSTEM: ADVANCED with integrated physiological monitor, Replacement Accessories
RT2-01-06a | A full set paediatric 3/5 lead ECG cable
RT2-01-06b | A full set adult 3/5 lead ECG cable
RT2-01-06c | Reusable paediatric cuff
RT2-01-06d | Reusable adult cuff
RT2-01-06e | Reusable neonate cuff
RT2-01-06f | Reusable obese cuff
RT2-01-06g | Air hose
RT2-01-06h | Finger sensor for paeds
RT2-01-06i | Finger sensor for adults
RT2-01-06j | Earlobe sensor for paeds
RT2-01-06k | Earlobe sensor for adults
RT2-01-06l | Finger wrap-around probe for neonates
RT2-01-06m | Sensor lead
RT2-01-06n | Reusable temperature probe
UPGRADES
RT2-01-06o | AIMS (Anaesthesia information management system)
RT2-01-06p | Auxilliary Delivery System
RT2-01-06q | Neuromuscular transmission monitoring
RT2-01-06r | Metabolic monitoring
RT2-01-06s | Depth of anaesthesia monitoring system
RT2-01-06t | Invasive cardiac output monitoring
RT2-01-06u | Non-invasive cardiac output monitoring
RT2-01-06v | Cerebral oximetry
Bidder to list any other Consumables or Accessories (Not listed above) for the model on offer below

## Anaesthetic Monitor
RT2-01-07 | ANAESTHESIA PHYSIOLOGICAL MONITOR
1 | Description
1a | Product Offered (Monitor)
1a.1 | Make:
1a.2 | Model:
1a.3 | Initial Manufacturing year of Model:
1a.4 | Radiation Control License number:
2 | Technical Specification
2.1 | The patient monitor must be modular and suitable for use for neonatal, paediatric and adult patients.
2.2 | The following paramaters must be monitored:
2.2.1 | - ECG: 3-lead and 5-lead. Must be capable of displaying 2 channels simultaneously.
2.2.2 | - ST-segment and arrhythmia analysis.
2.2.3 | - NIBP
2.2.4 | - Oxygen saturation (SpO2)
2.2.5 | - Plethysmography/respiratory rate
2.2.6 | - Temperature x 2 (minimum)
2.2.7 | - Invasive pressure x 4 (minimum)
2.3 | The patient monitor must be fully integrated with the workstation.
2.4 | Display must be no less than 400 mm diagonally.
2.5 | Minimum 6 waveforms must be able to be displayed simultaneously.
2.6 | Full colour display.
2.7 | The monitor must be capable of displaying trends for at least 24 hours in tabular and graphical format.
2.8 | The trend resolution must be user-adjustable and have a minimum resolution of 1 min.
2.9 | The front panel must be easy to clean and disinfect.
2.10 | The display must be flicker free and easy to read.
2.11 | Brightness control must be possible.
2.12 | In addition to the waveform display section, the monitor must have alphanumeric display of parameter identification, measured values, alarm and status messages.
2.13 | Audible and visual alarms must be available for all measured parameters.
2.14 | Default alarm limits if the limits are not set by the user
2.15 | Alarm indication: Visual and audible with mute
2.16 | Adjustable audio alarm volume
2.17 | Low battery: Alarm at least 10 minutes before battery power is exhausted
2.18 | It must be possible to mute the audible alarm for a short period (no longer than 2 minutes), but no override of the visual alarm must be possible.
2.19 | It must be possible to suspend the alarms of certain parameters during clinical procedures, but the alarm must reactivate automatically after a suitable time delay. Any silencing or disabling of alarms must be clearly indicated on the monitor.
2.20 | Input impedance of the system > 10 MΩ.
2.21 | Maximum allowable leakage current to the patient must be less than 10 µA.
2.22 | The monitor must be protected against defibrillator and electrosurgical interference.
2.23 | The monitor must be of low power consumption, and must not require fan or other forced ventilation.
2.24 | All cables and sockets must be colour coded for easy recognition, and must conform to international standards
2.25 | ECG
2.25.1 | A full set (adult and paediatric) of 3/5 lead ECG cables must be included in the offer.
2.25.2 | The ECG amplifier must employ a differential input with high noise suppression.
2.25.3 | Sweep speed: at least 25 to 50 mm/sec.
2.25.4 | Freeze facility for ECG.
2.25.5 | The input impedance must be > 10 MΩ.
2.25.6 | The heart rate measuring range: 20 to 300 bpm, with an accuracy < 5 bpm.
2.25.7 | An 50 Hz filter must be provided.
2.25.8 | Alarm limits must be adjustable in the following range:
2.25.8.1 | Low: 40 to 150
2.25.8.2 | High: 70 to 250.
2.26 | Non Invasive Blood Pressure
2.26.1 | Must be suitable for adults, paeds and neonates
2.26.2 | The unit must use the oscillometric method of measurement.
2.26.3 | Modes: Manual and automatic
2.26.4 | Reusable neonate, paediatric, adult and obese cuffs must be supplied
2.26.5 | The cuffs must be robust, interchangeable and easy to clean.
2.26.6 | Must display systolic, diastolic and mean pressures for each measurement
2.26.6.1 | Adult: Systolic: 30 - 270 mm Hg, Diastolic: 20 - 240 mm Hg, MAP: 20 - 250 mm Hg
2.26.6.2 | Paeds: Systolic: 30 - 160 mm Hg, Diastolic: 20 - 120, MAP: 20 - 120 mm Hg
2.26.7 | Values must be displayed numerically
2.26.8 | Auto zero when unit is switched on
2.26.9 | The interval between measurements must be selectable from 1 – 120 minutes.
2.26.10 | Adjustable high and low alarm limits for systolic, diastolic, mean arterial pressure and heart rate.
2.26.11 | Over pressure protection: Automatic deflation at 300 mm Hg adult, 150 mm Hg neonate
2.26.12 | Average measurement time: approximately 30 s
2.26.13 | If measurement exceeds two minutes the cuff must deflate, an alarm must sound and likely faults indicated.
2.26.14 | The air hose must be at least two metres long.
2.27 | Oxygen Saturation.
2.27.1 | Continuous monitoring of oxygen saturation levels by means of the photometric principle.
2.27.2 | Must be suitable for adults, paeds and neonates
2.27.3 | The unit must continuously display the SaO2 % waveform
2.27.4 | Oxygen saturation (SaO2 %): Range of measurement: 0 - 100%
2.27.5 | The accuracy must be better than 2,5 % (70 - 100%) and 5% (50 - 70%).
2.27.6 | Pulse rate (beats/min): Range of measurement: 20 - 300 bpm
2.27.7 | The pulse rate accuracy must be better than ± 2 %.
2.27.8 | An audible bleep should sound for each heartbeat, the pitch of which should vary with the oxygen saturation.
2.27.9 | Automatic rejection of motion artefacts
2.28 | SpO2 Sensor:
2.28.1 | It must be the latest technology with clinical research support based on accuracy and real-time monitoring. Please submit the published peer reviewed/researched paper as supporting documents.
2.28.2 | The following reusable sensors must be provided:
2.28.2.1 | Finger sensor for adults and paeds.
2.28.2.2 | Earlobe sensor for adults and paeds.
2.28.2.3 | Finger wrap-around probe for neonates
2.28.3 | The sensor lead must be approximately 3 metres long.
2.29 | Temperature
2.29.1 | Temperature range must be approximately 15º to 45º C.
2.29.2 | Two reusable temperature probes must be included
3 | The following accessories must be supplied with the unit:
3.1 | A full set paediatric 3/5 lead ECG cables
3.2 | A full set adult 3/5 lead ECG cables
3.3 | Reusable paediatric cuffs
3.4 | Reusable adult cuffs
3.5 | Reusable neonate cuffs
3.6 | Reusable obese cuffs
3.7 | Air hose
3.8 | Finger sensor for paeds
3.9 | Finger sensor for adults
3.10 | Earlobe sensor for paeds
3.11 | Earlobe sensor for adults
3.12 | Finger wrap-around probe for neonates
3.13 | Sensor lead
3.14 | Two reusable temperature probes
4 | Conformity Compliance (Please attach certificate)
4.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
4.2 | The unit must comply with electrical safety and performance standard such as IEC 60601-2-49 for Patient Monitoring, attached certification
4.2.1 | Please provide unique ref number of the IEC ceriticate:
4.3 | OEM must comply and certified on ISO 9001 quality standards, attach proof of compliance
4.3.1 | Please provide unique ref number of the ISO 9001 ceriticate:
4.4 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
4.4.1 | Please provide unique ref number of the ISO 13485 ceriticate:
4.5 | Model quoted for must be EC certified. Attach a copy of certification
4.5.1 | Please provide unique ref number of the EC ceriticate:
4.6 | Hazardous Substance Act: If this type of equipment / apparatus appears on the schedule of Hazardous substances issued by the Directorate: Health Technology of the Department of Health, a license in terms of the Act on Hazardous Substances (Act 15/1973) must be submitted with the bid / quotation documents. The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
4.7 | All electrical/electronic medical equipment must be licensed by Radiation Control, where applicable a copy of the license must be submitted The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
4.7.1 | Radiation Control License number:
4.8 | The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade (rubber to with 2 screws) type and it must be a minimum length of three (3) meters. NB the mains cable of the unit tendered for must be SABS colour coded.
4.9 | The equipment quoted for must be protected against electromagnetic interference IEC 60601-1-2
5 | COMPREHENSIVE MAINTENANCE
5.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) must be included in the bid price
5.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries, Flow sensors and Oxygen sensor), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
5.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
5.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries, Flow sensors and Oxygen sensor), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
5.3 | The bidder to provide the OEM recommended service interval of the unit:
5.4 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
6 | General
6.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
6.2 | The unit on offer must be the latest technology.
ANAESTHESIA PHYSIOLOGICAL MONITOR ACCESSORIES and CONSUMABLES
RT2-01-07a | A full set paediatric 3/5 lead ECG cable
RT2-01-07b | A full set adult 3/5 lead ECG cable
RT2-01-07c | Reusable paediatric cuff
RT2-01-07d | Reusable adult cuff
RT2-01-07e | Reusable neonate cuff
RT2-01-07f | Reusable obese cuff
RT2-01-07g | Air hose
RT2-01-07h | Finger sensor for paeds
RT2-01-07i | Finger sensor for adults
RT2-01-07j | Earlobe sensor for paeds
RT2-01-07k | Earlobe sensor for adults
RT2-01-07l | Finger wrap-around probe for neonates
RT2-01-07m | Sensor lead
RT2-01-07n | Reusable temperature probe
Bidder to list any other Consumables or Accessories (Not listed above) for the model on offer below

## Anaesthetic, MRI compatible
RT2-01-08 | ANAESTHETIC DELIVERY SYSTEM MRI Compatible, WITHOUT Physiological monitor
1 | DESCRIPTION
1.1 | An advanced anaesthesia delivery system comprising a completely integrated electronic anaesthesia workstation and can be used in MRI scanner rooms with magnets of 1.5 and 3 Tesla with a field strength between 300-400 gauss.
1a | Product Offered (Anaesthetic Machine)
1a.1 | Make:
1a.2 | Model:
1a.3 | Initial Manufacturing year of Model:
1a.4 | Radiation Control License number:
2 | TECHNICAL SPECIFICATION
2.1 | The unit must include electronicaly integrated oxygen, nitrous oxide and air delivery and monitoring.The Unit must have a physically integrated magnetic field strength monitor(Gauss Alarm)
2.2 | Ventilation and anaesthesia vapour monitoring must be fully integrated within the unit.
2.3 | The anaesthesia unit must be able to communicate/interface with a physiological monitor
2.4 | The anaesthesia system must comprise:
2.5 | Anaesthesia system suitable for neonatal (down to 500g), paediatric and adult anaesthesia.
2.6 | Ventilator suitable for ventilating neonatal (down to 500g), paediatric, adult and morbid/obese patients
3 | Multigas analyser: Must be available on Anaesthetic machine
3.1 | Oxygen
3.2 | Nitrous oxide
3.3 | Carbon dioxide
3.4 | Anaesthetic vapour monitoring:
3.4.1 | Halothane
3.4.2 | Isoflurane
3.4.3 | Sevoflurane
3.4.4 | Desflurane
3.5 | Complete and automated funtional test/check requiring minimal user intervention or disassembly.
3.6 | The functional test/check must provide the measurements and values for automatic leakage and compliance.
3.7 | Automatic calibration of sensors for anaesthetic gases and vapours must form part of the self test and require no user intervention.
3.8 | It must be possible to override the functional test/check facility. The system must display the override status
3.9 | The unit must provide conventional, low-flow (down to 200 ml/min) anaesthesia.
3.1 | The unit must return analysed gas back to the circle system
3.11 | A safety oxygen function is required for the delivery of oxygen through the anaesthesia unit in the event of an electronic failure of the gas delivery system.
3.12 | The unit must be supplied with a minimum of 1 storage drawer
3.13 | The unit must be supplied with a minimum of 3 additional 220 V power sockets. These sockets must be excluded from the battery backup circuit in case of a power failure
3.14 | The unit must have a working area or a retractable shelf. Width: not less than 300 mm. Depth: not less than 200 mm
3.15 | The unit must have a light to enable the anaesthetist to work in a darkened theatre.
4 | Trolley Made out of MRI compatable material
4.1 | The unit must be mounted on a strong , robust and compact trolley manufactured from corrosion free material. The material must be MRI compatible.
4.2 | The trolley must have no sharp corners and be easy to clean.
4.3 | The trolley must have at least four antistatic castors of diameter 125 mm or greater.
4.4 | Trolley must have two lockable castors or a central locking brake.
4.5 | The unit must accommodate an oxygen cylinder. Cylinder size according to user requirement. Size to be compatible with current state vendor of oxygen and cylinder must be MRI compatible (made out of alluminium).
4.6 | Trolley must have accessible handles for ease of movement
4.7 | The trolley, with all components mounted, must be stable without toppling over when placed on a 30° slope.
4.8 | The unit must be supplied with a castor guard system, preventing the entaglement and trapping of cables within the wheels.
4.9 | The trolley must have a mounting system suitable for a physiologial monitor.
4.1 | Pipeline and Cylinder Gas Supplies All made out of MRI compatible material. The anaesthesia system must be designed to use external piped medical gasses (O2, N2O, medical air), vacuum (scavenging) from medical service outlet sockets.
4.11 | Connection hoses for above must be provided and must be fitted with Heyer type non-interchangeable probes for the existing medical gas outlets.
4.12 | Hoses must be fitted with Heyer type non-interchangeable probes for the existing medical gas outlets.
4.13 | The medical air hose connection must include a water trap
4.14 | A safety pressure relief valve of 60 - 80 cm H2O must be provided to protect the system from excessive pressure .
4.15 | The unit must function safely and correctly within the pipeline supply pressure range of: 270 - 600 kPa
4.16 | The system must preferentially utilise the piped oxygen supply over the cylinder supply, if pipeline pressure is > 270 kPa
5 | Connection hoses must be Flexible and at least 6 metres long:
5.1 | Antistatic
5.2 | Colour coded as follows:
5.3 | O2: white
5.4 | N2O: blue
5.5 | Vacuum: yellow
5.6 | Medical air: black
5.7 | Scavenge low vacuum: orange or yellow
5.8 | Where relevant, fitted with outlet probes complying with SANS 1409:2014
5.9 | Permanently bonded to the outlet probe and the machine connector
5.1 | Hoses for gases above atmospheric pressure, must be designed for pressures of at least 1 000 kPa.
5.11 | Vacuum hoses may not collapse or kink at 90 kPa below atmospheric pressure, when subjected to 180° bend of inside radius equal to the outside diameter of the hose.
5.12 | All connection hoses must be fed to a central manifold.
5.13 | All high pressure gas and vacuum pipe connections (other than hoses) and all pipes connected to pressure regulators and pressure gauges must be stainless steel (18/8 grade) pipes that are oil-free and colour coded as for hoses (see above) over their entire lengths
5.14 | An in-line non-return valve must be fitted to each gas pipeline on the high pressure (inlet) line
5.15 | An in-line non-return valve must be fitted to each gas cylinder yoke
5.16 | A replaceable gas filter, to block particles greater than 65 µm, must be fitted on the upstream side of the non-return valve
5.17 | Each gas pipeline inlet must have a pressure gauge or electronic display of pressure. The pressure range (415 kPa ± 75 kPa) must be clearly indicated in colour
5.18 | Must be fitted with a gas cylinder yoke to accommodate one oxygen gas cylinder with a pin index system, MRI compatible
5.19 | Each gas cylinder yoke must have a pressure gauge or electronic display of pressure
5.2 | The pressure gauges or electronic displays of pressure for both pipeline & cylinder supplies, calibrated in kPa for each line, must be labelled and visible from the front of the unit.
5.21 | O2 and N2O regulators must be provided complying with CKS 605:1987.
5.22 | A pressure regulator complying with CKS 605 must be supplied for each gas inlet and must reduce the gas pressure to 415 kPa without freezing up, creeping, chattering or whistling
5.23 | Outlet sockets and probes must comply with SANS 1409:2014.
5.24 | Sockets with hoses must be colour-coded as per SANS 1409:2014:
5.25 | Oxygen: White background with black letters
5.26 | Nitrous oxide: Blue background with white letters
5.27 | Medical air: One half black with white letters, the other half white with black letters
5.28 | Vacuum: Yellow background with black letters
5.29 | Continuous display, visible from the front of the machine, of both pipeline and cylinder pressures even during power failure
5.3 | Pipeline and cylinder pressures must be differentially displayed, and clearly indicated
5.31 | Safety pressure relief valves must protect the entire pneumatic system from excessive input pressures from pipeline and cylinder supplies. Safety relief valve for all gases must be supplied and set at 600kPa
5.32 | When a cylinder is detached from the pin-indexed yoke no gas leak must occur at the yoke.
5.33 | Airway pressure must be monitored and displayed
5.34 | Scavenging-MRI compatible material
5.35 | Scavenging-MRI compatible material
5.36 | The unit must have a scavenging system with a hose at least 5 metres long with the suitable connector as per the institutions scavenging installation.
5.37 | Air-break device to prevent negative pressure on the breathing system
5.38 | If an active sytem is supplied a pressure indicator must be present
5.39 | The scavenging outlet must be a 30mm ISO connector.
6 | Patient Gas Delivery System
6.1 | Patient gas delivery must be electronically controlled\Rotameters
6.2 | Must display the delivered patient gas flow using an ISO colour standard for each gas.
6.3 | The unit must be able to deliver a total patient gas flow of 200 ml/min to 18 l/min (Minimum). The unit must have the ability to deliver at least 10 l/min of oxygen. These must be incremental adjustments
6.4 | Gas supply must conform to:
6.5 | O2 minimum flow range of 200 ml/min to 10 l/min. Accuracy ± 5%
6.6 | medical air minimum flow range 0 to 10 l/min. Accuracy ± 5%
6.7 | nitrous oxide minimum flow range 0 to 10 l/min. Accuracy ± 5%
6.8 | Flowmeter sequence to be according to standards approved by the South Africal Society of Anaesthesiologists - Oxygen on the right
6.9 | The delivery system must prevent simultaneous delivery of two anaesthetic vapours.
6.1 | A separate common gas outlet to accept a standard 22 mm outside diameter and 15 mm internal diameter.
7 | Hypoxic guard
7.1 | The unit must include a safety mechanism that ensures a minimum concentration of 25% oxygen in any oxygen/nitrous oxide gas mixture across all flow rates.
7.2 | The unit must include a safety mechanism that ensures a minimum concentration of 21% oxygen in any oxygen/air gas mixture across all flow rates.
7.3 | Accidental simultaneous delivery of N2O and Air into the breathing system, must either be prevented by means of an interlock switch or clearly indicated so that the anaesthetist can rectify the error. Delivery indicated purely by mechanical rotameters will not achieve the requirements of this clause.
7.4 | The unit must include a high-flow oxygen flush mechanism with a flow rate of at least 35 l/min
7.5 | Oxygen flush must be non-lockable, recessed and self closing
7.6 | This system must bypass the flowmeters and vaporisers
8 | Breathing System
8.1 | Total circuit volume, including absorber, must not exceed 2,7 litres
8.2 | With APL valve 5 - 75 cm H2O
8.3 | With ventilator to manual switch. This switch when directed to the ventilator must isolate the APL valve from the circle.
8.4 | Breathing system must be fully autoclavable to 134°C
8.5 | There must be a mechanism to remove water from the breathing system
8.6 | Breathing system must be latex free
8.7 | Bag-to-ventilator switch automatically begins in the ventilator position
8.8 | Reusable corrugated breathing circuit to be supplied
8.9 | Bidders to supply the breathing circuit as required by the user
9 | Integrated Ventilator
9.1 | Ventilator suitable for ventilating adults, paediatric and neonates (> 500g) patients without changing ventilator components (e.g. bellows)
9.2 | The ventilator must be electronically controlled
9.3 | The ventilation technology offered may be turbine, volume reflector, piston or ascending bellows
9.4 | The minimum basic modes to be supplied:
9.5 | Manual/Spontaneous Ventilation
9.6 | Volume Control Ventilation
9.7 | Pressure Control Ventilation
9.8 | Pressure Support with Mandatory Apnoea back-up ventilation
9.1 | Volume Controlled Ventilation + SIMV + Pressure Support
9.11 | Pressure Control Ventilation + SIMV + Presssure Support
9.12 | CPAP
9.13 | Additional modes of ventilation may be offered and quoted for separately
9.14 | The ventilator must be integrated with the workstation and user friendly
10 | The ventilator must meet the following minimum criteria:
10.1 | Frequency: minimum range of 5 – 30 bpm.
10.2 | Inspiratory time: minimum range of 0.2 – 10.0 seconds.
10.3 | Minimum Inspiration/Expiration Ratio: 4:1 – 1:4.
10.4 | Adjustable inspiratory pause/plateau time (volume control ventilation): 0-30% of inspiratory time
10.5 | Peak Inspiratory Pressure: minimum range of: PEEP + 5 to 30 cm H2O.
10.6 | Positive End Expiratory Pressure (PEEP): Integrated, electronically controlled. Minimum range: 0 – 30 cm H2O.
10.7 | Tidal Volume: 20ml – 1500ml per breath. Minimum accuracy: 15%
10.8 | Peak Inspiratory Flow: At least 120 /pm
10.9 | The ventilator must be capable of flow or pressure triggering.
10.1 | Adjustable Inspiratory Rise time/Ramp with a range of 0.1 - 2 sec
10.11 | Compliance compensation/Tidal volume compensation/Compressible volume compensation or equivalent.
10.12 | Patient gas flow compensation must be automatic so that alterations in patient gas flow does not affect delivered tidal volumes
10.13 | It must be possible to set CPAP in manual/spontaneous mode within a range of: 1 to 10 cm H2O
10.14 | In event of power or technical failure, manual ventilation must be possible.
11 | Display:
11.1 | All displays must be in English
11.2 | Colour display must include numerical and graphic display of parameters measured
11.3 | Graphic display to feature at least 3 waveforms: - pressure, flow and volume.
11.4 | Must also display flow-volume and pressure-volume loops with a reference loop
12 | Alarms : The ventilator must have the following user adjustable alarms
12.1 | Apnoea alarm
12.2 | High airway pressure alarm
12.3 | Continuous pressure alarm
12.4 | High and low expiratory tidal volume alarm
12.5 | Minute volume, high/low alarm
12.6 | Inspired oxygen (Fi02), high/low alarm
12.7 | PEEP, low alarm
12.8 | Leakage / Check breathing circuit alarm or equivalent
12.9 | Alarm profile must be user definable or default to patient appropriate settings
12.1 | The ventilator must have a mute alarm function, active only for a limited time (maximum 2 minutes)
12.11 | It must not be possible to cancel the visual alarm until the problem has been rectified.
12.12 | All components of the ventilator and breathing system in contact with the patient’s airway must be latex free.
12.13 | All components of the ventilator and breathing system in contact with the patient’s airway must be steam autoclavable up to 134ᴼ C
Integrated Gas Monitoring
13 | Oxygen Monitoring
13.1 | The system must display both inspired and end-tidal oxygen concentration.
13.2 | The oxygen analyser must have user adjustable low and high alarms – visual and audible. The low alarm must automatically default to no less than 21%.
13.3 | The system must be accurate (+/- 2% in 21-100%) range, and have a rapid response time (≤ 500ms).
13.4 | The unit must sound a continuous alarm if the oxygen pressure falls below 190 kPa, then all nitrous oxide must be automatically cut from the patient circuit.
14 | Anaesthetic vapour, carbon dioxide and nitrous oxide monitoring:
14.1 | Anaesthetic vapour, N2O monitoring and capnography must be electronically integrated into the anaesthetic workstation.
14.2 | The monitoring must be of the side-stream type.
14.3 | Sampling: 100 - 200 ml/min
14.4 | Sampled gases must be internally returned to the breathing system.
14.5 | CO2 must be presented as a numerical display of inspired and end-tidal concentration and a capnography waveform.The displayed units must be user-selectable: kPa or mmHg.
14.6 | The monitoring system must have automatic identification of anaesthetic vapours.
14.7 | The system must have a graphical display indicating the type of anaesthetic vapour/s currently in use and the concentration thereof. The colour coding must be as per international colour coding for anaesthetic vapours.
14.8 | The anaestheric vapour must be displayed as vol % and must include inspiratory and expiratory concentrations.
14.9 | The monitoring system must be able to simultaneously detect, identify and display mixed anaesthetic vapours.
14.1 | High and low anaesthetic vapour alarms must be present and these must be user-definable.
14.11 | There must be automatic calculation of minimum alveolar concentration.
14.12 | The system must be able to calculate anaesthetic gas utilization in ml per case
14.13 | The unit must have an indication of the level of vapour in the vaporiser
14.14 | Integrated vapour analysis must deliver an audible "low vapour " alarm and a critical alarm for an empty tank
15 | Circle System
15.1 | The circle absorber system must be integrated as part of the ventilation system.
15.2 | The circle absorber system must have an effective system to eliminate excess water created by the chemical reactions either by heating or water traps.
15.3 | The circle absorber system must be autoclavable at 134ºC.
15.4 | The circle absorber system must be supplied with a carbon dioxide absorber canister.
15.5 | The absorber canister may be reusable/refillable
15.6 | It must be easy to remove and replace the absorber canister while the workstation is in operation.
15.7 | Misalignment of the canister must be readily apparent to the user.
15.8 | The system must have a circle bypass facility while changing/refilling the absorber canister.
15.9 | The breathing circuit must contain an integrated adjustable pressure limiting (APL) valve with a minimum range of 0-70 cm H2O.
15.1 | The APL valve must have a tactile or electronic indicator when a value of between 20-30 cm H2O is exceeded.
15.11 | The circle absorber system must be latex free.
15.12 | The bag-to-ventilator switching mechanism must be clearly labeled and user-friendly. Mechanical or electronic
16 | General
16.1 | The unit must operate from both the 220 V, 50 Hz mains supply and internal rechargeable batteries
16.2 | The batteries must be charged while the unit is operating from the mains supply.
16.3 | The batteries must automatically take over in the event of a mains power failure.
16.4 | The batteries must have the capacity for full operation of the anaesthetic machine and ventilator for at least 90 minutes with batteries fully charged, under typical operating conditions.
16.5 | Graphic battery power indicator
16.6 | Recharge time < 8 hrs
16.7 | Mains cable at least six metres long with SABS approved 15A three prong plug.
16.8 | The unit must have a rail for attaching a suction regulator.
17 | Data management:
17.1 | The workstation must be capable of downloading data to a USB device or memory card.
17.2 | The Unit must have an intergated field strength monitor with acoustic, audible and visual alarms for Advisory,Caution and Warning.
17.3 | A laser printer must be incorporated in the unit and be MRI compatible
18 | Accessories and Consumables
18.1 | A full set of basic accessories for adult/paediatric/neonate must be supplied with the system,appropriate for use in an MRI suite
18.2 | List all the consumables required for full operation of the equipment, with prices separately
18.3 | All cables to be supplied as per user requirement
19 | STANDARD ACCESSORIES
19.1 | The unit must be supplied with a starter pack for full functionality upon commissioning.
19.2 | The unit must be delivered with a set of adult (A), paediatric (P) and neonatal (N) consumables:
19.2.1 | Patient circuit (A, P, N) - box of 10 each
19.2.2 | Filters (box of 10)
19.2.3 | Inflatable reusable face mask (A, P, N) - 3 full set of each
19.2.4 | Spirometery (box of 10 at least)
19.2.5 | Sampling line (box of 10 atleast)
19.2.6 | Flow sensors (box of 10 taleast)
19.2.7 | CO2 absorber canister disposable reusable x1 and disposable (box of 10 atleast)
19.2.8 | Vaporizer key-fillers (Sevo, & other choice below)
19.3 | The unit must be supplied with 2x MRI compatible vaporizers (one must sevoflurane, the other must be of the choice of the buying institution)
19.4 | Water traps (box of 10 at least)
21 | Vaporisers\MRI compatible(Tec 5 or 7)
21.1 | Vaporisers must be tilt, pressure and temperature compensated
21.2 | Mechanical, electronic or injection technology vaporisers are acceptable
21.3 | Filling mechanism:
21.3.1 | Vaporiser filling mechanism must be compatible with the current state approved vendor of inhalational volatile vapours.
21.3.2 | Each vaporiser must be supplied with 2 compatible filling adapters as per user requirement
21.3.3 | Minimum accuracy of the supplied volatile vapour delivery system must be at least 20% of set value
21.3.4 | Visible level indicator
22 | Auxilliary Delivery System
22.1 | An external auxilliary flowmeter suitable for delivery of oxygen via a face mask, supplied with a 6mm tapered connection. Flow range 0-10 l/min and above
23 | Trending:
23.1 | Trend information and event recording for:
23.1.1 | peak pressure
23.1.2 | plateau pressure
23.1.3 | PEEP
23.1.4 | expiratory tidal volume
23.1.5 | respiratory rate
23.2 | This must be provided in graphical and tabular formats.
23.3 | The trend and event information must cover at least the last 24 hours of operation.
23.4 | Resolution must be adjustable, with a minimum resolution of 1 minute intervals.
23.5 | The adjustment of the resolution must not affect the display of parameters which were sampled at different times to the resolution setting in such a way as to display a null reading, when one was in fact obtained.
24 | Software
24.1 | Any software for upgrades may be offered
25 | Suction Regulator
25.1 | A suction regulator capable of -50 kPa (minimum)
25.2 | The suction regulator must have a built-in diaphragm with a 0 – 100 kPa vacuum gauge and an “on-off” switch.
25.3 | It must aspirate 1 litre water in 45 seconds at a pressure of at least - 50 kPa
25.4 | A collection jar calibrated to at least 1 500 ml must be included.
25.5 | The collection jar must accommodate a disposable liner, as per user requirement
25 | Conformity Compliance (Please attach certificate)
25.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
25.2 | The unit must comply with electrical safety and performance standard such as IEC 60601-2-13 for Anaesthetic systems, attached certification
25.2.1 | Please provide unique ref number of the IEC ceriticate:
25.3 | OEM must comply and certified on ISO 9001 quality standards, attach proof of compliance
25.3.1 | Please provide unique ref number of the ISO 9001 ceriticate:
25.4 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
25.4.1 | Please provide unique ref number of the ISO 13485 ceriticate:
25.5 | Model quoted for must be EC certified. Attach a copy of certification
25.5.1 | Please provide unique ref number of the EC ceriticate:
25.6 | Hazardous Substance Act: If this type of equipment / apparatus appears on the schedule of Hazardous substances issued by the Directorate: Health Technology of the Department of Health, a license in terms of the Act on Hazardous Substances (Act 15/1973) must be submitted with the bid / quotation documents. The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
25.7 | All electrical/electronic medical equipment must be licensed by Radiation Control, where applicable a copy of the license must be submitted The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
25.7.1 | Radiation Control License number:
25.8 | The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade (rubber to with 2 screws) type and it must be a minimum length of three (3) meters. NB the mains cable of the unit tendered for must be SABS colour coded.
25.9 | The equipment quoted for must be protected against electromagnetic interference IEC 60601-1-2
26 | COMPREHENSIVE MAINTENANCE
26.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) must be included in the bid price
26.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries, Flow sensors and Oxygen sensor), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
26.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
26.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries, Flow sensors and Oxygen sensor), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
26.3 | The bidder to provide the OEM recommended service interval of the unit:
26.4 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
27 | General
27.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
27.2 | The unit on offer must be the latest technology.
ANAESTHETIC DELIVERY SYSTEM MRI Compatible, WITHOUT Physiological monitor- Accessories and Consumables
RT2-01-08a | Anaesthesia circuit (Adult, single/double limb design) Disposable. 1.8m (Complete with Y-piece, Elbow, Catheter mount, Rebreathing bag)
RT2-01-08b | Anaesthesia circuit (Paediatric, single/double limb design) Disposable. 1.8m (Complete with Y-piece, Elbow, Catheter mount, Rebreathing bag)
RT2-01-08c | Anaesthesia circuit (neonatal, single/double limb design) Disposable. 1.8m (Complete with Y-piece, Elbow, Catheter mount, Rebreathing bag)
RT2-01-08d | Re-breathing circuit complete
RT2-01-08e | Non-re-breathing circuit complete
RT2-01-08f | Flow sensor (pack of 2)
RT2-01-08g | Spirometry line
RT2-01-08h | Sample line (pack of 10)
RT2-01-08i | Face masks: sizes (00,0,1,2,3,4, 5)- End-user to specify size
RT2-01-08j | Bacterial/viral filter
RT2-01-08k | Y connector
RT2-01-08l | Manual ventilation hose
RT2-01-08m | Fresh gas tubing
RT2-01-08n | Seal for back-up cylinder connection on the unit
RT2-01-08o | Scavenger tubing (as per institution's scavenging outlet)
RT2-01-08p | Inspiratory and expiratory unidirectional valves
RT2-01-08q | Side stream CO2 kit (adult/paed)
RT2-01-08r | Disposable Adult Spirometry Set complete with spirometry tube, sample line and flow sensor, 3 m, box of 20 pieces
RT2-01-08s | Disposable Paediatric Spirometry Set complete with spirometry tube, sample line and flow sensor, 3 m, box of 20 pieces
RT2-01-08t | Gas sample tube with luer lock connection, 3m, disposable (10 pcs)
RT2-01-08u | Suction Tubing Set inc connector
RT2-01-08v | Resuscitator Adult, 1500mL, autoclavable, silicone
RT2-01-08w | Resuscitator Paediatric, 550mL, autoclavable, silicone
RT2-01-08x | Tubing set Paediatric for circle systems
RT2-01-08y | MRI compatible Sevoflurane vaporizer
RT2-01-08z | MRI compatible Isoflurane vaporizer
RT2-01-08zz | MRI compatible Desflurane vaporizer
Bidder to list any other Consumables or Accessories (Not listed above) for the model on offer below

## Anaesthetic, MRI comp + Monitor
RT2-01-09 | ANAESTHETIC DELIVERY SYSTEM WITH PHYSIOLOGICAL MONITOR MRI COMPATABLE
1 | DESCRIPTION
1.1 | An advanced anaesthesia delivery system comprising a completely integrated electronic anaesthesia workstation and can be used in MRI scanner rooms with magnets of 1.5 and 3 Tesla with a field strength between 300-400 gauss with compatible wireless patient monitor.
1a | Product Offered (Anaesthetic Machine)
1a.1 | Make:
1a.2 | Model:
1a.3 | Initial Manufacturing year of Model:
1a.4 | Radiation Control License number:
1b | Product Offered (Monitor)
1b.1 | Make:
1b.2 | Model:
1b.3 | Initial Manufacturing year of Model:
1b.4 | Radiation Control License number:
2 | TECHNICAL SPECIFICATION
2.1 | The unit must include electronicaly integrated oxygen, nitrous oxide and air delivery and monitoring.The Unit must have a physically integrated magnetic field strength monitor(Gauss Alarm)
2.2 | Ventilation and anaesthesia vapour monitoring must be fully integrated within the unit.
2.3 | The anaesthesia unit must be able to communicate/interface with a physiological monitor
2.4 | The anaesthesia system must comprise:
2.5 | Anaesthesia system suitable for neonatal (down to 500g), paediatric and adult anaesthesia.
2.6 | Ventilator suitable for ventilating neonatal (down to 500g), paediatric, adult and morbid/obese patients
3 | Multigas analyser: Must be available either on Anaesthetic machine or Physiological monitor
3.1 | Oxygen
3.2 | Nitrous oxide
3.3 | Carbon dioxide
3.4 | Anaesthetic vapour monitoring:
3.4.1 | Halothane
3.4.2 | Isoflurane
3.4.3 | Sevoflurane
3.4.4 | Desflurane
3.5 | Complete and automated functional test/check requiring no user intervention or disassembly.
3.6 | The functional test/check must provide the measurements and values for automatic leakage and compliance.
3.7 | Automatic calibration of sensors for anaesthetic gases and vapours must form part of the self test and require no user intervention.
3.8 | It must be possible to override the functional test/check facility. The system must display the override status
3.9 | The unit must provide conventional, low-flow (down to 200 ml/min) anaesthesia.
3.1 | The unit must return analysed gas back to the circle system
3.11 | A safety oxygen function is required for the delivery of oxygen through the anaesthesia unit in the event of an electronic failure of the gas delivery system.
3.12 | The unit must be supplied with a minimum of 1 storage drawer
3.13 | The unit must be supplied with a minimum of 3 additional 220 V power sockets. These sockets must be excluded from the battery backup circuit in case of a power failure
3.14 | The unit must have a working area or a retractable shelf. Width: not less than 300 mm. Depth: not less than 200 mm
3.15 | The unit must have a light to enable the anaesthetist to work in a darkened theatre.
4 | Trolley Made out of MRI compatable material
4.1 | The unit must be mounted on a strong , robust and compact trolley manufactured from corrosion free material. The material must be MRI compatible.
4.2 | The trolley must have no sharp corners and be easy to clean.
4.3 | The trolley must have at least four antistatic castors of diameter 125 mm or greater.
4.4 | Trolley must have two lockable castors or a central locking brake.
4.5 | The unit must accommodate an oxygen cylinder. Cylinder size according to user requirement. Size to be compatible with current state vendor of oxygen and cylinder must be MRI compatible (made out of alluminium).
4.6 | Trolley must have accessible handles for ease of movement
4.7 | The trolley, with all components mounted, must be stable without toppling over when placed on a 30° slope.
4.8 | The unit must be supplied with a castor guard system, preventing the entaglement and trapping of cables within the wheels.
4.9 | The trolley must have a mounting system suitable for a physiologial monitor.
4.1 | Pipeline and Cylinder Gas Supplies All made out of MRI compatible material. The anaesthesia system must be designed to use external piped medical gasses (O2, N2O, medical air), vacuum (scavenging) from medical service outlet sockets.
4.11 | Connection hoses for above must be provided and must be fitted with Heyer type non-interchangeable probes for the existing medical gas outlets.
4.12 | Hoses must be fitted with Heyer type non-interchangeable probes for the existing medical gas outlets.
4.13 | The medical air hose connection must include a water trap
4.14 | A safety pressure relief valve of 60 - 80 cm H2O must be provided to protect the system from excessive pressure .
4.15 | The unit must function safely and correctly within the pipeline supply pressure range of: 270 - 600 kPa
4.16 | The system must preferentially utilise the piped oxygen supply over the cylinder supply, if pipeline pressure is > 270 kPa
5 | Connection hoses must be Flexible and at least 6 metres long:
5.1 | Antistatic
5.2 | Colour coded as follows:
5.2.1 | O2: white
5.2.2 | N2O: blue
5.2.3 | Vacuum: yellow
5.2.4 | Medical air: black
5.2.5 | Scavenge low vacuum: orange or yellow
5.3 | Where relevant, fitted with outlet probes complying with SANS 1409:2014
5.4 | Permanently bonded to the outlet probe and the machine connector
5.5 | Hoses for gases above atmospheric pressure, must be designed for pressures of at least 1 000 kPa.
5.6 | Vacuum hoses may not collapse or kink at 90 kPa below atmospheric pressure, when subjected to 180° bend of inside radius equal to the outside diameter of the hose.
5.7 | All connection hoses must be fed to a central manifold.
5.8 | All high pressure gas and vacuum pipe connections (other than hoses) and all pipes connected to pressure regulators and pressure gauges must be stainless steel (18/8 grade) pipes that are oil-free and colour coded as for hoses (see above) over their entire lengths
5.9 | An in-line non-return valve must be fitted to each gas pipeline on the high pressure (inlet) line
5.1 | An in-line non-return valve must be fitted to each gas cylinder yoke
5.11 | A replaceable gas filter, to block particles greater than 65 µm, must be fitted on the upstream side of the non-return valve
5.12 | Each gas pipeline inlet must have a pressure gauge or electronic display of pressure. The pressure range (415 kPa ± 75 kPa) must be clearly indicated in colour
5.13 | Must be fitted with a gas cylinder yoke to accommodate one oxygen gas cylinder with a pin index system, MRI compatible
5.14 | Each gas cylinder yoke must have a pressure gauge or electronic display of pressure
5.15 | The pressure gauges or electronic displays of pressure for both pipeline & cylinder supplies, calibrated in kPa for each line, must be labelled and visible from the front of the unit.
5.16 | O2 and N2O regulators must be provided complying with CKS 605:1987.
5.17 | A pressure regulator complying with CKS 605 must be supplied for each gas inlet and must reduce the gas pressure to 415 kPa without freezing up, creeping, chattering or whistling
5.18 | Outlet sockets and probes must comply with SANS 1409:2014.
5.19 | Sockets with hoses must be colour-coded as per SANS 1409:2014:
5.2 | Oxygen: White background with black letters
5.21 | Nitrous oxide: Blue background with white letters
5.22 | Medical air: One half black with white letters, the other half white with black letters
5.23 | Vacuum: Yellow background with black letters
5.24 | Continuous display, visible from the front of the machine, of both pipeline and cylinder pressures even during power failure
5.25 | Pipeline and cylinder pressures must be differentially displayed, and clearly indicated
5.26 | Safety pressure relief valves must protect the entire pneumatic system from excessive input pressures from pipeline and cylinder supplies. Safety relief valve for all gases must be supplied and set at 600kPa
5.27 | When a cylinder is detached from the pin-indexed yoke no gas leak must occur at the yoke.
5.28 | Airway pressure must be monitored and displayed
5.29 | Scavenging-MRI compatible material
5.3 | Scavenging-MRI compatible material
5.31 | The unit must have a scavenging system with a hose at least 5 metres long with the suitable connector as per the institutions scavenging installation.
5.32 | Air-break device to prevent negative pressure on the breathing system
5.33 | If an active sytem is supplied a pressure indicator must be present
5.34 | The scavenging outlet must be a 30mm ISO connector.
6 | Patient Gas Delivery System
6.1 | Patient gas delivery must be electronically controlled\Rotameters
6.2 | Must display the delivered patient gas flow using an ISO colour standard for each gas.
6.3 | The unit must be able to deliver a total patient gas flow of 200 ml/min to 18 l/min (Minimum). The unit must have the ability to deliver at least 10 l/min of oxygen. These must be incremental adjustments
6.4 | Gas supply must conform to:
6.5 | O2 minimum flow range of 200 ml/min to 10 l/min. Accuracy ± 5%
6.6 | medical air minimum flow range 0 to 10 l/min. Accuracy ± 5%
6.7 | nitrous oxide minimum flow range 0 to 10 l/min. Accuracy ± 5%
6.8 | Flowmeter sequence to be according to standards approved by the South Africal Society of Anaesthesiologists - Oxygen on the right
6.9 | The delivery system must prevent simultaneous delivery of two anaesthetic vapours.
6.1 | A separate common gas outlet to accept a standard 22 mm outside diameter and 15 mm internal diameter.
7 | Hypoxic guard
7.1 | The unit must include a safety mechanism that ensures a minimum concentration of 25% oxygen in any oxygen/nitrous oxide gas mixture across all flow rates.
7.2 | The unit must include a safety mechanism that ensures a minimum concentration of 21% oxygen in any oxygen/air gas mixture across all flow rates.
7.3 | Accidental simultaneous delivery of N2O and Air into the breathing system, must either be prevented by means of an interlock switch or clearly indicated so that the anaesthetist can rectify the error. Delivery indicated purely by mechanical rotameters will not achieve the requirements of this clause.
7.4 | The unit must include a high-flow oxygen flush mechanism with a flow rate of at least 35 l/min
7.5 | Oxygen flush must be non-lockable, recessed and self closing
7.6 | This system must bypass the flowmeters and vaporisers
8 | Breathing System
8.1 | Total circuit volume, including absorber, must not exceed 2,7 litres
8.2 | With APL valve 5 - 75 cm H2O
8.3 | With ventilator to manual switch. This switch when directed to the ventilator must isolate the APL valve from the circle.
8.4 | Breathing system must be fully autoclavable to 134°C
8.5 | There must be a mechanism to remove water from the breathing system
8.6 | Breathing system must be latex free
8.7 | Bag-to-ventilator switch automatically begins in the ventilator position
8.8 | Reusable corrugated breathing circuit to be supplied
8.9 | Bidders to supply the breathing circuit as required by the user
9 | Integrated Ventilator
9.1 | Ventilator suitable for ventilating adults, paediatric and neonates (> 500g) patients without changing ventilator components (e.g. bellows)
9.2 | The ventilator must be electronically controlled
9.3 | The ventilation technology offered may be turbine, volume reflector, piston or ascending bellows
9.4 | The minimum basic modes to be supplied:
9.5 | Manual/Spontaneous Ventilation
9.6 | Volume Control Ventilation
9.7 | Pressure Control Ventilation
9.8 | Pressure Support with Mandatory Apnoea back-up ventilation
9.1 | Volume Controlled Ventilation + SIMV + Pressure Support
9.11 | Pressure Control Ventilation + SIMV + Presssure Support
9.12 | CPAP
9.13 | Additional modes of ventilation may be offered and quoted for separately
9.14 | The ventilator must be integrated with the workstation and user friendly
10 | The ventilator must meet the following minimum criteria:
10.1 | Frequency: minimum range of 5 – 30 bpm.
10.2 | Inspiratory time: minimum range of 0.2 – 10.0 seconds.
10.3 | Minimum Inspiration/Expiration Ratio: 4:1 – 1:4.
10.4 | Adjustable inspiratory pause/plateau time (volume control ventilation): 0-30% of inspiratory time
10.5 | Peak Inspiratory Pressure: minimum range of: PEEP + 5 to 30 cm H2O.
10.6 | Positive End Expiratory Pressure (PEEP): Integrated, electronically controlled. Minimum range: 0 – 30 cm H2O.
10.7 | Tidal Volume: 20ml – 1500ml per breath. Minimum accuracy: 15%
10.8 | Peak Inspiratory Flow: At least 120 /pm
10.9 | The ventilator must be capable of flow or pressure triggering.
10.1 | Adjustable Inspiratory Rise time/Ramp with a range of 0.1 - 2 sec
10.11 | Compliance compensation/Tidal volume compensation/Compressible volume compensation or equivalent.
10.12 | Patient gas flow compensation must be automatic so that alterations in patient gas flow does not affect delivered tidal volumes
10.13 | It must be possible to set CPAP in manual/spontaneous mode within a range of: 1 to 10 cm H2O
10.14 | In event of power or technical failure, manual ventilation must be possible.
11 | Display:
11.1 | All displays must be in English
11.2 | Colour display must include numerical and graphic display of parameters measured
11.3 | Graphic display to feature at least 3 waveforms: - pressure, flow and volume.
11.4 | Must also display flow-volume and pressure-volume loops with a reference loop
12 | Alarms : The ventilator must have the following user adjustable alarms
12.1 | Apnoea alarm
12.2 | High airway pressure alarm
12.3 | Continuous pressure alarm
12.4 | High and low expiratory tidal volume alarm
12.5 | Minute volume, high/low alarm
12.6 | Inspired oxygen (Fi02), high/low alarm
12.7 | PEEP, low alarm
12.8 | Leakage / Check breathing circuit alarm or equivalent
12.9 | Alarm profile must be user definable or default to patient appropriate settings
12.1 | The ventilator must have a mute alarm function, active only for a limited time (Maximum 2minutes)
12.11 | It must not be possible to cancel the visual alarm until the problem has been rectified.
12.12 | All components of the ventilator and breathing system in contact with the patient’s airway must be latex free.
12.13 | All components of the ventilator and breathing system in contact with the patient’s airway must be steam autoclavable up to 134ᴼ C
Integrated Gas Monitoring
13 | Oxygen Monitoring
13.1 | The system must display both inspired and end-tidal oxygen concentration.
13.2 | The oxygen analyser must have user adjustable low and high alarms – visual and audible. The low alarm must automatically default to no less than 21%.
13.3 | The system must be accurate (+/- 2% in 21-100%) range, and have a rapid response time (≤ 500ms).
13.4 | The unit must sound a continuous alarm if the oxygen pressure falls below 190 kPa, then all nitrous oxide must be automatically cut from the patient circuit.
14 | Anaesthetic vapour, carbon dioxide and nitrous oxide monitoring:
14.1 | Anaesthetic vapour, N2O monitoring and capnography must be electronically integrated into the anaesthetic workstation.
14.2 | The monitoring must be of the side-stream type.
14.3 | Sampling: 100 - 200 ml/min
14.4 | Sampled gases must be internally returned to the breathing system.
14.5 | CO2 must be presented as a numerical display of inspired and end-tidal concentration and a capnography waveform.The displayed units must be user-selectable: kPa or mmHg.
14.6 | The monitoring system must have automatic identification of anaesthetic vapours.
14.7 | The system must have a graphical display indicating the type of anaesthetic vapour/s currently in use and the concentration thereof. The colour coding must be as per international colour coding for anaesthetic vapours.
14.8 | The anaestheric vapour must be displayed as vol % and must include inspiratory and expiratory concentrations.
14.9 | The monitoring system must be able to simultaneously detect, identify and display mixed anaesthetic vapours.
14.1 | High and low anaesthetic vapour alarms must be present and these must be user-definable.
14.11 | There must be automatic calculation of minimum alveolar concentration.
14.12 | The system must be able to calculate anaesthetic gas utilization in ml per case
14.13 | The unit must have an indication of the level of vapour in the vaporiser
14.14 | Integrated vapour analysis must deliver an audible "low vapour " alarm and a critical alarm for an empty tank
15 | Circle System
15.1 | The circle absorber system must be integrated as part of the ventilation system.
15.2 | The circle absorber system must have an effective system to eliminate excess water created by the chemical reactions either by heating or water traps.
15.3 | The circle absorber system must be autoclavable at 134ºC.
15.4 | The circle absorber system must be supplied with a carbon dioxide absorber canister.
15.5 | The absorber canister may be reusable/refillable
15.6 | It must be easy to remove and replace the absorber canister while the workstation is in operation.
15.7 | Misalignment of the canister must be readily apparent to the user.
15.8 | The system must have a circle bypass facility while changing/refilling the absorber canister.
15.9 | The breathing circuit must contain an integrated adjustable pressure limiting (APL) valve with a minimum range of 0-70 cm H2O.
15.1 | The APL valve must have a tactile or electronic indicator when a value of between 20-30 cm H2O is exceeded.
15.11 | The circle absorber system must be latex free.
15.12 | The bag-to-ventilator switching mechanism must be clearly labeled and user-friendly. Mechanical or electronic
16 | General
16.1 | The unit must operate from both the 220 V, 50 Hz mains supply and internal rechargeable batteries
16.2 | The batteries must be charged while the unit is operating from the mains supply.
16.3 | The batteries must automatically take over in the event of a mains power failure.
16.4 | The batteries must have the capacity for full operation of the anaesthetic machine and ventilator for at least 90 minutes with batteries fully charged, under typical operating conditions.
16.5 | Graphic battery power indicator
16.6 | Recharge time < 8 hrs
16.7 | Mains cable at least six metres long with SABS approved 15A three prong plug.
16.8 | The unit must have a rail for attaching a suction regulator.
17 | Data management:
17.1 | The workstation must be capable of downloading data to a USB device or memory card.
17.2 | The Unit must have an intergated field strength monitor with acoustic, audible and visual alarms for Advisory,Caution and Warning.
17.3 | A laser printer must be incorporated in the unit and be MRI compatible
18 | Patient Physiological Monitor, MRI and Anaesthetic workstation compatible
18.1 | The patient monitor must be compatible with the Anaesthesia work station and the MRI equipment, wireless components where possible.
18.2 | The patient monitor must be modular or pre-configured and suitable for use for neonatal, paediatric and adult patients.
18.3 | The following paramaters must be monitored:
18.3.1 | - ECG: 3-lead and 5-lead. Must be capable of displaying 2 channels simultaneously.
18.3.2 | - ST-segment and arrhythmia analysis.
18.3.3 | - NIBP
18.3.4 | - Oxygen saturation (SpO2)
18.3.5 | - Plethysmography/respiratory rate
18.3.6 | - 2 x Temperature
18.3.7 | - At least 2 x Invasive pressure
18.4 | Display must be no less than 15" diagonally.
18.5 | Minimum 6 waveforms must be able to be displayed simultaneously.
18.6 | Full colour display.
18.7 | The front panel must be easy to clean and disinfect.
18.8 | The display must be flicker free and easy to read.
18.9 | Brightness control must be possible.
18.1 | In addition to the waveform display section, the monitor must have alpha-numeric display of parameter identification, measured values, alarm and status messages.
18.11 | Audible and visual alarms must be available for all measured parameters.
18.12 | Default alarm limits if the limits are not set by the user
18.13 | Alarm indication: Visual and audible with mute
18.14 | Adjustable audio alarm volume
18.15 | Low battery: Alarm at least 10 minutes before battery power is exhausted
18.16 | It must be possible to mute the audible alarm for a short period, but no override of the visual alarm must be possible.
18.17 | It must be possible to suspend the alarms of certain parameters during clinical procedures, but the alarm must reactivate automatically after a suitable time delay. Any silencing or disabling of alarms must be clearly indicated on the monitor.
18.18 | Input impedance of the system > 10 MΩ.
18.19 | Maximum allowable leakage current to the patient must be less than 10 µA.
18.2 | The monitor must be protected against defibrillator and electrosurgical interference.
18.21 | The monitor must be of low power consumption, and must not require fan or other forced ventilation.
18.22 | All cables and sockets must be colour coded for easy recognition, and must conform to international standards
18.23 | Display of trends for at least 8 hours in tabular and graphical format.
18.24 | The trend resolution must be user-adjustable and have a minimum resolution of 1 min.
19 | ECG
19.1 | A full set of 3/5 lead ECG cables must be included in the offer.
19.2 | The ECG amplifier must employ a differential input with high noise suppression.
19.3 | Sweep speed: 25 to 50 mm/sec.
19.4 | Freeze facility for ECG.
19.5 | The input impedance must be > 10 MΩ.
19.6 | The heart rate measuring range: 20 to 300 bpm, with an accuracy < 5 bpm.
19.7 | A 50 Hz filter must be provided.
19.8 | Alarm limits must be adjustable in the following range:
19.8.1 | Low: 40 to 150
19.8.2 | High: 70 to 250.
20 | Non Invasive Blood Pressure
20.1 | Must be suitable for adults, paeds and neonates
20.2 | The unit must use the oscillometric method of measurement.
20.3 | Modes: Manual and automatic
20.4 | Reusable neonate, paediatric, and adult cuffs must be supplied
20.5 | The cuffs must be robust, interchangeable and easy to clean.
20.6 | Must display systolic, diastolic and mean pressures for each measurement
20.7 | Adult: Systolic: 30 - 270 mm Hg, Diastolic: 20 - 240 mm Hg, MAP: 20 - 250 mm Hg
20.8 | Paeds: Systolic: 30 - 160 mm Hg, Diastolic: 20 - 120, MAP: 20 - 120 mm Hg
20.9 | Values must be displayed numerically
20.1 | Auto zero when unit is switched on
20.11 | The interval between measurements must be selectable from 1 – 120 minutes.
20.12 | Adjustable high and low alarm limits for systolic, diastolic, mean arterial pressure and heart rate.
20.13 | Over pressure protection: Automatic deflation at 300 mm Hg adult, 150 mm Hg neonate
20.14 | Average measurement time: approximately 30 s
20.15 | If measurement exceeds two minutes the cuff must deflate, an alarm must sound and likely faults indicated.
20.16 | The NIBP hose must be at least two metres long.
21 | Oxygen Saturation.
21.1 | Continuous monitoring of oxygen saturation levels by means of the photometric principle and wireless technology
21.2 | Must be suitable for adults, paeds and neonates
21.3 | The unit must continuously display the SpO2 % waveform
21.4 | Oxygen saturation (SpO2 %): Range of measurement: 0 - 100%
21.5 | The accuracy must be better than 2,5 % (70 - 100%) and 5% (50 - 70%).
21.6 | Pulse rate (beats/min): Range of measurement: 20 - 300 bpm
21.7 | The pulse rate accuracy must be better than ± 2 %.
21.8 | An audible bleep should sound for each heartbeat, the pitch of which should vary with the oxygen saturation.
21.9 | Automatic rejection of variuos artefacts
22 | SpO2 Sensor:
22.1 | Wireless Telemetry
22.2 | It must be the latest technology with clinical research support based on accuracy and real-time monitoring. Please submit the published peer reviewed/researched paper as supporting documents.
22.3 | The following reusable sensors must be provided:
22.4 | Finger sensor for adults and paeds.
22.5 | Earlobe sensor for adults and paeds.
22.6 | Finger wrap-around probe for neonates
22.7 | The sensor lead must be approximately 3 metres long if not wireless
23 | Temperature
23.1 | Temperature range must be approximately 15º C to 45º C.
23.2 | Two reusable temperature probes must be included
24 | Accessories and Consumables
24.1 | A full set of basic accessories for adult/paediatric/neonate must be supplied with the system,appropriate for use in an MRI suite
25 | STANDARD ACCESSORIES
25.1 | The unit must be supplied with a starter pack for full functionality upon commissioning.
25.2 | The unit must be delivered with a set of adult (A), paediatric (P) and neonatal (N) accessories:
25.2.0 | Invasive pressure cables compatible with the transducer system available at the institution
25.2.1 | SpO2 probe complete (A, P, N) x2
25.2.2 | ECG cable complete (A, P, N) x2
25.2.3 | Temperature probe x2
25.2.4 | NIBP hose and cuff (A, P, N) x2
25.2.5 | Patient circuit (A, P, N) - box of 10 each
25.2.6 | Filters (box of 10)
25.2.7 | Inflatable reusable face mask (A, P, N) - 3 full set of each
25.2.8 | Spirometery (box of 10 at least)
25.2.9 | Sampling line (box of 10 atleast)
25.2.10 | Flow sensors (box of 10 taleast)
25.2.11 | CO2 absorber canister disposable reusable x1 and disposable (box of 10 atleast)
25.2.12 | Vaporizer key-fillers (Sevo, & other choice below)
25.3 | The unit must be supplied with two vaporizers (one must sevoflurane, the other must be of the choice of the buying institution)
25.4 | Water traps (box of 10 at least)
27 | Vaporisers\MRI compatible(Tec 5 or 7)
27.1 | Vaporisers must be tilt, pressure and temperature compensated
27.2 | Mechanical, electronic or injection technology vaporisers are acceptable
27.3 | Filling mechanism:
27.3.1 | Vaporiser filling mechanism must be compatible with the current state approved vendor of inhalational volatile vapours.
27.3.2 | Each vaporiser must be supplied with 2 compatible filling adapters as per user requirement
27.3.3 | Minimum accuracy of the supplied volatile vapour delivery system must be at least 20% of set value
27.3.4 | Visible level indicator
28 | Auxilliary Delivery System
28.1 | An external auxilliary flowmeter suitable for delivery of oxygen via a face mask, supplied with a 6mm tapered connection. Flow range 0-10 l/min and above
29 | Trending:
29.1 | Trend information and event recording for:
29.1.1 | peak pressure
29.1.2 | plateau pressure
29.1.3 | PEEP
29.1.4 | expiratory tidal volume
29.1.5 | respiratory rate
29.2 | This must be provided in graphical and tabular formats.
29.3 | The trend and event information must cover at least the last 24 hours of operation.
29.4 | Resolution must be adjustable, with a minimum resolution of 1 minute intervals.
29.5 | The adjustment of the resolution must not affect the display of parameters which were sampled at different times to the resolution setting in such a way as to display a null reading, when one was in fact obtained.
30 | Software
30.1 | Any software for upgrades may be offered
31 | Suction Regulator
31.1 | A suction regulator capable of -50 kPa (minimum)
31.2 | The suction regulator must have a built-in diaphragm with a 0 – 100 kPa vacuum gauge and an “on-off” switch.
31.3 | It must aspirate 1 litre water in 45 seconds at a pressure of at least - 50 kPa
31.4 | A collection jar calibrated to at least 1 500 ml must be included.
31.5 | The collection jar must accommodate a disposable liner, as per user requirement
32 | Conformity Compliance (Please attach certificate)
32.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
32.2 | The unit must comply with electrical safety and performance standard such as IEC 60601-2-13 for Anaesthetic systems, attached certification
32.2.1 | Please provide unique ref number of the IEC ceriticate:
32.3 | The unit must comply with electrical safety and performance standard such as IEC 60601-2-49 for Patient Monitoring, attached certification
32.3.1 | Please provide unique ref number of the IEC ceriticate:
32.4 | OEM must comply and certified on ISO 9001 quality standards, attach proof of compliance
32.4.1 | Please provide unique ref number of the ISO 9001 ceriticate:
32.5 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
32.5.1 | Please provide unique ref number of the ISO 13485 ceriticate:
32.6 | Model quoted for must be CE certified. Attach a copy of certification
32.6.1 | Please provide unique ref number of the CE ceriticate:
32.7 | Hazardous Substance Act: If this type of equipment / apparatus appears on the schedule of Hazardous substances issued by the Directorate: Health Technology of the Department of Health, a license in terms of the Act on Hazardous Substances (Act 15/1973) must be submitted with the bid / quotation documents. The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
32.8 | All electrical/electronic medical equipment must be licensed by Radiation Control, where applicable a copy of the license must be submitted The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
32.8.1 | Radiation Control License number:
32.9 | The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade (rubber to with 2 screws) type and it must be a minimum length of three (3) meters. NB the mains cable of the unit tendered for must be SABS colour coded.
32.10 | The equipment quoted for must be protected against electromagnetic interference IEC 60601-1-2
33 | COMPREHENSIVE MAINTENANCE
33.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) must be included in the bid price
33.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries, Flow sensors and Oxygen sensor), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
33.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
33.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries, Flow sensors and Oxygen sensor), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
33.3 | The bidder to provide the OEM recommended service interval of the unit:
33.4 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
34 | General
34.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
34.2 | The unit on offer must be the latest technology.
ACCESSORIES FOR ANAESTHETIC DELIVERY SYSTEM WITH PHYSIOLOGICAL MONITOR MRI COMPATABLE- Accessories and Consumables
RT2-01-09a | A full set paediatric 3/5 lead ECG cable
RT2-01-09b | A full set adult 3/5 lead ECG cable
RT2-01-09c | Reusable paediatric cuff
RT2-01-09d | Reusable adult cuff
RT2-01-09e | Reusable neonate cuff
RT2-01-09f | Reusable obese cuff
RT2-01-09g | Air hose
RT2-01-09h | Finger sensor for paeds
RT2-01-09i | Finger sensor for adults
RT2-01-09j | Earlobe sensor for paeds
RT2-01-09k | Earlobe sensor for adults
RT2-01-09l | Finger wrap-around probe for neonates
RT2-01-09m | Sensor lead
RT2-01-09n | Reusable temperature probe
Bidder to list any other Consumables or Accessories (Not listed above) for the model on offer below

## Monitor, MRI Compatible
RT2-01-10 | PATIENT MONITOR COMPATABLE TO MRI, with Multigas analysis
1 | DESCRIPTION
1.1 | Patient physiological haemodynamic monitor compitable with MRI and Anaesthetic workstation
1a | Product Offered:
1a.1 | Make:
1a.2 | Model:
1a.3 | Initial Manufacturing year of Model:
1a.4 | Radiation Control License number:
2 | TECHNICAL SPECIFICATION
2.1 | Patient Physiological Monitor, MRI and Anaesthetic workstation compatible
2.2 | The patient monitor must be compatible with the Anaesthesia work station and the MRI equipment, wireless components where possible.
2.3 | The patient monitor must be modular or pre-configured and suitable for use for neonatal, paediatric and adult patients.
2.4 | The following paramaters must be monitored:
2.4.1 | - ECG: 3-lead and 5-lead. Must be capable of displaying 2 channels simultaneously.
2.4.2 | - ST-segment and arrhythmia analysis.
2.4.3 | - NIBP
2.4.4 | - Oxygen saturation (SpO2)
2.4.5 | - Plethysmography/respiratory rate
2.4.6 | - 2 x Temperature
2.4.7 | - At least 2 x Invasive pressure
3 | Multigas analyser
3.1 | Oxygen
3.2 | Nitrous oxide
3.3 | Carbon dioxide
3.4 | Anaesthetic vapour monitoring:
3.4.1 | Halothane
3.4.2 | Isoflurane
3.4.3 | Sevoflurane
3.4.4 | Desflurane
3.5 | Display must be no less than 15" diagonally.
3.6 | Minimum 6 waveforms must be able to be displayed simultaneously.
3.7 | Full colour display.
3.8 | The front panel must be easy to clean and disinfect.
3.9 | The display must be flicker free and easy to read.
3.1 | Brightness control must be possible.
3.11 | In addition to the waveform display section, the monitor must have alpha-numeric display of parameter identification, measured values, alarm and status messages.
3.12 | Audible and visual alarms must be available for all measured parameters.
3.13 | Default alarm limits if the limits are not set by the user
3.14 | Alarm indication: Visual and audible with mute
3.15 | Adjustable audio alarm volume
3.16 | Low battery: Alarm at least 10 minutes before battery power is exhausted
3.17 | It must be possible to mute the audible alarm for a short period, but no override of the visual alarm must be possible.
3.17.1 | It must be possible to suspend the alarms of certain parameters during clinical procedures, but the alarm must reactivate automatically after a suitable time delay. Any silencing or disabling of alarms must be clearly indicated on the monitor.
3.18 | Input impedance of the system > 10 MΩ.
3.19 | Maximum allowable leakage current to the patient must be less than 10 µA.
3.2 | The monitor must be protected against defibrillator and electrosurgical interference.
3.21 | The monitor must be of low power consumption, and must not require fan or other forced ventilation.
3.22 | All cables and sockets must be colour coded for easy recognition, and must conform to international standards
3.23 | Display of trends for at least 72 hours in tabular and graphical format.
3.24 | The trend resolution must be user-adjustable and have a minimum resolution of 1 min.
4 | ECG
4.1 | A full set of 3/5 lead ECG cables must be included in the offer.
4.2 | The ECG amplifier must employ a differential input with high noise suppression.
4.3 | Sweep speed: 25 to 50 mm/sec.
4.4 | Freeze facility for ECG.
4.5 | The input impedance must be > 10 MΩ.
4.6 | The heart rate measuring range: 20 to 300 bpm, with an accuracy < 5 bpm.
4.7 | A 50 Hz filter must be provided.
4.8 | Alarm limits must be adjustable in the following range:
4.8.1 | Low: 40 to 150
4.8.2 | High: 70 to 250.
5 | Non Invasive Blood Pressure
5.1 | Must be suitable for adults, paeds and neonates
5.2 | The unit must use the oscillometric method of measurement.
5.3 | Modes: Manual and automatic
5.4 | Reusable neonate, paediatric, adult and obese cuffs must be supplied
5.5 | The cuffs must be robust, interchangeable and easy to clean.
5.6 | Must display systolic, diastolic and mean pressures for each measurement
5.7 | Adult: Systolic: 30 - 270 mm Hg, Diastolic: 20 - 240 mm Hg, MAP: 20 - 250 mm Hg
5.8 | Paeds: Systolic: 30 - 160 mm Hg, Diastolic: 20 - 120, MAP: 20 - 120 mm Hg
5.9 | Values must be displayed numerically
5.1 | Auto zero when unit is switched on
5.11 | The interval between measurements must be selectable from 1 – 120 minutes.
5.12 | Adjustable high and low alarm limits for systolic, diastolic, mean arterial pressure and heart rate.
5.13 | Over pressure protection: Automatic deflation at 300 mm Hg adult, 150 mm Hg neonate
5.14 | Average measurement time: approximately 30 s
5.15 | If measurement exceeds two minutes the cuff must deflate, an alarm must sound and likely faults indicated.
5.16 | The NIBP hose must be at least two metres long.
6 | Oxygen Saturation.
6.1 | Continuous monitoring of oxygen saturation levels by means of the photometric principle and wireless technology
6.2 | Must be suitable for adults, paeds and neonates
6.3 | The unit must continuously display the SpO2 % waveform
6.4 | Oxygen saturation (SpO2 %): Range of measurement: 0 - 100%
6.5 | The accuracy must be better than 2,5 % (70 - 100%) and 5% (50 - 70%).
6.6 | Pulse rate (beats/min): Range of measurement: 20 - 300 bpm
6.7 | The pulse rate accuracy must be better than ± 2 %.
6.8 | An audible bleep should sound for each heartbeat, the pitch of which should vary with the oxygen saturation.
6.9 | Automatic rejection of variuos artefacts
7 | SpO2 Sensor:
7.1 | Wireless Telemetry
7.2 | It must be the latest technology with clinical research support based on accuracy and real-time monitoring. Please submit the published peer reviewed/researched paper as supporting documents.
7.3 | The following reusable sensors must be provided:
7.4 | Finger sensor for adults and paeds.
7.5 | Earlobe sensor for adults and paeds.
7.6 | Finger wrap-around probe for neonates
7.7 | The sensor lead must be approximately 3 metres long if not wireless
8 | Temperature
8.1 | Temperature range must be approximately 15º to 45º C.
8.2 | Two reusable temperature probes must be included
9 | STANDARD ACCESSORIES
9.1 | The unit must be supplied with a starter pack for full functionality upon commissioning.
9.2 | The unit must be delivered with a set of adult (A), paediatric (P) and neonatal (N) consumables:
9.2.1 | SpO2 probe complete (A, P, N) x2
9.2.2 | ECG cable complete (A, P, N) x2
9.2.3 | Temperature probe x2
9.2.4 | NIBP hose and cuff (A, P, N) x2
11 | Conformity Compliance (Please attach certificate)
11.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
11.2 | The unit must comply with electrical safety and performance standard such as IEC 60601-2-49 for Patient Monitoring, attached certification
11.2.1 | Please provide unique ref number of the IEC ceriticate:
11.3 | OEM must comply and certified on ISO 9001 quality standards, attach proof of compliance
11.3.1 | Please provide unique ref number of the ISO 9001 ceriticate:
11.4 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
11.4.1 | Please provide unique ref number of the ISO 13485 ceriticate:
11.5 | Model quoted for must be EC certified. Attach a copy of certification
11.5.1 | Please provide unique ref number of the EC ceriticate:
11.6 | Hazardous Substance Act: If this type of equipment / apparatus appears on the schedule of Hazardous substances issued by the Directorate: Health Technology of the Department of Health, a license in terms of the Act on Hazardous Substances (Act 15/1973) must be submitted with the bid / quotation documents. The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
11.7 | All electrical/electronic medical equipment must be licensed by Radiation Control, where applicable a copy of the license must be submitted The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
11.7.1 | Radiation Control License number:
11.8 | The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade (rubber to with 2 screws) type and it must be a minimum length of three (3) meters. NB the mains cable of the unit tendered for must be SABS colour coded.
11.9 | The equipment quoted for must be protected against electromagnetic interference IEC 60601-1-2
12 | COMPREHENSIVE MAINTENANCE
12.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) must be included in the bid price
12.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
12.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
12.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
12.3 | The bidder to provide the OEM recommended service interval of the unit:
12.4 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
13 | General
13.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
13.2 | The unit on offer must be the latest technology.
ACCESSORIES FOR PATIENT MONITOR COMPATABLE TO MRI, with Multigas analysis
RT2-01-10a | A full set paediatric 3/5 lead ECG cable
RT2-01-10b | A full set adult 3/5 lead ECG cable
RT2-01-10c | Reusable paediatric cuff
RT2-01-10d | Reusable adult cuff
RT2-01-10e | Reusable neonate cuff
RT2-01-10f | Reusable obese cuff
RT2-01-10g | Air hose
RT2-01-10h | Finger sensor for paeds
RT2-01-10i | Finger sensor for adults
RT2-01-10j | Earlobe sensor for paeds
RT2-01-10k | Earlobe sensor for adults
RT2-01-10l | Finger wrap-around probe for neonates
RT2-01-10m | Sensor lead
RT2-01-10n | Reusable temperature probe
UPGRADES
RT2-01-10o | Neuromuscular transmission monitoring
RT2-01-10p | Metabolic monitoring
RT2-01-10q | Electroencephalography (EEG) and depth of anaesthesia monitoring system
RT2-01-10r | Invasive cardiac output monitoring
RT2-01-10s | Non-invasive cardiac output monitoring
RT2-01-10t | Cerebral oximetry
Bidder to list any other Consumables or Accessories (Not listed above) for the model on offer below

## Video Laryngoscope
RT2-01-11 | INTUBATION VIDEO LARYNGOSCOPE
1 | DESCRIPTION
1.1 | A robust video laryngoscopy system to facilitate visualization of the larynx and then tracheal intubation by means of Intubation laryngoscope
1a | Product Offered:
1a.1 | Make:
1a.2 | Model:
1a.3 | Initial Manufacturing year of Model:
1a.4 | Radiation Control License number:
2 | TECHNICAL SPECIFICATION
2.1 | The unit must be mounted on a mobile stand.
2.2 | The unit must have facility to be used on all patients ranging from preterm neonatal (from 600g) to morbidly obese adults
2.3 | The handheld unit will be attached to an electronic display to facilitate visualization of the larynx
2.4 | The device should incorporate NTSC colour video or latest technology.
2.5 | Slimline durable handheld intubation scope similar in design to a laryngoscope handle and blade system
3 | The Blade should have:
3.1 | The acutely angled blade must have an angulation of at least 60⁰ and be available for adult and paediatrics sizes
3.2 | An automatic anti fogging mechanism and anti glare technology in order to ensure a clear image is displayed on the display unit during use. This anti fogging mechanism will not require any action on the part of the user to function.
3.3 | An automatic autofocus mechanism in order to ensure a clear image is displayed on the display unit during use.
3.4 | Fully reusable and disposable blades will be accepted.
3.5 | The device should be supplied with one full set of at least three blades (neonatal or paediatric or adult) of the choice of the end user.
3.6 | The full video laryngoscope blade, meaning both the length of the blade and the blade handle, must be able to be sterilised. The Bidder to clearly state the different sterilization methods that may be performed on the offered laryngoscope type blade and other components. Suitable wall cleaning charts are to be provided to ensure appropriate cleaning and unit maintenance of the device.
4 | The offered video intubation system shall consist of the following components which are fully intergrated into the system (not limitted to):
4.1 | Slimline durable handheld intubation scope similar in design to a laryngoscope handle and blade system.
4.2 | Display and Control unit monitor; 7" - 10" TFT LCD or LED panel monitor will be acceptable.
The video screen monitor must be separate from the laryngoscope blade, must offer user friendly and allow for immediate use when switched on
4.4 | Mobile stand on castor wheels (at least two of them with brakes)
4.5 | IV pole mount for semi-permanent fixation
4.6 | Mounting adaptor to secure the monitor onto the mobile stand or IV pole mount
4.7 | Suitable blade holder capable of storing a minimum of 2 blades. The holder must be compatible for attachment to the mobile stand and IV pole mount
4.8 | The offered system shall be fully operational on delivery and will not require any additional accessories or consumable items (at least one month of consumable items must be provided) in order to function completely.
4.9 | The control panel and all control switches must be well sealed preventing entry of fluids to the internal workings and circuits of the unit.
4.1 | All menus and key indicators shall be in English and/or internationally recognised symbols.
4.11 | The display shall be separate from the laryngoscope blade and shall be of a reasonable size and have a clear image display.
4.12 | The display must have the capability to connect via a video cable for monitor slaving option
4.13 | The unit must have the ability to record the intubation process either directly or with an attachment
4.14 | The Control panel must indicate whether the unit is operating off the 220 Volt, 50Hz A.C. line power or from its own internal battery power supply. The bidder should state the capacity and type of internal battery employed
4.15 | The battery power should last for 2 hrs or more.
5 | STANDARD ACCESSORIES
5.1 | Mcgill forceps x1 (size as per End-user's choice)
5.2 | 3 x blades (As per clause 3.5 above)
5.3 | Guide
5.4 | Intubation stylet/introducer
5.5 | Acutely bent blade
6 | Conformity Compliance (Please attach certificate)
6.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
6.2 | OEM must comply and certified on ISO 9001 quality standards, attach proof of compliance
6.2.1 | Please provide unique ref number of the ISO 9001 ceriticate:
6.3 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
6.3.1 | Please provide unique ref number of the ISO 13485 ceriticate:
6.4 | Model quoted for must be EC certified. Attach a copy of certification
6.4.1 | Please provide unique ref number of the EC ceriticate:
6.5 | Hazardous Substance Act: If this type of equipment / apparatus appears on the schedule of Hazardous substances issued by the Directorate: Health Technology of the Department of Health, a license in terms of the Act on Hazardous Substances (Act 15/1973) must be submitted with the bid / quotation documents. The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
6.7 | All electrical/electronic medical equipment must be licensed by Radiation Control, where applicable a copy of the license must be submitted The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
6.7.1 | Radiation Control License number:
6.8 | The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade (rubber to with 2 screws) type and it must be a minimum length of three (3) meters. NB the mains cable of the unit tendered for must be SABS colour coded.
6.9 | The equipment quoted for must be protected against electromagnetic interference IEC 60601-1-2
7 | COMPREHENSIVE WARRANTY COVER
7.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) must be included in the bid price
7.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
8 | General
8.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
8.2 | The unit on offer must be the latest technology.
INTUBATION VIDEO LARYNGOSCOPE ACCESSORIES & CONSUMABLES
RT2-01-11a | Mcgill forceps x1 (size as per End-user's choice)
RT2-01-11b | Guide
RT2-01-11c | Intubation stylet/introducer
RT2-01-11d | Acutely bent blade
RT2-01-11e | Reusable Bariatric blade
RT2-01-11f | Reusable Miller blade size 00
RT2-01-11g | Reusable Miller blade size 0
RT2-01-11h | Reusable Miller blade size 1
RT2-01-11i | Reusable Miller blade size 2
RT2-01-11j | Reusable McIntosh / Macintosh blade size 0
RT2-01-11k | Reusable McIntosh / Macintosh blade size 1
RT2-01-11l | Reusable McIntosh / Macintosh blade size 2
RT2-01-11m | Reusable McIntosh / Macintosh blade size 3
RT2-01-11n | Reusable McIntosh / Macintosh blade size 4
RT2-01-11o | Reusable Laryngoscope with 60 degrees angulation size 1
RT2-01-11p | Reusable Laryngoscope with 60 degrees angulation size 2
RT2-01-11q | Reusable Laryngoscope with 60 degrees angulation size 3
RT2-01-11r | Reusable Laryngoscope with 60 degrees angulation size 4
RT2-01-11s | A connecting cable (monitor and the blades).
RT2-01-11t | External power supply (where applicable)
RT2-01-11u | Disposable Laryngoscope with 60 degrees angulation size 1
RT2-01-11v | Disposable Laryngoscope with 60 degrees angulation size 2
RT2-01-11w | Disposable Laryngoscope with 60 degrees angulation size 3
RT2-01-11x | Disposable Laryngoscope with 60 degrees angulation size 4
Bidder to list any other Consumables or Accessories (Not listed above) for the model on offer below

## Video Laryngoscope + Broncho
RT2-01-12 | INTUBATION VIDEO LARYNGOSCOPE with Flexible bronchscope
1 | DESCRIPTION
1.1 | A robust video laryngoscopy system to facilitate visualization of the larynx and then tracheal intubation by means of an Intubation laryngoscope and / or video flexible bronchoscope.
1a | Product Offered:
1a.1 | Make:
1a.2 | Model:
1a.3 | Initial Manufacturing year of Model:
1a.4 | Radiation Control License number:
2 | TECHNICAL SPECIFICATION
2.1 | The unit must be mounted on a mobile stand.
2.2 | The unit must have facility to be used on all patients ranging from preterm neonatal (from 600g) to morbidly obese adults
2.3 | The handheld unit will be attached to an electronic display to facilitate visualization of the larynx
2.4 | The device should incorporate NTSC colour video or latest technology.
2.5 | Slimline durable handheld intubation scope similar in design to a laryngoscope handle and blade system
3 | The offered video intubation system shall consist of the following components which are fully intergrated into the system (not limitted to):
3.1 | Slimline durable handheld intubation scope similar in design to a laryngoscope handle and blade system.
3.2 | Display and Control unit monitor; at least 7" - 10" TFT LCD or LED panel monitor will be acceptable.
3.3 | Power supply (internal and external);
3.4 | Mobile stand on castor wheels (at least two of them with brakes)
3.5 | IV pole mount for semi-permanent fixation
3.6 | Mounting adaptor to secure the monitor onto the mobile stand or IV pole mount
3.7 | Suitable blade holder / basket capable of storing a minimum of two blades. The holder/ basket must be compatible for attachment to the mobile stand and IV pole mount
3.8 | The offered system shall be fully operational on delivery and will not require any additional accessories or consumable items (at least one month of consumable items) in order to function completely.
3.9 | The control panel and all control switches must be well sealed preventing entry of fluids to the internal workings and circuits of the unit.
3.10 | All menus and key indicators shall be in English and/or internationally recognised symbols.
3.11 | The display shall be separate from the laryngoscope blade and shall be of a reasonable size and have a clear image display.
3.12 | The display must have the capability to connect via a video cable for monitor slaving option
3.13 | Device must have the ability to record and capture the intubation process either directly or with an attachment
3.14 | The Control panel must indicate whether the unit is operating off the 220 Volt, 50Hz A.C. line power or from its own internal battery power supply. The bidder should state the capacity and type of internal battery employed
3.15 | The battery power should last for 2 hrs or more.
3.16 | The unit should be able to accommodate a flexible reusable bronchoscope (size as per end user request).
3.17 | Unit should be able to accommodate the video laryngoscope blades
4 | The Blade should have:
4.1 | The acutely angled blade must have an angulation of at least 60⁰ and be available for adult and paediatrics sizes
4.2 | An automatic anti fogging mechanism and anti glare technology in order to ensure a clear image is displayed on the display unit during use. This anti fogging mechanism will not require any action on the part of the user to function.
4.3 | An automatic autofocus mechanism in order to ensure a clear image is displayed on the display unit during use.
5.1 | Fully reusable and disposable blades will be accepted.
4.4 | The device should be supplied with one full set of three blades (neonatal or paediatric or adult) of the choice of the end user.
4.5 | The full video laryngoscope blade, meaning both the length of the blade and the blade handle, must be able to be sterilised. The Bidder to clearly state the different sterilization methods that may be performed on the offered laryngoscope type blade and other components. Suitable wall cleaning charts are to be provided to ensure appropriate cleaning and unit maintenance of the device.
5 | BRONCHOSCOPE
5.1 | Must have a connector cable to allow view to be projected to the monitor
5.2 | Minimum tip deflection should be 100º degrees up and 100º degrees down
5.3 | The image produced should be clearly defined without pixelation
5.4 | Must have a viewing angle of > 90º
5.5 | Must have a direction of view of 0 º
5.6 | Must have a minimum working length of ± 60 cm
5.7 | Must have an instrument channel with an ID of ± 1.2 mm
5.8 | A single channel should be available for suction with an appropriate suction controller
5.9 | The suction controller should be removable to allow alternative uses such as insufflation of oxygen or administration of medication
5.1 | Must have a distal tip OD of ± 4.0 mm
5.11 | State if any other size flexible bronchoscopes are available
6 | STANDARD ACCESSORIES
6.1 | mcgyl forceps x1
6.2 | guide
6.3 | Acutely bent blade
6.4 | Flexible bronchoscope size 4 (outer diametre 4mm, length 65cm), if there are any other size available please state.
7 | Conformity Compliance (Please attach certificate)
7.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
7.2 | OEM must comply and certified on ISO 9001 quality standards, attach proof of compliance
7.2.1 | Please provide unique ref number of the ISO 9001 ceriticate:
7.3 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
7.3.1 | Please provide unique ref number of the ISO 13485 ceriticate:
7.4 | Model quoted for must be CE certified. Attach a copy of certification
7.4.1 | Please provide unique ref number of the CE ceriticate:
7.5 | Hazardous Substance Act: If this type of equipment / apparatus appears on the schedule of Hazardous substances issued by the Directorate: Health Technology of the Department of Health, a license in terms of the Act on Hazardous Substances (Act 15/1973) must be submitted with the bid / quotation documents. The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
7.7 | All electrical/electronic medical equipment must be licensed by Radiation Control, where applicable a copy of the license must be submitted The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
7.7.1 | Radiation Control License number:
7.8 | The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade (rubber to with 2 screws) type and it must be a minimum length of three (3) meters. NB the mains cable of the unit tendered for must be SABS colour coded.
7.9 | The equipment quoted for must be protected against electromagnetic interference IEC 60601-1-2
8 | COMPREHENSIVE WARRANTY COVER
8.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) must be included in the bid price
8.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
9 | General
9.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
9.2 | The unit on offer must be the latest technology.
Accessories for INTUBATION VIDEO LARYNGOSCOPE with Flexible bronchscope
RT2-01-12a | Flexible bronchoscope size 4 (outer diametre 4mm, length 65cm), if there are any other size available please state.
Bidder to list any other Consumables or Accessories (Not listed above) for the model on offer below

## Single use flexible broncho
RT2-01-13 | INTUBATION VIDEO FLEXIBLE BRONCHOSCOPE (with disposable scopes)
1 | DESCRIPTION
1.1 | A robust video bronchoscopy system to facilitate visualization of the larynx and the trachea achieving intubation by means of a video flexible, disposable bronchoscope .
1a | Product Offered:
1a.1 | Make:
1a.2 | Model:
1a.3 | Initial Manufacturing year of Model:
1a.4 | Radiation Control License number:
2 | TECHNICAL SPECIFICATION
2.1 | The unit must be mounted on a mobile stand.
2.2 | The unit must have facility to be used on all patients ranging from preterm neonatal (from 600g) to morbidly obese adults
2.3 | The handheld unit will be attached to an electronic display to facilitate visualization of the larynx and trachea
2.4 | The device should incorporate NTSC colour video or latest technology.
2.5 | Slimline durable handheld intubation scope
3 | The offered video intubation system shall consist of the following components which are fully intergrated into the system (not limitted to):
3.1 | Slimline durable handheld intubation scope
3.3 | Display and Control unit monitor; at least 7" - 10" TFT LCD or LED panel monitor will be acceptable.
3.4 | Power supply (internal and external);
3.2 | Mobile stand on castor wheels (at least two of them with brakes)
3.2 | IV pole mount for semi-permanent fixation
3.2 | Mounting adaptor to secure the monitor onto the mobile stand or IV pole mount
3.2 | Suitable scope holder capable of storing the scope. The holder must be compatible for attachment to the mobile stand and IV pole mount
3.2 | The offered system shall be fully operational on delivery and will not require any additional accessories or consumable items in order to function completely.
3.2 | The control panel and all control switches must be well sealed preventing entry of fluids to the internal workings and circuits of the unit.
3.2 | All menus and key indicators shall be in English and/or internationally recognised symbols.
3.2 | The display shall be separate from the laryngoscope blade and shall be of a reasonable size and have a clear image display.
3.2 | The display must have the capability to connect via a video cable for monitor slaving option
3.2 | Device must have the ability to record and capture the intubation process either directly or with an attachment
3.2 | The Control panel must indicate whether the unit is operating off the 220 Volt, 50Hz A.C. line power or from its own internal battery power supply. The bidder should state the capacity and type of internal battery employed
3.2 | The battery power should last for 2 hrs or more.
3.2 | The unit should be able to accommodate a flexible reusable bronchoscope (size as per end user request).
5 | BRONCHOSCOPE
5.1 | Must have a connector cable to allow view to be projected to the monitor
5.2 | Minimum tip deflection should be 100º degrees up and 100º degrees down
5.3 | The image produced should be clearly defined without pixelation
5.4 | Must have a viewing angle of > 90º
5.5 | Must have a direction of view of 0 º
5.6 | Must have a minimum working length of ± 60 cm
5.11 | State the different sizes of flexible bronchoscopes available
6 | STANDARD ACCESSORIES
6.1 | 3 flexible bronchoscopes as per end user request
7 | OPTIONAL ACCESSORIES
7.1 | List all the accessories available with prices
8 | Conformity Compliance (Please attach certificate)
8.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
8.2 | OEM must comply and certified on ISO 9001 quality standards, attach proof of compliance
8.2.1 | Please provide unique ref number of the ISO 9001 ceriticate:
8.3 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
8.3.1 | Please provide unique ref number of the ISO 13485 ceriticate:
8.4 | Model quoted for must be CE certified. Attach a copy of certification
8.4.1 | Please provide unique ref number of the CE ceriticate:
8.5 | Hazardous Substance Act: If this type of equipment / apparatus appears on the schedule of Hazardous substances issued by the Directorate: Health Technology of the Department of Health, a license in terms of the Act on Hazardous Substances (Act 15/1973) must be submitted with the bid / quotation documents. The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
8.6 | All electrical/electronic medical equipment must be licensed by Radiation Control, where applicable a copy of the license must be submitted The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
8.6.1 | Radiation Control License number:
8.7 | The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade (rubber to with 2 screws) type and it must be a minimum length of three (3) meters. NB the mains cable of the unit tendered for must be SABS colour coded.
8.7.1 | The equipment quoted for must be protected against electromagnetic interference IEC 60601-1-2
9 | COMPREHENSIVE WARRANTY COVER
9.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) must be included in the bid price
9.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
10 | General
10.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
10.2 | The unit on offer must be the latest technology.

## Syringe Driver, TCI
RT2-01-15 | SYRINGE DRIVERS TCI
1 | DESCRIPTION
1.1 | A programmable infusion pump capable of infusing IVI administration via single channel. Suitable for adult, child, and neonatal infusions
1a | Product Offered:
1a.1 | Make:
1a.2 | Model:
1a.3 | Initial Manufacturing year of Model:
1a.4 | Radiation Control License number:
2 | TECHNICAL SPECIFICATION
2.1 | A Target Controlled Infusion (TCI) Syringe Driver programmable with computer controlled infusion capabilities that comprises of the latest technology must be supplied with a priced list of the syringes and extension sets that the unit has been calibrated on.
2.2 | The bidder must be prepared to provide testing/certification documentation to prove the compatibility of generic syringes available in the country and must be prepared to provide certification of the compitibility or incompatibility of various generic syringes for the life span of the TCI pump
2.3 | The unit must automatically sense the syringe size and give a continuous visble indication, on the front panel, of the syringe size fitted.
2.4 | The unit offered must:
2.4.1 | Be user friendly
2.4.2 | Be lightweight – State Weight
2.4.3 | Be portable
2.4.4 | Be able to be mounted/ attached to a RAIL or to a mobile drip stand with a universal adaptor
2.5 | Have a carry handle
2.6 | The unit offered must be well constructed with robust impact resistant casing and panels. The control panel, all push buttons and switches must be of a flush membrane type or an equivalent, to provide resistance for the entry of fluids in to the internal workings and circuits of the unit.
3 | Power and Batteries
3.1 | The monitor must operate off both the 220V-240V , 50Hz single phase AC power supply and an internally fitted rechargeable battery. The pump must be fused in both the LIVE and NEUTRAL.
3.2 | The unit offered must be supplied with an internal rechargeable battery.
3.3 | The internal rechargeable battery must be automatically charged when the unit is connected to a live A.C. supply
3.4 | In the event of A.C. mains failure the battery must automatically provide continuous operation
3.5 | The unit must be provided with visible indication to reflect whether the unit is being operated off the AC power or the internal rechargeable battery power
3.6 | The unit offered must be provided with circuitry, which must ensure that the internal rechargeable battery will be protected against over-charging and over-discharging
3.7 | Unit that are quoted on and employ external battery chargers and power supply will not be considered.
3.8 | The internal rechargeable battery must be of a reasonable capacity, such that with battery power the unit must be able to operate continuously for a minimum of four (4)hours at user set infusion rate of 125 ml/h
3.9 | Please state the following:
3.9.1 | Capacity
3.9.2 | Voltage
3.9.3 | Type of battery used
3.9.4 | Length of battery operating time
4 | Modes of operation:
4.1 | The unit offered must be capable of functioning in two different modes.
4.2 | Standard mode:
4.2.1 | Volume or drug delivery in an appropriate per unit volume and per unit weight
4.2.2 | Must be capable of an adjustable bolus administration up to rate of 1200 ml/hour for TIVA
4.2.3 | A drug library must be available that, when selected, allows the name, concentration, and infusion rate of the selected drug to be displayed.
5 | Pharmacokinetic mode:
5.1 | Delivery of a pre-specified drug using standardized pharmacokinetic models.
5.2 | When in the pharmacokinetic mode the unit must clearly display:
5.2.1 | Drug name
5.2.2 | Concentration
5.2.3 | Plasma/effect site concentration
5.2.4 | Plasma/effect site target.
5.2.5 | Initial induction dose
5.2.6 | Initial induction rate
5.2.7 | Initial induction volume
5.2.8 | Time of induction
5.2.9 | Initial maintenance rate
5.2.10 | Patient age
5.2.11 | Patient Length/height
5.2.12 | Patient gender
5.2.13 | Patient weight
5.2.14 | BMI
5.2.15 | Drug model being used
5.2.16 | Decrement time
5.2.17 | Decrement concentration
5.2.18 | Elapsed time
5.2.19 | Volume infused
5.2.20 | Dose infused
5.3 | The unit must be provided as standard with the following pharmacokinetic models:
5.3.1 | Propofol – Marsh – plasma targeting
5.3.2 | Propofol – Schnider – plasma and effect site targeting
5.3.3 | Propofol – Kataria – paediatric model
5.3.4 | Remifentanil – Minto – plasma targeting
5.3.5 | Remifentanil – effect site targeting
5.3.6 | Sufentanil – Gepts – plasma targeting
5.3.7 | Sufentanil – effect site targeting
5.3.8 | The Eleveld universal modal is highly desirable
5.4 | The supplier must provide performance verification data for the unit that demonstrates the validity of their pharmacokinetic models
5.5 | As a safety mechanism - when functioning in the pharmacokinetic mode the user should not be able to exit the mode without first specifically terminating the programme
5.6 | The pharmacokinetic mode should stay engaged even when the pump is switched off and restarted
5.7 | The unit being tendered on must be equipped with computer interface technology. E.g RS232, USB or Wifi. Please state the interface tech
6 | Rate settings
6.1 | The unit offered must also be capable of delivering a high flow rate and the flow rate must be user selectable through the whole range.
6.2 | The unit must be capable of an adjustable bolus administration up to rate of 1200 ml/hour when required during normal infusion delivery
6.3 | The unit must be able to revert to the previously selected normal infusion delivery after administering a bolus.
6.4 | The unit must have a pump run indicator, which will provide the operator with visual indication that the pump is carrying out infusion delivery
6.5 | There must be separate switch controls and indicators for starting and stopping infusion delivery
6.6 | When the unit is turned off, the last settings such as volume and dose delivery must be retained in memory and displayed at switch on.
6.7 | There must be fast purge mode of operation, which will allow priming of IV line. It must not be possible to operate the unit continuously in this mode
6.8 | It must not be possible to change the flow rate on the pump during delivery. The pump must first be stopped, flow rate changed and then restarted.
6.9 | If no flow rate/zero flow rate is selected, it must not be possible to start the unit
6.1 | The unit must provide an audible and visible alarm within a set period should the pump not be restarted. State this time interval.
6.11 | There must be adequate safety precautions against tampering of settings by unauthorized persons. Provide details.
6.12 | The bolus following release of occlusion at 1 ml/h and all higher rates (at alarm), must not exceed 0,6 ml.
7 | Accuracy
7.1 | The accuracy (long term) measured over 60 min at 1ml/h and all higher rates must be better than ± 5% for the unit being offered
7.2 | The accuracy (short term) at 5 ml/h must be better than ± 5% of mean on 2 minutes observation window for the unit being offered.
7.3 | State the actual accuracy of the unit offered.
7.4 | Bidders must submit a written statement / report on the performance of the offered unit in the presence of a working electro-surgery unit
7.5 | Time duration from start-up to attain 95% of “rate set” at 1ml/h and all higher rate < 10 min for the unit being offered.
8 | Display
8.1 | The following must be available and clearly displayed on the front panel under all lighting conditions:
8.2 | A.C. mains power supply operation
8.3 | Battery power supply operation
8.4 | Drug name
8.5 | Plasma/effect site concentration and / or Plasma/effect site target
8.6 | Infusion RATE selected
8.7 | Total volume infused. Any bolus administered by the unit during normal infusion must be added to the total “volume infused” and displayed.
8.8 | Elapsed time
8.9 | Alarm condition and possible alarm / error messages
8.1 | Alarms
8.10.1 | All alarm conditions must be accompanied by an audible and visual warnings
8.10.2 | It must be possible to silence the audible alarm for a maximum of two minutes after which if the fault condition still exists the unit must automatically provide an audible alarm again
8.11 | Trends of drug administration should be recallable
9 | The unit offered must activate alarms for the following minimum conditions
9.1 | LOW battery. State the duration of battery power operation left when this alarm sounds
9.2 | Infusion set not properly loaded
9.3 | Syringe removed
9.4 | Syringe barrel clamp alarm (or equivalent)
9.5 | Syringe plunger disengaged
9.6 | Internal malfunction.
9.7 | Malfunctions detected during self-test at power up
9.8 | Continuous audible and visual alarm for end of travel of syringe plunger.
9.9 | An audible and visible pre-alarm to indicate that the syringe travel plunger is nearing its end and also when the set “Volume to be infused” is approaching final completion
9.1 | Occlusion during infusion administration
9.11 | Maximum occlusion pressure alarm: < 500 mmHg with the use of a 50 ml syringe at ≥1 ml/h The time to alarm following occlusion at 1 ml/h and all higher rates < 36 min for the unit being tendered on. State time to alarm
10 | Conformity Compliance (Please attach certificate)
10.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
10.2 | The unit must comply with electrical safety and performance standard such as IEC 60601-2-24, attached certification
10.2.1 | Please provide unique ref number of the IEC ceriticate:
10.3 | OEM must comply and certified on ISO 9001 quality standards, attach proof of compliance
10.3.1 | Please provide unique ref number of the ISO 9001 ceriticate:
10.4 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
10.4.1 | Please provide unique ref number of the ISO 13485 ceriticate:
10.5 | Model quoted for must be EC certified. Attach a copy of certification
10.5.1 | Please provide unique ref number of the EC ceriticate:
10.6 | Hazardous Substance Act: If this type of equipment / apparatus appears on the schedule of Hazardous substances issued by the Directorate: Health Technology of the Department of Health, a license in terms of the Act on Hazardous Substances (Act 15/1973) must be submitted with the bid / quotation documents. The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
10.7 | All electrical/electronic medical equipment must be licensed by Radiation Control, where applicable a copy of the license must be submitted The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
10.7.1 | Radiation Control License number:
10.8 | The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade (rubber to with 2 screws) type and it must be a minimum length of three (3) meters. NB the mains cable of the unit tendered for must be SABS colour coded.
10.9 | The equipment quoted for must be protected against electromagnetic interference IEC 60601-1-2
11 | COMPREHENSIVE MAINTENANCE
11.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) must be included in the bid price
11.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
11.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
11.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
11.3 | The bidder to provide the OEM recommended service interval of the unit:
11.4 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
12 | General
12.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
12.2 | The unit on offer must be the latest technology.
SYRINGE DRIVER TCI -ACCESSORIES and CONSUMABLES
RT2-01-15a | Extension sets Clear Fluid (Straight) - 100/box
RT2-01-15b | Extension sets (Spiral)- 100/box
RT2-01-15c | Extension sets (Light resistant/sensitive)- 100/box
RT2-01-15d | Extension sets Low Sorb- 100/box
RT2-01-15e | Extension sets Blood- 100/box
RT2-01-15f | Neonatal Closed blood set with Filter - 100/box
RT2-01-15g | Compatible Syringes LUER LOCK Size: 2ml (Pk of 10)
RT2-01-15h | Compatible Syringes LUER LOCK Size: 3ml (Pk of 10)
RT2-01-15i | Compatible Syringes LUER LOCK Size: 5ml (Pk of 10)
RT2-01-15j | Compatible Syringes LUER LOCK Size: 10ml (Pk of 10)
RT2-01-15k | Compatible Syringes LUER LOCK Size: 20ml (Pk of 10)
RT2-01-15l | Compatible Syringes LUER LOCK Size: 30ml (Pk of 10)
RT2-01-15m | Compatible Syringes LUER LOCK Size: 50ml (Pk of 10)
RT2-01-15n | Compatible Syringes LUER LOCK Size: 60ml (Pk of 10)

## Syringe Driver, PCA
RT2-01-16 | SYRINGE DRIVERS PCA
1 | DESCRIPTION
1.1 | A programmable infusion pump capable of administering drugs with patient control
1a | Product Offered:
1a.1 | Make:
1a.2 | Model:
1a.3 | Initial Manufacturing year of Model:
1a.4 | Radiation Control License number:
2 | TECHNICAL SPECIFICATION
2.1 | An infusion pump (including patient hand set and locking facility) that are re-usable and programmable to provide patient controlled intravenous and epidural analgesia as well as regional analgesia infusions.
2.2 | Must be usable for intravenous, epidural, regional analgesia and wound infiltration infusions
2.3 | It should pe progammable to provide continuous infusion only, bolus only, continuous pus bolus and auto bolus infusion
2.4 | Weight based dosing software with ceiling ptotection must be available.
2.5 | A searchable drug library should be avaialble
2.6 | it should be possible to store programmed information
2.7 | At least 2 levels of security acces to prevent unauthorised programme changes.
2.8 | Flow rate range:
2.8.1 | Basal rate 0 to 100 ml/hr in 0.1 ml increments
2.8.2 | KVO Rate 0 (OFF) - 5 ml/hr
2.8.3 | Flow rate accuracy - ± 6 %
2.8.4 | Volume to be infused: 0.1 – 1000 ml
2.9 | Infusion safety limitations
2.9.1 | ultrasonic Air Sensor Air bubble size 0.04 – 1 ml
2.9.2 | Occlusion Pressure adjustable 250-1500 mmHg (adjustable depending on infusion mode)
2.10 | The unit should be light weight and small for easy mobility of patients. State weight
2.11 | It should have a carry handle
2.12 | The unit offered must be well constructed with robust impact resistant casing and panels. The control panel, all push buttons and switches must be of a flush membrane type or an equivalent, to provide resistance for the entry of fluids in to the internal workings and circuits of the unit.
2.13 | The unit must be user friendly for patients and doctors
2.14 | Disposables/accesories must be clearly described and compatibility (or not) with generic disposables should be confirmed with documentary evidence
2.15 | If syringes is part of the equipment, the size should be automatically sensed and displayed
2.16 | Non-freeflow or backflow disposable administration sets should be offered (colour coded for different applications eg IV and epidural)
2.18 | The unit, including reservoir of administered drug should be lockable to prevent tampering. State if additional accessories are necessary for this.
1.18 | An electronic safety programme is desireable
3 | Display
3.1 | The following must be available and clearly displayed on the front panel under all lighting conditions, preferably full colour:
3.2 | A.C. mains power supply operation
3.3 | Battery power supply operation
3.4 | Drug name & concentration
3.5 | Infusion rate selected
3.6 | Bolus volume
3.7 | Lock out time
3.8 | A history must be logged for at least 72 hours
3.9 | The history of bolus administration and total dose must be able to be called to view on screen
3.10 | I should be possible to export the history to computer
3.11 | Alarm condition and possible alarm / error messages
4 | Alarms
4.1 | All alarm conditions must be accompanied by an audible and visual warnings
4.2 | It must be possible to silence the audible alarm for a maximum of two minutes after which if the fault condition still exists the unit must automatically provide an audible alarm again
4.3 | Minimum alarms to be diplayed
4.3.1 | Air in line
4.3.2 | Up occlusion, Down occlusion
4.3.3 | Pump unattended
4.3.4 | End of infusion
4.3.5 | Near end of Infusion
4.3.6 | Low Battery, End Battery
4.3.7 | Pump Unattended
4.3.8 | Door open
4.3.9 | Missing set key
4.3.10 | System error
5 | Power and Batteries
5.1 | The monitor must operate off both the 220V-240V , 50Hz single phase AC power supply and an internally fitted rechargeable battery. The pump must be fused in both the LIVE and NEUTRAL.
5.2 | The unit offered must be supplied with an internal rechargeable battery.
5.3 | The internal rechargeable battery must be automatically charged when the unit is connected to a live A.C. supply
5.4 | In the event of A.C. mains failure the battery must automatically provide continuous operation
5.5 | The unit must be provided with visible indication to reflect whether the unit is being operated off the AC power or the internal rechargeable battery power
5.6 | The unit offered must be provided with circuitry, which must ensure that the internal rechargeable battery will be protected against over-charging and over-discharging
5.7 | Unit that are quoted on and employ external battery chargers and power supply will not be considered.
5.8 | The internal rechargeable battery must be of a reasonable capacity, such that with battery power the unit must be able to operate continuously for a minimum of twelve (12) hours at al settings
5.9 | Please state the following:
5.9.1 | Capacity
5.9.2 | Voltage
5.9.3 | Type of battery used
5.9.4 | Length of battery operating time
6 | Operating conditions
6.1 | Must function between 18 and 40⁰C
6.2 | Must function between 15 and 90 % humidity
6.3 | Must function at all SA altitudes
7 | Conformity Compliance (Please attach certificate)
7.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
7.2 | The unit must comply with electrical safety and performance standard such as IEC 60601-2-24, attached certification
7.2.1 | Please provide unique ref number of the IEC ceriticate:
7.3 | OEM must comply and certified on ISO 9001 quality standards, attach proof of compliance
7.3.1 | Please provide unique ref number of the ISO 9001 ceriticate:
7.4 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
7.4.1 | Please provide unique ref number of the ISO 13485 ceriticate:
7.5 | Model quoted for must be EC certified. Attach a copy of certification
7.5.1 | Please provide unique ref number of the EC ceriticate:
7.6 | Hazardous Substance Act: If this type of equipment / apparatus appears on the schedule of Hazardous substances issued by the Directorate: Health Technology of the Department of Health, a license in terms of the Act on Hazardous Substances (Act 15/1973) must be submitted with the bid / quotation documents. The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
7.7 | All electrical/electronic medical equipment must be licensed by Radiation Control, where applicable a copy of the license must be submitted The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
7.7.1 | Radiation Control License number:
7.8 | The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade (rubber to with 2 screws) type and it must be a minimum length of three (3) meters. NB the mains cable of the unit tendered for must be SABS colour coded.
7.9 | The equipment quoted for must be protected against electromagnetic interference IEC 60601-1-2
8 | COMPREHENSIVE MAINTENANCE
8.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) must be included in the bid price
8.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
8.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
8.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
8.3 | The bidder to provide the OEM recommended service interval of the unit:
8.4 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
9 | General
9.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
9.2 | The unit on offer must be the latest technology.
SYRINGE DRIVER PCA -ACCESSORIES and CONSUMABLES
RT2-01-16a | Extension sets Clear Fluid (Straight) - 100/box
RT2-01-16b | Extension sets (Spiral)- 100/box
RT2-01-16c | Extension sets (Light resistant/sensitive)- 100/box
RT2-01-16d | Extension sets Low Sorb- 100/box
RT2-01-16e | Extension sets Blood- 100/box
RT2-01-16f | Neonatal Closed blood set with Filter - 100/box
RT2-01-16g | Compatible Syringes LUER LOCK Size: 2ml (Pk of 10)
RT2-01-16h | Compatible Syringes LUER LOCK Size: 3ml (Pk of 10)
RT2-01-16i | Compatible Syringes LUER LOCK Size: 5ml (Pk of 10)
RT2-01-16j | Compatible Syringes LUER LOCK Size: 10ml (Pk of 10)
RT2-01-16k | Compatible Syringes LUER LOCK Size: 20ml (Pk of 10)
RT2-01-16l | Compatible Syringes LUER LOCK Size: 30ml (Pk of 10)
RT2-01-16m | Compatible Syringes LUER LOCK Size: 50ml (Pk of 10)
RT2-01-16n | Compatible Syringes LUER LOCK Size: 60ml (Pk of 10)

## Blood-Fluid Warmer
RT2-01-17 | BLOOD/FLUID WARMER
1 | DESCRIPTION
1.1 | A device that raises refrigerated blood or intravenous fluids to a desired temperature.
1a | Product Offered:
1a.1 | Make:
1a.2 | Model:
1a.3 | Initial Manufacturing year of Model:
1a.4 | Radiation Control License number:
2 | TECHNICAL SPECIFICATION
2.1 | A blood solution warmer must be lightweight, robustly constructed and its design must be such that it must be possible for it to be attached to a normal infusion stand and a rail or equivalent.
2.2 | The unit on offer must be fitted and supplied with a universal attachment for enabling attachment of the unit to both an infusion stand and a Gabler rail, and the cost of which must be included in the final bid price.
2.3 | The unit on offer must use universal or dedicated disposables.
2.4 | The unit on offer should be compatible with standard blood administration IV lines, a starter pack must be delivered for the full functionality of the unit.
2.5 | If the unit on offer uses any other dedicated disposables or consumable items, the bidder must list them with the prices including V.A.T on the schedule of optional accessories.
2.6 | The heat exchanger control circuit of the unit offered must provide well-controlled uniform heat distribution, preventing any hot spots developing that may damage blood.
2.7 | The design of the unit on offer must be such that there will be minimal cooling down of blood and fluid that is being administered to the patient after warming.
2.8 | The unit must have reliable safety protection against over temperature. Both an audible and visual alarm must be provided on the quoted unit should the temperature rise to a preset temperature, which is over and above the maximum user selectable temperature. The over temperature protection must also additionally turn off the power to the heater should the temperature exceed this preset temperature. Bidder must state the preset temperature at which this safety protection against over temperature will turn of the power to the heater.
Unit must warm 150ml/min of blood from 4 degrees Celcius to 38 degrees Celcius
2.9 | The unit on offer must operate off the 220V to 240V 50Hz single phase a.c. supply and be fused in both the LIVE and NEUTRAL.
2.1 | The unit offered must be provided with a suitable display of the temperature control set point and must be clear viewing under all lighting conditions.
2.11 | State what measures are used to prevent or eliminate air bubbles from the blood and fluid circuit.
4 | STANDARD ACCESSORIES
4.1 | Mounting bracket
4.2 | A starter pack (pack of 10) of all disposables/accessories (controller, cartridge, cassette, extensions) must be supplied with the unit
4.3 | Universal or dedicated sets (Pack of 10)
5 | Conformity Compliance (Please attach certificate)
5.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
5.2 | The unit must comply with electrical safety and performance standard such as IEC 60601-2-35, attached certification
5.2.1 | Please provide unique ref number of the IEC ceriticate:
5.3 | OEM must comply and certified on ISO 9001 quality standards, attach proof of compliance
5.3.1 | Please provide unique ref number of the ISO 9001 ceriticate:
5.4 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
5.4.1 | Please provide unique ref number of the ISO 13485 ceriticate:
5.5 | Model quoted for must be EC certified. Attach a copy of certification
5.5.1 | Please provide unique ref number of the EC ceriticate:
5.6 | Hazardous Substance Act: If this type of equipment / apparatus appears on the schedule of Hazardous substances issued by the Directorate: Health Technology of the Department of Health, a license in terms of the Act on Hazardous Substances (Act 15/1973) must be submitted with the bid / quotation documents. The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.7 | All electrical/electronic medical equipment must be licensed by Radiation Control, where applicable a copy of the license must be submitted The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.7.1 | Radiation Control License number:
5.8 | The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade (rubber to with 2 screws) type and it must be a minimum length of three (3) meters. NB the mains cable of the unit tendered for must be SABS colour coded.
5.9 | The equipment quoted for must be protected against electromagnetic interference IEC 60601-1-2
6 | COMPREHENSIVE MAINTENANCE
6.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) must be included in the bid price
6.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
6.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
6.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
6.3 | The bidder to provide the OEM recommended service interval of the unit:
6.4 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
7 | General
7.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
7.2 | The unit on offer must be the latest technology.
BLOOD/FLUID WARMER ACCESSORIES & CONSUMABLES
RT2-01-17a | Cartridge/Cassette (Pk of 10)
RT2-01-17b | Dedicated sets/tubing (100/box)
Bidder to list any other Consumables or Accessories (Not listed above) for the model on offer below

## Patient Warmer, Forced-Air
RT2-01-18 | PATIENT WARMER (FORCED AIR)
1 | DESCRIPTION
1.1 | A temperature management system using forced-air, utilized within a hospital to maintain a patient's core body temperature
1a | Product Offered:
1a.1 | Make:
1a.2 | Model:
1a.3 | Initial Manufacturing year of Model:
1a.4 | Radiation Control License number:
2 | TECHNICAL SPECIFICATION
2.1 | The unit must be able to operate off the 220V – 240V, 50Hz single phase A.C. power supply.
3 | The physical properties of the unit should fulfill the following:
3.1 | Weight should not exceed 8kg
3.2 | The unit must come with sturdy mounting capability
3.3 | It should be no larger than 38cm (height) X 41cm (width) X 33cm (depth)
3.4 | The operating noise level should be less than 60db
3.5 | The Control Panel must have LED or equivalent display and all switches must preferably be touch activated.
4 | The Control panel must offer a choice of at least 3 selectable temperature settings as well as ambient temperature setting. The temperature settings must be typically as follows:
4.1 | Low: set value between 30-34 degrees Celsius with an accuracy of ± 1.5 degrees Celsius
4.2 | Medium: set value between 35-39 degrees Celsius with an accuracy of ± 1.5 degrees Celsius
4.3 | High: set value between 39-43 degrees Celsius with an accuracy of ± 1.5 degrees Celsius
4.4 | The unit must be able to circulate room ambient temperature air
4.5 | The unit must indicate when the selected set temperature has been reached and thereafter maintain the selected set temperature
4.6 | The unit must reach the selected temperature within 90 seconds of the selection of that temperature.
4.7 | The unit must incorporate visual and audible temperature discrepancy alarms. These should be activated when the temperature measured at the hose-end exceeds or falls sort of the set temperature by more than 1 degree Celsius. Bidder to demonstrate these alarms during evaluation
4.8 | An over-temperature state must interrupt the heating function of the unit. The over-temperature alarm must activate when the measure hose-end temperature exceeds the set temperature by more than 3 degrees Celsius. Bidder to demostrate these alarms during evaluation
4.9 | The air supply inlet must incorporate a 0,3 micron (or smaller) HEPA filter.
4.10 | The unit must generate airflow of at least 19 litres per second.
5 | THE OUTLET HOSE : The outlet hose must be of the re-usable type and it must be easily removable for cleaning or replacement.
5.1 | The hose must have the following features:
5.2 | It must be able to swivel freely
5.3 | It must have securing devices for to prevent unintended disconnection
5.4 | Its minimum length must be 1.5 metre and suitable for both paediatrics and adults.
6 | THE BLANKETS
6.1 | These must be disposable, lightweight, soft, radiolucent and latex-free.
6.2 | The unit must utilize blankets that allow even heat distribution to the patient without the creation of hotspots.
7 | STANDARD ACCESSORIES
7.1 | A starter pack of the following blankets to be provided: 2 x Boxes (10 ea) of full body . (i.e 1 box paeds and 1box adults).
7.2 | Mobile stand with at least 4 castors and 2 fitted with brakes. (Complete for mounting the unit securely)
8 | Conformity Compliance (Please attach certificate)
8.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
8.2 | The unit must comply with electrical safety and performance standard such as IEC 60601-2-35, attached certification
8.2.1 | Please provide unique ref number of the IEC ceriticate:
8.3 | OEM must comply and certified on ISO 9001 quality standards, attach proof of compliance
8.3.1 | Please provide unique ref number of the ISO 9001 ceriticate:
8.4 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
8.4.1 | Please provide unique ref number of the ISO 13485 ceriticate:
8.5 | Model quoted for must be EC certified. Attach a copy of certification
8.5.1 | Please provide unique ref number of the EC ceriticate:
8.6 | Hazardous Substance Act: If this type of equipment / apparatus appears on the schedule of Hazardous substances issued by the Directorate: Health Technology of the Department of Health, a license in terms of the Act on Hazardous Substances (Act 15/1973) must be submitted with the bid / quotation documents. The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
8.7 | All electrical/electronic medical equipment must be licensed by Radiation Control, where applicable a copy of the license must be submitted The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
8.7.1 | Radiation Control License number:
8.8 | The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade (rubber to with 2 screws) type and it must be a minimum length of three (3) meters. NB the mains cable of the unit tendered for must be SABS colour coded.
8.9 | The equipment quoted for must be protected against electromagnetic interference IEC 60601-1-2
9 | COMPREHENSIVE MAINTENANCE
9.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) must be included in the bid price
9.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
9.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
9.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
9.3 | The bidder to provide the OEM recommended service interval of the unit:
9.4 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
10 | General
10.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
10.2 | The unit on offer must be the latest technology.
PATIENT WARMER (FORCED AIR), REPLACEMENT ACCESSORIES & CONSUMABLES
RT2-01-18a | Upper body blanket
RT2-01-18b | Lower body blanket
RT2-01-18c | Surgical Access blanket
RT2-01-18d | Underbody Peadiatric blanket
RT2-01-18e | Underbody Adult blanket
RT2-01-18f | Full body Peaditric blanket
RT2-01-18g | Full body Adult blanket
RT2-01-18h | Air Hose
RT2-01-18i | Mobile stand
Bidder to list any other Consumables or Accessories (Not listed above) for the model on offer below

## Nerve Stimulator
RT2-01-20 | NERVE STIMULATOR
1 | DESCRIPTION
1.1 | A device which utilises electrical current to stimulate the nerves for therapeutic/diagnostic purposes
1a | Product Offered:
1a.1 | Make:
1a.2 | Model:
1a.3 | Initial Manufacturing year of Model:
1a.4 | Radiation Control License number:
2 | TECHNICAL SPECIFICATION
2.1 | This specification establishes the requirements, supply, delivery, end user training, demonstration, commission and installation of a HAND HELD NERVE STIMULATOR, that is:
2.2 | Compact and light
2.3 | Robust
2.4 | User friendly
2.5 | Microprocessor controlled with the latest up to date technology.
2.6 | The unit offered must offer a full range of BOTH conventional and recently developed nerve stimulation patterns for monitoring of neuromuscular blockade during surgery or in the intensive care setting.
2.7 | The unit must have a constant current variable voltage feature, which will allow delivery of comparable stimuli throughout the whole operation, despite changes in impedance.
2.8 | The pulse stimulus must be of a monophasic, square type to ensure a constant current is maintained throughout the stimulus.
2.9 | The unit must offer the user all the conventional patterns of stimulation which must be easy to identify and select on the unit.
2.1 | The patterns of stimulation must include:
2.11
2.13 | Train of four: Four successive 2Hz pulses per 2 seconds (one pulse every 0.5 seconds) with a 10 second interval between trains,on demand or twelve-second repeat
2.15 | Tetanic stimulation: on demand at 50Hz or 100Hz for up to a minimum of five seconds
2.16 | Post -Tetanic count.
2.17 | Double burst stimulation: The burst repeats must be separated by 10 seconds. Selectable 3,3 or 3,2 patterns
2.19 | The unit must be provided with a digital alphanumeric display, which must provide clear viewing of all parameters under all lighting conditions.
2.2 | The unit offered must provide display of the current output
2.21 | The unit must be able to deliver a constant current output which must be selectable in the range of 0-80mA
2.22 | The output pulse must be:
2.23 | A monophasic square wave form
2.24 | Longer than 0.1milliseconds and less than 0.5milliseconds
2.25 | Longer duration pulses and other waveforms if available must be clearly indicated when selected.
2.26 | The bidder must clearly state the output pulse forms available on the unit offered.
3 | BATTERIES
3.1 | The unit must be battery operated.
3.2 | The battery cover must be easy to open and must require no tools to open and close the cover.
3.3 | It is important that the unit offered is provided with audible and visual indication of pulse delivery
3.4 | The unit must be provided with a Low Battery Indicator.
3.5 | The unit must have an Automatic Power Shutdown feature in order to conserve battery power.
3.6 | The units should have a programming mode whereby users may define and store any combination of pulse patterns and time intervals.
4 | STANDARD ACCESSORIES
4.1 | Electrodes set with crocodile clips or equivalent
4.2 | Battery
5 | Conformity Compliance (Please attach certificate)
5.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
5.2 | The unit must comply with electrical safety and performance standard such as IEC 60601-2-10, attached certification
5.2.1 | Please provide unique ref number of the IEC ceriticate:
5.3 | OEM must comply and certified on ISO 9001 quality standards, attach proof of compliance
5.3.1 | Please provide unique ref number of the ISO 9001 ceriticate:
5.4 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
5.4.1 | Please provide unique ref number of the ISO 13485 ceriticate:
5.5 | Model quoted for must be EC certified. Attach a copy of certification
5.5.1 | Please provide unique ref number of the EC ceriticate:
5.6 | Hazardous Substance Act: If this type of equipment / apparatus appears on the schedule of Hazardous substances issued by the Directorate: Health Technology of the Department of Health, a license in terms of the Act on Hazardous Substances (Act 15/1973) must be submitted with the bid / quotation documents. The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.7 | All electrical/electronic medical equipment must be licensed by Radiation Control, where applicable a copy of the license must be submitted The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.7.1 | Radiation Control License number:
6 | COMPREHENSIVE WARRANTY COVER
6.1 | A 2-YEAR warranty inclusive maintenance and service (Year 1 to Year 2) must be included in the bid price
6.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 5-year warranty will commence after formal Commissioning and handover of the equipment.
7 | General
7.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
7.2 | The unit on offer must be the latest technology.
NERVE STIMULATOR ACCESSORIES
RT2-01-20a | Electrodes set with crocodile clips or equivalent
RT2-01-20b | Battery
Bidder to list any other Consumables or Accessories (Not listed above) for the model on offer below

## Ultrasound, Anaesthetsia, PoC
RT2-01-23 | Ultrasound Machine, Anaesthesia, PoC
1 | DESCRIPTION
1.1 | A machine for superficial vascular and nerve imaging, as well as transthoracic and transabdominal application including appropriate ultrasound probes
2 | Product Offered:
2.1 | Make:
2.2 | Model:
2.3 | Initial Manufacturing year of Model:
2.4 | Radiation Control License number:
3 | TECHNICAL SPECIFICATION
3.1 | Startup time to be specifide. Faster times <60s form cold start will have prefference. Bidder to state
3.2 | Portable touch screen ultrasound machine, on a mounted cart
3.3 | Thermal Index display(Ultrasound immision)
3.4 | Touchscreen to function with user wearing gloves
3.5 | Fully Integrated,interactive, touch screen display/monitor
3.6 | Digital broadband beam-former
3.7 | Digital channels more than 1000 - 60000.
3.8 | 2D Imaging
3.9 | Tissue harmonics should be standard
3.10 | Dynamic range 160-250 Db
3.11 | Real-Time Compounding
3.12 | Adaptive Image Processing for noise and artifact reduction
3.13 | Ergonomic transducers designs with lightweight flexible cables
3.14 | High-resolution interactive graphical color screen with adjustment for various ambient light conditions
3.15 | Fully articulating flicker-free 15" LCD,integrated,interactive touch screen monitor
3.16 | At least 1 active transducer connectors and one non Imaging transducer port, maximum 3 active transducer ports
3.17 | Intelligent one-button optimization in 2D and Doppler modes
3.18 | Integrated Battery to last at least 1 hour when fully charged.
3.19 | CLINICAL APPLICATION IN SOFTWARES
3.19.1 | The Ultrasound machine must have the following software applications:
3.19.2 | Anaesthesia,Nerve location and blocking, Critical care, Emergency, vascular,Small parts, Abdomen, Basic Cardiology
3.19.3 | Needle Optimisation / visualisation Software
3.19.4 | Vascular software
3.20 | IMAGING MODES AND IMAGE DISPLAY
3.20.1 | The operating modes should include: 2D, M, 2D/M, PWD, Color Doppler, Power Doppler, Duplex & Triplex
3.20.2 | Cardiac measurements included standard on the sysytem should be stated
3.20.3 | M mode, Color, Pulsed wave, Power Doppler
3.20.4 | Annotation presets available
3.20.5 | Dual Live display with BW reference image and color Doppler image
3.20.6 | Trapezoid imaging on linear probes.
3.20.7 | High Definition Write Zoom and Read Zoom with pan feature
3.20.8 | Scanning depth minimum ranges of 30 cm
3.20.9 | Measurement tools including: distance, depth, area, and circumference
3.20.10 | Volume Flow Measurements
3.20.11 | On-board workstation-class data management with thumbnail previews and storage of images, loops, and reports
3.20.12 | Retrospective and prospective clip capture to internal drive or removable media
3.20.13 | USB and Ethernet connectivity ports must be available
3.20.14 | Internal HDD capacity range of 120GB - 500GB
3.20.15 | PERIPHERALS Standard
3.21 | Wheelbased cart:
3.21.1 | The cart must have 4 castors, two of which must have brakes
3.21.2 | Height Adjustable cart
3.21.3 | The cart or machine must have minimum two transducer holders
3.21.4 | Must be able to add Black and White printer on request with integrated AC onto the Ultrasound unit
3.21.5 | The cart including 1 shelf
4 | STANDARD ACCESSORIES
4.1 | High Frequency Linear Transducer 3-12MHz
4.2 | Adult Phased Array Transducer 2-4MHz
5 | Conformity Compliance (Please attach certificate)
5.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
5.2 | The unit must comply with electrical safety and performance standard such as IEC 60601-2-37, attached certification
5.2.1 | Please provide unique ref number of the IEC ceriticate:
5.3 | OEM must comply and certified on ISO 9001 quality standards, attach proof of compliance
5.3.1 | Please provide unique ref number of the ISO 9001 ceriticate:
5.4 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
5.4.1 | Please provide unique ref number of the ISO 13485 ceriticate:
5.5 | Model quoted for must be CE certified. Attach a copy of certification
5.5.1 | Please provide unique ref number of the CE ceriticate:
5.6 | Hazardous Substance Act: If this type of equipment / apparatus appears on the schedule of Hazardous substances issued by the Directorate: Health Technology of the Department of Health, a license in terms of the Act on Hazardous Substances (Act 15/1973) must be submitted with the bid / quotation documents. The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.7 | All electrical/electronic medical equipment must be licensed by Radiation Control, where applicable a copy of the license must be submitted The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.7.1 | Radiation Control License number:
5.8 | The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade (rubber to with 2 screws) type and it must be a minimum length of three (3) meters. NB the mains cable of the unit tendered for must be SABS colour coded.
5.9 | The equipment quoted for must be protected against electromagnetic interference IEC 60601-1-2
6 | COMPREHENSIVE MAINTENANCE
6.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) must be included in the bid price
6.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
6.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
6.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
6.3 | The bidder to provide the OEM recommended service interval of the unit:
6.4 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
7 | General
7.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
7.2 | The unit on offer must be the latest technology.
Ultrasound Machine, Anaesthesia, Portable
RT2-01-23a | High Frequency Linear Transducer 3-12MHz
RT2-01-23b | High Frequency Linear Transducer 5-18MHz
RT2-01-23c | Adult Phased Array Transducer 2-4MHz
RT2-01-23d | Convex transducer 3-5MHz
RT2-01-23e | Paediatric Phased array Transducer 3-7MHz
RT2-01-23f | Linear Transducer 4-10MHz
RT2-01-23g | Adult TEE tranducer (In storage case)
RT2-01-23h | Paeds TEE transducer (In storage case)
RT2-01-23i | TEE software (UPGRADE
Bidder to list any other Consumables or Accessories (Not listed above) for the model on offer below

## Category-2 Vent & Respiratory
VENTILATION & RESPIRATORY EQUIPMENT
# | Item number | Item Description
1 | RT2-02-01 | Ventilator, Neonatal-Paeds
2 | RT2-02-02 | Ventlator, Paeds-Adult
5 | RT2-02-05 | Ventilator, CPAP
6 | RT2-02-06 | Ventilator, Transport, Paeds-Adult
7 | RT2-02-07 | Ventilator, Transport, Neonatal-Paeds
8 | RT2-02-08 | High Frequency Oscilator
9 | RT2-02-09 | Resuscitator, Gas-powered, Infant
10 | RT2-02-10 | Humidifier, Servo control
11 | RT2-02-11 | Air-O2 Blender, Low Flow
12 | RT2-02-12 | Air-O2 Blender, High Flow
13 | RT2-02-13 | Pulmonary Function Unit
14 | RT2-02-14 | Oxygen Flowmeter, Single
15 | RT2-02-15 | Oxygen Flowmeter, Double
16 | RT2-02-16 | Nebulizer, Ultrasonic

## Ventilator, Neonatal-Paeds
RT2-02-01 | Ventilator: Neonatal/Paediatric
1 | Description
1.1 | Ventilator with the latest technology to ventilate neonates and paediatric patients (500g to 20kg) in the Neonatal (premature and full term newborns) and Paediatric Intensive Care Units
1a | Product Offered:
1a.1 | Make:
1a.2 | Model:
1a.3 | Initial Manufacturing year of Model:
1a.4 | Radiation Control License number:
2 | Technical Specification
2.1 | The ventilator must be pneumatically powered
2.2 | The ventilator must be microprocessor controlled
2.3 | The ventilator must be a Possitive Pressure ventilator
2.4 | Multi-mode state of the art touch screen or equivalent graphical interface
2.5 | Ability to display Airway graphics and Respiratory/Lung Function parameters. (This can be done by paging through the screen)
2.6 | Breath triggering: Pressure or Flow
2.7 | Breath termination: Time or Flow
2.8 | Adjustable I:E Ratio Normal 1:2 (Range 1:1-1:4) Inspiratory Time: 0.1 to 3 seconds (Range 0.11-10s)
2.9 | Gas flow: Decelerating and Constant (Range 10 - 200 L/min)
2.10 | The ventilator must provide the following ventilation modes (Invasively):
2.10.1 | Pressure Limited Mode (Pressure Control) + SIMV
2.10.2 | Volume Limited Mode ( Volume Control) + SIMV
2.10.3 | PEEP
2.10.4 | CPAP (0-40)
2.10.5 | Pressure Support Ventilation (PSV) + PEEP
2.10.6 | APRV
2.10.7 | PRVC/Volume Target Ventilation (Optional)
2.11 | The ventilator must provide the following ventilation modes (Non-Invasively):
2.11.1 | Nasal CPAP
2.11.2 | Mask CPAP
2.11.3 | Bi-PAP
2.11.4 | Hi-Flow Oxygen
2.12 | Integrated Nebulizer (Optional)- Ultrasonic
2.13 | Inspiratory Hold capability up to 30 sec
2.14 | Expiratory Hold capability up to 30 sec
2.15 | Back up Mode if Apnea occurs in all modes
2.16 | Back up/ Apnea ventilation in all modes
2.17 | Pre-Oxygenation capability with 100%
2.18 | Respiratory rate/min range 1-160/min
2.19 | Peak Inspired Pressure: 0 to at least 60cm H2O
2.20 | Inspratory Flow: 0,4 to 100 L/min
2.21 | Pause adjustable from 0 to 2 sec
2.22 | Flow Trigger: 0,1 to 5 L/min
2.23 | Proximal trigger option for neonate/infant.
2.24 | Pressure Trigger: 0.5 to 10 mbar below PEEP/CPAP.
2.25 | Pressure High: 0 to 70 mbar.
2.26 | Pressure Low: 0 to 45 mbar.
2.27 | Oxygen Concentration from 21% to 100%.
2.28 | Bias flow.
2.29 | Leak Compensation: Up to 8 L/min.
2.30 | Alarms:
2.30.1 | Both audible and visual alarms.
2.30.2 | Patient disconnect
2.30.3 | Oxygen concentration variation
2.30.4 | Low inspiratory pressure.
2.30.5 | High breath rates
2.30.6 | Incompatible settings.
2.30.7 | Low PEEP.
2.30.8 | Over and under patient pressure
2.30.9 | Over and under patient volume
2.30.10 | High continuous pressure
2.30.11 | Inverse I:E ratio
2.30.12 | Air/Oxygen % deviation
2.30.13 | Power failure
2.30.14 | Air and Oxygen supply failure
2.30.15 | Air and Oxygen low pressure.
2.30.16 | APNEA.
2.30.17 | Alarm mute. (Resets after 60 seconds.)
2.30.18 | Low battery.
2.31 | The unit must have automatic switch over to a built in battery with a capacity to ensure full functionality of the unit for not less than 2 hours.
2.32 | The unit must have automatic switch over that will switch between the piped air and oxygen supply should one of the supplies fail.
2.33 | The unit must have a circuit compliance compensation 0.0 to 7.5 ml/cmH2O and must be automatically measured.
2.34 | Proximal flow and pressure sensing technology for neonates.
2.36 | Graphics Display: The following parameters and measurements must be displayed:
2.36.1 | Inspired tidal volume in ml
2.36.2 | Expired tidal volume in ml
2.36.3 | Peak inspiratory pressure.
2.36.4 | Proximal airway pressure.
2.36.5 | Inspired tidal volume per kg in ml/kg
2.36.6 | Expired tidal volume per kg in ml/kg
2.36.7 | Spontaneous tidal volume in ml
2.36.8 | Total breath rate in bpm
2.36.9 | Spontaneous breath rate in bpm
2.36.10 | Minute volume in litres
2.36.11 | Spontaneous minute volume in litres
2.36.12 | Peak airway pressure in cmH2O
2.36.13 | Mean airways pressure in cmH2O
2.36.14 | Plateau pressure in cmH2O
2.36.15 | I:E ratio
2.36.16 | Set I:E ratio
2.36.17 | Inspired time in seconds
2.36.18 | Expired time in seconds
2.36.19 | Static airways compliance in ml/cmH2O
2.36.20 | Static airways compliance per kg in ml/cmH2O
2.36.21 | Dynamic airways compliance in ml/cmH2O
2.36.22 | Dynamic airways compliance per kg in ml/cmH2O
2.36.23 | Peak inspired flow rate in l/m
2.36.24 | Peak expired flow rate in l/m
2.36.25 | Inspired O2 in %
2.36.26 | Tube leak in %
2.37 | At least 5 selectable waveforms must be displayed at any time on the graphic display screen.
2.38 | The graphics display system must have freeze capabilities.
2.39 | Trending of at least 24 hours on at least 8 parameters of pressure-, flow- and tidal volumes.
2.40 | The unit must be able to calculate and display pressure- and flow/volume loops.
2.41 | The ventilator must be mounted on a stable trolley with at least two of the castors fitted with brakes that conform to SANS 601.
2.42 | The unit must have circuit holding arm
2.43 | The ventilator trolley must be suitable for mounting of the dual servo humidifier
2.44 | The unit must be supplied a pole or universal clamp unit to mount the exiting humidifiers in use onto the unit.
2.45 | Mains cable at least three metres long with SABS approved 15A three prong plug.
2.46 | The ventilator must come complete with colour-coded air and O2 hoses 3-metre length. (O2 and air Heyer type keys, SABS 1409, must also be supplied.)
2.47 | The unit must have a rail and circuit hanger of a suitable strength on both sides of the ventilator trolley.
2.48 | The unit must have an integrated blender with O2 concentration 21 to 100 % ± 3 %.
2.49 | The unit must be handed over in full operating order
3 | A starter-pack of consumables must be supplied with the unit comprising at least the following:
3.1 | Bacteria filter (pk of 10)
3.2 | Neonatal patient circuit (Pk of 10)
3.3 | Paeds patient circuits (Pk of 10)
3.4 | Test lung for neonates
3.5 | Humidifier (complete with its probes)
3.6 | Circuit support arm
4 | Conformity Compliance (Please attach certificate)
4.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
4.2 | The unit must comply with electrical safety and performance standard such as IEC 60601-2-12, attached certification
4.2.1 | Please provide unique ref number of the IEC ceriticate:
4.3 | OEM must comply and certified on ISO 9001 quality standards, attach proof of compliance
4.3.1 | Please provide unique ref number of the ISO 9001 ceriticate:
4.4 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
4.4.1 | Please provide unique ref number of the ISO 13485 ceriticate:
4.5 | Model quoted for must be EC certified. Attach a copy of certification
4.5.1 | Please provide unique ref number of the EC ceriticate:
4.6 | Hazardous Substance Act: If this type of equipment / apparatus appears on the schedule of Hazardous substances issued by the Directorate: Health Technology of the Department of Health, a license in terms of the Act on Hazardous Substances (Act 15/1973) must be submitted with the bid / quotation documents. The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
4.7 | All electrical/electronic medical equipment must be licensed by Radiation Control, where applicable a copy of the license must be submitted The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
4.7.1 | Radiation Control License number:
4.8 | The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade (rubber to with 2 screws) type and it must be a minimum length of three (3) meters. NB the mains cable of the unit tendered for must be SABS colour coded.
4.9 | The equipment quoted for must be protected against electromagnetic interference IEC 60601-1-2
5 | COMPREHENSIVE MAINTENANCE
5.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) must be included in the bid price
5.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries, Flow sensors and Oxygen sensor), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
5.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
5.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries, Flow sensors and Oxygen sensor), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
5.3 | The bidder to provide the OEM recommended service interval of the unit:
5.4 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
6 | General
6.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
6.2 | The unit on offer must be the latest technology.
Neonatal/Paediatric Ventilator Accessories, Consumables & Additional options
RT2-02-01a | Addition of HFOV Mode (UPGRADE)
RT2-02-01b | Bacteria filter (pk of 10)
RT2-02-01c | Neonatal patient circuit (Pk of 10)
RT2-02-01d | Paeds patient circuits (Pk of 10)
RT2-02-01e | Test lung for neonates
RT2-02-01f | Flow sensors
RT2-02-01g | Oxygen cell
RT2-02-01h | Medical Air compressor (Mountable to ventilator trolley)
Bidder to list any other Consumables or Accessories (Not listed above) for the model on offer below

## Ventilator, Paeds-Adult
RT2-02-02 | VENTILATOR PAEDIATRIC/ADULT
1 | Description
2 | Product Offered:
2.1 | Make:
2.2 | Model:
2.3 | Initial Manufacturing year of Model:
2.4 | Radiation Control License number:
3 | Technical Specification
3.1 | General details
3.1.1 | The unit must be able to ventilate and monitor paediatric and adult patients
3.1.2 | The ventilator must be a time-cycled, volume- and pressure-controlled unit.
3.1.3 | The ventilator must have an integrated, electronically controlled air compressor or turbine
3.1.4 | The ventilator must have an oxygen port for use with an oxygen cylinder during transport
3.1.5 | The unit must have automatic switch-over unit that will switch between the piped air and turbine/oxygen supply should one of the supplies fail.
3.1.6 | Self diagnosis with each start-up and regular integrity testing of all system parameters
3.1.7 | Ventilator must have "Standby" mode which suspends alarms, breath cycling and trends.
3.1.8 | "Standby" mode must retain most current ventilator settings and alarms, resume normal ventilation with one touch
3.1.9 | The unit must provide for leak test in the beginning and compensate the leak through out operation
3.1.10 | The ventilator must be mounted on a trolley with at least 4 antistatic swivel castors of at least 125mm diameter of which 2 must be lockable.
3.1.11 | Trolley must have a push-handle
3.1.12 | The trolley must have an oxygen cylinder holder
3.1.13 | The trolley must have a rail on both sides to fit the circuit arm
3.1.14 | Removeable arm with circuit hanger to fit the side rails
3.1.15 | Expiratory inlet to include integral water trap, valved to allow emptying without interruption of normal patient
3.1.16 | Air filter capacity at inlet: 99 % (for > 0.5 um)
3.1.17 | The ventilator must have a colour LCD screen with at least three simultaneous real-time waveforms
3.1.18 | The ventilator must measure and display inspiratory/expiratory loops.
3.1.19 | Waveforms and loops must be colour coded to denote inspiratory and expiratory breath phases
3.1.20 | The graphics display system must have a freeze capability.
3.1.21 | Loops: Flow/Volume Loop, Pressure/Volume Loop
3.1.22 | Reference loops must be able to be stored and displayed in the background in different colours
3.1.23 | All primary controls must be visible and changeable from all screens without entering a menu
3.1.24 | Monitored values can be arranged to match sequence of hospital's ventilator check sheet
3.2 | The following breath types must be selectable:
3.2.1 | Volume
3.2.2 | Pressure
3.2.3 | Pressure regulated volume control
3.2.4 | Pressure support ventilation
3.2.5 | Spontaneous
3.3 | The ventilator must have at the least the following modes of ventilation (Invasive and non-invasive ventilation must be possible):
3.3.1 | Volume ventilation with:
3.3.1a | Assist/control
3.3.1b | SIMV (Synchronous Intermittent Mandatory Ventilation)
3.3.1c | PSV (Pressure Support Ventilation)
3.3.1d | Option to disable demand flow
3.3.1e | Option of square or decelerating flow waveform
3.3.2 | Pressure ventilation (IPPV) with:
3.3.2a | Assist/control
3.3.2b | SIMV (Synchronous Intermittent Mandatory Ventilation)
3.3.2c | PSV (Pressure Support Ventilation)
3.3.2d | CPAP
3.3.2e | NPPV (Non-invasive positive pressure ventilation) with leak compensation and apnea back-up
3.3.3 | PRVC (Pressure Regulated Volume Control)
3.3.4 | APRV (Airway Pressure Release Ventilation) /BiPAP ( Biphasic , Duo Level, BiLevel)
3.3.5 | Spontaneous ventilation
3.3.6 | Apnea monitoring must be active in A/C, SIMV, CPAP and APRV
3.3.6a | Apnea mode: user settable time duration before ventilation onset
3.4 | BREATH CONTROL
3.4.1 | Breath rate: 1 - 120 bpm
3.4.2 | Tidal volume: 20 - 2000 ml
3.4.3 | Peak flow: 5 - 100 litres/min
3.4.4 | PEEP: 0 - 50 cm H2O
3.4.5 | Inspiratory flow trigger: 0.5 - 15 litres/min
3.4.6 | Pressure trigger: 0,5 - 10 mbar below PEEP/CPAP
3.4.7 | Bias flow: 0.4 - 5 litres/min
3.4.8 | Expiratory hold: 0 - 5 sec
3.4.9 | Inspiratory hold: 0 - 5 sec
3.4.10 | Inspiratory pause: 0 - 2 sec
3.4.11 | Inspiratory time: 0,2 - 5 sec
3.4.12 | Manual breath - add single control breath
3.4.13 | Inspiratory waveform: square and decelerating waves
3.4.14 | Inspiratory pressure control level: 1 - 80 cm H2O
3.4.15 | In ARPV mode:
3.4.15a | Pressure high: 0 - 50 cm H2O
3.4.15b | Time high: 0,2 - 30 sec
3.4.15c | Pressure low: 0 - 40 cm H2O
3.4.15d | Time low: 0,3 - 30 sec
3.5 | Leakage compensation in all non-invasive ventilation modes must be possible.
3.6 | The unit must have a circuit compliance compensation 0 to 7.5 ml/cmH2O and must be automatically measured.
3.7 | I:E ratio: 1:10 to 4:1
3.8 | PEEP/CPAP must be possible from 0 to approximately 50 cm H2O
3.9 | OXYGEN CONTROL
3.9.1 | FiO2: 21 – 100%
3.9.2 | Increase % oxygen: single key activation to deliver up to100% O2 for 2 min. FiO2 to be adjustable by user
3.10 | MONITOR
3.10.1 | Volume monitoring:
3.10.1.1 | Monitored values to include Tidal and Minute Volumes divided by patent's weight (ml/kg)
3.10.1.2 | Tidal volume
3.10.1.3 | Minute volume
3.10.1.4 | Mandatory expired tidal volume
3.10.1.5 | Spontaneous expired tidal volume
3.10.1.6 | Peak Inspiratory Flow Rate (PIFR): 0 to 180 litres/min
3.10.1.7 | Peak Expiratory Flow Rate (PEFR): 0 to 180 litres/min
3.10.2 | Pressure monitoring:
3.10.2a | Peak inspiratory pressure: 0 - 120 cm H2O
3.10.2b | Mean airway pressure: 0 - 120 cm H2O
3.10.2c | Plateau pressure: 0 - 100 cm H2O
3.10.2d | PEEP: 0 - 50 cm H2O
3.10.2e | Oxygen inlet pressure: 40 - 95 psig
3.10.2f | Delivered O2: 21 - 100 %
3.10.3 | Time cycle monitoring:
3.10.3a | Breath rate: 0 - 120 bpm
3.10.3b | Inspiratory time: 0 - 60 sec
3.10.3c | Exhalation time: 0 - 60 sec
3.10.3d | Rapid Mustow breathing index 0 - 500 bpm/litre
3.11 | The following parameters must be available as time-base curves:
3.11.1 | Pressure
3.11.2 | Flow
3.11.3 | Tidal volume
3.12 | Display of the following patient parameters must be user selectable and at least three selected parameters must be displayed as digital values
3.12.1 | Spontaneous minute volume
3.12.2 | Inspiratory flow
3.12.3 | Peak airway pressure
3.12.4 | Mean airway pressure
3.12.5 | Plateau pressure
3.12.6 | PEEP
3.12.7 | Expired tidal volume
3.12.8 | Expired minute volume
3.12.9 | Inspiratory time
3.12.10 | Breath rate
3.12.11 | Spontaneous rate
3.12.12 | Internal oxygen inlet pressure (pipeline inlet pressure)
3.12.13 | Battery charge status
3.12.14 | I:E ratio
3.12.15 | Set I:E ratio
3.12.16 | Static airways compliance in ml/cmH2O
3.12.17 | Static airways compliance per kg in ml/cmH2O
3.12.18 | Dynamic airways compliance in ml/cmH2O
3.12.19 | Dynamic airways compliance per kg in ml/cmH2O
3.12.20 | Peak inspired flow rate in l/m
3.12.21 | Peak expired flow rate in l/m
3.12.22 | Inspired O2 in %
3.12.23 | Tube leak in %
3.13 | SETUP
3.13.1 | Patient details
3.13.2 | Patient size: Adult / Paediatric
3.13.3 | Patient weight
3.14 | TRENDS
3.14.1 | Trends: at least 8 parameters for at least 24 hours
3.14.2 | Trends must be configurable to display at least 4 values as histograms with ability to review with cursor/time stamp
3.14.3 | Automatically recorded events:
3.14.3.1 | New patient
3.14.3.2 | Settings Change
3.14.3.3 | Increase 02
3.14.3.4 | Power On
3.14.3.5 | Insp Hold
3.14.3.6 | Power Off
3.14.3.7 | Exp Hold
3.14.3.8 | Power Recovery
3.14.3.9 | Nebulizer
3.14.3.10 | Enter Standby
3.14.3.11 | Manual Breath
3.14.3.12 | Exit Standby
3.15 | ALARMS: The ventilator must have the following audible and visual alarms:
3.15.1 | Low/high airway pressure
3.15.2 | Low/high minute volume
3.15.3 | Low/high tidal volume
3.15.4 | Low/high breath rate:1 to 200 bpm
3.15.5 | Low/high FiO2
3.15.6 | Loss of PEEP
3.15.7 | Low/high PEEP
3.15.8 | Low/high PIP
3.15.9 | Leakage alarm
3.15.10 | Apnea: 10 - 60 sec
3.15.11 | Patient disconnect
3.15.12 | Oxygen concentration variation (± 5%)
3.15.13 | Air/oxygen % deviation
3.15.14 | Power failure
3.15.15 | Air and oxygen supply failure
3.15.16 | Air and oxygen low pressure
3.15.17 | Battery low (When approximately 15 min operation left).
3.15.18 | Apnea
3.15.19 | Alarm limits must be user-adjustable
3.15.20 | Default alarm limits if user does not set alarm limits
3.15.21 | Alarm tone differentiates high, medium and low priority alarms
3.15.22 | Adjustable alarm volume
3.15.23 | Temporary mute for audible alarm: 0 - 2 min
3.15.24 | Alarm Reset. Deactivates visual indicators for resolved alarm conditions
3.16 | Machine Alarms:
3.16.1 | Safety Valve Open
3.16.2 | Ventilator Inoperative
3.16.3 | Fan Failure
3.16.4 | Loss of Air -Air Supply pressure < 18.0 psig
3.16.5 | Loss of O2 – O2 Supply pressure <18.0 psig (when % O2 > 21)
3.17 | GAS SUPPLY
3.17.1 | Oxygen pipeline supply: Pressure: 200 to 600 kPa) Minimum flow: 80 L/min at 1.5 bar
3.17.2 | Air pipeline supply: Pressure: 200 to 600 kPa) Minimum flow: 80 L/min at 1.5 bar
3.17.3 | Medical air and oxygen hoses, 3m long, colour coded, mounted with Heyer oxygen and medical air probes. The Heyer probes must be compatible with the wall points used in the instutution.
3.18 | POWER SUPPLY
3.18.1 | The unit must be powered by 220V, 50 Hz mains supply
3.18.2 | Internal rechargeable battery to operate the unit for approximately 2.5 hours
3.18.3 | Automatic switch to battery in case of power failure,
3.18.4 | Automatic battery recharge when connected to mains
3.18.5 | Power cord at least 3 m long with SABS approved 15A three prong plug
4 | The following Accessories and Consumables must be supplied with the unit:
4.1 | Patient circuit
4.2 | Reusable, sterilizable Paeds Mask
4.3 | Reusable, sterilizable Adult Mask
4.4 | Flow sensor.
4.5 | Test lung
Test circuit
4.6 | Bacteria filter (pk of 5)
4.9 | Air filters (pk of 5)
5 | Conformity Compliance (Please attach certificate)
5.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
5.2 | OEM must comply and certified on ISO 9001 quality standards, attach proof of compliance
5.2.1 | Please provide unique ref number of the ISO 9001 ceriticate:
5.3 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
5.3.1 | Please provide unique ref number of the ISO 13485 ceriticate:
5.4 | Model quoted for must be EC certified. Attach a copy of certification
5.4.1 | Please provide unique ref number of the EC ceriticate:
5.5 | Hazardous Substance Act: If this type of equipment / apparatus appears on the schedule of Hazardous substances issued by the Directorate: Health Technology of the Department of Health, a license in terms of the Act on Hazardous Substances (Act 15/1973) must be submitted with the bid / quotation documents. The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.6 | All electrical/electronic medical equipment must be licensed by Radiation Control, where applicable a copy of the license must be submitted The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.6.1 | Radiation Control License number:
5.7 | The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade (rubber to with 2 screws) type and it must be a minimum length of three (3) meters. NB the mains cable of the unit tendered for must be SABS colour coded.
5.8 | The equipment quoted for must be protected against electromagnetic interference IEC 60601-1-2
5.9 | The unit must comply with electrical safety and performance standard such as IEC 60601-2-12, attached certification
5.9.1 | Please provide unique ref number of the IEC ceriticate:
6 | COMPREHENSIVE MAINTENANCE
6.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) must be included in the bid price
6.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
6.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
6.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
6.3 | The bidder to provide the OEM recommended service interval of the unit:
6.4 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
7 | General
7.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
7.2 | The unit on offer must be the latest technology.
VENTILATOR PAEDIATRIC/ADULT
RT2-02-02a | Reusable patient circuit (Box of 10)
RT2-02-02b | Reusable, sterilizable Paeds Mask (Inflatable)
RT2-02-02c | Reusable, sterilizable Adult Mask (Inflatable)
RT2-02-02d | Flow sensor.
RT2-02-02e | Test lung
RT2-02-02f | Test circuit
RT2-02-02g | Bacteria filter (pk of 5)
RT2-02-02h | Air filters (pk of 5)
RT2-02-02i | Adult disposable patient circuit (Pk of 10)
RT2-02-02j | Paeds disposable patient circuits (Pk of 10)
RT2-02-02k | Temperature probes for humidifier
UPGRADES
RT2-02-02l | Built-in nebulizer
RT2-02-02m | Compressor air: 0.2 to 0.7 bar
RT2-02-02n | Humidifier
The humidifier must be dual servo controlled with an alarm system
All accessories for the humidifier must be included
Humidifier must have a hose heating element, dual temperature probe and electrical adapter.
RT2-02-02o | End Tidal CO2 curve
RT2-02-02p | Printer
RT2-02-02q | MANEUVERS
Ventilator must provide separate manoeuvres for MIP/NIF P100, AutoPEEP airway, Slow Flow Loop Inflection Points
Ventilator must return to normal ventilation at end of manoeuvre without any changes required to primary controls
Ventilator must have "Sensitivity" setting for maneuvre only, separate from primary "Sensitivity" setting. Allowing Patient to abort maneuvre.
Inflection point maneuvre must calculate upper and lower inflection points and display on Pressure/Volume loop
Stored inflection point loops must be able to be superimposed on current loop for comparison

## Ventilator, CPAP
RT2-02-05 | VENTILATOR CPAP
Description
1a | Product Offered:
1a.1 | Make:
1a.2 | Model:
1a.3 | Initial Manufacturing year of Model:
1a.4 | Radiation Control License number:
1 | Technical Specification
1.1 | An electronically controlled & monitored CPAP flow system to safely and effectively deliver fresh gas via a nasal interface or mask to the airways of neonate and paediatric patients.
1.2 | Electrically powered and microprocessor driven, monitored and controlled.
1.3 | The CPAP system must consist of the following:
1.3.1 | Blender
1.3.2 | O2 analyser
1.3.3 | Flowmeter
1.3.4 | Electronic pressure gauge
1.3.5 | Alarm system
1.4 | Manual breath inspiration/hold feature with ability to ensure sustained lung inflation which is used as a recruitment manoeuvre on infants with atelectasis
1.5 | Manual time for a manual breath adjustable from 2 to 10 sec for this manoeuvre
1.6 | Manual breath delivered in any mode adjustable 5 - 15 cm H2O
1.7 | Oxygen flush to ensure good oxygenation during suctioning etc 23 - 100%
1.8 | Must have leak compensation to allow for various cannulae and prongs to be used
1.9 | Able to deliver Demand flow in nCPAP thereby reducing the work of breathing and not have pre-set limited flow in nCPAP
1.10 | CPAP pressure range: 0 to 35 cm H2O
1.11 | nCPAP 2 - 13 cm H2O
1.12 | O2 flush time adjustable 30 - 120 sec
1.13 | The unit must be able to supply two independent pressure levels via 2 flow meters.
1.14 | Correctly coded high pressure gas supply hoses for oxygen and air of three metre length must be supplied with the unit. Appropriate type oxygen and air probes to be fitted. Probes to be in accordance with SABS 1409.
1.15 | Supply air/oxygen mixture to the patient at a flow rate between 0 – 15 lpm (95% accuracy)
1.18 | The control unit must automatically go through a self diagnosis and checking process.
1.19 | The control unit must have direct input connections for oxygen and medical air
1.20 | The source for oxygen either from a central wall outlet (gas bank) or an oxygen cylinder.
1.21 | The source for medical air from a central wall outlet (gas bank) and medical air cylinder.
1.22 | The input supply pressure for oxygen and medical air: 200 – 500 kPa.
1.23 | The unit must automatically switch between the piped air and oxygen supply should one of the supplies fail. (Alarm condition should be triggered in any event)
1.24 | Must provide compensation for leaks
1.25 | The unit must be user friendly with all circuit connections, front panel controls and displays ergonomically grouped by function and clearly marked.
1.26 | Oxygen concentration setting variable between 21 % and 100 %.
1.30 | Have a manometer with a range of 0 – 12 cm H2O
1.31 | Breath rate: 1 - 120 bpm
1.32 | Timed inspiration: 0.1 to 3.0 sec.
1.33 | Apnea backup ventilation: 1 - 80 bpm.
1.34 | Apnea interval: 10 - 30 seconds.
1.35 | Maximum safety pressure set to ± 3 cm H20 above measured airway pressure.
1.36 | Unit must have safety features to prevent:
1.36.1 | Volutrauma
1.36.2 | Barotrauma
1.36.3 | Over stretching of the lungs
1.37 | The unit must have a proximal pressure monitoring line.
1.38 | Must have a method to prevent it becoming occluded by water vapour.
1.39 | Must monitor the following parameters:
1.39.1 | nCPAP
1.19.2 | MAP
1.39.2 | PIP
1.19.3 | PEEP
1.39.3 | Fi02 (21% to 100%)
1.40 | Graphics Display
1.40.1 | Measured values to be displayed:
1.40.1.1 | CPAP
1.40.1.2 | PIP
1.40.1.3 | MAP
1.40.2 | The graphics display system must have freeze capabilities.
1.41 | The offered unit must display trending of at least 72 hours on at least 20 parameters of pressure-, flow- and tidal volumes.
1.42 | Alarms:
1.42.1 | Patient disconnection
1.42.2 | High airway pressure
1.42.3 | Low airway pressure.
1.42.4 | High oxygen alarm
1.42.5 | Low oxygen alarm
1.42.6 | Low nCPAP alarm: -2 to -20 cm H2O
1.42.7 | High nCPAP alarm: +2 to + 20 cm H2O
1.42.8 | Oxygen supply failure
1.42.9 | Air supply failure
1.42.10 | Apnoea: off; 10 - 20 sec
1.42.11 | Electrical supply failure
1.42.12 | Low battery
1.43 | A safety mechanical pressure release (overpressure) valve must be present
1.44 | The above alarm parameters must auto set themselves within 2 minutes of power up of drive gas flows and oxygen concentrations or by means of a single key setting.
1.45 | Violation of any of the set alarm limits must result in both audible and visual alarms.
1.46 | It must be possible to silence the alarm for a maximum of 120 seconds.
1.47 | The unit must be mounted on a mobile stand/trolley, with 4 to 5 castors (all lockable) with a facility to connect a humidifier.
1.48 | The unit must have a rail and circuit hanger on both sides of the trolley.
1.49 | The unit must operate from a 220 V, 50Hz supply.
1.50 | Mains cable at least three metres long with SABS approved 15A three prong plug.
1.51 | The offered unit shall have automatic switch over to a built in battery unit with a capacity to ensure full functionality of the unit for not less than 2 hour.
1.52 | The unit must be fitted with an internal rechargable battery.
1.53 | Battery must be capable of at least 2 hours operation
1.54 | Battery charger to be incorporated in the unit
1.55 | Must display battery status and give an alarm when battery is low
1.57 | The unit must be handed over in full operating order
2 | Standard Accessories and Consumables
2.1 | 5 sets of silicone nasal prongs (each set must include 3 different neonatal sizes)
2.2 | 5 sets of CPAP bonnets (each set to include 3 different neonatal sizes)
2.3 | 5 neonatal circuits
2.4 | Humidifier
2.5 | Temperature probes for humidifier
2.6 | Circuit holding arm
3 | Conformity Compliance (Please attach certificate)
3.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
3.2 | OEM must comply and certified on ISO 9001 quality standards, attach proof of compliance
3.2.1 | Please provide unique ref number of the ISO 9001 ceriticate:
3.3 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
3.3.1 | Please provide unique ref number of the ISO 13485 ceriticate:
3.4 | Model quoted for must be EC certified. Attach a copy of certification
3.4.1 | Please provide unique ref number of the EC ceriticate:
3.5 | Hazardous Substance Act: If this type of equipment / apparatus appears on the schedule of Hazardous substances issued by the Directorate: Health Technology of the Department of Health, a license in terms of the Act on Hazardous Substances (Act 15/1973) must be submitted with the bid / quotation documents. The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
3.6 | All electrical/electronic medical equipment must be licensed by Radiation Control, where applicable a copy of the license must be submitted The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
3.6.1 | Radiation Control License number:
3.7 | The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade (rubber to with 2 screws) type and it must be a minimum length of three (3) meters. NB the mains cable of the unit tendered for must be SABS colour coded.
3.8 | The equipment quoted for must be protected against electromagnetic interference IEC 60601-1-2
3.9 | The unit must comply with electrical safety and performance standard such as IEC 60601-2-12, attached certification
3.9.1 | Please provide unique ref number of the IEC ceriticate:
4 | COMPREHENSIVE MAINTENANCE
4.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) must be included in the bid price
4.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries, Flow sensors and oxygen cell), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
4.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
4.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries, Flow sensors and oxygen cell), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
4.3 | The bidder to provide the OEM recommended service interval of the unit:
4.4 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
5 | General
5.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
5.2 | The unit on offer must be the latest technology.
VENTILATOR CPAP ACCESSORIES and CONSUMABLES
RT2-02-05a | 5 sets of silicone nasal prongs (each set must include 3 different neonatal sizes)
RT2-02-05b | 5 sets of CPAP bonnets (each set to include 3 different neonatal sizes)
RT2-02-05c | 5 neonatal circuits
RT2-02-05d | Temperature probes for humidifier
RT2-02-05e | Circuit holding arm
Bidder to list any other Consumables or Accessories (Not listed above) for the model on offer below

## Ventilator, Transport, Adult
RT2-02-06 | Transport Adult / Paediatric
1 | Description
2 | Product Offered:
2.1 | Make:
2.2 | Model:
2.3 | Initial Manufacturing year of Model:
2.4 | Radiation Control License number:
3 | Technical Specification
3.1 | High-end Transport ventilator, suitable for adult / paediatric patients, that integrates ventilation with comprehensive monitoring capability,with intergrated Air supply via compressor or scroll pump. Hot swap battery function .An optional colour display of ventilation parameters, graphics and trends must be available.
3.2 | High pressure medical oxygen connection must be interchangeable by the clinician for Oxygen with appropriate CGA gas fitting from wall supply to Oxygen Cylinder.
3.3 | Ventilator must have intergrated Air compressor/ scroll pump able to deliver +/- 120 lpm
3.4 | Ventilator shall have leak compensation, able to compensate for baseline leaks which may occur at the patient/mask interface or around artificial airways
3.5 | Ventilator shall have "Standby" mode, which suspends alarms and breath cycling, suspends trends.
3.6 | "Standby" mode shall retain current ventilator settings and alarms, resume normal ventilation with one touch
3.7 | CONFIGURATION AND SETUP
3.7.1 | User interface should allow for entering or selecting:
3.7.2 | Current patient or new patient
3.7.3 | Patient size range – Infant/Adult / Pediatric
3.8 | VENTILATOR MODE CONTROL and TYPES
3.8.1 | BREATH TYPES
3.8.1a | Volume
3.8.1b | Pressure
3.8.1c | Pressure support ventilation
3.8.1d | Spontaneous
3.8.2 | MODE TYPES
3.8.2.1 | Assist/control (A/C) – pressure and volume:
3.8.2.1a | Volume Controlled mode shall have option to disable demand flow and shall offer option of square or decelerating flow waveform
3.8.2.2 | Synchronous Intermittent Mandatory Ventilation (SIMV) – Pressure, Volume and PRVC
3.8.2.3 | Continuous Positive Airway Pressure / Pressure Support Ventilation (CPAP/PSV)
3.8.2.3a | Variable Rise Time in all modes.
3.8.2.3b | Variable breath Time Termination.
3.8.2.3c | Pressure Control: Variable Flow Termination.
3.8.2.4 | Non-invasive positive pressure ventilation (NPPV)
3.8.2.4a | Leak Compensation.
3.8.2.4b | Apnea backup should be active in A/C, SIMV, CPAP and Non Invasive.
3.8.2.4c | Apnea modes shall include user settable time duration before Apnea ventilation onset
3.9 | BREATH CONTROLS
3.9.1 | Breath rate: 1 – 80bpm
3.9.2 | Inspiratory flow trigger: Off, 1 – 9.0 L/min
3.9.3 | Manual breath – add single control breath
3.9.4 | Expiratory hold: max 6.0 sec
3.9.5 | Inspiratory hold: max 6.0 sec
3.9.6 | PC / PS / Spont. Flow 160 lpm
3.9.7 | Tidal Volume: 50 – 2000 ml
3.9.8 | Inspiratory pressure control level: 0 – 90cmH2O
3.9.9 | Inspiratory pressure Support level: 0 – 60cmH2O
3.9.10 | Inspiratory Time: 0.3 – 9.9 sec
3.9.11 | Calculated I:E ratio: 1:4 and 4:1
3.9.12 | PEEP: 0 – 20 cmH2O
3.10 | OXYGEN CONTROL
3.10.1 | % Oxygen: 21 – 100% ( 1% adjustments)
3.10.2 | Use of Low pressure Oxygen on/off ( Calculated Time left in Cylinder )
3.11 | PATIENT MONITORING (Volume, Pressure, Rate & Time)
3.11.1 | Ventilator shall have a variable orifice flow sensor for all patient ranges
3.11.2 | Monitored values to include Tidal and Minute Volumes.
3.11.3 | Tidal Volume (Vt) -Absolute range: 0 to 4 L (All patients) Normalized range: 0.0 to 30.0 ml / Kg (All patients)
3.11.4 | Exhaled Tidal Volume
3.11.5 | Inspired Tidal Volume
3.11.6 | Minute Volume - Absolute range:O to 99.9 L (All patients) Normalized Range: 0 to 999 (All patients)
3.11.7 | Patient Effort Indicator:
3.11.8 | Breath Rate -0 to 80 BPM
3.11.9 | Inspiratory Time (Tj)
3.11.10 | I:E Ratio (I:E) -1 :4 to 4:1
3.11.11 | Rapid Shallow Breathing Index (f / Vt) -0 to 500 b2 / min / L
3.11.12 | Peak Inspiratory Pressure (PPEAK) -0 to 120 cmH20
3.11.13 | Mean Airway Pressure (PMEAN) -0 to 99 cmH20
3.11.14 | PEEP -0 to 50 cmH20
3.11.15 | Oxygen Inlet Pressure (02 Inlet) 140 to 550 kpa (1.4-5.5 bar)
3.11.16 | Delivered O2 (Fi02) -0 to 100%
3.12 | ALARMS
3.12.1 | Alarms shall have distinctive audible tone
3.12.2 | Audible alarm volume
3.12.3 | Alarm Silence -120 sec. max
3.12.4 | Alarm Reset. Deactivates visual indicators for resolved alarm conditions
3.12.5 | Volume Alarms
3.12.5a | Low Exhaled Minute Volume (Low Ve) -Off (indicated by 0), 1 to 99 L
3.12.6 | Pressure Alarms
3.12.6a | Low PPEAK Off to 60 cmH20
3.12.6b | High PPEAK Alarm 5 to 100 cmH20
3.12.6c | High PEEP: Off, 3 to 40 cmH20
3.12.6d | Low PEEP Off -3 to 20 cmH20
3.12.6e | Maximum Circuit Pressure Limit -Maximum pressure at the y-connection: 125 cmH20
3.12.7 | Rate / Time Alarms
3.12.7a | Apnoea Interval. 10 to 60 seconds
3.12.7b | High Rate:
3.12.7c | Auto Apnoea Backup:
3.12.8 | Machine Alarms
3.12.8a | Safety Valve Open (SAFETY VALVE)
3.12.8b | Ventilator Inoperative (VENT INOP)
3.12.8c | Fan Failure
3.13 | Display
3.13.1 | Graphical display screen must be available
3.14 | MANEUVRES
3.14.1 | Spontaneous Breathing Trial (SBT)
3.14.2 | Ventilator shall return to normal ventilation at end of maneuver without any changes required to primary controls
3.15 | POWER SUPPLY
3.15.1 | Input voltage: 115 – 230 V and 11 –15 VDC
3.15.2 | Operating frequency: 50-60 Hz
3.15.3 | Power cord > 3 m fitted with standard 3 point plug
3.15.4 | Internal Battery back-up: Re-chargeable batteries with minimum duration of 1 hour.
3.15.5 | State duration of battery operation in minutes
3.15.6 | Maximum time to full charge 8 to 12 hours
3.16 | GAS SUPPLY
3.16.1 | Oxygen Pipeline Supply: Pressure Range 40 to 80 psig (2.8 to 5.5 bar)
3.16.2 | Low Pressure Source: <80 lpm <10 PSIG
3.17 | DATA INPUT I OUTPUT
3.17.1 | Digital Communication -RS-232 Ch1 (input for software updates/enhancements) RS-232 Ch2 (Undefined)
4 | STANDARD ACCESSORIES
4.1 | Adult breathing circuit
4.2 | External charger
4.3 | Carry bag and bed clamp
5 | Conformity Compliance (Please attach certificate)
5.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
5.2 | OEM must comply and certified on ISO 9001 quality standards, attach proof of compliance
5.2.1 | Please provide unique ref number of the ISO 9001 ceriticate:
5.3 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
5.3.1 | Please provide unique ref number of the ISO 13485 ceriticate:
5.4 | Model quoted for must be EC certified. Attach a copy of certification
5.4.1 | Please provide unique ref number of the EC ceriticate:
5.5 | Hazardous Substance Act: If this type of equipment / apparatus appears on the schedule of Hazardous substances issued by the Directorate: Health Technology of the Department of Health, a license in terms of the Act on Hazardous Substances (Act 15/1973) must be submitted with the bid / quotation documents. The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.6 | All electrical/electronic medical equipment must be licensed by Radiation Control, where applicable a copy of the license must be submitted The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.6.1 | Radiation Control License number:
5.7 | The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade (rubber to with 2 screws) type and it must be a minimum length of three (3) meters. NB the mains cable of the unit tendered for must be SABS colour coded.
5.8 | The equipment quoted for must be protected against electromagnetic interference IEC 60601-1-2
5.9 | The unit must comply with electrical safety and performance standard such as IEC 60601-2-12, attached certification
5.9.1 | Please provide unique ref number of the IEC ceriticate:
611 | COMPREHENSIVE MAINTENANCE
6.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) must be included in the bid price
6.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries, Flow sensors and oxygen cell), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
6.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
6.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries, Flow sensors and oxygen cell), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
6.3 | The bidder to provide the OEM recommended service interval of the unit:
6.4 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
7 | General
7.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
7.2 | The unit on offer must be the latest technology.
Transport Adult / Paediatric Accessories and Consumables
RT2-02-06a | Adult breathing circuit
RT2-02-06b | External charger
RT2-02-06c | Carry bag and
RT2-02-06d | Bed clamp
Bidder to list any other Consumables or Accessories (Not listed above) for the model on offer below

## Ventilator, Transport, Neonatal
RT2-02-07 | Ventilator ,Transport , Neonatal
Description
1a | Product Offered:
1a.1 | Make:
1a.2 | Model:
1a.3 | Initial Manufacturing year of Model:
1a.4 | Radiation Control License number:
1 | Technical Specification
1.0 | A Basic Transport ventilator, suitable for Neonatal patients, that integrates ventilation with basic monitoring capability and a miniminum battery life of 4 hours.
1.1 | High pressure medical oxygen connection must be interchangeable by the clinician for Oxygen with appropriate CGA gas fitting from wall supply to Oxygen Cylinder.
2 | CONFIGURATION AND SETUP
2.1 | User interface should allow for entering or selecting:
2.1.1 | Oxygen /Air mixture 21% to 100%
2.1.2 | Mode of Ventilation
2.2 | Must be able to measure mainstream eCo2
3 | VENTILATOR MODE CONTROL and TYPES
BREATH TYPES
3.1 | Volume
3.2 | Pressure
3.3 | Spontaneous /CPAP
MODE TYPES
3.4 | Assist/control (A/C) – pressure and volume
3.6 | Continuous Positive Airway Pressure (CPAP)
3.7 | Variable breath Time Termination.
3.8 | Apnea Alarm should be active in A/C, SIMV, CPAP
4 | BREATH CONTROLS
4.1 | Breath rate: 1 - 80 bpm
4.3 | Tidal Volume: 5 ml
4.4 | Inspiratory pressure control level: 0 – 60cmH2O
4.5 | Inspiratory Time: 0.25 – 2 sec
4.6 | Calculated I:E ratio: 1:4 and 4:1
4.7 | PEEP: 0 – 20 cmH2O
5 | OXYGEN CONTROL
5.1 | % Oxygen: 21% – 100%
6 | PATIENT MONITORING
6.1 | Rate & Time Monitors
6.1.1 | Breath Rate -0 to 150 BPM
6.2 | Pressure Monitors
6.2.1 | Peak Inspiratory Pressure (PPEAK) -0 to 100 cmH20
6.2.2 | Mean Airway Pressure (PMEAN) -0 to 99 cmH20
6.2.3 | PEEP -0 to 50 cmH20
7 | ALARMS
7.1 | Alarms shall have distinctive audible tone
7.2 | Alarm Silence -120 sec. max
7.3 | Alarm Reset. Deactivates visual indicators for resolved alarm conditions
7.4 | Low Paw cmH2O
7.5 | High Paw cmH20
7.6 | Maximum Circuit Pressure Limit -Maximum pressure at the y-connection:60 cmH20
7.7 | Apnoea Interval
7.8 | Safety Valve Open (SAFETY VALVE)
7.9 | Ventilator Inoperative (VENT INOP)
7.10 | Fan Failure
7.11 | Loss of O2 – O2 Supply pressure and Duration Calculation:
8 | MANEUVERS
8.1 | Manual Breath
9 | POWER SUPPLY
9.1 | Input voltage: 115 – 230 V and 11 –15 VDC
9.2 | Operating frequency: 50-60 Hz
9.3 | Power cord > 3 m fitted with standard 3 point plug
9.4 | Internal Battery back-up: Re-chargeable batteries with minimum duration of 4 hour.
9.5 | State duration of battery operation in minutes
9.6 | Maximum time to full charge 4 to 6 hours
10 | GAS SUPPLY
10.1 | Oxygen Pipeline Supply: Pressure Range 280 to 550 kpa (2.8 to 5.5 bar)
10.2 | Low Pressure Source: <80 lpm <68 kpa
11 | ACCESSORIES
11.1 | Neonatal breathing circuits
11.2 | External charger
11.3 | Carry bag and bed clamp
12 | Conformity Compliance (Please attach certificate)
12.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
12.2 | OEM must comply and certified on ISO 9001 quality standards, attach proof of compliance
12.2.1 | Please provide unique ref number of the ISO 9001 ceriticate:
12.3 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
12.3.1 | Please provide unique ref number of the ISO 13485 ceriticate:
12.4 | Model quoted for must be EC certified. Attach a copy of certification
12.4.1 | Please provide unique ref number of the EC ceriticate:
12.5 | Hazardous Substance Act: If this type of equipment / apparatus appears on the schedule of Hazardous substances issued by the Directorate: Health Technology of the Department of Health, a license in terms of the Act on Hazardous Substances (Act 15/1973) must be submitted with the bid / quotation documents. The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
12.6 | All electrical/electronic medical equipment must be licensed by Radiation Control, where applicable a copy of the license must be submitted The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
12.6.1 | Radiation Control License number:
12.7 | The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade (rubber to with 2 screws) type and it must be a minimum length of three (3) meters. NB the mains cable of the unit tendered for must be SABS colour coded.
12.8 | The equipment quoted for must be protected against electromagnetic interference IEC 60601-1-2
12.9 | The unit must comply with electrical safety and performance standard such as IEC 60601-2-12, attached certification
12.9.1 | Please provide unique ref number of the IEC ceriticate:
13 | COMPREHENSIVE MAINTENANCE
13.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) must be included in the bid price
13.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries,Flow sensors & Oxygen cell), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
13.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
13.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries, Flow sensors and oxygen cell), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
13.3 | The bidder to provide the OEM recommended service interval of the unit:
13.4 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
14 | General
14.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
14.2 | The unit on offer must be the latest technology.
Ventilator ,Transport , Neonatal Accessories and Consumables
RT2-02-07a | Neonatal breathing circuit
RT2-02-07b | External charger
RT2-02-07c | Carry bag and
RT2-02-07d | Bed clamp
Bidder to list any other Consumables or Accessories (Not listed above) for the model on offer below

## HF Oscillator
RT2-02-08 | HF Oscillator, Neonatal/Paediatric
Description
1a | Product Offered:
1a.1 | Make:
1a.2 | Model:
1a.3 | Initial Manufacturing year of Model:
1a.4 | Radiation Control License number:
2 | Technical Specification
2.1 | The unit must be a piston driven oscillation with no need for centering manually
2.2 | The unit must have at least 35cm colour LCD screen measured diagonally
2.3 | Power requirements: 220 V AC, 50Hz
2.4 | Exhaled volume must be measured through a Proximal Flow sensor
2.5 | It must be possible to select Flow trigger or Pressure trigger
2.6 | The trigger must be proximally detected at the patient end
2.7 | The volume delivered in every stroke must be according to the setting in any respiratory conditions
2.8 | The following parameters must be set independently:
2.8.1 | Mean Airway Pressure (MAP)
2.8.2 | Stroke Volume (SV)
2.8.3 | Frequency
3 | Range of controls
3.1 | Mean airway pressure 5-40 cmH2O
3.2 | Stroke volume 0 to 160 ml
3.3 | Frequency 5-20 Hz
3.4 | SI Pressure 3-50 cmH2O
3.5 | Base flow 10-30L/min
3.6 | Max mean pressure 10-50 cmH2O
3.7 | FiO2 21-100%
3.8 | It must be possible to suspend alarms for 2 mins
4 | Alarms
4.1 | Oxygen supply source low
4.2 | Air supply source source
4.3 | Power failure.
5 | Monitored variables
5.1 | Both waveforms and data for Pressure, Flow, Volume must be displayed
5.2 | It must be possible to display loops
6 | Conformity Compliance (Please attach certificate)
6.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
6.2 | OEM must comply and certified on ISO 9001 quality standards, attach proof of compliance
6.2.1 | Please provide unique ref number of the ISO 9001 ceriticate:
6.3 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
6.3.1 | Please provide unique ref number of the ISO 13485 ceriticate:
6.4 | Model quoted for must be EC certified. Attach a copy of certification
6.4.1 | Please provide unique ref number of the EC ceriticate:
6.5 | Hazardous Substance Act: If this type of equipment / apparatus appears on the schedule of Hazardous substances issued by the Directorate: Health Technology of the Department of Health, a license in terms of the Act on Hazardous Substances (Act 15/1973) must be submitted with the bid / quotation documents. The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
6.6 | All electrical/electronic medical equipment must be licensed by Radiation Control, where applicable a copy of the license must be submitted The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
6.6.1 | Radiation Control License number:
6.7 | The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade (rubber to with 2 screws) type and it must be a minimum length of three (3) meters. NB the mains cable of the unit tendered for must be SABS colour coded.
6.8 | The equipment quoted for must be protected against electromagnetic interference IEC 60601-1-2
6.9 | The unit must comply with electrical safety and performance standard such as IEC 60601-2-12, attached certification
6.9.1 | Please provide unique ref number of the IEC ceriticate:
7 | COMPREHENSIVE MAINTENANCE
7.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) must be included in the bid price
7.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries,Flow sensors & Oxygen cell), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
7.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
7.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries, Flow sensors and oxygen cell), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
7.3 | The bidder to provide the OEM recommended service interval of the unit:
7.4 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
8 | General
8.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
8.2 | The unit on offer must be the latest technology.

## Resus, Gas-powered, Infant
RT2-02-09 | Resuscitator, infant, gas powered
Description
1a | Product Offered:
1a.1 | Make:
1a.2 | Model:
1a.3 | Initial Manufacturing year of Model:
1 | Technical Specification
1.1 | Must be able to provide 100% oxygen for infants up to 10 kg
1.2 | The gas which drives the resuscitator must be the oxygen
1.3 | Oxygen flow regulator: 5 - 15 l/min
1.4 | Adjustable Peak Inspiratory Pressure (PIP): 2 - 75 cm H2O
1.5 | Adjustable Positive End Expiratory Pressure (PEEP): 1 - 10 cm H2O
1.6 | Resuscitator must provide controlled pressure settings by means of a pressure limiting device
1.7
1.8 | Pressure gauge: -10 to 80 cm H2O
1.9 | Maximum pressure relief: 20 - 70cmH²O @ 8lpm.
1.10 | The unit must be handed over in full operating order
1.11 | A starter-pack of consumables must be supplied with the unit.
2 | Conformity Compliance (Please attach certificate)
2.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
2.2 | OEM must comply and certified on ISO 9001 quality standards, attach proof of compliance
2.2.1 | Please provide unique ref number of the ISO 9001 ceriticate:
2.3 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
2.3.1 | Please provide unique ref number of the ISO 13485 ceriticate:
2.4 | Model quoted for must be EC certified. Attach a copy of certification
2.4.1 | Please provide unique ref number of the EC ceriticate:
3 | COMPREHENSIVE MAINTENANCE
3.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) must be included in the bid price
3.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
3.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
3.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
3.3 | The bidder to provide the OEM recommended service interval of the unit:
3.4 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
4 | General
4.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
4.2 | The unit on offer must be the latest technology.
Resuscitator, infant, gas powered, Additional Options
RT2-02-09a | Air/O2 blender (Upgrade)
RT2-02-09b | Suction availability (Upgrade)
Bidder to list any other Consumables or Accessories (Not listed above) for the model on offer below

## Humidifier, Servo control
RT2-02-10 | Heated Humidifier, servo control type
1 | DESCRIPTION
1.1 | a device design to adjust the amount of moisture in the respiratory device
1a | Product Offered:
1a.1 | Make:
1a.2 | Model:
1a.3 | Initial Manufacturing year of Model:
2 | Technical Specification
2.1 | The Air Humidification System needs to provide optimal humidity to all intensive and high care patients that is suitable for neonates to adults invansive and non inavsive, including nasal prong O2
2.2 | The unit will be used to heat and humidify passing air, replacing the function of the vital nasal and throat passages in patients
2.3 | The units shall have silent operation
2.4 | Proximal and Distal temperature servo feedback loop must be able to maintain the distal setup temperature constant. While in operation the set temperature on the unit must be able to adjusted
2.5 | The unit shall have no moving parts that require maintenance.
2.6 | The unit must have a mounting capability
2.7 | Bidders to state maximum temperature deviation from set temperature
2.8 | It shall be easy to refill the fluid, i.e. Autofill chamber with constant compressible volume.
2.9 | The unit shall require less than 30 minutes to warm up
2.1 | The unit shall operate from a standard 220V 50Hz electrical supply
2.11 | The unit shall operate with a flow range: Invasive 3 - 60 L/min, Non Invasive 3 -40 L/min
3 | STANDARD ACCESSORIES to be added
3.1 | Mounting bracket
3.2 | Heater wire x 5
3.3 | Temperature probe x 5
3.4 | Humidifier chamber
4 | Conformity Compliance (Please attach certificate)
4.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
4.2 | OEM must comply and certified on ISO 9001 quality standards, attach proof of compliance
4.2.1 | Please provide unique ref number of the ISO 9001 ceriticate:
4.3 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
4.3.1 | Please provide unique ref number of the ISO 13485 ceriticate:
4.4 | Model quoted for must be EC certified. Attach a copy of certification
4.4.1 | Please provide unique ref number of the EC ceriticate:
4.5 | The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade (rubber to with 2 screws) type and it must be a minimum length of three (3) meters. NB the mains cable of the unit tendered for must be SABS colour coded.
4.6 | The equipment quoted for must be protected against electromagnetic interference IEC 60601-1-2
5 | COMPREHENSIVE MAINTENANCE
5.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) must be included in the bid price
5.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS, Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
6 | General
6.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
6.2 | The unit on offer must be the latest technology.
Heated Humidifier, servo control type Accessories and Consumables
RT2-02-10a | Mounting bracket
RT2-02-10b | Heater wire
RT2-02-10c | Temperature probe
RT2-02-10d | Humidifier chamber
Bidder to list any other Consumables or Accessories (Not listed above) for the model on offer below

## Air-O2 Blender, Low Flow
RT2-02-11 | BLENDER - AIR/OXYGEN, LOW FLOW
1 | DESCRIPTION
1.1 | A medical device used to mix Medical Air and hospital grade Oxygen (O2) into a gas source ranging from 21% (normal) to 100% (max) oxygen
1a | Product Offered:
1a.1 | Make:
1a.2 | Model:
1a.3 | Initial Manufacturing year of Model:
2 | TECHNICAL SPECIFICATION
2.1 | The blender must be able to operate accurately on gas supply of between 2.5 and 5 bar.
2.2 | The Blender must be accurate up to 3%.
2.3 | The blender must have 2 intergrated flow meters:
2.4 | Flow meters with a range of 0 to 15 L.P.M flow and 0 to 1 or 3.5 L.P.M, with the latter calibrated in increments of at least 0.1 LPM.
2.5 | The O2 concentration control adjustment range must be between 21% - 100%.
2.7 | The alarm must be audible from 5 meters and must be high pitched. The alarm must be activated when there is a drop in wall O2 or Air pressure. The alarm must be demonstrated to an evaluation committee.
2.8 | The alarm bypass/ reset must operate when inlet pressure is corrected
2.9 | The pressure drop must be less than 45kpa at 3,5 bar inlet pressures and 40 L.P.M. flow rate.
2.1 | The weight must not exceed 3kg, please state the weight in kg
The unit must have 2 outlet ports
2.11 | The blender must come with a bracket suitable to the end user facility e.g. rail Bracket/ wall mounted bracket.
2.12 | The blender must come complete with colour-coded air and oxygen hoses and “S.I.S.S.” type O2 gas outlet connector and “N.I.S.T” type air and O2 connectors and must be supplied with colour- coded air and oxygen hoses with Heyer- type connectors that are compatible with the wall O2 points at end user institution.
4 | STANDARD ACCESSORIES
4.1 | Mounting bracket (Rail or pole bracket suitable for the institution)
4.2 | At least 1.5m Colour-coded hoses complete with heyer connectors
5 | Conformity Compliance (Please attach certificate)
5.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
5.2 | OEM must comply and certified on ISO 9001 quality standards, attach proof of compliance
5.2.1 | Please provide unique ref number of the ISO 9001 ceriticate:
5.3 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
5.3.1 | Please provide unique ref number of the ISO 13485 ceriticate:
5.4 | Model quoted for must be EC certified. Attach a copy of certification
5.4.1 | Please provide unique ref number of the EC ceriticate:
6 | COMPREHENSIVE MAINTENANCE
6.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) must be included in the bid price
6.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries, Flow sensors and Oxygen sensor), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
6.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
6.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS, labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
6.3 | The bidder to provide the OEM recommended service interval of the unit:
6.4 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
BLENDER - AIR/OXYGEN, LOW FLOW ACCESSORIES
RT2-02-11a | Mobile stand
RT2-02-11b | Mounting bracket
RT2-02-11c | Flow meter
RT2-02-11d | Control knob
Bidder to list any other Consumables or Accessories (Not listed above) for the model on offer below

## Air-O2 Blender, High Flow
RT2-02-12 | BLENDER - AIR/OXYGEN, HIGH FLOW
1 | DESCRIPTION
1.1 | A medical device used to mix Medical Air and hospital grade Oxygen (O2) into a gas source ranging from 21% (normal) to 100% (max) oxygen
1a | Product Offered:
1a.1 | Make:
1a.2 | Model:
1a.3 | Initial Manufacturing year of Model:
2 | TECHNICAL SPECIFICATION
2.1 | The blender must be able to operate accurately on gas supply of between 2.5 bars (29 psi) and 5 bars (72.5 psi).
2.2 | The blender must have a primary and an auxiliary outlet:
2.3 | Primary 15 – 20 L.P.M. flow range.
2.4 | Auxiliary 2 – 90 L.P.M. flow range with a bleed of no more than 10 – 12 L.P.M.
2.5 | The oxygen concentration control adjustment range must be between 21% and 100%.
2.6 | The maximum total flow at 60% on the control knob at 3,5 bar inlet pressure settings must be at least 120 L.P.M.
2.7 | Accuracy must be (with inlet gasses within 2 (two) bar and each gas pressure greater than 2 (two) bar but less than 5 (five) bar) + 10% of setting from 21% to 50% or + 5% oxygen which ever is smaller, accuracy maintained up to the audible alarm point.
2.8 | The alarm bypass activation must work when the inlet gas pressures differ between 1,37 bar (20 psi) and 2 bars (29 psi).
2.9 | The alarm must be audible from 5 meters. The alarm must be activated when there is a drop in wall O2 or Air pressure. It must be high pitched.
2.1 | The alarm / bypass reset must operate when inlet pressure differential is 0,5 bar (7.25 psi) or less.
2.11 | The pressure drop must be less than 45kpa at 3,5 bar (50.77 psi) inlet pressures and 40 L.P.M. flow rate.
2.12 | The weight must not exceed 3 kg. (please state the weight in Kg)
2.13 | The blender must come complete with colour-coded air and oxygen hoses and “S.I.S.S.” type O2 gas outlet connector and “N.I.S.T” type air and O2 connectors and must be supplied with colour- coded air and oxygen hoses with Heyer- type connectors that are compatible with the wall O2 points at end user institution.
2.14 | Mechanical Medical Oxygen / Air Blender
2.15 | Requirements :
2.16 | Deliver Oxygen Range 21 % to 100 %
2.17 | Flow Range 2 to 120 lpm.
2.18 | Maximum Bleed Flow range 0-12 lpm.
2.19 | 2 x Outlet ports
2.2 | The unit must alarm when the inlet pressures differ more than 1.37 bars (20 psi)
2.21 | 3m Colour Coded Oxygen and Air Hoses must be supplied.
2.22 | Appropriate Oxygen and Air Hose probes must be fitted.
3 | STANDARD ACCESSORIES
3.1 | Consumables used for day to day operations should be specified.
3.2 | Price and planned inflationary increase on consumable costs.
3.3 | Price of Optional Accessories ( Brackets or Flow Meters)
5 | Conformity Compliance (Please attach certificate)
5.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
5.2 | OEM must comply and certified on ISO 9001 quality standards, attach proof of compliance
5.2.1 | Please provide unique ref number of the ISO 9001 ceriticate:
5.3 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
5.3.1 | Please provide unique ref number of the ISO 13485 ceriticate:
5.4 | Model quoted for must be EC certified. Attach a copy of certification
5.4.1 | Please provide unique ref number of the EC ceriticate:
6 | COMPREHENSIVE MAINTENANCE
6.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) must be included in the bid price
6.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries, Flow sensors and Oxygen sensor), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
6.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
6.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS, labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
6.3 | The bidder to provide the OEM recommended service interval of the unit:
6.4 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
BLENDER - AIR/OXYGEN, HIGH FLOW ACCESSORIES
RT2-02-12a | Mobile stand
RT2-02-12b | Mounting bracket
RT2-02-12c | Flow meter
RT2-02-12d | Control knob
Bidder to list any other Consumables or Accessories (Not listed above) for the model on offer below

## Pulmonary Function Unit
RT2-02-13 | PULMONARY FUNCTION UNIT
1 | DESCRIPTION
1.1 | Device used in diagnostic procedures to detect lung disease.
1a | Product Offered:
1a.1 | Make:
1a.2 | Model:
1a.3 | Initial Manufacturing year of Model:
2 | SPECIFICATIONS
2.1 | The unit offered must be a portable pulmonary function unit, which must provide measuring as follows:
2.1.1 | Slow vital capacity (Vt; ERV; Vcin; Vcex and others. Forced spirometry (inspiratory and expiratory flow-volume curves). Maximal voluntary ventilation (MVV).
3 | The structure of the portable unit being offered must be as follows:
3.1 | The unit must be portable, compact and lightweight for easy transport.
3.2 | Must be easy to clean, disassemble and sterilize all parts in direct contact with patient.
3.3 | The display of the unit must be backlit LCD type, minimum resolution of 320 x 200 pixels.
3.4 | Must have an up to date communication and printer port. Bidders to state type output device
3.5 | The unit must be AC / DC with built-in battery backup, sufficient for a minimum of two (2) hours continuous use.
4 | The signal processing on the unit offered must be as follows:
4.1 | Highly linear, open circuit, differential pressure pneumotachograph, or a bi-directional digital turbine flowmeter.
4.2 | Flow: Range: 0-18 L/sec (bi-directional).
4.3 | Accuracy: + 3% over the full scale.
4.4 | Resolution: 10ml/sec.
4.5 | Volume: Range: 0-10 litre.
4.6 | Accuracy: 3% or 50ml, whichever is greater, over the full range.
4.7 | Resolution: 1ml.
4.8 | Resistance: The resistance of the pneumotachograph and tubing should be less than 1.5cm H2O/L/sec at 12L/sec.
4.9 | Time: The unit should be capable of reading time for at least 15 seconds with an accuracy of 1%.
5 | The data management on the unit offered must be as follows:
5.1 | Calculation of user selectable universal predicted standards. Graphical and numerical display during measurement and in measurement reports is essential.
5.2 | Very fast access to all patient data, including previously performed “pre” data for simultaneous pre and post testing.
5.3 | Backup of at least 10 sets of patient data, together with pre and post graphics in internal database.
5.4 | Software must facilitate adjustments to measurement routines, selectable protocols with a choice of parameters, predicted values and graphics in a pre / post comparison.
5.5 | Saving of measured values and curves in the database.
5.6 | At least three manoevures each for pre and post bronchodilator studies must be obtainable.
5.7 | Automatic saving of the best curve is essential, but must be user selectable as well.
6 | CALIBRATION:
6.1 | Computer assisted calibration of pneumotachograph and automatic Body Temperature and Pressure Saturated (BTPS) correction by means of a 2 or 3 litre calibration syringe.
7 | The parameters offered on the quoted unit must be as follows:
7.1 | Slow spirometry: Vt; BF; MV; ERV; ERV; Vcin; Vcex and others. Flow-volume curve: FVC; FEV1; FEV0,5; FEF2; FEV1% / Fvmax; PEF; MEF95; MEF50; MEF25; MMEF; REF75-85; PIF; MIF50; MEF50% / MIF50 and others. Maximal voluntary ventilation: MVV; RR; Vt and others.
8 | Reports must be provided and available on the unit offered as follows:
8.1 | Selection of defined protocols must be available with individual selection of parameters, predicted values and graphics in a pre / post comparison with output of measuring reports on printer or display. It must be possible to overlay pre and post curves in the printout and on screen.
8.2 | Re-reading of test results for output together with possibility to edit reports.
8.3 | Built-in, high resolution thermal printer, at least 15cm wide.
9 | STANDARD ACCESSORIES
9.1 | spirometry sensor
9.2 | calibration syringe
9.3 | Weather station
9.4 | mouth pieces
9.5 | paper
10 | Conformity Compliance (Please attach certificate)
10.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
10.2 | OEM must comply and certified on ISO 9001 quality standards, attach proof of compliance
10.2.1 | Please provide unique ref number of the ISO 9001 ceriticate:
10.3 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
10.3.1 | Please provide unique ref number of the ISO 13485 ceriticate:
10.4 | Model quoted for must be EC certified. Attach a copy of certification
10.4.1 | Please provide unique ref number of the EC ceriticate:
10.5 | Hazardous Substance Act: If this type of equipment / apparatus appears on the schedule of Hazardous substances issued by the Directorate: Health Technology of the Department of Health, a license in terms of the Act on Hazardous Substances (Act 15/1973) must be submitted with the bid / quotation documents. The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
10.6 | All electrical/electronic medical equipment must be licensed by Radiation Control, where applicable a copy of the license must be submitted The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
10.6.1 | Radiation Control License number:
10.7 | The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade (rubber to with 2 screws) type and it must be a minimum length of three (3) meters. NB the mains cable of the unit tendered for must be SABS colour coded.
10.8 | The equipment quoted for must be protected against electromagnetic interference IEC 60601-1-2
11 | COMPREHENSIVE MAINTENANCE
11.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) must be included in the bid price
11.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS, Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
11.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
11.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS, labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
11.3 | The bidder to provide the OEM recommended service interval of the unit:
11.4 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
12 | General
12.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
12.2 | The unit on offer must be the latest technology.
PULMONARY FUNCTION UNIT ACCESSORIES
RT2-02-13a | spirometry sensor
RT2-02-13a | calibration syringe
RT2-02-13a | Weather station
RT2-02-13a | mouth pieces
RT2-02-13a | paper
Bidder to list any other Consumables or Accessories (Not listed above) for the model on offer below

## O2 Flowmeter, Single
RT2-02-14 | Oxygen Flowmeter, single
Description
1 | Product Offered:
1.1 | Make:
1.2 | Model:
1.3 | Initial Manufacturing year of Model:
2 | Technical Specification
2.1 | Flowmeter must be of the pressure compensated type and fit directly into the Oxygen outlet.
2.2 | The flowmeter must have a stainless steel float in a calibrated flow tube
2.3 | Must have a flow accuracy within 10%.
2.4 | The range of 0 – 15 l/m of the flowmeter must be permanent and easy to read.
2.5 | Connector must be Institution specific, 1 flow meter with 1 outlet
2.6 | The supplier must confirm the connection before delivery.
2.7 | Must have a chrome-plated brass body and adjustment knob
2.8 | The flowmeter must operate accurately on a 4 bar pipeline source.
2.9 | The flowmeter hood must be made of crack–resistant polycarbonate.
2.10 | Each flowmeter must be supplied with oxygen outlet ‘tail’ fitting.
2.11 | The flow meter must be easy to service and maintain
2.12 | The unit must be handed over in full operating order
3 | Conformity Compliance (Please attach certificate)
3.1 | International OEM must comply and certified on ISO 9001 quality standards, attach proof of compliance
3.1.1 | Please provide unique ref number of the ISO 9001 ceriticate:
3.2 | International OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
3.2.1 | Please provide unique ref number of the ISO 13485 ceriticate:
3.3 | Model quoted for must be EC or SABS certified. Attach a copy of certification
3.3.1 | Please provide unique ref number of the EC or SABS ceriticate:
4 | COMPREHENSIVE MAINTENANCE
4.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) must be included in the bid price
4.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries, Flow sensors and Oxygen sensor), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
4.2 | The bidder to provide the OEM recommended service interval of the unit:
4.3 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
Oxygen Flowmeter, single - Accessories and Consumables
RT2-02-14a | Outer/flow tube (Pk of 10)
RT2-02-14b | Inner tube (Pk of 10)
RT2-02-14c | Flow ball (pk of 10)
RT2-02-14d | Tail piece (Xmas tree-plastic) Pk of 100
RT2-02-14e | Flow Knob (Pk of 10)
Bidder to list any other Consumables or Accessories (Not listed above) for the model on offer below

## O2 Flowmeter, Double
RT2-02-15 | Oxygen Flowmeter, double
Description
1 | Product Offered:
1.1 | Make:
1.2 | Model:
1.3 | Initial Manufacturing year of Model:
2 | Technical Specification
2.1 | Flowmeter must be of the pressure compensated type and fit directly into the Oxygen outlet.
2.2 | The flowmeter must have a stainless steel float in a calibrated flow tube
2.3 | Must have a flow accuracy within 10%.
2.4 | The range of 0 – 15 l/m of the flowmeter must be permanent and easy to read.
2.5 | Connector must be Institution specific, 1 flow meter with 1 outlet
2.6 | The supplier must confirm the connection before delivery.
2.7 | Must have a chrome-plated brass body and adjustment knob
2.8 | The flowmeter must operate accurately on a 4 bar pipeline source.
2.9 | The flowmeter hood must be made of crack–resistant polycarbonate.
2.10 | Each flowmeter must be supplied with oxygen outlet ‘tail’ fitting.
2.11 | The flow meter must be easy to service and maintain
2.12 | The unit must be handed over in full operating order
3 | Conformity Compliance (Please attach certificate)
3.1 | International OEM must comply and certified on ISO 9001 quality standards, attach proof of compliance
3.1.1 | Please provide unique ref number of the ISO 9001 ceriticate:
3.2 | International OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
3.2.1 | Please provide unique ref number of the ISO 13485 ceriticate:
3.3 | Model quoted for must be EC or SABS certified. Attach a copy of certification
3.3.1 | Please provide unique ref number of the EC or SABS ceriticate:
4 | COMPREHENSIVE MAINTENANCE
4.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) must be included in the bid price
4.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries, Flow sensors and Oxygen sensor), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
4.2 | The bidder to provide the OEM recommended service interval of the unit:
4.3 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
Oxygen Flowmeter, Double- Accessories and Consumables
RT2-02-15a | Outer/flow tube (Pk of 10)
RT2-02-15b | Inner tube (Pk of 10)
RT2-02-15c | Flow ball (pk of 10)
RT2-02-15d | Tail piece (Xmas tree-plastic) Pk of 100
RT2-02-15e | Flow Knob (Pk of 10)
Bidder to list any other Consumables or Accessories (Not listed above) for the model on offer below

## Nebulizer, Ultrasonic
RT2-02-16 | NEBULIZER, ULTRASONIC
1 | DESCRIPTION
1.1 | a humidifier using high-frequency electricity to power a transducer that vibrates and divides liquid into particles in its nebulizing chamber; used in inhalation therapy.
1a | Product Offered:
1a.1 | Make:
1a.2 | Model:
1a.3 | Initial Manufacturing year of Model:
1a.4 | Radiation Control License number:
2 | TECHNICAL SPECIFICATION
2.1 | This unit must be the ultrasonic oscillator type.
2.2 | The unit must operate off 220V-240V 50hz A.C. single phase power supply.
2.3 | Most of the droplets must be within the 1 to 6 micron range.
2.4 | Output volumes must be a minimum of 2,5ml of aerosol per minute at the maximum output setting.
2.5 | The aerosol flow must be adjustable.
2.6 | The unit must have a self-contained air supply delivering of upto 20l/min.
2.7 | The nebulizer must be manufactured from rust free or plastic material.
2.8 | The unit must have a feed system in the reservoir for continuous operation.
2.9 | The nebulizer must not resist any patient coughing.
2.1 | The operating frequency must be between 1,6 and 1,8MHz.
2.11 | The nebulizer must not weigh more than 5Kg.
2.12 | Circuit hoses and chamber must come complete with the nebulizer.
3 | STANDARD ACCESSORIES
3.1 | Adult and Paeds Mask
3.2 | Mouth piece
3.3 | Hose
3.4 | Medication cups (pk of 5)
4 | Conformity Compliance (Please attach certificate)
4.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
4.2 | OEM must comply and certified on ISO 9001 quality standards, attach proof of compliance
4.2.1 | Please provide unique ref number of the ISO 9001 ceriticate:
4.3 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
4.3.1 | Please provide unique ref number of the ISO 13485 ceriticate:
4.4 | Model quoted for must be EC certified. Attach a copy of certification
4.4.1 | Please provide unique ref number of the EC ceriticate:
4.5 | Hazardous Substance Act: If this type of equipment / apparatus appears on the schedule of Hazardous substances issued by the Directorate: Health Technology of the Department of Health, a license in terms of the Act on Hazardous Substances (Act 15/1973) must be submitted with the bid / quotation documents. The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
4.6 | All electrical/electronic medical equipment must be licensed by Radiation Control, where applicable a copy of the license must be submitted The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
4.6.1 | Radiation Control License number:
4.7 | The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade (rubber to with 2 screws) type and it must be a minimum length of three (3) meters. NB the mains cable of the unit tendered for must be SABS colour coded.
4.8 | The equipment quoted for must be protected against electromagnetic interference IEC 60601-1-2
5 | COMPREHENSIVE MAINTENANCE
5.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) must be included in the bid price
5.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS, Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
5.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
5.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS, labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
5.3 | The bidder to provide the OEM recommended service interval of the unit:
5.4 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
6 | General
6.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
6.2 | The unit on offer must be the latest technology.
NEBULIZER -ACCESSORIES and CONSUMABLES
RT2-02-16a | Adult and Paeds Mask
RT2-02-16b | Mouth piece
RT2-02-16c | Hose
RT2-02-16d | Medication cups (pk of 5)
Bidder to list any other Consumables or Accessories (Not listed above) for the model on offer below

## Category-3 Paeds
PAEDIATRIC CARE EQUIPMENT
# | Item number | Item Description
3 | RT2-03-03 | Apnea Monitor, Airflow sensor
4 | RT2-03-04 | Bilirubin Meter, Xenon
7 | RT2-03-07 | ICU Crib
8 | RT2-03-08 | Incubator, ICU
9 | RT2-03-09 | Incubator, Standard
10 | RT2-03-10 | Incubator, Transport
11 | RT2-03-11 | Infant Radiant Warmer
12 | RT2-03-12 | Phototherapy Unit, LED
13 | RT2-03-13 | Phototherapy Unit, Blanket, LED
15 | RT2-03-15 | Ultrasound, Neonatal

## Apnea Monitor, Airflow Sensor
RT2-03-03 | APNEA MONITOR (Airflow/CO2 sensor)
1 | Description
1.1 | A monitoring device, for continuous respiratory surveillance in neonates and infants, using an airflow/CO2 sensor.
2 | Product Offered:
2.1 | Make:
2.2 | Model:
2.3 | Initial Manufacturing year of Model:
2.4 | Radiation Control License number:
3 | Technical Specification
3.1 | The unit must be a monitoring device, for continuous respiratory surveillance in neonates and infants.
3.2 | The monitor must trigger both audible and visible alarms after a preselected delay of 10, 15 or 20 sec following the onset of apnea.
3.3 | The audible alarm must be audible at a distance of at least 10 metres
3.4 | The monitor must indicate every breathing motion detected.
3.5 | Bidder must describe the working principle of the breathing sensor/transducer.
3.6 | Digital display for breath rate of the infant. The display must be clearly visible under all lighting conditions
3.7 | The unit must operate from 220V, 50 Hz single phase supply.
3.8 | Mains cable at least three metres long with SABS approved 15A three prong plug.
3.9 | The unit must operate from an internally mounted rechargeable battery in case of mains failure.
3.10 | Battery charger must be incorporated
3.11 | Battery must be supplied
3.12 | Low battery: audible and visual alarms
3.13 | The sensor/transducer must be safely isolated from the 220V, 50 Hz supply.
3.14 | The unit must be well sealed from liquids
3.15 | The unit must be handed over in full operating order
4 | Standard Accessories
4.1 | Sensor
4.2 | Rechargeable battery
5 | Conformity Compliance (Please attach certificate)
5.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
5.2 | OEM must comply and certified on ISO 9001 quality standards, attach proof of compliance
5.2.1 | Please provide unique ref number of the ISO 9001 ceriticate:
5.3 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
5.3.1 | Please provide unique ref number of the ISO 13485 ceriticate:
5.4 | Model quoted for must be EC certified. Attach a copy of certification
5.4.1 | Please provide unique ref number of the EC ceriticate:
5.5 | Hazardous Substance Act: If this type of equipment / apparatus appears on the schedule of Hazardous substances issued by the Directorate: Health Technology of the Department of Health, a license in terms of the Act on Hazardous Substances (Act 15/1973) must be submitted with the bid / quotation documents. The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.7 | All electrical/electronic medical equipment must be licensed by Radiation Control, where applicable a copy of the license must be submitted The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.7.1 | Radiation Control License number:
5.8 | The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade (rubber to with 2 screws) type and it must be a minimum length of three (3) meters. NB the mains cable of the unit tendered for must be SABS colour coded.
5.8 | The equipment quoted for must be protected against electromagnetic interference IEC 60601-1-2
6 | COMPREHENSIVE MAINTENANCE
6.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) must be included in the bid price
6.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
6.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
6.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
6.3 | The bidder to provide the OEM recommended service interval of the unit:
6.4 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
7 | General
7.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
7.2 | The unit on offer must be the latest technology.
APNEA MONITOR (Airflow sensor) ACCESSORIES & CONSUMABLES
RT2-03-03a | Sensor
RT2-03-03b | Rechargeable battery
Bidder to list any other Consumables or Accessories (Not listed above) for the model on offer below

## Bilirubin Meter, Xenon
RT2-03-04 | Billirubin meter Xenon
1 | Description
1.1 | Non-invasive transcutaneous bilirubin measuring device to assess the risk of neonatal hyperbilirubinemia using Xenon technology
2 | Product Offered:
2.1 | Make:
2.2 | Model:
2.3 | Initial Manufacturing year of Model:
3 | Technical Specification
3.1 | Effective regardless of skin pigmentation
3.2 | Suitable for neonates of gestational age > 24 weeks
3.3 | Suitable for neonates of post-natal age 0 - 20 days
3.4 | The unit must be handheld
3.5 | The unit must be easy to operate
3.6 | Measurement by placing the unit on the skin of the newborn and pressing it lightly
3.7 | Measurements must be done by means of a white light measured at specific wavelengths between 300 and 900 nm
3.8 | The measuring method must be based on the determination of the yellowness of the subcutaneous tissue by using two optical paths to measure optical density difference at two wavelengths
3.9 | The light source: Xenon
3.10 | The unit must take the measurement, process and display the information without need of any peripheral devices.
3.11 | Must not be sensitive to user-applied pressure or angle of device
3.12 | Must be able to display result of a single measurement or the average of a number of measurements
3.13 | Must not be sensitive to motion artifacts (or movements) of the baby
3.14 | Can be used with or without consumable tips
3.15 | Must not require routine calibration
3.16 | The light source life: > 150 000 measurements
3.17 | The unit must be capable of being used before, during and post phototherapy.
3.18 | The unit must incorporate a LCD display to display results, operating parameters, errors, date and time.
3.19 | The unit must report results in mg/dl or µmol/L
3.20 | The total serum billirubin range must be from 0 to 20 mg/dl or 0 to 340 µmol/L.
3.21 | Repeatability must be ±1.5mg/dl or 25 µmol/L.
3.22 | Accuracy during and post phototherapy must be ±1.75mg/dl or 30µmol/L.
3.23 | The following error signals must be displayed on the unit.
3.23.1 | Unacceptable calibration
3.23.2 | Unacceptable patient measurement
3.23.3 | Too much movement during measurement
3.23.4 | Measurement is out of clinically proven range
3.23.5 | Light bulb faulty
3.23.6 | System error
3.23.7 | Low battery
3.24 | The unit must incorporate a rechargeable battery, and must operate only from the battery power.
3.25 | Battery must be supplied
3.26 | The battery must allow at least 100 measurements with a full charge.
3.27 | Expected battery life: at least 1 year
3.28 | A battery charger, operating from 220 Volt, 50Hz mains power must be supplied
3.29 | The battery charger must have a battery overcharge protection circuit.
3.30 | The unit must indicate that the battery is being charged when the unit is plugged into mains power.
3.31 | Must indicate either the charge or the number of tests left in the battery
3.32 | A protective carry case must be supplied
3.33 | Physical dimensions: approximately 50mm(W ) X 150mm(H) X 35mm(D)
3.34 | Weight: approximately 150 grams
3.35 | The unit must be handed over in full operating order
4 | Standard Accessories
4.1 | Protective carry case
4.2 | Rechargeable battery
5 | Conformity Compliance (Please attach certificate)
5.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
5.2 | OEM must comply and certified on ISO 9001 quality standards, attach proof of compliance
5.2.1 | Please provide unique ref number of the ISO 9001 ceriticate:
5.3 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
5.3.1 | Please provide unique ref number of the ISO 13485 ceriticate:
5.4 | Model quoted for must be EC certified. Attach a copy of certification
5.4.1 | Please provide unique ref number of the EC ceriticate:
5.5 | The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade (rubber to with 2 screws) type and it must be a minimum length of three (3) meters. NB the mains cable of the unit tendered for must be SABS colour coded.
5.6 | The equipment quoted for must be protected against electromagnetic interference IEC 60601-1-2
6 | COMPREHENSIVE MAINTENANCE
6.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) must be included in the bid price
6.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
6.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
6.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
6.3 | The bidder to provide the OEM recommended service interval of the unit:
6.4 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
7 | General
7.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
7.2 | The unit on offer must be the latest technology.

## ICU Crib
RT2-03-07 | ICU Crib
1 | Description
1.1 | A crib used to provide patient support and relief in the intesive care unit.
2 | Product Offered:
2.1 | Make:
2.2 | Model:
2.3 | Initial Manufacturing year of Model:
2.4 | Radiation Control License number:
3 | Technical Specification
3.1 | The unit must offer long-term intensive care of a neonate
3.2 | The ICU crib must function off 220 V, 50 Hz single phase power supply.
3.3 | Platform and mattress:
3.3.1 | Adjustable tilt to Trendellenburg and reverse Trendellenburg positions
3.3.2 | Locking mechanism to keep the mattress platform locked in desired positions.
3.3.3 | Must accommodate X-ray film cassettes beneath the mattress.
3.3.4 | The mattress must be waterproof and must be X-ray transparent.
3.3.5 | Mattress size: approximately 500 x 650 mm.
3.3.6 | Mattress thickness: at least 25 mm
3.3.7 | Mattress must support patient weight up to 4,5 kg.
3.3.8 | All four sides of the platform must be fitted with transparent guards/panels.
3.3.9 | The front and two side panels must be hinged.
3.3.10 | Facility for holding patient cables, tubes, drains etc to prevent then being trapped or dislodged even when sides are opened
3.4 | Heater and Controls
3.4.1 | Radiant heater housed in an overhead canopy, with internal reflector, mounted on the rear mainframe.
3.4.2 | Indicator to show heater is switched ON
3.4.3 | The overhead canopy must swivel left and right and must lock in the desired position.
3.4.4 | Microprocessor control with automatic self-test when switched on.
3.4.5 | The heat controller must provide the following modes for temperature control:
3.4.5.1 | Pre-heat
3.4.5.2 | Manual heat control: adjustable heater power from 0 - 100%.
3.4.5.3 | Skin temperature control: servo-control using a skin temperature probe
3.4.6 | A reusable skin temperature probe must be supplied
3.4.7 | Skin temperature probe cable length: 1m
3.4.8 | The skin temperature set point must be selectable in the range of 34 °C to 38 °C, in 0,1 °C intervals.
3.4.9 | The skin temperature measurement range: 30 °C to 40 °C.
3.4.10 | Skin temperature measurement accuracy: 0,5 °C
3.4.11 | If a malfunction occurs affecting safety or performance the heater must be automatically turned off
3.5 | Displays:
3.5.1 | Measured temperature: digital display
3.5.2 | Heater power
3.5.3 | Indicator of the selected mode (manual/control)
3.6 | Alarms
3.6.1 | Audible and visual alarms when the 220 V power is lost during operation.
3.6.2 | Audible and visual alarms if the skin temperature deviates by ± 1 °C from the set temperature.
3.6.3 | Audible and visual alarms if the skin temperature exceeds 39°C
3.6.4 | Audible and visual alarms if skin temperature probe is disconnected or fails during control mode.
3.6.5 | Should the skin temperature probe become disconnected or fail, the unit must alarm (alarm must not be silenced more than 2 munites) and automatically switch to manual mode.
3.6.6 | Alarm test facility
3.6.7 | Temporary mute of audible alarms.
3.6.8 | Apgar timer, with audible alarm, with digital display: 1, 5 and 10 minutes elapsed time.
3.6.9 | Apgar timer with on/off switch and reset button
3.7 | Air/Oxygen Blender
3.7.1 | The unit must have intergrated Air/Oxygen blender
3.7.2 | Deliver Oxygen Range 21 % to 100 %
3.7.3 | Flow Range 2 to 15 lpm.
3.8 | Other
3.8.1 | The overhead canopy must be provided with an examination lamp, with on/off switch.
3.8.2 | The ICU crib must be supplied with an oxygen cylinder holder.
3.8.3 | The ICU crib must have a drawer
3.8.4 | The minimum acceptable height of the heater canopy must be 1.85 metres above floor
3.8.5 | The minimum acceptable height of the mattress platform must 900 mm above floor.
3.8.6 | The mains cable of the unit must be at least three metres long and be fitted with an SABS approved 15A three pin plug
3.8.7 | Instrument tray, approximately 200 x 300 mm, 2 kg capacity
3.8.8 | Swivel mounted monitor shelf, approximately 300 x 400 mm, 10 kg capacity
3.8.9 | IV pole
3.8.10 | Accessory rail must be provided on the rear mainframe.
3.8.11 | The unit must be mobile, but stable during movement
3.9 | Resuscitator
3.9.1 | The crib must be equipped with an infant resuscitator integrated into the crib
3.9.2 | Unit must be gas powered via the wall oxygen supply
3.9.3 | Oxygen flow regulator: 5 - 15 l/min
3.9.4 | Adjustable Peak Inspiratory Pressure (PIP): 2 - 75 cm H2O
3.9.5 | Adjustable Positive End Expiratory Pressure (PEEP): 1 - 25 cm H2O
3.9.6 | Resuscitator must provide controlled pressure settings by means of a pressure limiting device
3.9.7 | Pressure gauge: -10 to 80 cm H2O
3.9.8 | Maximum pressure relief: 5 - 70cmH²O @ 8lpm.
3.9.9 | The unit must be handed over in full operating order
3.10 | A starter-pack of consumables must be supplied with the unit.
4 | The following accessories must be supplied with the unit:
4.1 | Temperature probe
4.2 | Disposable hydrophobic bacterial filters, box of 10
4.3 | IV pole
4.4 | Shelf for monitor
4.5 | Instrument tray
4.6 | Battery
5 | Conformity Compliance (Please attach certificate)
5.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
5.2 | OEM must comply and certified on ISO 9001 quality standards, attach proof of compliance
5.2.1 | Please provide unique ref number of the ISO 9001 ceriticate:
5.3 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
5.3.1 | Please provide unique ref number of the ISO 13485 ceriticate:
5.4 | Model quoted for must be EC certified. Attach a copy of certification
5.4.1 | Please provide unique ref number of the EC ceriticate:
5.5 | Hazardous Substance Act: If this type of equipment / apparatus appears on the schedule of Hazardous substances issued by the Directorate: Health Technology of the Department of Health, a license in terms of the Act on Hazardous Substances (Act 15/1973) must be submitted with the bid / quotation documents. The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.7 | All electrical/electronic medical equipment must be licensed by Radiation Control, where applicable a copy of the license must be submitted The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.7.1 | Radiation Control License number:
5.8 | The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade (rubber to with 2 screws) type and it must be a minimum length of three (3) meters. NB the mains cable of the unit tendered for must be SABS colour coded.
5.8 | The equipment quoted for must be protected against electromagnetic interference IEC 60601-1-2
5.9 | The unit must comply with electrical safety and performance standard such as IEC 60601-2-21, attached certification
5.9.1 | Please provide unique ref number of the IEC ceriticate:
6 | COMPREHENSIVE MAINTENANCE
6.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) must be included in the bid price
6.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
6.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
6.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
6.3 | The bidder to provide the OEM recommended service interval of the unit:
6.4 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
7 | General
7.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
7.2 | The unit on offer must be the latest technology.
RT2-03-07 | ICU Crib
RT2-03-07a | Temperature probe
RT2-03-07c | IV pole
RT2-03-07d | Shelf for monitor
RT2-03-07e | Instrument tray
RT2-03-07f | Battery
RT2-03-07g | Air/oxygen blender (UPGRADE)
RT2-03-07h | Built-in scale (UPGRADE)
RT2-03-07i | LED Phototherapy (UPGRADE)
RT2-03-07j | Pulse Oxymeter (UPGRADE)
RT2-03-07k | Disposable covers for skin temperature probe (box of 10)
RT2-03-07l | Intergrated suction (UPGRADE)
Bidder to list any other Consumables or Accessories (Not listed above) for the model on offer below

## Incubator, ICU
RT2-03-08 | Incubator ICU
1 | Description
1.1 | A unit that provides a closed, controlled environment that warms an infant by circulating heated air around the infant´s skin, and is to be used in the neonatal intensive care unit
2 | Product Offered:
2.1 | Make:
2.2 | Model:
2.3 | Initial Manufacturing year of Model:
2.4 | Radiation Control License number:
3 | Technical Specification
3.1 | The control module must provide an even distribution of heat and humidity within the hood
3.2 | The unit must perform an automatic self test when switched on.
3.3 | The incubator frame and chassis must be robust
3.4 | Chassis must not detach from the frame
3.5 | Must be fitted with a robust transparent double hood. Approximate dimensions (LxWxH): 900 x 450 x 430 mm
3.6 | The front of the hood must have a full width transparent access panel, to allow the mattress tray to be extended outside the hood enclosure for examination and treatment of the infant without raising the complete hood.
3.7 | When the mattress tray is extended outside the hood enclosure, it must remain attached to the incubator and support the weight of the baby.
3.8 | The hood must be hinged to allow easy access for bedding the infant.
3.9 | The hood must have a safety latch to prevent accidental opening
3.10 | The hood must remain in a stable position when opened
3.11 | The hood must have at least four arm openings with robust latches that are elbow operated. Two of these openings must be on the front and two at the rear of the hood
3.12 | The hood must also have two iris openings, on the left and right hand sides
3.13 | The hood must have 2 small self-sealing ports, for IV tubes etc.
3.14 | The hood must seal well onto the chassis frame to avoid any form of heat loss.
3.15 | It must be possible to easily manipulate the bed into the Trendellenberg and anti-Trendellenberg positions.
3.16 | The unit must have an IV pole and a monitor stand.
3.17 | The monitor stand must support a mass of at least 10 kg
3.18 | The incubator must be easy to clean, disinfect, and/or sterilize
3.19 | Switches and controls must be protected against accidental setting changes
3.20 | TEMPERATURE CONTROL:
3.20.1 | The control unit must be microprocessor-based.
3.20.2 | The control module must have servo heat control mode
3.20.3 | The unit must not take more than ± 30 minutes from power-on to reach the desired temperature setting.
3.20.4 | User selectable air temperature and skin temperatures.
3.20.5 | Air temperature must be selectable from ± 25°C to ± 37°C
3.20.6 | Air temperature must be selectable to 39°C with override confirmation
3.20.7 | Skin temperature must be selectable from ± 34°C to ± 37°C.
3.20.8 | Skin temperature must be selectable to 39°C with override confirmation
3.20.9 | Temperature variation less than +/- 0.2°C
3.20.10 | A temperature sensor must be available in the hood for temperature control and monitoring.
3.20.11 | In addition to the primary sensor for controlling the temperature, the unit must have a secondary fail-safe sensor which will act as a safety back-up and activate audible and visible alarms and shut off power to the heater if the temperature exceeds 40oC.
3.20.12 | The supply to the heater may only be restored after a manual reset and if the incubator’s secondary temperature monitoring device senses that the internal temperature has decreased to a safe level.
3.20.13 | At switch-on the incubator must automatically heat to 34°C +/- 1°C
3.20.14 | When the hood’s front access panel is opened, it will have a minimal effect on the set temperature within the incubator.
3.20.15 | Reusable skin temperature probe must be supplied
3.21 | DISPLAY
3.21.1 | Display for the following:
3.21.1a | Skin temperature.
3.21.1b | Air temperature.
3.21.1c | Set temperature
3.21.1d | Relative humidity
3.21.1e | Details of all alarms
3.21.2 | Display must be large and clear enough to be viewed from three metres under all lighting conditions.
3.22 | ALARMS
3.22.1 | Default alarms must be enabled when the incubator is turned on
3.22.2 | Audible and visual alarms with heater power cut-off in the event of:
3.22.2a | over temperature
3.22.2b | failure of air circulation
3.23 | Audible and visual alarms in the event of:
3.23.1 | skin temperature differs from set temperature by + 3°C
3.23.2 | air temperature differs from set temperature by + 3°C
3.23.3 | power failure
3.23.4 | heater failure
3.23.5 | skin or air probe disconnection or failure.
3.23.6 | The audible alarm must have a mute option. Visual alarm may not be reset until the cause of the alarm has been rectified
3.23.7 | Test facility for audible and visual alarms.
3.23.8 | Low water alarm
3.23.9 | The unit must be supplied with an oxygen regulator and dial-in flowmeter
3.24 | GENERAL
3.24.1 | Exhaled air must continuously be removed to prevent carbon dioxide build-up inside the hood.
3.24.2 | Fresh air drawn to the hood must be filtered by an easily replaceable bacterial filter
3.24.3 | The mattress must be approximately 630 x 350 x 30 mm and must be very durable, comfortable and have a waterproof cover.
3.24.4 | The sound level within the baby compartment must not exceed 50 dB(A) except if an audible alarm is sounding, then the sound level within the baby compartment must not exceed 80dB(A).
3.24.5 | The air velocity above the mattress must not exceed 0.35 m/s.
3.24.6 | The incubator must operate off 220V, 50Hz single phase supply
3.24.7 | Mains cable at least three metres long with SABS approved 15A three prong plug.
3.24.8 | All metal parts of the incubator must be well earthed.
3.24.9 | The incubator’s control module must be well designed to guard against entry of liquids
3.24.10 | All parts on the incubator and trolley must be rust resistant
3.24.11 | X-Ray tray incorporated without need to open incubator
3.24.12 | Rail for accessories on head side
3.24.13 | Switches and controls must be protected against accidental setting changes
3.24.14 | Labels and markings must be durable enough to withstand routine cleaning and normal wear
3.24.15 | Appropriate warning legends must be provided on the unit.
3.24.16 | The unit must be simple to use and maintain.
3.24.17 | The unit must be designed for easy access to serviceable parts
3.24.18 | The incubator must have no sharp edges.
3.24.19 | The unit must be mounted on a robust trolley
3.24.20 | Trolley must have 100 mm diameter (minimum) antistatic swivel castors.
3.24.21 | At least two of the castors must have brakes.
3.24.22 | Handle (for pushing)
3.24.23 | Holder for a standard 1 kg oxygen cylinder must be mounted to the trolley.
3.24.24 | Switches, knobs, and other controls must be designed for conditions of heavy use.
3.24.25 | Storage space (drawers)
3.25 | The unit must be handed over in full operating order
3.26 | A starter-pack of consumables must be supplied with the unit.
4 | The following accessories must be supplied with the unit:
4.1 | Bacterial filter (X 4)
4.4 | Skin temperature probe (reusable)
4.5 | Covers for temperature probe
4.6 | Internal battery for alarms
5 | Conformity Compliance (Please attach valid certificate with ref number and expiry date)
5.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
5.2 | OEM must comply and certified on ISO 9001 quality standards, attach proof of compliance
5.2.1 | Please provide unique ref number of the ISO 9001 ceriticate:
5.3 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
5.3.1 | Please provide unique ref number of the ISO 13485 ceriticate:
5.4 | Model quoted for must be EC certified. Attach a copy of certification
5.4.1 | Please provide unique ref number of the EC ceriticate:
5.5 | Hazardous Substance Act: If this type of equipment / apparatus appears on the schedule of Hazardous substances issued by the Directorate: Health Technology of the Department of Health, a license in terms of the Act on Hazardous Substances (Act 15/1973) must be submitted with the bid / quotation documents. The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.6 | All electrical/electronic medical equipment must be licensed by Radiation Control, where applicable a copy of the license must be submitted The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.6.1 | Radiation Control License number:
5.7 | The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade (rubber to with 2 screws) type and it must be a minimum length of three (3) meters. NB the mains cable of the unit tendered for must be SABS colour coded.
5.8 | The equipment quoted for must be protected against electromagnetic interference IEC 60601-1-2
5.9 | The unit must comply with electrical safety and performance standard such as IEC 60601-2-19, attached certification
5.9.1 | Please provide unique ref number of the IEC ceriticate:
6 | COMPREHENSIVE MAINTENANCE
6.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) must be included in the bid price
6.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
6.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
6.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
6.3 | The bidder to provide the OEM recommended service interval of the unit:
6.4 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
7 | General
7.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
7.2 | The unit on offer must be the latest technology.
Incubator, ICU Accessories
RT2-03-08a | Bacterial filter (Pk of 4)
RT2-03-08b | Iris port sleeves (Pk of 10)
RT2-03-08c | Oval port sleeves (Pk of 10)
RT2-03-08d | Skin temperature probe (reusable)
RT2-03-08e | Covers for temperature probe (Pk of 25)
RT2-03-08f | Internal battery for alarms
RT2-03-08g | Scale (UPGRADE)
RT2-03-08h | Observation lamp (UPGRADE)
Bidder to list any other Consumables or Accessories (Not listed above) for the model on offer below

## Incubator, Standard
RT2-03-09 | Incubator (Standard)
1 | Description
1.1 | A unit that provides a closed, controlled environment that warms an infant by circulating heated air around the infant´s skin.
2 | Product Offered:
2.1 | Make:
2.2 | Model:
2.3 | Initial Manufacturing year of Model:
2.4 | Radiation Control License number:
3 | Technical Specification
3.1 | The control module must provide an even distribution of heat and humidity within the hood
3.2 | The unit must perform an automatic self test when switched on.
3.3 | The incubator frame and chassis must be robust
3.4 | Chassis must not detach from the frame
3.5 | Must be fitted with a robust transparent double hood. Approximate dimensions (LxWxH): 900 x 450 x 430 mm
3.6 | The front of the hood must have a full width transparent access panel, to allow the mattress tray to be extended outside the hood enclosure for examination and treatment of the infant without raising the complete hood.
3.7 | When the mattress tray is extended outside the hood enclosure, it must remain attached to the incubator and support the weight of the baby.
3.8 | The hood must be hinged to allow easy access for bedding the infant.
3.9 | The hood must have a safety latch to prevent accidental opening
3.10 | The hood must remain in a stable position when opened
3.11 | The hood must have at least four arm openings with robust latches that are elbow operated. Two of these openings must be on the front and two at the rear of the hood
3.12 | The hood must also have two iris openings, on the left and right hand sides
3.13 | The hood must have at least 4 small self-sealing ports, for IV tubes etc.
3.14 | The hood must seal well onto the chassis frame to avoid heat loss.
3.15 | Easy manipulation of the bed into Trendellenberg and anti-Trendellenberg positions.
3.16 | The unit must have an IV pole
3.17 | The incubator must be easy to clean, disinfect, and/or sterilize
3.18 | Switches and controls must be protected against accidental setting changes
3.19 | TEMPERATURE CONTROL:
3.19.1 | The control unit must be microprocessor-based.
3.19.2 | The control module must have servo heat control modes
3.19.3 | The unit must not take more than 30 minutes from power-on to reach the desired temperature setting.
3.19.4 | User selectable air temperature and skin temperatures.
3.19.5 | Air temperature must be selectable from ± 25°C to ± 37°C
3.19.6 | Air temperature must be selectable to 39°C with override confirmation
3.19.7 | Skin temperature must be selectable from 34°C to 37°C.
3.19.8 | Skin temperature must be selectable to 39°C with override confirmation
3.19.9 | Temperature variation less than +/- 0.2°C
3.19.10 | A temperature sensor must be available in the hood for temperature control and monitoring.
3.19.11 | In addition to the primary sensor for controlling the temperature, the unit must have a secondary fail-safe sensor which will act as a safety back-up and activate audible and visible alarms and shut off power to the heater if the temperature exceeds 40oC.
3.19.12 | The supply to the heater may only be restored after a manual reset and if the incubator’s secondary temperature monitoring device senses that the internal temperature has decreased to a safe level.
3.19.13 | At switch-on the incubator must automatically heat to 34°C +/- 1°C
3.19.14 | When the hood’s front access panel is opened, it must have a minimal effect on the set temperature within the incubator.
3.19.15 | Reusable skin temperature probe must be supplied
3.20 | DISPLAY
3.20.1 | LCD or LED display for the following:
3.20.1a | Skin temperature.
3.20.1b | Air temperature.
3.20.1c | Set temperature
3.20.1d | Details of all alarms
3.20.2 | Display must be large and clear enough to be viewed from three metres under all lighting conditions.
3.21 | ALARMS
3.21.1 | Default alarms must be enabled when the incubator is turned on
3.21.2 | Audible and visual alarms with heater power cut-off in the event of:
3.21.2a | Over temperature
3.21.2b | Failure of air circulation
3.21.2c | Audible and visual alarms in the event of:
3.21.2d | skin temperature differs from set temperature by - 1°C
3.21.2e | air temperature differs from set temperature by - 1°C
3.21.2f | low humidity
3.21.2g | Power failure
3.21.2h | Heater failure
3.21.2i | Skin or air probe disconnection or failure.
3.21.3 | The audible alarm must have a mute option. Visual alarm may not be reset until the cause of the alarm has been rectified
3.21.4 | Visual alarm may not be reset until the cause of the alarm has been rectified
3.21.5 | Test facility for audible and visual alarms.
3.21.6 | Low water alarm
3.22 | GENERAL
3.22.1 | The air velocity above the mattress must not exceed 0.35 m/s.
3.22.2 | Exhaled air must continuously be removed to prevent carbon dioxide build-up inside the hood.
3.22.3 | Fresh air drawn to the hood must be filtered by an easily replaceable bacterial filter
3.22.4 | The mattress must be approximately 630 x 350 x 30 mm and must be very durable, comfortable and have a waterproof cover.
3.22.5 | The sound level within the baby compartment must not exceed 45 dB(A) except if an audible alarm is sounding, then the sound level within the baby compartment must not exceed 80 dB(A).
3.22.6 | The incubator must operate from a 220V, 50Hz single phase supply
3.22.7 | Mains cable at least three metres long with SABS approved 15A three prong plug.
3.22.8 | All metal parts of the incubator must be well earthed.
3.22.9 | The incubator’s control module must be well designed to guard against entry of liquids
3.22.10 | All parts on the incubator and trolley must be rust resistant
3.22.11 | Labels and markings must be durable enough to withstand routine cleaning and normal wear
3.22.12 | The unit must be simple to use and maintain.
3.22.13 | The unit must be mounted on a robust trolley
3.22.14 | Trolley must have 100 mm diameter (minimum) antistatic swivel castors.
3.22.15 | At least two of the castors must have brakes.
3.22.16 | Handles, for pushing, on trolley
3.22.17 | The unit must be designed for easy access to serviceable parts
3.22.18 | The incubator must be easy to clean, disinfect, and/or sterilize
3.22.19 | The incubator must have no sharp edges.
3.22.20 | Switches, knobs, and other controls must be designed for conditions of heavy use.
3.22.21 | Storage drawer
3.23 | The unit must be handed over in full operating order
3.24 | A starter-pack of consumables must be supplied with the unit.
4 | The following accessories must be supplied with the unit:
4.1 | Bacterial filter (Pk of 4)
4.2 | Skin temperature probe (reusable)
5 | Conformity Compliance (Please attach valid certificate with ref number and expiry date)
5.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
5.2 | OEM must comply and certified on ISO 9001 quality standards, attach proof of compliance
5.2.1 | Please provide unique ref number of the ISO 9001 ceriticate:
5.3 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
5.3.1 | Please provide unique ref number of the ISO 13485 ceriticate:
5.4 | Model quoted for must be EC certified. Attach a copy of certification
5.4.1 | Please provide unique ref number of the EC ceriticate:
5.5 | Hazardous Substance Act: If this type of equipment / apparatus appears on the schedule of Hazardous substances issued by the Directorate: Health Technology of the Department of Health, a license in terms of the Act on Hazardous Substances (Act 15/1973) must be submitted with the bid / quotation documents. The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.6 | All electrical/electronic medical equipment must be licensed by Radiation Control, where applicable a copy of the license must be submitted The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.6.1 | Radiation Control License number:
5.7 | The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade (rubber to with 2 screws) type and it must be a minimum length of three (3) meters. NB the mains cable of the unit tendered for must be SABS colour coded.
5.8 | The equipment quoted for must be protected against electromagnetic interference IEC 60601-1-2
5.9 | The unit must comply with electrical safety and performance standard such as IEC 60601-2-19, attached certification
5.9.1 | Please provide unique ref number of the IEC ceriticate:
6 | COMPREHENSIVE MAINTENANCE
6.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) must be included in the bid price
6.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
6.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
6.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
6.3 | The bidder to provide the OEM recommended service interval of the unit:
6.4 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
7 | General
7.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
7.2 | The unit on offer must be the latest technology.
Incubator, Standard Accessories
RT2-03-09a | Bacterial filter (Pk of 4)
RT2-03-09b | Iris port sleeves (Pk of 10)
RT2-03-09c | Oval port sleeves (Pk of 10)
RT2-03-09d | Skin temperature probe (reusable)
RT2-03-09e | Covers for temperature probe (Pk of 25)
RT2-03-09f | Internal battery for alarms
RT2-03-09g | Scale (UPGRADE)
RT2-03-09h | Observation lamp (UPGRADE)
RT2-03-09i | Blanket cover
RT2-03-09j | X-ray tray
RT2-03-09k | Accessory rail
Bidder to list any other Consumables or Accessories (Not listed above) for the model on offer below

## Incubator, Transport
RT2-03-10 | Incubator, Transport
1 | Description
1.1 | A unit used to transport infants safely, while providing a closed, controlled environment that warms an infant by circulating heated air around the infant´s skin.
2 | Product Offered:
2.1 | Make:
2.2 | Model:
2.3 | Initial Manufacturing year of Model:
2.4 | Radiation Control License number:
3 | Technical Specification
3.1 | Closed Controlled environment that warms an infant by circulating heated air over the skin.
3.2 | Capable of providing thermal support during transportation even in adverse weather conditions maintaining air temperature for at least 4 hours when operating on battery power.
3.3 | Double wall to decrease radiant heat loss
3.4 | Control through microprcessed panel with environmental temperature wide view displays
3.5 | Air Controlled mode temperature range 20 - 39⁰ C
3.6 | Servo Controlled mode temperature range 34- 37⁰ C
3.7 | Manual modes should have 15 second interval alarms on overrides.
3.8 | Temperature measurements displayed in increments of 0.1⁰ C
3.9 | Required Alarms :
3.9.1 | High Air Temperature
3.9.2 | Low Air Temperature
3.9.3 | High Skin Temperature
3.9.4 | Low Skin Temperature
3.9.5 | Fan Failure
3.9.6 | Sensor Failure
3.9.7 | Power Failure
3.9.8 | Battery power level and when recharge is required.
3.10 | Battery Requirements:
3.10.1 | Battery should not take longer than 16 hours to recharge
3.10.2 | Unit should automatically switch over to battery power ones removed from line power.
3.10.3 | The unit should Maintain temperature at least 4 hours working on battery power.
3.10.4 | Indicate type of battery and life expectancy
3.11 | Minimum required ports :
3.11.1 | 2 Hand Ports
3.11.2 | 6 Tubing Ports
3.11.3 | 1 Oxygen inlet port
3.11.4 | Internal noise of less than 60 dB
3.12 | Caster Requirements
3.12.1 | Caster Size ≥ 12.7 cm
3.12.2 | At least 2 of the cater wheels should have locks,
3.12.3 | Casters should be conductive and swivel.
3.12.4 | Space should be provided to store and secure two gas Cylinders
3.12.5 | Aluminum stand with three positions system coupliing for in-hospital transport purposes.
4 | Standard Accessories
4.1 | Bacterial filter (Pk of 4)
4.4 | Skin temperature probe (reusable)
4.5 | Covers for temperature probe (Pk of 25)
5 | Conformity Compliance (Please attach certificate)
5.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
5.2 | OEM must comply and certified on ISO 9001 quality standards, attach proof of compliance
5.2.1 | Please provide unique ref number of the ISO 9001 ceriticate:
5.3 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
5.3.1 | Please provide unique ref number of the ISO 13485 ceriticate:
5.4 | Model quoted for must be EC certified. Attach a copy of certification
5.4.1 | Please provide unique ref number of the EC ceriticate:
5.5 | Hazardous Substance Act: If this type of equipment / apparatus appears on the schedule of Hazardous substances issued by the Directorate: Health Technology of the Department of Health, a license in terms of the Act on Hazardous Substances (Act 15/1973) must be submitted with the bid / quotation documents. The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.6 | All electrical/electronic medical equipment must be licensed by Radiation Control, where applicable a copy of the license must be submitted The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.6.1 | Radiation Control License number:
5.7 | The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade (rubber to with 2 screws) type and it must be a minimum length of three (3) meters. NB the mains cable of the unit tendered for must be SABS colour coded.
5.8 | The equipment quoted for must be protected against electromagnetic interference IEC 60601-1-2
5.9 | The unit must comply with electrical safety and performance standard such as IEC 60601-2-20, attached certification
5.9.1 | Please provide unique ref number of the IEC ceriticate:
6 | COMPREHENSIVE MAINTENANCE
6.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) must be included in the bid price
6.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
6.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
6.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
6.3 | The bidder to provide the OEM recommended service interval of the unit:
6.4 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
7 | General
7.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
7.2 | The unit on offer must be the latest technology.
Incubator, Transport Accessories and Consumables
RT2-03-10a | Bacterial filter (Pk of 4)
RT2-03-10b | Iris port sleeves (Pk of 10)
RT2-03-10c | Oval port sleeves (Pk of 10)
RT2-03-10d | Skin temperature probe (reusable)
RT2-03-10e | Covers for temperature probe (Pk of 25)
Bidder to list any other Consumables or Accessories (Not listed above) for the model on offer below

## Infant Radiant Warmer
RT2-03-11 | Infant radiant warmer
1 | Description
1.1 | A warming device to provide uniformal heat to the body using an overhead heating source
2 | Product Offered:
2.1 | Make:
2.2 | Model:
2.3 | Initial Manufacturing year of Model:
2.4 | Radiation Control License number:
3 | Technical Specification
3.1 | The unit must offer long-term care of a neonate
3.2 | The infant radiant warmer must function off 220 V, 50 Hz single phase power supply.
3.3 | Platform and mattress:
3.3.1 | Adjustable tilt to Trendellenburg and reverse Trendellenburg positions
3.3.2 | Locking mechanism to keep the mattress platform locked in all positions.
3.3.3 | The mattress must be waterproof and must be X-ray transparent.
3.3.4 | Mattress size: approximately 500 x 650 mm.
3.3.5 | Mattress thickness: at least 25 mm
3.3.6 | Mattress must support patient weight up to 4,5 kg.
3.3.7 | All four sides of the platform must be fitted with transparent guards/panels.
3.3.8 | The front and two side panels must be hinged.
3.4 | Heater
3.4.1 | Radiant heater housed in an overhead canopy, with internal reflector, mounted on the rear mainframe.
3.4.2 | Indicator to show heater is switched ON
3.4.3 | The overhead canopy must swivel left and right and must lock in the desired position.
3.5 | Controls
3.5.1 | Heat control: adjustable heater power from 0 - 100%.
3.5.2 | If a malfunction occurs affecting safety or performance the heater must be automatically turned off
3.5.3 | Skin temperature probe to monitor infant temperature.
3.5.4 | Heat output must be controlled manually or through baby mode (feedback mode) for thermoregulation.
3.6 | Displays:
3.6.1 | Heater power
3.7 | Alarms
3.7.1 | Visual and audio alarms must be present for safety
3.7.2 | Apgar timer, with audible alarm, with digital display: 1, 5 and 10 minutes elapsed time.
3.7.3 | Apgar timer with on/off switch and reset button
3.8 | Other
3.8.1 | The overhead canopy must be provided with an examination lamp, with on/off switch.
3.8.2 | The infant radiant warmer must have a drawer
3.8.3 | The minimum acceptable height of the heater canopy must be 1.85 metres above floor
3.8.4 | Adjustable mattress platform height with a minimum of 900 mm
3.8.5 | The mains cable of the unit must be at least three metres long and be fitted with an SABS approved 15A three pin plug
3.8.6 | The unit must be mobile, but stable during movement
3.8.7 | 4 swivel castors of minimum diameter 100 mm
3.8.8 | Two of the castors must be lockable.
3.8.9 | The unit must have oxygen cylinder holder
3.9 | The unit must be handed over in full operating order
3.10 | A starter-pack of consumables must be supplied with the unit.
4 | The following accessories must be supplied with the unit:
4.1 | Reusable Temperature probe
4.2 | Mattress
1 | Conformity Compliance (Please attach certificate)
5.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
5.2 | OEM must comply and certified on ISO 9001 quality standards, attach proof of compliance
5.2.1 | Please provide unique ref number of the ISO 9001 ceriticate:
5.3 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
5.3.1 | Please provide unique ref number of the ISO 13485 ceriticate:
5.4 | Model quoted for must be EC certified. Attach a copy of certification
5.4.1 | Please provide unique ref number of the EC ceriticate:
5.5 | Hazardous Substance Act: If this type of equipment / apparatus appears on the schedule of Hazardous substances issued by the Directorate: Health Technology of the Department of Health, a license in terms of the Act on Hazardous Substances (Act 15/1973) must be submitted with the bid / quotation documents. The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.7 | All electrical/electronic medical equipment must be licensed by Radiation Control, where applicable a copy of the license must be submitted The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.7.1 | Radiation Control License number:
5.8 | The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade (rubber to with 2 screws) type and it must be a minimum length of three (3) meters. NB the mains cable of the unit tendered for must be SABS colour coded.
5.9 | The equipment quoted for must be protected against electromagnetic interference IEC 60601-1-2
5.10 | The unit must comply with electrical safety and performance standard such as IEC 60601-2-21, attached certification
5.10.1 | Please provide unique ref number of the IEC ceriticate:
6 | COMPREHENSIVE MAINTENANCE
6.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) must be included in the bid price
6.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
6.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
6.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
6.3 | The bidder to provide the OEM recommended service interval of the unit:
6.4 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
7 | General
7.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
7.2 | The unit on offer must be the latest technology.
Infant radiant warmer Accessories
RT2-03-11a | Reusable Temperature probe
RT2-03-11b | Mattress
RT2-03-11c | LED Phototherapy (UPGRADE)
RT2-03-11d | IV pole
RT2-03-11e | Disposable temperature probe (Pk of 10)
RT2-03-11f | Accessory rail
RT2-03-11g | Monitor stand
Bidder to list any other Consumables or Accessories (Not listed above) for the model on offer below

## Phototherapy Lamp, LED
RT2-03-12 | Phototherapy Light (LED)
1 | Description
2 | Product Offered:
2.1 | Make:
2.2 | Model:
2.3 | Initial Manufacturing year of Model:
3 | Technical Specification
3.1 | The phototherapy light system must be suitable for the treatment of neonatal hyperbilirubinemia.
3.2 | Microprocessor control with LCD display of set treatment time, lamp usage time, intensity level
3.3 | The light system must be suitable for babies in a bassinet, incubator, open incubator or radiant warmer.
3.4 | The light source must be LEDs, emiting light in the range of 450 – 475nm
3.5 | Target light to ensure correct placement of lamp over patient
3.6 | The unit must have at least two intensity settings, high and low.
3.7 | The high setting must give an output of 30-45 µW/cm²/nm at a distance of 300 mm from the light enclosure.
3.8 | The low setting must give an output of 12-15 µW/cm²/nm at a distance of 300 mm from the light enclosure.
3.9 | The LED light source must ensure a uniform illumination pattern at the skin surface
3.10 | The light source housing must tilt at least to 30o in either direction to provide effective phototherapy
3.11 | The lamp effective surface area must be at least 400 x 200 mm at 300 mm from the light source
3.12 | The unit must be height adjustable : approximately 1200 - 1600 mm from light source to floor
3.13 | The light enclosure must be protected from the ingress of fluids.
3.14 | The light unit must be mounted on a roll stand.
3.15 | The roll stand must have lockable castors.
3.16 | The base of the roll stand must not tip, even when the light enclosure is tilted.
3.17 | Must have a re-settable treatment timer that will switch off the light when the set time has expired
3.18 | The unit must have a time recorder that operates only when the phototherapy light is switched on. It must be possible to reset the time recorder to zero when the light source is replaced.
3.19 | It must not be possible to accidently reset the time recorder
3.20 | LED array replacement must be a simple operation
3.21 | It must be possible to replace only a section of the LED array
3.22 | The unit must operate off the 220 Volt, 50Hz power supply.
3.23 | Mains cable at least three metres long with SABS approved 15A three prong plug.
3.24 | Warranty period for LED light source of no less than 20000 hrs or 3yrs
3.25 | The unit must be handed over in full operating order
4 | Conformity Compliance (Please attach certificate)
4.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
4.2 | OEM must comply and certified on ISO 9001 quality standards, attach proof of compliance
4.2.1 | Please provide unique ref number of the ISO 9001 ceriticate:
4.3 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
4.3.1 | Please provide unique ref number of the ISO 13485 ceriticate:
4.4 | Model quoted for must be EC certified. Attach a copy of certification
4.4.1 | Please provide unique ref number of the EC ceriticate:
4.5 | The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade (rubber to with 2 screws) type and it must be a minimum length of three (3) meters. NB the mains cable of the unit tendered for must be SABS colour coded.
4.6 | The equipment quoted for must be protected against electromagnetic interference IEC 60601-1-2
4.7 | The unit must comply with electrical safety and performance standard such as IEC 60601-2-50, attached certification
4.7.1 | Please provide unique ref number of the IEC ceriticate:
5 | COMPREHENSIVE MAINTENANCE
5.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) must be included in the bid price
5.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
5.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
5.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
5.3 | The bidder to provide the OEM recommended service interval of the unit:
5.4 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
6 | General
6.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
6.2 | The unit on offer must be the latest technology.
Phototherapy Light (LED) Accessories and Consumables
RT2-03-12a | Monitoring of baby temperature (UPGRADE)
RT2-03-12b | LED examination light/Bar replacement
Bidder to list any other Consumables or Accessories (Not listed above) for the model on offer below

## Phototherapy Blanket, LED
RT2-03-13 | Phototherapy Blanket LED
1 | Description
2 | Product Offered:
2.1 | Make:
2.2 | Model:
2.3 | Initial Manufacturing year of Model:
3 | Technical Specification
3.1 | The phototherapy light system must be suitable for the treatment of neonatal hyperbilirubinemia.
3.2 | Microprocessor control with LCD display of set treatment time, therapy time, lamp usage time, intensity level, infant skin temp
3.3 | The light system must be suitable for babies in a bassinet, incubator, or radiant warmer.
3.4 | The light source must be LEDs, emiting light in the range of 450 – 475nm
3.5 | The unit must have at least two intensity settings, high and low.
3.6 | Intensity: Maximum 40-50uW/cm2/nm and Minimum 30-35uW/cm2/nm
3.7 | Effective treatment area: Small pad of 250-350cm2
3.8 | The unit must not emit light in the Ultraviolet or Infrared radiation range
3.9 | The unit must consist of a soft, flexible, oval design that conforms to the shape of the baby to deliver phototherapy via an LED light source
3.10 | The unit must be able to shut off automatically in the event of the blanket overheating
3.11 | Must have a re-settable treatment timer that will switch off the light when the set time has expired
3.12 | The unit must have a time recorder that operates only when the phototherapy light is switched on. It must be possible to reset the time recorder to zero when the light source is replaced.
3.13 | It must not be possible to accidently reset the time recorder
3.14 | LED array replacement must be a simple operation
3.15 | It must be possible to replace only a section of the LED array
3.16 | The unit must operate off the 220 Volt, 50Hz power supply.
3.17 | Mains cable at least three metres long with SABS approved 15A three prong plug.
3.18 | Warranty period for LED light source of no less than 20000 hrs or 3yrs
3.19 | The unit must be handed over in full operating order
4 | Conformity Compliance (Please attach certificate)
4.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
4.2 | OEM must comply and certified on ISO 9001 quality standards, attach proof of compliance
4.2.1 | Please provide unique ref number of the ISO 9001 ceriticate:
4.3 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
4.3.1 | Please provide unique ref number of the ISO 13485 ceriticate:
4.4 | Model quoted for must be EC certified. Attach a copy of certification
4.4.1 | Please provide unique ref number of the EC ceriticate:
4.5 | The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade (rubber to with 2 screws) type and it must be a minimum length of three (3) meters. NB the mains cable of the unit tendered for must be SABS colour coded.
4.6 | The equipment quoted for must be protected against electromagnetic interference IEC 60601-1-2
4.7 | The unit must comply with electrical safety and performance standard such as IEC 60601-2-50, attached certification
4.7.1 | Please provide unique ref number of the IEC ceriticate:
5 | COMPREHENSIVE MAINTENANCE
5.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) must be included in the bid price
5.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
5.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
5.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
5.3 | The bidder to provide the OEM recommended service interval of the unit:
5.4 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
6 | General
6.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
6.2 | The unit on offer must be the latest technology.
Phototherapy Blanket LED
RT2-03-13a | Monitoring of baby temperature (UPGRADE)
RT2-03-13b | Effective treatment area: Large pad of 450-550cm2
Bidder to list any other Consumables or Accessories (Not listed above) for the model on offer below

## Ultrasound, Neonatal
RT2-03-15 | ULTRASOUND UNIT, NEONATAL
1 | Description
2 | Product Offered:
2.1 | Make:
2.2 | Model:
2.3 | Initial Manufacturing year of Model:
2.4 | Radiation Control License number:
3 | Technical Specification
3.1 | The unit shall be a fully digital unit with a memory beam former and processor to generate sharp defined ultrasound images, reduce artefacts and improve image quality.
3.2 | It shall utilise the latest software to perform investigations.
3.3 | High image quality will be one of the main criteria that will be used to evaluate the offers.
3.4 | Although offered units might comply with all the specifications, image quality shall be a critical evaluation criterion, specifically with increased BMI cases.
3.5 | The unit must be of the latest technology.
3.6 | It must be a state of the art fully digital phase array ultrasound unit.
3.7 | IMAGING MODALITIES
3.7.1 | The unit must be latest technology; please state year in which this unit was first manufactured.
3.7.2 | The unit shall offer excellent high image quality in B mode and fully digital colour flow Doppler. Colour Doppler should have no delay after activation.
3.7.3 | The system Dynamic Range shall not be less than 172 dB. The Dynamic range should be able to be changed in 5dB increments.
3.7.4 | The unit shall have direct storage on hard drive of not less than 80 Gig. State all other storage facility that is included in the offered price. (Such as CD, DVD, MOD). Recorded CD roms shall be compatible with any standard PC, without requiring extra software. USB download of images from the system must be possible
3.7.5 | A digital black and white thermal printer shall be included in the offer for the machine.
3.7.6 | The system shall offer a Digital Continuous Beam Former to generate sharp defined ultrasound beams, to reduce artefacts and to provide more consistent overall image quality.
3.7.7 | The unit shall offer multi-frequency on all probes to provide sharp reproducible image quality in both superficial and deep anatomy.
3.7.8 | The unit shall be equipped with a non-interlaced colour monitor with a screen size of not less than 15 inch, measured from corner to corner. The monitor shall be of the flat screen TFT LCD type.
3.7.9 | The unit should have at least three active probe connectors with electronic switching.
3.7.10 | The unit shall be able to offer scanning methods for Electronic Convex and Linear technology.
3.7.11 | The unit shall offer depth gain in not less than 8 steps. Unit must be programmable to start up automatically at a depth of operator’s choice and programmable to start up at different depths for specific examinations.
3.7.12 | Image processing shall be done via a rotary dial or similar technology with at least 5 combinations for optimising the best image quality for each patient’s physique (user programmable). Image processing should visibly improve image and not just change it.
3.7.13 | Pan-Zoom enlargement shall be in real and frozen mode up to 10X zoom.
3.7.14 | A colour pallet shall be available for the colourization of the B mode and M mode and PW Doppler trace.
3.7.15 | Independent M mode and PW Doppler gain adjustment shall be available.
3.7.16 | Sweep speed shall be adjustable in at least 4 steps.Unit shall have selectable sweep rates of 1,2,3,4,8 s/screen
3.7.17 | The unit shall offer Tissue Harmonics as standard in the system.
3.7.18 | Grey scale level shall be at least 256 shades.
3.7.19 | The unit shall have not less than 1024 processing channels.
3.7.20 | The unit shall be housed in a robust mobile cart. The cart shall house all elements of the unit, and shall have enough dedicated storage space for the probes, and all space for all disposables needed on the machine. The unit to be used for mobile ultrasound in different areas. a) Width of system in centimetres b) Depth of system in centimetres.
3.7.21 | The frame rate shall not be less than 394 frames per second. Indicate the frame rate that the system offers
3.7.22 | Image depth shall be not less than 30cm.
3.7.23 | The unit should offer quick scan to automatically adjusts STC and gain through one-touch operation. This must visibly improve the image
3.7.24 | Post processing must be available on frozen images
3.7.25 | Unit must offer user programmable annotation for quick labelling. User must be able to program approximately 40 words of their own choice per selected examination type for annotation purposes
3.8 | DOPPLER
3.8.1 | PW Doppler shall be offered on all imaging probes.
3.8.2 | Max detectable velocity for PW shall be stated.
3.8.3 | Pulse Repetition Frequency shall vary between 0,7 – 19,8 kHz pulsed wave and 2,1 – 40kHz continuous wave
3.8.4 | Independent angle steering up to ±20° for colour and Doppler shall be offered.
3.8.5 | Independent gain control for PW shall be available.
3.8.6 | Colour pallet offering 7 types independent of B mode and Pulse Doppler pallet to be available.
3.8.7 | Sampling with Pulse Doppler shall vary between 1 and 20 mm.
3.8.8 | Gate size must vary at least between 1 and 8mm.
3.8.9 | Base line shift must be possible in at least 8 steps.
3.8.10 | Spectral invert must be possible.
3.8.11 | At least 6 wall filter selections shall be available. Bidder to state minimum.
3.8.12 | Accurate autotrace function for both arterial and venous waveforms shall be available.
3.9 | COLOUR DOPPLER
3.9.1 | A minimum of 8 colour maps shall be available for colour display.
3.9.2 | Colour maps shall be available to be selected in real time and freeze mode.
3.9.3 | Independent colour gain adjustment shall be from 1 to 16 dB in 0,5 dB steps. Indicate range
3.9.4 | Maximum colour frame rate shall not be less than 285 frames per second.
3.9.5 | Colour flow shall operate to the full depth of 30 cm.
3.9.6 | Colour maps and processing selections shall be able to be selected by the user and stored in the presets.
3.9.7 | Quad signal processing and parallel signal processing shall be available. Bidders to supply details.
3.9.8 | Power Doppler on all transducers shall be available.
3.9.9 | The unit offered shall be able of real time triplex mode.
3.9.10 | Colour mapping must be instantaneous – no time delay /lag.
3.10 | IMAGE MEMORY
3.10.1 | Image memory to operate in both dual and quad screen format.
3.10.2 | ECG synchronised recording shall be possible.
3.10.3 | It shall be possible to record B loop up to 500 frames.
3.10.4 | Cine loop and measurements must be active simultaneously.
3.10.5 | M/D review of a minimum of 15 seconds shall be possible.
3.10.6 | B Mode Imaging on linear transducer to be electronically steerable.
3.11 | TRANSDUCERS
3.11.1 | The following probes shall be included for the machine:
3.11.1.1 | Phased Array probe - 2.0 to 7.5 Mhz)
3.11.1.2 | Linear probe - 5 to 10 Mhz
3.11.1.3 | Neonatal Cardiac probe – 3 to 9 Mhz
3.11.2 | All imaging transducers shall be multi-frequency.
3.11.3 | Tissue Harmonics shall be available on all offered transducers.
3.11.4 | The selected frequency of the imaging transducer shall be displayed on the screen.
3.11.5 | All transducers shall be digitally controlled so continuous beam forming is possible. State number of focal zones.
3.11.6 | A minimum of 2 active probes shall be available. It shall be possible to electronically switch between these 2 probes.
3.12 | MEASUREMENTS
3.12.1 | Patient data MUST offer the option to enter previous ultrasound date and measurements.
3.12.2 | Gestational age according to LMP must be displayed on working screen, not just not patient information page.
3.12.3 | On-screen measurements must be possible from real time. 16 callipers (8 channel)each individually numbered for measuring distance/volume/angle and area must be offered
3.12.4 | The unit shall offer a storage facility for up to 300 patient measurement data. It shall be possible to retrieve the data for follow-ups.
3.12.5 | The unit shall offer 50th percentile growth graphs of the Doppler measurements RI/RP. It shall be possible to retrieve the data for follow-ups.
3.12.6 | The unit shall offer user programmable annotation for quick labelling.
3.12.7 | State if the unit comply to the following: 1. Panoramic view. 2. Quad signal processing
3.12.8 | All software upgrades shall be offered optionally, with full description what improvement the software upgrade will bring to the offered unit.
3.12.9 | The unit shall be supplied with an uninterrupted power supply (UPS), to be able to keep the unit running for at least two hours in case of power failure.
3.13 | DICOM COMPLIANCE
3.13.1 | The offered unit shall be fully DICOM compliant.
3.13.2 | The DICOM licence, as well as connecting the ultrasound to the envisaged PACS shall be included in the offer.
3.13.3 | Please note that installation of the PACS at False Bay Hospital is planned for this year, and that the successful contractor shall assist in setting up the machine with the PACS.
4 | Standard Accessories
4.1 | Phased Array probe - 2.0 to 7.5 Mhz)
4.2 | Linear probe - 5 to 10 Mhz
4.3 | Neonatal Cardiac probe – 3 to 9 Mhz
4.4 | Uninterrupted Power Supply (UPS)
4.5 | Gel
4.6 | Printer
4.7 | Printing paper
5 | Conformity Compliance (Please attach certificate)
5.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
5.2 | OEM must comply and certified on ISO 9001 quality standards, attach proof of compliance
5.2.1 | Please provide unique ref number of the ISO 9001 ceriticate:
5.3 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
5.3.1 | Please provide unique ref number of the ISO 13485 ceriticate:
5.4 | Model quoted for must be EC certified. Attach a copy of certification
5.4.1 | Please provide unique ref number of the EC ceriticate:
5.5 | Hazardous Substance Act: If this type of equipment / apparatus appears on the schedule of Hazardous substances issued by the Directorate: Health Technology of the Department of Health, a license in terms of the Act on Hazardous Substances (Act 15/1973) must be submitted with the bid / quotation documents. The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.7 | All electrical/electronic medical equipment must be licensed by Radiation Control, where applicable a copy of the license must be submitted The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.7.1 | Radiation Control License number:
5.8 | The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade (rubber to with 2 screws) type and it must be a minimum length of three (3) meters. NB the mains cable of the unit tendered for must be SABS colour coded.
5.9 | The equipment quoted for must be protected against electromagnetic interference IEC 60601-1-2
5.10 | The unit must comply with electrical safety and performance standard such as IEC 60601-2-37, attached certification
5.10.1 | Please provide unique ref number of the IEC ceriticate:
6 | COMPREHENSIVE MAINTENANCE
6.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) must be included in the bid price
6.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
6.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
6.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
6.3 | The bidder to provide the OEM recommended service interval of the unit:
6.4 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
7 | General
7.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
7.2 | The unit on offer must be the latest technology.
ULTRASOUND UNIT, NEONATAL
RT2-03-15a | Phased Array probe - 2.0 to 7.5 Mhz)
RT2-03-15b | Linear probe - 5 to 10 Mhz
RT2-03-15c | Neonatal Cardiac probe – 3 to 9 Mhz
RT2-03-15d | Uninterrupted Power Supply (UPS)
RT2-03-15e | Gel
RT2-03-15f | Printer
RT2-03-15g | Printing paper
Bidder to list any other Consumables or Accessories (Not listed above) for the model on offer below

## Category-4 O&G Equipment
OBSTETRIC & GYNAECOLOGY EQUIPMENT
# | Item number | Item Description
1 | RT2-04-01 | Cardiotocograph, Single
2 | RT2-04-02 | Cardiotocograph, Single + MHR
3 | RT2-04-03 | Cardiotocograph, Twins
4 | RT2-04-04 | Cardiotocograph, Twins + MHR
5 | RT2-04-05 | Fetal Doppler
6 | RT2-04-06 | Video Colposcope
7 | RT2-04-07 | Vacuum Extractor
8 | RT2-04-08 | Ultrasound Machine, O&G

## Cardiotocograph, single
RT2-04-01 | Cardiotocograph, Singleton
1 | Description
2 | Product Offered:
2.1 | Make:
2.2 | Model:
2.3 | Initial Manufacturing year of Model:
2.4 | Radiation Control License number:
3 | Technical Specification
3.1 | Required for external monitoring of fetal heart rate (FHR), and external monitoring of uterine activity of pregnant patients
3.2 | Singleton monitoring is required
3.3 | Must perform automatic self-check when switched on
3.4 | There must a visible Fetal Heart Beat indicator which will flash with each detected heart beat
3.5 | The unit must provide an audio signal, with variable volume, for every heart beat detected
3.6 | The fetal heart rate must be measured by means of wide beam pulsed Doppler ultrasound, with auto correlation processing.
3.7 | Must be able to detect and display FHR
3.8 | The oscillator frequency < 1,2 MHz.
3.9 | Ultrasound intensity must be less than 10 mW/cm2
3.10 | The transducer must be manufactured of high-impact plastic that is robust and watertight.
3.11 | The belts and buckles supplied for positioning of the transducer must allow easy angling for optimum positioning.
3.12 | Adjustable baseline position
3.13 | A test button must be provided that produces a series of pulses for calibration which must be recorded and displayed.
3.14 | Display: > 16 cm (corner to corner).
3.15 | Displays must be clearly visible under all lighting conditions.
3.16 | Must allow Maternal Heart Rate (MHR) and Fetal Heart Rate(FHR) coincidence detection?
3.17 | The following parameters must be displayed:
3.17.1 | Fetal heart rate approximately 50 to 200 beats per minute
3.17.2 | Uterine activity: approximately 0 to 100 relative units.
3.18 | Fetal heart rate must update once per second when FHR is greater than 60 and once each beat when FHR is less than 60.
3.19 | The monitor must have the following controls:
3.19.1 | Volume Up/Down
3.19.2 | TOCO/IUP Zero
3.19.3 | Printer On/Off
3.19.4 | Printer fast forward
3.20 | Must have audible and visual alarms for low/high FHR
3.21 | Electronic storage data capacity.
3.22 | RECORDER
3.22.1 | Integrated dual channel thermal recorder
3.22.2 | Heart rate recording
3.22.3 | All dot print test
3.22.4 | The strip chart paper must be easy to load
3.23 | Fetal heart rate:
3.23.1 | Range: approximately 50 to 210 beats per minute.
3.23.2 | Vertical chart: at least 7 cm.
3.23.3 | Vertical scale sensitivity: approximately 20 bpm per cm.
3.23.4 | Resolution: 1 bpm
3.24 | Uterine activity:
3.24.1 | Range: 0 – 100 units
3.24.2 | Vertical chart: at least 4 cm.
3.24.3 | Vertical scale sensitivity: approximately 25 relative units per cm
3.24.4 | Resolution: 1 relative unit
3.25 | Paper chart speed must be selectable for 1, 2 and 3 cm per minute.
3.26 | The time and date must be printed by the recorder at least every 10 minutes.
3.27 | Recording time must be at least 24 hours at 1 cm per minute.
3.28 | The unit must have a facility for remote event marking
3.29 | TRANSDUCERS
3.29.1 | Ultrasound and toco transducers must be supplied
3.29.2 | All transducers must be of very robust design. Bidders must guarantee shock resistance on the transducers e.g. will withstand a drop from one metre high onto a hard surface
3.29.3 | Bidders must state whether the ultrasound transducer is repairable and which parts are replaceable or whether it is a sealed unit that must be replaced when faulty.
3.29.4 | The ultrasound transducer cable must be a minimum of 2 m long
3.29.5 | Connecting plug must be keyed and colour coded to avoid it being plugged or forced into the wrong receptacle on the unit.
3.29.6 | Transducer holders must be provided
3.30 | General
3.30.1 | The unit must operate off 220V, 50Hz power supply
3.30.2 | All metal parts must be earthed.
3.30.3 | Mains cable at least three metres long with SABS approved 15A three prong plug.
3.30.4 | Must be fitted with rechargeable batteries
3.30.5 | The unit must be fitted with an internal battery charger
3.30.6 | Indicator for battery status
3.30.7 | Must store trends for at least 8 hours
3.30.8 | A cart/trolley must be supplied
3.30.9 | The cart must have four castor wheels, two of them with brakes
3.30.10 | The cart must have a basket/holder for accessories
3.30.11 | The unit must be handed over in full operating order
3.30.12 | A starter-pack of consumables must be supplied with the unit.
4 | Standard Accessories and Consumables
4.1 | US transducer
4.2 | Toco transducer
4.3 | Thermal printing paper (Pk of 5)
4.4 | Belts x2 (For each transducer)
4.5 | Marker
5 | Conformity Compliance (Please attach certificate)
5.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
5.2 | OEM must comply and certified on ISO 9001 quality standards, attach proof of compliance
5.2.1 | Please provide unique ref number of the ISO 9001 ceriticate:
5.3 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
5.3.1 | Please provide unique ref number of the ISO 13485 ceriticate:
5.4 | Model quoted for must be EC certified. Attach a copy of certification
5.4.1 | Please provide unique ref number of the EC ceriticate:
5.5 | Hazardous Substance Act: If this type of equipment / apparatus appears on the schedule of Hazardous substances issued by the Directorate: Health Technology of the Department of Health, a license in terms of the Act on Hazardous Substances (Act 15/1973) must be submitted with the bid / quotation documents. The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.6 | All electrical/electronic medical equipment must be licensed by Radiation Control, where applicable a copy of the license must be submitted The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.6.1 | Radiation Control License number:
5.7 | The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade (rubber to with 2 screws) type and it must be a minimum length of three (3) meters. NB the mains cable of the unit tendered for must be SABS colour coded.
5.8 | The equipment quoted for must be protected against electromagnetic interference IEC 60601-1-2
6 | COMPREHENSIVE MAINTENANCE
6.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) must be included in the bid price
6.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
6.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
6.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
6.3 | The bidder to provide the OEM recommended service interval of the unit:
6.4 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
7 | General
7.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
7.2 | The unit on offer must be the latest technology.
Cardiotocograph Accessories and Consumables
RT2-04-01a | US transducer
RT2-04-01a | Toco transducer
RT2-04-01a | Thermal printing paper (Pk of 5)
RT2-04-01a | Belt
RT2-04-01a | Marker
Bidder to list any other Consumables or Accessories (Not listed above) for the model on offer below

## Cardiotocograph with MHR
RT2-04-02 | Cardiotocograph, singleton with MRH monitoring
1 | Description
2 | Product Offered:
2.1 | Make:
2.2 | Model:
2.3 | Initial Manufacturing year of Model:
2.4 | Radiation Control License number:
3 | Technical Specification
3.1 | Required for external monitoring of fetal heart rate (FHR), and external monitoring of uterine activity of pregnant patients + must include monitoring and display of maternal heart rate
3.2 | Singelton monitoring is required
3.3 | Must perform automatic self-check when switched on
3.4 | There must a visible Fetal Heart Beat indicator which will flash with each detected heart beat
3.5 | The unit must provide an audio signal, with variable volume, for every heart beat detected
3.6 | The fetal heart rate must be measured by means of wide beam pulsed Doppler ultrasound, with auto correlation processing.
3.7 | Must be able to detect and display multiple FHR while discriminating between different foetuses
3.8 | The oscillator frequency < 1,2 MHz.
3.9 | Ultrasound intensity must be less than 10 mW/cm2
3.10 | The transducer must be manufactured of high-impact plastic that is robust and watertight.
3.11 | The belts and buckles supplied for positioning of the transducer must allow easy angling for optimum positioning.
3.12 | Adjustable baseline position
3.13 | A test button must be provided that produces a series of pulses for calibration which must be recorded and displayed.
3.14 | Display: > 16 cm (corner to corner).
3.15 | Displays must be clearly visible under all lighting conditions.
3.16 | Must allow Maternal Heart Rate (MHR) and Fetal Heart Rate(FHR) coincidence detection?
3.17 | The following parameters must be displayed:
3.18 | Fetal heart rate approximately 50 to 200 beats per minute
3.19 | Uterine activity: approximately 0 to 100 relative units.
3.20 | Heart Rate monitoring
3.21 | Fetal heart rate must update once per second when FHR is greater than 60 and once each beat when FHR is less than 60.
3.22 | Maternal heart rate must be obtained, and displayed, reliably using preferably ECG or plethysmography
3.23 | The monitor must have the following controls:
3.23.1 | Volume Up/Down
3.23.2 | TOCO/IUP Zero
3.23.3 | Printer On/Off
3.23.4 | Printer fast forward
3.24 | Must have audible and visual alarms for low/high FHR
3.25 | Electronic storage data capacity.
3.26 | RECORDER
3.26.1 | Integrated dual channel thermal recorder
3.26.2 | Heart rate recording must be available.
3.26.3 | All dot print test
3.26.4 | The strip chart paper must be easy to load
3.27 | Fetal heart rate:
3.27.1 | Range: approximately 50 to 210 beats per minute.
3.27.2 | Vertical chart: at least 7 cm.
3.27.3 | Vertical scale sensitivity: approximately 20 bpm per cm.
3.27.4 | Resolution: 1 bpm
3.28 | Uterine activity:
3.28.1 | Range: 0 – 100 units
3.28.2 | Vertical chart: at least 4 cm.
3.28.3 | Vertical scale sensitivity: approximately 25 relative units per cm
3.28.4 | Resolution: 1 relative unit
3.28.5 | Paper chart speed must be selectable for 1, 2 and 3 cm per minute.
3.28.6 | The time and date must be printed by the recorder at least every 10 minutes.
3.28.7 | Recording time must be at least 24 hours at 1 cm per minute.
3.28.8 | The unit must have a facility for remote event marking
3.29 | TRANSDUCERS
3.29.1 | Ultrasound and toco transducers must be supplied
3.29.2 | All transducers must be of very robust design. Bidders must guarantee shock resistance on the transducers e.g. will withstand a drop from one metre high onto a hard surface
3.29.3 | Bidders must state whether the ultrasound transducer is repairable and which parts are replaceable or whether it is a sealed unit that must be replaced when faulty.
3.29.4 | The ultrasound transducer cable must be a minimum of 2 m long
3.29.5 | Connecting plug must be keyed and colour coded to avoid it being plugged or forced into the wrong receptacle on the unit.
3.29.6 | Transducer holders must be provided
3.30 | General
3.30.1 | The unit must operate off 220V, 50Hz power supply
3.30.2 | All metal parts must be earthed.
3.30.3 | Mains cable at least three metres long with SABS approved 15A three prong plug.
3.30.4 | Must be fitted with rechargeable batteries
3.30.5 | The unit must be fitted with an internal battery charger
3.30.6 | Indicator for battery status
3.30.7 | Must store trends for at least 8 hours
3.30.8 | A cart/trolley must be supplied
3.30.9 | The cart must have four castor wheels, two of them with brakes
3.30.10 | The cart must have a basket/holder for accessories
3.30.11 | The unit must be handed over in full operating order
3.30.12 | A starter-pack of consumables must be supplied with the unit.
4 | Standard Accessories and Consumables
4.1 | US transducer
4.2 | Toco transducer
4.3 | Thermal printing paper (Pk of 5)
4.4 | Belts x2 (For each transducer)
4.5 | Marker
5 | Conformity Compliance (Please attach certificate)
5.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
5.2 | OEM must comply and certified on ISO 9001 quality standards, attach proof of compliance
5.2.1 | Please provide unique ref number of the ISO 9001 ceriticate:
5.3 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
5.3.1 | Please provide unique ref number of the ISO 13485 ceriticate:
5.4 | Model quoted for must be CE certified. Attach a copy of certification
5.4.1 | Please provide unique ref number of the CE ceriticate:
5.5 | Hazardous Substance Act: If this type of equipment / apparatus appears on the schedule of Hazardous substances issued by the Directorate: Health Technology of the Department of Health, a license in terms of the Act on Hazardous Substances (Act 15/1973) must be submitted with the bid / quotation documents. The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.6 | Must comply with 2.5 Occupational Health and Safety Act, Act no 85 of 1993 and Regulations
5.7 | All electrical/electronic medical equipment must be licensed by Radiation Control, where applicable a copy of the license must be submitted The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.7.1 | Radiation Control License number:
5.8 | The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade (rubber to with 2 screws) type and it must be a minimum length of three (3) meters. NB the mains cable of the unit tendered for must be SABS colour coded.
5.8 | The equipment quoted for must be protected against electromagnetic interference IEC 60601-1-2
6 | COMPREHENSIVE MAINTENANCE
6.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) must be included in the bid price
6.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
6.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
6.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
6.3 | The bidder to provide the OEM recommended service interval of the unit:
6.4 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
7 | General
7.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
7.2 | The unit on offer must be the latest technology.
Cardiotocograph Accessories and Consumables
RT2-04-02a | US transducer
RT2-04-02b | Toco transducer
RT2-04-02c | Thermal printing paper (Pk of 5)
RT2-04-02d | Belt
RT2-04-02e | Marker
Bidder to list any other Consumables or Accessories (Not listed above) for the model on offer below

## Cardiotocograph, Twins
RT2-04-03 | Cardiotocograph (twins) on Trolley
1 | Description
2 | Product Offered:
2.1 | Make:
2.2 | Model:
2.3 | Initial Manufacturing year of Model:
2.4 | Radiation Control License number:
3 | Technical Specification
3.1 | Required for external monitoring of fetal heart rate (FHR), and external monitoring of uterine activity of pregnant patients
3.2 | Twins monitoring is required
3.3 | Must perform automatic self-check when switched on
3.4 | There must a visible Fetal Heart Beat indicator which will flash with each detected heart beat
3.5 | The unit must provide an audio signal, with variable volume, for every heart beat detected
3.6 | The fetal heart rate must be measured by means of wide beam pulsed Doppler ultrasound, with auto correlation processing.
3.7 | Must be able to detect and display multiple FHR while discriminating between different foetuses
3.8 | The oscillator frequency < 1,2 MHz.
3.9 | Ultrasound intensity must be less than 10 mW/cm2
3.10 | The transducer must be manufactured of high-impact plastic that is robust and watertight.
3.11 | The belts and buckles supplied for positioning of the transducer must allow easy angling for optimum positioning.
3.12 | Adjustable baseline position
3.13 | A test button must be provided that produces a series of pulses for calibration which must be recorded and displayed.
3.14 | Display: > 16 cm (corner to corner).
3.15 | Displays must be clearly visible under all lighting conditions.
3.16 | Must allow Maternal Heart Rate (MHR) and Fetal Heart Rate(FHR) coincidence detection?
3.17 | The following parameters must be displayed:
3.17.1 | Fetal heart rate approximately 50 to 200 beats per minute
3.17.2 | Uterine activity: approximately 0 to 100 relative units.
3.18 | Fetal heart rate must update once per second when FHR is greater than 60 and once each beat when FHR is less than 60.
3.19 | The monitor must have the following controls:
3.19.1 | Volume Up/Down
3.19.2 | TOCO/IUP Zero
3.19.3 | Printer On/Off
3.19.4 | Printer fast forward
3.20 | Must have audible and visual alarms for low/high FHR
3.21 | Electronic storage data capacity.
3.22 | RECORDER
2.22.1 | Integrated dual channel thermal recorder
2.22.2 | Dual heart rate recording for twins must be available.
2.22.3 | All dot print test
2.22.4 | The strip chart paper must be easy to load
3.23 | Fetal heart rate:
3.23.1 | Range: approximately 50 to 210 beats per minute.
3.23.2 | Vertical chart: at least 7 cm.
3.23.3 | Vertical scale sensitivity: approximately 20 bpm per cm.
3.23.4 | Resolution: 1 bpm
3.24 | Uterine activity:
3.24.1 | Range: 0 – 100 units
3.24.2 | Vertical chart: at least 4 cm.
3.24.3 | Vertical scale sensitivity: approximately 25 relative units per cm
3.24.4 | Resolution: 1 relative unit
3.25 | Paper chart speed must be selectable for 1, 2 and 3 cm per minute.
3.26 | The time and date must be printed by the recorder at least every 10 minutes.
3.27 | Recording time must be at least 24 hours at 1 cm per minute.
3.28 | The unit must have a facility for remote event marking
3.29 | TRANSDUCERS
3.29.1 | 2x Ultrasound and toco transducers must be supplied
3.29.2 | All transducers must be of very robust design. Bidders must guarantee shock resistance on the transducers e.g. will withstand a drop from one metre high onto a hard surface
3.29.3 | Bidders must state whether the ultrasound transducer is repairable and which parts are replaceable or whether it is a sealed unit that must be replaced when faulty.
3.29.4 | The ultrasound transducer cable must be a minimum of 2 m long
3.29.5 | Connecting plug must be keyed and colour coded to avoid it being plugged or forced into the wrong receptacle on the unit.
3.29.6 | Transducer holders must be provided
3.30 | Other system requirements
3.30.1 | The unit must operate off 220V, 50Hz power supply
3.30.2 | All metal parts must be earthed.
3.30.3 | Mains cable at least three metres long with SABS approved 15A three prong plug.
3.30.4 | Must be fitted with rechargeable batteries
3.30.5 | The unit must be fitted with an internal battery charger
3.30.6 | Indicator for battery status
3.30.7 | Must store trends for at least 8 hours
3.30.8 | A cart/trolley must be supplied
3.30.9 | The cart must have four castor wheels, two of them with brakes
3.30.10 | The cart must have a basket/holder for accessories
3.30.11 | The unit must be handed over in full operating order
3.30.12 | A starter-pack of consumables must be supplied with the unit.
4 | Standard Accessories and Consumables
4.1 | US transducers x2
4.2 | Toco transducer
4.3 | Thermal printing paper (Pk of 5)
4.4 | Belts x3 (For each transducer)
5 | Conformity Compliance (Please attach certificate)
5.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
5.2 | OEM must comply and certified on ISO 9001 quality standards, attach proof of compliance
5.2.1 | Please provide unique ref number of the ISO 9001 ceriticate:
5.3 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
5.3.1 | Please provide unique ref number of the ISO 13485 ceriticate:
5.4 | Model quoted for must be EC certified. Attach a copy of certification
5.4.1 | Please provide unique ref number of the EC ceriticate:
5.5 | Hazardous Substance Act: If this type of equipment / apparatus appears on the schedule of Hazardous substances issued by the Directorate: Health Technology of the Department of Health, a license in terms of the Act on Hazardous Substances (Act 15/1973) must be submitted with the bid / quotation documents. The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.6 | All electrical/electronic medical equipment must be licensed by Radiation Control, where applicable a copy of the license must be submitted The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.6.1 | Radiation Control License number:
5.7 | The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade (rubber to with 2 screws) type and it must be a minimum length of three (3) meters. NB the mains cable of the unit tendered for must be SABS colour coded.
5.8 | The equipment quoted for must be protected against electromagnetic interference IEC 60601-1-2
6 | COMPREHENSIVE MAINTENANCE
6.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) must be included in the bid price
6.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
6.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
6.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
6.3 | The bidder to provide the OEM recommended service interval of the unit:
6.4 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
7 | General
7.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
7.2 | The unit on offer must be the latest technology.
Cardiotocograph Accessories and Consumables
RT2-04-03a | US transducer
RT2-04-03b | Toco transducer
RT2-04-03c | Thermal printing paper (Pk of 5)
RT2-04-03d | Belt
RT2-04-03e | Trolley
Bidder to list any other Consumables or Accessories (Not listed above) for the model on offer below

## Cardiotocographs, Twins + MHR
RT2-04-04 | Cardiotocograph (Twins) with MHR, on Trolley
1 | Description
2 | Product Offered:
2.1 | Make:
2.2 | Model:
2.3 | Initial Manufacturing year of Model:
2.4 | Radiation Control License number:
3 | Technical Specification
3.1 | Required for external monitoring of fetal heart rate (FHR), and external monitoring of uterine activity of pregnant patients + must include monitoring and display of maternal heart rate
3.2 | Twins monitoring is required
3.3 | Must perform automatic self-check when switched on
3.4 | There must a visible Fetal Heart Beat indicator which will flash with each detected heart beat
3.5 | The unit must provide an audio signal, with variable volume, for every heart beat detected
3.6 | The fetal heart rate must be measured by means of wide beam pulsed Doppler ultrasound, with auto correlation processing.
3.7 | Must be able to detect and display multiple FHR while discriminating between different foetuses
3.8 | The oscillator frequency < 1,2 MHz.
3.9 | Ultrasound intensity must be less than 10 mW/cm2
3.10 | The transducer must be manufactured of high-impact plastic that is robust and watertight.
3.11 | The belts and buckles supplied for positioning of the transducer must allow easy angling for optimum positioning.
3.12 | Adjustable baseline position
3.13 | A test button must be provided that produces a series of pulses for calibration which must be recorded and displayed.
3.14 | Display: > 16 cm (corner to corner).
3.15 | Displays must be clearly visible under all lighting conditions.
3.16 | Must allow Maternal Heart Rate (MHR) and Fetal Heart Rate(FHR) coincidence detection?
3.17 | The following parameters must be displayed:
3.18 | Fetal heart rate approximately 50 to 200 beats per minute
3.19 | Uterine activity: approximately 0 to 100 relative units.
3.20 | Heart Rate monitoring
3.21 | Fetal heart rate must update once per second when FHR is greater than 60 and once each beat when FHR is less than 60.
3.22 | Maternal heart rate must be obtained, and displayed, reliably using preferably ECG or plethysmography
3.23 | The monitor must have the following controls:
3.23.1 | Volume Up/Down
3.23.2 | TOCO/IUP Zero
3.23.3 | Printer On/Off
3.23.4 | Printer fast forward
3.24 | Must have audible and visual alarms for low/high FHR
3.25 | Electronic storage data capacity.
3.26 | RECORDER
3.26.1 | Integrated dual channel thermal recorder
3.26.2 | Heart rate recording must be available.
3.26.3 | All dot print test
3.26.4 | The strip chart paper must be easy to load
3.27 | Fetal heart rate:
3.27.1 | Range: approximately 50 to 210 beats per minute.
3.27.2 | Vertical chart: at least 7 cm.
3.27.3 | Vertical scale sensitivity: approximately 20 bpm per cm.
3.27.4 | Resolution: 1 bpm
3.28 | Uterine activity:
3.28.1 | Range: 0 – 100 units
3.28.2 | Vertical chart: at least 4 cm.
3.28.3 | Vertical scale sensitivity: approximately 25 relative units per cm
3.28.4 | Resolution: 1 relative unit
3.28.5 | Paper chart speed must be selectable for 1, 2 and 3 cm per minute.
3.28.6 | The time and date must be printed by the recorder at least every 10 minutes.
3.28.7 | Recording time must be at least 24 hours at 1 cm per minute.
3.28.8 | The unit must have a facility for remote event marking
3.29 | TRANSDUCERS
3.29.1 | 2x Ultrasound and toco transducers must be supplied
3.29.2 | All transducers must be of very robust design. Bidders must guarantee shock resistance on the transducers e.g. will withstand a drop from one metre high onto a hard surface
3.29.3 | Bidders must state whether the ultrasound transducer is repairable and which parts are replaceable or whether it is a sealed unit that must be replaced when faulty.
3.29.4 | The ultrasound transducer cable must be a minimum of 2 m long
3.29.5 | Connecting plug must be keyed and colour coded to avoid it being plugged or forced into the wrong receptacle on the unit.
3.29.6 | Transducer holders must be provided
3.30 | General
3.30.1 | The unit must operate off 220V, 50Hz power supply
3.30.2 | All metal parts must be earthed.
3.30.3 | Mains cable at least three metres long with SABS approved 15A three prong plug.
3.30.4 | Must be fitted with rechargeable batteries
3.30.5 | The unit must be fitted with an internal battery charger
3.30.6 | Indicator for battery status
3.30.7 | Must store trends for at least 8 hours
3.30.8 | A cart/trolley must be supplied
3.30.9 | The cart must have four castor wheels, two of them with brakes
3.30.10 | The cart must have a basket/holder for accessories
3.30.11 | The unit must be handed over in full operating order
3.30.12 | A starter-pack of consumables must be supplied with the unit.
4 | Standard Accessories
4.1 | US transducers x2
4.2 | Toco transducer
4.3 | Thermal printing paper (Pk of 5)
4.4 | Belts x3 (For each transducer)
5 | Conformity Compliance (Please attach certificate)
5.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
5.2 | OEM must comply and certified on ISO 9001 quality standards, attach proof of compliance
5.2.1 | Please provide unique ref number of the ISO 9001 ceriticate:
5.3 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
5.3.1 | Please provide unique ref number of the ISO 13485 ceriticate:
5.4 | Model quoted for must be EC certified. Attach a copy of certification
5.4.1 | Please provide unique ref number of the EC ceriticate:
5.5 | Hazardous Substance Act: If this type of equipment / apparatus appears on the schedule of Hazardous substances issued by the Directorate: Health Technology of the Department of Health, a license in terms of the Act on Hazardous Substances (Act 15/1973) must be submitted with the bid / quotation documents. The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.6 | All electrical/electronic medical equipment must be licensed by Radiation Control, where applicable a copy of the license must be submitted The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.6.1 | Radiation Control License number:
5.7 | The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade (rubber to with 2 screws) type and it must be a minimum length of three (3) meters. NB the mains cable of the unit tendered for must be SABS colour coded.
5.8 | The equipment quoted for must be protected against electromagnetic interference IEC 60601-1-2
6 | COMPREHENSIVE MAINTENANCE
6.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) must be included in the bid price
6.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
6.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
6.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
6.3 | The bidder to provide the OEM recommended service interval of the unit:
6.4 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
7 | General
7.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
7.2 | The unit on offer must be the latest technology.
Cardiotocograph Accessories and Consumables
RT2-04-04a | US transducer
RT2-04-04b | Toco transducer
RT2-04-04c | Thermal printing paper (Pk of 5)
RT2-04-04d | Belt
RT2-04-04e | Trolley
Bidder to list any other Consumables or Accessories (Not listed above) for the model on offer below

## Fetal Doppler
RT2-04-05 | Doppler Fetal Heart Rate Monitor
1 | Description
1.1
2 | Product Offered:
2.1 | Make:
2.2 | Model:
2.3 | Initial Manufacturing year of Model:
2.4 | Radiation Control License number:
3 | Technical Specification
3.1 | The unit must be a portable Doppler for measuring fetal heart rate.
3.2 | The unit must be able to accommodate two types of obstetric probes – 2 MHz and 3 MHz
3.3 | The 2 MHz probe must be supplied with the unit
3.4 | The probes must be robust and well constructed and the crystals of the transducer must be well protected against accidental droppage.
3.5 | Fetal heart rate measuring range: approximately 60 - 220 beats per minute.
3.6 | Resolution: 1 bpm
3.7 | Accuracy +- 2 bpm
3.8 | The following must be displayed:
3.8.1 | Heart rate
3.8.2 | Heartbeat synchronous indicator
3.9 | Low battery indicator
3.10 | Battery charging indicator
3.11 | The power output of the obstetric probes must not exceed 10 mW/cm².
3.12 | The unit must operate with rechargable batteries
3.13 | Battery and charger must be supplied.
3.14 | Battery charger must operate from a 220 V, 50 Hz power supply
3.15 | Batteries must have capacity for more than 250 examinations of 1 minute each
3.16 | Mains cable at least three metres long with SABS approved 15A three prong plug.
3.17 | The handheld unit must be lightweight, approximately 2 kg
3.18 | The unit must be equipped with a high quality speaker
3.19 | The unit must accommodate an external headset. Headset may be offered as an option
3.20 | If the headset is used for the audio output, the internal loudspeaker of the unit must be disabled.
3.21 | The volume of the audio output must be adjustable.
3.22 | The unit must have a clear LED or LCD display showing fetal heart rate
3.23 | Low battery indication
3.24 | The unit must automatically shut down after one minute of inactivity
3.25 | The unit must have a storage holder for the Doppler probes.
3.26 | The unit must be handed over in full operating order
3.27 | A starter-pack of consumables must be supplied with the unit.
4 | Standard Accessories
4.1 | 2 MHz probe
4.2 | Carry bag
5 | Conformity Compliance (Please attach certificate)
5.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
5.2 | OEM must comply and certified on ISO 9001 quality standards, attach proof of compliance
5.2.1 | Please provide unique ref number of the ISO 9001 ceriticate:
5.3 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
5.3.1 | Please provide unique ref number of the ISO 13485 ceriticate:
5.4 | Model quoted for must be EC certified. Attach a copy of certification
5.4.1 | Please provide unique ref number of the EC ceriticate:
5.5 | Hazardous Substance Act: If this type of equipment / apparatus appears on the schedule of Hazardous substances issued by the Directorate: Health Technology of the Department of Health, a license in terms of the Act on Hazardous Substances (Act 15/1973) must be submitted with the bid / quotation documents. The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.6 | All electrical/electronic medical equipment must be licensed by Radiation Control, where applicable a copy of the license must be submitted The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.6.1 | Radiation Control License number:
6 | COMPREHENSIVE MAINTENANCE
6.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) must be included in the bid price
6.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
6.2 | The bidder to provide the OEM recommended service interval of the unit:
6.3 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
7 | General
7.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
7.2 | The unit on offer must be the latest technology.
Doppler Fetal Heart Rate Monitor Accessories
RT2-04-05a | 2 MHz probe
RT2-04-05b | Carry bag
RT2-04-05c | A 3 MHz probe may be offered as an option
RT2-04-05d | Headset
Bidder to list any other Consumables or Accessories (Not listed above) for the model on offer below

## Video Colposcope
RT2-04-06 | VIDEO COLPOSCOPE
1 | DESCRIPTION
1.1 | The unit must be designed for gynecological applications and it must aid in the detection of carcinogenic conditions
2 | Product Offered:
2.1 | Make:
2.2 | Model:
2.3 | Initial Manufacturing year of Model:
3 | TECHNICAL SPECIFICATION
3.1 | The unit offered must be supplied complete with camera on a vertical mobile stand, monitor mounted on a trolley, Image storage with a minimum of 120GB, Keyboard and mouse, and Printer
3.2 | Camera
3.2.1 | 3 level electronic green filter LED cold lighting system to enhance vascular image display without compromise of illumination
3.2.2 | The unit must have a fast auto focusing system to ensure clear images
3.2.3 | The unit must have a remote capture control using a mouse or keyboard or foot paddle
3.2.4 | The unit must have a real time display of magnification
3.2.5 | Must have an angle adjustment function
3.2.6 | Must be equipped with Image freeze control button
3.2.7 | Precise vertical and horizontal positioning with adjustable height
3.2.8 | Adjustable zoom, focus and field view
3.3 | Monitor
3.3.1 | It should be atleast 23" Flat LCD or LED medical grade monitor with the following specifications:
3.3.1a | Resolution of at least 1920 x 1080 pixels
3.3.1b | The unit offered should have a minimum of 3 inputs which include; Video, Y/C, RGB, DVI, HDMI
3.3.1c | High contrast and wide viewing angle
3.4 | Image Management System
3.4.1 | The unit offered should record both video and still images in full HD
3.4.2 | Internal memory should be a minimum of 120GB
3.4.3 | button to record and capture functions
3.4.4 | The system offered should be compatible with the printer offered so that the recorded items can be printed out.
3.4.5 | The video colpscope device must be equipped with atleast intergrated USB port
3.5 | Power requirements
3.5.1 | The unit must operate from a 220 V, 50 Hz. supply
3.5.2 | The mains cable at least three metres long with SABS approved 15A three prong plug.
3.6 | Trolley
3.6.1 | The offered unit should come in a purpose built Colposcope System trolley with a minimum of one drawer for storage
3.6.2 | The trolley offered should have a tray for the keyboard and mouse.
3.6.3 | The trolley should have antistatic castors with at least 2 equipped with locking system/brake.
3.6.4 | The trolley should have a handle for ease access and transportation and a cable management arm.
3.6.5 | The trolley must be supplied with a minimum of 3 additional 220 V power sockets.
3.7 | IT NETWORKING AND CONNECTIVITY
3.7.1 | An RS232/RJ45 or any other latest data output port with a freely available data format for connection to hospital information system shall be provided
4 | STANDARD ACCESSORIES and CONSUMABLES
4.1 | Mouse
4.2 | Printer
4.3 | Keyboard
4.4 | A4 or A6 paper (1000 sheets)
4.5 | PC software
4.6 | Dust cover must b included and supplied for the colposcope system offered
5 | Conformity Compliance (Please attach certificate)
5.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
5.2 | OEM must comply and certified on ISO 9001 quality standards, attach proof of compliance
5.2.1 | Please provide unique ref number of the ISO 9001 ceriticate:
5.3 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
5.3.1 | Please provide unique ref number of the ISO 13485 ceriticate:
5.4 | Model quoted for must be EC certified. Attach a copy of certification
5.4.1 | Please provide unique ref number of the EC ceriticate:
5.5 | Hazardous Substance Act: If this type of equipment / apparatus appears on the schedule of Hazardous substances issued by the Directorate: Health Technology of the Department of Health, a license in terms of the Act on Hazardous Substances (Act 15/1973) must be submitted with the bid / quotation documents. The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.7 | All electrical/electronic medical equipment must be licensed by Radiation Control, where applicable a copy of the license must be submitted The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.7.1 | Radiation Control License number:
5.8 | The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade (rubber to with 2 screws) type and it must be a minimum length of three (3) meters. NB the mains cable of the unit tendered for must be SABS colour coded.
5.8 | The equipment quoted for must be protected against electromagnetic interference IEC 60601-1-2
6 | COMPREHENSIVE MAINTENANCE
6.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) must be included in the bid price
6.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
6.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
6.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
6.3 | The bidder to provide the OEM recommended service interval of the unit:
6.4 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
7 | General
7.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
7.2 | The unit on offer must be the latest technology.

## Vacuum Extractor
RT2-04-07 | VACUUM EXTRACTOR, OBSTETRIC - PORTABLE ELECTRIC
1 | DESCRIPTION
1.1 | A device for applying traction to the fetus during delivery by using a suction cup attached to the fetal head
2 | Product Offered:
2.1 | Make:
2.2 | Model:
2.3 | Initial Manufacturing year of Model:
3 | SPECIFICATIONS
3.1 | The unit must be designed for performing vacuum (ventonse) extraction of a baby in order to assist with childbirth.
3.2 | The suction unit being tendered on must be lightweight, easily portable, it must employ a reliable robust electric vacuum pump and it must be supplied with all the necessary accessories as part and parcel of the tendered system. The pump must be of the oil-less maintenance free piston type.
3.3 | The unit must have a strong durable A.B.S. or equivalent corrosion resistant case with a handle.
3.4 | There must be a Jar Holder or Clip for a Jar to be fixed to the suction unit cage.
3.5 | The unit must have a graduated, polycarbonate collection jar with a minimum capacity of 1000cc with a float assembly in the jar lid. The jar must be suitable for use with disposable liners
3.6 | The collection jar must be autoclavable or other means of sterilization, please specify if not autoclavable
3.7 | The unit must have a vacuum control to adjust the vacuum over a range of minus 20Kpa to minus 90Kpa.
3.8 | The maximum vacuum must not be less than minus 90Kpa with a flow rate of not less than 14 to 15Lpm. N.B. The bidder must state the maximum vacuum and flow rate on the unit offered.
3.9 | The unit must have anti-bacterial filters fitted at the inlet side and exhaust side of the pump. These filters must be fitted where it is conveniently located such that they are easily accessible to the clinical staff for changing.
3.10 | In addition to the two normal filters specified above, in line shut-off filter must be supplied and fitted to the unit offered inside the housing of the unit and just before the inlet to the pump.
3.11 | The unit must have an ON / OFF switch fitted to the case.
3.12 | The vacuum gauge must be fitted to the pump case near the vacuum control knob.
3.13 | The motor must be 220 Volt + 10%, 50 Hz single phase a.c.
3.14 | The motor must be protected by an inline circuit interrupter
3.15 | The noise level of the unit offered must not exceed 60dB.
3.16 | A complete set of re-usable silicone extractor cups with stainless steel handles and fittings such as chains, plates, filters and appropriate length tubing for each with cup sizes and quantity as follows: 1 x 70mm, 2 x 60mm, 2 x 50mm and 1 x 40mm must be supplied at no extra cost to the final bid price.
3.17 | A completed set of silicone extractor cups with all fittings and cup sizes and quantity as follows: 1 x 40mm, 2 x 50mm, 2 x 60mm must be supplied at no extra cost to the final bid price
3.18 | Please state the sterilization method that must be used to sterilize the abovementioned accessories in (2.16 and 2.17).
3.19 | All parts must be of a durable material construction and they must be corrosion resistant.
4 | STANDARD ACCESSORIES
4.1 | 1 x 40mm extractor cup
4.2 | 2 x 50mm extractor cups
4.3 | 2 x 60mm extractor cups
4.4 | Polycarbonate Jar
5 | Conformity Compliance (Please attach certificate)
5.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
5.2 | OEM must comply and certified on ISO 9001 quality standards, attach proof of compliance
5.2.1 | Please provide unique ref number of the ISO 9001 ceriticate:
5.3 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
5.3.1 | Please provide unique ref number of the ISO 13485 ceriticate:
5.4 | Model quoted for must be EC certified. Attach a copy of certification
5.4.1 | Please provide unique ref number of the EC ceriticate:
5.5 | Hazardous Substance Act: If this type of equipment / apparatus appears on the schedule of Hazardous substances issued by the Directorate: Health Technology of the Department of Health, a license in terms of the Act on Hazardous Substances (Act 15/1973) must be submitted with the bid / quotation documents. The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.6 | All electrical/electronic medical equipment must be licensed by Radiation Control, where applicable a copy of the license must be submitted The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.6.1 | Radiation Control License number:
5.7 | The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade (rubber to with 2 screws) type and it must be a minimum length of three (3) meters. NB the mains cable of the unit tendered for must be SABS colour coded.
5.8 | The equipment quoted for must be protected against electromagnetic interference IEC 60601-1-2
6 | COMPREHENSIVE MAINTENANCE
6.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) must be included in the bid price
6.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
6.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
6.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
6.3 | The bidder to provide the OEM recommended service interval of the unit:
6.4 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
7 | General
7.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
7.2 | The unit on offer must be the latest technology.
VACUUM EXTRACTOR, OBSTETRIC ACCESSORIES
RT2-04-07a | 1 x 40mm extractor cup
RT2-04-07b | 2 x 50mm extractor cups
RT2-04-07c | 2 x 60mm extractor cups
RT2-04-07d | Mobile trolley
Bidder to list any other Consumables or Accessories (Not listed above) for the model on offer below

## Ultrasound, Obstetric
RT2-04-08 | ULTRASOUND UNIT , OBSTETRIC
1 | Description
2 | Product Offered:
2.1 | Make:
2.2 | Model:
2.3 | Initial Manufacturing year of Model:
2.4 | Radiation Control License number:
3.1 | Technical Specification
3.1.1 | The unit shall be a fully digital unit with a memory beam former and processor to generate sharp defined ultrasound images, reduce artefacts and improve image quality.
3.1.2 | It shall utilise the latest software to perform investigations.
3.1.3 | High image quality will be one of the main criteria that will be used to evaluate the offers.
3.1.4 | Although offered units might comply with all the specifications, image quality shall be a critical evaluation criterion, specifically with increased BMI cases.
3.1.5 | The unit must be of the latest technology.
3.1.6 | It must be a state of the art fully digital phase array ultrasound unit.
3.1.7 | The unit will be used for general applications as a multi-purpose unit for Abdominal, Vascular, Obstetric, Gynaecology, Musculo-Skeletal, Small Parts and Paediatric.
3.2 | IMAGING MODALITIES
3.2.1 | The unit must be latest technology; please state year in which this unit was first manufactured.
3.2.2 | The unit shall offer excellent high image quality in B mode and fully digital colour flow Doppler. Colour Doppler should have no delay after activation.
3.2.3 | The system Dynamic Range shall not be less than 172 dB. The Dynamic range should be able to be changed minimum of 4dB increments.
3.2.4 | The unit shall have direct storage on hard drive of not less than 80 Gig. State all other storage facility that is included in the offered price. (Such as CD, DVD, MOD). Recorded CD roms shall be compatible with any standard PC, without requiring extra software. USB download of images from the system must be possible
3.2.5 | A digital black and white thermal printer shall be included in the offer for the machine.
3.2.6 | The system shall offer a Digital Continuous Beam Former to generate sharp defined ultrasound beams, to reduce artefacts and to provide more consistent overall image quality.
3.2.7 | The unit shall offer multi-frequency on all probes to provide sharp reproducible image quality in both superficial and deep anatomy.
3.2.8 | The unit shall be equipped with a non-interlaced colour monitor with a screen size of not less than 15 inch, measured from corner to corner. The monitor shall be of the flat screen TFT LCD type.
3.2.9 | The unit should have at least two (2) active probe connectors with electronic switching.
3.2.10 | The unit shall be able to offer scanning methods for Electronic Convex and Linear technology.
3.2.11 | The unit shall offer depth gain in not less than 8 steps. Unit must be programmable to start up automatically at a depth of operator’s choice and programmable to start up at different depths for specific examinations.
3.2.12 | Image processing shall be done via a rotary dial or similar technology with at least 5 combinations for optimising the best image quality for each patient’s physique (user programmable). Image processing should visibly improve image and not just change it.
3.2.13 | Pan-Zoom enlargement shall be in real and frozen mode up to 10X zoom.
3.2.14 | A colour pallet shall be available for the colourization of the B mode and M mode and PW Doppler trace.
3.2.15 | Independent M mode and PW Doppler gain adjustment shall be available.
3.2.16 | Sweep speed shall be adjustable in at least 4 steps.Unit shall have selectable sweep rates of 1,2,3,4,8 s/screen
3.2.17 | The unit shall offer Tissue Harmonics as standard in the system.
3.2.18 | Grey scale level shall be at least 256 shades.
3.2.19 | The unit shall have not less than 1024 processing channels.
3.2.20 | The unit shall be housed in a robust mobile cart. The cart shall house all elements of the unit, and shall have enough dedicated storage space for the probes, and all space for all disposables needed on the machine. The unit to be used for mobile ultrasound in different areas. a) Width of system in centimetres b) Depth of system in centimetres.
3.2.21 | The frame rate shall not be less than 394 frames per second. Indicate the frame rate that the system offers
3.2.22 | Image depth shall be not less than 30cm.
3.2.23 | The unit should offer quick scan to automatically adjusts STC and gain through one-touch operation. This must visibly improve the image
3.2.24 | Post processing must be available on frozen images
3.2.25 | Unit must offer user programmable annotation for quick labelling. User must be able to program approximately 40 words of their own choice per selected examination type for annotation purposes
3.2.26 | Gestational age according to LMP must be displayed on the working screen (not just on the patient information page)
3.3 | DOPPLER
3.3.1 | PW Doppler shall be offered on all imaging probes.
3.3.2 | Max detectable velocity for PW shall be stated.
3.3.3 | Pulse Repetition Frequency shall vary between 0,7 – 19,8 kHz pulsed wave and 2,1 – 40kHz continuous wave
3.3.4 | Independent angle steering up to ±20° for colour and Doppler shall be offered.
3.3.5 | Independent gain control for PW shall be available.
3.3.6 | Colour pallet offering 7 types independent of B mode and Pulse Doppler pallet to be available.
3.3.7 | Sampling with Pulse Doppler shall vary between 1 and 20 mm.
3.3.8 | Gate size must vary at least between 1 and 8mm.
3.3.9 | Base line shift must be possible in at least 8 steps.
3.3.10 | Spectral invert must be possible.
3.3.11 | At least 6 wall filter selections shall be available. Bidder to state minimum.
3.3.12 | Accurate autotrace function for both arterial and venous waveforms shall be available.
3.4 | COLOUR DOPPLER
3.4.1 | A minimum of 8 colour maps shall be available for colour display.
3.4.2 | Colour maps shall be available to be selected in real time and freeze mode.
3.4.3 | Independent colour gain adjustment shall be from 1 to 16 dB in 0,5 dB steps. Indicate range
3.4.4 | Maximum colour frame rate shall not be less than 285 frames per second.
3.4.5 | Colour flow shall operate to the full depth of 30 cm.
3.4.6 | Colour maps and processing selections shall be able to be selected by the user and stored in the presets.
3.4.7 | Quad signal processing and parallel signal processing shall be available. Bidders to supply details.
3.4.8 | Power Doppler on all transducers shall be available.
3.4.9 | The unit offered shall be able of real time triplex mode.
3.4.10 | Colour mapping must be instantaneous – no time delay /lag.
3.5 | IMAGE MEMORY
3.5.1 | Image memory to operate in both dual and quad screen format.
3.5.2 | ECG synchronised recording shall be possible. (optional)
3.5.3 | It shall be possible to record B loop up to 500 frames.
3.5.4 | Cine loop and measurements must be active simultaneously.
3.5.5 | M/D review of a minimum of 15 seconds shall be possible.
3.5.6 | B Mode Imaging on linear transducer to be electronically steerable.
3.6 | TRANSDUCERS
3.6.1 | The following probes shall be included for the machine: o Curvi-linear probe (Abdominal, vascular, O&G applications- 2.0 to 7.0 Mhz) o Linear probe - 4.5 to 16 Mhz o Transvaginal probe(1600 FOV) – 3.9 to 15.0 Mhz A quote that clearly shows that all the above probes are included in the offer, shall form part of the offer. Failure of showing this detail will render the offer not compliant.
3.6.2 | All imaging transducers shall be multi-frequency.
3.6.3 | Tissue Harmonics shall be available on all offered transducers.
3.6.4 | The selected frequency of the imaging transducer shall be displayed on the screen.
3.6.5 | All transducers shall be digitally controlled so continuous beam forming is possible. State number of focal zones.
3.6.6 | A minimum of 2 active probes shall be available. It shall be possible to electronically switch between these 2 probes.
3.7 | MEASUREMENTS
3.7.1 | Patient data MUST offer the option to enter previous ultrasound date and measurements, (not just LMP from previous examination).
3.7.2 | Gestational age according to LMP must be displayed on working screen, not just not patient information page.
3.7.3 | On-screen measurements must be possible from real time. 16 callipers (8 channel)each individually numbered for measuring distance/volume/angle and area must be offered
3.7.4 | Full range of obstetrics measurements and weight charts which shall include Hadlock, Chitty and others.
3.7.5 | The unit shall offer a storage facility for up to 300 patient foetal growth measurement data. It shall be possible to retrieve the data for follow-ups.
3.7.6 | The unit shall offer 50th percentile growth graphs of the Doppler measurements RI/RP. It shall be possible to retrieve the data for follow-ups.
3.7.7 | 50th Percentile growth graphs shall be offered on all measurements for example BPD,HC, AC, FL Fetal weight . It shall be possible to retrieve the data for follow-ups.
3.7.8 | The unit shall offer user programmable annotation for quick labelling.
3.7.9 | State if the unit comply to the following: 1. Panoramic view. 2. Quad signal processing
3.7.10 | All software upgrades shall be offered optionally, with full description what improvement the software upgrade will bring to the offered unit.
3.7.11 | The unit shall be supplied with an uninterrupted power supply (UPS), to be able to keep the unit running for at least two hours in case of power failure.
3.8 | DICOM COMPLIANCE
3.8.1 | The offered unit shall be fully DICOM compliant.
3.8.2 | The DICOM licence, as well as connecting the ultrasound to the envisaged PACS shall be included in the offer.
3.8.3 | Please note that installation of the PACS at False Bay Hospital is planned for this year, and that the successful contractor shall assist in setting up the machine with the PACS.
4 | Standard Accessories and Consumables
4.1 | Curvi-linear probe (Abdominal, vascular, O&G applications - 2.0 to 7.0 MHz)
4.2 | Linear probe - 4.5 to 16 MHz
4.3 | Transvaginal probe - 3.9 to 15.0 MHz
4.4 | Printer
4.5 | Ultrasound Gel
4.6 | Uninterrupted Power Supply (UPS)
4.7 | Printing paper (1 box)
5 | Conformity Compliance (Please attach certificate)
5.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
5.2 | OEM must comply and certified on ISO 9001 quality standards, attach proof of compliance
5.2.1 | Please provide unique ref number of the ISO 9001 ceriticate:
5.3 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
5.3.1 | Please provide unique ref number of the ISO 13485 ceriticate:
5.4 | Model quoted for must be EC certified. Attach a copy of certification
5.4.1 | Please provide unique ref number of the EC ceriticate:
5.5 | Hazardous Substance Act: If this type of equipment / apparatus appears on the schedule of Hazardous substances issued by the Directorate: Health Technology of the Department of Health, a license in terms of the Act on Hazardous Substances (Act 15/1973) must be submitted with the bid / quotation documents. The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.6 | All electrical/electronic medical equipment must be licensed by Radiation Control, where applicable a copy of the license must be submitted The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.6.1 | Radiation Control License number:
5.7 | The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade (rubber to with 2 screws) type and it must be a minimum length of three (3) meters. NB the mains cable of the unit tendered for must be SABS colour coded.
5.8 | The equipment quoted for must be protected against electromagnetic interference IEC 60601-1-2
5.9 | The unit must comply with electrical safety and performance standard such as IEC 60601-2-37, attached certification
5.9.1 | Please provide unique ref number of the IEC ceriticate:
6 | COMPREHENSIVE MAINTENANCE
6.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) must be included in the bid price
6.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
6.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
6.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
6.3 | The bidder to provide the OEM recommended service interval of the unit:
6.4 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
7 | General
7.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
7.2 | The unit on offer must be the latest technology.
ULTRASOUND UNIT ACCESSORIES
RT2-04-08a | Curvi-linear probe (Abdominal, vascular, O&G applications - 2.0 to 7.0 MHz)
RT2-04-08b | Linear probe - 4.5 to 16 MHz
RT2-04-08c | Transvaginal probe - 3.9 to 15.0 MHz
RT2-04-08d | Printer
RT2-04-08e | Ultrasound Gel
RT2-04-08f | Uninterrupted Power Supply (UPS)
Bidder to list any other Consumables or Accessories (Not listed above) for the model on offer below

## Category-5 Monitors & General
GENERAL MEDICAL EQUIPMENT
# | Item number | Item Description
4 | RT2-05-04 | Defibrilator, ECG&Pacing
5 | RT2-05-05 | Defibrilator, ECG&SpO2
6 | RT2-05-06 | Defibrilator, ECG,SpO2&NIBP
7 | RT2-05-07 | ECG Machine, 12-lead
9 | RT2-05-09 | Examination Light, LED, Mobile
10 | RT2-05-10 | Examination Light, LED, Mobile with battery
11 | RT2-05-11 | Examination Light, LED, Wall-mount
14 | RT2-05-14 | Monitor, Basic
15 | RT2-05-15 | Monitor, ICU
16 | RT2-05-16 | Monitor, ICU with Transport Module
17 | RT2-05-17 | Central Monitoring, 8-bed package
18 | RT2-05-18 | Central Monitoring Station, ICU
19 | RT2-05-19 | Vital Signs Monitor with oximetery & Temp
20 | RT2-05-20 | Pulse Oximeter, Tabletop
21 | RT2-05-21 | Suction Unit, Wall-mount
22 | RT2-05-22 | Suction Machine, Mobile
23 | RT2-05-23 | Theater Light, Mobile
25 | RT2-05-25 | Fluid Warmer, Cabinet, 40-60L
26 | RT2-05-26 | Fluid Warmer, Cabinet, 80-150L
27 | RT2-05-27 | Fluid Warmer, Cabinet, 200-300L
28 | RT2-05-28 | Fridge, Vaccine, 100-150L
29 | RT2-05-29 | Fridge, Vaccine, 200-250L
30 | RT2-05-30 | Fridge, Vaccine, 300-350L
31 | RT2-05-31 | Fridge, Medicine, 200-270L
32 | RT2-05-32 | Fridge, Medicine, 900-1100L
33 | RT2-05-33 | Ultrasound, Cardiology

## Defib, wit ECG&Pacing
RT2-05-04 | Defibrillator, with ECG & Pacing
1 | Description
1.1 | Defibrillator to reverse fibrillation to a normal rhythm with pacing capabilties and ECG monitoring
2 | Product Offered:
2.1 | Make:
2.2 | Model:
2.3 | Initial Manufacturing year of Model:
2.4 | Radiation Control License number:
3 | Technical Specification
3.1 | The defibrillator offered must provide ECG monitoring, a printer and external pacing
3.2 | The defibrillator, monitor, printer, pacing unit, battery and battery-charger must be integrated into a single unit.
3.3 | The defibrillator must deliver a bi-phasic impulse
3.4 | The Defib paddles must be detacheable from the cable or machine to exchange with AED pads
3.5 | The defibrillator must have manual and automatic (AED) modes
3.6 | Automatic mode must measure patient impedance and deliver an appropriate impulse automatically generated, between 0 and 250 J
3.7 | The unit must offer audible/voice and on-screen messages for use during AED mode
3.8 | Manual mode energy selection: 5 - 200 J
3.9 | Selection of manual energy must be simple with a clear unambiguous indication of the energy selected
3.10 | The defibrillator must have the following selection of defibrillation:
3.11 | a. Non-synchronous defibrillation
3.12 | b. Synchronous defibrillation
3.13 | The defibrillator must indicate if the synchronous mode has been selected.
3.14 | The defibrillator must default to non-synchronous mode at initial switch on.
3.15 | Synchronous defibrillation detection must be possible via ECG patient cable and via paddles
3.16 | The synch marker must be clearly indicated on the ECG trace.
3.17 | The defibrillator must operate from a 220 V, 50 Hz. supply
3.18 | The defibrillator must contain a rechargeable battery which charges automatically when the defibrillator is plugged into the 220 V, 50 Hz supply.
3.19 | The battery must charge whether the defibrillator is switched on or not
3.20 | The defibrillator must have a visual indicator to indicate that the mains supply is switched on.
3.21 | The internal rechargeable battery must be able to deliver at least 50 defibrillations at the maximum setting and provide a minimum of 2 hours ECG monitoring in the event of a mains power failure
3.22 | Battery status indicator with audible and visual alarm for low battery
3.23 | Audible tone while charging the capacitor
3.24 | Audible tone, different from above, to indicate full charge of capacitor
3.25 | Visual indication of fully charged state
3.26 | Defibrillating charge time to maximum setting must not exceed 10 seconds.
3.27 | In the event of a completely depleted battery, the defibrillator must be fully operable on the 220 V, 50 Hz supply.
3.28 | The defibrillator must do an automatic self-test when switched on.
3.29 | Override of self test must be possible
3.30 | Operation of the defibrillator must be user-friendly with clear indications for:
3.30.1 | Switch on;
3.30.2 | Select energy;
3.30.3 | Charge;
3.30.4 | Discharge.
3.31 | The defibrillator must be portable and weigh less than 7 kg with all accessories attached
3.22 | The unit must have a carry handle
3.13 | The defibrillator must incorporate an internal 50 ohms test load to allow a check for satisfactory functioning.
3.4 | The defibrillator must incorporate an emergency dump circuit that will automatically provide a safe current path for unused energy when the paddles are charged but not discharged within sixty seconds.
3.5 | The defibrillator must be able to do both internal and external defibrillation.
3.14 | Energy selection must accommodate patients from paeds to adults for both internal and external defibrillation.
3.23 | The defibrillator must automatically sense when internal paddles are in use and restrict the maximum selectable energy to 50 Joules.
3.32 | Apex paddle to incorporate a push button to charge the defibrillator.
3.41 | Energy discharge via defibrillator paddles to patient or test load must be activated by pressing buttons on both apex and sternum paddles simultaneously.
3.50 | Energy discharge via AED pads or internal electrodes by means of a button on the panel
3.59 | ECG monitoring must be possible through either the ECG leads or the paddles and unit must display the ECG lead selected (I, II or III)/(aVl, aVR or aVF) or indicate whether the paddles are the source of the ECG
3.59.1 | Size of ECG: 0,5; 1; 2 and 4 cm/mV
3.59.2 | Sweep rate: 25 mm/s
3.59.3 | 50 Hz filter
3.59.4 | Heart rate: 30 - 300 bpm
3.59.5 | Analysis time: maximum 15 s
3.59.6 | Artefact detection
3.59.7 | Pacemaker detection
3.60 | The following alarms must be displayed:
3.60.1 | low battery,
3.60.2 | no paper (printer)
3.60.3 | ECG lead off/fault.
3.61 | Device must alarm audibly and visually on violation of adjustable setting of LOW and HIGH heart rate
3.62 | Alarm limit range: adjustable from 30 to 300 bpm (approximately) and 4 seconds of asystole.
3.63 | The audible alarm should have a mute function to temporarily silence the audible alarm.
3.64 | Audible alarm-silencing period: 120 s (max)
3.65 | The screen must be visible under all lighting conditions.
3.66 | Display size: Approximately 115 x 85 mm.
3.67 | Display resolution: Approximately 320 x 240 pixels
3.68 | Display to indicate the following:
3.68.1 | Alarm messages
3.68.2 | Heart rate
3.68.3 | Systole blinker
3.68.4 | Alarm limits
3.68.5 | ECG, with freeze capability
3.69 | A high-resolution, approximately 200 dpi, printer must be incorporated in the unit
3.70 | Two channel printer for ECG and SpO2 waveform
3.71 | Printer sensitivity range 1 mV = 5 mm to 1 mV = 25 mm
3.72 | Muscle artifact filter
3.73 | Print-out of the following:
3.73.1 | ECG trace
3.73.2 | SpO2 wavefrom
3.73.3 | Date
3.73.4 | Time
3.73.5 | Energy selected
3.73.6 | Energy delivered
3.74 | Printer paper: 50 -100mm
3.75 | Paper speed: 25 to 50 mm/s .
3.76 | Manual and automatic printer modes.
3.77 | Automatic mode must print ECG trace before and after alarm/defibrillation conditions
3.78 | Memory/storage: At least 1 hour of ECG recording with all significant events
3.79 | External pacemaker
3.79.1 | Modes:
3.79.1.1 | Fixed
3.79.1.2 | Demand
3.79.2 | 30 - 180 paces/min (Fixed, demand)
3.79.3 | Impulse current: 10 - 180 mA, in 1 mA steps
3.79.4 | Impulse shape: Square: 20 -40 ms
3.79.5 | Refractory period (Demand): 240 ms (HR > 100 bpm), 240 ms (HR < 100 bpm)
3.79.6 | QRS marker in Demand mode
3.79.7 | There must be a continuous display of the pacing mode, the selected current and the selected pacing rate.
3.80 | The unit must be handed over in full operating order
3.81 | A starter-pack of consumables must be supplied with the unit.
4 | Standard Accessories and Consumables
4.1 | 3-Lead ECG patient cable (x1)
4.2 | ECG electrodes (25/pack)
4.3 | AED pads (x2 pairs- Paeds & Adults)
4.4 | Cable for AED pads (x1)
4.5 | Pacing electrodes
4.6 | Pacing cable
4.7 | Electrode gel (x1: 250ml)
4.8 | Printer paper (x 10 rolls)
5 | Conformity Compliance (Please attach certificate)
5.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
5.2 | OEM must comply and certified on ISO 9001 quality standards, attach proof of compliance
5.2.1 | Please provide unique ref number of the ISO 9001 ceriticate:
5.3 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
5.3.1 | Please provide unique ref number of the ISO 13485 ceriticate:
5.4 | Model quoted for must be EC certified. Attach a copy of certification
5.4.1 | Please provide unique ref number of the EC ceriticate:
5.5 | Hazardous Substance Act: If this type of equipment / apparatus appears on the schedule of Hazardous substances issued by the Directorate: Health Technology of the Department of Health, a license in terms of the Act on Hazardous Substances (Act 15/1973) must be submitted with the bid / quotation documents. The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.6 | All electrical/electronic medical equipment must be licensed by Radiation Control, where applicable a copy of the license must be submitted The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.6.1 | Radiation Control License number:
5.7 | The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade (rubber to with 2 screws) type and it must be a minimum length of three (3) meters. NB the mains cable of the unit tendered for must be SABS colour coded.
5.8 | The equipment quoted for must be protected against electromagnetic interference IEC 60601-1-2
5.9 | The unit must comply with electrical safety and performance standard such as IEC 60601-2-4, attached certification
5.9.1 | Please provide unique ref number of the IEC ceriticate:
6 | COMPREHENSIVE MAINTENANCE
6.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) must be included in the bid price
6.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
6.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
6.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
6.3 | The bidder to provide the OEM recommended service interval of the unit:
6.4 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
7 | General
7.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
7.2 | The unit on offer must be the latest technology.
Defibrillator Accessories and Consumables
RT2-05-04a | Reusable external adult/paediatric defibrillator paddles
RT2-05-04b | 3-Lead ECG patient cable (x1)
RT2-05-04c | ECG electrodes (x25)
RT2-05-04d | AED pads (x2 pairs)
RT2-05-04e | Cable for AED pads (x1)
RT2-05-04f | Pacing electrodes
RT2-05-04g | Pacing cable
RT2-05-04h | Electrode gel (x1: 250ml)
RT2-05-04i | Printer paper (x 10 rolls)
RT2-05-04j | Dedicated trolley
RT2-05-04k | Internal paddles (pair)
Bidder to list any other Consumables or Accessories (Not listed above) for the model on offer below

## Defib, ECG&SPO2
RT2-05-05 | Defibrillator, with ECG & SpO2
1 | Description
1.1 | Defibrillator to reverse fibrillation to a normal rhythm with ECG and peripheral oxygen saturation monitoring
2 | Product Offered:
2.1 | Make:
2.2 | Model:
2.3 | Initial Manufacturing year of Model:
2.4 | Radiation Control License number:
3 | Technical Specification
3.1 | The defibrillator offered must provide ECG monitoring, a printer and SpO2
3.2 | The defibrillator, monitor, printer, pacing unit, SpO2 unit, battery and battery-charger must be integrated into a single unit.
3.3 | The defibrillator must deliver a bi-phasic impulse
3.4 | The Defib paddles must be detacheable from the cable or machine to exchange with AED pads
3.5 | The defibrillator must have manual and automatic (AED) modes
3.6 | Automatic mode must measure patient impedance and deliver an appropriate impulse automatically generated, between 0 and 250 J
3.7 | The unit must offer audible/voice and on-screen messages for use during AED mode
3.8 | Manual mode energy selection: 5 - 200 J
3.9 | Selection of manual energy must be simple with a clear unambiguous indication of the energy selected
3.10 | The defibrillator must have the following selection of defibrillation:
3.10.1 | a. Non-synchronous defibrillation
3.10.2 | b. Synchronous defibrillation
3.11 | The defibrillator must indicate if the synchronous mode has been selected.
3.12 | The defibrillator must default to non-synchronous mode at initial switch on.
3.13 | Synchronous defibrillation detection must be possible via ECG patient cable and via paddles
3.14 | The synch marker must be clearly indicated on the ECG trace.
3.15 | The defibrillator must operate from a 220 V, 50 Hz. supply
3.16 | The defibrillator must contain a rechargeable battery which charges automatically when the defibrillator is plugged into the 220 V, 50 Hz supply.
3.17 | The battery must charge whether the defibrillator is switched on or not
3.18 | The defibrillator must have a visual indicator to indicate that the mains supply is switched on.
3.19 | The internal rechargeable battery must be able to deliver at least 50 defibrillations at the maximum setting and provide a minimum of 2 hours ECG monitoring in the event of a mains power failure
3.20 | Battery status indicator with audible and visual alarm for low battery
3.21 | Audible tone while charging capacitor
3.22 | Audible tone, different from above, to indicate full charged capacitor
3.23 | Visual indication of fully charged state
3.24 | Defibrillating charge time to maximum setting must not exceed 10 seconds.
3.25 | In the event of a completely depleted battery, the defibrillator must be fully operable on the 220 V, 50 Hz supply.
3.26 | The defibrillator must do an automatic self-test when switched on.
3.27 | Override of self test must be possible
3.28 | Operation of the defibrillator must be user-friendly with clear indications for:
3.28.1 | Switch on;
3.28.2 | Select energy;
3.28.3 | Charge;
3.28.4 | Discharge.
3.29 | The defibrillator must be portable and weigh less than 7 kg with all accessories attached
3.30 | The unit must have a carry handle
3.31 | The defibrillator must incorporate an internal 50 ohms test load to allow a check for satisfactory functioning.
3.32 | The defibrillator must incorporate an emergency dump circuit that will automatically provide a safe current path for unused energy when the paddles are charged but not discharged within sixty seconds.
3.33 | The defibrillator must be able to do both internal and external defibrillation.
3.34 | Energy selection must accommodate patients from paeds to adults for both internal and external defibrillation.
3.35 | The defibrillator must automatically sense when internal paddles are in use and restrict the maximum selectable energy to 50 Joules.
3.36 | Apex paddle to incorporate a push button to charge the defibrillator.
3.37 | Energy discharge via defibrillator paddles to patient or test load must be activated by pressing buttons on both apex and sternum paddles simultaneously.
3.38 | Energy discharge via AED pads or internal electrodes by means of a button on the panel
3.39 | ECG monitoring must be possible through either the ECG leads or the paddles and unit must display the ECG lead selected (I, II or III)/(aVl, aVR or aVF) or indicate whether the paddles are the source of the ECG
3.40 | Size of ECG: 0,5; 1; 2 and 4 cm/mV
3.41 | Sweep rate: 25 mm/s
3.42 | 50 Hz filter
3.43 | Heart rate: 30 - 300 bpm
3.44 | Analysis time: maximum 15 s
3.45 | Artefact detection
3.46 | Pacemaker detection
3.47 | The following alarms must be displayed:
3.47.1 | low battery,
3.47.2 | no paper (printer)
3.47.3 | ECG lead off/fault.
3.48 | Device must alarm audibly and visually on violation of adjustable setting of LOW and HIGH heart rate
3.49 | Alarm limit range: adjustable from 30 to 300 bpm (approximately) and 4 seconds of asystole.
3.50 | The audible alarm should have an mute function to temporarily silence the audible alarm.
3.51 | Audible alarm-silencing period: 120 s (max)
3.52 | The screen must be visible under all lighting conditions.
3.53 | Display size: Approximately 115 x 85 mm.
3.54 | Display resolution: Approximately 320 x 240 pixels
3.55 | Display to indicate the following:
3.55.1 | Alarm messages
3.55.2 | Heart rate
3.55.3 | Systole blinker
3.55.4 | Alarm limits
3.55.5 | ECG, with freeze capability
3.56 | A high-resolution, approximately 200 dpi, printer must be incorporated in the unit
3.57 | Two channel printer for ECG and SpO2 waveform
3.58 | Printer sensitivity range 1 mV = 5 mm to 1 mV = 25 mm
3.59 | Muscle artifact filter
3.60 | Print-out of the following:
3.60.1 | ECG trace
3.60.2 | SpO2 wavefrom
3.60.3 | Date
3.60.4 | Time
3.60.5 | Energy selected
3.60.6 | Energy delivered
3.61 | Printer paper: 50 - 100 mm
3.62 | Paper speed: 25 to 50 mm/s .
3.63 | Manual and automatic printer modes.
3.64 | Automatic mode must print ECG trace before and after alarm/defibrillation conditions
3.65 | Memory/storage: At least 1 hour of ECG recording with all significant events
3.66 | SpO2
3.66.1 | Alarm for max and min SpO2
3.66.2 | Saturation range: 80 - 100%, accuracy within 2% to 3% for readings between 70 - 90%
3.66.3 | Pulse rate range: 30 - 300 ppm, +- 5 ppm
3.67 | The unit must be handed over in full operating order
3.68 | A starter-pack of consumables must be supplied with the unit.
4 | Standard Accessories and Consumables
4.1 | 3-Lead ECG patient cable (x1)
4.2 | ECG electrodes (x25)
4.3 | AED pads (x2 pairs)
4.4 | Cable for AED pads (x1)
4.5 | Electrode gel (x1: 250ml)
4.6 | Printer paper (x 10 rolls)
4.7 | SpO2 transducer
4.8 | SpO2 cable
5 | Conformity Compliance (Please attach certificate)
5.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
5.2 | OEM must comply and certified on ISO 9001 quality standards, attach proof of compliance
5.2.1 | Please provide unique ref number of the ISO 9001 ceriticate:
5.3 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
5.3.1 | Please provide unique ref number of the ISO 13485 ceriticate:
5.4 | Model quoted for must be EC certified. Attach a copy of certification
5.4.1 | Please provide unique ref number of the EC ceriticate:
5.5 | Hazardous Substance Act: If this type of equipment / apparatus appears on the schedule of Hazardous substances issued by the Directorate: Health Technology of the Department of Health, a license in terms of the Act on Hazardous Substances (Act 15/1973) must be submitted with the bid / quotation documents. The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.6 | All electrical/electronic medical equipment must be licensed by Radiation Control, where applicable a copy of the license must be submitted The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.6.1 | Radiation Control License number:
5.7 | The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade (rubber to with 2 screws) type and it must be a minimum length of three (3) meters. NB the mains cable of the unit tendered for must be SABS colour coded.
5.8 | The equipment quoted for must be protected against electromagnetic interference IEC 60601-1-2
5.9 | The unit must comply with electrical safety and performance standard such as IEC 60601-2-4, attached certification
5.9.1 | Please provide unique ref number of the IEC ceriticate:
6 | COMPREHENSIVE MAINTENANCE
6.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) must be included in the bid price
6.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries, Flow sensors and Oxygen sensor), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
6.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
6.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries, Flow sensors and Oxygen sensor), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
6.3 | The bidder to provide the OEM recommended service interval of the unit:
6.4 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
7 | General
7.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
7.2 | The unit on offer must be the latest technology.
Defibrillator, with ECG & SpO2- - Accessories and Consumables
RT2-05-05a | 3-Lead ECG patient cable (x1)
RT2-05-05b | ECG electrodes (25/Pack)
RT2-05-05c | AED pads (x2 pairs)
RT2-05-05d | Cable for AED pads (x1)
RT2-05-05e | Electrode gel (x1: 250ml)
RT2-05-05f | Printer paper (x 10 rolls)
RT2-05-05g | SpO2 transducer
RT2-05-05h | SpO2 cable
RT2-05-05i | Dedicated trolley
RT2-05-05j | Internal paddles
RT2-05-05k | Reusable external adult/paediatric defibrillator paddles
Bidder to list any other Consumables or Accessories (Not listed above) for the model on offer below

## Defib, ECG,SpO2&NIBP
RT2-05-06 | Defibrillator, with ECG, SpO2, and NIBP
1 | Description
1.1 | Defibrillator to reverse fibrillation to a normal rhythm with ECG, peripheral oxygen saturation and NIBP monitoring
2 | Product Offered:
2.1 | Make:
2.2 | Model:
2.3 | Initial Manufacturing year of Model:
2.4 | Radiation Control License number:
3 | Technical Specification
3.1 | The defibrillator offered must provide ECG monitoring, a printer and SpO2, and NIBP
3.2 | The defibrillator, monitor, printer, SpO2 unit, NIBP, battery and battery-charger must be integrated into a single unit.
3.3 | The defibrillator must deliver a bi-phasic impulse
3.4 | The Defib paddles must be detacheable from the cable or machine to exchange with AED pads
3.5 | The defibrillator must have manual and automatic (AED) modes
3.6 | Automatic mode must measure patient impedance and deliver an appropriate impulse automatically generated, between 0 and 250 J
3.7 | The unit must offer audible/voice and on-screen messages for use during AED mode
3.8 | Manual mode energy selection: 5 - 200 J
3.9 | Selection of manual energy must be simple with a clear unambiguous indication of the energy selected
3.10 | The defibrillator must have the following selection of defibrillation:
3.10.1 | a. Non-synchronous defibrillation
3.10.2 | b. Synchronous defibrillation
3.11 | The defibrillator must indicate if the synchronous mode has been selected.
3.12 | The defibrillator must default to non-synchronous mode at initial switch on.
3.13 | Synchronous defibrillation detection must be possible via ECG patient cable and via paddles
3.14 | The synch marker must be clearly indicated on the ECG trace.
3.15 | The defibrillator must operate from a 220 V, 50 Hz. supply
3.16 | The defibrillator must contain a rechargeable battery which charges automatically when the defibrillator is plugged into the 220 V, 50 Hz supply.
3.17 | The battery must charge whether the defibrillator is switched on or not
3.18 | The defibrillator must have a visual indicator to indicate that the mains supply is switched on.
3.19 | The internal rechargeable battery must be able to deliver at least 50 defibrillations at the maximum setting and provide a minimum of 2 hours ECG monitoring in the event of a mains power failure
3.20 | Battery status indicator with audible and visual alarm for low battery
3.21 | Audible tone while charging capacitor
3.22 | Audible tone, different from above, to indicate full charged capacitor
3.23 | Visual indication of fully charged state
3.24 | Defibrillating charge time to maximum setting must not exceed 10 seconds.
3.25 | In the event of a completely depleted battery, the defibrillator must be fully operable on the 220 V, 50 Hz supply.
3.26 | The defibrillator must do an automatic self-test when switched on.
3.27 | Override of self test must be possible
3.28 | Operation of the defibrillator must be user-friendly with clear indications for:
3.28.1 | Switch on;
3.28.2 | Select energy;
3.28.3 | Charge;
3.28.4 | Discharge.
3.29 | The defibrillator must be portable and weigh less than 7 kg with all accessories attached
3.30 | The unit must have a carry handle
3.31 | The defibrillator must incorporate an internal 50 ohms test load to allow a check for satisfactory functioning.
3.32 | The defibrillator must incorporate an emergency dump circuit that will automatically provide a safe current path for unused energy when the paddles are charged but not discharged within sixty seconds.
3.33 | The defibrillator must be able to do both internal and external defibrillation.
3.34 | Energy selection must accommodate patients from paeds to adults for both internal and external defibrillation.
3.35 | The defibrillator must automatically sense when internal paddles are in use and restrict the maximum selectable energy to 50 Joules.
3.36 | Apex paddle to incorporate a push button to charge the defibrillator.
3.37 | Energy discharge via defibrillator paddles to patient or test load must be activated by pressing buttons on both apex and sternum paddles simultaneously.
3.38 | Energy discharge via AED pads or internal electrodes by means of a button on the panel
3.39 | ECG monitoring must be possible through either the ECG leads or the paddles and unit must display the ECG lead selected (I, II or III)/(aVl, aVR or aVF) or indicate whether the paddles are the source of the ECG
3.40 | Size of ECG: 0,5; 1; 2 and 4 cm/mV
3.41 | Sweep rate: 25 mm/s
3.42 | 50 Hz filter
3.43 | Heart rate: 30 - 300 bpm
3.44 | Analysis time: maximum 15 s
3.45 | Artefact detection
3.46 | Pacemaker detection
3.47 | The following alarms must be displayed:
3.47.1 | low battery,
3.47.2 | no paper (printer)
3.47.3 | ECG lead off/fault.
3.48 | Device must alarm audibly and visually on violation of adjustable setting of LOW and HIGH heart rate
3.49 | Alarm limit range: adjustable from 30 to 300 bpm (approximately) and 4 seconds of asystole.
3.50 | The audible alarm should have a mute function to temporarily silence the audible alarm.
3.51 | Audible alarm-silencing period: 120 s (max)
3.52 | The screen must be visible under all lighting conditions.
3.53 | Display size: Approximately 115 x 85 mm.
3.54 | Display resolution: Approximately 320 x 240 pixels
3.55 | Display to indicate the following:
3.55.1 | Alarm messages
3.55.2 | Heart rate
3.55.3 | Systole blinker
3.55.4 | Alarm limits
3.55.5 | ECG, with freeze capability
3.56 | A high-resolution, approximately 200 dpi, printer must be incorporated in the unit
3.57 | Two channel printer for ECG and SpO2 waveform
3.58 | Printer sensitivity range 1 mV = 5 mm to 1 mV = 25 mm
3.59 | Muscle artifact filter
3.60 | Print-out of the following:
3.60.1 | ECG trace
3.60.2 | SpO2 wavefrom
3.60.3 | Date
3.60.4 | Time
3.60.5 | Energy selected
3.60.6 | Energy delivered
3.61 | Printer paper: 50 - 100 mm
3.62 | Paper speed: 25 to 50 mm/s .
3.63 | Manual and automatic printer modes.
3.64 | Automatic mode must print ECG trace before and after alarm/defibrillation conditions
3.65 | Memory/storage: At least 1 hour of ECG recording with all significant events
3.66 | SpO2
3.66.1 | Alarm for max and min SpO2
3.66.2 | Saturation range: 80 - 100%, accuracy within 2% to 3% for readings between 70 - 90%
3.66.3 | Pulse rate range: 30 - 300 ppm, +- 5 ppm
3.67 | NIBP
3.67.1 | Oscillometric measurement method
3.67.2 | Ranges (approximate) :
3.67.2.1 | Systolic : 30 - 270 mmHg. Resolution : 1 mmHg
3.67.2.2 | Diastolic : 10 - 150 mmHg. Resolution : 1 mmHg
3.67.2.3 | Mean : 20 - 250 mmHg. Resolution : 1 mmHg
3.67.3 | Cuff pressure : 0 - 280 mmHg
3.67.4 | Cuff inflation time < 10 s
3.67.5 | All available adult reusable cuffs must be supplied with the monitor.
3.68 | The unit must be handed over in full operating order
3.69 | A starter-pack of consumables must be supplied with the unit.
4 | Standard Accessories and Consumbles
4.1 | 3-Lead ECG patient cable (x1)
4.2 | ECG electrodes (25/Pack)
4.3 | AED pads (x2 pairs - Paeds & Adults)
4.4 | Cable for AED pads (x1)
4.5 | Electrode gel (x1: 250ml)
4.6 | Printer paper (x 10 rolls)
4.7 | SpO2 sensor
4.8 | SpO2 cable
4.9 | BP hose
4.10 | BP cuff (Adult)
5 | Conformity Compliance (Please attach certificate)
5.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
5.2 | OEM must comply and certified on ISO 9001 quality standards, attach proof of compliance
5.2.1 | Please provide unique ref number of the ISO 9001 ceriticate:
5.3 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
5.3.1 | Please provide unique ref number of the ISO 13485 ceriticate:
5.4 | Model quoted for must be EC certified. Attach a copy of certification
5.4.1 | Please provide unique ref number of the EC ceriticate:
5.5 | Hazardous Substance Act: If this type of equipment / apparatus appears on the schedule of Hazardous substances issued by the Directorate: Health Technology of the Department of Health, a license in terms of the Act on Hazardous Substances (Act 15/1973) must be submitted with the bid / quotation documents. The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.6 | All electrical/electronic medical equipment must be licensed by Radiation Control, where applicable a copy of the license must be submitted The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.6.1 | Radiation Control License number:
5.7 | The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade (rubber to with 2 screws) type and it must be a minimum length of three (3) meters. NB the mains cable of the unit tendered for must be SABS colour coded.
5.8 | The equipment quoted for must be protected against electromagnetic interference IEC 60601-1-2
5.9 | The unit must comply with electrical safety and performance standard such as IEC 60601-2-4, attached certification
5.9.1 | Please provide unique ref number of the IEC ceriticate:
6 | COMPREHENSIVE MAINTENANCE
6.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) must be included in the bid price
6.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
6.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
6.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
6.3 | The bidder to provide the OEM recommended service interval of the unit:
6.4 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
7 | General
7.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
7.2 | The unit on offer must be the latest technology.
Defibrillator, with ECG, SpO2, and NIBP
RT2-05-06a | 3-Lead ECG patient cable (x1)
RT2-05-06b | ECG electrodes (25/Pack)
RT2-05-06c | AED pads (x2 pairs - Paeds & Adults)
RT2-05-06d | Cable for AED pads (x1)
RT2-05-06e | Electrode gel (x1: 250ml)
RT2-05-06f | Printer paper (x 10 rolls)
RT2-05-06g | SpO2 sensor
RT2-05-06h | SpO2 cable
RT2-05-06i | BP hose
RT2-05-06j | BP cuff (Adult)
RT2-05-06k | Dedicated trolley
RT2-05-06l | Internal paddles
RT2-05-06m | Reusable external adult/paediatric defibrillator paddles
Bidder to list any other Consumables or Accessories (Not listed above) for the model on offer below

## ECG, 12-Lead
RT2-05-07 | ELECTROCARDIOGRAPH (12 CHANNEL) on trolley.
1 | Description
1.1
2 | Product Offered:
2.1 | Make:
2.2 | Model:
2.3 | Initial Manufacturing year of Model:
2.4 | Radiation Control License number:
3 | Technical Specification
3.1 | Twelve channel portable electrocardiograph
3.2 | The unit must contain a recorder/printer.
3.3 | The unit must have a colour TFT display of approximately 115 x 85 mm
3.4 | High resolution display: at least 800 x 600
3.5 | The unit must do a self-test at power-on
3.6 | The following must be displayed:
3.6.1 | Heart rate
3.6.2 | Patient ID
3.6.3 | 12 lead waveforms
3.6.4 | Lead identification
3.6.5 | ECG gain
3.6.6 | Paper speed
3.6.7 | Filters
3.6.8 | Electrode connection indicator
3.6.9 | Messages and prompts
3.6.10 | Clock with date and time
3.6.11 | Battery power status
3.7 | R-R analysis for 30 to 300 sec
3.8 | Pacemaker detection
3.9 | The display screen must be capable of displaying two lines of text and a minimum of three ECG waveforms.
3.10 | The unit must be supplied with a 10 lead ECG cable and be able to record 3, 6 and 12 lead configurations.
3.11 | 12 lead configuration:
3.11.1 | Lead I, Lead II, Lead III
3.11.2 | AVR, AVL, AVF
3.11.3 | V1, V2, V3, V4, V5, V6
3.12 | Manual and automatic modes
3.13 | The unit must be capable of ECG interpretation for paediatric to adult.
3.14 | Full alphanumeric keyboard or touchscreen
3.15 | The unit must have facility for entering patient data and text
3.16 | The unit must offer a minimum of six pre-set ECG programs.
3.17 | Event markers
3.18 | Automatic indication of disconnected leads/electrodes
3.19 | Automatic check for incorrect lead placement
3.20 | Full screen print preview of ECG waveforms prior to printing
3.21 | Heart rate recognition: peak-to peak detection. 30 - 300 bpm, accuracy 1 bpm
3.22 | Electrode connection quality check.
3.23 | ECG amplifier specs
3.23.1 | Simultaneous acquisition of 12 leads
3.23.2 | Frequency response: 0.05 - 150 Hz
3.23.3 | Sensitivity: Automatic or selectable: 2.5 - 20 mm/mV
3.23.4 | Automatic amplitude adjustment in the range of 2.5 to 40 mm/mV.
3.23.5 | Sensitivity threshold: 20 µV
3.23.6 | Input impedance: > 50 MΩ
3.23.7 | Patient leakage current: < 10 µA
3.23.8 | Calibration pulse: 1 mV
3.23.9 | Filters:
3.23.9a | 50 Hz
3.23.9b | Low pass: 25, 35, 75, 100, 150 Hz
3.23.10 | CMRR: > 100 dB
3.23.11 | Sampling frequency: 1 kHz
3.23.12 | Artifact rejection
3.23.13 | Baseline stabiliser
3.24 | Printer
3.24.1 | Automatic and manual printout mode.
3.24.2 | In addition to the required ECG waveforms the printer must be able to print:
3.24.2a | device settings
3.24.2b | heart rate
3.24.2c | date and time
3.24.2d | lead identification
3.24.2e | calibration pulse
3.24.2f | patient data
3.24.3 | The ECG printout must be available on A4 size, Z fold paper.
3.24.4 | Printer quality: 8 dots/mm (vertical), 20 dots/mm (horiz, 25 mm/s)
3.24.5 | Paper speed must be selectable as follows: 25 and 50 mm/s.
3.25 | Internal memory capable of storing at least 200 patient files.
3.26 | External storage by means of SD card (or other) for 2 000 patient files
3.27 | Internal protection against damage from 360 J defibrillator discharge.
3.28 | The unit must have an internal battery with integrated charger
3.29 | Battery capacity: at least 2 hours of continuous recording
3.30 | The unit must be fully operable from mains power regardless of the internal battery condition.
3.31 | A sturdy trolley to house the ECG machine and accessories must be supplied.
3.32 | Trolley must have at least one drawer
3.33 | Trolley must have at least four castors, two with brakes
3.34 | Mains cable at least four metres long with SABS approved 15A three prong plug .
3.35 | The unit must be handed over in full operating order
3.36 | A starter-pack of consumables must be supplied with the unit.
4 | Standard Accessories and Consumables
4.1 | ECG cable with adaptor clips
4.2 | ECG electrodes (25/Pack)
4.3 | Dedicated trolley
5 | Conformity Compliance (Please attach certificate)
5.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
5.2 | OEM must comply and certified on ISO 9001 quality standards, attach proof of compliance
5.2.1 | Please provide unique ref number of the ISO 9001 ceriticate:
5.3 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
5.3.1 | Please provide unique ref number of the ISO 13485 ceriticate:
5.4 | Model quoted for must be EC certified. Attach a copy of certification
5.4.1 | Please provide unique ref number of the EC ceriticate:
5.5 | Hazardous Substance Act: If this type of equipment / apparatus appears on the schedule of Hazardous substances issued by the Directorate: Health Technology of the Department of Health, a license in terms of the Act on Hazardous Substances (Act 15/1973) must be submitted with the bid / quotation documents. The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.6 | All electrical/electronic medical equipment must be licensed by Radiation Control, where applicable a copy of the license must be submitted The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.6.1 | Radiation Control License number:
5.7 | The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade (rubber to with 2 screws) type and it must be a minimum length of three (3) meters. NB the mains cable of the unit tendered for must be SABS colour coded.
5.8 | The equipment quoted for must be protected against electromagnetic interference IEC 60601-1-2
5.9 | The unit must comply with electrical safety and performance standard such as IEC 60601-2-25, attached certification
5.9.1 | Please provide unique ref number of the IEC ceriticate:
5.10 | The unit must comply with electrical safety and performance standard such as IEC 60601-2-51, attached certification
5.10.1 | Please provide unique ref number of the IEC ceriticate:
6 | COMPREHENSIVE MAINTENANCE
6.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) must be included in the bid price
6.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
6.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
6.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
6.3 | The bidder to provide the OEM recommended service interval of the unit:
6.4 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
7 | General
7.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
7.2 | The unit on offer must be the latest technology.
ELECTROCARDIOGRAPH (12 CHANNEL) ACCESSORIES and CONSUMABLES
RT2-05-07a | ECG cable with adaptor clips
RT2-05-07b | ECG electrodes (25/Pack)
RT2-05-07c | Dedicated trolley
RT2-05-07d | Battery
Bidder to list any other Consumables or Accessories (Not listed above) for the model on offer below

## Exam light, LED, Mobile
RT2-05-09 | Examination light, LED, Mobile
1 | Description
1.1
2 | Product Offered:
2.1 | Make:
2.2 | Model:
2.3 | Initial Manufacturing year of Model:
3 | Technical Specification
3.1 | The examination light must be a single head LED with suitable reflector.
3.2 | The lamp must be mounted on a mobile stand.
3.3 | The lamp must have sufficient movement to ensure effective lighting over the bed with a minimum length of 1m
3.4 | The extension must be spring balanced with smooth and easy manoeuvring with one hand
3.5 | No movement (drift) must take place after positioning the lamp.
3.6 | Must have sturdy upright pole at least 1,5 m with enough counter-weight to support arm at full reach at 90o
3.7 | Must have friction joints to allow for mobility and positioning.
3.8 | Must have a handle for positioning the light
3.9 | Power supply must be hidden in the base of the unit, or neatly installed out of sight.
3.10 | Light intensity of the lamp must not be less than 30 000 lux measured at 1,0 metre.
3.11 | The luminaire must be cool to touch, even during multiple examinations.
3.12 | The light must be separately switched either on the column or on the luminaire.
3.13 | Colour temp, approximately 4 000K
3.14 | Colour rendering index: > 90
3.15 | Lamp life: > 30 000 continuous hours
3.16 | Spot size (maximum):
3.16.1 | Working distance of 50 cm = 15 cm
3.16.2 | Working distance of 100 cm = 25 cm
3.17 | Base diameter: at least 35 cm
3.18 | Must operate from 230V, 50 Hz mains power
3.19 | Mains cable at least four metres, with 15A three pin plug, SABS approved
3.20 | Mobile stand: lockable castors + hook for mains cable
3.21 | Durable castors
3.22 | Smooth easily cleaned surface
3.23 | The unit must be handed over in full operating order
4 | Conformity Compliance (Please attach certificate)
4.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
4.2 | OEM must comply and certified on ISO 9001 quality standards, attach proof of compliance
4.2.1 | Please provide unique ref number of the ISO 9001 ceriticate:
4.3 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
4.3.1 | Please provide unique ref number of the ISO 13485 ceriticate:
4.4 | Model quoted for must be EC certified. Attach a copy of certification
4.4.1 | Please provide unique ref number of the EC ceriticate:
4.5 | The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade (rubber to with 2 screws) type and it must be a minimum length of three (3) meters. NB the mains cable of the unit tendered for must be SABS colour coded.
4.6 | The equipment quoted for must be protected against electromagnetic interference IEC 60601-1-2
5 | COMPREHENSIVE MAINTENANCE
5.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) must be included in the bid price
5.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
6 | General
6.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
6.2 | The unit on offer must be the latest technology.
Examination light, LED, Mobile - Accessories
RT2-05-09a | Bidders to state the replacement cost per LED or range of LEDs.
Bidder to list any other Consumables or Accessories (Not listed above) for the model on offer below

## Exam light, LED, Mobile+Battery
RT2-05-10 | Examination light, LED, mobile, with battery
1 | Description
1.1 | A mobile LED examination light with battery power backup
2 | Product Offered:
2.1 | Make:
2.2 | Model:
2.3 | Initial Manufacturing year of Model:
3 | Technical Specification
3.1 | The examination light must be a single head LED with suitable reflector.
3.2 | The lamp must be mounted on a mobile stand
3.3 | The lamp must have sufficient movement to ensure effective lighting over the bed with a minimum length of 1m
3.4 | The extension must be spring balanced with smooth and easy manoeuvring with one hand
3.5 | No movement (drift) must take place after positioning the lamp.
3.6 | Must have sturdy upright pole at least 1,5 m with enough counter-weight to support arm at full reach at 90o
3.7 | Must have friction joints to allow for mobility and positioning.
3.8 | Must have a handle for positioning the light
3.9 | Power supply and battery must be hidden in the base of the unit, or neatly installed out of sight.
3.10 | Light intensity of the lamp must not be less than 30 000 lux measured at 1,0 metre.
3.11 | The luminaire must be cool to touch, even during multiple examinations.
3.12 | The light must be separately switched either on the column or on the luminaire.
3.13 | Colour temp, approximately 4 000K
3.14 | Colour rendering index: > 90
3.15 | Lamp life: > 30 000 continuous hours
3.16 | Spot size (maximum):
3.16.1 | Working distance of 50 cm = 15 cm
3.16.2 | Working distance of 100 cm = 25 cm
3.17 | Base diameter: at least35 cm
3.18 | Must operate from 230V, 50 Hz mains power
3.19 | Must contain a battery charger
3.20 | Battery must be capable for at least 2 hours of operation
3.21 | Mains cable at least four metres, with 15A three pin plug, SABS approved
3.22 | Mobile stand: lockable castors + hook for mains cable
3.23 | Durable castors
3.24 | Smooth easily cleaned surface
3.25 | The unit must be handed over in full operating order
4 | Conformity Compliance (Please attach certificate)
4.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
4.2 | OEM must comply and certified on ISO 9001 quality standards, attach proof of compliance
4.2.1 | Please provide unique ref number of the ISO 9001 ceriticate:
4.3 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
4.3.1 | Please provide unique ref number of the ISO 13485 ceriticate:
4.4 | Model quoted for must be EC certified. Attach a copy of certification
4.4.1 | Please provide unique ref number of the EC ceriticate:
4.5 | The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade (rubber to with 2 screws) type and it must be a minimum length of three (3) meters. NB the mains cable of the unit tendered for must be SABS colour coded.
4.6 | The equipment quoted for must be protected against electromagnetic interference IEC 60601-1-2
5 | COMPREHENSIVE MAINTENANCE
5.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) must be included in the bid price
5.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
6 | General
6.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
6.2 | The unit on offer must be the latest technology.
Examination light, LED, Mobile + Battery - Accessories
RT2-05-10a | Bidders to state the replacement cost per LED or range of LEDs.
RT2-05-10b | Battery
Bidder to list any other Consumables or Accessories (Not listed above) for the model on offer below

## Exam light, LED, Wall-mount
RT2-05-11 | Examination light, LED, wall mount
1 | Description
1.1 | A wall mounted LED examination light
2 | Product Offered:
2.1 | Make:
2.2 | Model:
2.3 | Initial Manufacturing year of Model:
3 | Technical Specification
3.1 | The examination light must be a single head LED with suitable reflector.
3.2 | The lamp must have sufficient movement to ensure effective lighting over the bed with a minimum length of 1m
3.3 | The extension must be spring balanced with smooth and easy manoeuvring with one hand. No "goose-necks" will be acceptable
3.4 | Must be able to push completely out of the way against the wall for storage.
3.5 | Restriction to prevent bashing against the wall
3.6 | No movement (drift) must take place after positioning the lamp.
3.7 | Must have friction joints to allow for mobility and positioning.
3.8 | Must have a handle for positioning the light
3.9 | Light intensity of the lamp must not be less than 30 000 lux measured at 1,0 metre.
3.10 | The luminaire must be cool to touch, even during multiple examinations.
3.11 | Colour temp, approximately 4 000K
3.12 | Colour rendering index: > 90
3.13 | Lamp life: > 30 000 continuous hours
3.14 | Spot size (maximum):
3.15 | Working distance of 50 cm = 15 cm
3.16 | Working distance of 100 cm = 25 cm
3.17 | Must operate from 230V, 50 Hz mains power
3.18 | Power supply to be built in and mounted in/on the wall
3.19 | Smooth easily cleaned surface
3.20 | Bidder must be responsible for installation/mounting and electrical connection at no extra cost - end user consulation
3.21 | The unit must be handed over in full operating order
4 | Conformity Compliance (Please attach certificate)
4.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
4.2 | OEM must comply and certified on ISO 9001 quality standards, attach proof of compliance
4.2.1 | Please provide unique ref number of the ISO 9001 ceriticate:
4.3 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
4.3.1 | Please provide unique ref number of the ISO 13485 ceriticate:
4.4 | Model quoted for must be EC certified. Attach a copy of certification
4.4.1 | Please provide unique ref number of the EC ceriticate:
4.5 | The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade (rubber to with 2 screws) type and it must be a minimum length of three (3) meters. NB the mains cable of the unit tendered for must be SABS colour coded.
4.6 | The equipment quoted for must be protected against electromagnetic interference IEC 60601-1-2
5 | COMPREHENSIVE MAINTENANCE
5.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) must be included in the bid price
5.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries, Flow sensors and Oxygen sensor), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
6 | General
6.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
Examination light, LED, wall mount- Accessories
RT2-05-11a | Bidders to state the replacement cost per LED or range of LEDs.
Bidder to list any other Consumables or Accessories (Not listed above) for the model on offer below

## Monitor, Basic
RT2-05-14 | MONITOR, BASIC
1 | Description
Basic physiologic monitor
2 | Product Offered:
2.1 | Make:
2.2 | Model:
2.3 | Initial Manufacturing year of Model:
2.4 | Radiation Control License number:
3 | Technical Specification
3.1 | Preconfigured, state of the art technology, suitable for neonates to adult including mounting solution suitable for the facility
3.2 | Microprocessor controlled with automatic self-test and calibration
3.3 | The following parameters must be measured and displayed :
3.3.1 | ECG waveform
3.3.2 | Oxygen saturation, waveform + digital
3.3.3 | Heart rate, digital
3.3.4 | Respiration, waveform + digital
3.3.5 | NIBP, digital
3.3.6 | Temperature, digital
3.4 | The monitor must operate off both 220 V, 50 Hz single phase AC power supply and an internal rechargeable battery.
3.5 | If mains power fails the monitor must operate off battery power for at least 2 hours.
3.6 | Visible indication to show whether the monitor is operating off the mains or battery power.
3.7 | All control keys must be splashproof and resistant to the entry of fluids into the unit.
3.8 | The front panel must be easy to clean and disinfect.
3.9 | The monitor must have a carry handle.
3.10 | High resolution LCD colour display, at least 35 cm diagonal
3.10.1 | Digits: large: 10 mm, small: 5 mm
3.10.2 | Sweep speeds : 25 and 50 mm/s
3.10.3 | It must be possible to freeze the display. This must not affect the QRS signal.
3.11 | Trending for 24 hours must be possible on the monitor
3.12 | Defibrillator-proof up to 360 Joules
3.13 | Protected against HF interference
3.14 | Patient isolated from all potentially harmful circuits
3.15 | Must comply with IEC 60601-1
3.16 | ECG
3.16.1 | Must comply with IEC 60601-2-27
3.16.2 | 3 lead ECG with arrhythmia detection
3.16.3 | Adult ECG lead must be supplied
3.16.4 | The amplitude of the ECG trace must be adjustable from 1 mV = 5 mm to 1 mV = 20 mm.
3.16.5 | Digital display of heart rate : 30 - 300 bpm
3.16.6 | Accuracy : 3 bpm
3.16.7 | Resolution : 1 bpm
3.16.8 | Differential input with CMRR > 80 dB
3.16.9 | Input impedance > 10 MOhm
3.16.10 | Respiration monitoring through ECG electrodes
3.16.11 | Pacemaker detection/rejection : 2 to 500 mV; 0,1 to 2 ms
3.17 | OXYGEN SATURATION
3.17.1 | It must be the latest technology with clinical research support based on accuracy and real-time monitoring. Please submit the published peer reviewed/researched paper as supporting documents.
3.17.2 | The unit should indicate whether the signal from the pulse oximeter sensor is adequate for reliable measurements. If the signal strength or quality is inadequate, this must be clearly displayed.
3.17.3 | An audible bleep must sound for each heart beat. The pitch must vary with the oxygen saturation.
3.17.4 | Measurement range : 0 - 100%
3.17.5 | Accuracy:
3.17.5a | 80 - 100% : 1,5%
3.17.5b | 60 - 80% : 3,5%
3.17.6 | Continuous waveform display
3.17.7 | Digital display of SpO2%
3.17.8 | SpO2 finger sensor to be supplied with the monitor
3.17.9 | The sensor lead must be at least 3 metres long.
3.17.10 | Automatic calibration for skin colour
3.17.11 | Pulse rate: 25 - 250 bpm
3.18 | NON-INVASIVE BLOOD PRESSURE
3.18.1 | Must comply with IEC 60601-2-30
3.18.2 | Oscillometric measurement method
3.18.3 | Automatic, manual and stat modes
3.18.4 | Measure and display :
3.18.4a | - systolic pressure
3.18.4b | - diastolic pressure
3.18.4c | - mean pressure
3.18.5 | The interval between measurements must be selectable from 1 to 60 minutes.
3.18.6 | Ranges (approximate) :
3.18.6a | Systolic : 30 - 270 mmHg. Resolution : 1 mmHg
3.18.6b | Diastolic : 10 - 150 mmHg. Resolution : 1 mmHg
3.18.6c | Mean : 20 - 250 mmHg. Resolution : 1 mmHg
18.7 | Cuff pressure : 0 - 280 mmHg
3.18.8 | Cuff inflation time < 10 s
3.18.9 | Automatic deflation if measurement exceeds 3 min or if mains power fails
3.19 | TEMPERATURE
3.19.1 | Range : 10 - 45 ˚C
3.19.2 | Digital display
3.19.3 | Surface and oral probe to be supplied
3.19.4 | Accuracy: 0.3 ˚C
3.19.5 | Resolution : 0.1 ˚C
3.20 | ALARMS
3.20.1 | Visible and audio alarms
3.20.2 | The monitor must automatically set default alarms when switched on.
3.20.3 | The alarms must be programmable
3.20.4 | Alarm limits must be displayed with the parameter digit
3.20.5 | Heart rate alarm : 30 - 300 bpm
3.20.6 | Asystole alarm
3.20.7 | Arrhythmia alarms
3.20.8 | Apnea alarm
3.20.9 | Alarms for faulty/not-connected electrode/probe
3.20.10 | User selectable alarm limits for systolic, diastolic and mean pressures
3.20.11 | Alarm for incorrect NIBPcuff connection
3.20.12 | Adjustable SpO2 high and low alarms:
3.20.12a | High : 70 - 100%
3.20.12b | Low : 50 - 95%
3.20.13 | Adjustable high and low temperature alarms
3.20.14 | Violation of limits must result in both audible and visible alarms and indicate which parameter is affected.
3.20.15 | The sound level of the audible alarm must be adjustable.
3.20.16 | It must be possible to temporarily silence the audible alarm but not remove the visible alarm.
3.20.17 | Duration of the silenced period: 2 minutes.
3.20.18 | Low battery alarm
3.21 | GENERAL
3.21.1 | Must operate from 220V 50 Hz mains power
3.21.2 | Mains cable at least four metres long with SABS approved 15A three prong plug.
3.21.3 | All cables and sockets must be colour coded, labelled or otherwise prevented from being incorrectly coupled
3.21.4 | The unit must be handed over in full operating order
3.21.5 | A starter-pack of consumables must be supplied with the unit.
4 | Standard Accessories and Consumables
4.1 | Paediatric 3/5 lead ECG cables
4.2 | Adult 3/5 lead ECG cables
4.3 | ECG interface cable
4.4 | Reusable latex-free paediatric cuff
4.5 | Reusable latex-free adult cuff
4.6 | Reusable latex-free neonate cuff
4.7 | Reusable latex-free obese cuff
4.8 | NIBP hose
4.9 | Finger sensor for paeds
4.10 | Finger sensor for adults
4.11 | Finger wrap-around probe for neonates
4.12 | Spo2 interface cable
4.13 | Two reusable temperature probes
5 | Conformity Compliance (Please attach certificate)
5.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
5.2 | OEM must comply and certified on ISO 9001 quality standards, attach proof of compliance
5.2.1 | Please provide unique ref number of the ISO 9001 ceriticate:
5.3 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
5.3.1 | Please provide unique ref number of the ISO 13485 ceriticate:
5.4 | Model quoted for must be EC certified. Attach a copy of certification
5.4.1 | Please provide unique ref number of the EC ceriticate:
5.5 | Hazardous Substance Act: If this type of equipment / apparatus appears on the schedule of Hazardous substances issued by the Directorate: Health Technology of the Department of Health, a license in terms of the Act on Hazardous Substances (Act 15/1973) must be submitted with the bid / quotation documents. The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.6 | All electrical/electronic medical equipment must be licensed by Radiation Control, where applicable a copy of the license must be submitted The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.6.1 | Radiation Control License number:
5.7 | The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade (rubber to with 2 screws) type and it must be a minimum length of three (3) meters. NB the mains cable of the unit tendered for must be SABS colour coded.
5.8 | The equipment quoted for must be protected against electromagnetic interference IEC 60601-1-2
5.9 | The unit must comply with electrical safety and performance standard such as IEC 60601-2-49 for Patient Monitoring, attached certification
5.9.1 | Please provide unique ref number of the IEC ceriticate:
6 | COMPREHENSIVE MAINTENANCE
6.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) must be included in the bid price
6.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
6.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
6.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
6.3 | The bidder to provide the OEM recommended service interval of the unit:
6.4 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
7 | General
7.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
7.2 | The unit on offer must be the latest technology.
BASIC MONITOR ACCESSORIES and CONSUMABLES
RT2-05-14a | Paediatric 3/5 lead ECG cables
RT2-05-14b | Adult 3/5 lead ECG cables
RT2-05-14c | ECG interface cable
RT2-05-14d | Reusable latex-free paediatric cuff
RT2-05-14e | Reusable latex-free adult cuff
RT2-05-14f | Reusable latex-free neonate cuff
RT2-05-14g | Reusable latex-free obese cuff
RT2-05-14h | NIBP hose
RT2-05-14i | Finger sensor for paeds
RT2-05-14j | Finger sensor for adults
RT2-05-14k | Finger wrap-around probe for neonates
RT2-05-14l | Spo2 interface cable
RT2-05-14m | Two reusable temperature probes
RT2-05-14n | Monitor Front bezel
RT2-05-14o | Battery
RT2-05-14p | NIBP pump module
RT2-05-14q | Wall bracket, height adjustable on which the monitor can be mounted.
RT2-05-14r | Roll stand
Bidder to list any other Consumables or Accessories (Not listed above) for the model on offer below

## Monitor, ICU
RT2-05-15 | MONITOR, ICU
1 | Description
A physiological monitor for use in intensive care settings
2 | Product Offered:
2.1 | Make:
2.2 | Model:
2.3 | Initial Manufacturing year of Model:
2.4 | Radiation Control License number:
3 | Technical Specification
3.1 | The monitor must be compatible with various brands of Central Station monitoring. Bidder to list brands compatible brands
3.2 | Modular, state of the art technology, suitable for neonates to adult
3.3 | Microprocessor controlled with automatic self-test and calibration
3.4 | The following parameters must be measured and displayed :
3.4.1 | ECG waveform
3.4.2 | Oxygen saturation, waveform + digital
3.4.3 | Heart rate, digital
3.4.4 | Respiration, waveform + digital
3.4.5 | 3X Invasive pressure, waveform + digital (systolic, diastolic, mean)
3.4.6 | NIBP, digital (systolic, diastolic, mean)
3.4.7 | Temperature, digital
3.4.8 | Side-stream Capnography highly desireable. Provide full information
3.5 | The monitor must operate off both 220 V, 50 Hz single phase AC power supply and an internal rechargeable battery.
3.6 | If mains power fails the monitor must operate off battery power for at least 2 hours.
3.7 | Visible indication to show whether the monitor is operating from the mains or battery power.
3.8 | All control keys must be splashproof and resistant to the entry of fluids into the unit.
3.9 | The front panel must be easy to clean and disinfect.
3.10 | High resolution LCD colour display, 45 cm diagonal minimum
3.11 | At least 6 waveforms must be displayed
3.11.1 | Digits: large: 10 mm, small: 5 mm
3.11.2 | Sweep speeds : 25 and 50 mm/s
3.11.3 | It must be possible to freeze the display. This must not affect the QRS signal.
3.12 | Trending for 72 hours must be possible on the monitor,
3.13 | Unit to provide at least 72 hours of full disclosure wave data for review
3.14 | Defibrillator-proof up to 360 Joules
3.15 | Protected against HF interference
3.16 | Patient isolated from all potentially harmful circuits
3.17 | Must comply with IEC 60601-1
3.18 | ECG
3.18.1 | Must comply with IEC 60601-2-27
3.18.2 | ECG 12 lead with arrhythmia detection
3.18.3 | Adult ECG lead must be supplied
3.18.4 | The amplitude of the ECG trace must be adjustable from 1 mV = 5 mm to 1 mV = 20 mm.
3.18.5 | Digital display of heart rate : 30 - 300 bpm
3.18.6 | Accuracy : 3 bpm
3.18.7 | Resolution : 1 bpm
3.18.8 | Differential input with CMRR > 80 dB
3.18.9 | Input impedance > 10 MOhm
3.18.10 | Respiration monitoring through ECG electrodes
3.18.11 | Pacemaker detection/rejection : 2 to 500 mV; 0,1 to 2 ms
3.18.12 | Synch pulse for defibrillator
3.19 | OXYGEN SATURATION
3.19.1 | It must be the latest technology with clinical research support based on accuracy and real-time monitoring. Please submit the published peer reviewed/researched paper as supporting documents.
3.19.2 | The unit should indicate whether the signal from the pulse oximeter sensor is adequate for reliable measurements. If the signal strength or quality is inadequate, this must be clearly displayed.
3.19.3 | An audible bleep must sound for each heart beat. The pitch must vary with the oxygen saturation.
3.19.4 | Measurement range : 0 - 100%
3.19.5 | Accuracy:
3.19.5a | 80 - 100% : 1,5%
3.19.5b | 60 - 80% : 3,5%
3.19.6 | Continuous waveform display
3.19.7 | Digital display of SpO2%
3.19.8 | The sensor lead must be at least 3 metres long.
3.19.9 | Automatic calibration for skin colour
3.19.10 | Pulse rate: 25 - 250 bpm
3.20 | NON-INVASIVE BLOOD PRESSURE
3.20.1 | Must comply with IEC 60601-2-30
3.20.2 | Oscillometric measurement method
3.20.3 | Automatic, manual and stat modes
3.20.4 | Measure and display :
3.20.4a | - systolic pressure
3.20.4b | - diastolic pressure
3.20.4c | - mean pressure
3.20.5 | The interval between measurements must be selectable from 1 to 60 minutes.
3.20.6 | Ranges (approximate) :
3.20.6a | Systolic : 30 - 270 mmHg. Resolution : 1 mmHg
3.20.6b | Diastolic : 10 - 150 mmHg. Resolution : 1 mmHg
3.20.6c | Mean : 20 - 250 mmHg. Resolution : 1 mmHg
3.20.7 | Cuff pressure : 0 - 280 mmHg
3.20.8 | Cuff inflation time < 10 s
3.20.9 | Automatic deflation if measurement exceeds 3 min or if mains power fails
3.21 | INVASIVE BLOOD PRESSURE
3.21.1 | Must comply with IEC 60601-2-34
3.21.2 | Measure and display :
3.21.2a | - systolic pressure
3.21.2b | - diastolic pressure
3.21.2c | - mean pressure
3.21.3 | Systolic : 0 - 300 mmHg. Resolution : 1 mmHg
3.21.4 | Diastolic : 0 - 250 mmHg. Resolution : 1 mmHg
3.21.5 | Mean : 0 - 250 mmHg. Resolution : 1 mmHg
3.22 | TEMPERATURE
3.22.1 | Range : 10 - 45 ˚C
3.22.2 | Digital display
3.22.3 | Surface and oral probe to be supplied
3.22.4 | Accuracy: 0.3 ˚C
3.22.5 | Resolution : 0.1 ˚C
3.23 | ALARMS
3.23.1 | Visible and audio alarms
3.23.2 | The monitor must automatically set default alarms when switched on.
3.23.3 | The alarms must be programmable
3.23.4 | Alarm limits must be displayed with the parameter digit
3.23.5 | Heart rate alarm : 30 - 300 bpm
3.23.6 | Asystole alarm
3.23.7 | Arrhythmia alarms
3.23.8 | Apnea alarm
3.23.9 | Alarms for faulty/not-connected electrode/probe
3.23.10 | User selectable alarm limits for systolic, diastolic and mean pressures
3.23.11 | Alarm for incorrect NIBPcuff connection
3.23.12 | Adjustable SpO2 high and low alarms:
3.23.12a | High : 70 - 100%
3.23.12b | Low : 50 - 95%
3.23.13 | Adjustable high and low temperature alarms
3.23.14 | Violation of limits must result in both audible and visible alarms and indicate which parameter is affected.
3.23.15 | The sound level of the audible alarm must be adjustable.
3.23.16 | It must be possible to temporarily silence the audible alarm but not remove the visible alarm.
3.23.17 | Duration of the silenced period: 2 minutes.
3.23.18 | Low battery alarm
3.24 | GENERAL
3.24.1 | Mains cable at least four metres long with SABS approved 15A three prong plug.
3.24.2 | All cables and sockets must be colour coded, labelled or otherwise prevented from being incorrectly coupled
3.24.3 | The unit must be handed over in full operating order
3.24.4 | A starter-pack of consumables must be supplied with the unit.
4 | Standard Accessories and Consumables
4.1 | Paediatric 3/5 lead ECG cables
4.2 | Adult 3/5 lead ECG cables
4.3 | ECG interface cable
4.4 | Reusable latex-free paediatric cuff
4.5 | Reusable latex-free adult cuff
4.6 | Reusable latex-free neonate cuff
4.7 | Reusable latex-free obese cuff
4.8 | NIBP hose
4.9 | Finger sensor for paeds
4.10 | Finger sensor for adults
4.11 | Finger wrap-around probe for neonates
4.12 | Spo2 interface cable
4.13 | Two reusable temperature probes
5 | Conformity Compliance (Please attach certificate)
5.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
5.2 | OEM must comply and certified on ISO 9001 quality standards, attach proof of compliance
5.2.1 | Please provide unique ref number of the ISO 9001 ceriticate:
5.3 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
5.3.1 | Please provide unique ref number of the ISO 13485 ceriticate:
5.4 | Model quoted for must be EC certified. Attach a copy of certification
5.4.1 | Please provide unique ref number of the EC ceriticate:
5.5 | Hazardous Substance Act: If this type of equipment / apparatus appears on the schedule of Hazardous substances issued by the Directorate: Health Technology of the Department of Health, a license in terms of the Act on Hazardous Substances (Act 15/1973) must be submitted with the bid / quotation documents. The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.6 | All electrical/electronic medical equipment must be licensed by Radiation Control, where applicable a copy of the license must be submitted The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.6.1 | Radiation Control License number:
5.7 | The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade (rubber to with 2 screws) type and it must be a minimum length of three (3) meters. NB the mains cable of the unit tendered for must be SABS colour coded.
5.8 | The equipment quoted for must be protected against electromagnetic interference IEC 60601-1-2
5.9 | The unit must comply with electrical safety and performance standard such as IEC 60601-2-49 for Patient Monitoring, attached certification
5.9.1 | Please provide unique ref number of the IEC ceriticate:
6 | COMPREHENSIVE MAINTENANCE
6.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) must be included in the bid price
6.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
6.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
6.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
6.3 | The bidder to provide the OEM recommended service interval of the unit:
6.4 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
7 | General
7.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
7.2 | The unit on offer must be the latest technology.
ICU MONITOR ACCESSORIES and CONSUMABLES
RT2-05-15a | Paediatric 3/5 lead ECG cables
RT2-05-15b | Adult 3/5 lead ECG cables
RT2-05-15c | ECG interface cable
RT2-05-15d | Reusable latex-free paediatric cuff
RT2-05-15e | Reusable latex-free adult cuff
RT2-05-15f | Reusable latex-free neonate cuff
RT2-05-15g | Reusable latex-free obese cuff
RT2-05-15h | NIBP hose
RT2-05-15i | Finger sensor for paeds
RT2-05-15j | Finger sensor for adults
RT2-05-15k | Finger wrap-around probe for neonates
RT2-05-15l | Spo2 interface cable
RT2-05-15m | Two reusable temperature probes
RT2-05-15n | Monitor Front bezel
RT2-05-15o | Battery
RT2-05-15p | NIBP pump module
RT2-05-15q | Wall bracket, height adjustable on which the monitor can be mounted.
RT2-05-15r | Roll stand
RT2-05-15s | EEG (UPGRADE)
RT2-05-15t | Intra-cranial pressure (UPGRADE)
RT2-05-15u | Cardiac output (UPGRADE)
RT2-05-15v | CO2/EtCO2 (UPGRADE)
RT2-05-15w | Built-in Printer (UPGRADE)
RT2-05-15x | Apnea monitor for neonatal use (UPGRADE)
Bidder to list any other Consumables or Accessories (Not listed above) for the model on offer below

## Monitor, ICU, wit Module
RT2-05-16 | MONITOR, ICU, with Transport Module
1 | Description
1.1 | A physiological monitor for use in intensive care settings with detachable module for transportation without losing trends
2 | Product Offered:
2.1 | Make:
2.2 | Model:
2.3 | Initial Manufacturing year of Model:
2.4 | Radiation Control License number:
3 | Technical Specification
3.1 | The monitor must be compatible with various brands of Central Station monitoring. Bidder to list brands compatible brands
3.2 | Modular, state of the art technology, suitable for neonates to adult
3.3 | Microprocessor controlled with automatic self-test and calibration
3.4 | The following parameters must be measured and displayed :
3.4.1 | ECG waveform
3.4.2 | Oxygen saturation, waveform + digital
3.4.3 | Heart rate, digital
3.4.4 | Respiration, waveform + digital
3.4.5 | 3X Invasive pressure, waveform + digital (systolic, diastolic, mean)
3.4.6 | NIBP, digital (systolic, diastolic, mean)
3.4.7 | Temperature, digital
3.4.8 | Side-stream Capnography highly desireable. Provide full information
3.5 | The monitor must operate off both 220 V, 50 Hz single phase AC power supply and an internal rechargeable battery.
3.6 | If mains power fails the monitor must operate off battery power for at least 2 hours.
3.7 | Visible indication to show whether the monitor is operating from the mains or battery power.
3.8 | All control keys must be splashproof and resistant to the entry of fluids into the unit.
3.9 | The front panel must be easy to clean and disinfect.
3.10 | High resolution LCD colour display, 45 cm diagonal minimum
3.11 | At least 6 waveforms must be displayed
3.11.1 | Digits: large: 10 mm, small: 5 mm
3.11.2 | Sweep speeds : 25 and 50 mm/s
3.11.3 | It must be possible to freeze the display. This must not affect the QRS signal.
3.12 | Trending for 72 hours must be possible on the monitor,
3.13 | Unit to provide at least 72 hours of full disclosure wave data for review
3.14 | Defibrillator-proof up to 360 Joules
3.15 | Protected against HF interference
3.16 | Patient isolated from all potentially harmful circuits
3.17 | The monitor must have detachable module
3.17.1 | Module must have a minimum 10cm screen
3.17.2 | Parameters ports must be located on the module to avoid any reconnection for transportation
3.17.3 | Automatic display parameters on the monitor screen once attached
3.17.4 | Module must have a battery capacity of at least 4 hours
3.18 | ECG
3.18.1 | Must comply with IEC 60601-2-27
3.18.2 | ECG 12 lead with arrhythmia detection
3.18.3 | Adult ECG lead must be supplied
3.18.4 | The amplitude of the ECG trace must be adjustable from 1 mV = 5 mm to 1 mV = 20 mm.
3.18.5 | Digital display of heart rate : 30 - 300 bpm
3.18.6 | Accuracy : 3 bpm
3.18.7 | Resolution : 1 bpm
3.18.8 | Differential input with CMRR > 80 dB
3.18.9 | Input impedance > 10 MOhm
3.18.10 | Respiration monitoring through ECG electrodes
3.18.11 | Pacemaker detection/rejection : 2 to 500 mV; 0,1 to 2 ms
3.18.12 | Synch pulse for defibrillator
3.19 | OXYGEN SATURATION
3.19.1 | It must be the latest technology with clinical research support based on accuracy and real-time monitoring. Please submit the published peer reviewed/researched paper as supporting documents.
3.19.2 | The unit should indicate whether the signal from the pulse oximeter sensor is adequate for reliable measurements. If the signal strength or quality is inadequate, this must be clearly displayed.
3.19.3 | An audible bleep must sound for each heart beat. The pitch must vary with the oxygen saturation.
3.19.4 | Measurement range : 0 - 100%
3.19.5 | Accuracy:
3.19.5a | 80 - 100% : 1,5%
3.19.5b | 60 - 80% : 3,5%
3.19.6 | Continuous waveform display
3.19.7 | Digital display of SpO2%
3.19.8 | The sensor lead must be at least 3 metres long.
3.19.9 | Automatic calibration for skin colour
3.19.10 | Pulse rate: 25 - 250 bpm
3.20 | NON-INVASIVE BLOOD PRESSURE
3.20.1 | Must comply with IEC 60601-2-30
3.20.2 | Oscillometric measurement method
3.20.3 | Automatic, manual and stat modes
3.20.4 | Measure and display :
3.20.4a | - systolic pressure
3.20.4b | - diastolic pressure
3.20.4c | - mean pressure
3.20.5 | The interval between measurements must be selectable from 1 to 60 minutes.
3.20.6 | Ranges (approximate) :
3.20.6a | Systolic : 30 - 270 mmHg. Resolution : 1 mmHg
3.20.6b | Diastolic : 10 - 150 mmHg. Resolution : 1 mmHg
3.20.6c | Mean : 20 - 250 mmHg. Resolution : 1 mmHg
3.20.7 | Cuff pressure : 0 - 280 mmHg
3.20.8 | Cuff inflation time < 10 s
3.20.9 | Automatic deflation if measurement exceeds 3 min or if mains power fails
3.21 | INVASIVE BLOOD PRESSURE
3.21.1 | Must comply with IEC 60601-2-34
3.21.2 | Measure and display :
3.21.2a | - systolic pressure
3.21.2b | - diastolic pressure
3.21.2c | - mean pressure
3.21.3 | Systolic : 0 - 300 mmHg. Resolution : 1 mmHg
3.21.4 | Diastolic : 0 - 250 mmHg. Resolution : 1 mmHg
3.21.5 | Mean : 0 - 250 mmHg. Resolution : 1 mmHg
3.22 | TEMPERATURE
3.22.1 | Range : 10 - 45 ˚C
3.22.2 | Digital display
3.22.3 | Surface and oral probe to be supplied
3.22.4 | Accuracy: 0.3 ˚C
3.22.5 | Resolution : 0.1 ˚C
3.23 | ALARMS
3.23.1 | Visible and audio alarms
3.23.2 | The monitor must automatically set default alarms when switched on.
3.23.3 | The alarms must be programmable
3.23.4 | Alarm limits must be displayed with the parameter digit
3.23.5 | Heart rate alarm : 30 - 300 bpm
3.23.6 | Asystole alarm
3.23.7 | Arrhythmia alarms
3.23.8 | Apnea alarm
3.23.9 | Alarms for faulty/not-connected electrode/probe
3.23.10 | User selectable alarm limits for systolic, diastolic and mean pressures
3.23.11 | Alarm for incorrect NIBPcuff connection
3.23.12 | Adjustable SpO2 high and low alarms:
3.23.12a | High : 70 - 100%
3.23.12b | Low : 50 - 95%
3.23.13 | Adjustable high and low temperature alarms
3.23.14 | Violation of limits must result in both audible and visible alarms and indicate which parameter is affected.
3.23.15 | The sound level of the audible alarm must be adjustable.
3.23.16 | It must be possible to temporarily silence the audible alarm but not remove the visible alarm.
3.23.17 | Duration of the silenced period: 2 minutes.
3.23.18 | Low battery alarm
3.24 | GENERAL
3.24.1 | Mains cable at least four metres long with SABS approved 15A three prong plug.
3.24.2 | All cables and sockets must be colour coded, labelled or otherwise prevented from being incorrectly coupled
3.24.3 | The unit must be handed over in full operating order
3.24.4 | A starter-pack of consumables must be supplied with the unit.
4 | Standard Accessories and Consumables
4.1 | Paediatric 3/5 lead ECG cables
4.2 | Adult 3/5 lead ECG cables
4.3 | ECG interface cable
4.4 | Reusable latex-free paediatric cuff
4.5 | Reusable latex-free adult cuff
4.6 | Reusable latex-free neonate cuff
4.7 | Reusable latex-free obese cuff
4.8 | NIBP hose
4.9 | Finger sensor for paeds
4.10 | Finger sensor for adults
4.11 | Finger wrap-around probe for neonates
4.12 | Spo2 interface cable
4.13 | Two reusable temperature probes
5 | Conformity Compliance (Please attach certificate)
5.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
5.2 | OEM must comply and certified on ISO 9001 quality standards, attach proof of compliance
5.2.1 | Please provide unique ref number of the ISO 9001 ceriticate:
5.3 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
5.3.1 | Please provide unique ref number of the ISO 13485 ceriticate:
5.4 | Model quoted for must be EC certified. Attach a copy of certification
5.4.1 | Please provide unique ref number of the EC ceriticate:
5.5 | Hazardous Substance Act: If this type of equipment / apparatus appears on the schedule of Hazardous substances issued by the Directorate: Health Technology of the Department of Health, a license in terms of the Act on Hazardous Substances (Act 15/1973) must be submitted with the bid / quotation documents. The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.6 | All electrical/electronic medical equipment must be licensed by Radiation Control, where applicable a copy of the license must be submitted The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.6.1 | Radiation Control License number:
5.7 | The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade (rubber to with 2 screws) type and it must be a minimum length of three (3) meters. NB the mains cable of the unit tendered for must be SABS colour coded.
5.8 | The equipment quoted for must be protected against electromagnetic interference IEC 60601-1-2
5.9 | The unit must comply with electrical safety and performance standard such as IEC 60601-2-49 for Patient Monitoring, attached certification
5.9.1 | Please provide unique ref number of the IEC ceriticate:
6 | COMPREHENSIVE MAINTENANCE
6.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) must be included in the bid price
6.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
6.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
6.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
6.3 | The bidder to provide the OEM recommended service interval of the unit:
6.4 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
7 | General
7.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
7.2 | The unit on offer must be the latest technology.
ICU MONITOR with Transport Module ACCESSORIES and CONSUMABLES
RT2-05-16a | Paediatric 3/5 lead ECG cables
RT2-05-16b | Adult 3/5 lead ECG cables
RT2-05-16c | ECG interface cable
RT2-05-16d | Reusable latex-free paediatric cuff
RT2-05-16e | Reusable latex-free adult cuff
RT2-05-16f | Reusable latex-free neonate cuff
RT2-05-16g | Reusable latex-free obese cuff
RT2-05-16h | NIBP hose
RT2-05-16i | Finger sensor for paeds
RT2-05-16j | Finger sensor for adults
RT2-05-16k | Finger wrap-around probe for neonates
RT2-05-16l | Spo2 interface cable
RT2-05-16m | Two reusable temperature probes
RT2-05-16n | Monitor Front bezel
RT2-05-16o | Battery
RT2-05-16p | NIBP pump module
RT2-05-16q | Wall bracket, height adjustable on which the monitor can be mounted.
RT2-05-16r | Roll stand
RT2-05-16s | EEG (UPGRADE)
RT2-05-16t | Intra-cranial pressure (UPGRADE)
RT2-05-16u | Cardiac output (UPGRADE)
RT2-05-16v | CO2/EtCO2 (UPGRADE)
RT2-05-16w | Built-in Printer (UPGRADE)
RT2-05-16x | Apnea monitor for neonatal use (UPGRADE)
Bidder to list any other Consumables or Accessories (Not listed above) for the model on offer below

## Central Monitoring, 8-bed
RT2-05-17 | Central Patient Monitoring, 8-Bed ICU Package
1 | Description
This package includes supply and installation of ICU central station, 8 ICU monitors
2a | Product Offered (Central Station):
2a.1 | Make:
2a.2 | Model:
2a.3 | Brief Description:
2b | Product Offered (Monitor):
2b.1 | Make:
2b.2 | Model:
2b.3 | Initial Manufacturing year of Model:
3 | Technical Specification
3a | Central Station
3a.1 | The central station must have a Single display, and 8 ICU monitors with complete accessories.
3a.2 | The unit must be network configurable and wireless.
3a.3 | Must have a viewing screen which will also do patient management and reviews and size must be at least 480 mm colour LCD
3a.4 | Touchscreen (Optional)
3a.5 | The unit must be able to remotely manage patient bedside monitors, including viewing active or historic data.
3a.6 | Password protection must be available to protect the critical menus and certain settings.
3a.7 | All waveforms and alphanumeric displays must be available at the central station depending on the parameters of the bedside monitor
3a.8 | Bi-directional comunication between monitors and central station
3a.9 | The central station must be able to download all the patient vitals after the patient has been disconnected and reconnected
3a.10 | Display at least 8 beds information
3a.11 | Full arrhythmia recall and ST segment analysis with 12 lead ECG interpretation
3a.12 | Screen configurations must be able to be saved as favourites for easy access
3a.13 | Up to to 10 arrhythmia events per patient monitoring session must be available
3a.14 | Must be able to review full disclosure on the central station for the duration of the disclosure time. At least 120 hours
3a.15 | The central station must allow access to patient demographic information; information arrhythmias, ST-limit violations, and ECG samples in event review; parameter data in graphic and numeric trends; full disclosure data in strip view, page view with zoom, or in calipers.
3a.16 | Alarms
3a.16.1 | All alarms must be able to be muted only at the bedside
3a.16.2 | 3-level alarm must be available audible & visual.
3a.16.3 | The sound levels and colour to differentiate alarm type
3a.16.4 | High and low limits of alarm must be offered
3a.16.5 | Monitor alarm limits must be adjustable from the central station. This setting must be enabled/disabled via a password protected menu.
3b | Monitor
3b.1 | Modular, state of the art technology, suitable for neonates to adult
3b.2 | Microprocessor controlled with automatic self-test and calibration
3b.3 | The following parameters must be measured and displayed :
3b.3.1 | ECG waveform
3b.3.2 | Oxygen saturation, waveform + digital
3b.3.3 | Heart rate, digital
3b.3.4 | Respiration, waveform + digital
3b.3.5 | Minimum 3 Invasive pressure, waveform + digital (systolic, diastolic, mean)
3b.3.6 | NIBP, digital (systolic, diastolic, mean)
3b.3.7 | Temperature, digital
3b.3.8 | Side-stream Capnography highly desireable. Provide full information
3b.4 | The monitor must operate off both 220 V, 50 Hz single phase AC power supply and an internal rechargeable battery.
3b.5 | If mains power fails the monitor must operate off battery power for at least 2 hours.
3b.6 | Visible indication to show whether the monitor is operating from the mains or battery power.
3b.7 | All control keys must be splashproof and resistant to the entry of fluids into the unit.
3b.8 | The front panel must be easy to clean and disinfect.
3b.9 | High resolution LCD colour display, 35 cm diagonal minimum
3b.10 | At least 6 waveforms must be displayed
3b.10.1 | Digits: large: 10 mm, small: 5 mm
3b.10.2 | Sweep speeds : 25 and 50 mm/s
3b.10.3 | It must be possible to freeze the display. This must not affect the QRS signal.
3b.11 | Trending for 72 hours must be possible on the monitor,
3b.12 | Unit to provide at least 72 hours of full disclosure wave data for review
3b.13 | Defibrillator-proof up to 360 Joules
3b.14 | Protected against HF interference
3b.15 | Patient isolated from all potentially harmful circuits
3b.16 | ECG
3b.16.1 | Must comply with IEC 60601-2-27
3b.16.2 | ECG 12 lead with arrhythmia detection
3b.16.3 | Adult ECG lead must be supplied
3b.16.4 | The amplitude of the ECG trace must be adjustable from 1 mV = 5 mm to 1 mV = 20 mm.
3b.16.5 | Digital display of heart rate : 30 - 300 bpm
3b.16.6 | Accuracy : 3 bpm
3b.16.7 | Resolution : 1 bpm
3b.16.8 | Differential input with CMRR > 80 dB
3b.16.9 | Input impedance > 10 MOhm
3b.16.10 | Respiration monitoring through ECG electrodes
3b.16.11 | Pacemaker detection/rejection : 2 to 500 mV; 0,1 to 2 ms
3b.16.12 | Synch pulse for defibrillator
3b.17 | OXYGEN SATURATION
3b.17.1 | It must be the latest technology with clinical research support based on accuracy and real-time monitoring. Please submit the published peer reviewed/researched paper as supporting documents.
3b.17.2 | The unit should indicate whether the signal from the pulse oximeter sensor is adequate for reliable measurements. If the signal strength or quality is inadequate, this must be clearly displayed.
3b.17.3 | An audible bleep must sound for each heart beat. The pitch must vary with the oxygen saturation.
3b.17.4 | Measurement range : 0 - 100%
3b.17.5 | Accuracy:
3b.17.5a | 80 - 100% : 1,5%
3b.17.5b | 60 - 80% : 3,5%
3b.17.6 | Continuous waveform display
3b.17.7 | Digital display of SpO2%
3b.17.8 | The sensor lead must be at least 3 metres long.
3b.17.9 | Automatic calibration for skin colour
3b.17.10 | Pulse rate: 25 - 250 bpm
3b.18 | NON-INVASIVE BLOOD PRESSURE
3b.18.1 | Must comply with IEC 60601-2-30
3b.18.2 | Oscillometric measurement method
3b.18.3 | Automatic, manual and stat modes
3b.18.4 | Measure and display :
3b.18.4a | - systolic pressure
3b.18.4b | - diastolic pressure
3b.18.4c | - mean pressure
3b.18.5 | The interval between measurements must be selectable from 1 to 60 minutes.
3b.18.6 | Ranges (approximate) :
3b.18.6a | Systolic : 30 - 270 mmHg. Resolution : 1 mmHg
3b.18.6b | Diastolic : 10 - 150 mmHg. Resolution : 1 mmHg
3b.18.6c | Mean : 20 - 250 mmHg. Resolution : 1 mmHg
3b.18.7 | Cuff pressure : 0 - 280 mmHg
3b.18.8 | Cuff inflation time < 10 s
3b.18.9 | Automatic deflation if measurement exceeds 3 min or if mains power fails
3b.19 | INVASIVE BLOOD PRESSURE
3b.19.1 | Must comply with IEC 60601-2-34
3b.19.2 | Measure and display :
3b.19.2a | - systolic pressure
3b.19.2b | - diastolic pressure
3b.19.2c | - mean pressure
3b.19.3 | Systolic : 0 - 300 mmHg. Resolution : 1 mmHg
3b.19.4 | Diastolic : 0 - 250 mmHg. Resolution : 1 mmHg
3b.19.5 | Mean : 0 - 250 mmHg. Resolution : 1 mmHg
3b.20 | TEMPERATURE
3b.20.1 | Range : 10 - 45 ˚C
3b.20.2 | Digital display
3b.20.3 | Surface and esophageal probe to be supplied
3b.20.4 | Accuracy: 0.3 ˚C
3b.20.5 | Resolution : 0.1 ˚C
3b.21 | ALARMS
3b.21.1 | Visible and audio alarms
3b.21.2 | The monitor must automatically set default alarms when switched on.
3b.21.3 | The alarms must be programmable
3b.21.4 | Alarm limits must be displayed with the parameter digit
3b.21.5 | Heart rate alarm : 30 - 300 bpm
3b.21.6 | Asystole alarm
3b.21.7 | Arrhythmia alarms
3b.21.8 | Apnea alarm
3b.21.9 | Alarms for faulty/not-connected electrode/probe
3b.21.10 | User selectable alarm limits for systolic, diastolic and mean pressures
3b.21.11 | Alarm for incorrect NIBPcuff connection
3b.21.12 | Adjustable SpO2 high and low alarms:
3b.21.12a | High : 70 - 100%
3b.21.12b | Low : 50 - 95%
3b.21.13 | Adjustable high and low temperature alarms
3b.21.14 | Violation of limits must result in both audible and visible alarms and indicate which parameter is affected.
3b.21.15 | The sound level of the audible alarm must be adjustable.
3b.21.16 | It must be possible to temporarily silence the audible alarm but not remove the visible alarm.
3b.21.17 | Duration of the silenced period: 2 minutes.
3b.21.18 | Low battery alarm
3b.22 | Printer
3b.22.1 | Paper size must be standard A4 off the shelf paper
3b.22.2 | With the capability of full disclosure waveform printing, multi-wave printing, tabular trend printing and trend graph printing.
3b.22.3 | Printing must be possible at both bedside monitor and central station
3b.23 | Data Storage and Review
3b.23.1 | Trends of at least 120 hours must be available
3b.23.2 | Trends must be displayed graphically and in tabular format
3b.23.3 | Full disclosure of at least 144 hours to be offered
3b.23.4 | Waveform storage must be available for at least 144 hours post discharge
3b.23.5 | Events must be able to be sorted by time and type
3b.23.6 | All patient events, graphical and numerical data and alarms must be available for viewing on a first in first out basis.
3b.24 | General
3b.24.1 | Network installation and cabling will be done by the bidder and included in the tender price.
3b.24.2 | Height adjustable brackets. Monitor to be adjustable to a lower level than the bottom of the wall mounted baseplate to accommodate shorter staff members.
3b.24.3 | The bidder shall provide fixtures and wall fittings for the entire system including network cables and switches.
3b.24.4 | The bid includes the license and any software upgrades
3b.24.5 | The unit must be handed over in full operating order
4 | The following accessories must be suitable for the patient size identified and supplied for each monitor:
4.1 | NIBP cuff with hose
4.2 | SpO2 finger probe with cable
4.3 | Reusable temperature sensor
4.4 | ECG cable, 3/5 lead
4.5 | Invasive blood pressure cable X2 with connection for tranceducers used by user
4.6 | The unit to be supplied with the UPS (uninterrupted power supply)
5 | Conformity Compliance (Please attach certificate)
5.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
5.2 | OEM must comply and certified on ISO 9001 quality standards, attach proof of compliance
5.2.1 | Please provide unique ref number of the ISO 9001 ceriticate:
5.3 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
5.3.1 | Please provide unique ref number of the ISO 13485 ceriticate:
5.4 | Model quoted for must be EC certified. Attach a copy of certification
5.4.1 | Please provide unique ref number of the EC ceriticate:
5.5 | Hazardous Substance Act: If this type of equipment / apparatus appears on the schedule of Hazardous substances issued by the Directorate: Health Technology of the Department of Health, a license in terms of the Act on Hazardous Substances (Act 15/1973) must be submitted with the bid / quotation documents. The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.6 | All electrical/electronic medical equipment must be licensed by Radiation Control, where applicable a copy of the license must be submitted The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.6.1 | Radiation Control License number:
5.7 | The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade (rubber to with 2 screws) type and it must be a minimum length of three (3) meters. NB the mains cable of the unit tendered for must be SABS colour coded.
5.8 | The equipment quoted for must be protected against electromagnetic interference IEC 60601-1-2
5.9 | The unit must comply with electrical safety and performance standard such as IEC 60601-2-49 for Patient Monitoring, attached certification
5.9.1 | Please provide unique ref number of the IEC ceriticate:
6 | COMPREHENSIVE MAINTENANCE
6.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) must be included in the bid price
6.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
6.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
6.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
6.3 | The bidder to provide the OEM recommended service interval of the unit:
6.4 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
7 | General
7.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
7.2 | The unit on offer must be the latest technology.
RT2-05-17a | Central Patient Monitoring, 12-Bed ICU Package
This package includes supply and installation of ICU central station, 12 ICU monitors
Product Offered (Central Station):
Make:
Model:
Brief Description:
Product Offered (Monitor):
Make:
Model:
Initial Manufacturing year of Model:
1 | Technical Specification
1.1 | The central station must have a Single display, and 12 ICU monitors with complete accessories.
1.2 | Must comply with the specification of item (8-bed ICU package)
RT2-05-17b | Central Patient Monitoring, 16-Bed ICU Package
This package includes supply and installation of ICU central station, 16 ICU monitors
Product Offered (Central Station):
Make:
Model:
Brief Description:
Product Offered (Monitor):
Make:
Model:
Initial Manufacturing year of Model:
1 | Technical Specification
1.1 | The central station must have two (2) screens, and 16 ICU monitors with complete accessories.
1.2 | Must comply with the specification of item (8-bed ICU package)

## Central Station
RT2-05-18 | Central Station [8 bed]
1 | Description
This includes supply and installation of monitor-ready central station
2 | Product Offered (Central Station):
Make:
Model:
Brief Description:
3 | Technical Specification
3.1 | The offer bidded for must be Universal/Vendor Neutral Central Station allowing end-user choice of Network configurable Monitor brands for the clinical area.
3.2 | The central station must have a 2 screens.
3.3 | The unit must be network configurable and wireless.
3.4 | Must have a viewing screen as well as a review screen for measurements and both screens sizes must be at least 480 mm colour LCD
3.5 | The second screen must be able to do patient management and reviews without interruption of the live patient view on the first screen
3.6 | Touchscreen (Optional)
3.7 | The unit must be able to remotely manage patient bedside monitors, including viewing active or historic data.
3.8 | Password protection must be available to protect the critical menus and certain settings.
3.9 | All waveforms and alphanumeric displays must be available at the central station depending on the parameters of the bedside monitor
3.10 | Bi-directional comunication between monitors and central station
3.11 | The central station must be able to download all the patient vitals after the patient has been disconnected and reconnected
3.12 | Display at least 8 beds information
3.13 | Full arrhythmia recall and ST segment analysis with 12 lead ECG interpretation
3.14 | Screen configurations must be able to be saved as favourites for easy access
3.15 | Up to to 10 arrhythmia events per patient monitoring session must be available
3.16 | Must be able to review full disclosure on the central station for the duration of the disclosure time. At least 120 hours
3.17 | The central station must allow access to patient demographic information; information arrhythmias, ST-limit violations, and ECG samples in event review; parameter data in graphic and numeric trends; full disclosure data in strip view, page view with zoom, or in calipers.
3.18 | The station must be provided complete with licence [once off], switch and installation (Monitor configuration ready)
3.19 | Alarms
3.19.1 | All alarms must be able to be muted only at the bedside
3.19.2 | 3-level alarm must be available audible & visual.
3.19.3 | The sound levels and colour to differentiate alarm type
3.19.4 | High and low limits of alarm must be offered
3.19.5 | Monitor alarm limits must be adjustable from the central station. This setting must be enabled/disabled via a password protected menu.
3.20 | Printer
3.20.1 | Paper size must be standard A4 off the shelf paper
3.20.2 | With the capability of full disclosure waveform printing, multi-wave printing, tabular trend printing and trend graph printing.
3.20.3 | Printing must be possible at both bedside monitor and central station
3.21 | Data Storage and Review
3.21.1 | Trends of at least 120 hours must be available
3.21.2 | Trends must be displayed graphically and in tabular format
3.21.3 | Full disclosure of at least 144 hours to be offered
3.21.4 | Waveform storage must be available for at least 144 hours post discharge
3.21.5 | Events must be able to be sorted by time and type
3.21.6 | All patient events, graphical and numerical data and alarms must be available for viewing on a first in first out basis.
3.22 | General
3.22.1 | Network installation and cabling will be done by the bidder and included in the tender price.
3.22.2 | Height adjustable brackets. Monitor to be adjustable to a lower level than the bottom of the wall mounted baseplate to accommodate shorter staff members.
3.22.3 | The bidder shall provide fixtures and wall fittings for the entire system including network cables and switches.
3.22.4 | The bid includes the license and any software upgrades
3.22.5 | The central station must be compatible with different brands.
The unit to be supplied with the UPS (uninterrupted power supply)
5 | Conformity Compliance (Please attach certificate)
5.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
5.2 | OEM must comply and certified on ISO 9001 quality standards, attach proof of compliance
5.2.1 | Please provide unique ref number of the ISO 9001 ceriticate:
5.3 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
5.3.1 | Please provide unique ref number of the ISO 13485 ceriticate:
5.4 | Model quoted for must be EC certified. Attach a copy of certification
5.4.1 | Please provide unique ref number of the EC ceriticate:
5.7 | The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade (rubber to with 2 screws) type and it must be a minimum length of three (3) meters. NB the mains cable of the unit tendered for must be SABS colour coded.
5.8 | The equipment quoted for must be protected against electromagnetic interference IEC 60601-1-2
6 | COMPREHENSIVE MAINTENANCE
6.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) must be included in the bid price
6.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
7 | General
7.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
7.2 | The unit on offer must be the latest technology.
RT2-05-18a | Central Station [12 bed]
1 | Technical Specification
1.1 | The unit must comply with the requirements and specification of item TR2-05-18
RT2-05-18b | Central Station [16 bed]
1 | Technical Specification
1.1 | The unit must comply with the requirements and specification of item TR2-05-18

## Vital Signs, NIBP+SpO2
RT2-05-19 | Vital Signs Monitor, on trolley
1 | Description
1.1 | A device used to monitor body temperature, pulse rate, respiratory rate and blood pressure of a patient
2 | Product Offered:
2.1 | Make:
2.2 | Model:
2.3 | Initial Manufacturing year of Model:
2.4 | Radiation Control License number:
3 | Technical Specification
3.1 | The unit must measure both NIBP and SpO2
3.2 | Self test procedure at switch-on
3.3 | The unit must have trending for at least 24 hours
3.4 | NIBP
3.4.1 | The unit must measure blood pressure non-invasively.
3.4.2 | Must be suitable for adults, paeds and neonates
3.4.3 | The unit must use the oscillometric method of measurement.
3.4.4 | Modes: Manual, automatic and STAT
3.4.5 | Pressure measurements:
3.4.5a | Adult: Systolic: 30 - 270 mm Hg, Diastolic: 20 - 240 mm Hg, MAP: 20 - 250 mm Hg
3.4.5b | Paeds: Systolic: 30 - 160 mm Hg, Diastolic: 20 - 120, MAP: 20 - 120 mm Hg
3.4.5c | Neonate: Systolic: 30 - 130 mm Hg, Diastolic: 20 - 100 mm Hg, MAP: 20 - 120 mm Hg
3.4.6 | Auto zero when unit is switched on
3.4.7 | The interval between measurements must be selectable from 1 – 120 minutes.
3.4.8 | Over pressure protection: Automatic deflation at 300 mm Hg adult, 150 mm Hg neonate
3.4.9 | Average measurement time: approximately 30 s
3.4.10 | If measurement exceeds two minutes the cuff must deflate, an alarm must sound and likely faults indicated.
3.4.11 | The air hose must be at least two metres long.
3.5 | PULSE OXIMETER
3.5.1 | Suitable for adult, paediatric and neonatal use.
3.5.2 | The unit must continuously measure and display SpO2 %
3.5.3 | Oxygen saturation (SpO2 %): Range of measurement: 0 - 100%
3.5.4 | The accuracy must be better than 2,5 % (70 - 100%) and 5% (50 - 70%).
3.5.5 | Pulse rate (beats/min): Range of measurement: 20 - 300 bpm
3.5.6 | The pulse rate accuracy must be better than ± 2 %.
3.5.7 | An audible bleep should sound for each heartbeat, the pitch of which should vary with the oxygen saturation.
3.5.8 | Automatic rejection of motion artefacts
3.6 | SpO2 SENSOR:
3.6.1 | Sensor: Dual LED, Infra-red/visible red, which does not heat the point of application to the patient.
3.6.2 | The sensor lead must be approximately 3 metres long.
3.7 | LCD Display:
3.7.1 | The display must offer clear viewing under all lighting conditions.
3.7.2 | The following must be displayed:
3.7.2a | Systolic, diastolic and mean arterial pressures (numeric)
3.7.2b | Oxygen saturation (SpO2 %)
3.7.2c | Pulse rate (beats/min)
3.7.2d | Alarm limits
3.7.2e | SpO2 waveform or bar-graph
3.7.2f | Trend information
3.7.2g | Battery status
3.7.2h | Indication of mains or battery operation
3.7.3 | The unit must indicate whether the signal to the oximeter sensor is adequate for reliable measurements. If the signal strength or quality is inadequate, this must be clearly displayed.
3.8 | ALARMS:
3.8.1 | All alarms must have default settings
3.8.2 | The alarm parameters must be user-adjustable:
3.8.3 | Adjustable high and low alarm limits for systolic, diastolic, mean arterial pressure and heart rate.
3.8.4 | SpO2 LOW: 50 % - 99 %
3.8.5 | Pulse rate LOW: 40 – 150 b.p.m.
3.8.6 | Pulse rate HIGH: 70 – 250 b.p.m.
3.8.7 | SpO2 probe not attached to patient or the pulse oximeter
3.8.8 | Violation of limits must result in both audible and visual alarms and indication of the parameters affected.
3.8.9 | Adjustable audio alarm volume
3.8.10 | Audio alarm mute: 2 minute.
3.8.11 | Low battery: Alarm at least 10 minutes before battery power is exhausted
3.9 | Other system requirements
3.9.1 | The unit must operate from 220 V, 50 Hz mains supply
3.9.2 | Mains cable at least four metres long with SABS approved 15A three prong plug.
3.9.3 | The unit must include an internal rechargeable battery
3.9.4 | The unit must be able to function from the mains supply regardless of the battery condition.
3.9.5 | Battery must have the capacity for the unit to operate fully for at least five hours
3.9.6 | The battery must be charged while the unit is operating from the mains supply.
3.9.7 | The battery must automatically take over in the event of a mains power failure.
3.9.8 | The power supply must be an integral part of the unit.
3.9.9 | The length of the mains cable must be a minimum of three metres.
3.9.10 | The case of the unit must be robust and impact resistant.
3.9.11 | The design must be such that spilled liquids do not gain access to the internal working of the unit.
3.9.12 | The monitor must be mounted on a trolley with basket for accessories
3.9.13 | The unit must be handed over in full operating order
3.9.14 | A starter-pack of consumables must be supplied with the unit.
4 | Standard Accessories and Consumables
4.1 | Reusable latex-free paediatric cuff
4.2 | Reusable latex-free adult cuff
4.3 | Reusable latex-free neonate cuff
4.4 | Reusable latex-free obese cuff
4.5 | NIBP hose
4.6 | Finger sensor for paeds
4.7 | Finger sensor for adults
4.8 | Finger wrap-around probe for neonates
4.9 | Spo2 interface cable
4.10 | Two reusable temperature probes
5 | Conformity Compliance (Please attach certificate)
5.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
5.2 | OEM must comply and certified on ISO 9001 quality standards, attach proof of compliance
5.2.1 | Please provide unique ref number of the ISO 9001 ceriticate:
5.3 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
5.3.1 | Please provide unique ref number of the ISO 13485 ceriticate:
5.4 | Model quoted for must be EC certified. Attach a copy of certification
5.4.1 | Please provide unique ref number of the EC ceriticate:
5.5 | Hazardous Substance Act: If this type of equipment / apparatus appears on the schedule of Hazardous substances issued by the Directorate: Health Technology of the Department of Health, a license in terms of the Act on Hazardous Substances (Act 15/1973) must be submitted with the bid / quotation documents. The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.6 | All electrical/electronic medical equipment must be licensed by Radiation Control, where applicable a copy of the license must be submitted The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.6.1 | Radiation Control License number:
5.7 | The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade (rubber to with 2 screws) type and it must be a minimum length of three (3) meters. NB the mains cable of the unit tendered for must be SABS colour coded.
5.8 | The equipment quoted for must be protected against electromagnetic interference IEC 60601-1-2
5.9 | The unit must comply with electrical safety and performance standard such as IEC 60601-2-30, attached certification
5.9.1 | Please provide unique ref number of the IEC ceriticate:
6 | COMPREHENSIVE MAINTENANCE
6.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) must be included in the bid price
6.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
6.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
6.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
6.3 | The bidder to provide the OEM recommended service interval of the unit:
6.4 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
7 | General
7.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
7.2 | The unit on offer must be the latest technology.
Vital Signs Monitor, on trolley
RT2-05-19a | Reusable latex-free paediatric cuff
RT2-05-19b | Reusable latex-free adult cuff
RT2-05-19c | Reusable latex-free neonate cuff
RT2-05-19d | Reusable latex-free obese cuff
RT2-05-19e | NIBP hose
RT2-05-19f | Finger sensor for paeds
RT2-05-19g | Finger sensor for adults
RT2-05-19h | Finger wrap-around probe for neonates
RT2-05-19i | Spo2 interface cable
RT2-05-19j | Two reusable temperature probes
RT2-05-19k | Front bezel
RT2-05-19l | Back cover
RT2-05-19m | Battery
RT2-05-19n | NIBP pump/module
Bidder to list any other Consumables or Accessories (Not listed above) for the model on offer below

## Pulse Oximeter, Tabletop
RT2-05-20 | PULSE OXIMETER, Tabletop
1 | Description
1.1 | A device that measures the oxygen saturation of arterial blood by utilizing a sensor attached typically to a finger, toe, or ear to determine the percentage of oxyhemoglobin in blood pulsating through a network of capillaries
2 | Product Offered:
2.1 | Make:
2.2 | Model:
2.3 | Initial Manufacturing year of Model:
2.4 | Radiation Control License number:
3 | Technical Specification
3.1 | Suitable for adult, paediatric and neonatal use.
3.2 | The design must be flat with rubber footing to provide stability from any disturbance or movements on flat surface or desktop, therefore any other design that could be confused as a Vital Signs will not be acceptable
3.3 | The unit must continuously measure and display SpO2 %
3.4 | Oxygen saturation (SpO2 %): Range of measurement: 0 - 100%
3.5 | The accuracy must be better than 2,5 % (70 - 100%) and 5% (50 - 70%).
3.6 | Pulse rate (beats/min): Range of measurement: 20 - 300 bpm
3.7 | The pulse rate accuracy must be better than ± 2 %.
3.8 | An audible bleep should sound for each heartbeat, the pitch of which should vary with the oxygen saturation.
3.9 | Automatic rejection of motion artefacts
3.10 | The unit must have trending for at least 72 hours
3.11 | Self test procedure at switch-on
3.12 | The display must offer clear viewing under all lighting conditions.
3.13 | The following must be displayed:
3.13.1 | Oxygen saturation (SpO2 %)
3.13.2 | Pulse rate (beats/min)
3.13.3 | Alarm limits
3.13.4 | Pulse waveform
3.13.5 | Trend information
3.13.6 | Battery status
3.13.7 | Indication of mains or battery operation
3.14 | The unit must indicate whether the signal to the oximeter sensor is adequate for reliable measurements. If the signal strength or quality is inadequate, this must be clearly displayed.
3.15 | SENSOR:
3.15.1 | Sensor: Dual LED, Infra-red/visible red, which does not heat the point of application to the patient.
3.15.2 | The sensor lead must be approximately 3 metres long.
3.16 | ALARMS:
3.16.1 | All alarms must have default settings
3.16.2 | The alarm parameters must be user-adjustable:
3.16.2a | SpO2 LOW: 50 % - 99 %
3.16.2b | Pulse rate LOW: 40 – 150 b.p.m.
3.16.2c | Pulse rate HIGH: 70 – 250 b.p.m.
3.16.2d | Probe not attached to patient or the pulse oximeter
3.16.3 | Violation of these limits must result in both audible and visual alarms and indication of the parameters affected.
3.16.4 | Adjustable audio alarm volume
3.16.5 | Audio alarm mute: 1 minute. The unit must resume normal alarm functions after the minute has expired.
3.16.6 | Low battery: Alarm at least 10 minutes before battery power is exhausted
3.17 | General
3.17.1 | The unit must operate from 220 V, 50 Hz mains supply
3.17.2 | Mains cable at least three metres long with SABS approved 15A three prong plug.
3.17.3 | The unit must include an internal rechargeable battery
3.17.4 | Battery charger to be incorporated in the unit
3.17.5 | The unit must be able to function from the mains supply regardless of the battery condition.
3.17.6 | Battery must have the capacity for the unit to operate fully for at least five hours
3.17.7 | The battery must be charged while the unit is operating from the mains supply.
3.17.8 | The battery must automatically take over in the event of a mains power failure.
3.17.9 | Automatic switch-off after a period of non-use
3.17.10 | The case of the unit must be robust and impact resistant.
3.17.11 | The design must be such that spilled liquids do not gain access to the internal working of the unit.
3.17.12 | The unit must be handed over in full operating order
3.17.13 | A starter-pack of consumables must be supplied with the unit.
4 | Standard Accessories and Consumables
4.1 | Finger sensor for paeds
4.2 | Finger sensor for adults
4.3 | Finger wrap-around probe for neonates
4.4 | Spo2 interface cable
4.5 | Battery
5 | Conformity Compliance (Please attach certificate)
5.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
5.2 | OEM must comply and certified on ISO 9001 quality standards, attach proof of compliance
5.2.1 | Please provide unique ref number of the ISO 9001 ceriticate:
5.3 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
5.3.1 | Please provide unique ref number of the ISO 13485 ceriticate:
5.4 | Model quoted for must be EC certified. Attach a copy of certification
5.4.1 | Please provide unique ref number of the EC ceriticate:
5.5 | Hazardous Substance Act: If this type of equipment / apparatus appears on the schedule of Hazardous substances issued by the Directorate: Health Technology of the Department of Health, a license in terms of the Act on Hazardous Substances (Act 15/1973) must be submitted with the bid / quotation documents. The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.6 | All electrical/electronic medical equipment must be licensed by Radiation Control, where applicable a copy of the license must be submitted The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.6.1 | Radiation Control License number:
5.7 | The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade (rubber to with 2 screws) type and it must be a minimum length of three (3) meters. NB the mains cable of the unit tendered for must be SABS colour coded.
5.8 | The equipment quoted for must be protected against electromagnetic interference IEC 60601-1-2
6 | COMPREHENSIVE MAINTENANCE
6.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) must be included in the bid price
6.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
6.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
6.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
6.3 | The bidder to provide the OEM recommended service interval of the unit:
6.4 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
7 | General
7.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
7.2 | The unit on offer must be the latest technology.
PULSE OXIMETER, Tabletop Accessories and Consumables
RT2-05-20a | Finger sensor for paeds
RT2-05-20b | Finger sensor for adults
RT2-05-20c | Finger wrap-around probe for neonates
RT2-05-20d | Spo2 interface cable
RT2-05-20e | Battery
RT2-05-20f | Front bezel
RT2-05-20g | Roll stand
Bidder to list any other Consumables or Accessories (Not listed above) for the model on offer below

## Suction Unit, Wall-mount, Doubl
RT2-05-21 | Suction unit, wall-mount
1 | Description
1.1 | An appliance used to remove substance such as blood, saliva, mucus and vomit from person's airway
2 | Product Offered:
2.1 | Make:
2.2 | Model:
2.3 | Initial Manufacturing year of Model:
3 | Technical Specification
3.1 | Must be a complete suction unit with an accessory rail mounting clip.
3.2 | The unit must have a Heyer connection
3.3 | The unit must be of robust design.
3.4 | Two receiver jars, each of 2 litre capacity, graduated in 100 ml steps
3.5 | Receiver jars must be high impact crack resistant polycarbonate or similar material (No glass)
3.6 | The receiving jar with lid must be fully autoclavable at 134o C.
3.7 | Receiver jar lid must have overflow protection
3.8 | The unit offered must be able to be used with liners.
3.9 | The suction inlets and connectors must be plastic.
3.10 | Suction power must be controllable by means of a vacuum control
3.11 | The unit must have a durable on/off switch
3.12 | Operating pressure: 0 to -80 kPa
3.13 | The gauge must be marked marked in kPa and mmHg
3.14 | The gauge must be fully removable for quick maintenance.
3.15 | The pressure hose must be colour coded yellow.
3.16 | Length of the pressure hose: 2 m
3.17 | The suction unit must be easy to service and maintain.
3.18 | The suction capacity must be at least 25 litres per minute.
3.19 | Water trap/safety trap
3.20 | The unit must be handed over in full operating order
3.21 | A starter-pack of consumables must be supplied with the unit.
4 | Standard Accessories and Consumables
4.1 | Silicone tubes (soft) set 8 x 14 mm
4.2 | Tube connectors
4.3 | Probe connector
4.4 | Aspiration probe.
4.5 | Particle filter
4.6 | Suction tube with canula, reusable
4.7 | Vacuum probe
4.8 | Hose vacuum probe
4.9 | Suction Regulator
4.10 | Suction Gauge
4.11 | 2 liter canister x2
5 | Conformity Compliance (Please attach certificate)
5.1 | OEM must comply and certified on ISO 9001 quality standards, attach proof of compliance
5.1.1 | Please provide unique ref number of the ISO 9001 ceriticate:
5.2 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
5.2.1 | Please provide unique ref number of the ISO 13485 ceriticate:
5.3 | Model quoted for must be EC or SABS certified. Attach a copy of certification
5.3.1 | Please provide unique ref number of the EC or SABS ceriticate:
6 | COMPREHENSIVE MAINTENANCE
6.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) must be included in the bid price
6.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS, Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
7 | General
7.1 | The unit on offer must be the latest technology.
Suction unit, wall-mount Accessories and Consumables
RT2-05-21a | Silicone tubes (soft) set 8 x 14 mm
RT2-05-21b | Tube connectors
RT2-05-21c | Probe connector
RT2-05-21d | Aspiration probe.
RT2-05-21e | Particle filter
RT2-05-21f | Suction tube with canula, reusable
RT2-05-21g | Vacuum probe
RT2-05-21h | Hose vacuum probe
RT2-05-21i | Suction Regulator
RT2-05-21j | Suction Gauge
RT2-05-21k | 3 liter canister
RT2-05-21l | Mobile stand (with double canister slots)
Bidder to list any other Consumables or Accessories (Not listed above) for the model on offer below

## Suction Unit, Mobile
RT2-05-22 | Suction Unit, Mobile with Battery
1 | Description
1.1 | An appliance used to remove substance such as blood, saliva, mucus and vomit from person's airway
2 | Product Offered:
2.1 | Make:
2.2 | Model:
2.3 | Initial Manufacturing year of Model:
3 | Technical Specification
3.1 | Unit must be mounted on mobile stand with at least 4 x castor wheels.
3.2 | Must contain a rechargable battery back-up with built-in charger
3.3 | Two of the castors must be lockable
3.4 | Operating pressure: 0 to -60 kPa
3.5 | The gauge must be marked marked in kPa and mmHg
3.6 | The gauge must be fully removable for maintenance.
3.7 | The suction capacity must be at least 25 litres per minute.
3.8 | On-off switch must be illuminated and have a splash-proof cover
3.9 | Must have suction control for accurate regulation of the vacuum
3.10 | Two receiver jars, each of 2 litre capacity, graduated in 100 ml steps
3.11 | Receiver jars must be high impact crack resistant polycarbonate or similar material (No glass)
3.12 | Receiver jars and lids must be removable and autoclavable at 134o C
3.13 | Receiver jar lid must have overflow protection
3.14 | Receiver jar must operate with and without liners.
3.15 | Vacuum pump regulator must be fitted with a hydrophobic bacterial filter to prevent aspirated fluid from entering the suction pump.
3.16 | Working noise less than 50 dB
3.17 | The suction outlets must be stainless steel.
3.18 | The pump must be oil free and maintenance free
3.19 | The mobile stand must have a handle for pushing
3.20 | The unit must operate from a 220 V AC supply.
3.21 | Must have a mains cable at least 4 m long fitted with a SABS approved three pin 15A plug
3.22 | Battery status indicator
3.23 | Battery capacity: at least 1 hour operation
3.24 | Battery must be sealed, maintenance free
3.25 | The suction unit must be easy to service and maintain.
3.26 | The unit must be handed over in full operating order
3.27 | A starter-pack of consumables must be supplied with the unit.
4 | Standard Accessories and Consumables
4.1 | Silicone tubes (soft) set 8 x 14 mm
4.2 | Tube connectors
4.3 | Probe connector
4.4 | Aspiration probe.
4.5 | Particle filter
4.6 | Suction tube with canula, reusable
4.7 | Vacuum probe
4.8 | Hose vacuum probe
4.9 | Suction Regulator
4.10 | Suction Gauge
5 | Conformity Compliance (Please attach certificate)
5.1 | OEM must comply and certified on ISO 9001 quality standards, attach proof of compliance
5.1.1 | Please provide unique ref number of the ISO 9001 ceriticate:
5.2 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
5.2.1 | Please provide unique ref number of the ISO 13485 ceriticate:
5.3 | Model quoted for must be EC or SABS certified. Attach a copy of certification
5.3.1 | Please provide unique ref number of the EC or SABS ceriticate:
6 | COMPREHENSIVE MAINTENANCE
6.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) must be included in the bid price
6.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS, Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
7 | General
7.1 | The unit on offer must be the latest technology.
Mobile Suction unit, Accessories and Consumables
RT2-05-22a | Silicone tubes (soft) set 8 x 14 mm
RT2-05-22b | Tube connectors
RT2-05-22c | Probe connector
RT2-05-22d | Aspiration probe.
RT2-05-22e | Particle filter
RT2-05-22f | Suction tube with canula, reusable
RT2-05-22i | Suction Regulator
RT2-05-22j | Suction Gauge
RT2-05-22k | Battery
Bidder to list any other Consumables or Accessories (Not listed above) for the model on offer below

## Theater Light, Mobile
RT2-05-23 | THEATER LIGHT MOBILE
1 | DESCRIPTION
1.1 | Mobile single head theatre light. The unit offered shall be designed with suitabilty to provide appropriate light during surgical operations.
2 | Product Offered:
2.1 | Make:
2.2 | Model:
2.3 | Initial Manufacturing year of Model:
3 | TECHNICAL SPECIFICATION
3.1 | Unit must be of LED technology.
3.2 | The unit must be rubust, constructed from a corrosion resistant material i.e. Alluminium or equivalent.
3.3 | The unit must comply with the laminar flow thus controlling contamination.
3.4 | Once positioned the light must not drift away.
3.5 | The unit must provide low heat radiation, efficient and stable illuminination.
3.6 | ILLUMINATION TO BE AS FOLLOWS:
3.6.1 | Concentrated spot: At 1 meter distance: The unit must have central illumination of 130 000 LUX - 160 000 LUX.
3.6.2 | The LED should perfom for 40 000 hours and above.
3.6.3 | Diameter of lighthead to be between 60 – 100 cm.
3.6.4 | Depth of the illumination to be ± 80 cm3.
3.6.5 | Colour rendering index not be less than 95 RA
3.7 | ELECTRICAL DATA:
3.7.1 | Main power system supply 220V to 240 V.
3.7.2 | The unit must have a built-in electrical protection system .
3.7.3 | The unit must have a built-in battery bank which will make it to work under full load for a minumum period of 4 hrs.
3.8 | MECHANICAL DATA:
3.8.1 | The unit must have a central handle to adjust the positioning of the headlight.
3.8.2 | Breaking system on all movement parts must be available.
3.8.3 | The unit must be balanced in all positions of the headlight.
3.8.4 | Handle must be sterilisable.
3.9 | CONTROLS AND ACCESSORIES:
3.9.1 | The intergrated lighthandle control must allow the user to control the light focus and its lighthead to any desirable user position.
4 | STANDARD ACCESSORIES
4.1 | light handle
5 | Conformity Compliance (Please attach certificate)
5.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
5.2 | OEM must comply and certified on ISO 9001 quality standards, attach proof of compliance
5.2.1 | Please provide unique ref number of the ISO 9001 ceriticate:
5.3 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
5.3.1 | Please provide unique ref number of the ISO 13485 ceriticate:
5.4 | Model quoted for must be EC certified. Attach a copy of certification
5.4.1 | Please provide unique ref number of the EC ceriticate:
5.7 | The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade (rubber to with 2 screws) type and it must be a minimum length of three (3) meters. NB the mains cable of the unit tendered for must be SABS colour coded.
5.8 | The equipment quoted for must be protected against electromagnetic interference IEC 60601-1-2
6 | COMPREHENSIVE MAINTENANCE
6.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) must be included in the bid price
6.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
6.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
6.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
6.3 | The bidder to provide the OEM recommended service interval of the unit:
6.4 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
7 | General
7.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
7.2 | The unit on offer must be the latest technology.
THEATER LIGHT MOBILE Accessories
RT2-05-23a | light handle
RT2-05-23b | Battery
RT2-05-23c | LED replacement
Bidder to list any other Consumables or Accessories (Not listed above) for the model on offer below

## Fluid Warmer, Cabinet, 40-60L
RT2-05-25 | FLUID WARMING CABINET, SMALL
1 | DESCRIPTION
1.1 | Stationary cabinet used to store and warm fluids and blood (from 40 - 60 Litres)
2 | Product Offered:
2.1 | Make:
2.2 | Model:
2.3 | Initial Manufacturing year of Model:
3 | TECHNICAL SPECIFICATION
3.1 | The offered unit shall have variable compartment sizes with detachable/interchangeable atleast 1 shelve at standard waist height
3.2 | The exterior of the unit must be stainless steel 304
3.3 | The chamber or inside of the unit offered must be stainless steel 304 including detachable shelves.
3.4 | The door of the offered unit should be built in with a toughened and tempered glass door/window to allow visualisation of chamber contents without opening
3.5 | The offered unit should have lighting for easy viewing as well as a light switch on the interface panel
3.6 | The unit should offer temperature ranges of not less than 30˚C and not above 42˚C. The working range should be protected by a locking mechanism which needs specific authority to prevent accidental overheating.
3.7 | The offered unit shall have an LED/LCD user interface panel to allow user to set and monitor temperatures
3.8 | Displayed temperatures should have an accuracy of ±2 ̊C
3.9 | Unit shall display distinctively SET Temperatures as well as ACTUAL chamber temperature at any given time
3.10 | Displayed temperature must be in ˚C
4 | Alarms and Safety features
4.1 | Over Temperature: High - ± 2°C
4.2 | Low Temperature: Low - ± 2°C
4.3 | Power loss
4.4 | Malfunction
4.5 | Unit shall be built with an over temperature shut off system
5 | Conformity Compliance (Please attach certificate)
5.1 | OEM must comply and certified on ISO 9001 quality standards, attach proof of compliance
5.1.1 | Please provide unique ref number of the ISO 9001 ceriticate:
5.2 | Model quoted for must be EC or SABS certified. Attach a copy of certification
5.2.1 | Please provide unique ref number of the EC or SABS ceriticate:
6 | COMPREHENSIVE MAINTENANCE
6.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) must be included in the bid price
6.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
6.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
6.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
6.3 | The bidder to provide the OEM recommended service interval of the unit:
6.4 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
7 | General
7.1 | The unit on offer must be the latest technology.

## Fluid Warmer, Cabinet, 80-150L
RT2-05-26 | FLUID WARMING CABINET, MEDIUM
1 | DESCRIPTION
1.1 | Stationary cabinet used to store and warm fluids and blood (from 80 - 150 Litres)
2 | Product Offered:
2.1 | Make:
2.2 | Model:
2.3 | Initial Manufacturing year of Model:
3 | TECHNICAL SPECIFICATION
3.1 | The offered unit shall have variable compartment sizes with detachable/interchangeable atleast 2 shelves at standard waist height
3.2 | The exterior of the unit must be stainless steel 304
3.3 | The chamber or inside of the unit offered must be stainless steel 304 including detachable shelves.
3.4 | The door of the offered unit should be built in with a toughened and tempered glass door/window to allow visualisation of chamber contents without opening
3.5 | The offered unit should have lighting for easy viewing as well as a light switch on the interface panel
3.6 | The unit should offer temperature ranges of not less than 30˚C and not above 42˚C. The working range should be protected by a locking mechanism which needs specific authority to prevent accidental overheating.
3.7 | The offered unit shall have an LED/LCD user interface panel to allow user to set and monitor temperatures
3.8 | Displayed temperatures should have an accuracy of ±2 ̊C
3.9 | Unit shall display distinctively SET Temperatures as well as ACTUAL chamber temperature at any given time
3.10 | Displayed temperature must be in ˚C
4 | Alarms and Safety features
4.1 | Over Temperature: High - ± 2°C
4.2 | Low Temperature: Low - ± 2°C
4.3 | Power loss
4.4 | Malfunction
4.5 | Unit shall be built with an over temperature shut off system
5 | Conformity Compliance (Please attach certificate)
5.1 | OEM must comply and certified on ISO 9001 quality standards, attach proof of compliance
5.1.1 | Please provide unique ref number of the ISO 9001 ceriticate:
5.2 | Model quoted for must be EC or SABS certified. Attach a copy of certification
5.2.1 | Please provide unique ref number of the EC or SABS ceriticate:
6 | COMPREHENSIVE MAINTENANCE
6.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) must be included in the bid price
6.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
6.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
6.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
6.3 | The bidder to provide the OEM recommended service interval of the unit:
6.4 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
7 | General
7.1 | The unit on offer must be the latest technology.

## Fluid Warmer, Cabinet, 200-300L
RT2-05-27 | FLUID WARMING CABINET, LARGE
1 | DESCRIPTION
1.1 | Stationary cabinet used to store and warm fluids and blood (from 200 - 300 Litres)
2 | Product Offered:
2.1 | Make:
2.2 | Model:
2.3 | Initial Manufacturing year of Model:
3 | TECHNICAL SPECIFICATION
3.1 | The offered unit shall have variable compartment sizes with detachable/interchangeable atleast 3 shelves at standard waist height
3.2 | The exterior of the unit must be stainless steel 304
3.3 | The chamber or inside of the unit offered must be stainless steel 304 including detachable shelves.
3.4 | The door of the offered unit should be built in with a toughened and tempered glass door/window to allow visualisation of chamber contents without opening
3.5 | The offered unit should have lighting for easy viewing as well as a light switch on the interface panel
3.6 | The unit should offer temperature ranges of not less than 30˚C and not above 42˚C. The working range should be protected by a locking mechanism which needs specific authority to prevent accidental overheating.
3.7 | The offered unit shall have an LED/LCD user interface panel to allow user to set and monitor temperatures
3.8 | Displayed temperatures should have an accuracy of ±2 ̊C
3.9 | Unit shall display distinctively SET Temperatures as well as ACTUAL chamber temperature at any given time
3.10 | Displayed temperature must be in ˚C
4 | Alarms and Safety features
4.1 | Over Temperature: High - ± 2°C
4.2 | Low Temperature: Low - ± 2°C
4.3 | Power loss
4.4 | Malfunction
4.5 | Unit shall be built with an over temperature shut off system
5 | Conformity Compliance (Please attach certificate)
5.1 | OEM must comply and certified on ISO 9001 quality standards, attach proof of compliance
5.1.1 | Please provide unique ref number of the ISO 9001 ceriticate:
5.2 | Model quoted for must be EC or SABS certified. Attach a copy of certification
5.2.1 | Please provide unique ref number of the EC or SABS ceriticate:
6 | COMPREHENSIVE MAINTENANCE
6.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) must be included in the bid price
6.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
6.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
6.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
6.3 | The bidder to provide the OEM recommended service interval of the unit:
6.4 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
7 | General
7.1 | The unit on offer must be the latest technology.

## Fridge, Vaccine, 100-150L
RT2-05-28 | FRIDGE, VACCINE, 200 - 250L
1 | DESCRIPTION
1.1 | Vaccine fridge: For vaccine storage and ice pack freezing area with a capacity from 200 to at least 250 litres
2 | Product Offered:
2.1 | Make:
2.2 | Model:
2.3 | Initial Manufacturing year of Model:
3 | TECHNICAL SPECIFICATION
3.1 | Fridge temperature: +2°C to +8°C
3.2 | Freezer temperature: -14°C to -10°C
3.3 | Fridge capacity: 100 to 150 Litres
3.4 | Freezer capacity: 30 -50 Litres
3.5 | Interior lighting should be On when the door opens
3.6 | Minimum of 3 adjustable stainless steel trays with perforated design
3.7 | The door with internal illumination and control panel
3.8 | Must have an integral digital control panel on the outside of the fridge to monitor and display both the fridge and freezer internal temperatures
3.9 | Must have a fan to deliver optimum temperature stability and a faster temperature recovery after door openings
3.10 | Must include an automatic defrost to ensures more efficient cooling
3.11 | Complete with door locking system
3.12 | The unit should include individual freezer door compartment
3.13 | Must incorporate a temperature alarm which sounds if the temperature deviates from the set parameters for more than 15 minutes
3.14 | Stainless steel inner liner
3.15 | Adjustable feet
3.16 | Must have atleast 48 hour holdover facility during power failures.
3.17 | Chlorofluorocarbon (CFC) and Hydro Chlorofluorocarbon (HCFC) Free
3.18 | Mains cable must be in excess of 3 meters, complete with a standard South African 220/230 V Ac plug top.
4 | The alarms must be audible and visual to signal the following:
4.1 | High and low temperature alarm
4.2 | Power failure
4.3 | Door open
5 | Conformity Compliance (Please attach certificate)
5.1 | OEM must comply and certified on ISO 9001 quality standards, attach proof of compliance
5.1.1 | Please provide unique ref number of the ISO 9001 ceriticate:
5.2 | Model quoted for must be EC or SABS certified. Attach a copy of certification
5.2.1 | Please provide unique ref number of the EC or SABS ceriticate:
6 | COMPREHENSIVE MAINTENANCE
6.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) must be included in the bid price
6.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
6.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
6.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
6.3 | The bidder to provide the OEM recommended service interval of the unit:
6.4 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
7 | General
7.1 | The unit on offer must be the latest technology.

## Fridge, Vaccine, 200-250L
RT2-05-29 | FRIDGE, VACCINE, 200 - 250L
1 | DESCRIPTION
1.1 | Vaccine fridge: For vaccine storage and ice pack freezing area with a capacity from 200 to at least 250 litres
2 | Product Offered:
2.1 | Make:
2.2 | Model:
2.3 | Initial Manufacturing year of Model:
3 | TECHNICAL SPECIFICATION
3.1 | Fridge temperature: +2°C to +8°C
3.2 | Freezer temperature: -14°C to -10°C
3.3 | Fridge capacity: 200 to 250 Litres
3.4 | Freezer capacity: 30 -50 Litres
3.5 | Interior lighting should be On when the door opens
3.6 | Minimum of 3 adjustable stainless steel trays with perforated design
3.7 | The door with internal illumination and control panel
3.8 | Must have an integral digital control panel on the outside of the fridge to monitor and display both the fridge and freezer internal temperatures
3.9 | Must have a fan to deliver optimum temperature stability and a faster temperature recovery after door openings. Bidder to state how consistant temperature is achieved throughtout the whole fridge
3.10 | Must include an automatic defrost to ensures more efficient cooling
3.11 | Complete with door locking system
3.12 | The unit should include individual freezer door compartment
3.13 | Must incorporate a temperature alarm which sounds if the temperature deviates from the set parameters for more than 15 minutes
3.14 | Stainless steel inner liner
3.15 | Adjustable feet
3.16 | Must have atleast 48 hour holdover facility during power failures at an ambient temperature of 32o C.
3.17 | Chlorofluorocarbon (CFC) and Hydro Chlorofluorocarbon (HCFC) Free
3.18 | Mains cable must be in excess of 3 meters, complete with a standard South African 220/230 V Ac plug top.
4 | The alarms must be audible and visual to signal the following:
4.1 | High and low temperature alarm
4.2 | Power failure
4.3 | Door open
5 | Conformity Compliance (Please attach certificate)
5.1 | OEM must comply and certified on ISO 9001 quality standards, attach proof of compliance
5.1.1 | Please provide unique ref number of the ISO 9001 ceriticate:
5.2 | Model quoted for must be EC or SABS certified. Attach a copy of certification
5.2.1 | Please provide unique ref number of the EC or SABS ceriticate:
6 | COMPREHENSIVE MAINTENANCE
6.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) must be included in the bid price
6.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
6.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
6.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
6.3 | The bidder to provide the OEM recommended service interval of the unit:
6.4 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
7 | General
7.1 | The unit on offer must be the latest technology.

## Fridge, Vaccine, 300-350L
RT2-05-30 | FRIDGE, VACCINE, 300 - 350 litres
1 | DESCRIPTION
1.1 | Vaccine fridge: For vaccine storage and ice pack freezing area with a capacity of 290 to 350 litres
2 | Product Offered:
2.1 | Make:
2.2 | Model:
2.3 | Initial Manufacturing year of Model:
3 | TECHNICAL SPECIFICATION
3.1 | Fridge temperature: +2°C to +8°C
3.2 | Freezer temperature: -14°C to -10°C
3.3 | Fridge capacity: 300 to 350 Litres
3.4 | Freezer capacity: 30 -50 Litres
3.5 | Interior lighting should be On when the door opens
3.6 | Minimum of 3 adjustable stainless steel trays with perforated design
3.7 | The door with internal illumination and control panel
3.8 | Must have an integral digital control panel on the outside of the fridge to monitor and display both the fridge and freezer internal temperatures
3.9 | Must have a fan to deliver optimum temperature stability and a faster temperature recovery after door openings. Bidder to state how consistant temperature is achieved throughtout the whole fridge
3.10 | Must include an automatic defrost to ensures more efficient cooling
3.11 | Complete with door locking system
3.12 | The unit should include individual freezer door compartment
3.13 | Must incorporate a temperature alarm which sounds if the temperature deviates from the set parameters for more than 15 minutes
3.14 | Stainless steel inner liner
3.15 | Adjustable feet
3.16 | Must have atleast 48 hour holdover facility during power failures at an ambient temperature of 32o C.
3.17 | Chlorofluorocarbon (CFC) and Hydro Chlorofluorocarbon (HCFC) Free
3.18 | Mains cable must be in excess of 3 meters, complete with a standard South African 220/230 V Ac plug top.
4 | The alarms must be audible and visual to signal the following:
4.1 | High and low temperature alarm
4.2 | Power failure
4.3 | Door open
5 | Conformity Compliance (Please attach certificate)
5.1 | OEM must comply and certified on ISO 9001 quality standards, attach proof of compliance
5.1.1 | Please provide unique ref number of the ISO 9001 ceriticate:
5.2 | Model quoted for must be EC or SABS certified. Attach a copy of certification
5.2.1 | Please provide unique ref number of the EC or SABS ceriticate:
6 | COMPREHENSIVE MAINTENANCE
6.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) must be included in the bid price
6.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
6.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
6.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
6.3 | The bidder to provide the OEM recommended service interval of the unit:
6.4 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
7 | General
7.1 | The unit on offer must be the latest technology.

## Fridge, Medicine, 200-270L
RT2-05-31 | FRIDGE, MEDICINE, 200-270L
1 | DESCRIPTION
1.1 | Single glass door fridge: To store Pharmaceutical stock correctly, especially in relation to those medicines which require storage at low temperatures.
2 | Product Offered:
2.1 | Make:
2.2 | Model:
2.3 | Initial Manufacturing year of Model:
3 | TECHNICAL SPECIFICATION
3.1 | Single door fridge: To store Pharmaceutical stock correctly, especially in relation to those medicines which require storage at low temperatures.
3.2 | Approximate size: DEPTH 540mm to 700mm ; WIDTH 580mm to 700mm; HEIGHT 2020mm to 2500mm
3.3 | Must display Temperature: +2°C to +8°C
3.4 | Capacity: 200 to 270 Litres
3.5 | Interior lighting should always be On
3.6 | Minimum of 3 adjustable stainless steel trays with perforated design
3.7 | Glass door with internal illumination and control panel
3.8 | Must have an integral digital control panel on the outside of the fridge to monitor the internal temperature
3.9 | Must have a fan to deliver optimum temperature stability and a faster temperature recovery after door openings. Bidder to state how consistant temperature is achieved throughtout the whole fridge
3.10 | Must include an automatic defrost to ensures more efficient cooling
3.11 | Complete with door locking system
3.12 | Must incorporate a temperature alarm which sounds if the temperature deviates from the set parameters for more than 15 minutes
3.13 | Stainless steel inner liner
3.14 | Adjustable feet
3.15 | Must have atleast 4-8 hour holdover facility during power failures.
3.16 | Chlorofluorocarbon (CFC) and Hydro Chlorofluorocarbon (HCFC) Free
3.17 | Mains cable must be in excess of 3 meters, complete with a standard South African 220/230 V Ac plug top.
4 | The alarms must be audible and visual to signal the following:
4.1 | High and low temperature alarm
4.2 | Power failure
4.3 | Door open
5 | Conformity Compliance (Please attach certificate)
5.1 | OEM must comply and certified on ISO 9001 quality standards, attach proof of compliance
5.1.1 | Please provide unique ref number of the ISO 9001 ceriticate:
5.2 | Model quoted for must be EC or SABS certified. Attach a copy of certification
5.2.1 | Please provide unique ref number of the EC or SABS ceriticate:
6 | COMPREHENSIVE MAINTENANCE
6.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) must be included in the bid price
6.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
6.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
6.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
6.3 | The bidder to provide the OEM recommended service interval of the unit:
6.4 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
7 | General
7.1 | The unit on offer must be the latest technology.

## Fridge, Medicine, 900-1100L
RT2-05-32 | FRIDGE, MEDICINE, LARGE (900-1100L)
1 | DESCRIPTION
1.1 | Double glass door fridge: To store Pharmaceutical stock correctly, especially in relation to those medicines which require storage at low temperatures.
2 | Product Offered:
2.1 | Make:
2.2 | Model:
2.3 | Initial Manufacturing year of Model:
3 | TECHNICAL SPECIFICATION
3.1 | Double glass door fridge: To store Pharmaceutical stock correctly, especially in relation to those medicines which require storage at low temperatures.
3.2 | Approximate size: DEPTH 540mm to 700mm ; WIDTH 1100mm to 1400mm; HEIGHT 1300mm to 2000mm
3.3 | Must display Temperature: +2°C to +8°C
3.4 | Capacity: 900 to 1100 Litres
3.5 | Interior lighting should always be On
3.6 | Minimum of 6 adjustable stainless steel trays with perforated design
3.7 | Glass doors with internal illumination and control panel
3.8 | Must have an integral digital control panel on the outside of the fridge to monitor the internal temperature
3.9 | Must have a fan to deliver optimum temperature stability and a faster temperature recovery after door openings. Bidder to state how consistant temperature is achieved throughtout the whole fridge
3.10 | Must include an automatic defrost to ensures more efficient cooling
3.11 | Complete with door locking system
3.12 | Must incorporate a temperature alarm which sounds if the temperature deviates from the set parameters for more than 15 minutes
3.13 | Stainless steel inner liner
3.14 | Adjustable feet
3.15 | Must have atleast 4-8 hour holdover facility during power failures.
3.16 | Chlorofluorocarbon (CFC) and Hydro Chlorofluorocarbon (HCFC) free
3.17 | Mains cable must be in excess of 3 meters, complete with a standard South African 220/230 V Ac plug top.
4 | The alarms must be audible and visual to signal the following:
4.1 | High and low temperature alarm
4.2 | Power failure
4.3 | Door open
5 | Conformity Compliance (Please attach certificate)
5.1 | OEM must comply and certified on ISO 9001 quality standards, attach proof of compliance
5.1.1 | Please provide unique ref number of the ISO 9001 ceriticate:
5.2 | Model quoted for must be EC or SABS certified. Attach a copy of certification
5.2.1 | Please provide unique ref number of the EC or SABS ceriticate:
6 | COMPREHENSIVE MAINTENANCE
6.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) must be included in the bid price
6.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
6.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
6.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
6.3 | The bidder to provide the OEM recommended service interval of the unit:
6.4 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
7 | General
7.1 | The unit on offer must be the latest technology.

## Ultrasound, Cardiology
RT2-05-33 | Ultrasound machine, 3D or better- Cardiology
1 | DESCRIPTION
1.1 | A Point-of-care Ultrasound machine for Adult and Paediatric Cardiac , including transoesofageal and transthoracic cardiac, as well as superficial vascular and nerve imaging, including appropriate ultrasound probes
1a | Product Offered:
1a.1 | Make:
1a.2 | Model:
1a.3 | Initial Manufacturing year of Model:
1a.4 | Radiation Control License number:
2 | TECHNICAL SPECIFICATION
2.1 | Basic Machine requirements
2.1.1 | A High-end Cardiac Ultrasound system needs to be offered
2.1.2 | Digital broadband beam-former
2.1.3 | Real-Time Compounding
2.1.4 | Adaptive Image Processing for noise and artefact reduction
2.1.5 | 2D as well as 3D imaging
2.1.6 | Real time 4D (Live 3D) TEE using matrix transducer technology for instantaneous volume rendering of cardiac anatomy .
2.1.7 | Live Biplane Imaging using Matrix transducers (infinite plane adjustments)
2.1.8 | New Single crystal technology or Multiple crystal on transthoracic and Transoesophageal transducers
2.1.9 | The equipment should be able to function in the physical surroundings of theatre in temperature range of at least 10 to 35 degrees Celsius
2.1.10 | Ergonomic transducers designs with lightweight flexible cables
2.1.11 | High-resolution interactive graphical colour touch panel with adjustment for various ambient light conditions
2.1.12 | 3 Active transducer connectors and one non Imaging transducer port
2.1.13 | Intelligent one-button optimization in 2D and Doppler modes
2.2 | IMAGING MODALITIES
2.2.1 | B/Mode 2D
2.2.1.1 | Scan methods:
· Linear scan (state if transducers can perform oblique scanning)
· Sector scan
· Convex scan
2.2.1.2 | Viewing depth:
· Maximum >30 cm, state scanning depth for specific probes
2.2.1.3 | Line density:
· The line density differs depending on the transducer used. Several line densities are available for selection for each transducer.
2.2.1.4 | Scan angle and scan width:
· Adjustment of scan angle (scan width) is possible.
· Adjustment of scan steering (scanning position adjustment) is possible.
· Adjustment of linear scan steering is possible.
· Trapezoid scanning providing convex view on linear probes .
2.2.1.5 | PAN/Zoom:
· Scale can be enlarged or reduced by the encoder.
· The panning position can be arbitrarily selected using the trackball.
· HD zoom: Available (real-time only)
· Spot zoom: Available
· High Definition Write Zoom and Read Zoom with pan feature
2.2.1.6 | Focus:
· Transmission: Maximum 8 steps
· Reception: Continuous
· Auto-focusing: Available
2.2.1.7 | Transmission frequency:
· Image frequency needs to be 0 to 20 MHz
· It differ for different Ultrasound probes- eg Multi Array Cardiac probe 1-5 MHz, Transoesophageal probe 2-7 MHz, Linear hockey stick-like probe 7-15 MHz
2.2.1.8 | Gain:
· Display brightness for B-mode display can be changed.
2.2.1.9 | Dynamic range above 270dB
2.2.1.10 | System Grey scale level needs to be± 256 shades
2.2.1.11 | 2 D loop up to 2 000 frames should be recorded
2.2.1.12 | Indicate maximum length of captured loop in heart beats and minutes
2.2.1.13 | Real time triplex mode must be standard
2.2.1.14 | Frame Rate: supports high frame rate
2.2.1.15 | Gain control one button
2.2.1.16 | STC (TGC ): 8-step slide controls (common to B and M modes) (Time Gain Compensation – lateral gain)
2.2.1.17 | Acoustic output adjustment: Adjustable up to 100%
2.2.1.18 | Colour Palette (2D MAP): Available
2.2.1.19 | THI (Tissue Harmonic Imaging):state if available on all probes
· Pulse Subtraction : Available
· Differential Harmonic Imaging: Available
2.2.1.20 | Compounding Imaging: state if supported
· Frequency and/or Spatial Compounding imaging in real-time is supported.
2.2.1.21 | Display orientation:
· Top/bottom reversal is possible
· Left/right reversal is possible
2.2.1.22 | Panoramic View: state if supported
· Panoramic View is supported
· Measurement is available
2.2.1.23 | One button optimization:
· Enables automatic gain and STC/TGC control for B mode as well as Velocity range and Base line shift control for Spectrum Doppler
2.2.2 | Spectral Doppler
2.2.2.1 | Doppler mode:
· PWD (Pulsed-Wave Doppler) and Colour Doppler available on all probes
· CWD (Continuous-Wave Doppler) available on all cardiac sector probes
2.2.2.2 | Doppler pulse repetition frequency (PRF)
· State PRF frequency
2.2.2.3 | Doppler scan:
· B/C/D simultaneous scan
· State if simultaneous display of Doppler image , side by side, in real-time is possible
2.2.2.4 | Doppler sampling volume:
· The Doppler range gate width can be changed.
· Sample volume size ± 0,5 – 20 mm
2.2.2.5 | Sample shift:
· From 0 cm to the maximum depth
2.2.2.5 | Doppler filter:
· Doppler cut-off frequency can be changed.
2.2.2.6 | Doppler gain:
· Doppler image display brightness can be changed
· Independent gain control for PW and CW
2.2.2.7 | Indication of Doppler spectrum direction:
· Display of the forward and backward velocity spectrum can be reversed
2.2.2.8 | Doppler baseline shift (zero-shift):
· The baseline (zero-velocity line) of Doppler images can be shifted.
· The baseline shift setting can also be adjusted when images that were frozen are displayed.
2.2.2.9 | Doppler audio:
· Stereo output (forward flow and backward flow)
2.2.2.10 | Colour Palette (Doppler MAP): Available to be selected when frozen
2.2.2.11 | Image memory to operate in both dual and quad screen format
2.2.2.12 | Display of Doppler scale:
· 2 Types (velocity and Doppler shift frequency)
2.2.2.13 | Doppler focus:
· Automatically follows the sample position
2.2.2.14 | Doppler angle display:
· Available
2.2.2.15 | Oblique Doppler scan:
· Oblique scan is possible with some linear transducers.
2.2.2.16 | Doppler multi-frequency:
· Doppler transmission and reception frequencies are selectable in PWD.
2.2.2.17 | Doppler sweep speed:
· Sweep speed can be changed
2.2.2.18 | Post processing must be available on frozen images
2.2.2.19 | Doppler Auto trace:
· Waveform Peak, Mean, Peak + Mean
· Trace area specified: Toward, Reverse, All
· Measurement item: Max, Min, Mean, Pl, RI, etc.
· One button optimization technology.
2.2.2.20 | Tissue Doppler- TDI
2.2.2.21 | M-mode should be available, with Anatomical M-Mode
2.2.2.22 | Live (on screen) automatic Doppler Analysis
2.2.3 | COLOUR DOPPLER
2.2.3.1 | Colour Doppler mode:
· Flow velocity
· Flow velocity/variance
2.2.3.2 | Power Imaging mode: state if available
· On all transducers
· Advanced Dynamic Flow mode (high-resolution flow imaging)
2.2.3.3 | Colour Doppler:
· State maximum colour frame rate
· M-mode should be available
2.2.3.4 | Doppler velocity range: to be stated
2.2.3. | Colour Doppler frame interpolation (C Frame Interpolation) Automatically set according to the transducer information and selected exam conditions .
2.2.3.5 | Colour Doppler baseline (C Baseline):
· The zero-velocity line of colour Doppler images can be shifted.
· Zero-shift setting can also be adjusted when an image is frozen or an image in the cine memory is displayed.
2.2.3.6 | Colour reversal is available
2.2.3.7 | Balance between B/W and colour display. Colour weighting to B/W image can be set by comparing the colour Doppler images and B/W images.
2.2.3.8 | Parallel colour Doppler processing. Simultaneous signal reception in a maximum of 4 directions (QSP) is possible.
2.2.3.9 | Colour gain:
· Display brightness of colour Doppler images can be changed.
2.2.3.10 | Colour Doppler multi-frequency
2.2.3.11 | Transmission frequency for colour Doppler image acquisition can be changed.
2.2.3.12 | Colour Doppler line density
2.2.3.13 | Line density of colour Doppler images can be changed.
2.2.3.14 | Adjustment of colour ROI position, size, and angle is possible. Must be fully operational through the total depth of image
2.2.3.15 | Colour Doppler transmission focus
2.2.3.16 | Automatically follows the colour Doppler ROI position.
2.2.3.17 | Colour Doppler filter
2.2.3.18 | Filter cut-off can be changed
2.2.3.19 | ROI filter function
2.2.3.20 | Simultaneous dual-screen display with B mode
2.2.3.21 | Cineloop review possible
2.2.3.22 | On screen measurements must be possible from real time examination
2.2.3.23 | Unit to offer programmable quick annotation
2.2.3.24 | Indicate if Directional Colour Angio is available
2.2.3.25 | Trapezoid scanning providing convex view on linear probes.
2.3 | MEASUREMENT and ANALYSIS
2.3.1 | ±16 callipers (8 channels) for measuring must be offered
2.3.2 | The operating modes should include: 2D, M-Mode, Anatomic M-mode, Colour Flow Doppler, Pulsed Wave Doppler (PW), High PRF PW, Continuous Wave Doppler, Chroma Imaging, Tissue Doppler Imaging, Pulse Inversion, Cineloop Image, M-mode and Doppler Review, 4D / Live 3D, triplex and Harmonic with pulse inversion technology.
2.3.3 | Basic measurements:
2.3.3.1 | · Measurement tools including: distance, depth, area, and circumference
2.3.3.2 | · B-mode measurements
2.3.3.3 | · Distance
2.3.3.4 | · Area: using ellipse or trace
2.3.3.5 | · Angle: intersection of two lines
2.3.3.6 | · Volume flow Applications measurements auto labelling on touch screen.
2.3.3.7 | · Comprehensive reports generated automatically.
2.3.3.8 | · Maximum Dynamic Range of 280dB
2.3.3.9 | · Real-Time Compounding
2.3.3.10 | · Adaptive Image Processing for noise and artifact reduction
2.3.3.11 | · The latest version of vascular, adult & child cardiac- (transthoracic and transoesophageal - preferably 4D / Live 3D Transesophageal Echocardiography), lung- and regional anaesthesia software packages should be included
2.3.4 | Spectral Doppler measurements:
2.3.4.1 | · Velocity: manual
2.3.4.2 | · Acceleration
2.3.4.3 | · Time /slope
2.3.4.4 | · Heart rate
2.3.4.5 | · PI
2.3.4.6 | · RI
2.3.4.7 | · Flow volume
2.3.4.8 | · Doppler trace
2.3.4.9 | · Automatic Doppler analysis
2.3.5 | Cardiac imaging & measurement
2.3.5.1 | · Tissue specific imaging software for adult and paediatric echo ultrasound applications
2.3.5.2 | · Display optimisation software with tissue specific pre-sets for adult and paediatric echo imaging and Doppler application
2.3.5.3 | · Analysis software for full package of adult and paediatric echo imaging protocol and report (to include standard and 3D quantification on TEE)
2.3.5.4 | · Tissue Harmonics and Pulse Inversion Harmonic Imaging
2.3.5.5 | · High Definition Write Zoom and Read Zoom with pan features
2.3.5.6 | · Chroma Imaging
2.3.5.7 | · Measurement tools including: distance, depth, area, and circumference
2.3.5.8 | · Allows operation with all the desired transducers
2.3.5.9 | · Speckle tracking should be offered as an optional functionality
2.3.5.10 | · For 3/4D:
2.3.5.11 | · 4D / Live 3D Volume Imaging using matrix transducer
2.3.5.12 | · 4D / Live 3D Zoom
2.3.5.13 | · 4D / 3D Full volume
2.3.5.14 | · 4D / Live 3D Quantification with MPR and islice.
2.3.5.15 | · 4D / Live 3D display of dynamic 3D rendering and LV true volumes.
2.3.5.16 | · Full Quantification package for 3D cardiac volume datasets
2.3.6 | Vascular Access & Regional anaesthesia
2.3.6.1 | · Tissue pre-sets for vascular access and regional anaesthesia should be available
2.3.6.2 | · Imaging software for vascular access (eg jugular vein, radial artery, femoral vessels.
2.3.6.3 | · The imaging should also be suitable for regional analgesia imaging
2.3.7 | Lung imaging
· Software should be included for optimising lung imaging
2.3.8 | Tissue Doppler:
2.3.8.1 | · Tissue Doppler with high frame rate
2.3.8.2 | · M mode colour Doppler and M mode tissue Doppler
2.3.9 | Worksheet/ Report functions:
2.3.9.1 | · Fully editable results datasheet
2.3.9.2 | · Integrated patient exam report, editable
2.3.9.3 | · Embed images in report
2.3.9.4 | · Customizable to workflow
2.3.9.5 | · On-screen recall, measurement, text editing
2.3.9.6 | · Exam directory
2.3.9.7 | · Selection between mean value display and latest value display is possible.
2.3.9.8 | · Comment entry is possible
2.3.9.9 | · The created reports can be printed
2.3.9.10 | · The created reports can be output as PDF files This should also be available for cardiac, lung, and regional anaesthesia reports
2.3.10 | Software. As described elsewhere, but to include:
2.3.10.1 | · Adult cardiac including TTE & TEE 3D & live 3D TEE as well as lung ultrasound
2.3.10.2 | · Paediatric cardiac, including TTE
2.3.10.3 | · Vascular access & regional anaesthesia
2.3.10.4 | · Dicom 3 software must be included
2.3.10.5 | Software should be ugradably and future upgrades included in the offer
2.3.10.6 | Note: Paediatric 3D cardiac echography is desirable and an alternative system with appropriate software and probes should also be tendered if available.
2.4 | MEMORY
2.4.1 | Cine/ storage:
2.4.1.1 | · Memory capacity: (Minimum 2200 frames)
2.4.1.2 | · Internal HDD of at least 512GB
2.4.1.3 | · Record/playback modes
2.4.1.4 | · Loop playback is possible
2.4.1.5 | · Frame-advance playback is possible
2.4.1.6 | · Cine playback is possible for Doppler images and M-mode images. (Minimum 30 s )
2.4.1.7 | · Clips: Live images can be recorded
2.4.2 | Archiving:
2.4.2.1 | · Hard disk drive to exceed 512 MB
2.4.2.2 | · Recorded data on CD\DVD must be readable on standard laptop\computer (with additional software from company included if necessary
2.4.2.3 | · Integrated DVD/CD burning capability for storage of DICOM images or export in JPEG and avi for PC compatibility
2.4.2.4 | · USB flash drive
2.4.2.5 | · Network: DICOM : must be compliant and must be included
2.4.2.6 | · Advanced DICOM with full work list must be offered and included
2.4.2.7 | · DICOM 3.0 Print and Store capability to internal drive or DVD/CD
2.4.2.8 | · Export of images from system must be possible, state methods
2.5 | TRANSDUCERS
All transducers required must be included in the main offer as part of the total price
2.5.1 | Transducer selection:
2.5.1.1 | · 4 Active imaging transducer connectors and one non Imaging transducer port
2.5.1.2 | · Electronic switching between transducers using at least 4 universal connectors
2.5.1.3 | · Automatic pre-set optimization of each transducer
2.5.1.4 | · High-resolution interactive graphical colour touch panel with adjustment for various ambient light conditions
2.5.2 | Transducer design:
2.5.2.1 | · Ergonomic design with lightweight flexible cables
2.5.2.2 | · Advanced penetration with less artefacts
2.5.2.3 | · Broadband frequency
2.5.2.4 | · Support for high frequencies
2.5.2.5 | · Selected frequency to be displayed on screen
2.5.2.6 | · All transducers be digitally controlled so that continuous beam forming is possible
2.5.2.7 | · Multi array cardiac transducer 1-5 MHz for Cardiac Echography as well as lung ultrasound
2.5.2.8 | · A 2 to 7 MHz Transoesophageal 3D echocardiography probe with live 3D software (adults)
2.5.2.9 | · Sector array with 48 elements
2.5.2.10 | · Manually rotatable 0-180°
2.5.2.11 | · SD steerable with PW Doppler, CW Doppler, colour Doppler, XRES and harmonic imaging
2.5.2.12 | · 5-12 MHz broadband phased array with small footprint for paediatric cardiac and general applications
2.5.2.13 | · SD steerable with PW Doppler, CW Doppler, colour Doppler, XRES and harmonic imaging
2.5.2.14 | · A 3 to 7 MHz paediatric transoesophagial (TEE) probe
2.5.2.15 | · Manually rotatable 0-180°
2.5.2.16 | · SD steerable with PW Doppler, CW Doppler, colour Doppler, XRES and harmonic imaging
2.5.3 | Broadband linear array transducer:
2.5.3.1 | · Hockey stick-like design
2.5.3.2 | · Supports vascular carotid, arterial, venous and regional anaesthesia superficial imaging applications.
2.5.3.3 | · Minimum 7 to 15 MHz frequency range
2.5.3.4 | · Supports PW, colour Doppler, power Doppler
2.5.3.5 | · Supports Strain & Strain Rate and Speckle Tracking
2.5.4 | Paediatric 3D Cardiac echocardiography is desirable and an alternative system capable of this should be tendered with appropriate probes
2.6 | SYSTEM CONTROLS\ DESIGN
2.6.1 | Fully digital, high end, colour Doppler , ultrasound unit
2.6.2 | On-board workstation-class data management with thumbnail previews and storage of images, loops, and reports
2.6.3 | System must have 12 “ or larger touch command screen
2.6.4 | Unit must be the latest technology. State year of manufacturing
2.6.5 | Display screen:
2.6.5.1 | · Nearly infinite positioning adjustments for optimum scanning, height, swivel, tilt
2.6.5.2 | · Up and down
2.6.5.3 | · Side to side movement
2.6.5.4 | · 20” wide
2.6.5.5 | · flicker free
2.6.6 | Control panel:
2.6.6.1 | · Retractable alphanumeric keyboard
2.6.6.2 | · Palm rest
2.6.6.3 | · Reduced number of hard controls
2.6.6.4 | · Primary controls clustered for convenience
2.6.6.5 | · Ambient lighting for different environments
2.6.6.6 | · Intelligent one-button optimization in 2D and Doppler modes eg Iscan
2.6.7 | Touch Screen:
2.6.7.1 | · Workflow related controls
2.6.7.2 | · Quick access to hidden controls
2.6.7.3 | · Supports alphanumeric keyboard option
2.6.7.4 | · High-resolution interactive graphical colour touch panel with adjustment for various ambient light conditions
2.6.8 | Pre-sets:
2.6.8.1 | · Single selection to choose pre-set, transducers
2.6.8.2 | · Completely editable pre-sets for specific exams
2.7 | PHYSICAL SPECIFICATIONS
2.7.1 | System Cart:
2.7.1.1 | · State of the art ergonomic design for comfort and convenience
2.7.1.2 | · Easy manoeuvrability and mobility
2.7.1.3 | · Transducer and gel bottle holders
2.7.1.4 | · High quality, shock absorbing casters with wheel lock facility and 4 wheel swivel
2.7.1.5 | · Automatic cooling capacity with minimal noise
2.7.2 | Monitor:
2.7.2.1 | · Flat screen LCD monitor at least 20 inch
2.7.2.2 | · Extended viewing angle
2.7.2.3 | · State of the art technology
2.7.2.4 | · Mounted on fully articulating extension arm
2.7.2.5 | · Nearly infinite positioning adjustment, height, swivel and tilt
2.7.3 | EKG:
2.7.3.1 | · A 3 lead ECG cable should be provided
2.7.3.2 | · Lead l is the standard connection
2.7.3.3 | · External input is possible
2.7.3.4 | · Lead selection is possible
2.7.3.5 | · Automatic heart rate calculation and display
2.7.4 | Peripherals:
2.7.4.1 | · System supports up to two on-board peripheral devices such as DVD, colour printer (USB), B/W printer (USB)
2.7.4.2 | · Support various colour printers externally mounted
2.7.4.3 | · Automated measurement tool for Speckle tracking (optional)
2.7.5 | Input/output ports:
2.7.5.1 | · Transducer Connectors: minimum 4 imaging
2.7.5.2 | · DVD input/output signals
2.7.5.3 | · USB ports
2.7.5.4 | · External video output signal
2.7.6 | Printer
3 | Power and Batteries
3.1 | The monitor must operate on both the 220V-240V , 50Hz single phase AC power supply and an internally fitted rechargeable battery. The apparatus must be fused in both the LIVE and NEUTRAL.
3.2 | The unit offered must be supplied with an rechargeable battery.
3.3 | The rechargeable battery must be automatically charged when the unit is connected to a live A.C. supply
3.4 | In the event of A.C. mains failure the battery must automatically provide continuous operation
3.5 | The unit must be provided with visible indication to reflect whether the unit is being operated off the AC power or the rechargeable battery power
3.6 | The unit offered must be provided with circuitry, which must ensure that the rechargeable battery will be protected against over-charging and over-discharging
3.7 | Unit that are quoted on and employ external battery chargers and power supply will not be considered.
3.8 | Battery back-up option should be provided (battery allowing unit to be placed in a sleep mode while being moved from one location to the next so as to provide rapid resumption of operation)
3.9 | The rechargeable battery must be of a reasonable capacity, such that with battery power the unit must be able to operate continuously for a minimum of four (4)hours
3.9 | Please state the following:
3.9.1 | Capacity
3.9.2 | Voltage
3.9.3 | Type of battery used
3.9.4 | Length of battery operating time
4 | STANDARD ACCESSORIES
4.1 | 3 lead ECG cable
4.2 | Multi array cardiac transducer 1-5 MHz for Cardiac Echography as well as lung ultrasound
4.3 | 2 to 7 MHz Transoesophageal 3D echocardiography probe with live 3D software (adults)
4.4 | 5-12 MHz broadband phased array with small footprint for paediatric cardiac and general applications
4.5 | Ultrasound Gel
4.6 | Printer
4.7 | Printing paper (1 box)
5 | Conformity Compliance (Please attach certificate)
5.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
5.2 | OEM must comply and certified on ISO 9001 quality standards, attach proof of compliance
5.2.1 | Please provide unique ref number of the ISO 9001 ceriticate:
5.3 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
5.3.1 | Please provide unique ref number of the ISO 13485 ceriticate:
5.4 | Model quoted for must be EC certified. Attach a copy of certification
5.4.1 | Please provide unique ref number of the EC ceriticate:
5.5 | Hazardous Substance Act: If this type of equipment / apparatus appears on the schedule of Hazardous substances issued by the Directorate: Health Technology of the Department of Health, a license in terms of the Act on Hazardous Substances (Act 15/1973) must be submitted with the bid / quotation documents. The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.6 | All electrical/electronic medical equipment must be licensed by Radiation Control, where applicable a copy of the license must be submitted The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.6.1 | Radiation Control License number:
5.7 | The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade (rubber to with 2 screws) type and it must be a minimum length of three (3) meters. NB the mains cable of the unit tendered for must be SABS colour coded.
5.8 | The equipment quoted for must be protected against electromagnetic interference IEC 60601-1-2
5.9 | The unit must comply with electrical safety and performance standard such as IEC 60601-2-37, attached certification
5.9.1 | Please provide unique ref number of the IEC ceriticate:
6 | COMPREHENSIVE MAINTENANCE
6.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) must be included in the bid price
6.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
6.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
6.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
6.3 | The bidder to provide the OEM recommended service interval of the unit:
6.4 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
7 | General
7.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
7.2 | The unit on offer must be the latest technology.
RT2-05-33 | Ultrasound machine, 3D or better- Cardiology Accessories and Consumables
RT2-05-33a | Uninterrupted Power Supply (UPS)
RT2-05-33b | Hockey stick-like broadband linear array transducer 7-15 MHz
RT2-05-33c | 3 to 7 MHz paediatric transoesophagial (TEE) probe
RT2-05-33d | 3 lead ECG cable
RT2-05-33e | Multi array cardiac transducer 1-5 MHz for Cardiac Echography as well as lung ultrasound
RT2-05-33f | 2 to 7 MHz Transoesophageal 3D echocardiography probe with live 3D software (adults)
RT2-05-33g | 5-12 MHz broadband phased array with small footprint for paediatric cardiac and general applications
RT2-05-33h | Printing paper
RT2-05-33i | Ultrasound Gel
Bidder to list any other Consumables or Accessories (Not listed above) for the model on offer below

## Category-6 Gen Surgery
GENERAL SURGERY EQUIPMENT
# | Item number | Item Description
1 | RT2-06-01 | Doppler, Vascular
2 | RT2-06-02 | Electrosurgical Unit, Basic
4 | RT2-06-04 | Electrosurgical Unit, Advanced
5 | RT2-06-05 | Electrosurgical Unit, Argon

## Doppler, Vascular
RT2-06-01 | DOPPLER, VASCULAR
1 | DESCRIPTION
1.1 | A handheld, battery operated vascular doppler for detection of bloodstream status of arteries/veins.
2 | Product Offered:
2.1 | Make:
2.1 | Model:
2.1 | Initial Manufacturing year of Model:
3 | TECHNICAL SPECIFICATION
3.1 | The unit offered should detect the blood flow average velocity.
3.1 | The unit offered should have audio and visual display(LCD or LED)
3.1 | The following must be displayed:
3.3.1 | Heart rate
3.3.2 | Heartbeat synchronous indicator
3.3.3 | Low battery indicator
3.3.4 | Accuracy +- 3 bpm
3.4 | Power source: Rechargerble battery and main supply of 220 volts/ 50Hz with adaptor.
3.5 | The unit offered should automatically shut off when not in use
3.6 | The unit offered should have interchangeable probes (Vascular and Foetal)
3.7 | The unit offered must be in a protective case
3.8 | Vascular Probe
3.8.1 | The probe offered must be waterproof
3.8.2 | The probe of the unit offered should fully operate at 4Mhz to 8Mhz for different vascular extremities.
4 | STANDARD ACCESSORIES
4.1 | Vascular probe
4.1 | Rechargeble Battery
4.1 | Power adapter
5 | Conformity Compliance (Please attach certificate)
5.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
5.2 | OEM must comply and certified on ISO 9001 quality standards, attach proof of compliance
5.2.1 | Please provide unique ref number of the ISO 9001 ceriticate:
5.3 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
5.3.1 | Please provide unique ref number of the ISO 13485 ceriticate:
5.4 | Model quoted for must be EC certified. Attach a copy of certification
5.4.1 | Please provide unique ref number of the EC ceriticate:
6 | COMPREHENSIVE MAINTENANCE
6.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) must be included in the bid price
6.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
7 | General
7.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
7.2 | The unit on offer must be the latest technology.
DOPPLER, VASCULAR
RT2-06-01a | Vascular probe
RT2-06-01b | Rechargeble Battery
RT2-06-01c | Power adapter
Bidder to list any other Consumables or Accessories (Not listed above) for the model on offer below

## Electrosurgery basic
RT2-06-02 | Electrosurgical unit basic
1 | DESCRIPTION
1.1 | A unit intended for surgical cutting and for controlling bleeding by coagulation at the surgical site.
2 | Product Offered:
2.1 | Make:
2.1 | Model:
2.1 | Initial Manufacturing year of Model:
2.1 | Radiation Control License number:
3 | TECHNICAL SPECIFICATION
3.1 | Electrosurgical cutting and coagulation unit, 100W, mobile
3.2 | Electrosurgical unit, for general surgery
3.3 | Microprocessor controlled generator with mono-polar and bipolar output
3.4 | The unit must employ solid state technology for the generation of the various outputs required
3.5 | The unit must do a self-test upon start-up
3.6 | The unit must be able to perform the following functions in monopolar mode:
3.6.1 | Pure cut
3.6.2 | Coagulation
3.6.3 | Blend
3.6.4 | Spray coagulation
3.7 | The output power selected must be digitally displayed
3.8 | Output frequency: approx 400 kHz
3.9 | Max output power, mono-polar: up to 100 W (cutting),
3.10 | Bipolar output power: 50 W (coagulation)
3.11 | Double function foot switch (mono and bipolar)
3.12 | Hand switch mode when button-activated probes are connected
3.13 | Visual and audible activation indicators
3.14 | Audible and visual warnings must be given when:
3.14.1 | Cut function is activated;
3.14.2 | Coagulation function is activated;
3.14.3 | Neutral electrode (patient plate) disconnected from the unit
3.14.4 | The unit has been connected with a faulty neutral electrode or cable;
3.15 | When an alarm associated within the neutral electrode occurs, the power output of the unit must be reduced to zero.
3.16 | Reusable neutral electrode to be supplied with the unit
3.17 | The unit must monitor patient contact with the neutral electrode
3.18 | The unit must fail-safe should the patient make insufficient contact with the neutral electrode.
3.19 | Return circuit sensing monitor that deactivates generator in case patient plate or cable fails
3.20 | Display: output power, system errors and electrode failure
3.21 | Power requirements: 220 V, 50 Hz
3.22 | Mounted on a sturdy trolley with 4 antistatic swivel castors, 2 with brakes
3.23 | Trolley must have a drawer and storage for foot pedal and cable
3.24 | Mains cable at least three metres long with SABS approved 15A three prong plug.
3.25 | The unit must be handed over in full operating order
3.26 | A starter-pack of consumables must be supplied with the unit.
4 | Standard Accessories and Consumables
4.1 | Foot switch, double, with cable (1x)
4.2 | Patient electrode, adult, stainless steel or conductive rubber, reusable
4.3 | Patient electrode cable
4.4 | Patient electrode, split, suitable for monitoring patient contact, reusable, sterlizable
4.5 | Cable for split patient electrode
4.6 | Electrode handle, mono-polar, reusable, foot switch controlled, (2x) with connecting cable
4.7 | Electrode handle, mono-polar, reusable, finger switch controlled, (2x) with connecting cable
4.8 | Set different mono-polar reusable electrodes (needle, blade, ball and loop) (1x)
4.9 | Bipolar forceps, reusable, foot switch controlled, with connecting cable (1x)
4.10 | Dedicated trolley
5 | Conformity Compliance (Please attach certificate)
5.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
5.2 | OEM must comply and certified on ISO 9001 quality standards, attach proof of compliance
5.2.1 | Please provide unique ref number of the ISO 9001 ceriticate:
5.3 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
5.3.1 | Please provide unique ref number of the ISO 13485 ceriticate:
5.4 | Model quoted for must be EC certified. Attach a copy of certification
5.4.1 | Please provide unique ref number of the EC ceriticate:
5.5 | Hazardous Substance Act: If this type of equipment / apparatus appears on the schedule of Hazardous substances issued by the Directorate: Health Technology of the Department of Health, a license in terms of the Act on Hazardous Substances (Act 15/1973) must be submitted with the bid / quotation documents. The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.6 | All electrical/electronic medical equipment must be licensed by Radiation Control, where applicable a copy of the license must be submitted The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.6.1 | Radiation Control License number:
5.7 | The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade (rubber to with 2 screws) type and it must be a minimum length of three (3) meters. NB the mains cable of the unit tendered for must be SABS colour coded.
5.8 | The equipment quoted for must be protected against electromagnetic interference IEC 60601-1-2
5.9 | The unit must comply with electrical safety and performance standard such as IEC 60601-2-2, attached certification
5.9.1 | Please provide unique ref number of the IEC ceriticate:
6 | COMPREHENSIVE MAINTENANCE
6.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) must be included in the bid price
6.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
6.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
6.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
6.3 | The bidder to provide the OEM recommended service interval of the unit:
6.4 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
7 | General
7.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
7.2 | The unit on offer must be the latest technology.
Electrosurgical unit basic - Accessories and Consumables
RT2-06-02a | Foot switch, double, with cable
RT2-06-02b | Patient electrode, adult, stainless steel or conductive rubber, reusable
RT2-06-02c | Patient electrode cable
RT2-06-02d | Patient electrode, split, suitable for monitoring patient contact, reusable, sterlizable
RT2-06-02e | Cable for split patient electrode
RT2-06-02f | Electrode handle, mono-polar, reusable, foot switch controlled, (2x) with connecting cable
RT2-06-02g | Electrode handle, mono-polar, reusable, finger switch controlled, (2x) with connecting cable
RT2-06-02h | Set different mono-polar reusable electrodes (needle, blade, ball and loop) (1x)
RT2-06-02i | Bipolar forceps, reusable, foot switch controlled, with connecting cable (1x)
RT2-06-02j | Disposable single neutral electrode pad
RT2-06-02k | Disposable split neutral electrode pad
Bidder to list any other Consumables or Accessories (Not listed above) for the model on offer below

## Electrosurgery advanced
RT2-06-04 | Electrosurgery unit, advanced
1 | DESCRIPTION
1.1 | A unit intended for surgical cutting and for controlling bleeding by coagulation at the surgical site.
2 | Product Offered:
2.1 | Make:
2.2 | Model:
2.3 | Initial Manufacturing year of Model:
2.4 | Radiation Control License number:
3 | TECHNICAL SPECIFICATION
3.1 | Must be programmable for the full range of electrosurgical requirements:
3.1.1 | General surgery
3.1.2 | Cardiology
3.1.3 | ENT
3.1.4 | Gastroenterology
3.1.5 | Gynaecology
3.1.6 | Laparoscopy
3.1.7 | Neurosurgery
3.1.8 | Thoracic surgery
3.1.9 | Urology (TURP, TUVP)
3.2 | User must be able to enter and save preference settings
3.3 | Power output: at least 400 W with output isolated from ground
3.4 | Automatic power adjustment for variation in tissue impedance
3.5 | Automatic recognition of the electrode used and adjustment of the output parameters accordingly
3.6 | The unit must employ solid state technology for the generation of the various outputs required
3.7 | The unit must do a self-test at switch-on
3.8 | The unit must have the following modes of operations:
3.8.1 | Monopolar mode
3.8.2 | Bipolar mode
3.8.3 | Suitable for two operators with active outputs simultaneously
3.8.4 | Two footswitches active simultaneously
3.8.5 | One footswitch and one handswitch active simultaneously
3.8.6 | Two handswitches active simultaneously
3.8.7 | Monopolar and bipolar active simultaneously
3.9 | The unit must be able to perform the following functions in monopolar mode:
3.9.1 | Pure cut
3.9.2 | Coagulation
3.9.3 | Blend
3.9.4 | Spray coagulation
3.10 | The pure cut function must have variable power from up to 400 W with a 500 Ohm load
3.11 | The output power selected must be digitally displayed
3.12 | The coagulation function must have the following output:
3.12.1 | Variable power from 0 to 120 W delivered into a 500 Ohm load in the coagulation pin point mode
3.12.2 | Variable power from 0 to 100 W delivered into a 500 Ohm load in the coagulation spray mode
3.13 | The blend function must have a variable control of the degree of haemostasis.
3.14 | The Bipolar mode must deliver variable power from 0 to 50 W into a 125 Ohm load in both the Bipolar Cut and Bipolar Coagulation modes.
3.15 | Frequency: approximately 500 kHz.
3.16 | Memory for at least 10 programs with their waveforms and power levels
3.17 | Visual and audible activation indicators
3.18 | Audible and visual warnings must be given when:
3.18.1 | Cut function is activated;
3.18.2 | Coagulation function is activated;
3.18.3 | Neutral electrode (patient plate) disconnected from the unit
3.18.4 | The unit has been connected with a faulty neutral electrode or cable;
3.19 | When an alarm associated within the neutral electrode occurs, the power output of the unit must be reduced to zero.
3.20 | Reusable neutral electrode to be supplied
3.21 | The unit must monitor patient contact with the neutral electrode
3.22 | The unit must fail-safe should the patient make insufficient contact with the neutral electrode.
3.23 | The output power to the active electrode must be activated by either foot pedal control or a finger-operated switch on the handpiece of the active electrode
3.24 | There must be no possibility of activating both hand and foot operated accessories simultaneously.
3.25 | The foot switch must have a barrier between the two pedals
3.26 | The foot pedal cable must be detachable from both the electrosurgical unit and the foot switch
3.27 | The foot pedal must be splash proof.
3.28 | The unit’s power on switch must be enclosed to prevent accidental switch On or Off;
3.29 | The equipment must function off 220 V, 50 Hz single phase mains supply
3.30 | Mains cable must be at least 3 metres long and have a SABS approved 15A three pin plug
3.31 | The unit must be supplied with a trolley on antistatic castors, with brakes and a drawer to store footswitch and accessories.
3.32 | The unit must be handed over in full operating order
3.33 | A starter-pack of consumables must be supplied with the unit.
4 | Standard Accessories and Consumables
4.1 | Foot switch, double, with cable (1x)
4.2 | Patient electrode, adult, stainless steel or conductive rubber, reusable
4.3 | Patient electrode cable
4.4 | Patient electrode, split, suitable for monitoring patient contact, reusable, sterlizable
4.5 | Cable for split patient electrode
4.6 | Electrode handles with cables, reusable, sterilizable (2x)
4.7 | Blade/knife electrodes, straight, reusable (2x)
4.8 | Needle electrodes (thick) (2x)
4.9 | Needle electrode (thin) (1x)
4.10 | Ball electrodes, straight, 4 mm diameter ((2x)
4.11 | Ball electrode, straight, 2,5 mm diameter (1x)
4.12 | Bipolar electrodes reusable, sterilizable (2x)
5 | Conformity Compliance (Please attach certificate)
5.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
5.2 | OEM must comply and certified on ISO 9001 quality standards, attach proof of compliance
5.2.1 | Please provide unique ref number of the ISO 9001 ceriticate:
5.3 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
5.3.1 | Please provide unique ref number of the ISO 13485 ceriticate:
5.4 | Model quoted for must be EC certified. Attach a copy of certification
5.4.1 | Please provide unique ref number of the EC ceriticate:
5.5 | Hazardous Substance Act: If this type of equipment / apparatus appears on the schedule of Hazardous substances issued by the Directorate: Health Technology of the Department of Health, a license in terms of the Act on Hazardous Substances (Act 15/1973) must be submitted with the bid / quotation documents. The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.6 | All electrical/electronic medical equipment must be licensed by Radiation Control, where applicable a copy of the license must be submitted The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.6.1 | Radiation Control License number:
5.7 | The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade (rubber to with 2 screws) type and it must be a minimum length of three (3) meters. NB the mains cable of the unit tendered for must be SABS colour coded.
5.8 | The equipment quoted for must be protected against electromagnetic interference IEC 60601-1-2
5.9 | The unit must comply with electrical safety and performance standard such as IEC 60601-2-2, attached certification
5.9.1 | Please provide unique ref number of the IEC ceriticate:
6 | COMPREHENSIVE MAINTENANCE
6.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) must be included in the bid price
6.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
6.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
6.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
6.3 | The bidder to provide the OEM recommended service interval of the unit:
6.4 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
7 | General
7.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
7.2 | The unit on offer must be the latest technology.
Electrosurgery unit, Advanced- Accessories and Consumables
RT2-06-04a | Foot switch, double, with cable
RT2-06-04b | Patient electrode, adult, stainless steel or conductive rubber, reusable
RT2-06-04c | Patient electrode cable
RT2-06-04d | Patient electrode, split, suitable for monitoring patient contact, reusable, sterlizable
RT2-06-04e | Cable for split patient electrode
RT2-06-04f | Electrode handle, mono-polar, reusable, foot switch controlled, (2x) with connecting cable
RT2-06-04g | Electrode handle, mono-polar, reusable, finger switch controlled, (2x) with connecting cable
RT2-06-04h | Set different mono-polar reusable electrodes (needle, blade, ball and loop) (1x)
RT2-06-04i | Bipolar forceps, reusable, foot switch controlled, with connecting cable (1x)
RT2-06-04j | Disposable single neutral electrode pad
RT2-06-04k | Disposable split neutral electrode pad
RT2-06-04l | Vessel sealing
RT2-06-04m | Memory for at least 10 programs with their waveforms and power levels
RT2-06-04n | Smoke evacuation
Bidder to list any other Consumables or Accessories (Not listed above) for the model on offer below

## Electrosurgery with Argon
RT2-06-05 | Electrosurgery unit, with Argon Module
1 | DESCRIPTION
1.1 | A unit intended for surgical cutting and for controlling bleeding by coagulation at the surgical site.
2 | Product Offered:
2.1 | Make:
2.2 | Model:
2.3 | Initial Manufacturing year of Model:
2.4 | Radiation Control License number:
3 | TECHNICAL SPECIFICATION
3.1 | Must be programmable for the full range of electrosurgical requirements:
3.1.1 | General surgery
3.1.2 | Cardiology
3.1.3 | ENT
3.1.4 | Gastroenterology
3.1.5 | Gynaecology
3.1.6 | Laparoscopy
3.1.7 | Neurosurgery
3.1.8 | Thoracic surgery
3.1.9 | Urology (TURP, TUVP)
3.2 | User must be able to enter and save preference settings
3.3 | Power output: at least 400 W with output isolated from ground
3.4 | Automatic power adjustment for variation in tissue impedance
3.5 | Automatic recognition of the electrode used and adjustment of the output parameters accordingly
3.6 | The unit must employ solid state technology for the generation of the various outputs required
3.7 | The unit must do a self-test at switch-on
3.8 | The unit must have the following modes of operations:
3.8.1 | Monopolar mode
3.8.2 | Bipolar mode
3.8.3 | Suitable for two operators with active outputs simultaneously
3.8.4 | Two footswitches active simultaneously
3.8.5 | One footswitch and one handswitch active simultaneously
3.8.6 | Two handswitches active simultaneously
3.8.7 | Monopolar and bipolar active simultaneously
3.9 | The unit must be able to perform the following functions in monopolar mode:
3.9.1 | Pure cut
3.9.2 | Coagulation
3.9.3 | Blend
3.9.4 | Spray coagulation
3.9.5 | Argon cutting & coagulation
3.9.6 | Argon pulse coagulation
3.10 | The pure cut function must have variable power from up to 400 W with a 500 Ohm load
3.11 | The output power selected must be digitally displayed
3.12 | The coagulation function must have the following output:
3.13 | Variable power from 0 to 120 W delivered into a 500 Ohm load in the coagulation pin point mode
3.14 | Variable power from 0 to 100 W delivered into a 500 Ohm load in the coagulation spray mode
3.15 | The blend function must have a variable control of the degree of haemostasis.
3.16 | The Bipolar mode must deliver variable power from 0 to 50 W into a 125 Ohm load in both the Bipolar Cut and Bipolar Coagulation modes.
3.17 | Frequency: approximately 500 kHz.
3.18 | Memory for at least 10 programs with their waveforms and power levels
3.19 | Visual and audible activation indicators
3.20 | Audible and visual warnings must be given when:
3.20.1 | Cut function is activated;
3.20.2 | Coagulation function is activated;
3.20.3 | Neutral electrode (patient plate) disconnected from the unit
3.20.4 | The unit has been connected with a faulty neutral electrode or cable;
3.21 | When an alarm associated within the neutral electrode occurs, the power output of the unit must be reduced to zero.
3.22 | Reusable neutral electrode to be supplied
3.23 | The unit must monitor patient contact with the neutral electrode
3.24 | The unit must fail-safe should the patient make insufficient contact with the neutral electrode.
3.25 | The output power to the active electrode must be activated by either foot pedal control or a finger-operated switch on the handpiece of the active electrode
3.26 | There must be no possibility of activating both hand and foot operated accessories simultaneously.
3.27 | The foot switch must have a barrier between the two pedals
3.28 | The foot pedal cable must be detachable from both the electrosurgical unit and the foot switch
3.29 | The foot pedal must be splash proof.
3.30 | The unit’s power on switch must be enclosed to prevent accidental switch On or Off;
3.31 | Argon
3.31.1 | The unit must utilize pure argon with inlet gas pressure 3 - 5 bar
3.31.2 | Argon cylinder must have pressure reducer
3.31.3 | Gas Flow range between 0.1 - 10 litre per minutes
3.31.4 | Gas filtration must be possible. Please state how it is achieved and filtration range
3.31.5 | The unit must have overpressure monitoring
3.32 | The equipment must function off 220 V, 50 Hz single phase mains supply
3.33 | Mains cable must be at least 3 metres long and have a SABS approved 15A three pin plug
3.34 | The unit must be supplied with a trolley on antistatic castors, with brakes and a drawer to store footswitch and accessories.
3.35 | The unit must be handed over in full operating order
3.36 | A starter-pack of consumables must be supplied with the unit.
4 | Standard Accessories and Consumables
4.1 | Foot switch, double, with cable (1x)
4.2 | Patient electrode, adult, stainless steel or conductive rubber, reusable
4.3 | Patient electrode cable
4.4 | Patient electrode, split, suitable for monitoring patient contact, reusable, sterlizable
4.5 | Cable for split patient electrode
4.6 | Electrode handles with cables, reusable, sterilizable (2x)
4.7 | Blade/knife electrodes, straight, reusable (2x)
4.8 | Needle electrodes (thick) (2x)
4.9 | Needle electrode (thin) (1x)
4.10 | Ball electrodes, straight, 4 mm diameter ((2x)
4.11 | Ball electrode, straight, 2,5 mm diameter (1x)
4.12 | Bipolar electrodes reusable, sterilizable (2x)
4.13 | Flexible argon probe
4.14 | Argon electrode connection
5 | Conformity Compliance (Please attach certificate)
5.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
5.2 | OEM must comply and certified on ISO 9001 quality standards, attach proof of compliance
5.2.1 | Please provide unique ref number of the ISO 9001 ceriticate:
5.3 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
5.3.1 | Please provide unique ref number of the ISO 13485 ceriticate:
5.4 | Model quoted for must be EC certified. Attach a copy of certification
5.4.1 | Please provide unique ref number of the EC ceriticate:
5.5 | Hazardous Substance Act: If this type of equipment / apparatus appears on the schedule of Hazardous substances issued by the Directorate: Health Technology of the Department of Health, a license in terms of the Act on Hazardous Substances (Act 15/1973) must be submitted with the bid / quotation documents. The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.6 | All electrical/electronic medical equipment must be licensed by Radiation Control, where applicable a copy of the license must be submitted The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.6.1 | Radiation Control License number:
5.7 | The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade (rubber to with 2 screws) type and it must be a minimum length of three (3) meters. NB the mains cable of the unit tendered for must be SABS colour coded.
5.8 | The equipment quoted for must be protected against electromagnetic interference IEC 60601-1-2
5.9 | The unit must comply with electrical safety and performance standard such as IEC 60601-2-2, attached certification
5.9.1 | Please provide unique ref number of the IEC ceriticate:
6 | COMPREHENSIVE MAINTENANCE
6.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) must be included in the bid price
6.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
6.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
6.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
6.3 | The bidder to provide the OEM recommended service interval of the unit:
6.4 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
7 | General
7.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
7.2 | The unit on offer must be the latest technology.
Electrosurgery unit, with Argon Module - Accessories and Consumables
RT2-06-05a | Foot switch, double, with cable
RT2-06-05b | Patient electrode, adult, stainless steel or conductive rubber, reusable
RT2-06-05c | Patient electrode cable
RT2-06-05d | Patient electrode, split, suitable for monitoring patient contact, reusable, sterlizable
RT2-06-05e | Cable for split patient electrode
RT2-06-05f | Electrode handles with cables, reusable, sterilizable
RT2-06-05g | Blade/knife electrodes, straight, reusable
RT2-06-05h | Needle electrodes (thick)
RT2-06-05i | Needle electrode (thin)
RT2-06-05j | Ball electrodes, straight, 4 mm diameter
RT2-06-05k | Ball electrode, straight, 2,5 mm diameter
RT2-06-05l | Bipolar electrodes reusable, sterilizable
RT2-06-05m | Flexible argon probe
RT2-06-05n | Argon electrode connection
RT2-06-05o | Disposable single and split neutral electrode pads
RT2-06-05p | Vessel sealing
RT2-06-05q | Smoke evacuation
Bidder to list any other Consumables or Accessories (Not listed above) for the model on offer below

## Category-7 Ortho
ORTHOPAEDIC EQUIPMENT
# | Item number | Item Description
1 | RT2-07-01 | Electric Drill, Small bones
2 | RT2-07-02 | Electric Drill, Large bones
3 | RT2-07-03 | Battery-Operated Drill, Small bones
4 | RT2-07-04 | Battery-Operated Drill, Large bones
5 | RT2-07-05 | Tourniquet
6 | RT2-07-06 | Plaster Saw
7 | RT2-07-07 | Plaster Saw with Vacuum Extractor
8 | RT2-07-08 | Dermatome, Cordless
9 | RT2-07-09 | Dermatome, Pneumatic
10 | RT2-07-10 | Skin Graft Mesher

## Drill, Electric, Small bones
RT2-07-01 | Drill, Orthopaedic, Electric, Small Bones
1 | DESCRIPTION
1.1 | A rotating cutting instrument for making holes in hard substances, such as bones
2 | Product Offered:
2.1 | Make:
2.1 | Model:
2.1 | Initial Manufacturing year of Model:
3 | SPECIFICATIONS
3.1 | The unit must have a handpiece Cord of a minimum length of 3 metres.
3.2 | The unit must be made out of a durable, robust light weight material for easy and smooth operating
3.3 | The unit must operate of 220 Volt + 10% - 50Hz – Power consumption 500VA
3.4 | The unit must be used for large bones
3.6 | The unit quoted for must be provided with a modular Handpiece with the following attachment (Drill / Ream Mode): SPECIFY torque, speed and cannel size where applicable
3.6.1 | Drill Attachments
3.6.1a | Angled
3.6.1b | Medium straight
3.6.1c | Long Straight
3.6.1d | Long Angled
3.6.2 | Reaming / High torque attachments
3.6.3 | Reaming / High attachments
3.6.4 | TAP mode with drill attachments only
3.6.5 | Screw mode with drill attachments only
4 | Standard Accessories. SPECIFY torque, speed and cannel size where applicable
4.1 | Trinkle / AO attachments, which must accept trinkle and ASIF / AO, type accessories
4.2 | Reaming attachment, which must accept large HEX. AO Reamers
4.3 | Sagittal saw attachment
4.4 | 4.0mm Jacobs chuck and a key
4.5 | Oscillating saw
4.6 | Pin driver
4.7 | Wire driver
5 | Conformity Compliance (Please attach certificate)
5.1 | The unit must comply with an acceptable international standard such as ISO 9714-1 for Orthopaedic surgical equipment, attached certification
5.1.1 | Please provide unique ref number of the ISO 9714-1 ceriticate:
5.2 | OEM must comply and certified on ISO 9001 quality standards, attach proof of compliance
5.2.1 | Please provide unique ref number of the ISO 9001 ceriticate:
5.3 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
5.3.1 | Please provide unique ref number of the ISO 13485 ceriticate:
5.4 | Model quoted for must be EC certified. Attach a copy of certification
5.4.1 | Please provide unique ref number of the EC ceriticate:
5.5 | The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade (rubber to with 2 screws) type and it must be a minimum length of three (3) meters. NB the mains cable of the unit tendered for must be SABS colour coded.
5.6 | The equipment quoted for must be protected against electromagnetic interference IEC 60601-1-2
5.7 | The unit must comply with ISO 5835 standard, attached certification
5.7.1 | Please provide unique ref number of the ISO 5835 ceriticate:
6 | COMPREHENSIVE MAINTENANCE
6.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) must be included in the bid price
6.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
6.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
6.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
6.3 | The bidder to provide the OEM recommended service interval of the unit:
6.4 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
7 | General
7.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
7.2 | The unit on offer must be the latest technology.
Drill, Orthopaedic, Electric, Small Bones
RT2-07-01a | Trinkle / AO attachments, which must accept trinkle and ASIF / AO, type accessories
RT2-07-01b | Reaming attachment, which must accept large HEX. AO Reamers
RT2-07-01c | Sagittal saw attachment
RT2-07-01d | 4.0mm Jacobs chuck and a key
RT2-07-01e | Oscillating saw
RT2-07-01f | Pin driver
RT2-07-01g | Wire driver
Bidder to list any other Consumables or Accessories (Not listed above) for the model on offer below

## Drill, Electric, Large bones
RT2-07-02 | Drill, Orthopaedic, Electric, Large Bones
1 | DESCRIPTION
1.1 | A rotating cutting instrument for making holes in hard substances, such as bones
2 | Product Offered:
2.1 | Make:
2.2 | Model:
2.3 | Initial Manufacturing year of Model:
3 | SPECIFICATIONS
3.1 | The unit must have a handpiece Cord of a minimum length of 3 metres.
3.2 | The unit must be made out of a durable robust material light weight and for easy and smooth operating
3.3 | The unit must operate of 220 Volt + 10% - 50Hz – Power consumption 500VA
3.4 | The unit must be used for large bones
3.6 | The unit quoted for must be provided with a modular Handpiece with the following attachment (Drill / Ream Mode): SPECIFY torque, speed and cannel size where applicable
3.6.1 | Drill Attachments
3.6.1a | Angled
3.6.1b | Medium straight
3.6.1c | Long Straight
3.6.1d | Long Angled
3.6.2 | Reaming / High torque attachments
3.6.3 | Reaming / High attachments
3.6.4 | TAP mode with drill attachments only
3.6.5 | Screw mode with drill attachments only
4 | Standard Accessories. SPECIFY torque, speed and cannel size where applicable
4.1 | Trinkle / AO attachments, which must accept trinkle and ASIF / AO, type accessories
4.2 | Reaming attachment, which must accept large HEX. AO Reamers
4.3 | Sagittal saw, attachments
4.4 | 4.0mm Jacobs chuck with key
4.5 | Oscillating saw
4.6 | Pin driver
4.7 | Wire driver
5 | Conformity Compliance (Please attach certificate)
5.1 | The unit must comply with an acceptable international standard such as ISO 9714-1 for Orthopaedic surgical equipment, attached certification
5.1.1 | Please provide unique ref number of the ISO 9714-1 ceriticate:
5.2 | OEM must comply and certified on ISO 9001 quality standards, attach proof of compliance
5.2.1 | Please provide unique ref number of the ISO 9001 ceriticate:
5.3 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
5.3.1 | Please provide unique ref number of the ISO 13485 ceriticate:
5.4 | Model quoted for must be EC certified. Attach a copy of certification
5.4.1 | Please provide unique ref number of the EC ceriticate:
5.5 | The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade (rubber to with 2 screws) type and it must be a minimum length of three (3) meters. NB the mains cable of the unit tendered for must be SABS colour coded.
5.6 | The equipment quoted for must be protected against electromagnetic interference IEC 60601-1-2
5.7 | The unit must comply with ISO 5835 standard, attached certification
5.7.1 | Please provide unique ref number of the ISO 5835 ceriticate:
6 | COMPREHENSIVE MAINTENANCE
6.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) must be included in the bid price
6.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
6.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
6.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
6.3 | The bidder to provide the OEM recommended service interval of the unit:
6.4 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
7 | General
7.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
7.2 | The unit on offer must be the latest technology.
Drill, Orthopaedic, Electric, Large Bones
RT2-07-02a | Trinkle / AO attachments, which must accept trinkle and ASIF / AO, type accessories
RT2-07-02b | Reaming attachment, which must accept large HEX. AO Reamers
RT2-07-02c | Sagittal saw attachment
RT2-07-02d | 4.0mm Jacobs chuck and a key
RT2-07-02e | Oscillating saw
RT2-07-02f | Pin driver
RT2-07-02g | Wire driver
Bidder to list any other Consumables or Accessories (Not listed above) for the model on offer below

## Drill, Battery-Op, Small bones
RT2-07-03 | Drill, Orthopaedic, Battery Operated, Small bones
1 | DESCRIPTION
1.1 | The Orthopeadic drill must be able to do the following functions:
1.1.1 | Cleaning of bones(Reaming)
1.1.2 | Tapping
1.1.3 | Screwing in and unscrewing of screws
1.1.4 | Small and large drill operation
1.1.5 | To insert intramedulary nails
1.1.6 | Hip replacement
1.1.7 | Truamatology
1.1.8 | Applying of Kirschner wires in bone surgery.
2 | Product Offered:
2.1 | Make:
2.2 | Model:
2.3 | Initial Manufacturing year of Model:
3 | SPECIFICATIONS
3.1 | The equipment must be made out of a durable robust material light weight and for easy and smooth operating.
3.2 | The unit must have two (2) triggers, it must allow for variable speed in both forward and reverse directions.
3.3 | The unit must have a pistol grip hand piece for single handed operation.
3.4 | All the attachments for the unit must be quick – coupling.
3.5 | The unit must be supplied with reciprocating saw.
3.6 | The unit must be provided with an Aseptic battery which must be easily attached and detached from the base of the hand-piece. If alternate batteries and methods of sterilization are available, PLEASE SPECIFY.
3.7 | The Orthopeadic drill must have the following:
3.7.1 | a) Built-in motor and rechargeable battery for power supply of the motor.
3.7.2 | b) Electronic circuitry for controlling its rotational speed. The motor speed must be adjusted continously.
3.7.3 | The housing of the battery must be attached to the handpiece.
3.7.4 | The battery compartment should have watertight cap that keeps out of germs.
3.7.5 | The battery charger must be a modular type and able to allow minimum of 4 batteries to charge at the same time.
3.7.6 | The battery charger must have a screen/indicator to allow users to assess the condition of the battery
3.7.7 | The unit must be supplied with a minimum of 2 Li-Ion rechargeable batteries and the price must be included in the bidder's price.
3.7.8 | NiCAd batteries will not be accepted.
3.7.9 | The unit must be supplied with battery transfer shield for transferring of unsterile battery to sterile drill.
4 | The drill must be supplied with the following accessories, SPECIFY torque, speed and cannel size where applicable
4.1 | Basic hand piece with attachment coupling that accept variety of attachements
4.2 | Jacob Chuck and a key
4.3 | AO/ASIF quick coupling
4.4 | Quick coupling for Kirschner wires
4.5 | Quick coupling for drill bits
4.6 | Large quick coupling
4.7 | Reduction drive unit
4.8 | Drill chuck keyless
4.9 | Instrument tray where all accessories can be placed for cleaning
4.1 | Oscillating/Sagittal saw attachement with different blades which are:
4.11 | Sterilization case
4.12 | Hudson Modified Trinkle
4.13 | Li-Ion battery x2
4.14 | Battery charger
4.15 | Battery trabsfer shield
5 | Conformity Compliance (Please attach certificate)
5.1 | The unit must comply with an acceptable international standard such as ISO 9714-1 for Orthopaedic surgical equipment, attached certification
5.1.1 | Please provide unique ref number of the ISO 9714-1 ceriticate:
5.2 | OEM must comply and certified on ISO 9001 quality standards, attach proof of compliance
5.2.1 | Please provide unique ref number of the ISO 9001 ceriticate:
5.3 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
5.3.1 | Please provide unique ref number of the ISO 13485 ceriticate:
5.4 | Model quoted for must be EC certified. Attach a copy of certification
5.4.1 | Please provide unique ref number of the EC ceriticate:
5.5 | The unit must comply with ISO 5835 standard, attached certification
5.5.1 | Please provide unique ref number of the ISO 5835 ceriticate:
6 | COMPREHENSIVE MAINTENANCE
6.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) must be included in the bid price
6.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
6.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
6.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
6.3 | The bidder to provide the OEM recommended service interval of the unit:
6.4 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
7 | General
7.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
7.2 | The unit on offer must be the latest technology.
Drill, Orthopaedic, Battery Operated, Small bones
RT2-07-03a | Basic hand piece with attachment coupling that accept variety of attachements
RT2-07-03b | Jacob Chuck and a key
RT2-07-03c | AO/ASIF quick coupling
RT2-07-03d | Quick coupling for Kirschner wires
RT2-07-03e | Quick coupling for drill bits
RT2-07-03f | Large quick coupling
RT2-07-03g | Reduction drive unit
RT2-07-03h | Drill chuck keyless
RT2-07-03i | Instrument tray where all accessories can be placed for cleaning
RT2-07-03j | Oscillating/Sagittal saw attachement with different blades which are:
RT2-07-03k | Sterilization case
RT2-07-03l | Hudson Modified Trinkle
RT2-07-03m | Li-Ion battery x2
RT2-07-03n | Battery charger
RT2-07-03o | Battery trabsfer shield
Bidder to list any other Consumables or Accessories (Not listed above) for the model on offer below

## Drill, Battery-Op, Large bones
RT2-07-04 | Drill, Orthopaedic, Battery Operated, Large Bones
1 | DESCRIPTION
1.1 | The Orthopeadic drill must be able to do the following functions:
1.1.1 | Cleaning of bones(Reaming)
1.1.2 | Tapping
1.1.3 | Screwing in and unscrewing of screws
1.1.4 | Small and large drill operation
1.1.5 | To insert intramedulary nails
1.1.6 | Hip replacement
1.1.7 | Truamatology
1.1.8 | Applying of Kirschner wires in bone surgery.
1a | Product Offered:
1a.1 | Make:
1a.2 | Model:
1a.3 | Initial Manufacturing year of Model:
2 | SPECIFICATIONS
2.1 | The equipment must be made out of a durable robust material light weight and for easy and smooth operating.
2.2 | The unit must have two (2) triggers, it must allow for variable speed in both forward and reverse directions.
2.3 | The unit must have a pistol grip hand piece for single handed operation.
2.4 | All the attachments for the unit must be quick – coupling.
2.5 | The unit must be supplied with reciprocating saw.
2.6 | The unit must be provided with an Aseptic battery which must be easily attached and detached from the base of the hand-piece. If alternate batteries and methods of sterilization are available, PLEASE SPECIFY.
3 | The Orthopeadic drill must have the following:
3.1 | a) Built-in electric motor and rechargeable battery for power supply of the motor.
3.2 | b) Electronic circuitry for controlling its rotational speed. The motor speed must be adjusted continously.
3.3 | The housing of the battery must be attached to the handpiece.
3.4 | The battery compartment should have watertight cap that keeps out of germs.
3.5 | The battery charger must be a modular type and able to allow minimum of 4 batteries to charge at the same time.
3.6 | The battery charger must have a screen/indicator to allow users to assess the condition of the battery
3.6 | The unit must be supplied with a minimum of 2 Li-Ion rechargeable batteries and the price must be included in the bidder's price.
3.7 | NiCAd batteries will not be accepted.
3.8 | The unit must be supplied with battery transfer shield for transferring of unsterile battery to sterile drill.
4 | The drill must be supplied with the following accessories, SPECIFY torque, speed and cannel size where applicable
4.1 | Basic hand piece with attachment coupling that accept variety of attachements
4.2 | Jacob Chuck key with key
4.3 | AO/ASIF quick coupling
4.4 | Quick coupling for Kirschner wires
4.5 | Quick copuling for drill bits
4.6 | Large quick coupling
4.7 | Reduction drive unit
4.8 | Drill chuck keyless
4.9 | Instrument tray where all accessories can be placed for cleaning
4.1 | Oscillating/Sagittal saw attachement with different blades which are:
4.11 | Sterilization case
4.12 | Hudson Modified Trinkle
5 | Conformity Compliance (Please attach certificate)
5.1 | The unit must comply with an acceptable international standard such as ISO 9714-1 for Orthopaedic surgical equipment, attached certification
5.1.1 | Please provide unique ref number of the ISO 9714-1 ceriticate:
5.2 | OEM must comply and certified on ISO 9001 quality standards, attach proof of compliance
5.2.1 | Please provide unique ref number of the ISO 9001 ceriticate:
5.3 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
5.3.1 | Please provide unique ref number of the ISO 13485 ceriticate:
5.4 | Model quoted for must be EC certified. Attach a copy of certification
5.4.1 | Please provide unique ref number of the EC ceriticate:
5.5 | The unit must comply with ISO 5835 standard, attached certification
5.5.1 | Please provide unique ref number of the ISO 5835 ceriticate:
6 | COMPREHENSIVE MAINTENANCE
6.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) must be included in the bid price
6.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
6.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
6.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
6.3 | The bidder to provide the OEM recommended service interval of the unit:
6.4 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
7 | General
7.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
7.2 | The unit on offer must be the latest technology.
Drill, Orthopaedic, Battery Operated, Large bones
RT2-07-04a | Basic hand piece with attachment coupling that accept variety of attachements
RT2-07-04b | Jacob Chuck and a key
RT2-07-04c | AO/ASIF quick coupling
RT2-07-04d | Quick coupling for Kirschner wires
RT2-07-04e | Quick coupling for drill bits
RT2-07-04f | Large quick coupling
RT2-07-04g | Reduction drive unit
RT2-07-04h | Drill chuck keyless
RT2-07-04i | Instrument tray where all accessories can be placed for cleaning
RT2-07-04j | Oscillating/Sagittal saw attachement with different blades which are:
RT2-07-04k | Sterilization case
RT2-07-04l | Hudson Modified Trinkle
RT2-07-04m | Li-Ion battery x2
RT2-07-04n | Battery charger
RT2-07-04o | Battery trabsfer shield
Bidder to list any other Consumables or Accessories (Not listed above) for the model on offer below

## Tourniquet
RT2-07-05 | TOURNIQUET SYSTEM
1 | DESCRIPTION
1.1 | Electric Tourniquet for occluding blood flow and obtaining a bloodless field during limb surgery.
2 | Product Offered:
2.1 | Make:
2.2 | Model:
2.3 | Initial Manufacturing year of Model:
2.4 | Radiation Control License number:
3 | TECHNICAL SPECIFICATION
3.1 | Cuff pressure ranging from 0 - 650 mmHg
3.2 | Cuffs should be washable, autoclavable with Velcro, single and double cuff for the dual port system.
3.3 | Device shall be a dual port system
3.4 | Deflate cuff without disconnecting cuff from machine
3.5 | Display:
3.5.1 | Set Pressure
3.5.2 | Measured Pressure
3.5.3 | Time Lapsed
3.5.4 | Count Down
3.6 | afety Features
3.6.1 | Self-check upon switching the unit on
3.6.2 | Retain cuff pressure (not released) when power fails
3.6.3 | The unit shall have a quick release/disconnecting facility of the cuff.
3.6.4 | The offered unit shall have Automatic Compensation in case of pressure loss.
3.7 | The unit should include visual and audible alarms:
3.7.1 | Exceeded time
3.7.2 | Internal failure
3.7.3 | Low battery
3.7.4 | Leakage
3.7.5 | High and low pressure
3.8 | Battery
3.8.1 | Intergrated rechargeable battery back-up.
3.8.2 | Back-up battery must last at least 6 hours
3.9 | Mobile Stand
3.9.1 | The offered unit shall be complete with roll stand and an accessory basket.
3.9.2 | The trolley shall have a minimum of 4 casters with central locking mechanism.
4 | STANDARD ACCESSORIES
4.1 | The trolley should have antistatic castors with at least 2 equipped with locking system/brake.
4.2 | Paediatric cuff
4.3 | Adult cuff
4.4 | Large Adult
4.5 | Reusable, Esmarch Bandages
5 | Conformity Compliance (Please attach certificate)
5.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
5.2 | OEM must comply and certified on ISO 9001 quality standards, attach proof of compliance
5.2.1 | Please provide unique ref number of the ISO 9001 ceriticate:
5.3 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
5.3.1 | Please provide unique ref number of the ISO 13485 ceriticate:
5.4 | Model quoted for must be EC certified. Attach a copy of certification
5.4.1 | Please provide unique ref number of the EC ceriticate:
5.5 | Hazardous Substance Act: If this type of equipment / apparatus appears on the schedule of Hazardous substances issued by the Directorate: Health Technology of the Department of Health, a license in terms of the Act on Hazardous Substances (Act 15/1973) must be submitted with the bid / quotation documents. The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.6 | All electrical/electronic medical equipment must be licensed by Radiation Control, where applicable a copy of the license must be submitted The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.6.1 | Radiation Control License number:
5.7 | The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade (rubber to with 2 screws) type and it must be a minimum length of three (3) meters. NB the mains cable of the unit tendered for must be SABS colour coded.
5.8 | The equipment quoted for must be protected against electromagnetic interference IEC 60601-1-2
6 | COMPREHENSIVE MAINTENANCE
6.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) must be included in the bid price
6.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
6.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
6.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
6.3 | The bidder to provide the OEM recommended service interval of the unit:
6.4 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
7 | General
7.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
7.2 | The unit on offer must be the latest technology.
TOURNIQUET SYSTEM -ACCESSORIES and CONSUMABLES
RT2-07-05a | Paediatric cuff
RT2-07-05b | Adult cuff
RT2-07-05c | Large Adult
RT2-07-05d | Reusable, Esmarch Bandages
Bidder to list any other Consumables or Accessories (Not listed above) for the model on offer below

## Plaster saw
RT2-07-06 | ELECTRICAL PLASTER OF PARIS CUTTER WITH ACCESSORIES
1 | SCOPE
1.1 | This specification establishes the requirements for the Supply and delivery of electrical cast cutter and accessories as a set as per specification
2 | Product Offered:
2.1 | Make:
2.2 | Model:
2.3 | Initial Manufacturing year of Model:
3 | SPECIFICATION
3.1 | Each set shall consist of the following components:
3.1.1 | Electrical plaster cutter as specified
3.1.2 | Stainless Steel cast scissors (manual cast cutter)
3.1.3 | Small cast spreader (stainless steel)
3.1.4 | Large cast spreader (stainless steel)
3.2 | The system shall have a dual speed control
3.3 | The unit shall be versatile and may be used with battery or plugged directly into the wall. Bidder to state if the unit offered is capable to be used as required.
3.4 | A dedicated five star mobile stand which provides stability and easy transportation.
3.5 | The unit shall have an impact resistant housing.
3.6 | The unit offered shall be of ergonomic design to ensure comfort and control while cutting.
3.7 | The unit offered shall have a noise reducing mechanism.
3.8 | The unit must be durable, long life expectancy
3.9 | Technical Detailas
3.1 | SIZE ;±10.0 in. (254mm) in length, 3.0 in. (76mm) in height.
3.2 | WEIGHT ; ±1.3Kg
3.3 | SPEED – ±High 15,900 CPM, Low 10,700 CPM
3.4 | NOISE – ±76dB (low speed), and 82dB (high speed)
3.7 | POWER – 230v 50-60Hz 0.75A
4 | ACCESSORIES
4.1 | POP blades (Pk of 5)
4.2 | Fibreglass material blades (Pk of 5)
4.3 | Stainless Steel cast scissors (manual cast cutter)
4.4 | Small cast spreader (stainless steel)
4.5 | Large cast spreader (stainless steel)
5 | Conformity Compliance (Please attach certificate)
5.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
5.2 | OEM must comply and certified on ISO 9001 quality standards, attach proof of compliance
5.2.1 | Please provide unique ref number of the ISO 9001 ceriticate:
5.3 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
5.3.1 | Please provide unique ref number of the ISO 13485 ceriticate:
5.4 | Model quoted for must be EC certified. Attach a copy of certification
5.4.1 | Please provide unique ref number of the EC ceriticate:
5.5 | The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade (rubber to with 2 screws) type and it must be a minimum length of three (3) meters. NB the mains cable of the unit tendered for must be SABS colour coded.
5.6 | The equipment quoted for must be protected against electromagnetic interference IEC 60601-1-2
6 | COMPREHENSIVE MAINTENANCE
6.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) must be included in the bid price
6.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
7 | General
7.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
7.2 | The unit on offer must be the latest technology.
RT2-07-06 | ELECTRICAL PLASTER OF PARIS CUTTER WITH ACCESSORIES
RT2-07-06a | POP blades (Pk of 5)
RT2-07-06b | Fibreglass material blades (Pk of 5)
RT2-07-06c | Stainless Steel cast scissors (manual cast cutter)
RT2-07-06d | Small cast spreader (stainless steel)
RT2-07-06e | Large cast spreader (stainless steel)
Bidder to list any other Consumables or Accessories (Not listed above) for the model on offer below

## Plaster saw with extractor
RT2-07-07 | ELECTRICAL PLASTER OF PARIS CUTTER WITH VACUUM AND ACCESSORIES SET
1 | SCOPE
1.1 | This specification establishes the requirements for the Supply and delivery of electrical cast cutter plus vacuum and accessories as a set as per specification
2 | Product Offered:
2.1 | Make:
2.2 | Model:
2.3 | Initial Manufacturing year of Model:
3 | SPECIFICATION
3.1 | Each set shall consist of the following components:
3.1.1 | Electrical plaster cutter as specified
3.1.2 | Mobile Vacuum Extractor to collect cast dust
3.1.3 | Stainless Steel cast scissors (manual cast cutter)
3.1.4 | Small cast spreader (stainless steel)
3.1.5 | Large cast spreader (stainless steel)
3.3 | The system shall have a dual speed control
3.4 | The offer shall include a vacuum to collect and filter any cast dust while cutting.
3.5 | The unit shall be versatile and may be used with vacuum or plugged directly into the wall. Bidder to state if the unit offered is capable to be used as required.
3.6 | A dedicated five star mobile stand which provides stability and easy transportation.
3.7 | The unit shall have an impact resistant housing.
3.8 | The unit offered shall be of ergonomic design to ensure comfort and control while cutting.
3.9 | The unit offered shall have a noise reducing mechanism.
3.10 | The unit shall have an integral vacuum housing of at least 2m length. Bidder to state the length of the vacuum housing of the product offered.
3.11 | The cast vacuum canister shall have a quick and easy method of emptying and cleaning of the dust collection filter.
3.12 | The hose mount shall allow for easy movement of the vacuum hose.
3.13 | The vacuum shall turn on automatically when the cast cutter is activated.
3.14 | Vacuum Extractor must have a Cast Cutter holder or flat surface to place it safely
3.15 | Technical Details
3.15.1 | SIZE – ±10.0 in. (254mm) in length, 3.0 in. (76mm) in height.
3.15.2 | WEIGHT – ±2.8 lbs. (1.3Kg)
3.15.3 | SPEED – ±High 15,900 CPM, Low 10,700 CPM
3.15.4 | NOISE – ±76dB (low speed), and 82dB (high speed)
3.15.5 | POWER – 230v 50-60Hz 0.75A
3.15.6 | VACUUM: 220-240V Vacuum
4 | STANDARD ACCESSORIES
4.1 | POP blades (Pk of 5)
4.2 | Fibreglass material blades (Pk of 5)
4.3 | Re-usable vacuum filters x10
4.4 | vacuum bags and liners (10 of each)
4.5 | Stainless Steel cast scissors (manual cast cutter)
4.6 | Small cast spreader (stainless steel)
4.7 | Large cast spreader (stainless steel)
5 | Conformity Compliance (Please attach certificate)
5.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
5.2 | OEM must comply and certified on ISO 9001 quality standards, attach proof of compliance
5.2.1 | Please provide unique ref number of the ISO 9001 ceriticate:
5.3 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
5.3.1 | Please provide unique ref number of the ISO 13485 ceriticate:
5.4 | Model quoted for must be EC certified. Attach a copy of certification
5.4.1 | Please provide unique ref number of the EC ceriticate:
5.5 | The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade (rubber to with 2 screws) type and it must be a minimum length of three (3) meters. NB the mains cable of the unit tendered for must be SABS colour coded.
5.6 | The equipment quoted for must be protected against electromagnetic interference IEC 60601-1-2
6 | COMPREHENSIVE MAINTENANCE
6.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) must be included in the bid price
6.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
7 | General
7.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
7.2 | The unit on offer must be the latest technology.
ELECTRICAL PLASTER OF PARIS CUTTER WITH VACUUM AND ACCESSORIES SET
RT2-07-07a | POP blades (Pk of 5)
RT2-07-07b | Fibreglass material blades (Pk of 5)
RT2-07-07c | Re-usable vacuum filters x10
RT2-07-07d | vacuum bags and liners (10 of each)
RT2-07-07e | Stainless Steel cast scissors (manual cast cutter)
RT2-07-07f | Small cast spreader (stainless steel)
RT2-07-07g | Large cast spreader (stainless steel)
Bidder to list any other Consumables or Accessories (Not listed above) for the model on offer below

## Dermatome, Cordless
RT2-07-08 | DERMATOME, CORDLESS
1 | SCOPE
1.1 | This specification establishes the requirements for the supply, delivery, installation, demonstration and commissioning of a dermatome for use in operating theatres.
2 | Product Offered:
2.1 | Make:
2.2 | Model:
2.3 | Initial Manufacturing year of Model:
3 | SPECIFICATION
3.1 | The unit supplied shall be a Dermatome with a cutting width of approximately 70mm.
3.2 | It shall be possible to select cutting widths of up to the abovementioned maximum width.
3.3 | It shall be possible to select different cutting depths (Thickness). The following thickness settings shall be possible:
3.3.1 | 0.08mm
3.3.2 | 0.13mm
3.3.3 | 0.18mm
3.3.4 | 0.23mm
3.3.5 | 0.28mm
3.4 | The unit shall be compact and lightweight, and easy to handle.
3.5 | It shall be easy to adjust different selections i.e. thickness and width.
3.6 | The unit shall be a proven design, and be robust for use in a theatre setting.
3.7 | Bidders shall prove that the unit is very accurate when different settings are selected.
3.8 | It shall be possible to sterilise the full head and handle of the offered unit.
3.10 | The offer shall include a sterilisation case to enable sterilising of the unit.
4 | Standard Accessories and Consumables
4.1 | 25mm (Pk of 10)
4.2 | 50mm (Pk of 10)
4.3 | 70mm (Pk of 10)
4.4 | Sterilizing Case
4.5 | Battery Charger (2-slots)
5 | Conformity Compliance (Please attach certificate)
5.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
5.2 | OEM must comply and certified on ISO 9001 quality standards, attach proof of compliance
5.2.1 | Please provide unique ref number of the ISO 9001 ceriticate:
5.3 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
5.3.1 | Please provide unique ref number of the ISO 13485 ceriticate:
5.4 | Model quoted for must be EC certified. Attach a copy of certification
5.4.1 | Please provide unique ref number of the EC ceriticate:
6 | COMPREHENSIVE MAINTENANCE
6.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) must be included in the bid price
6.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
7 | General
7.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
7.2 | The unit on offer must be the latest technology.
RT2-07-08 | DERMATOME, CORDLESS
RT2-07-08a | 25mm (Pk of 10)
RT2-07-08b | 50mm (Pk of 10)
RT2-07-08c | 70mm (Pk of 10)
RT2-07-08d | Sterilizing Case
RT2-07-08e | Battery Charger (2-slots)
Bidder to list any other Consumables or Accessories (Not listed above) for the model on offer below

## Dermatome, Pneumatic
RT2-07-09 | DERMATOME, PNEUMATIC
1 | SCOPE
1.1 | This specification establishes the requirements for the supply, delivery, installation, demonstration and commissioning of a dermatome for use in operating theatres.
2 | Product Offered:
2.1 | Make:
2.2 | Model:
2.3 | Initial Manufacturing year of Model:
3 | SPECIFICATION
3.1 | The unit supplied shall be a Dermatome with a cutting width of approximately 70mm.
3.2 | It shall be possible to select cutting widths of up to the abovementioned maximum width.
3.3 | It shall be possible to select different cutting depths (Thickness). The following thickness settings shall be possible:
3.3.1 | 0.08mm
3.3.2 | 0.13mm
3.3.3 | 0.18mm
3.3.4 | 0.23mm
3.3.5 | 0.28mm
3.4 | The unit shall be compact and lightweight, and easy to handle.
3.5 | It shall be easy to adjust different selections i.e. thickness and width.
3.6 | The unit shall be a proven design, and be robust for use in a theatre setting.
3.7 | Bidders shall prove that the unit is very accurate when different settings are selected.
3.8 | It shall be possible to sterilise the full head and handle of the offered unit.
3.10 | The offer shall include a sterilisation case to enable sterilising of the unit.
4 | Standard Accessories and Consumables
4.1 | 25mm (Pk of 10)
4.2 | 50mm (Pk of 10)
4.3 | 70mm (Pk of 10)
4.4 | Sterilizing Case
4.5 | 3m Hose
5 | Conformity Compliance (Please attach certificate)
5.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
5.2 | OEM must comply and certified on ISO 9001 quality standards, attach proof of compliance
5.2.1 | Please provide unique ref number of the ISO 9001 ceriticate:
5.3 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
5.3.1 | Please provide unique ref number of the ISO 13485 ceriticate:
5.4 | Model quoted for must be EC certified. Attach a copy of certification
5.4.1 | Please provide unique ref number of the EC ceriticate:
6 | COMPREHENSIVE MAINTENANCE
6.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) must be included in the bid price
6.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
7 | General
7.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
7.2 | The unit on offer must be the latest technology.
RT2-07-09 | DERMATOME, PNEUMATIC
RT2-07-09a | 25mm (Pk of 10)
RT2-07-09b | 50mm (Pk of 10)
RT2-07-09c | 70mm (Pk of 10)
RT2-07-09d | Sterilizing Case
RT2-07-09e | 3m Hose
Bidder to list any other Consumables or Accessories (Not listed above) for the model on offer below

## Skin Mesher
RT2-07-10 | SKIN MESHER
1 | SCOPE
1.1 | This specification establishes the requirements for the supply, delivery, installation, demonstration and commissioning of a skin mesher for use in operating theatres.
2 | Product Offered:
2.1 | Make:
2.2 | Model:
2.3 | Initial Manufacturing year of Model:
3 | SPECIFICATION FOR MESHER
3.1 | The unit shall be a skin mesher used for the expansion of harvested skin.
3.2 | It shall be a mechanical unit designed for simple operation, utilising disposable carriers to carry the skin through the mesher operation.
3.3 | The unit shall be manufactured from high quality, durable stainless steel.
3.4 | It shall be a simple operation to change the cutters. Four (4) different cutters and four (4) sets of meshers shall be included in the offer
3.5 | The unit must have the following expansion ratios:
3.5.1 | 1.5:1
3.5.2 | 2:1
3.5.3 | 3:1
3.5.4 | 4:1
3.5 | The skin shall be guided through the mesher by a ratchet system, ensuring smooth transport of the skin graft carrier.
3.6 | The mesher instrument shall consist of a ratchet handle, feed roller, cutting roller and a guidance platform. The harvested skin shall be guided along the platform by means of the carriers and the ratchet handle.
4 | Standard Accessories
4.1 | Small carriers (Pack of 10)
4.2 | Large carriers (Pack of 10)
5 | Conformity Compliance (Please attach certificate)
5.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
5.2 | OEM must comply and certified on ISO 9001 quality standards, attach proof of compliance
5.2.1 | Please provide unique ref number of the ISO 9001 ceriticate:
5.3 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
5.3.1 | Please provide unique ref number of the ISO 13485 ceriticate:
5.4 | Model quoted for must be EC certified. Attach a copy of certification
5.4.1 | Please provide unique ref number of the EC ceriticate:
6 | COMPREHENSIVE MAINTENANCE
6.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) must be included in the bid price
6.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
7 | General
7.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
7.2 | The unit on offer must be the latest technology.
SKIN MESHER
RT2-07-10a | Small carriers (Pack of 10)
RT2-07-10b | Large carriers (Pack of 10)
Bidder to list any other Consumables or Accessories (Not listed above) for the model on offer below

## Category-8 Ophthalmo
OPHTHALMOLOGY EQUIPMENT
# | Item number | Item Description
1 | RT2-08-01 | A&B Scanner Combo Unit
2 | RT2-08-02 | Auto Ref-Keratometer
3 | RT2-08-03 | Laser Unit, Argon
8 | RT2-08-08 | Slit-Lamp, Basic
9 | RT2-08-09 | Slit-Lamp, Advanced
10 | RT2-08-10 | Tonometer, Hand-held
11 | RT2-08-11 | Visual Field Analyser
12 | RT2-08-12 | Vitrectomy Unit
13 | RT2-08-13 | Microscope, Ophthamoscope, Basic
14 | RT2-08-14 | Microscope, Ophthamoscope, Advanced

## A&B Scanner Combo
RT2-08-01 | A& B Scanner with Pachymetry
1 | Description
1.1 | The specification establishes the requirements, supply, delivery, end user training, demonstration, commissioning and installation of A&B Scanner with Pachymetry
2 | Product Offered:
2.1 | Make:
2.2 | Model:
2.3 | Initial Manufacturing year of Model:
2.4 | Radiation Control License number:
3 | Technical Specification
3.1 | The unit must be able to do the following functions:
3.1.1 | Biometry
3.1.2 | IOL calculation
3.1.3 | Standardised A-diagnostic with S-mode
3.1.4 | Pachymetry with central and map function
3.1.5 | Internal database
3.2 | The unit must include the following :
3.2.1 | New generation of annular array probes
3.2.2 | Probe holders(preferable on the back of the unit)
3.2.3 | High resolution touch-screen operation: minimum of 10 inches, colour touch screen
3.2.4 | Data communication must be possible via USB or LAN
3.2.5 | Modular configurable with Pachymetry
3.2.6 | Foot switch
3.3 | Include the following probes
3.3.1 | Internal database UBM 40MHz B-probe
3.3.2 | Biometer A-scan 10 MHz A-diagnostic probe
3.3.3 | 10 MHz 2-ring array B-probe
3.3.4 | Pachymetry 20 MHz solid state probe tip 1.5 mm with an angle of at least 45°
3.3.5 | Biometry 10 MHz solid state probe with built in red fixation LED
3.3.6 | A-scan diagnosis 10 MHz solid state probe
3.4 | 10 MHz B-probe , 2-Ring sector scanning
3.4.1 | State the range
3.4.2 | Frame rate basic mode
3.4.3 | Maximum number of pages in a movie
3.4.4 | Image display range Standard
3.4.5 | Wide
3.5 | Axial & lateral resolution / accuracy
3.6 | Colour scale
3.7 | The B-scan must be able to do zooming,report/export image,sampling rate,scanning angle and adjustable gamma
3.8 | Biometry 10 Mhz solid state probe with built in red fixation LED
3.8.1 | State the range
3.8.2 | Measurement range
3.8.3 | Measurement resolution / accuracy
3.8.4 | IOL power calculation
3.9 | A-scan diagnosis 10 MHz solid state probe
3.9.1 | Measurement range
3.9.2 | Measurement resolution / accuracy
3.9.3 | Capture:automatic and manual
3.9.4 | Custom eye type settings
3.10 | Pachymetry 20 MHz solid state probe tip 1.5 mm with an angle of at least 45°
3.10.1 | State the range
3.10.2 | Measurement range
3.10.3 | Measurement resolution / accuracy
3.10.4 | Percent bias
3.10.5 | Plus minus bias
3.10.6 | Velocity range / standard
3.10.7 | IOP adjustment as well as sampling rate
3.11 | The following connectors/ports must be available
3.11.1 | B-scan for 10 or 40 MHz,
3.11.2 | A-scan / pachymeter
3.11.3 | Foot switch
3.11.4 | LAN, USB 2 x standard 2 x OA-2000/PicBrigde
3.12 | It must be possible on the unit to save the data on the internal database or on the PC. Intuitive(Spontaneous) software guides all measurement results, which can be printed out through built-in printer or external printer.
3.13 | If the unit does not have a built-in printer , the bidder must provide an external printer which must be included in the bid price.
4 | Standard Accessories
4.1 | Internal database UBM 40MHz B-probe
4.2 | Biometer A-scan 10 MHz A-diagnostic probe
4.3 | 10 MHz 2-ring array B-probe
4.4 | Pachymetry 20 MHz solid state probe tip 1.5 mm with an angle of at least 45°
4.5 | Biometry 10 MHz solid state probe with built in red fixation LED
4.6 | A-scan diagnosis 10 MHz solid state probe
4.7 | Printing paper (Pk of 10)
5 | Conformity Compliance (Please attach certificate)
5.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
5.2 | OEM must comply and certified on ISO 9001 quality standards, attach proof of compliance
5.2.1 | Please provide unique ref number of the ISO 9001 ceriticate:
5.3 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
5.3.1 | Please provide unique ref number of the ISO 13485 ceriticate:
5.4 | Model quoted for must be EC certified. Attach a copy of certification
5.4.1 | Please provide unique ref number of the EC ceriticate:
5.5 | Hazardous Substance Act: If this type of equipment / apparatus appears on the schedule of Hazardous substances issued by the Directorate: Health Technology of the Department of Health, a license in terms of the Act on Hazardous Substances (Act 15/1973) must be submitted with the bid / quotation documents. The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.6 | All electrical/electronic medical equipment must be licensed by Radiation Control, where applicable a copy of the license must be submitted The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.6.1 | Radiation Control License number:
5.7 | The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade (rubber to with 2 screws) type and it must be a minimum length of three (3) meters. NB the mains cable of the unit tendered for must be SABS colour coded.
5.8 | The equipment quoted for must be protected against electromagnetic interference IEC 60601-1-2
6 | COMPREHENSIVE MAINTENANCE
6.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) must be included in the bid price
6.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
6.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
6.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
6.3 | The bidder to provide the OEM recommended service interval of the unit:
6.4 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
7 | General
7.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
7.2 | The unit on offer must be the latest technology.
A& B Scanner with Pachymetry- Accessories and Consumables
RT2-08-01a | Internal database UBM 40MHz B-probe
RT2-08-01b | Biometer A-scan 10 MHz A-diagnostic probe
RT2-08-01c | 10 MHz 2-ring array B-probe
RT2-08-01d | Pachymetry 20 MHz solid state probe tip 1.5 mm with an angle of at least 45°
RT2-08-01d | Biometry 10 MHz solid state probe with built in red fixation LED
RT2-08-01e | A-scan diagnosis 10 MHz solid state probe
RT2-08-01f | Printing paper (Pk of 10)
Bidder to list any other Consumables or Accessories (Not listed above) for the model on offer below

## Auto Ref-Keratometer
RT2-08-02 | OPHTHALMIC AUTO REF-KERATOMETER
1 | DESCRIPTION
1.1 | An instrument for measuring the curvature of the anterior corneal surface
2 | Product Offered:
2.1 | Make:
2.2 | Model:
2.3 | Initial Manufacturing year of Model:
2.4 | Radiation Control License number:
3 | SPECIFICATIONS
3.1 | The system offered must have the following test modes:
3.1.1 | Refraction
3.1.2 | Keratometry.
3.1.3 | Refraction and Keratometry
3.2 | The measuring range (Diopter-D) for HYPEROPIA must be at least 0D to + 25D in a minimum increments of 0,01D steps.
3.3 | The measuring range (Diopter-D) for MYOPIA must be at least 0D to -30D in a minimum increments of 0,01D steps.
3.4 | The measuring range (Diopter-D) for ASTIGMATISM must be at least 0D to ±12D in a minimum increments of 0,01D steps.
3.5 | The AXIS measuring range must be at least 1° to 180° in steps of one deg (1°) .
3.6 | Testing of the naked eye, the pseudophakic eye and over refraction with contact lenses must be available and selectable.
3.7 | The vertex distance for the refraction must be displayed.
3.8 | The pupillary distance (PD) must be calculated and shown on the printout.
3.9 | PD measurement adjustment with 1mm step must be available. Bidder to specify the range
3.10 | KERATOMETRY must meet the following specifications:
3.10.1 | The CORNEAL CURVATURE RADIUS measuring range must be adjustable with minimum 0,01mm steps. Bidder to specify the range
3.10.2 | The CORNEAL REFRACTION measuring range must be adjustable with minimum 0,01D steps. Bidder to specify the range
3.10.3 | The CORNEAL ASTIGMATISM ANGLE measuring range must be at 1° to 180° in steps of 1°.
3.10.4 | The readout display must be selectable between MILLIMETER and DIOPTER.
3.10.5 | An TFT colour display of at least 16cm must be incorporated in the system to provide ease of operation and accurate alignment.
3.10.6 | Reading must be displayed on the video monitor for all three modes as follows: REFRACTION; KERATOMETRY; REFRACTION and KERATOMETRY.
3.10.7 | Approximately 10 (ten) or more readings of both REFRACTION and KERATOMETRY must be retained in memory for averaging.
3.11 | Built-in Printer
3.11.1 | PRINTOUT must be obtained at the end of the procedure and the printout must provide the following information:
3.12 | A printout for refraction, when selected, must provide the following:
3.12.1 | Vertex distance.
3.12.2 | PD.
3.12.3 | The readings taken for each eye.
3.12.4 | The average of the readings for each eye.
3.13 | A printout for keratometry, when selected, must provide the following
3.13.1 | Corneal curvature radius
3.13.2 | Corneal refraction.
3.13.3 | The average of the keratometry values.
3.13.4 | Corneal astigmatism value in D and axis.
3.14 | A following modes and printout, when selected, must provide the following:
3.14.1 | Pupil size from 1mm – 10mm in 0,1mm increments.
3.14.2 | Corneal size from 1mm – 14mm in 0,1mm increments.
3.14.3 | The system offered must have an automatic shut down / off facility to conserve energy and the life of the system components when not in use.
3.14.4 | The system must be controlled by means of a joystick or any other reliable latest technology.
3.14.5 | The system must include a motorised chin support device capable of vertical adjustment for the patients head.
3.14.6 | The system must produce quality results under normal room lighting conditions.
3.14.7 | The system must include a mobile sturdy table with lockable castors and with an easily adjustable electrical height mechanism, at no extra cost to the final bid price.
4 | STANDARD ACCESSORIES (Any other cables or connectors required for the functionality of the unit must be provided)
4.1 | Printing paper (pack of 10)
4.2 | Dust cover
5 | Conformity Compliance (Please attach certificate)
5.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
5.2 | OEM must comply and certified on ISO 9001 quality standards, attach proof of compliance
5.2.1 | Please provide unique ref number of the ISO 9001 ceriticate:
5.3 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
5.3.1 | Please provide unique ref number of the ISO 13485 ceriticate:
5.4 | Model quoted for must be EC certified. Attach a copy of certification
5.4.1 | Please provide unique ref number of the EC ceriticate:
5.5 | Hazardous Substance Act: If this type of equipment / apparatus appears on the schedule of Hazardous substances issued by the Directorate: Health Technology of the Department of Health, a license in terms of the Act on Hazardous Substances (Act 15/1973) must be submitted with the bid / quotation documents. The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.6 | All electrical/electronic medical equipment must be licensed by Radiation Control, where applicable a copy of the license must be submitted The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.6.1 | Radiation Control License number:
5.7 | The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade (rubber to with 2 screws) type and it must be a minimum length of three (3) meters. NB the mains cable of the unit tendered for must be SABS colour coded.
5.8 | The equipment quoted for must be protected against electromagnetic interference IEC 60601-1-2
6 | COMPREHENSIVE MAINTENANCE
6.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) must be included in the bid price
6.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
6.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
6.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
6.3 | The bidder to provide the OEM recommended service interval of the unit:
6.4 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
7 | General
7.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
7.2 | The unit on offer must be the latest technology.
OPHTHALMIC AUTO REF-KERATOMETER -Consumables
RT2-08-02a | Printing paper (pack of 10)
Bidder to list any other Consumables or Accessories (Not listed above) for the model on offer below

## Laser Unit, Argon
RT2-08-03 | ARGON LASER-MULTISPOT WITH SLIT LAMP AND INDIRECT OPTHALMOSCOPE
1 | Description
2 | Product Offered:
2.1 | Make:
2.2 | Model:
2.3 | Initial Manufacturing year of Model:
2.4 | Radiation Control License number:
3 | Technical Specification
3.1 | The specification establishes the requirement of supply, delivery, end user training, demonstration, commissioning and installation of an Argon Laser-Multisport with Slit lamp and Indirect ophthalmoscope which must be supplied as a system as per the under mentioned Clauses.
3.2 | The unit is intended to be used for the following:
3.2.1 | In the Posterior segment to perform retinal photocoagulation, Panretinal photocoagulation, focal photocoagulation, and grid photocoagulation for vascular and structural abnormalities of the retina and choroid.
3.2.2 | In the treatment of ocular pathology in the anterior segment including the Iridotomy and Trabeculoplasty.
3.3 | The design should be such that the multi-spot approach uses shorter pulse duration of individual pulses to allow for manageable time for multi-spot delivery.
3.4 | The unit must be designed such that the Pan retinal photocoagulation (PRP) give rise to lower levels of discomfort and distress to the patient compared to the single spot PRP treatment.
3.5 | For safety of the patient the unit should be designed such that the laser is safe, effective and other several advantages associated with shorter exposure including reduced pain, reduced inner retinal damage and reduced scarring.
3.6 | The device must perform single session treatment which help to induce the lower levels of anxiety, headache pain and photophobia.
3.7 | Units that offers conventional laser will not be accepted as these require multiple sessions for treatment.
3.8 | The device must provide complete control over all customary settings such as pulse duration, laser power, pattern selection and pattern size.
3.9 | All parameters must be clearly marked with exceptionally easy access to the standby/ ready-modes.
3.10 | It must be possible to control aspects such as pattern size, pattern orientation and laser power without removing your view from the patient's fundus.
3.11 | It must be possible to display different types of patterns such as:
3.11.1 | single spot;
3.11.2 | array;
3.11.3 | triple ring;
3.11.4 | arc;
3.11.5 | line circle;
3.11.6 | circle;
3.11.7 | macular grid;
3.11.8 | grid;
3.11.9 | and hexagon
3.12 | The controls and displays must be touchscreen (that allows easy selection of arrays, lines and rectangles) user interface and integrated slit lamp.
3.13 | The system should be equipped with a three position key switch that enables laser and controls, as well as an emergency off switch.
3.14 | A system must have a 3D remote control that allows physician control over lasting positioning, power settings, adjust and rotate laser patterns.
3.15 | A footswitch activator must be provided with a shrouded housing
3.16 | For patient that are unable to sit on a slit lamp, a laser Indirect Ophthalmoscope must be provided. This must be included in the total bid price.
3.17 | For the performance of the argon laser the following features should be possible:
3.17.1 | Power: 0-2000mW
3.17.2 | Wavelength: 577nm
3.17.3 | Pulse duration (single spot): 5-1000ms
3.17.4 | Aim beam: 620nm diode 650nm adjustable
3.17.5 | Delivery spot size: 50, 100, 200, 400um etc.
3.17.6 | Exposure time: 2milli-seconds-1000milli-seconds
3.17.7 | Cooling system: Thermoelectric or Air cooled
3.17.8 | Pulse counter: 0-99999 with automatic reset available
3.17.9 | Repetition rate: Available with single spot pattern
3.17.10 | 5 inch or above touchscreen monitor display
4 | Standard Accessories and Cosumables
Dust cover
5 | Conformity Compliance (Please attach certificate)
5.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
5.2 | OEM must comply and certified on ISO 9001 quality standards, attach proof of compliance
5.2.1 | Please provide unique ref number of the ISO 9001 ceriticate:
5.3 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
5.3.1 | Please provide unique ref number of the ISO 13485 ceriticate:
5.4 | Model quoted for must be EC certified. Attach a copy of certification
5.4.1 | Please provide unique ref number of the EC ceriticate:
5.5 | Hazardous Substance Act: If this type of equipment / apparatus appears on the schedule of Hazardous substances issued by the Directorate: Health Technology of the Department of Health, a license in terms of the Act on Hazardous Substances (Act 15/1973) must be submitted with the bid / quotation documents. The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.6 | All electrical/electronic medical equipment must be licensed by Radiation Control, where applicable a copy of the license must be submitted The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.6.1 | Radiation Control License number:
5.7 | The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade (rubber to with 2 screws) type and it must be a minimum length of three (3) meters. NB the mains cable of the unit tendered for must be SABS colour coded.
5.8 | The equipment quoted for must be protected against electromagnetic interference IEC 60601-1-2
6 | COMPREHENSIVE MAINTENANCE
6.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) must be included in the bid price
6.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
6.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
6.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
6.3 | The bidder to provide the OEM recommended service interval of the unit:
6.4 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
7 | General
7.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
7.2 | The unit on offer must be the latest technology.

## Slit-Lamp, Basic
RT2-08-08 | OPHTHALMIC SLIT LAMP
1 | DESCRIPTION
1.1 | The binocular slit-lamp examination provides a stereoscopic magnified view of the eye structures in detail, enabling anatomical diagnoses to be made for a variety of eye conditions
2 | Product Offered:
2.1 | Make:
2.2 | Model:
2.3 | Initial Manufacturing year of Model:
3 | SPECIFICATIONS
2.1 | The specification requires a complete slit lamp installed on stable table with Halogen illumination
2.2 | The slit lamp must have the following features:
2.2.1 | Range for Magnification: from 6x to 40x or above
2.2.2 | Field of view:5mm to 40mm
2.2.3 | Rotation of slit lamp: 180 degrees with reference scale
2.2.4 | Swivel range of still projector +-180
2.2.5 | Angle of incidence: various steps
2.2.6 | Filters:Blue ; Green and White
2.2.7 | Light source: Halogen/LED. Bidder to state the light source
2.2.8 | Travel of instrument base: Vertical; lateral and axial
2.2.9 | Brighteness: continous adjustable
2.2.10 | The Slit lamp be supllied with applanation Tonometer.
2.2.11 | Provision for mounting the applanation Tonometer must be provided
2.2.12 | The slit lamp must be provided with chinrest
3 | STANDARD ACCESSORIES
3.1 | dust cover
5 | Conformity Compliance (Please attach certificate)
5.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
5.2 | OEM must comply and certified on ISO 9001 quality standards, attach proof of compliance
5.2.1 | Please provide unique ref number of the ISO 9001 ceriticate:
5.3 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
5.3.1 | Please provide unique ref number of the ISO 13485 ceriticate:
5.4 | Model quoted for must be EC certified. Attach a copy of certification
5.4.1 | Please provide unique ref number of the EC ceriticate:
6 | COMPREHENSIVE MAINTENANCE
6.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) must be included in the bid price
6.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
7 | General
7.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
7.2 | The unit on offer must be the latest technology.
OPHTHALMIC SLIT LAMP
RT2-08-08a | LED replacement bulb
RT2-08-08b | Halogen replacement bulb
Bidder to list any other Consumables or Accessories (Not listed above) for the model on offer below

## Slit-Lamp, Advanced
RT2-08-09 | OPHTHALMIC SLIT LAMP, intergrated imaging module
1 | DESCRIPTION
1.1 | The binocular slit-lamp examination provides a stereoscopic magnified view of the eye structures in detail, enabling anatomical diagnoses to be made for a variety of eye conditions
2 | Product Offered:
2.1 | Make:
2.2 | Model:
2.3 | Initial Manufacturing year of Model:
3 | SPECIFICATIONS
3.1 | The specification requires a complete slit lamp installed on electronic controlled table with LED illumination and a monitor
3.2 | The slit lamp must have the following features:
3.2.1 | Range for Magnification: from 6x to 40x or above
3.2.2 | Field of view:5mm to 40mm
3.2.3 | Rotation of slit lamp: 180 degrees with reference scale
3.2.4 | Swivel range of still projector +-180
3.2.5 | Angle of incidence: various steps
3.2.6 | Filters:Blue ; Green and White
3.2.7 | Light source: LED
3.2.8 | Travel of instrument base: Vertical; lateral and axial
3.2.9 | Brighteness: continous adjustable
3.2.10 | The Slit lamp be supllied with applanation Tonometer.
3.2.11 | Provision for mounting the applanation Tonometer must be provided
3.2.12 | The slit lamp must be provided with chinrest
3.3 | Imaging module
3.3a | Make:
3.3b | Model:
3.3.1 | The slit lamp should have an intergrated digital camera system which can allow enabling of images (imaging software) and video recording of all examination
3.3.2 | Preferance will be given to units with 8 -to 10 megapixel camera or above.
3.3.3 | For easy capturing of images a joystick push button must be provided on the slit lamp
3.3.4 | For teaching and viewing purposes the unit must be provided with monitor of at least 21" with minimum 1280 x 1024 resolution
3.3.5 | All setting must be software driven to make capturing and manipulation of images easier and faster
3.3.6 | The software must allow the user to be able to store and retrieve images.
3.3.7 | The monitor must be mounted on the table
3.3.8 | The unit must supplied with storage space of at least 120GB
3.3.9 | The unit must have an intergrated User chair.
4 | STANDARD ACCESSORIES
4.1 | dust cover
5 | Conformity Compliance (Please attach certificate)
5.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
5.2 | OEM must comply and certified on ISO 9001 quality standards, attach proof of compliance
5.2.1 | Please provide unique ref number of the ISO 9001 ceriticate:
5.3 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
5.3.1 | Please provide unique ref number of the ISO 13485 ceriticate:
5.4 | Model quoted for must be EC certified. Attach a copy of certification
5.4.1 | Please provide unique ref number of the EC ceriticate:
6 | COMPREHENSIVE MAINTENANCE
6.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) must be included in the bid price
6.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
7 | General
7.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
7.2 | The unit on offer must be the latest technology.
OPHTHALMIC SLIT LAMP, intergrated imaging module
RT2-08-09a | LED replacement
RT2-08-09b | Monitor
Bidder to list any other Consumables or Accessories (Not listed above) for the model on offer below

## Tonometer, Hand-held
RT2-08-10 | APPLANATION TONOMETER, HAND-HELD
1 | DESCRIPTION
1.1 | An instrument for determining pressure or tension, especially an instrument for determining ocular tension
2 | Product Offered:
2.1 | Make:
2.2 | Model:
2.3 | Initial Manufacturing year of Model:
3 | SPECIFICATIONS
3.1 | The unit must use LED illumination
3.2 | With measurement range of 0 - 60mmHg
3.3 | The unit must be of light weight not exceeding 1kg
3.4 | The unit should be powered from a rechargeable battery
3.5 | The unit must have an indicator of the battery status
3.6 | The battery must be able to last up to at least 10 000 readings
3.7 | The unit must be supplied with a battery charger and hard case for storage.
4 | STANDARD ACCESSORIES
4.1 | Reusable prism (pk of 10)
4.2 | Tip covers (pk of 10)
5 | Conformity Compliance (Please attach certificate)
5.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
5.2 | OEM must comply and certified on ISO 9001 quality standards, attach proof of compliance
5.2.1 | Please provide unique ref number of the ISO 9001 ceriticate:
5.3 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
5.3.1 | Please provide unique ref number of the ISO 13485 ceriticate:
5.4 | Model quoted for must be EC certified. Attach a copy of certification
5.4.1 | Please provide unique ref number of the EC ceriticate:
6 | COMPREHENSIVE MAINTENANCE
6.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) must be included in the bid price
6.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
7 | General
7.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
7.2 | The unit on offer must be the latest technology.
RT2-08-10 | APPLANATION TONOMETER - Accessories and Consumables
RT2-08-10a | Reusable prism (pk of 10)
RT2-08-10b | Tip covers (pk of 10)
Bidder to list any other Consumables or Accessories (Not listed above) for the model on offer below

## Visual Field Analyser
RT2-08-11 | VISUAL FIELD ANALYZER
1 | Description
2 | Product Offered:
2.1 | Make:
2.2 | Model:
2.3 | Initial Manufacturing year of Model:
2.4 | Radiation Control License number:
3 | Specification
3.1 | The specification establishes the requirement of supply, delivery, end user training, demonstration, commissioning and installation of a Visual Field Analyzer and which must be supplied as a system as per the under mentioned clauses.
3.2 | The unit must be used for measuring the visual field by detecting monocular visual field and also for screeining, monitoring and assiting in the diagnosis of certain conditions in a variety of opthalmic and neurologic diseases
3.3 | The unit must provide information regarding the location of any disease processes or lesion throughout the visual pathway
3.4 | The design of the unit should be such that the paient is presented with a series of light stimuli of varying intensities in different visual field locations and the threshold sensitivity of the retina Is assessed based on the patient ability to consciously detect and respond the these stimuli.
3.5 | The unit must able to distinguish focal from diffuse vision loss.
3.6 | The unit tendered must be designed such that it can be used on patient who are wheelchair users, hearing impaired, have postural and fixation problems and low visual acuity
3.7 | Touch-screen and menu-driven interfaces must simplifies and expedite instrument operation
3.8 | The unit must offer agronomic design for maximum comfort, access and versatility.
3.9 | Complete Archiving system. State the method(s) used.
3.10 | The unit must be provided with a software that can be used to reduce blue-yellow threshold test time and also provide a clinically practical tool for early detection of glucoma.
3.11 | The unit must able to operate in two modes:
3.11.1 | Manual:With patient response registered with a mouse click
3.11.2 | Visual grasp: where the eye tracker senses change in gaze direction as evidence of target acquisition.
3.12 | The sensitivity range is from 0-40dB
3.13 | The following hardware components must be provided with the unit.
3.13.1 | Organic light-emiting diode(OLED) as opposed to LCD's
3.13.2 | Binocular Viewpoint Eye Tracker system which uses infrared video tracking with temporal resolution of 60 Hz.
3.13.3 | For patient requiring far field correction standard trial lenses must be accomondated by lense holders
3.13.4 | Intraocular distance may be adjusted with push pull mechanism that permits independent movement of each display assembly in the horizontal direction.
3.14 | The device must be operated through a portable windows computer that provides access to the four essential functionalities:
3.14.1 | Test setup, input and display of patient information
3.14.2 | Color-coded display of visual field measurement
3.14.3 | VF Expert automatic interpretation of visual field measurement
3.14.4 | Access to achived data
3.15. | The Monitor capability should include but not limited to the following
3.15.1 | Social Security Disability testing
3.15.2 | Monocular testing
3.15.3 | Binocular testing
3.15.4 | Kinetic Testing
3.15.5 | Custom Testing
3.16 | The unit must display rate of vision loss relative to patient age for individual patient care
3.17 | Streamlines clinical interpretation and simplifies patient education
3.18 | A mechanism to limit the need for trial lense correction in patients
3.19 | The visual field testing distance must be infinity
3.20 | The minimum stimulus duration should be 200-300ms
3.21 | Fixation control:
3.21.1 | Blind spot monitoring
3.21.2 | Video eye monitoring
3.21.3 | Gaze tracking
3.21.4 | Head tracking
3.21.5 | Vertex tracking
3.21.6 | Remote eye monitor capability
3.22 | The operator interface must be an touch screen with keyboard
3.23 | For data storage, retrieval and analysis the following must be available
3.23.1 | PC-based
3.23.2 | Minimum of 30GB hard drive or above
3.23.3 | USB
3.24 | A colour printer with printing papers must be provided with the unit.
3.25 | The software features must include the following :
3.25.1 | Single field analysis
3.25.2 | Visual field index
3.25.3 | Guided pregression Analysis
3.25.4 | Serial field overview
3.25.5 | Networking
3.25.6 | When the patient's head is placed on the machine ,it must be possible to adjust the patient head using directional key buttons or any machanism of the latest technollogy. This allows the user not to manualy adjust the patient head for the focal point(s) as some patient maybe critical ill and minimise time. for the whole process.
4 | Standard Accessories and Consumables
4.1 | Dust Cover
5 | Conformity Compliance (Please attach certificate)
5.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
5.2 | OEM must comply and certified on ISO 9001 quality standards, attach proof of compliance
5.2.1 | Please provide unique ref number of the ISO 9001 ceriticate:
5.3 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
5.3.1 | Please provide unique ref number of the ISO 13485 ceriticate:
5.4 | Model quoted for must be EC certified. Attach a copy of certification
5.4.1 | Please provide unique ref number of the EC ceriticate:
5.5 | Hazardous Substance Act: If this type of equipment / apparatus appears on the schedule of Hazardous substances issued by the Directorate: Health Technology of the Department of Health, a license in terms of the Act on Hazardous Substances (Act 15/1973) must be submitted with the bid / quotation documents. The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.6 | All electrical/electronic medical equipment must be licensed by Radiation Control, where applicable a copy of the license must be submitted The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.6.1 | Radiation Control License number:
5.7 | The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade (rubber to with 2 screws) type and it must be a minimum length of three (3) meters. NB the mains cable of the unit tendered for must be SABS colour coded.
5.8 | The equipment quoted for must be protected against electromagnetic interference IEC 60601-1-2
6 | COMPREHENSIVE MAINTENANCE
6.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) must be included in the bid price
6.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
6.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
6.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
6.3 | The bidder to provide the OEM recommended service interval of the unit:
6.4 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
7 | General
7.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
7.2 | The unit on offer must be the latest technology.

## Vitrectomy Unit
RT2-08-12 | VITRECTOMY UNIT
1 | DESCRIPTION
1.1 | Ophthalmic surgical units designed to remove vitreous gel from the eye (vitrectomy)
1.2 | This facility on the unit offered must include the following selectable features:
1.2.1 | Phacoemulsification
1.2.2 | Irrigation
1.2.3 | Irrigation / Aspiration
1.2.4 | Diathermy
1.2.5 | Anterior Vitrectomy
2 | Product Offered:
2.1 | Make:
2.2 | Model:
2.3 | Initial Manufacturing year of Model:
2.4 | Radiation Control License number:
3 | PHACOEMULSIFICATION:
3.1 | This feature on the unit offered must include the following selectable modes and specifications:
3.2 | A fixed mode in which phacoemulsification power and aspiration level are preselected on the console must be available.
3.3 | A linear mode in which phacoemulsification power is fixed while aspiration level is linearly controlled via the footswitch must be available.
3.4 | A linear mode in which phacoemulsification power is footswitch controlled while aspiration level is preselected on the console must be available.
3.5 | A pulsed mode must be available and cover a range of 1 (one) pulse to at least 12 (twelve) pulses per second and it must be adjustable in increments of 1 pulse per second.
3.6 | Phacoemulsification power must be displayed on the console as a percentage of maximum available power.
3.7 | Both the average phaco power and the current power when phaco is in use must be displayed on the console.
3.8 | The elapsed phaco time must be displayed on the console in minutes and seconds.
3.9 | The unit must perform phacoemulsification via a fully autoclavable piezo-electric hand piece.
3.10 | The phaco frequency must be between 20 and 40 kilohertz. Bidder must state the exact frequency of the unit offered.
3.11 | Self-calibration in order to ensure that hand piece and the unit are both in tune must be provided on the unit offered.
4.9 | The unit offered must be able to provide self-diagnosis and trouble shooting with user-friendly display indication on the console.
3.12 | The unit offered must permit customised surgeon setting via independent access keys.
3.13 | The unit must be capable of performing phacoemulsification through and incision in the range of 1.8-3mm. Bidder to specify
5.9 | The tubing packs supplied with the system must be re-usable and re-sterilisable and the cost of which must be included in the final bid price.
3.14 | The fluid collection container supplied with the system must be re-usable and re-sterilisable and the cost of which must be included in the final bid price.
3.15 | Irrigation and Irrigation / Aspiration on the unit offered must meet the following specifications:
3.15.1 | This feature must be controlled by means of a footswitch and the footswitch must be supplied and the cost of which must be included in the final bid price.
3.15.2 | The footswitch on the unit offered must have at least 3 (three) positions permitting switching as follows:
3.15.2a | Irrigation only.
3.15.2b | Irrigation and aspiration
3.15.2c | Irrigation and aspiration with phacoemulsification
3.15.3 | Aspiration on the unit offered must be available both in a preset fixed mode and a linearly variable mode.
3.15.4 | Aspiration must be able to generate suction pressure in a minimum range of 0 to 500mmHg on the unit offered.
3.15.5 | The unit offered must utilise a venturi pump aspiration system.
3.15.6 | The fluid collection cassette must be measureable and capable of holding a minimum of 150cc on the unit offered.
3.15.7 | The unit offered must deliver a clear audio warning signal when the aspiration cassette is nearly full.
3.15.8 | A footswitch-controlled facility must be available on the unit offered to initiate reflux of aspirated contents.
3.15.9 | The audible tone indicating the level of aspiration power being used must be selectable on the unit offered.
3.16 | The diathermy on the system offered must meet the following specifications:
3.16.1 | A fully isolated, footswitch controlled bipolar diathermy pencil must be included.
3.16.2 | Diathermy power must be adjustable from 0 to 100% in reasonable increments. Bidder must state the selectable increments available on the diathermy of the quoted system.Diathermy power must be adjustable from 0 to 100% in reasonable increments. Bidder must state the selectable increments available on the diathermy of the quoted system.
3.16.3 | A volume adjustable audible tone must indicate when diathermy is active.A volume adjustable audible tone must indicate when diathermy is active.
3.17 | Anterior Vitrectomy specification must be as follows on the system offered:
3.17.1 | This facility must include vitreous cutting, irrigation and aspiration capability under footswitch control.
3.17.2 | Vitreous cutting rate must be adjustable from 0 to at least 2500cuts per minute.
3.17.3 | The anterior vitrectomy probe must be of the guillotine variety.
3.18 | Posterior segment facility specifications must be as follows on the system offered:
3.18.1 | The unit must perform posterior vitrectomy with a guillotine type cutting probe.
3.18.2 | Cutting rate must be variable from 0 to 2500 cuts per minute.
3.18.3 | Fluid aspiration must be via linear footswitch control with precise maintenance of low end aspiration.
3.18.4 | The unit must perform intra-ocular illumination via an autoclavable, infrared filtered, fibre optic cable capable of producing a variable intensity while light.
3.18.5 | There must be an automatic switch over facility to a spare bulb in the event of bulb failure
3.18.6 | The unit offered must include an intra-ocular scissors drive with single cut, multi cut variable rate and proportional cut capabilities. These must be preset and footswitch operated.
3.18.7 | The unit offered must include an intra-ocular fully isolated bipolar diathermy facility.
3.18.8 | The unit offered must include an air infusion facility which can operate constantly or variably to maintain intra-ocular pressure (IOP) to preset levels.
3.18.9 | The unit offered must include a phaco fragmentation facility capable of performing fixed or pulsed fragmentation with linear aspiration control.
4 | STANDARD ACCESSORIES
4.1 | Phacoemulsification hand pieces x2.
4.2 | 30-degree phaco needles (6/box).
4.3 | Irrigation / aspiration hand pieces 23 gauge with 0.3mm aspiration port and one must be with a straight tip and the other must be with a curved tip (x2)
4.4 | Irrigation sleeves (6/box).
4.5 | Bipolar eraser cautery with connecting cable (6/box)
4.6 | Anterior vitrectomy set (6/box).
4.7 | Irrigation and phaco pack (6/box)
4.8 | Posterior vitrectomy probe with re-usable and disposable tubing and cassette (6/box)
4.9 | Microscissors driver tubing and cable
4.10 | Fibre optic cable and probe
4.11 | Intra-ocular eraser type diathermy probe with cable
4.12 | Three infusion cannulas 2.5; 4 and 6mm.
4.13 | Pheumatic vitreous scissors and membrane peeler
4.14 | Machmyer vitreous lens.
4.15 | Posterior intra-ocular pressure control tubing sets (6/box)
4.16 | Dust cover
5 | Conformity Compliance (Please attach certificate)
5.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
5.2 | OEM must comply and certified on ISO 9001 quality standards, attach proof of compliance
5.2.1 | Please provide unique ref number of the ISO 9001 ceriticate:
5.3 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
5.3.1 | Please provide unique ref number of the ISO 13485 ceriticate:
5.4 | Model quoted for must be EC certified. Attach a copy of certification
5.4.1 | Please provide unique ref number of the EC ceriticate:
5.5 | Hazardous Substance Act: If this type of equipment / apparatus appears on the schedule of Hazardous substances issued by the Directorate: Health Technology of the Department of Health, a license in terms of the Act on Hazardous Substances (Act 15/1973) must be submitted with the bid / quotation documents. The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.6 | All electrical/electronic medical equipment must be licensed by Radiation Control, where applicable a copy of the license must be submitted The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.6.1 | Radiation Control License number:
5.7 | The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade (rubber to with 2 screws) type and it must be a minimum length of three (3) meters. NB the mains cable of the unit tendered for must be SABS colour coded.
5.8 | The equipment quoted for must be protected against electromagnetic interference IEC 60601-1-2
6 | COMPREHENSIVE MAINTENANCE
6.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) must be included in the bid price
6.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
6.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
6.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
6.3 | The bidder to provide the OEM recommended service interval of the unit:
6.4 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
7 | General
7.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
7.2 | The unit on offer must be the latest technology.
VITRECTOMY UNIT -Accessories and Consumables
RT2-08-12a | Phacoemulsification hand pieces
RT2-08-12b | 30-degree phaco needles (6/box).
RT2-08-12c | Irrigation / aspiration hand pieces 23 gauge with 0.3mm aspiration port and one must be with a straight tip and the other must be with a curved tip (x2)
RT2-08-12d | Irrigation sleeves (6/box).
RT2-08-12e | Bipolar eraser cautery with connecting cable (6/box)
RT2-08-12f | Anterior vitrectomy set (6/box).
RT2-08-12g | Irrigation and phaco pack (6/box)
RT2-08-12h | Posterior vitrectomy probe with re-usable and disposable tubing and cassette (6/box)
RT2-08-12i | Microscissors driver tubing and cable
RT2-08-12j | Fibre optic cable and probe
RT2-08-12k | Intra-ocular eraser type diathermy probe with cable
RT2-08-12l | Three infusion cannulas 2.5; 4 and 6mm.
RT2-08-12m | Pheumatic vitreous scissors and membrane peeler
RT2-08-12n | Machmyer vitreous lens.
RT2-08-12o | Posterior intra-ocular pressure control tubing sets (6/box)
RT2-08-12p | Dust cover
Bidder to list any other Consumables or Accessories (Not listed above) for the model on offer below

## Microscope, Ophthal, Basic
RT2-08-13 | MICROSCOPE, OPHTHALMOLOGY, BASIC
1 | DESCRIPTION
1.1 | A surgical microscope suitable for opthalmology surgery with an 180 degrees variable angle binocular tube configuration with foot pedal.
2 | Product Offered:
2.1 | Make:
2.2 | Model:
2.3 | Initial Manufacturing year of Model:
2.4 | Radiation Control License number:
3 | TECHNICAL SPECIFICATION
3.1 | The unit offered should include focusing, zooming and magnification controls.
3.2 | The unit must have a motorized zoom system with apochromatic lens, zoom ratio 1:6
3.3 | The unit offered should have either an LED light source and/or Xenon (180w) illumination system via a fibre optic cable.
3.4 | A floor stand shall have atleast 4 castors to ensure base stability and a minimum of 2 brakes on castors for smooth operation of the microscope.
3.5 | The microscope stand must have a stand mounted grip for easy movement of the unit
3.6 | The maximum load of the accessories to the microscope should not affect stability of the microscope
3.7 | The base of the stand shall be compact and have 360 degrees rotatable wheels to allow easy positioning of the microscope.
3.8 | The foot switch should be wireless or wired and waterproof
3.9 | The microscope offered should have pedal and/or hand-switch controls for ease of operation
3.10 | The pedal must be waterproof and inclinable.
3.11 | The surface of the microscope shall be anti-bacterial and have anti-mycotic protection.
4.9 | The optic carrier shall have an electric system to balance the three gravitational angles.
3.12 | Handles must be autoclave-able
3.13 | The microscope should have binocular tubes for the main surgeon and possible for the assistant (as and when needed)
3.14 | Surgeons shall be able to store their personal preferences for magnification, focus and illumination.
3.15 | System Functionality
3.15.1 | Magnification must be atleast 3 step with 1.6 to 20 x range or more
3.15.2 | Optics: Apochromatic with magnification changer
3.15.3 | Stereo base: 22mm - 25mm
3.15.4 | The trolley should have antistatic castors with at least 2 equipped with locking system/brake.
3.15.5 | Eyepiece:
3.15.5a | · 10x wide-field eyepieces
3.15.5b | · 12.5x wide-field eyepieces
3.15.5c | · Dioptric setting ±5 with adjustable eye cup
3.15.6 | Lateral tilt: 0 degrees - 360 degrees
3.15.7 | Inclination angle: 200 degrees
3.15.8 | Binocular tubes: 0 - 180 degrees
3.15.9 | Field of view diameter: 8 mm to 55 mm
3.15.10 | Focal length ranging from 150 - 400mm
3.16 | Illumination
3.16.1 | Two LED or Xenon lamps source with emergency back-up lamp.
3.16.2 | Protection filter: UV & IR filter
3.17 | X-Y Coupling
3.17.1 | Should have motorised X-Y movement
3.17.2 | Motorised focusing and zooming
3.17.3 | Should have a key for initial position of X-Y coupling and focus
4 | STANDARD ACCESSORIES
4.1 | Binocular tubes
4.2 | Autoclavable handle
4.3 | Eyepiece cups
4.4 | Dustcover
4.5 | Second observer
4.6 | Objective lense: 200mm - 400mm
5 | Conformity Compliance (Please attach certificate)
5.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
5.2 | OEM must comply and certified on ISO 9001 quality standards, attach proof of compliance
5.2.1 | Please provide unique ref number of the ISO 9001 ceriticate:
5.3 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
5.3.1 | Please provide unique ref number of the ISO 13485 ceriticate:
5.4 | Model quoted for must be EC certified. Attach a copy of certification
5.4.1 | Please provide unique ref number of the EC ceriticate:
5.5 | Hazardous Substance Act: If this type of equipment / apparatus appears on the schedule of Hazardous substances issued by the Directorate: Health Technology of the Department of Health, a license in terms of the Act on Hazardous Substances (Act 15/1973) must be submitted with the bid / quotation documents. The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.6 | All electrical/electronic medical equipment must be licensed by Radiation Control, where applicable a copy of the license must be submitted The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.6.1 | Radiation Control License number:
5.7 | The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade (rubber to with 2 screws) type and it must be a minimum length of three (3) meters. NB the mains cable of the unit tendered for must be SABS colour coded.
5.8 | The equipment quoted for must be protected against electromagnetic interference IEC 60601-1-2
6 | COMPREHENSIVE MAINTENANCE
6.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) must be included in the bid price
6.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
6.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
6.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
6.3 | The bidder to provide the OEM recommended service interval of the unit:
6.4 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
7 | General
7.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
7.2 | The unit on offer must be the latest technology.
MICROSCOPE, OPHTHALMOLOGY- Accessories and Consumables
RT2-08-13a | Binocular tubes
RT2-08-13b | Autoclavable handle
RT2-08-13c | Eyepiece cups
RT2-08-13d | Dustcover
RT2-08-13e | Second observer
RT2-08-13f | Xenon illumination system
RT2-08-13g | Stereo co-observation tube/Teaching piece
Bidder to list any other Consumables or Accessories (Not listed above) for the model on offer below

## Microscope, Ophthal, Advanced
RT2-08-14 | OPHTHALMIC SURGICAL MICROSCOPE, ADVANCED
1 | DESCRIPTION
1.1 | A surgical microscope suitable for opthalmology surgery with an 180 degrees variable angle binocular tube configuration with foot pedal and an external monitor. The microscope must have image storage and a 3 CCD high definition camera.
1.2 | An electronic surgical microscope with full display medical grade monitor, minimum 2TB Image storage and a 3 CCD high definition camera.
2 | Product Offered:
2.1 | Make:
2.2 | Model:
2.3 | Initial Manufacturing year of Model:
2.4 | Radiation Control License number:
3 | TECHNICAL SPECIFICATION
3.1 | The unit offered should include focusing, zooming and magnification controls.
3.2 | The unit must have a motorized zoom system with apochromatic lens, zoom ratio 1:6
3.3 | The unit must be integrated with both slit illuminator and keratoscope
3.4 | The unit offered should have either an LED light source and/or Xenon (180w) illumination system via a fibre optic cable.
3.5 | A floor stand shall have atleast 4 castors to ensure base stability and all 4 brakes on castors for smooth operation of the microscope.
3.6 | The microscope stand must have a stand mounted grip for easy movement of the unit
3.7 | The maximum load of the accessories to the microscope should not affect stability of the microscope
3.8 | The base of the stand shall be compact and have 360 degrees rotatable wheels to allow easy positioning of the microscope.
3.9 | The foot switch should be wireless or wired and waterproof
3.10 | The microscope offered should have pedal and/or hand-switch controls for ease of operation
3.11 | The pedal must be waterproof and inclinable, with a minimum of 15 programmable functions to control all the microscope's functions.
4.9 | The surface of the microscope shall be anti-bacterial and have anti-mycotic protection.
3.12 | The optic carrier shall have an electric system to balance the three gravitational angles.
3.13 | The microscope shall have a function to reduce the light intensity at a short distance during examination
5.9 | Handles must be autoclave-able
3.14 | The microscope offered shall have an integrated monitor, of at least 21" with HD resolution.
3.15 | The microscope should have binocular tubes for the main surgeon and possible for the assistant (as and when needed)
6.9 | The control module shall have a screen with touch pad to select different user profiles and settings.
3.16 | Surgeons shall be able to store their personal preferences for magnification, focus and illumination.
3.17 | System Functionality
3.17.1 | Motorized magnification must be atleast 3 step with 1.6 to 20 x range or more
3.17.2 | Optics: Apochromatic with magnification changer
3.17.3 | Stereo base: 22mm - 25mm
3.17.4 | The trolley should have antistatic castors with at least 4 equipped with locking system/brake.
3.17.5 | Eyepiece:
3.17.5a | · 10x wide-field eyepieces
3.17.5b | · 12.5x wide-field eyepieces
3.17.5c | · Dioptric setting ±5 with adjustable eye cup
3.17.6 | Lateral tilt: 0 degrees - 360 degrees
3.17.7 | Inclination angle: 200 degrees
3.17.8 | Binocular tubes: 0 - 180 degrees
3.17.9 | Field of view diameter: 8 mm to 55 mm
3.17.10 | Focal length ranging from 150 - 400mm
3.18 | Illumination
3.18.1 | Two LED or Xenon lamps source with emergency back-up lamp.
3.18.2 | Protection filter: UV & IR filter
3.19 | X-Y Coupling
3.19.1 | Should have motorised X-Y movement
3.19.2 | Motorised focusing and zooming
3.19.3 | Should have a key for initial position of X-Y coupling and focus
4 | STANDARD ACCESSORIES
4.1 | Binocular tubes
4.2 | Autoclavable handle
4.3 | Eyepiece cups
4.5 | Dustcover
4.6 | Second observer
4.7 | Video adapter
4.8 | Objective lense: 200mm - 400mm
5 | Conformity Compliance (Please attach certificate)
5.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
5.2 | OEM must comply and certified on ISO 9001 quality standards, attach proof of compliance
5.2.1 | Please provide unique ref number of the ISO 9001 ceriticate:
5.3 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
5.3.1 | Please provide unique ref number of the ISO 13485 ceriticate:
5.4 | Model quoted for must be EC certified. Attach a copy of certification
5.4.1 | Please provide unique ref number of the EC ceriticate:
5.5 | Hazardous Substance Act: If this type of equipment / apparatus appears on the schedule of Hazardous substances issued by the Directorate: Health Technology of the Department of Health, a license in terms of the Act on Hazardous Substances (Act 15/1973) must be submitted with the bid / quotation documents. The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.6 | All electrical/electronic medical equipment must be licensed by Radiation Control, where applicable a copy of the license must be submitted The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.6.1 | Radiation Control License number:
5.7 | The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade (rubber to with 2 screws) type and it must be a minimum length of three (3) meters. NB the mains cable of the unit tendered for must be SABS colour coded.
5.8 | The equipment quoted for must be protected against electromagnetic interference IEC 60601-1-2
6 | COMPREHENSIVE MAINTENANCE
6.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) must be included in the bid price
6.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
6.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
6.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
6.3 | The bidder to provide the OEM recommended service interval of the unit:
6.4 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
7 | General
7.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
7.2 | The unit on offer must be the latest technology.
MICROSCOPE, OPHTHALMOLOGY- Accessories and Consumables
RT2-08-14a | Binocular tubes
RT2-08-14b | Autoclavable handle
RT2-08-14c | Eyepiece cups
RT2-08-14d | Dustcover
RT2-08-14e | Second observer
RT2-08-14f | Xenon illumination system
RT2-08-14g | Stereo co-observation tube/Teaching piece
RT2-08-14h | Monitor
Bidder to list any other Consumables or Accessories (Not listed above) for the model on offer below

## Category-9 Neuro
NEURO EQUIPMENT (Neurlogy & Neurosurgery)
# | Item number | Item Description
1 | RT2-09-01 | EEG, Portable
2 | RT2-09-02 | EEG, Desktop
3 | RT2-09-03 | Ultrasonic Surgical Aspirator
4 | RT2-09-04 | Neuronavigation System
5 | RT2-09-05 | Microscope, Neurosurgery
6 | RT2-09-06 | Microscope, Neurosurgery, Advanced

## EEG, Portable
RT2-09-01 | EEG PORTABLE
1 | DESCRIPTION
1.1 | an electronic monitoring device to measure and record electrical activity in the brain.
2 | Product Offered:
2.1 | Make:
2.2 | Model:
2.3 | Initial Manufacturing year of Model:
2.4 | Radiation Control License number:
3 | TECHNICAL SPECIFICATION
3.1 | The EEG unit quoted for must be user friendly, it must comprise the latest technology and it must be a 32 Channel computerized Portable EEG – 17" Notebook with the following minimum spec: Core i7, 8GB of RAM, 256GB SSD or better, CD/DVD RW.
3.2 | The Portable EEG Unit must be upgradeable to Polysomnographic facilities for respiratory assessments and sleep staging.
3.3 | The Portable EEG Unit must upgradeable to Sleep analysis software, including detailed analysis for respiratory disturbances and sleep stages.
3.4 | The EEG unit must be provided with the following software requirements: spike and seizure detection, EEG mapping, polysmith analysis and report software.
3.5 | The system offered must be specifically based on the easy to use Operating System as per end-user’s requirements.
3.6 | The system offered must be able to display a minimum of 8 user definable channels of DSA trendgraphs (Density Spectral Array) during recording and Review.
3.7 | The system offered must be able to display 64 channels plus one mark channel.
3.8 | An interfacing between the electrode junction box and computer is essential on the unit offered. (USB or latest equivalent)
3.9 | The system offered must provide for a standard Networking capability as per the end-user’s requirements.
3.10 | Automatic photic stimulation with up to 30 steps must be provided on the unit offered.
3.11 | A 3D Voltage mapping feature with customizable views must be provided as standard on the unit offered.
3.12 | Data review during acquisition must be available on the unit offered.
3.13 | Simultaneous acquisition and review of the current recording in a configurable split screen for comparison purpose with freeze function must be provided on the unit offered.
3.14 | A file append feature that allows separate EEG’s to be combined into one file must be provided on the unit offered.
3.15 | A facility to copy any part of the waveform window by just dragging and dropping for comparison functionally must be provided on the unit offered.
3.16 | A facility to save EEG data to CD-R with integrated embedded review software for review on any PC without the need of additional software must be provided on the unit offered.
3.17 | A montage map displaying the electrode and montage positioning with print function must be provided on the unit offered.
3.18 | A high resolution 1200Dpi laser printer or more advanced must be provided on the unit offered.
3.19 | The headbox offered must be mountable and provide for a minimum of 38 simultaneously recorded inputs.
3.20 | The system offered must provide simultaneous inputs for at least: 25 EEG Inputs, 14 Bipolar Inputs, 3 Respiratory Inputs, 4 DC Inputs, 1 x SpO2 and 1 CO2 Inputs.
3.21 | The Input impedance must be 100 MΩ or greater on the unit offered.
3.22 | The Input leakage current must be less than 5nA on the unit offered.
3.23 | The Internal noise level must be less than 1.5 µVp-p (0.53 Hz to 60 Hz) on the unit offered.
3.24 | The CMRR must be 105 dB or greater at 50Hz on the unit offered.
3.25 | The low-cut hardware filter for non DC inouts must be 0.08 Hz or less on the unit offered.
3.26 | The High-cut hardware filter must be 300 Hz with a (-18 dB/oct) roll off or better on the unit offered.
3.27 | The A/D conversion must be a simultaneous sampling of all inputs at 16bits per channel on the unit offered.
3.28 | There must be a selectable sampling rate per input of 100Hz to 1000Hz on the unit offered.
3.29 | The Display Resolution must range from 1024 X 768 to 1600 x 1200 on the unit offered.
3.30 | Overwrite and page by page display modes must be provided on the unit offered.
3.31 | Waveform display position feature must be provided on the unit offered
3.32 | The waveform sweep speed must be from 5seconds to 5 minutes per page on the unit offered.
3.33 | Timing marks. Timing scale and event marks must be provided on screen on the unit offered.
3.34 | An EEG input sensitivity of: 1,2,3,(2.5), 5,7,10,15,20,30,50,75,100,150,200 µVmm must be provided on the unit offered.
3.35 | A DC input sensibility of: off, 10, 15, 20, 30, 50, 70, 100, 150, 200 mV/mm must be provided on the unit offered.
3.36 | The time constant must be variable in steps of; 0.001, 0.003, 0.03, 0.1, 0.3, 0.6, 1.0, 2.0 on the unit offered.
3.37 | The low-cut filter on the unit offered must be variable in steps of: 0.08, 0.16, 0.27, .053, 1.6, 5.3, 53, 159 Hz with a roll off of (-6 dB/ oct).
3.38 | The High-cut filter on the unit offered must be variable with the following standard options: 15, 30, 35, 60, 70, 120 Hz (-12 dB/oct) 50 Hz (RAPID), 3000 Hz (-18 dB/oct).
3.39 | The AC interference filter: at 50 or 60 Hz must have a rejection ratio of 1/25 or more on the unit offered.
3.40 | A calibration waveform on the unit offered must have the option for a Step of sine wave and variable voltages of: 2, 5, 10, 20, 50, 100, 200, 500, 1000 µV (x 1000 for DC input signals).
3.41 | An ECG artifact elimination filter must be available in acquisition and review modes on the unit offered.
3.42 | An impedance check function, on screen with values and at the headbox with simultaneous display using threshold method must be provided on the unit offered.
3.43 | The ability to program at least 36 sets of programmable montages combined with programmable individual amplifier settings must be provided on the unit offered.
3.44 | The reference electrode on the unit offered must be selectable between A1→A2, A1←A2, A1↔A2, A1+A2, VX, AV (with unsuitable electrode deletion function), Aav, Org, SD and OFF.
3.45 | The following events on the unit offered must be automatically marked: Start stop Recording, montage change with montage name, patient, button press, photic stimulation pulse and frequency, hyperventilation with time stamp.
3.46 | The display resolution on the unit offered must support up to 1600 x 1280.
3.47 | The ability to display up to 64 simultaneous waveform channels and 1 mark channels must be provided on the unit offered.
3.48 | Functions to adjust the waveform colour, position, width, number of channels must be provided on the unit offered.
3.49 | Multiple paging modes such as page overwrite and page-by-page must be provided on the unit offered.
3.50 | A waveform Freeze function must be provided on the unit offered.
3.51 | The waveform sweep speed should be adjustable; 5, 10, 20, 30, 60 s/ page or 5 min/page on the unit offered.
3.52 | The displayed waveforms on the unit offered should include user selectable time markings of 0.1 and 1 seconds or both.
3.53 | Pre-programmable and real-time user event markings must be provided on the unit offered.
3.54 | EEG scales and cursor measurement scales for amplitude, time and frequency must be provided on the unit offered.
3.55 | The review program on the unit offered must allow adjustment of the following parameters montage, sensitivity, high–cut filter, time constant, reference electrode and display speed.
3.56 | Jump functions, to allow quick access to: Specified event, page by page, and specified time must be provided on the unit offered.
3.57 | Review display modes on the unit offered must include the following for productive EEG scoring: continuous, high speed, high speed with pause, manually page by page, manually second by seconds, and manually waveform centering.
4 | The following sensors and accessories must be included
4.1 | Oral and nasal airflow
4.2 | Respiratory effort sensors for chest and abdomen.
4.3 | Pulse oximetery sensor (finger type)
4.4 | Snoring sensor.
4.5 | Body Position sensor.
4.6 | Automatic CPAP Titration System with manual remote override ability and analogue signal out for connection to the recording system to monitor flow, CPAP pressure leak etc.
4.7 | Carrying case
5 | Conformity Compliance (Please attach certificate)
5.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
5.2 | OEM must comply and certified on ISO 9001 quality standards, attach proof of compliance
5.2.1 | Please provide unique ref number of the ISO 9001 ceriticate:
5.3 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
5.3.1 | Please provide unique ref number of the ISO 13485 ceriticate:
5.4 | Model quoted for must be EC certified. Attach a copy of certification
5.4.1 | Please provide unique ref number of the EC ceriticate:
5.5 | Hazardous Substance Act: If this type of equipment / apparatus appears on the schedule of Hazardous substances issued by the Directorate: Health Technology of the Department of Health, a license in terms of the Act on Hazardous Substances (Act 15/1973) must be submitted with the bid / quotation documents. The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.6 | All electrical/electronic medical equipment must be licensed by Radiation Control, where applicable a copy of the license must be submitted The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.6.1 | Radiation Control License number:
5.7 | The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade (rubber to with 2 screws) type and it must be a minimum length of three (3) meters. NB the mains cable of the unit tendered for must be SABS colour coded.
5.8 | The equipment quoted for must be protected against electromagnetic interference IEC 60601-1-2
6 | COMPREHENSIVE MAINTENANCE
6.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) must be included in the bid price
6.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
6.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
6.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
6.3 | The bidder to provide the OEM recommended service interval of the unit:
6.4 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
7 | General
7.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
7.2 | The unit on offer must be the latest technology.
EEG PORTABLE- Accessories, Consumables and Upgrades
RT2-09-01a | Oral and nasal airflow
RT2-09-01b | Respiratory effort sensors for chest and abdomen.
RT2-09-01c | Pulse oximetery sensor (finger type)
RT2-09-01d | Snoring sensor.
RT2-09-01e | Body Position sensor.
RT2-09-01f | Automatic CPAP Titration System with manual remote override ability and analogue signal out for connection to the recording system to monitor flow, CPAP pressure leak etc.
RT2-09-01g | Carrying case
RT2-09-01h | SLEEP ANALYSIS SOFTWARE (UPGRADE) -specifications below
Dedicated software for the analysis of sleep disordered breathing & periodic limb movements.
Display and analysis of sleep stages and respiratory data
Generation of comprehensive reports.
Re-formatting of individual channel filters and sensitivity.
Re-formatting of montages.
Time linked trend graphs superimposed on waveform data
Automatic analysis data types must include the following: EEG1, EEG2, EOG1, EOG2, ECG, Chin, EMG, LEG, EMG, Airflow, Chest Effort, Abdomen Effort, SpO2, CPAP, BIPAP, Snoring, Body Position, Pulse Rate, Airflow2, ETCO2, R-R, pH.
Support a dual monitor and enhanced multitasking interface.
Must allow easy, fast, and accurate scoring and editing.
Graphical trending of all scanned and scored parameters
Monitor multiple recordings from a single station.
On-line scanning, scoring and editing
Separate adult and infant data analysis settings.
Single click scoring and editing.
Multiple synchronized time bases.
Navigate data by graphical interface, comment query, event query and scroll bar.
Customized event display.
View AHI any time in the study or by treatment level.
Automatic event association.
Customized workspaces – manual scored changes are stored separately for each scorer allowing QA comparisons and preserving manual changes even if the test is re-analyzed.
Automatic frame insertion into report from digital video.
Complete post-study remontaging and refiltering.
Score data over LAN, WAN or Internet.
Export and Import of European Data Format (EDF).
Dedicated Infant data analysis must support.
Active, Quiet, Transitional and Wake scoring
Infant staging, trending and graphics
Periodic breathing analysis
Ph analysis and reporting (requires PH probe).
Complete infant database with correlations of all types of clinical events with infant staging.
Flexible and easy-to-use DC channel set up and calibration
RT2-09-01i | REPORTING FUNCTIONALITY (UPGRADE):
Auto text complete field searches that allow you to find the values or fields you want to insert to into your reports.
Auto pop-up field descriptions that give you easy to understand field descriptions
Customizable graphs
Comparison reports give your lab epoch-to-epoch comparisons of the different scores in the lab
Create customized impression and recommendation templates.
Seamless interface over the LAN, WAN, or Internet for remote interpretation from anywhere.
HTML report format compatible with most word processing software.
Direct compatibility with MicrosoftTM Internet Explore
Auto e-mail directly from the report Generator.
Three re-usable EEG paediatric head caps of the following sizes: 1x 42cm to 46cm; 1x 38cm to 42cm; 1x 34cm to 38cm must be included and supplied with the unit offered at no extra cost to the final bid price. These head caps must include a complete set of re-usable electrodes and electrode board adaptor and these must also be supplied at no extra cost to the final bid pricel

## EEG, Desktop
RT2-09-02 | EEG Machine, Desktop
1 | DESCRIPTION
1.1 | an electronic monitoring device to measure and record electrical activity in the brain.
2 | Product Offered:
2.1 | Make:
2.2 | Model:
2.3 | Initial Manufacturing year of Model:
2.4 | Radiation Control License number:
3 | TECHNICAL SPECIFICATION
3.1 | The EEG unit complete with software for acquisition and review and compatible computer with necessary interface and printer as required.
3.2 | The EEG unit quoted for must be user friendly, it must comprise the latest technology and it must be a 32 Channel computerized EEG – PC based system with the following minimum spec: Core i7 CPU, 8GB of RAM, 256GB SSD or better, CD/DVD RW, 19" LCD TFT display, Keyboard, Mouse and UPS. All placed neatly on a trolley.
3.3 | The EEG Unit must be upgradeable to Polysomnographic facilities for respiratory assessments and sleep staging.
3.4 | The EEG Unit must upgradeable to Sleep analysis software, including detailed analysis for respiratory disturbances and sleep stages.
3.5 | The EEG unit must be provided with the following software requirements: spike and seizure detection, EEG mapping, polysmith analysis and report software.
3.6 | The system offered must be specifically based on the easy to use Operating System as per end-user’s requirements.
3.7 | The system offered must be able to display a minimum of 8 user definable channels of DSA trendgraphs (Density Spectral Array) during recording and Review.
3.8 | The system offered must be able to display 64 channels plus one mark channel.
3.9 | An interfacing between the electrode junction box and computer is essential on the unit offered. (USB or latest equivalent)
3.10 | The electrode junction box must be mounted on a EEG trolley or provided on a separate stand but both option must offer accessory basket.
3.11 | The system offered must provide for a standard Networking capability as per the end-user’s requirements.
3.12 | Automatic photic stimulation with up to 30 steps must be provided on the unit offered.
3.13 | A 3D Voltage mapping feature with customizable views must be provided as standard on the unit offered.
3.14 | Data review during acquisition must be available on the unit offered.
3.15 | Simultaneous acquisition and review of the current recording in a configurable split screen for comparison purpose with freeze function must be provided on the unit offered.
3.16 | A file append feature that allows separate EEG’s to be combined into one file must be provided on the unit offered.
3.17 | A facility to copy any part of the waveform window by just dragging and dropping for comparison functionally must be provided on the unit offered.
3.18 | A facility to save EEG data to CD-R with integrated embedded review software for review on any PC without the need of additional software must be provided on the unit offered.
3.19 | A montage map displaying the electrode and montage positioning with print function must be provided on the unit offered.
3.20 | A high resolution 1200Dpi laser printer or more advanced must be provided on the unit offered.
3.21 | The headbox offered must provide for a minimum of 38 simultaneously recorded inputs.
3.22 | The system offered must provide simultaneous inputs for at least: 25 EEG Inputs, 14 Bipolar Inputs, 3 Respiratory Inputs, 4 DC Inputs, 1 x SpO2 and 1 CO2 Inputs.
3.23 | The Input impedance must be 100 MΩ or greater on the unit offered.
3.24 | The Input leakage current must be less than 5nA on the unit offered.
3.25 | The Internal noise level must be less than 1.5 µVp-p (0.53 Hz to 60 Hz) on the unit offered.
3.26 | The CMRR must be 105 dB or greater at 50Hz on the unit offered.
3.27 | The low-cut hardware filter for non DC inouts must be 0.08 Hz or less on the unit offered.
3.28 | The High-cut hardware filter must be 300 Hz with a (-18 dB/oct) roll off or better on the unit offered.
3.29 | The A/D conversion must be a simultaneous sampling of all inputs at 16bits per channel on the unit offered.
3.30 | There must be a selectable sampling rate per input of 100Hz to 1000Hz on the unit offered.
3.31 | The Display Resolution must range from 1024 X 768 to 1600 x 1200 on the unit offered.
3.32 | Overwrite and page by page display modes must be provided on the unit offered.
3.33 | Waveform display position feature must be provided on the unit offered
3.34 | The waveform sweep speed must be from 5seconds to 5 minutes per page on the unit offered.
3.35 | Timing marks. Timing scale and event marks must be provided on screen on the unit offered.
3.36 | An EEG input sensitivity of: 1,2,3,(2.5), 5,7,10,15,20,30,50,75,100,150,200 µVmm must be provided on the unit offered.
3.37 | A DC input sensibility of: off, 10, 15, 20, 30, 50, 70, 100, 150, 200 mV/mm must be provided on the unit offered.
3.38 | The time constant must be variable in steps of; 0.001, 0.003, 0.03, 0.1, 0.3, 0.6, 1.0, 2.0 on the unit offered.
3.39 | The low-cut filter on the unit offered must be variable in steps of: 0.08, 0.16, 0.27, .053, 1.6, 5.3, 53, 159 Hz with a roll off of (-6 dB/ oct).
3.40 | The High-cut filter on the unit offered must be variable with the following standard options: 15, 30, 35, 60, 70, 120 Hz (-12 dB/oct) 50 Hz (RAPID), 300 Hz (-18 dB/oct).
3.41 | The AC interference filter: at 50 or 60 Hz must have a rejection ratio of 1/25 or more on the unit offered.
3.42 | A calibration waveform on the unit offered must have the option for a Step of sine wave and variable voltages of: 2, 5, 10, 20, 50, 100, 200, 500, 1000 µV (x 1000 for DC input signals).
3.43 | An ECG artifact elimination filter must be available in acquisition and review modes on the unit offered.
3.44 | An impedance check function, on screen with values and at the headbox with simultaneous display using threshold method must be provided on the unit offered.
3.45 | The ability to program at least 36 sets of programmable montages combined with programmable individual amplifier settings must be provided on the unit offered.
3.46 | The reference electrode on the unit offered must be selectable between A1→A2, A1←A2, A1↔A2, A1+A2, VX, AV (with unsuitable electrode deletion function), Aav, Org, SD and OFF.
3.47 | The following events on the unit offered must be automatically marked: Start stop Recording, montage change with montage name, patient, button press, photic stimulation pulse and frequency, hyperventilation with time stamp.
3.48 | The display resolution on the unit offered must support up to 1600 x 1280.
3.49 | The ability to display up to 64 simultaneous waveform channels and 1 mark channels must be provided on the unit offered.
3.50 | Functions to adjust the waveform colour, position, width, number of channels must be provided on the unit offered.
3.51 | Multiple paging modes such as page overwrite and page-by-page must be provided on the unit offered.
3.52 | A waveform Freeze function must be provided on the unit offered.
3.53 | The waveform sweep speed should be adjustable; 5, 10, 20, 30, 60 s/ page or 5 min/page on the unit offered.
3.54 | The displayed waveforms on the unit offered should include user selectable time markings of 0.1 and 1 seconds or both.
3.55 | Pre-programmable and real-time user event markings must be provided on the unit offered.
3.56 | EEG scales and cursor measurement scales for amplitude, time and frequency must be provided on the unit offered.
3.57 | The review program on the unit offered must allow adjustment of the following parameters montage, sensitivity, high–cut filter, time constant, reference electrode and display speed.
3.58 | Jump functions, to allow quick access to: Specified event, page by page, and specified time must be provided on the unit offered.
3.59 | Review display modes on the unit offered must include the following for productive EEG scoring: continuous, high speed, high speed with pause, manually page by page, manually second by seconds, and manually waveform centering.
4 | The following Sensors and Accessories must be included
4.1 | Oral and nasal airflow
4.2 | Respiratory effort sensors for chest and abdomen.
4.3 | Pulse oximetery sensor (finger type)
4.4 | Snoring sensor.
4.5 | Body Position sensor.
4.6 | Automatic CPAP Titration System with manual remote override ability and analogue signal out for connection to the recording system to monitor flow, CPAP pressure leak etc.
4.7 | Mobile trolley
4.8 | UPS
4.9 | Laser printer
4.10 | Printing paper for atleast 1000 EEG tests
5 | Conformity Compliance (Please attach certificate)
5.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
5.2 | OEM must comply and certified on ISO 9001 quality standards, attach proof of compliance
5.2.1 | Please provide unique ref number of the ISO 9001 ceriticate:
5.3 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
5.3.1 | Please provide unique ref number of the ISO 13485 ceriticate:
5.4 | Model quoted for must be EC certified. Attach a copy of certification
5.4.1 | Please provide unique ref number of the EC ceriticate:
5.5 | Hazardous Substance Act: If this type of equipment / apparatus appears on the schedule of Hazardous substances issued by the Directorate: Health Technology of the Department of Health, a license in terms of the Act on Hazardous Substances (Act 15/1973) must be submitted with the bid / quotation documents. The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.6 | All electrical/electronic medical equipment must be licensed by Radiation Control, where applicable a copy of the license must be submitted The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.6.1 | Radiation Control License number:
5.7 | The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade (rubber to with 2 screws) type and it must be a minimum length of three (3) meters. NB the mains cable of the unit tendered for must be SABS colour coded.
5.8 | The equipment quoted for must be protected against electromagnetic interference IEC 60601-1-2
6 | COMPREHENSIVE MAINTENANCE
6.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) must be included in the bid price
6.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
6.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
6.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
6.3 | The bidder to provide the OEM recommended service interval of the unit:
6.4 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
7 | General
7.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
7.2 | The unit on offer must be the latest technology.
EEG Machine, Desktop - Accessories, Consumables and Upgrades
RT2-09-02a | Oral and nasal airflow
RT2-09-02b | Respiratory effort sensors for chest and abdomen.
RT2-09-02c | Pulse oximetery sensor (finger type)
RT2-09-02d | Snoring sensor.
RT2-09-02e | Body Position sensor.
RT2-09-02f | Automatic CPAP Titration System with manual remote override ability and analogue signal out for connection to the recording system to monitor flow, CPAP pressure leak etc.
RT2-09-02g | Mobile trolley
RT2-09-02h | UPS
RT2-09-02i | Laser printer
RT2-09-02j | Printing paper for atleast 1000 EEG tests
RT2-09-02k | SLEEP ANALYSIS SOFTWARE (UPGRADE)– Specification below
Dedicated software for the analysis of sleep disordered breathing & periodic limb movements.
Display and analysis of sleep stages and respiratory data
Generation of comprehensive reports.
Re-formatting of individual channel filters and sensitivity.
Re-formatting of montages.
Time linked trend graphs superimposed on waveform data
Automatic analysis data types must include the following: EEG1, EEG2, EOG1, EOG2, ECG, Chin, EMG, LEG, EMG, Airflow, Chest Effort, Abdomen Effort, SpO2, CPAP, BIPAP, Snoring, Body Position, Pulse Rate, Airflow2, ETCO2, R-R, pH.
Support a dual monitor and enhanced multitasking interface.
Must allow easy, fast, and accurate scoring and editing.
Graphical trending of all scanned and scored parameters
Monitor multiple recordings from a single station.
On-line scanning, scoring and editing
Separate adult and infant data analysis settings.
Single click scoring and editing.
Multiple synchronized time bases.
Navigate data by graphical interface, comment query, event query and scroll bar.
Customized event display.
View AHI any time in the study or by treatment level.
Automatic event association.
Customized workspaces – manual scored changes are stored separately for each scorer allowing QA comparisons and preserving manual changes even if the test is re-analyzed.
Automatic frame insertion into report from digital video.
Complete post-study remontaging and refiltering.
Score data over LAN, WAN or Internet.
Export and Import of European Data Format (EDF).
Dedicated Infant data analysis must support.
Active, Quiet, Transitional and Wake scoring
Infant staging, trending and graphics
Periodic breathing analysis
Ph analysis and reporting (requires PH probe).
Complete infant database with correlations of all types of clinical events with infant staging.
Flexible and easy-to-use DC channel set up and calibration
RT2-09-02l | REPORTING FUNCTIONALITY (UPGRADE):
Auto text complete field searches that allow you to find the values or fields you want to insert to into your reports.
Auto pop-up field descriptions that give you easy to understand field descriptions
Customizable graphs
Comparison reports give your lab epoch-to-epoch comparisons of the different scores in the lab
Create customized impression and recommendation templates.
Seamless interface over the LAN, WAN, or Internet for remote interpretation from anywhere.
HTML report format compatible with most word processing software.
Direct compatibility with MicrosoftTM Internet Explore
Auto e-mail directly from the report Generator.
Three re-usable EEG paediatric head caps of the following sizes: 1x 42cm to 46cm; 1x 38cm to 42cm; 1x 34cm to 38cm must be included and supplied with the unit offered at no extra cost to the final bid price. These head caps must include a complete set of re-usable electrodes and electrode board adaptor and these must also be supplied at no extra cost to the final bid pricel

## Ultrasonic Surgical Aspirator
RT2-09-03 | ULTRASONIC SURGICAL ASPIRATOR
1 | DESCRIPTION
1.1 | A medical instrument that uses suction to withdraw fluids from other parts of the body during surgery
1.2 | The Surgical Aspirator must use the latest technology (magnetostrictive or better), incorporating the CEM Nose cone or equivalent and suitable for use in various surgical procedures.
1.3 | The unit must be supplied with a trolley and it should have at least 4 antistatic castors with at least 2 equipped with locking system/brake.
2 | Product Offered:
2.1 | Make:
2.2 | Model:
2.3 | Initial Manufacturing year of Model:
2.4 | Radiation Control License number:
3 | TECHNICAL SPECIFICATION
3.1 | It must be suitable for the following applications for soft tissue:
3.1.1 | Fragmentation
3.1.2 | Emulsification
3.1.3 | Aspiration
3.2 | It must be suitable for the following applications for hard tissue:
3.2.1 | Fragmentation
3.2.2 | Emulsification
3.2.3 | Aspiration
3.3 | The Surgical Aspirator must be suitable for the following theatre procedures:
3.3.1 | Neurosurgery
3.3.2 | Gastro-intestinal and affiliated organs surgery
3.3.3 | Urology
3.3.4 | Plastic & reconstructive surgery
3.3.5 | General surgery
3.3.6 | Orthopaedic surgery
3.3.7 | Gynaecological surgery
3.3.8 | Thoracic surgery
3.3.9 | Laparoscopic surgery
3.3.10 | Thorascopic surgery
3.4 | Both audible and visual alarms must be available on the following:
3.4.1 | Occlusion
3.4.2 | Hand piece disconnections
3.4.3 | Power failure
3.4.4 | Safety features
3.5 | The Surgical Aspirator must have a built-in cooling fan for the electronics in the unit.
4 | STANDARD ACCESSORIES
4.1 | It must be supplied with a total of 12 tips comprising of:
4.1.1 | Shear tip
4.1.2 | Micro tip
4.1.3 | Micro tip plus
4.2 | Suction cannister
4.3 | Suction tubing
4.4 | Intenal vacuum filter
4.5 | External vaccum filter
4.6 | Cleaning brush
4.7 | foot switch
5 | Conformity Compliance (Please attach certificate)
5.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
5.2 | OEM must comply and certified on ISO 9001 quality standards, attach proof of compliance
5.2.1 | Please provide unique ref number of the ISO 9001 ceriticate:
5.3 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
5.3.1 | Please provide unique ref number of the ISO 13485 ceriticate:
5.4 | Model quoted for must be EC certified. Attach a copy of certification
5.4.1 | Please provide unique ref number of the EC ceriticate:
5.5 | Hazardous Substance Act: If this type of equipment / apparatus appears on the schedule of Hazardous substances issued by the Directorate: Health Technology of the Department of Health, a license in terms of the Act on Hazardous Substances (Act 15/1973) must be submitted with the bid / quotation documents. The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.6 | All electrical/electronic medical equipment must be licensed by Radiation Control, where applicable a copy of the license must be submitted The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.6.1 | Radiation Control License number:
5.7 | The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade (rubber to with 2 screws) type and it must be a minimum length of three (3) meters. NB the mains cable of the unit tendered for must be SABS colour coded.
5.8 | The equipment quoted for must be protected against electromagnetic interference IEC 60601-1-2
6 | COMPREHENSIVE MAINTENANCE
6.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) must be included in the bid price
6.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
6.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
6.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
6.3 | The bidder to provide the OEM recommended service interval of the unit:
6.4 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
7 | General
7.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
7.2 | The unit on offer must be the latest technology.
ULTRASONIC SURGICAL ASPIRATOR- Accessories and Conusmables
RT2-09-03a | Shear tip
RT2-09-03b | Micro tip
RT2-09-03c | Micro tip plus
RT2-09-03d | Suction cannister
RT2-09-03e | Suction tubing
RT2-09-03f | Intenal vacuum filter
RT2-09-03g | External vaccum filter
RT2-09-03h | Cleaning brush
RT2-09-03i | foot switch

## Neuronavigation system
RT2-09-04 | NEURONAVIGATION SYSTEM
1 | DESCRIPTION
1.1 | An intra-operative brain mapping navigation system specifically designed for neurosurgery and not a general navigation system
2 | Product Offered:
2.1 | Make:
2.2 | Model:
2.3 | Initial Manufacturing year of Model:
2.4 | Radiation Control License number:
3 | TECHNICAL SPECIFICATION
3.1 | The system must offer the following:
3.1.1 | High performance Image-Guided-Surgery station.
3.1.2 | Must be integrated in a mobile display cart with separate mobile camera cart for flexible positioning of the camera
3.1.3 | Must feature advanced optical, wireless passive marker tracking technology for pre-operative planning and surgical visualization.
3.1.4 | Must be supplied with a UPS.
3.1.5 | Must offer optical and Electromagnetic (EM) tracking capabilities, and integrates with external devices like microscopes, ultrasound, and a broad array of instruments.
3.2 | The neuronavigation system must be used for all types of emergency and scheduled cranial surgeries, Deep Brain Stimulation (DBS) surgeries, microscopic procedures and scheduled spinal surgeries.
3.3 | The neuronavigation system must have the ability to relate the position of a real surgical instrument in the surgeon's hand or the microscope's focal point to the location of the imaged pathology, updated in real time in an integrated operating room, highlights the modern version of this set of technologies.
3.4 | The unit must be able to vizualize the scenario for surgery in a 3-D model of manipulable computer data. In this way the physician can "practice and check" the surgery, try alternative approaches, assess possible difficulties, etc., before the real surgery takes place.
3.5 | The System must include the following:
3.5.1 | Cranial Surgery software and hardware.
3.5.2 | Complete package of DBS software and hardware (enhanced target localization by audiovisual integration of physiological signal, simultaneous display of recording position, advanced 3-D modeling capabilities and display of three orthogonal views).
3.5.3 | Image merging software to merge CT, MRI and Microscope interface.
3.5.4 | 3-D Navigation and volume rendering software for brain and spine
3.5.5 | Spine surgery software and hardware allowing seamless interface to a variety of C-arm Images.
3.6 | The System must be supplied with the following features:
3.6.1 | Neuronavigation systems must offer accurate planning, targeting, and monitoring solutions for Transcranial Magnetic Stimulation (TMS) and/or Transcranial Electrical Stimulation (tES) studies, in addition to 3-D localisation of Neuroimaging systems, such as Electroencephalogram (EEG) and Near-Infrared Spectroscopy (NIRS).
3.6.2 | Neuroimaging devices to offer researchers an ability to view the structure, function and pharmacology of the brain.
3.6.3 | Neurosensory devices for use in neuroscience research.
3.6.4 | Neuromodulation devices to enhance or suppress activities of the brain or nervous system in the research of brain function and disease.
3.6.5 | Monitor with the following minimum specifications:
3.6.5.1 | Built in touch screen monitor 24 inch HD with minimum 1920 x 1200 pixels resolution.
3.6.5.2 | Must accept DICOM format Images from intra-operative MRI, intra-operative CT, and C-Arm.
3.6.5.3 | Must accept standard DICOM format Images from MRI and CT Scanning Departments via CD/USB enabling surgeon to plan surgery in advance on main navigation system.
3.6.5.4 | Surgeon must have control of the navigation system through touch screen and patient frame without compromising the sterility
3.7 | Mandatory accessories to be supplied with the System:
3.7.1 | Small, medium and large passive trackers
3.7.2 | Small, medium and large mounts
3.7.3 | Mount driver
3.7.4 | Sterilization tray
3.7.5 | Cranial biopsy instrument to be supplied with the following:
3.7.5.1 | Articulated arm
3.7.5.2 | Mount, double star adapter mayfield
3.7.5.3 | Biopsy guide
3.7.5.4 | Probe
3.8 | Microscope integration interface only
3.8.1 | Display of the microscope focal point on anatomical images
3.8.2 | Display of Region of interest (ROI) and Navigation symbols in oculars
3.8.3 | Display of navigation monitor in oculars 1x tracking hardware (cables & bracket)
3.8.4 | Tracking interface hardware
3.9 | The system must be capable of networking from modular operation theatre to CT Room and Neurosurgery ward
3.10 | Planning station with the following:
3.10.1 | Separate, latest and advanced computer for planning of navigated surgical procedures with the following minimum specifications:
3.10.1.1 | ≥8GB Ram, ≥256 SSD, ≥17 inch LED screen, the LED screen and Core i7 CPU must be integrated in one casing, software for processing all captured images, a cordless keyboard and mouse, with windows 10 or upper version
3.10.1.2 | Preloaded planning software for brain and spine
4 | Standard Accessories and Consumables
4.1 | Small, medium and large passive trackers
4.2 | Small, medium and large mounts
4.3 | Mount driver
4.4 | Sterilization tray
5 | Conformity Compliance (Please attach certificate)
5.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
5.2 | OEM must comply and certified on ISO 9001 quality standards, attach proof of compliance
5.2.1 | Please provide unique ref number of the ISO 9001 ceriticate:
5.3 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
5.3.1 | Please provide unique ref number of the ISO 13485 ceriticate:
5.4 | Model quoted for must be EC certified. Attach a copy of certification
5.4.1 | Please provide unique ref number of the EC ceriticate:
5.5 | Hazardous Substance Act: If this type of equipment / apparatus appears on the schedule of Hazardous substances issued by the Directorate: Health Technology of the Department of Health, a license in terms of the Act on Hazardous Substances (Act 15/1973) must be submitted with the bid / quotation documents. The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.6 | All electrical/electronic medical equipment must be licensed by Radiation Control, where applicable a copy of the license must be submitted The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.6.1 | Radiation Control License number:
5.7 | The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade (rubber to with 2 screws) type and it must be a minimum length of three (3) meters. NB the mains cable of the unit tendered for must be SABS colour coded.
5.8 | The equipment quoted for must be protected against electromagnetic interference IEC 60601-1-2
6 | COMPREHENSIVE MAINTENANCE
6.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) must be included in the bid price
6.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
6.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
6.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
6.3 | The bidder to provide the OEM recommended service interval of the unit:
6.4 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
7 | General
7.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
7.2 | The unit on offer must be the latest technology.

## Microscope, Neurosurgery
RT2-09-05 | MICROSCOPE, NEUROSURGERY, MOBILE
1 | DESCRIPTION
1.1 | A surgical microscope suitable for neurosurgery and spinal surgery with an 540 degrees variable angle binocular tube configuration with foot pedal and an external monitor. The microscope must have image storage and a 3 CCD high definition camera.
1.2 | An electronic surgical microscope with full display monitor, image storage and a 3 CCD high definition camera.
2 | Product Offered:
2.1 | Make:
2.2 | Model:
2.3 | Initial Manufacturing year of Model:
2.4 | Radiation Control License number:
3 | TECHNICAL SPECIFICATION
3.1 | The unit offered should include focusing, zooming and magnification controls.
3.2 | Magnification must be atleast 6 step with 1.6 to 20 x range
3.3 | Optics: Apochromatic with magnification changer
3.4 | Stereo base: 22 - 25mm
3.5 | The trolley should have antistatic castors with at least 2 equipped with locking system/brake.
3.6 | Eyepiece:
3.6.1 | · 10x wide-field eyepieces
3.6.2 | · 12.5x wide-field eyepieces
3.6.3 | · Dioptric setting ±5 with adjustable eye cup
3.7 | Lateral tilt: -30 degrees - 50 degrees
3.8 | Inclination angle: -30 degrees - 120 degrees
3.9 | Illumination: two LED or Xenon lamps source with emergency back-up lamp.
3.10 | Field of view: 16.5 – 210 mm
3.11 | Focal length ranging from 150 - 400mm
3.12 | Protection filter: UV & IR filter
3.13 | A floor stand shall have atleast 4 castors to ensure base stability and a minimum of 2 brakes on castors for smooth operation of the microscope.
3.14 | The microscope stand must have a stand mounted grip for easy movement of the unit
3.15 | The maximum load of the accessories to the microscope should be atleast 5kg
3.16 | The base of the stand shall be compact and have 360 degrees rotatable wheels to allow easy positioning of the microscope.
3.17 | The foot switch should be wireless or wired and waterproof
3.18 | The microscope offered should have pedal and/or hand-switch controls for ease of operation
3.19 | The pedal must be waterproof and inclinable, with a minimum of 10 programmable functions to control all the microscope's functions.
3.20 | The surface of the microscope shall be anti-bacterial and have anti-mycotic protection.
3.21 | The optic carrier shall have an electric system to balance the three gravitational angles.
3.22 | The microscope shall have a function to automatically reduce the light intensity at a short distance during examination
3.23 | Handles must be autoclave-able
3.24 | The microscope offered shall have an integrated monitor, of at least 21" with HD resolution.
3.25 | The microscope should have binocular tubes for the main surgeon and possible for the assistant (as and when needed)
3.26 | The control module shall have a screen with touch pad to select different user profiles and settings.
3.27 | Surgeons shall be able to store their personal preferences for magnification, focus and illumination.
4 | Standard Accessories and Consumables
4.1 | Binocular tubes
4.2 | Autoclavable handle
4.3 | Eyepiece cups
4.4 | Wireless or wired foot pedal
4.5 | Dustcover
4.6 | Second observer
4.7 | Video adapter
4.8 | Objective lense: 200mm - 600mm
5 | Conformity Compliance (Please attach certificate)
5.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
5.2 | OEM must comply and certified on ISO 9001 quality standards, attach proof of compliance
5.2.1 | Please provide unique ref number of the ISO 9001 ceriticate:
5.3 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
5.3.1 | Please provide unique ref number of the ISO 13485 ceriticate:
5.4 | Model quoted for must be EC certified. Attach a copy of certification
5.4.1 | Please provide unique ref number of the EC ceriticate:
5.5 | Hazardous Substance Act: If this type of equipment / apparatus appears on the schedule of Hazardous substances issued by the Directorate: Health Technology of the Department of Health, a license in terms of the Act on Hazardous Substances (Act 15/1973) must be submitted with the bid / quotation documents. The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.6 | All electrical/electronic medical equipment must be licensed by Radiation Control, where applicable a copy of the license must be submitted The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.6.1 | Radiation Control License number:
5.7 | The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade (rubber to with 2 screws) type and it must be a minimum length of three (3) meters. NB the mains cable of the unit tendered for must be SABS colour coded.
5.8 | The equipment quoted for must be protected against electromagnetic interference IEC 60601-1-2
6 | COMPREHENSIVE MAINTENANCE
6.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) must be included in the bid price
6.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
6.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
6.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
6.3 | The bidder to provide the OEM recommended service interval of the unit:
6.4 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
7 | General
7.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
7.2 | The unit on offer must be the latest technology.
MICROSCOPE, NEUROSURGERY, MOBILE - Accessories and Consumables
RT2-09-05a | Binocular tubes
RT2-09-05b | Autoclavable handle
RT2-09-05c | Eyepiece cups
RT2-09-05d | Wireless or wired foot pedal
RT2-09-05e | Dustcover
RT2-09-05f | Second observer
RT2-09-05g | Video adapter
RT2-09-05h | Objective lense: 200mm - 600mm
RT2-09-05i | Xenon illumination system
RT2-09-05j | Stereo co-observation tube/Teaching piece

## Microscope, Neuro, Advanced
RT2-09-06 | OPERATING MICROSCOPE SYSTEM FOR NEURO-SURGERY
1 | DESCRIPTION
1.1 | A surgical microscope suitable for neurosurgery and spinal surgery with an 540 degrees variable angle binocular tube configuration with foot pedal and an external monitor. The microscope must have image storage and a 3 CCD high definition camera.
1.2 | An electronic surgical microscope with full display monitor, image storage and a 3 CCD high definition camera.
2 | Product Offered:
2.1 | Make:
2.2 | Model:
2.3 | Initial Manufacturing year of Model:
2.4 | Radiation Control License number:
3 | TECHNICAL SPECIFICATION
3.1 | The surgical microscope must be specifically designed for neurosurgery and not a general microscope
3.2 | Floor Stand Mobile Operating Microscope customized for Neurosurgical Procedures
3.3 | Microscope degree of freedom must allow work in difficult angles to reach areas of interest
3.4 | Microscope robotics must have at least 6 axis of movement
3.5 | MAGNIFICATION SYSTEM
3.5.1 | The system shall have a motorized zoom; 1:6 zoom ratio
3.5.2 | The magnification shall be displayed by means of a display
3.5.3 | The system shall have a user specific start position
3.5.4 | 10x to 12.5x magnetic widefield eyepieces with intergrated eyecups
3.5.5 | tubes that allow rapid switch between large magnification and maximum overview of the region of interest must be available
3.5.6 | magnification speed is adjustable
3.6 | FOCUSING SYSTEM
3.6.1 | Autofocus with 2 visible laser spots
3.6.2 | The focusing system shall have a motorised working range of at least 225mm to 575mm
3.6.3 | The system shall have a internal motorised continuos adjustment
3.6.4 | The magnification shall be linked to the adjustment of the focus speed
3.6.5 | The system shall have a high-speed auto focus, with an accuracy of ± 0.5 mm
3.6.6 | The system must be able to electronically adjust the depth of field
3.6.7 | the microscope must be able to offer integrated tube rotation for better ergonomic observation
3.6.8 | focus speed is adjustable
3.7 | ILLUMINATION
3.7.1 | At least 2x300W xenon lamps
3.7.2 | illumination field diameter able to vary automatically
3.7.3 | Display of remaining lamp life on Touch screen
3.7.4 | adjustable light threshold settings
3.8 | SYSTEM OPERATION
3.8.1 | Multifunctional programmable hand grips
3.8.2 | central user interface
3.8.3 | automatic microscope balancing must happen in any position of the microscopes'tilt and swivel range and entire microscope must balance all system axes without any manual input or axis adjustment
3.9 | TUBES AND CO-OBSERVATION
3.9.1 | must include stereo-co-observation tube
3.9.2 | Stereo co-observation tube shall remain fixed when tilting microscope head
3.9.3 | The main tube shall be 0° - 180° rotatable
3.9.4 | face to face surgery must be possible with both surgeon and assistant in a standing position
3.10 | SYSTEM SETUP
3.10.1 | Autobalance of microscope or entire system shall be by means of pushing a button
3.10.2 | offered system must be draped by means of a integrated vacuum system to remove air from sterile drape for fast and easy draping
3.10.3 | Mouth switch for fine balance must be available
3.10.4 | system must have a central touch screen user interface that programs and allows for operation of the entire system that includes the microscope, patient data management, video camera and recording
3.10.5 | surgeon must be able to see parameters such as magnification, working distance and light intesity within the field of view of the occulars
3.11 | VIDEO
3.11.1 | At least HD Video Camera Unit with recording and storing capability in both 2D and 3D
3.11.2 | At least 22 inch High Definition video touchscreen on arm extension to allow visualisation of microscope image
3.11.3 | integrated video still image capturing on HDD and USB stick
3.11.4 | must be possible to record still images from the inegrated HD video camera without requiring additional recording devices
3.12 | CONNECTIVITY
3.12.1 | Video-in for external SD video sources
3.12.2 | Microscope interface integration with endoscope and neuronavigation
3.12.3 | DICOM module for data transfer from or to PACS system
3.13 | INTRAOPERATIVE FLUORESCENCE AND ICG
3.13.1 | the system must offer an option for fluoresecne guided glioma resection
3.13.2 | system must offer fluorescence based video angiography
3.13.3 | system must have an automatic analysis tool for fluoresecne video angiography
3.13.4 | System must offer automated setup for focus, zoom and light for optimum fluoroscene visualisation
3.13.5 | system must allow parallel automatic recording of floroscene video and standard white light video
3.13.6 | fluroscene camera system must have manual and automatic gain control for adjusting fluoroscene image quality
3.14 | ELECTRICAL DATA
3.14.1 | The Voltage supply shall be 230V (Range: 220-240V)
3.14.2 | The frequency shall be 50-60 Hz
4 | Standard Accessories and Consumables
4.1 | Dust Cover
5 | Conformity Compliance (Please attach certificate)
5.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
5.2 | OEM must comply and certified on ISO 9001 quality standards, attach proof of compliance
5.2.1 | Please provide unique ref number of the ISO 9001 ceriticate:
5.3 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
5.3.1 | Please provide unique ref number of the ISO 13485 ceriticate:
5.4 | Model quoted for must be EC certified. Attach a copy of certification
5.4.1 | Please provide unique ref number of the EC ceriticate:
5.5 | Hazardous Substance Act: If this type of equipment / apparatus appears on the schedule of Hazardous substances issued by the Directorate: Health Technology of the Department of Health, a license in terms of the Act on Hazardous Substances (Act 15/1973) must be submitted with the bid / quotation documents. The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.6 | All electrical/electronic medical equipment must be licensed by Radiation Control, where applicable a copy of the license must be submitted The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.6.1 | Radiation Control License number:
5.7 | The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade (rubber to with 2 screws) type and it must be a minimum length of three (3) meters. NB the mains cable of the unit tendered for must be SABS colour coded.
5.8 | The equipment quoted for must be protected against electromagnetic interference IEC 60601-1-2
6 | COMPREHENSIVE MAINTENANCE
6.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) must be included in the bid price
6.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
6.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
6.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
6.3 | The bidder to provide the OEM recommended service interval of the unit:
6.4 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
7 | General
7.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
7.2 | The unit on offer must be the latest technology.

## Category 10 - Sterilization Equ
STERILIZATION & DISINFECTION EQUIPMENT (Sterilizers & Instrument Washers)
# | Item number | Item Description
1 | RT2-10-01 | Autoclave, Tabletop, 24-30L
2 | RT2-10-02 | Autoclave, Tabletop, 32-38L
3 | RT2-10-03 | Autoclave, Tabletop, 40-50L
4 | RT2-10-04 | Ultrasonic Washer (30L)
5 | RT2-10-05 | Heat Sealer
6 | RT2-10-06 | Plasma Sterilizer 35-50L
7 | RT2-10-07 | Plasma Sterilizer 85-110L
8 | RT2-10-08 | Plasma Sterilizer 145-160L
9 | RT2-10-09 | Endoscope Washer

## Autoclave, Tabletop, 24-30L
RT2-10-01 | AUTOCLAVE - TABLETOP 24-30 liters
1 | DESCRIPTION
1.1 | A strong sealed chamber in which surgical instruments, towels, dressings, and other items can be sterilized by steam under raised pressure
2 | Product Offered:
2.1 | Make:
2.2 | Model:
2.3 | Initial Manufacturing year of Model:
3 | TECHNICAL SPECIFICATION
3.1 | The autoclave must be fully automatic with a computerized control system, which monitors all the physical parameters and must be equipped with a built-in printer to provide a traceable reference documentation of all the sterilization parameters.
3.2 | The chamber volume must be at least 24 to 30 Litre capacity.
3.3 | The autoclave must be fitted with a high vacuum silent running vacuum pump. To ensure perfect pre and post vacuum conditions. (SPLIT vacuum).
3.4 | CLEAN WATER SUPPLY: An external water supply or an internal RESERVOIR of sufficient capacity will be acceptable.Please state capacity of RESERVOIR.
3.5 | WASTE WATER DRAINS: An external waste water drain or an internal RESERVOIR of sufficient capacity will be acceptable.Please state capacity of RESERVOIR.
3.6 | The autoclave offered must be capable of performing the following sterilization cycles: Six sterilization cycles i.e. C1-wrapped,C2-wrapped,C3-unwrapped,C4-unwrapped, C5-Prion cycle (Creutzfeldt-Jakob disease), C6-Flash Cycle.
3.7 | The autoclave offered must be capable of performing the following 3 test cycles:
3.7.1 | 1)Vacuum Test: Test to check if the pressure is correct.
3.7.2 | 2) Bowie & Dick Test: Test to check if the steam correctly penetrates a porous load.
3.7.3 | 3) Helix Test: Test to check if the steam correctly penetrates a hollow load.
3.7.4 | Sterilizing temperature: C1-134⁰ , C2-121⁰, C3-134⁰, C4-121⁰, C5-134⁰, C6-134⁰
3.7.5 | Programmable cycles : Wrapped & Unwrapped automatic.
3.7.6 | Temperature range : 100- 134⁰ with a heat shutoff function.
3.7.7 | Controls should be sealed against fluid penetration.
3.7.8 | Recorder/Printer
3.7.9 | Bidder to state sterilazation time and drying time offered.
3.7.10 | The unit should have chamber temperature monitoring, accurate to ± 1⁰C over the operating range.
3.7.11 | A suitable circuit interrupter to match the load must be fitted.
3.7.12 | Power consumption not to exceed 3,5kW (16 Amps).
3.7.13 | Supply voltage 230 V.A.C. 50Hz.
3.8 | Please specify other safety devices incorporated in the autoclave offered.
3.9 | The autoclave’s drum trays and door must be constructed from stainless steel. (State grade of stainless steel used).
3.10 | State the maximum safe working pressure of the autoclave (in Kpa).
3.11 | State the test pressure of the autoclave drum.
4.9 | The body of the autoclave must be finished off in a baked enamel or be epoxy powder coated.
3.12 | A bacteria filter must be fitted to the vacuum break circuit.
3.19 | Alarms
3.19.1 | Water supply failure
3.19.2 | Cycle failure
3.20 | Safety
3.20.1 | Pressure safety valve.
3.20.2 | Door safety catch.
3.20.3 | Error display warning indicator.
3.20.4 | Overheat protection with warning indicator.
3.20.5 | Door interlock switch.
3.20.6 | Pressure door lock.
3.21 | Water Conditioning System
3.21.1 | The bidder must indicate whether they can offer a water conditioning system for the autoclave for example: An R.O. system, Filter system, Water distiller.
4 | STANDARD ACCESSORIES and CONSUMABLES
4.1 | DIN tray
4.2 | Bacteria filters
4.3 | Printing paper
5 | Conformity Compliance (Please attach certificate)
5.1 | The unit must comply with an acceptable international safety standard such as EN 61010-2-041 for Medical Autoclaves, attached certification
5.1.1 | Please provide unique ref number of the EN 61010-2-041 ceriticate:
5.2 | OEM must comply and certified on ISO 4126-1 quality standards, attach proof of compliance
5.2.1 | Please provide unique ref number of the ISO 4126-1 ceriticate:
5.3 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
5.3.1 | Please provide unique ref number of the ISO 13485 ceriticate:
5.4 | Model quoted for must be EC certified. Attach a copy of certification
5.4.1 | Please provide unique ref number of the EC ceriticate:
5.5 | Model quoted for must comply with EN 50081-2 and EN 50082-2. Attach a copy of certification
5.5.1 | Please provide unique ref number of the EN 50081-2 and EN 50082-2 ceriticate:
5.6 | Model quoted for must comply with EN 13060/1-4. Attach a copy of certification
5.6.1 | Please provide unique ref number of the EN 13060/1-4 ceriticate:
5.7 | Local Manufacturers must provide certification that unit bidded for comply with SANS 982: 2009
6 | COMPREHENSIVE MAINTENANCE
6.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) must be included in the bid price
6.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
6.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
6.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
6.3 | The bidder to provide the OEM recommended service interval of the unit:
6.4 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
7 | General
7.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
7.2 | The unit on offer must be the latest technology.
AUTOCLAVE - TABLETOP 24-30 liters - Accessories and Consumables
RT2-10-01a | DIN tray
RT2-10-01b | Bacteria filters
RT2-10-01c | Printing paper

## Autoclave, Tabletop, 32-38L
RT2-10-02 | AUTOCLAVE - TABLETOP 32-38 liters
1 | DESCRIPTION
1.1 | A strong sealed chamber in which surgical instruments, towels, dressings, and other items can be sterilized by steam under raisedpressure
2 | Product Offered:
2.1 | Make:
2.2 | Model:
2.3 | Initial Manufacturing year of Model:
3 | TECHNICAL SPECIFICATION
3.1 | The autoclave must be fully automatic with a computerized control system, which monitors all the physical parameters and must be equipped with a built-in printer to provide permanent documentation of all the sterilization parameters.
3.2 | The chamber volume must be at least 32 to 38 Litre capacity.
3.3 | The autoclave must be fitted with a high vacuum silent running vacuum pump. To ensure perfect pre and post vacuum conditions. (SPLIT vacuum).
3.4 | CLEAN WATER SUPPLY: An external water supply or an internal RESERVOIR of sufficient capacity will be acceptable.Please state capacity of RESERVOIR.
3.5 | WASTE WATER DRAINS: An external waste water drain or an internal RESERVOIR of sufficient capacity will be acceptable.Please state capacity of RESERVOIR.
3.6 | The autoclave offered must be capable of performing the following sterilization cycles: Six sterilization cycles i.e. C1-wrapped,C2-wrapped,C3-unwrapped,C4-unwrapped, C5-Prion cycle (Creutzfeldt-Jakob disease), C6-Flash Cycle.
3.7 | The autoclave offered must be capable of performing the following 3 test cycles:
3.7.1 | 1)Vacuum Test: Test to check if the pressure is correct.
3.7.2 | 2) Bowie & Dick Test: Test to check if the steam correctly penetrates a porous load.
3.7.3 | 3) Helix Test: Test to check if the steam correctly penetrates a hollow load.
3.7.4 | Sterilizing temperature: C1-134⁰ , C2-121⁰, C3-134⁰, C4-121⁰, C5-134⁰, C6-134⁰
3.7.5 | Programmable cycles : Wrapped & Unwrapped automatic.
3.7.6 | Temperature range : 100- 134⁰ with a heat shutoff function.
3.7.7 | Controls should be sealed against fluid penetration.
3.7.8 | Recorder/Printer
3.7.9 | Bidder to state sterilazation time and drying time offered.
3.7.10 | The unit should have chamber temperature monitoring, accurate to ± 1⁰C over the operating range.
3.7.11 | A suitable circuit interrupter to match the load must be fitted.
3.7.12 | Power consumption not to exceed 3,5kW (16 Amps).
3.7.13 | Supply voltage 230 V.A.C. 50Hz.
3.8 | Please specify other safety devices incorporated in the autoclave offered.
3.9 | The autoclave’s drum trays and door must be constructed from stainless steel. (State grade of stainless steel used).
3.10 | State the maximum safe working pressure of the autoclave (in Kpa).
3.11 | State the test pressure of the autoclave drum.
4.9 | The body of the autoclave must be finished off in a baked enamel or be epoxy powder coated.
3.12 | A bacteria filter must be fitted to the vacuum break circuit.
3.19 | Alarms
3.19.1 | Water supply failure
3.19.2 | Cycle failure
3.20 | Safety
3.20.1 | Pressure safety valve.
3.20.2 | Door safety catch.
3.20.3 | Error display warning indicator.
3.20.4 | Overheat protection with warning indicator.
3.20.5 | Door interlock switch.
3.20.6 | Pressure door lock.
3.21 | Water Conditioning System
3.21.1 | The bidder must indicate whether they can offer a water conditioning system for the autoclave for example: An R.O. system, Filter system, Water distiller.
4 | STANDARD ACCESSORIES and CONSUMABLES
4.1 | DIN tray
4.2 | Bacteria filters
4.3 | Printing paper
5 | Conformity Compliance (Please attach certificate)
5.1 | The unit must comply with an acceptable international safety standard such as EN 61010-2-041 for Medical Autoclaves, attached certification
5.1.1 | Please provide unique ref number of the EN 61010-2-041 ceriticate:
5.2 | OEM must comply and certified on ISO 4126-1 quality standards, attach proof of compliance
5.2.1 | Please provide unique ref number of the ISO 4126-1 ceriticate:
5.3 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
5.3.1 | Please provide unique ref number of the ISO 13485 ceriticate:
5.4 | Model quoted for must be EC certified. Attach a copy of certification
5.4.1 | Please provide unique ref number of the EC ceriticate:
5.5 | Model quoted for must comply with EN 50081-2 and EN 50082-2. Attach a copy of certification
5.5.1 | Please provide unique ref number of the EN 50081-2 and EN 50082-2 ceriticate:
5.6 | Model quoted for must comply with EN 13060/1-4. Attach a copy of certification
5.6.1 | Please provide unique ref number of the EN 13060/1-4 ceriticate:
5.7 | Local Manufacturers must provide certification that unit bidded for comply with SANS 982: 2009
6 | COMPREHENSIVE MAINTENANCE
6.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) must be included in the bid price
6.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
6.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
6.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
6.3 | The bidder to provide the OEM recommended service interval of the unit:
6.4 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
7 | General
7.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
7.2 | The unit on offer must be the latest technology.
AUTOCLAVE - TABLETOP 32-38 liters - Accessories and Consumables
RT2-10-02a | DIN tray
RT2-10-02b | Bacteria filters
RT2-10-02c | Printing paper

## Autoclave, Tabletop, 40-50L
RT2-10-03 | AUTOCLAVE - TABLETOP 40-50 liters
1 | DESCRIPTION
1.1 | A strong sealed chamber in which surgical instruments, towels, dressings, and other items can be sterilized by steam under raisedpressure
2 | Product Offered:
2.1 | Make:
2.2 | Model:
2.3 | Initial Manufacturing year of Model:
3 | TECHNICAL SPECIFICATION
3.1 | The autoclave must be fully automatic with a computerized control system, which monitors all the physical parameters and must be equipped with a built-in printer to provide permanent documentation of all the sterilization parameters.
3.2 | The chamber volume must be at least 40 to 50 Litre capacity.
3.3 | The autoclave must be fitted with a high vacuum silent running vacuum pump. To ensure perfect pre and post vacuum conditions. (SPLIT vacuum).
3.4 | CLEAN WATER SUPPLY: An external water supply or an internal RESERVOIR of sufficient capacity will be acceptable.Please state capacity of RESERVOIR.
3.5 | WASTE WATER DRAINS: An external waste water drain or an internal RESERVOIR of sufficient capacity will be acceptable.Please state capacity of RESERVOIR.
3.6 | The autoclave offered must be capable of performing the following sterilization cycles: Six sterilization cycles i.e. C1-wrapped,C2-wrapped,C3-unwrapped,C4-unwrapped, C5-Prion cycle (Creutzfeldt-Jakob disease), C6-Flash Cycle.
3.7 | The autoclave offered must be capable of performing the following 3 test cycles:
3.7.1 | 1)Vacuum Test: Test to check if the pressure is correct.
3.7.2 | 2) Bowie & Dick Test: Test to check if the steam correctly penetrates a porous load.
3.7.3 | 3) Helix Test: Test to check if the steam correctly penetrates a hollow load.
3.7.4 | Sterilizing temperature: C1-134⁰ , C2-121⁰, C3-134⁰, C4-121⁰, C5-134⁰, C6-134⁰
3.7.5 | Programmable cycles : Wrapped & Unwrapped automatic.
3.7.6 | Temperature range : 100- 134⁰ with a heat shutoff function.
3.7.7 | Controls should be sealed against fluid penetration.
3.7.8 | Recorder/Printer
3.7.9 | Bidder to state sterilazation time and drying time offered.
3.7.10 | The unit should have chamber temperature monitoring, accurate to ± 1⁰C over the operating range.
3.7.11 | A suitable circuit interrupter to match the load must be fitted.
3.7.12 | Power consumption not to exceed 3,5kW (16 Amps).
3.7.13 | Supply voltage 230 V.A.C. 50Hz.
3.8 | Please specify other safety devices incorporated in the autoclave offered.
3.9 | The autoclave’s drum trays and door must be constructed from stainless steel. (State grade of stainless steel used).
3.10 | State the maximum safe working pressure of the autoclave (in Kpa).
3.11 | State the test pressure of the autoclave drum.
4.9 | The body of the autoclave must be finished off in a baked enamel or be epoxy powder coated.
3.12 | A bacteria filter must be fitted to the vacuum break circuit.
3.19 | Alarms
3.19.1 | Water supply failure
3.19.2 | Cycle failure
3.20 | SAFETY
3.20.1 | Pressure safety valve.
3.20.2 | Door safety catch.
3.20.3 | Error display warning indicator.
3.20.4 | Overheat protection with warning indicator.
3.20.5 | Door interlock switch.
3.20.6 | Pressure door lock.
3.21 | Water Conditioning System
3.21.1 | The bidder must indicate whether they can offer a water conditioning system for the autoclave for example: An R.O. system, Filter system, Water distiller.
4 | STANDARD ACCESSORIES and CONSUMABLES
4.1 | DIN tray
4.2 | Bacteria filters
4.3 | Printing paper
5 | Conformity Compliance (Please attach certificate)
5.1 | The unit must comply with an acceptable international safety standard such as EN 61010-2-041 for Medical Autoclaves, attached certification
5.1.1 | Please provide unique ref number of the EN 61010-2-041 ceriticate:
5.2 | OEM must comply and certified on ISO 4126-1 quality standards, attach proof of compliance
5.2.1 | Please provide unique ref number of the ISO 4126-1 ceriticate:
5.3 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
5.3.1 | Please provide unique ref number of the ISO 13485 ceriticate:
5.4 | Model quoted for must be EC certified. Attach a copy of certification
5.4.1 | Please provide unique ref number of the EC ceriticate:
5.5 | Model quoted for must comply with EN 50081-2 and EN 50082-2. Attach a copy of certification
5.5.1 | Please provide unique ref number of the EN 50081-2 and EN 50082-2 ceriticate:
5.6 | Model quoted for must comply with EN 13060/1-4. Attach a copy of certification
5.6.1 | Please provide unique ref number of the EN 13060/1-4 ceriticate:
5.7 | Local Manufacturers must provide certification that unit bidded for comply with SANS 982: 2009
6 | COMPREHENSIVE MAINTENANCE
6.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) must be included in the bid price
6.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
6.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
6.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
6.3 | The bidder to provide the OEM recommended service interval of the unit:
6.4 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
7 | General
7.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
7.2 | The unit on offer must be the latest technology.
AUTOCLAVE - TABLETOP 40-50 liters -Accessories and Consumables
RT2-10-03a | DIN tray
RT2-10-03b | Bacteria filters
RT2-10-03c | Printing paper

## Ultrasonic cleaner 30L
RT2-10-04 | ULTRASONIC CLEANER
1 | SCOPE
1.1 | This specification establishes the requirements for the supply, delivery, installation, demonstration and commissioning of an ultrasonic cleaner for use in theatres and CSSD's
2 | Product Offered:
2.1 | Make:
2.2 | Model:
2.3 | Initial Manufacturing year of Model:
3 | TECHNICAL SPECIFICATION
3.1 | Intended use: The device will be used for pre-cleaning of instruments, especially hollow instruments, and other necessities before sterilization.
3.2 | The ultrasonic cleaner shall be used for the cleaning of tough dirt, blood, bacteria etc on instruments used during surgical procedures.
3.3 | The unit shall be equipped with a temperature controller to set water temperature up to approximately 70 degrees. (Although it is not advisable to heat solution in medical cleaning, to prevent protein coagulation, therefore it shall also be possible to use with solution at ambient temperature)
3.4 | The unit shall be equipped with a mechanical or electronic timer to enable setting of ultrasonic output time from 0 to 20 minutes.
3.5 | The ultrasound frequency shall be approximately 42 KHz. Bidders to state ultrasound frequency, as well as the advantage of the frequency the unit operates at.
3.6 | The tank internal capacity shall be approximately 30 litres and the size approximately 600 X 300 X 200 (LXWXH) – Longest instrument is 580mm, therefore the requirement for length. The closest offer that can satisfy the requirement for the long instruments will be favourably taken into account.
3.7 | The offered unit shall have an ultrasonic output power of approximately 500Watt.
3.8 | The unit shall have a wire basket and a lid.
3.9 | The unit shall be manufactured of good quality stainless steel 304L, and have a smooth, polished surface.
3.10 | The unit shall be equipped with a bottom mounted drain tap to enable draining of solution for cleaning purposes.
4 | Standard Accessories
4.1 | Wire basket
4.2 | Lid
5 | Conformity Compliance (Please attach certificate)
5.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
5.2 | OEM must comply and certified on ISO 9001 quality standards, attach proof of compliance
5.2.1 | Please provide unique ref number of the ISO 9001 ceriticate:
5.3 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
5.3.1 | Please provide unique ref number of the ISO 13485 ceriticate:
5.4 | Model quoted for must be EC certified. Attach a copy of certification
5.4.1 | Please provide unique ref number of the EC ceriticate:
5.5 | The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade (rubber to with 2 screws) type and it must be a minimum length of three (3) meters. NB the mains cable of the unit tendered for must be SABS colour coded.
5.6 | The equipment quoted for must be protected against electromagnetic interference IEC 60601-1-2
6 | COMPREHENSIVE MAINTENANCE
6.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) must be included in the bid price
6.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
7 | General
7.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
7.2 | The unit on offer must be the latest technology.

## Heat sealer
RT2-10-05 | HEAT SEALING MACHINE
1 | SCOPE
1.1 | This specification establishes the requirements for the supply, delivery, installation, demonstration and commissioning of a heat sealing machine
2 | Product Offered:
2.1 | Make:
2.2 | Model:
2.3 | Initial Manufacturing year of Model:
3 | TECHNICAL SPECIFICATION
3.1 | Intended use: The device will be used for packaging of instruments and other necessities before sterilization and maintaining this sterility for a fixed period of time after the sterilization process.
3.2 | The device offered shall be trustworthy, provide a solid seal and operational reliability.
3.3 | The heat sealer shall be both internally and externally made of stainless steel in order to reduce maintenance to a minimum and enhance hygiene and infection control.
3.4 | The device shall be suitable for laminated pouches as well as paper pouches
3.5 | Bidder to state type of seal which the device will produce.
3.6 | The device shall have an automatic transportation start as well as a reverse transportation feature which ensures safety and prevents the instruments from being damaged.
3.7 | An accurate sealing temperature setting by means of a keyboard which has an LCD display with a good visual read out.
3.8 | Temperature shall range between 100-199°C
3.9 | Temperature tolerance: ± 2°C
3.10 | The machine shall have a built-in mechanism to protect it from overheating.
3.11 | It shall be possible to do Seal edge adjustment.
3.12 | Seal width shall at least be 15mm
3.13 | The gradual feeder adjustment shall be 10-35mm
3.14 | Bidder to state the transportation speed (m/min).
3.15 | The device shall be able to self-adjust the pressure needed during the sealing process
3.16 | The device shall be equipped with a production counter with audible alarm
3.17 | The device shall be equipped with a safety measure on the auto insert feed which stops the machine from sealing when there is a temperature difference of 5°C from the set point.
3.18 | The unit shall be controlled by a microprocessor. Bidder to state control mechanism.
3.19 | Bidder to indicate whether the device offers a standby system and specify the type of standby to be provided.
3.20 | The device shall be resettable
3.21 | Dimensions: ± 610 x 180mm
3.22 | Mains voltage: 220/240 VAC
3.23 | Capacity: Max: 330W
3.24 | Weight: ± 12 kg
4 | standard Accessories
4.1 | Front table / roller front table
5 | Conformity Compliance (Please attach certificate)
5.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
5.2 | OEM must comply and certified on ISO 9001 quality standards, attach proof of compliance
5.2.1 | Please provide unique ref number of the ISO 9001 ceriticate:
5.3 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
5.3.1 | Please provide unique ref number of the ISO 13485 ceriticate:
5.4 | Model quoted for must be EC certified. Attach a copy of certification
5.4.1 | Please provide unique ref number of the EC ceriticate:
5.5 | The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade (rubber to with 2 screws) type and it must be a minimum length of three (3) meters. NB the mains cable of the unit tendered for must be SABS colour coded.
5.6 | The equipment quoted for must be protected against electromagnetic interference IEC 60601-1-2
6 | COMPREHENSIVE MAINTENANCE
6.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) must be included in the bid price
6.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
7 | General
7.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
7.2 | The unit on offer must be the latest technology.

## Plasma Sterilizer 35-50L
RT2-10-06 | PLASMA STERILIZER, LOW TEMPERATURE, 35- 50L
1 | DESCRIPTION
1.1 | Low Temperature Sterilizer that deliver rapid and safe sterilization for delicate intruments that are sensitive to heat and moisture
2 | Product Offered:
2.1 | Make:
2.2 | Model:
2.3 | Initial Manufacturing year of Model:
3 | TECHNICAL SPECIFICATION
3.1 | The unit offered must be fully automatic, microprocessor controlled and feature advanced technology for safe low-temperature plasma sterilization.
3.2 | The unit's frames, external panels, valves, piping and internal shelving must be constructed from 316 grade stainless steel material
3.3 | The Plasma Sterilizer shall use hydrogen peroxide (H2O2) as sterilization agent
3.4 | The unit shall have processing cycles, which are pre-programmed with the recommended sterilizing cycles as follows (+5 mins)
3.4.1 | Fast Cycle time: 40 mins
3.4.2 | Standard Cycle time: 50 Mins
3.4.3 | Advanced Cycle time: 60 Mins
3.5 | The sterilization temperature must be within + 45OC to 50OC
3.6 | The unit offered shall have auto interlock and open pressure monitoring system that ensures all chamber pressure has been released prior to allowing the door to open
3.7 | The chamber must be able to withstand pressure from vacuum up to 3.5 bars
3.8 | The sterilizing chamber shall be in stainless steel and rectangular with stainless steel instrument loading trays to hold pouches in the correct position.
3.9 | The usable volume of the chamber shall be between 35 Litres to 50 Litres
3.10 | The unit offered shall have a built in thermal printer that automatically and continuously monitor and record dates, time of day, load, identification no. alarms, error display and operating parameters.
3.11 | The front panel on the unit shall have user interface control panel, clear menu system icons, visual messages and warning alarms.
3.12 | Display
3.12.1 | It must be at least 7" touch screen display
3.12.2 | The follwing fuctions must visible:
3.12.3 | Date and time
3.12.4 | Temperature
3.12.5 | Set point follow up
3.12.6 | Cycle selected
3.12.7 | Phase
3.12.8 | System errors
3.13 | The unit shall have a built in wheels to provide mobility
3.14 | The following alarms must be available in both audio and visual:
3.14.1 | Power failure
3.14.2 | Low/High Temp
3.14.3 | Door lock
3.14.4 | Vacuum & Pressure alarms
3.15 | The unit must be a single manual door
3.16 | The unit must have a built-in thermal printer
3.17 | USB and Ethernet connectivity must be available in the unit
3.18 | Power
3.18.1 | The unit must utilize 230v mains supply with 50Hz
3.18.2 | Power consumption of the unit must not exceed 2.5 kVA
4 | STANDARD ACCESSORIES and CONSUMABLES
4.1 | 10 x thermal printing paper rolls
4.2 | 2 x Instrument loading trays
4.3 | 10 x Hydrogen Peroxide cassettes or Cartridges
4.4 | Biological and Chemical Indicators
5 | Conformity Compliance (Please attach certificate)
5.1 | The unit must comply with an acceptable international safety standard such as EN 61010-2-041 for Medical Autoclaves, attached certification
5.1.1 | Please provide unique ref number of the EN 61010-2-041 ceriticate:
5.2 | OEM must comply and certified on ISO 4126-1 quality standards, attach proof of compliance
5.2.1 | Please provide unique ref number of the ISO 4126-1 ceriticate:
5.3 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
5.3.1 | Please provide unique ref number of the ISO 13485 ceriticate:
5.4 | Model quoted for must be EC certified. Attach a copy of certification
5.4.1 | Please provide unique ref number of the EC ceriticate:
5.5 | Model quoted for must comply with EN 50081-2 and EN 50082-2. Attach a copy of certification
5.5.1 | Please provide unique ref number of the EN 50081-2 and EN 50082-2 ceriticate:
5.6 | Model quoted for must comply with EN 13060/1-4. Attach a copy of certification
5.6.1 | Please provide unique ref number of the EN 13060/1-4 ceriticate:
5.7 | Model quoted for must comply with ISO 17665-1. Attach a copy of certification
5.7.1 | Please provide unique ref number of the ISO 17665-1 ceriticate:
5.8 | Local Manufacturers must provide certification that unit bidded for comply with SANS 982: 2009
6 | COMPREHENSIVE MAINTENANCE
6.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) must be included in the bid price
6.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
6.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
6.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
6.3 | The bidder to provide the OEM recommended service interval of the unit:
6.4 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
7 | General
7.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
7.2 | The unit on offer must be the latest technology.
PLASMA STERILIZER, LOW TEMPERATURE, 35- 50L - Accessories and Consumables
RT2-10-06a | Thermal printing paper rolls (Pk of 10)
RT2-10-06b | Instrument loading tray
RT2-10-06c | Hydrogen Peroxide cassettes or Cartridges (Pk of 5)
RT2-10-06d | Biological Indicators
RT2-10-06e | Chemical Indicators

## Plasma Sterilizer 85-110L
RT2-10-07 | PLASMA STERILIZER, LOW TEMPERATURE, 85-110L
1 | DESCRIPTION
1.1 | Low Temperature Sterilizer that deliver rapid and safe sterilization for delicate intruments that are sensitive to heat and moisture
2 | Product Offered:
2.1 | Make:
2.2 | Model:
2.3 | Initial Manufacturing year of Model:
3 | TECHNICAL SPECIFICATION
3.1 | The unit offered must be fully automatic, microprocessor controlled and feature advanced technology for safe low-temperature plasma sterilization.
3.2 | The unit's frames, external panels, valves, piping and internal shelving must be constructed from 316 grade stainless steel material
3.3 | The Plasma Sterilizer shall use hydrogen peroxide (H2O2) as sterilization agent
3.4 | The unit shall have processing cycles, which are pre-programmed with the recommended sterilizing cycles as follows (+5 mins)
3.4.1 | Fast Cycle time: 40 mins
3.4.2 | Standard Cycle time: 50 Mins
3.4.3 | Advanced Cycle time: 60 Mins
3.5 | The sterilization temperature must be within + 45OC to 50OC
3.6 | The unit offered shall have auto interlock and open pressure monitoring system that ensures all chamber pressure has been released prior to allowing the door to open
3.7 | The chamber must be able to withstand pressure from vacuum up to 3.5 bars
3.8 | The sterilizing chamber shall be in stainless steel and rectangular with stainless steel instrument loading trays to hold pouches in the correct position.
3.9 | The usable volume of the chamber shall be between 85 Litres to 110 Litres
3.10 | The unit offered shall have a built in thermal printer that automatically and continuously monitor and record dates, time of day, load, identification no. alarms, error display and operating parameters.
3.11 | The front panel on the unit shall have user interface control panel, clear menu system icons, visual messages and warning alarms.
3.12 | Display
3.12.1 | It must be at least 7" touch screen display
3.12.2 | The follwing fuctions must visible:
3.12.3 | Date and time
3.12.4 | Temperature
3.12.5 | Set point follow up
3.12.6 | Cycle selected
3.12.7 | Phase
3.12.8 | System errors
3.13 | The unit shall have a built in wheels to provide mobility
3.14 | The following alarms must be available in both audio and visual:
3.14.1 | Power failure
3.14.2 | Low/High Temp
3.14.3 | Door lock
3.14.4 | Vacuum & Pressure alarms
3.15 | The unit must be a Single auto sliding door and two shelves
3.16 | The unit must have a built-in thermal printer
3.17 | USB and Ethernet connectivity must be available in the unit
3.18 | Power
3.18.1 | The unit must utilize 230v mains supply with 50Hz
3.18.2 | Power consumption of the unit must not exceed 2.5 kVA
4 | STANDARD ACCESSORIES
4.1 | 10 x thermal printing paper rolls
4.2 | 4 x Instrument loading trays
4.3 | 10 x Hydrogen Peroxide cassettes or Cartridges
4.4 | Biological and Chemical Indicators
5 | Conformity Compliance (Please attach certificate)
5.1 | The unit must comply with an acceptable international safety standard such as EN 61010-2-041 for Medical Autoclaves, attached certification
5.1.1 | Please provide unique ref number of the EN 61010-2-041 ceriticate:
5.2 | OEM must comply and certified on ISO 4126-1 quality standards, attach proof of compliance
5.2.1 | Please provide unique ref number of the ISO 4126-1 ceriticate:
5.3 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
5.3.1 | Please provide unique ref number of the ISO 13485 ceriticate:
5.4 | Model quoted for must be EC certified. Attach a copy of certification
5.4.1 | Please provide unique ref number of the EC ceriticate:
5.5 | Model quoted for must comply with EN 50081-2 and EN 50082-2. Attach a copy of certification
5.5.1 | Please provide unique ref number of the EN 50081-2 and EN 50082-2 ceriticate:
5.6 | Model quoted for must comply with EN 13060/1-4. Attach a copy of certification
5.6.1 | Please provide unique ref number of the EN 13060/1-4 ceriticate:
5.7 | Model quoted for must comply with ISO 17665-1. Attach a copy of certification
5.7.1 | Please provide unique ref number of the ISO 17665-1 ceriticate:
5.8 | Local Manufacturers must provide certification that unit bidded for comply with SANS 982: 2009
6 | COMPREHENSIVE MAINTENANCE
6.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) must be included in the bid price
6.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
6.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
6.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
6.3 | The bidder to provide the OEM recommended service interval of the unit:
6.4 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
7 | General
7.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
7.2 | The unit on offer must be the latest technology.
PLASMA STERILIZER, LOW TEMPERATURE, 85- 110L - Accessories and Consumables
RT2-10-07a | Thermal printing paper rolls (Pk of 10)
RT2-10-07b | Instrument loading tray
RT2-10-07c | Hydrogen Peroxide cassettes or Cartridges (Pk of 5)
RT2-10-07d | Biological Indicators
RT2-10-07e | Chemical Indicators

## Plasma Sterilizer 145-160L
RT2-10-08 | PLASMA STERILIZER, LOW TEMPERATURE, 145-160L
1 | DESCRIPTION
1.1 | Low Temperature Sterilizer that deliver rapid and safe sterilization for delicate intruments that are sensitive to heat and moisture
2 | Product Offered:
2.1 | Make:
2.2 | Model:
2.3 | Initial Manufacturing year of Model:
3 | TECHNICAL SPECIFICATION
3.1 | The unit offered must be fully automatic, microprocessor controlled and feature advanced technology for safe low-temperature plasma sterilization.
3.2 | The unit's frames, external panels, valves, piping and internal shelving must be constructed from 316 grade stainless steel material
3.3 | The Plasma Sterilizer shall use hydrogen peroxide (H2O2) as sterilization agent
3.4 | The unit shall have processing cycles, which are pre-programmed with the recommended sterilizing cycles as follows (+5 mins)
3.4.1 | Fast Cycle time: 40 mins
3.4.2 | Standard Cycle time: 50 Mins
3.4.3 | Advanced Cycle time: 60 Mins
3.5 | The sterilization temperature must be within + 45OC to 50OC
3.6 | The unit offered shall have auto interlock and open pressure monitoring system that ensures all chamber pressure has been released prior to allowing the door to open
3.7 | The chamber must be able to withstand pressure from vacuum up to 3.5 bars
3.8 | The sterilizing chamber shall be in stainless steel and rectangular with stainless steel instrument loading trays to hold pouches in the correct position.
3.9 | The usable volume of the chamber shall be between 145 Litres to 160 Litres
3.10 | The unit offered shall have a built in thermal printer that automatically and continuously monitor and record dates, time of day, load, identification no. alarms, error display and operating parameters.
3.11 | The front panel on the unit shall have user interface control panel, clear menu system icons, visual messages and warning alarms.
3.12 | Display
3.12.1 | It must be at least 7" touch screen display
3.12.2 | The follwing fuctions must visible:
3.12.3 | Date and time
3.12.4 | Temperature
3.12.5 | Set point follow up
3.12.6 | Cycle selected
3.12.7 | Phase
3.12.8 | System errors
3.13 | The unit shall have a built in wheels to provide mobility
3.14 | The following alarms must be available in both audio and visual:
3.14.1 | Power failure
3.14.2 | Low/High Temp
3.14.3 | Door lock
3.14.4 | Vacuum & Pressure alarms
3.15 | The unit must be a Single auto sliding door and two shelves
3.16 | The unit must have a built-in thermal printer
3.17 | USB and Ethernet connectivity must be available in the unit
3.18 | Power
3.18.1 | The unit must utilize 230v mains supply with 50Hz
3.18.2 | Power consumption of the unit must not exceed 2.5 kVA
4 | STANDARD ACCESSORIES
4.1 | 10 x thermal printing paper rolls
4.2 | 4 x Instrument loading trays
4.3 | 10 x Hydrogen Peroxide cassettes or Cartridges
4.4 | Biological and Chemical Indicators
5 | Conformity Compliance (Please attach certificate)
5.1 | The unit must comply with an acceptable international safety standard such as EN 61010-2-041 for Medical Autoclaves, attached certification
5.1.1 | Please provide unique ref number of the EN 61010-2-041 ceriticate:
5.2 | OEM must comply and certified on ISO 4126-1 quality standards, attach proof of compliance
5.2.1 | Please provide unique ref number of the ISO 4126-1 ceriticate:
5.3 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
5.3.1 | Please provide unique ref number of the ISO 13485 ceriticate:
5.4 | Model quoted for must be EC certified. Attach a copy of certification
5.4.1 | Please provide unique ref number of the EC ceriticate:
5.5 | Model quoted for must comply with EN 50081-2 and EN 50082-2. Attach a copy of certification
5.5.1 | Please provide unique ref number of the EN 50081-2 and EN 50082-2 ceriticate:
5.6 | Model quoted for must comply with EN 13060/1-4. Attach a copy of certification
5.6.1 | Please provide unique ref number of the EN 13060/1-4 ceriticate:
5.7 | Model quoted for must comply with ISO 17665-1. Attach a copy of certification
5.7.1 | Please provide unique ref number of the ISO 17665-1 ceriticate:
5.8 | Local Manufacturers must provide certification that unit bidded for comply with SANS 982: 2009
6 | COMPREHENSIVE MAINTENANCE
6.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) must be included in the bid price
6.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
6.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
6.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
6.3 | The bidder to provide the OEM recommended service interval of the unit:
6.4 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
7 | General
7.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
7.2 | The unit on offer must be the latest technology.
PLASMA STERILIZER, LOW TEMPERATURE, 145- 160L - Accessories and Consumables
RT2-10-08a | Thermal printing paper rolls (Pk of 10)
RT2-10-08b | Instrument loading tray
RT2-10-08c | Hydrogen Peroxide cassettes or Cartridges (Pk of 5)
RT2-10-08d | Biological Indicators
RT2-10-08e | Chemical Indicators

## Scope Washer
RT2-10-09 | ENDOSCOPE WASHER
1 | DESCRIPTION
1.1 | An automatic machine used for the reprocessing of flexible endospes
2 | Product Offered:
2.1 | Make:
2.2 | Model:
2.3 | Initial Manufacturing year of Model:
3 | TECHNICAL SPECIFICATION
3.1 | The unit is required for the washing and disinfectant of flexible fiber- and videoscopes. The system shall be able to accommodate two scopes per cycle
3.2 | Bidders to state the functions of the washer as far as amount of scopes is concerned. Preference will be given to units that can accommodate two scopes per cycle, but have the option run two separate, independent cycles
3.3 | It shall be possible to wash and disinfect all makes and all types (gastro- colono- etc) of scopes currently available on the SA market
3.4 | The unit offered shall be able to use all makers’ disinfectant and detergent
3.5 | Bidders shall clearly, on a separate page, show what detergents and disinfectants should be used per cycle, and show the cost of these detergents and disinfectants. All couplers and tubing if required shall also be included in the calculations. It shall be possible to determine the running cost of the washer from these calculations, and these running costs shall be taken into account when adjudicating the offers
3.6 | The washer shall disinfect all channels of the scopes to ensure acceptable international standards of hygiene
3.7 | Bidders to clearly and fully describe the functioning of the washer, and how the abovementioned standards of hygiene are achieved with the functioning of the offered unit
3.8 | The washer shall have audio and visual alarms for all monitored parameters. The alarms shall be accurate, and no nuisance alarms shall be sounded
3.9 | The washer shall monitor the presence as well as volumes of detergent and disinfectant
3.1 | The washer shall be able to perform a fully controlled thermal self disinfectant cycle
3.11 | The washer shall monitor fluid pathways with continuous leak testing to prevent fluid ingress during any part of the cleansing cycle
3.12 | The washer shall monitor for channel blockages, and also indicate when sufficient fluid is pumped through the channels
3.13 | The washer shall incorporate high fluid circulation speed to enhance disinfectant removal and prevent bio film production
3.14 | The washer shall be able to flush air through the separate channels to remove water
3.15 | The washer shall be supplied with all accessories, ready for operation, as well as a starter pack consisting of enough consumables to last for a month when 6 cycles are done per day
3.16 | A trough that can be used, and is dedicated for the brushing of the scopes shall be included in the offer. Bidders to give full data of the offered unit
3.17 | Bidders to clearly indicate the water and electricity requirements for the offered unit
4 | Conformity Compliance (Please attach certificate)
4.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
4.2 | The unit must comply and certified on ISO 15883-4:2018 or latest, attach proof of compliance
4.2.1 | Please provide unique ref number of the ISO 15883-4:2018 ceriticate:
4.3 | The unit must comply and certified on ISO 14937, attach proof of compliance
4.3.1 | Please provide unique ref number of the ISO 14937 ceriticate:
4.4 | OEM must comply and certified on ISO 9001 quality standards, attach proof of compliance
4.4.1 | Please provide unique ref number of the ISO 9001 ceriticate:
4.5 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
4.5.1 | Please provide unique ref number of the ISO 13485 ceriticate:
4.6 | Model quoted for must be EC certified. Attach a copy of certification
4.6.1 | Please provide unique ref number of the EC ceriticate:
4.7 | The equipment quoted for must be protected against electromagnetic interference IEC 60601-1-2
4.7.1 | Please provide unique ref number of the IEC ceriticate:
5 | COMPREHENSIVE MAINTENANCE
5.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) must be included in the bid price
5.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
5.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
5.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
5.3 | The bidder to provide the OEM recommended service interval of the unit:
5.4 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
6 | General
6.1 | Has the product on offer or any of its components being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
13.2 | The unit on offer must be the latest technology.

## Category13-summary
INSTALLATION ALTERATIONS REQUIRED
# | Item number | Item Description
1 | RT2-13-01 | Instrument Washer, Under-counter
2 | RT2-13-02 | Instrument Washer, Single door
3 | RT2-13-03 | Instrument Washer, Double door
4 | RT2-13-04 | Theater Light, Ceiling-Mount

## Instrument Washer, Undercounter
RT2-13-01 | INSTRUMENT WASHER, UNDER COUNTER
1 | DESCRIPTION
1.1 | Under counter Intrument Washer for cleaning, thermal disinfecting and drying surgical instruments, glassware and anesthesia material
1a | Product Offered:
1a.1 | Make:
1a.2 | Model:
1a.3 | Initial Manufacturing year of Model:
2 | TECHNICAL SPECIFICATION
2.1 | The unit offered must be fully automatic and microprocessor controlled.
2.2 | The unit shall have processing cycles, which are pre-programmed with the recommended treatment parameters for pre-washing, cleaning, rinsing, thermal disinfection and drying.
2.3 | The instrument washer shall have detergents and rinse agents, which are automatically dispensed during cycle
2.4 | The instrument washer shall have glass window on the load chamber
2.5 | The exterior height of the washer should be designed to fit under counter/bench
2.6 | The instrument chamber shall be fitted with at least two (2) removable stainless steel shelves/levels each able to accommodate at least 2 full size DIN trays of (WxHxD) 480mm x 250mm x 64mm.
2.8 | The instrument washer shall have rotary spray arms at the top and bottom of the chamber.
2.9 | The front panel on the instrument washer shall have a display, displaying all cycles and errors.
2.1 | The spray system and recirculation piping of the instrument washer shall be stainless steel.
2.11 | The cabinet casing of the instrument washer should be constructed of 316 L stainless steel grade.
2.12 | The instrument washer shall have a water heating system.
2.13 | The instrument washer shall have temperature control and monitoring.
2.14 | The drain capacity of the instrument washer shall be no less than 30 L/min.
2.15 | The instrument washer electrical sump heater shall be no less than 7.4kW.
2.16 | The instrument washer electrical dryer heater shall be no less than 1.5kW.
2.17 | The noise level of the instrument washer shall be no greater than 80 db in the washing and drying cycle.
2.18 | The unit must be fitted with an intergral thermal printer.
2.19 | The instrument washer pump capacities shall be as follows:
2.19.1 | a) Circulation pump – no less than 350L/min
2.19.2 | b) Pump purified water (Optional)
2.19.3 | c) Pump alkaline detergent : +_ 20 L/min.
2.2 | The instrument washer water consumption shall be no less than 12 L/min in the wash phase.
2.21 | The instrument washer shall have an audible alarm.
2.22 | The instrument washer shall have safety interlocking system to prevent operation when the door is open
2.23 | The instrument washer shall have safety interlocking system to prevent door from opening when the washer is in operation
2.24 | The instrument washer shall be supplied with drainage hose, hot and cold water hoses with approximately 1.5m in length
3 | STANDARD ACCESSORIES
3.1 | Heat resistant gloves
3.2 | 10 x thermal printing paper rolls
3.3 | 4x DIN Instrument trays
3.4 | 2L cleaning detergent
3.5 | 2L Rinse
3.6 | 1x Loading trolley
4 | Conformity Compliance (Please attach certificate)
4.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
4.2 | The unit must comply and certified on ISO 15883-4:2018 or latest, attach proof of compliance
4.2.1 | Please provide unique ref number of the ISO 15883-4:2018 ceriticate:
4.3 | OEM must comply and certified on ISO 9001 quality standards, attach proof of compliance
4.3.1 | Please provide unique ref number of the ISO 9001 ceriticate:
4.4 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
4.4.1 | Please provide unique ref number of the ISO 13485 ceriticate:
4.5 | Model quoted for must be EC certified. Attach a copy of certification
4.5.1 | Please provide unique ref number of the EC ceriticate:
4.6 | The equipment quoted for must be protected against electromagnetic interference IEC 60601-1-2
4.6.1 | Please provide unique ref number of the IEC ceriticate:
5 | COMPREHENSIVE MAINTENANCE
5.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) must be included in the bid price
5.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
5.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
5.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
5.3 | The bidder to provide the OEM recommended service interval of the unit:
5.4 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
6 | General
6.1 | Has the product on offer or any of its components being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
6.2 | The unit on offer must be the latest technology.
INSTRUMENT WASHER, UNDER COUNTER REPLACEMENT ACCESSORIES
RT2-13-01a | Water Pressure booster
RT2-13-01b | Water Purifier
RT2-13-01c | Heat resistant gloves (Pair)
RT2-13-01d | Thermal printing paper rolls (Pk of 10)
RT2-13-01e | DIN Instrument tray
RT2-13-01f | 2L cleaning detergent
RT2-13-01g | 2L Rinse
RT2-13-01h | Loading trolley

## Instrument Washer, Single door
RT2-13-02 | INSTRUMENT WASHER, SINGLE DOOR
1 | DESCRIPTION
1.1 | Free standing intrument washer for cleaning, thermal disinfecting and drying surgical instruments, glassware and anesthesia material
1a | Product Offered:
1a.1 | Make:
1a.2 | Model:
1a.3 | Initial Manufacturing year of Model:
2 | TECHNICAL SPECIFICATION
2.1 | The unit offered must be fully automatic and microprocessor controlled.
2.2 | The unit shall have processing cycles, which are pre-programmed with the recommended treatment parameters for pre-washing, cleaning, rinsing, thermal disinfection and drying.
2.3 | The instrument washer shall have detergents and rinse agents, which are automatically dispensed during cycle
2.4 | The instrument washer shall have glass window on the load chamber
2.5 | The exterior height of the washer should be sufficient to accomodate the load height and load chamber height with drying option
2.6 | The instrument chamber shall be fitted with at least four (4) removable stainless steel shelves/levels each able to accommodate at least 2 full size DIN trays of (WxHxD) 480mm x 250mm x 64mm.
2.7 | The instrument washer must have a load height of + 780mm (tolerance 50mm)
2.8 | The instrument washer shall have rotary spray arms at the top and bottom of the chamber.
2.9 | The front panel on the instrument washer shall have a display, displaying all cycles and errors.
2.1 | The spray system and recirculation piping of the instrument washer shall be stainless steel.
2.11 | The cabinet casing of the instrument washer should be constructed of stainless steel.
2.12 | The instrument washer shall have a water heating system.
2.13 | The instrument washer shall have temperature control and monitoring.
2.14 | The drain capacity of the instrument washer shall be no less than 30 L/min.
2.15 | The instrument washer electrical sump heater shall be no less than 7.4kW.
2.16 | The instrument washer electrical dryer heater shall be no less than 1.5kW.
2.17 | The noise level of the instrument washer shall be no greater than 80 db in the washing and drying cycle.
2.18 | The unit must be fitted with an intergral thermal printer.
2.19 | The instrument washer pump capacities shall be as follows:
2.19.1 | a) Circulation pump – no less than 350L/min
2.19.2 | b) Pump purified water (Optional)
2.19.3 | c) Pump alkaline detergent : +_ 20 L/min.
2.2 | The instrument washer water consumption shall be no less than 12 L/min in the wash phase.
2.21 | The instrument washer shall have an audible alarm.
2.22 | The instrument washer shall have safety interlocking system to prevent operation when the door is open
2.23 | The instrument washer shall have safety interlocking system to prevent door from opening when the washer is in operation
2.24 | The instrument washer shall be supplied with drainage hose, hot and cold water hoses with approximately 1.5m in length
2.25 | The trolley should have antistatic castors with at least 2 equipped with locking mechanism/brakes
3 | STANDARD ACCESSORIES
3.1 | Heat resistant gloves
3.2 | 10 x thermal printing paper rolls
3.3 | 8x DIN Instrument trays
3.4 | 10L cleaning detergent
3.5 | 10L Rinse
3.6 | 1x Loading trolley
4 | Conformity Compliance (Please attach certificate)
4.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
4.2 | The unit must comply and certified on ISO 15883-4:2018 or latest, attach proof of compliance
4.2.1 | Please provide unique ref number of the ISO 15883-4:2018 ceriticate:
4.3 | OEM must comply and certified on ISO 9001 quality standards, attach proof of compliance
4.3.1 | Please provide unique ref number of the ISO 9001 ceriticate:
4.4 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
4.4.1 | Please provide unique ref number of the ISO 13485 ceriticate:
4.5 | Model quoted for must be EC certified. Attach a copy of certification
4.5.1 | Please provide unique ref number of the EC ceriticate:
4.6 | The equipment quoted for must be protected against electromagnetic interference IEC 60601-1-2
4.6.1 | Please provide unique ref number of the IEC ceriticate:
5 | COMPREHENSIVE MAINTENANCE
5.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) must be included in the bid price
5.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
5.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
5.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
5.3 | The bidder to provide the OEM recommended service interval of the unit:
5.4 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
6 | General
6.1 | Has the product on offer or any of its components being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
6.2 | The unit on offer must be the latest technology.
INSTRUMENT WASHER, SINGLE DOOR REPLACEMENT ACCESSORIES
RT2-13-02a | Water Pressure booster
RT2-13-02b | Water Purifier
RT2-13-02c | Heat resistant gloves (Pair)
RT2-13-02d | Thermal printing paper rolls (Pk of 10)
RT2-13-02e | DIN Instrument tray
RT2-13-02f | 5L cleaning detergent
RT2-13-02g | 10L cleaning detergent
RT2-13-02h | 5L Rinse
RT2-13-02i | 10L Rinse
RT2-13-02j | Loading trolley

## Instrument Washer, Double door
RT2-13-03 | INSTRUMENT WASHER, DOUBLE DOOR
1 | DESCRIPTION
1.1 | Free standing intrument washer for cleaning, thermal disinfecting and drying surgical instruments, glassware and anesthesia material
1a | Product Offered:
1a.1 | Make:
1a.2 | Model:
1a.3 | Initial Manufacturing year of Model:
2 | TECHNICAL SPECIFICATION
2.1 | The unit offered must be fully automatic and microprocessor controlled.
2.2 | The unit shall have processing cycles, which are pre-programmed with the recommended treatment parameters for pre-washing, cleaning, rinsing, thermal disinfection and drying.
2.3 | The instrument washer shall have detergents and rinse agents, which are automatically dispensed during cycle
2.4 | The instrument washer shall have glass windows on both doors
2.5 | The exterior height of the washer should be sufficient to accomodate the load height and load chamber height with drying option
2.6 | The instrument chamber shall be fitted with at least four (4) removable stainless steel shelves/levels each able to accommodate at least 2 full size DIN trays of (WxHxD) 480mm x 250mm x 64mm.
2.7 | The instrument washer must have a load height of + 780mm (tolerance 50mm)
2.8 | The instrument washer shall have rotary spray arms at the top and bottom of the chamber.
2.9 | The front panel on the instrument washer shall have a display, displaying all cycles and errors.
2.1 | The spray system and recirculation piping of the instrument washer shall be stainless steel.
2.11 | The cabinet casing of the instrument washer should be constructed of stainless steel.
2.12 | The instrument washer shall have a water heating system.
2.13 | The instrument washer shall have temperature control and monitoring.
2.14 | The drain capacity of the instrument washer shall be no less than 30 L/min.
2.15 | The instrument washer electrical sump heater shall be no less than 7.4kW.
2.16 | The instrument washer electrical dryer heater shall be no less than 1.5kW.
2.17 | The noise level of the instrument washer shall be no greater than 80 db in the washing and drying cycle.
2.18 | The unit must be fitted with an intergral thermal printer.
2.19 | The instrument washer pump capacities shall be as follows:
2.19.1 | a) Circulation pump – no less than 350L/min
2.19.2 | b) Pump purified water (Optional)
2.19.3 | c) Pump alkaline detergent : +_ 20 L/min.
2.2 | The instrument washer water consumption shall be no less than 12 L/min in the wash phase.
2.21 | The instrument washer shall have an audible alarm.
2.22 | The instrument washer shall have safety interlocking system to prevent operation when the door is open
2.23 | The instrument washer shall have safety interlocking system to prevent door from opening when the washer is in operation
2.24 | The instrument washer shall be supplied with drainage hose, hot and cold water hoses with approximately 1.5m in length
2.25 | The trolley should have antistatic castors with at least 2 equipped with locking mechanism/brakes
3 | STANDARD ACCESSORIES
3.1 | Heat resistant gloves
3.2 | 10 x thermal printing paper rolls
3.3 | 8x DIN Instrument trays
3.4 | 10L cleaning detergent
3.5 | 10L Rinse
3.6 | 2x Loading trolleys (Clean & Dirty side)
4 | Conformity Compliance (Please attach certificate)
4.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
4.2 | The unit must comply and certified on ISO 15883-4:2018 or latest, attach proof of compliance
4.2.1 | Please provide unique ref number of the ISO 15883-4:2018 ceriticate:
4.3 | OEM must comply and certified on ISO 9001 quality standards, attach proof of compliance
4.3.1 | Please provide unique ref number of the ISO 9001 ceriticate:
4.4 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
4.4.1 | Please provide unique ref number of the ISO 13485 ceriticate:
4.5 | Model quoted for must be EC certified. Attach a copy of certification
4.5.1 | Please provide unique ref number of the EC ceriticate:
4.6 | The equipment quoted for must be protected against electromagnetic interference IEC 60601-1-2
4.6.1 | Please provide unique ref number of the IEC ceriticate:
5 | COMPREHENSIVE MAINTENANCE
5.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) must be included in the bid price
5.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
5.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
5.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
5.3 | The bidder to provide the OEM recommended service interval of the unit:
5.4 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
6 | General
6.1 | Has the product on offer or any of its components being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
6.2 | The unit on offer must be the latest technology.
INSTRUMENT WASHER, DOUBLE DOOR REPLACEMENT ACCESSORIES
RT2-13-03a | Water Pressure booster
RT2-13-03b | Water Purifier
RT2-13-03c | Heat resistant gloves (Pair)
RT2-13-03d | Thermal printing paper rolls (Pk of 10)
RT2-13-03e | DIN Instrument tray
RT2-13-03f | 5L cleaning detergent
RT2-13-03g | 10L cleaning detergent
RT2-13-03h | 5L Rinse
RT2-13-03i | 10L Rinse
RT2-13-03j | Loading trolley

## Theater Light, Ceiling-mount
RT2-13-04 | Theater Light, Ceiling-Mount
1 | DESCRIPTION
1.1
1a | Product Offered:
1a.1 | Make:
1a.2 | Model:
1a.3 | Initial Manufacturing year of Model:
2 | TECHNICAL SPECIFICATION
2.1 | The requirement is for a double head system, with a main lamp with output of at least 160 000 lux, and a satellite lamp system with output of at least 110 000 lux. All measurements at 1 metre
2.2 | The lamps shall be ceiling suspended, and all movements required during a normal surgical procedure shall be possible. During the site visit to an installed base the movements of the lamps shall be used as evaluation criteria for the adjudication of the offers
2.3 | The ceiling mounted suspension shall be so suspended that any position may be selected with ease, for illuminating any required operating area on the patient, with virtually unlimited rotatability
2.4 | The lamp shall be spring balanced in the supporting arm, be easy to move, have no backlash, and shall have friction clutches or similar devices for adjustment
2.5 | The lamp shall be capable of shining both vertically downwards for normal use, and horizontally for lithotomy position operations
2.6 | Good, positive yet easy flexibility of height, lateral and tilt movement is required
2.7 | The lamp head shall be equipped with scratch-resistant material to protect the LED’s from accidental damage. The lamp head shall be of modern design combined in one housing
2.8 | The unit shall have easy, positive manoeuvrability
2.9 | The unit shall have excellent laminar flow features, and smooth surfaces that offer low resistance to airflow
2.1 | The light head and the positioning of the LED’s shall ensure a shadowless working area, regardless of the surgeon and supports staff’s position
2.11 | The focusable size of the light field for the main lamp shall be approximately 200 to 300 mm, and for the second lamp approximately 180 to 230mm. Bidders to state the way focusing is achieved. It shall not be necessary to move the lamp away from the patient to achieve a bigger focal spot
2.12 | Depth of illumination shall be more than 1 000mm without refocusing
2.13 | LED life shall be guaranteed to be at least 40 000 hours
2.14 | The lamp head shall have smooth, flat surfaces for easy cleaning
2.15 | The lights shall be electrically dimmable from approximately 30 to 100% of full illumination. The lamps shall have an “Endoscopy” setting to dim the lights to low light condition. Bidders to state the % of full light achieved during “Endoscopy” dimming
2.16 | The light shall have a colour temperature of approximately 4 300K and a colour rendering index of 95%.
2.17 | Bidders to state the composition of the LED’s in the lamphead, and what replacement of the LED’s entail – per each, per block etc, as well as the current cost for replacement of LED’s
2.18 | UPS backup is required for each lamp. The UPS shall keep the lamp functional on full load for at least one hour after power failure
2.19 | A removable sterilisable handle on the lamp head shall enable the surgeon to easily adjust the position of the light, as well as setting the light field or focus. A spare handle shall be supplied with each unit
2.2 | If ceilings or any other construction is in any way damaged during the installation, repair and making good shall be done immediately
3 | STANDARD ACCESSORIES
3.1 | Handles
4 | Conformity Compliance (Please attach certificate)
4.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
4.2 | OEM must comply and certified on ISO 9001 quality standards, attach proof of compliance
4.2.1 | Please provide unique ref number of the ISO 9001 ceriticate:
4.3 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
4.3.1 | Please provide unique ref number of the ISO 13485 ceriticate:
4.4 | Model quoted for must be EC certified. Attach a copy of certification
4.4.1 | Please provide unique ref number of the EC ceriticate:
4.5 | The equipment quoted for must be protected against electromagnetic interference IEC 60601-1-2
4.5.1 | Please provide unique ref number of the IEC ceriticate:
5 | COMPREHENSIVE MAINTENANCE
5.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) must be included in the bid price
5.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
5.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
5.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
5.3 | The bidder to provide the OEM recommended service interval of the unit:
5.4 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
6 | General
6.1 | Has the product on offer or any of its components being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
6.2 | The unit on offer must be the latest technology.
Theater Light, Ceiling-Mount- Accessories and Consumables
RT2-13-04a | LED replacement

## Category 14- Summary
Multi-disciplinary Specialized Equipment
# | Item number | Item Description
1 | RT2-14-01 | Theater Table, Standard
2 | RT2-14-02 | Theater Table, Bariatric
3 | RT2-14-03 | Laparascope, Mid-range
4 | RT2-14-04 | Laparascope, High-end (3D)
5 | RT2-14-06 | Laparascope-Exoscope, Neuro
6 | RT2-14-07 | Video Endoscopic System

## Theater Table, Standard
RT2-14-01 | Theater Table, Standard
1 | DESCRIPTION
1.1 | Theatre table on which the patient lies during a surgical operation, universal electronically operated
2 | Product Offered:
2.1 | Make:
2.2 | Model:
2.3 | Initial Manufacturing year of Model:
3 | TECHNICAL SPECIFICATION
3.1 | The table must have a maximum load capacity of atleast 250kg
3.2 | Radiotranslucent table top for free C-arm access
3.3 | Decompression mattress with non-slip, antistatic, waterproof and seamless design
3.4 | Longitudinal shift of at least 25cm and still maintaining the balance and stability with maximum load capacity
3.5 | Four section operation
3.6 | Trendelenberg and reverse trendelenberg function to at least 25 degrees
3.7 | The maximum height of the table measured without mattress must be more than 1000mm from the floor and minimum height of the table must be less than 710mm from the floor
3.8 | Table controls on remote handpiece as well as on the table column with positions clearly indicated.
3.9 | Table column base covers must be at least 304 stainless steel grade
3.1 | The table base must be stable, mobile and lockable/brakes through both the hand remote piece and the coloumn panel and the unit must have manual override for the brakes should the electronic fails
3.11 | Separate 2 leg sections
3.12 | Battery backup of at least 2 hours operation when fully charded with internal battery charger
3.13 | Lateral tilt of at least 18 degrees
3.14 | Reflex position of at least 100 degrees
4 | STANDARD ACCESSORIES
4.1 | Anaesthetics arm board x 2
4.2 | Set of padded Lithotomy poles (manual) with their own universal clamps to allow resting of the calf
4.3 | 2 sets of bedside rail clamps
4.4 | Anaesthetic screen
4.5 | Accessory storage container
5 | Conformity Compliance (Please attach certificate)
5.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
5.2 | The unit must comply with electrical safety and performance standard such as IEC 60601-2-46, attached certification
5.2.1 | Please provide unique ref number of the IEC ceriticate:
5.3 | OEM must comply and certified on ISO 9001 quality standards, attach proof of compliance
5.3.1 | Please provide unique ref number of the ISO 9001 ceriticate:
5.4 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
5.4.1 | Please provide unique ref number of the ISO 13485 ceriticate:
5.5 | Model quoted for must be EC certified. Attach a copy of certification
5.5.1 | Please provide unique ref number of the EC ceriticate:
5.6 | The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade (rubber to with 2 screws) type and it must be a minimum length of three (3) meters. NB the mains cable of the unit tendered for must be SABS colour coded.
5.7 | The equipment quoted for must be protected against electromagnetic interference IEC 60601-1-2
5.7.1 | Please provide unique ref number of the IEC ceriticate:
6 | COMPREHENSIVE MAINTENANCE
6.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) must be included in the bid price
6.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
6.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
6.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
6.3 | The bidder to provide the OEM recommended service interval of the unit:
6.4 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
7 | General
7.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
7.2 | The unit on offer must be the latest technology.
RT2-14-01a | Orthopaedic Accessories Set (Compatible with RT2-14-01 & RT2-14-02)
1.1 | Neuro Head Support, Horseshoe Head Support
1.2 | Orthopaedic extension unit (complete to accommodate upper & lower limp surgery)
1.3 | Back section extention
1.4 | Arm hand table
1.5 | Shoulder plate
1.6 | Orthopaedic Accessory cart
1.7 | Patient strapping
1.8 | Foot stop/board
RT2-14-01b | Neurosurgery Accessories Set (Compatible with RT2-14-01 & RT2-14-02)
1.1 | General purpose head support
1.2 | Neuro – Skull Clamp
1.3 | Neuro – Horse Shoe
1.4 | Neuro – Base Unit
1.5 | Neuro – Swivel adaptor with clamps
1.6 | Neuro – Crossbar Adaptor
1.7 | Neuro – Wall Storage Unit
1.8 | Neuro – Head Rest adaptor
RT2-14-01c | Spinal Accessories Set (Compatible with RT2-14-01 & RT2-14-02)
1.1 | General purpose head support
1.2 | Neuro – Skull Clamp
1.3 | Neuro – Horse Shoe
1.4 | Neuro – Base Unit
1.5 | Neuro – Swivel adaptor with clamps
1.6 | Neuro – Crossbar Adaptor
1.7 | Neuro – Wall Storage Unit
1.8 | Neuro – Head Rest adaptor
1.9 | Posterior spinal attachment consisting of a self-contained transport system, a completely radiolucent table top made up of a head support system as well as an abdominal support system all of which must be able to move longitudinally to accommodate adult and pediatric size patients
Theater Table Replacement Accessories (Compatible with RT2-14-01 & RT2-14-02)
RT2-14-01d | Anaesthetics arm board (Pair)
RT2-14-01e | Padded Lithotomy poles (manual) with their own universal clamps to allow resting of the calf (Pair)
RT2-14-01f | Bedside rail clamps (Pair)
RT2-14-01g | Anaesthetic screen (Complete with clamps)
RT2-14-01h | Accessory storage container
RT2-14-01i | Disposable bin for Urology and Gynecology
RT2-14-01j | urology adaptor
RT2-14-01k | General purpose head support
RT2-14-01l | Neuro – Skull Clamp
RT2-14-01m | Neuro – Horse Shoe
RT2-14-01n | Neuro – Base Unit
RT2-14-01o | Neuro – Swivel adaptor with clamps
RT2-14-01p | Neuro – Crossbar Adaptor
RT2-14-01q | Neuro – Wall Storage Unit
RT2-14-01r | Neuro – Head Rest adaptor

## Theater Table, Bariatric
RT2-14-02 | THEATER TABLE, BARIATRIC
1 | DESCRIPTION
1.1 | Theatre table on which the patient lies during a surgical operation, universal electronically operated
2 | Product Offered:
1.1 | Make:
1.2 | Model:
1.3 | Initial Manufacturing year of Model:
3 | TECHNICAL SPECIFICATION
3.1 | The table must have a maximum load capacity of 450kg
3.2 | Radiotranslucent table top for free C-arm access
3.3 | Decompression mattress with non-slip, antistatic, waterproof and seamless design
3.4 | Longitudinal shift of at least 25cm and still maintaining the balance and stability with maximum load capacity
3.5 | Four section operation
3.6 | Trendelenberg and reverse trendelenberg function to at least 25 degrees
3.7 | The maximum height of the table measured without mattress must be more than 1000mm from the floor and minimum height of the table must be less than 710mm from the floor
3.8 | Table controls on remote handpiece as well as on the table column with positions clearly indicated.
3.9 | Table column base covers must be at least 304 stainless steel grade
3.1 | The table base must be stable, mobile and lockable/brakes through both the hand remote piece and the coloumn panel and the unit must have manual override for the brakes should the electronic fails
3.11 | Separate 2 leg sections
3.12 | Battery backup of at least 2 hours operation when fully charded with internal battery charger
3.13 | Lateral tilt of at least 18 degrees
3.14 | Reflex position of at least 100 degrees
4 | STANDARD ACCESSORIES
4.1 | Anaesthetics arm board x 2
4.2 | Set of padded Lithotomy poles (manual) with their own universal clamps to allow resting of the calf
4.3 | 2 sets of bedside rail clamps
4.4 | Anaesthetic screen
4.5 | Accessory storage container
5 | Conformity Compliance (Please attach certificate)
5.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
5.2 | The unit must comply with electrical safety and performance standard such as IEC 60601-2-46, attached certification
5.2.1 | Please provide unique ref number of the IEC ceriticate:
5.3 | OEM must comply and certified on ISO 9001 quality standards, attach proof of compliance
5.3.1 | Please provide unique ref number of the ISO 9001 ceriticate:
5.4 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
5.4.1 | Please provide unique ref number of the ISO 13485 ceriticate:
5.5 | Model quoted for must be EC certified. Attach a copy of certification
5.5.1 | Please provide unique ref number of the EC ceriticate:
5.6 | Hazardous Substance Act: If this type of equipment / apparatus appears on the schedule of Hazardous substances issued by the Directorate: Health Technology of the Department of Health, a license in terms of the Act on Hazardous Substances (Act 15/1973) must be submitted with the bid / quotation documents. The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.7 | All electrical/electronic medical equipment must be licensed by Radiation Control, where applicable a copy of the license must be submitted The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.7.1 | Radiation Control License number:
5.8 | Must comply with 2.5 Occupational Health and Safety Act, Act no 85 of 1993 and Regulations
5.9 | The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade (rubber to with 2 screws) type and it must be a minimum length of three (3) meters. NB the mains cable of the unit tendered for must be SABS colour coded.
5.8 | The equipment quoted for must be protected against electromagnetic interference IEC 60601-1-2
5.8.1 | Please provide unique ref number of the IEC ceriticate:
6 | COMPREHENSIVE MAINTENANCE
6.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) must be included in the bid price
6.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
6.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
6.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
6.3 | The bidder to provide the OEM recommended service interval of the unit:
6.4 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
7 | General
7.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
7.2 | The unit on offer must be the latest technology.
BARIATRIC THEATER TABLE ACCESSORIES (Compatible & Series award with RT2-14-02)
RT2-14-02a | Table extensions to 710mm
RT2-14-02b | Reverse trendelenburg straps
RT2-14-02c | Foot stop/board
RT2-14-02d | 710mm wide Foot extension
Bidder to provide any other Bariatric accessories not listed above

## Laparascope, Mid-range
RT2-14-03 | VIDEO SYSTEM, LAPAROSCOPIC (Mid-range)
1 | DESCRIPTION
1.1 | Laparoscopic camera stack used to examine the organs inside the abdomen
1a | Product Offered:
1a.1 | Make:
1a.2 | Model:
1a.3 | Initial Manufacturing year of Model:
2 | TECHNICAL SPECIFICATION
2.1 | The stack system should come complete with monitor, video processor, insufflator, camera console, data management system, laparoscopic rigid scopes, fiber optic light cable and the appropriate trolley.
3 | Video Processor
3a | Make:
3b | Model:
3.1 | An Autoclavable light guide cable of at least 5mm diameter and atleast 2.3 meter length.
3.2 | Should have digital zoom facility and white balancing
3.3 | Video control unit must have 1 or more camera head inputs with image display settings available
3.4 | The unit offered should have the following output configurations Composite, S-video,Digital, DVI, HDMI, BNC and AVI
4 | HD camera head (Urology specific Camera Head must be substituted and supplied for Urology)
4a | Make:
4b | Model:
4.1 | Controlled on camera console and head with programmable buttons for capture, zoom, white balancing.
4.2 | Latest available HD Camera Head with a "clip-on" scope mounting design shall be provided.
4.3 | Camera head should have a minimum of three programmable switches to allow user remote control
4.4 | Capturing facility (Programmable Button)
4.5 | White balancing function must be available on the camera
4.6 | The length of the camera head cable must be between 3m and 5 m
4.7 | It must be autoclavable or other means sterilization, please state the method
5 | CO2 Insufflator
5a | Make:
5b | Model:
5.1 | Insufflator must have both low and high flow capacity
5.2 | The unit should have display to indicate pressure in abdomen
5.3 | Visual and audible alarms are required to indicate overpressure.
5.4 | There should be an indication of the gas level in the cylinder.
5.5 | Automatic pressure relief mechanism when abdominal pressure exceeds set pressure
5.6 | Universal wrench
6 | Light Source
6a | Make:
6b | Model:
6.1 | Lamp design must be LED.
6.2 | Light surce must have adjustable light intensity
6.3 | An autoclavable light guide cable of at least 5mm diameter and 2.3m length should be provided.
7 | Monitor
7a | Make:
7b | Model:
7.1 | It should be at least 26" Flat LCD or LED medical grade monitor with the following specifications:
7.2 | The trolley should have antistatic castors with at least 2 equipped with locking system/brake.
7.3 | The unit offered should have a minimum of 3 inputs which include; Video, Y/C, RGB, DVI, HDMI
7.4 | High contrast and wide viewing angle
8 | Data Management System
8a | Make:
8b | Model:
8.1 | The unit offered should record both video and still images in full HD
8.2 | Internal memory should be a minimum of 500GB
8.3 | Two-channel simultaneous recording
8.4 | Scope button to record and capture functions
8.5 | The system offered should be compatible with the printer offered so that the recorded items can be printed out.
9 | Trolley
9.1 | The offered unit should come in a purpose built Endoscopic System trolley with a minimum of one drawer for storage
9.2 | The trolley offered should have a tray for the keyboard and mouse.
9.3 | The trolley should have antistatic castors with at least 2 equipped with locking system/brake.
9.4 | The trolley should have a handle for ease access and transportation and a cable management arm.
9.5 | The trolley should have a cover at the rear for safety of cables.
9.6 | Trolley must have a secure gas cylinder holder
9.7 | The trolley must be supplied with a minimum of 3 additional 220 V power sockets.
10 | Printer
10a | Make:
10b | Model:
10.1 | A4 or A6 (minimum resolution of 423dpi) colour printer
11 | Conformity Compliance (Please attach certificate)
11.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
11.2 | The unit must comply with electrical safety and performance standard such as IEC 60601-2-18, attached certification
11.2.1 | Please provide unique ref number of the IEC ceriticate:
11.3 | OEM must comply and certified on ISO 9001 quality standards, attach proof of compliance
11.3.1 | Please provide unique ref number of the ISO 9001 ceriticate:
11.4 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
11.4.1 | Please provide unique ref number of the ISO 13485 ceriticate:
11.5 | Model quoted for must be EC certified. Attach a copy of certification
11.5.1 | Please provide unique ref number of the EC ceriticate:
11.6 | The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade (rubber to with 2 screws) type and it must be a minimum length of three (3) meters. NB the mains cable of the unit tendered for must be SABS colour coded.
11.7 | The equipment quoted for must be protected against electromagnetic interference IEC 60601-1-2
11.7.1 | Please provide unique ref number of the IEC ceriticate:
12 | COMPREHENSIVE MAINTENANCE
12.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) must be included in the bid price
12.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
12.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
12.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
12.3 | A 5-year service exchange for the laparoscopic rigid scopes should be quoted for separately
12.4 | The bidder to provide the OEM recommended service interval of the unit:
12.5 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
13 | General
13.1 | Has the product on offer or any of its components being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
13.2 | The unit on offer must be the latest technology.

## Laparascope, High-End
RT2-14-04 | VIDEO SYSTEM, LAPAROSCOPIC (3D High-End)
1 | DESCRIPTION
1.1 | Laparoscopic camera stack used to examine the organs inside the abdomen
1a | Product Offered:
1a.1 | Make:
1a.2 | Model:
1a.3 | Initial Manufacturing year of Model:
2 | TECHNICAL SPECIFICATION
2.1 | The stack system should come complete with monitor, video processor, insufflator, camera console, data management system, laparoscopic rigid scopes, fiber optic light cable and the appropriate trolley.
3 | Video Processor
3a | Make & Model:
3.1 | An Autoclavable light guide cable of at least 5mm diameter and atleast 2.3meter length.
3.2 | Should have digital zoom facility and automatic white balancing
3.3 | Video control unit must have 1 or more camera head inputs with image display settings available
3.4 | The unit offered should have the following output configurations Composite, S-video,Digital, DVI, HDMI, BNC and AVI
4 | HD camera head (Urology specific Camera Head must be substituted and supplied for Urology)
4a | Make & Model:
4.1 | Controlled on camera console and head with programmable buttons for capture, zoom, white balancing.
4.2 | A 3 Chip CCD or CMOS, HD Camera Head with a "clip-on" scope mounting design shall be provided.
4.3 | Camera head should have a minimum of three programmable switches to allow user remote control
4.4 | Freezing facility (Programmable Button)
4.5 | White balancing function must be available on the camera
4.6 | The length of the camera head cable must be between 3m and 5 m
4.7 | It must be autoclavable or other means sterilization, please state the method
5 | CO2 Insufflator
5a | Make & Model:
5.1 | Insufflator must have both low and high flow capacity
5.2 | The unit should have display to indicate pressure in abdomen
5.3 | Visual and audible alarms are required to indicate overpressure.
5.4 | There should be an indication of the gas level in the cylinder.
5.5 | Automatic pressure relief mechanism when abdominal pressure exceeds set pressure
5.6 | Universal wrench
6 | Light Source
6a | Make & Model:
6.1 | Lamp design must be LED.
6.2 | Light surce must have adjustable light intensity
6.3 | An autoclavable light guide cable of at least 5mm diameter and 2.3m length should be provided.
7 | Monitor
7a | Make & Model:
7.1 | It should be at least 26" Flat LCD or LED medical grade monitor with the following specifications:
7.2 | The trolley should have antistatic castors with at least 2 equipped with locking system/brake.
7.3 | The unit offered should have a minimum of 3 inputs which include; Video, Y/C, RGB, DVI, HDMI
7.4 | High contrast and wide viewing angle
8 | Data Management System
8a | Make & Model:
8.1 | The unit offered should record both video and still images in full HD
8.2 | Internal memory should be a minimum of 500GB
8.3 | Two-channel simultaneous recording
8.4 | Scope button to record and capture functions
8.5 | The system offered should be compatible with the printer offered so that the recorded items can be printed out.
9 | Trolley
9.1 | The offered unit should come in a purpose built Endoscopic System trolley with a minimum of one drawer for storage
9.2 | The trolley offered should have a tray for the keyboard and mouse.
9.3 | The trolley should have antistatic castors with at least 2 equipped with locking system/brake.
9.4 | The trolley should have a handle for ease access and transportation and a cable management arm.
9.5 | The trolley should have a cover at the rear for safety of cables.
9.6 | Trolley must have a secure gas cylinder holder
9.7 | The trolley must be supplied with a minimum of 3 additional 220 V power sockets.
10 | Printer
10a | Make & Model:
10.1 | A4 or A6 (minimum resolution of 423dpi) colour printer
11 | Conformity Compliance (Please attach certificate)
11.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
11.2 | The unit must comply with electrical safety and performance standard such as IEC 60601-2-18, attached certification
11.2.1 | Please provide unique ref number of the IEC ceriticate:
11.3 | OEM must comply and certified on ISO 9001 quality standards, attach proof of compliance
11.3.1 | Please provide unique ref number of the ISO 9001 ceriticate:
11.4 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
11.4.1 | Please provide unique ref number of the ISO 13485 ceriticate:
11.5 | Model quoted for must be EC certified. Attach a copy of certification
11.5.1 | Please provide unique ref number of the EC ceriticate:
11.6 | Must comply with 2.5 Occupational Health and Safety Act, Act no 85 of 1993 and Regulations
11.7 | The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade (rubber to with 2 screws) type and it must be a minimum length of three (3) meters. NB the mains cable of the unit tendered for must be SABS colour coded.
11.8 | The equipment quoted for must be protected against electromagnetic interference IEC 60601-1-2
11.8.1 | Please provide unique ref number of the IEC ceriticate:
12 | COMPREHENSIVE MAINTENANCE
12.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) must be included in the bid price
12.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
12.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
12.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
12.3 | A 5-year service exchange for the laparoscopic rigid scopes should be quoted for separately
12.4 | The bidder to provide the OEM recommended service interval of the unit:
12.5 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
13 | General
13.1 | Has the product on offer or any of its components being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
13.2 | The unit on offer must be the latest technology.
RT2-14-04a | LAP CHOLE SET
12.1.1 | Atraumatic Dissecting and Grasping Forceps, rotating, with connector pin for unipolar coagulation, size 5 mm, length 36 cm, atraumatic, fenestrated, double action jaws, consisting of: Plastic Handle with disengageable ratchet, with larger contact area, Outer Tube insulated and Forceps Insert X2
12.1.2 | Coagulating and Dissecting Electrode, L-shaped, with connector pin for unipolar coagulation, size 5 mm, working length 36 cm X1
12.1.3 | Claw Forceps, rotating, size 5mm, length 36 cm, 2 x 3 teeth, single action jaws, consisting of handle with ratchet, Outer Tube, insulated, and Forceps Insert X1
12.1.4 | Suction and Irrigation Tube,with lateral holes, size 5 mm, length 36 cm, for use with suction and irrigation handles X1
12.1.5 | Suction irrigation handle X1
12.1.6 | Fascial Closure Instrument for subcutaneous ligature of trocar incisions, size 2.8 mm, length 17 cm X1
12.1.7 | Unipolar High Frequency Cord, with 8 mm plug, length 300 cm X1
12.1.8 | Curved long, Dissecting and Grasping Forceps, rotating, dismantling, insulated, with connector pin for unipolar coagulation, with LUER-Lock connector for cleaning, double action jaws, long, size 5 mm, length 36 cm, consisting of: Plastic Handle, without ratchet Metal Outer Sheath, insulated and Forceps Insert X1 OR
12.1.9 | Metzenbaum Scissors, rotating, dismantling, with connector pin for unipolar coagulation, with LUER-Lock irrigation connector for cleaning, double action jaws, spoon-shaped jaws, serrated, curved, length of jaws 20 mm, size 5 mm, length 36 cm consisting of: Plastic Handle, without ratchet, Metal Outer Sheath, insulated. and Scissors Insert, for use with trocars size 6 mm X1
RT2-14-04b | LAP NISSEN SET
12.2.1 | Atraumatic Dissecting and Grasping Forceps, rotating, with connector pin for unipolar coagulation, size 5 mm, length 36 cm, atraumatic, fenestrated, double action jaws, consisting of: Plastic Handle with disengageable ratchet, with larger contact area, Outer Tube insulated and Forceps Insert X2
12.2.2 | Curved long, Dissecting and Grasping Forceps, rotating, dismantling, insulated, with connector pin for unipolar coagulation, with LUER-Lock connector for cleaning, double action jaws, long, size 5 mm, length 36 cm, consisting of: Plastic Handle, without ratchet Metal Outer Sheath, insulated and Forceps Insert X1
12.2.3 | Metzenbaum Scissors, rotating, dismantling, with connector pin for unipolar coagulation, with LUER-Lock irrigation connector for cleaning, double action jaws, spoon-shaped jaws, serrated, curved, length of jaws 20 mm, size 5 mm, length 36 cm consisting of: Plastic Handle, without ratchet, Metal Outer Sheath, insulated. and Scissors Insert, for use with trocars size 6 mm X1
12.2.4 | Bowel Grasper, fenestrated, double action jaws, size 5mm,length 36 cm X1
12.2.5 | Liver retractor X1
12.2.6 | Suction and Irrigation Tube,with lateral holes, size 5 mm, length 36 cm, for use with suction and irrigation handles X1
12.2.7 | Suction irrigation handle X1
12.2.8 | Fascial Closure Instrument for subcutaneous ligature of trocar incisions, size 2.8 mm, length 17 cm X1
12.2.9 | Unipolar High Frequency Cord, with 8 mm plug, length 300 cm X1
RT2-14-04c | LAP HERNIA SET
12.3.1 | Atraumatic Grasping Forceps rotating, dismantling, insulated, with connector pin for unipolar coagulation, single action jaws, with especially fine serration, fenestrated, size 5 mm, length 36 cm consisting of:PlasticHandle,one with and one without ratchet, Metal OuterSheath, insulated and Forceps Insert X2
12.3.2 | Blunt Nose Dissecting and Grasping Forceps, rotating, with connector pin for unipolar coagulation, size 5 mm, length 36 cm, heavy, double action jaws, consisting of: Plastic Handle, without ratchet, Outer Tube, insulated and Forceps Insert X1
12.3.3 | BABCOCK Clamp, rotating, size 10 mm, length 36 cm, double action jaws, consisting of: Metal Handle, with hemostat style ratchet, Outer Tube, insulated and Forceps Insert X1
12.3.4 | Right Angled Dissecting Forceps, rotating, size 5mm , length 36 cm, 2 x 3 teeth, single action jaws, consisting of: Metal Handle, Outer Tube, insulated and Forceps Insert X1
12.3.5 | METZENBAUM Scissors, rotating, dismantling, insulated, with connector pin for unipolar coagulation, with LUER-Lock connector for cleaning, double action jaws, curved, length of blades 15 mm, size 5 mm, length 36 cm consisting of: Plastic Handle, insulated, without ratchet Metal Outer Sheath, insulated and Scissors Insert for use with trocars size 6 mm X1
12.3.6 | Curved Bipolar Grasping Forceps, rotating, dismantling, with connector pin for bipolar coagulation, especially suitable for dissection, double action jaws, size 5 mm, length 36 cm, consisting of:Plastic Handle, without ratchet, Outer Sheath and Forceps Insert X1
12.3.7 | Coagulating and Dissecting Electrode, with channel, L-shaped size 5 mm, length 36 cm X1
12.3.8 | Suction and Irrigation Tube, with lateral holes, with two-way stopcock for single-hand control, size 5 mm, length 36 cm X1
12.3.9 | Suction and Irrigation Handle X1
12.3.10 | Knot Tier, size 5 mm, length 36 cm, for extracorporeal knotting X1
12.3.11 | Fascial Closure Instrument for subcutaneous ligature of trocar incisions, size 2.8 mm, length 17 cm X1
12.3.12 | Unipolar High Frequency Cord, with 4 mm plug, length 300 cm X1
12.3.13 | Bipolar Cable X1
RT2-14-04d | COLORECTAL SET
12.4.1 | Atraumatic Grasping Forceps rotating, dismantling, insulated, with connector pin for unipolar coagulation, single action jaws, with especially fine serration, fenestrated, size 5 mm, length 36 cm consisting of:PlasticHandle, without ratchet, Metal OuterSheath, insulated and Forceps Insert X2
12.4.2 | Atraumatic Grasping Forceps, rotating, with connector pin for unipolar coagulation, size 5 mm, length 36 cm, atraumatic, fenestrated, single action jaws, consisting of: Plastic Handle with disengageable ratchet, with larger contact area, Outer Tube, insulated and Forceps Insert X1
12.4.3 | Atraumatic Dissecting and Grasping Forceps, rotating, with connector pin for unipolar coagulation, size 5 mm, length 36 cm, atraumatic, fenestrated, double action jaws, consisting of: Plastic Handle, with larger contact area, Outer Tube, insulated and Forceps Insert X1
12.4.4 | Atraumatic Grasping Forceps, rotating, size 5 mm, length 36 cm, double action,consisting of: Plastic-Handle with disengageable ratchet, with larger contact area, Outer Tube, insulated and Forceps Insert X1
12.4.5 | BABCOCK Grasping Forceps, rotating, with connector pin for unipolar coagulation, size 10mm, length 36 cm, atraumatic, jaws with multiple teeth, fenestrated, single action jaws, consisting of: Plastic Handle with disengageable ratchet, with larger contact area, Outer Tube, insulated and Forceps Insert X1
12.4.6 | Right Angled Dissecting Forceps, rotating, with connector pin for unipolar coagulation, size 5 mm, length 36 cm, atraumatic, jaws with multiple teeth, fenestrated, single action jaws, consisting of: Plastic Handle with disengageable ratchet, with larger contact area, Outer Tube, insulated and Forceps Insert X1
12.4.7 | BABCOCK Clamp, rotating, size 10 mm, length 36 cm, double action jaws, consisting of: Metal Handle, with hemostat style ratchet, Outer Tube, insulated and Forceps Insert X1
12.4.8 | Claw Forceps, rotating, size 10 mm, length 36 cm, 2 x 3 teeth, single action jaws, consisting of: Metal Handle, with hemostat style ratchet, Outer Tube, insulated and Forceps Insert X1
12.4.9 | METZENBAUM Scissors, rotating, dismantling, insulated, with connector pin for unipolar coagulation, with LUER-Lock connector for cleaning, double action jaws, curved, length of blades 15 mm, size 5 mm, length 36 cm consisting of: Plastic Handle, insulated, without ratchet Metal Outer Sheath, insulated and Scissors Insert for use with trocars size 6 mm X1
12.4.10 | METZENBAUM Scissors, curved, length of blades, double action jaws, size 5 mm, length 36 cm X1
12.4.11 | Bowel Grasper, fenestrated, double action jaws, length of jaws 18mm, size 5mm, length 36 cm X2
12.4.12 | MANHES Grasping Forceps,rotating, dismantling, insulated, with connector pin for unipolar coagulation, with LUER-Lock irrigation connector for cleaning, single action jaws, with multiple teeth, width of jaws 4.8 mm, for atraumatic and accurate grasping, size 5 mm,length 36 cm,consisting of Plastic Handle, with ratchet, with larger contact area,Metal Outer Sheath,Forceps Insert X1
12.4.12 | Knot Tier, size 5 mm, length 36 cm, for extracorporeal knotting X1
12.4.13 | Fascial Closure Instrument for subcutaneous ligature of trocar incisions, size 2.8 mm, length 17 cm X1
12.4.14 | Unipolar High Frequency Cord, with 4 mm plug, length 300 cm, X1
12.4.15 | Bipolar Cable X2
RT2-14-04e | GYNAE SET
12.18 | VERESS Pneumoperitoneum needle, with spring -loaded blunt stylet, LUER-Lock, length 10cm X1
12.19 | Trocar and Cannula 6mm X3
12.2 | Trocar and Cannula 11mm X2
12.21 | Reducer 11mm to 5mm X1
12.22 | Curved long, Dissecting and Grasping Forceps, rotating, dismantling, insulated, with connector pin for unipolar coagulation, with LUER-Lock connector for cleaning, double action jaws, long, size 5 mm, length 36 cm, consisting of: Plastic Handle, without ratchet Metal Outer Sheath, insulated and Forceps Insert X1
12.23 | Curved short, Dissecting and Grasping Forceps, rotating, dismantling, insulated, with connector pin for unipolar coagulation, with LUER-Lock connector for cleaning, double action jaws, long, size 5 mm, length 36 cm, consisting of: Plastic Handle, without ratchet, Metal Outer Sheath, insulated and Forceps Insert X1
12.24 | Atraumatic Grasping Forceps rotating, dismantling, insulated, with connector pin for unipolar coagulation, single action jaws, with especially fine serration, fenestrated, size 5 mm, length 36 cm consisting of:PlasticHandle, without ratchet, Metal OuterSheath, insulated and Forceps Insert X1
12.25 | Atraumatic Dissecting and Grasping Forceps, rotating, with connector pin for unipolar coagulation, size 5 mm, length 36 cm, atraumatic, fenestrated, double action jaws, consisting of: Plastic Handle with disengageable ratchet, with larger contact area, Outer Tube insulated and Forceps Insert X1
12.26 | Lymph node Grasping Forceps, rotating, with connector pin for unipolar coagulation, size 5 mm, length 36 cm, atraumatic, single action jaws, consisting of: Plastic Handle with disengageable ratchet, with larger contact area Outer Tube, insulated, and Forceps Insert X1
12.27 | Grasping Forceps, rotating, with connector pin for unipolar coagulation, size 5 mm, length 36 cm, jaws with multiple teeth, width of jaws 4,8 mm, single action jaws, consisting of: 33156 Plastic-Handle with disengageable ratchet, with larger contact area 33300 Outer Tube, insulated 33310 ME Forceps Insert X1
12.28 | BABCOCK Grasping Forceps, rotating, with connector pin for unipolar coagulation, size 5 mm, length 36 cm, atraumatic, jaws with multiple teeth, fenestrated, single action jaws, consisting of: Plastic Handle with disengageable ratchet, with larger contact area, Outer Tube, insulated and Forceps Insert X1
12.29 | Claw Forceps, rotating, size 5mm, length 36 cm, 2 x 3 teeth, single action jaws, consisting of: Metal Handle, with hemostat style ratchet, Outer Tube, insulated, and Forceps Insert X1
12.3 | Metzenbaum Scissors, rotating, dismantling, with connector pin for unipolar coagulation, with LUER-Lock irrigation connector for cleaning, double action jaws, spoon-shaped jaws, serrated, curved, length of jaws 20 mm, size 5 mm, length 36 cm consisting of: Plastic Handle, without ratchet, Metal Outer Sheath, insulated. and Scissors Insert for use with trocars size 6 mm X1
12.31 | Biopsy Forceps, rotating, dismantling, insulated, with connector pin for unipolar coagulation, with LUER lock adaptor for cleaning, large jaws, single action jaws, size 5 mm, length 36 cm consisting of: Plastic Handle, without ratchet, Outer Sheath, insulated, and Forceps Insert X1
12.32 | Curved Bipolar Grasping Forceps, rotating, dismantling, with connector pin for bipolar coagulation, especially suitable for dissection, double action jaws, size 5 mm, length 36 cm, consisting of:Plastic Handle, without ratchet, Metal Outer Sheath and Forceps Insert X1
12.33 | Fenestrated Bipolar Grasping Forceps, rotating, dismantling, with connector pin for bipolar coagulation, with especially fine atraumatic serration, fenestrated jaws, double action jaws, size 5 mm, length 36 cm, consisting of: Plastic Handle, without ratchet Metal Outer Sheath and Forceps Insert X1
12.34 | Coagulating and Dissecting Electrode, spatula-sheaped, blunt with connector pin for unipolar coagulation, size 5 mm, working length 36 cm X1
12.35 | Coagulating and Dissecting Electrode, L-shaped, with connector pin for unipolar coagulation, size 5 mm, working length 36 cm X1
12.36 | Palpation Probe X1
12.37 | Myoma Fixation Instrument, screw shaped, ø 5 mm X1
12.38 | Suction and Irrigation Tube, with lateral holes, with two-way stopcock for single-hand control, size 5 mm, length 36 cm X1
12.39 | Needle Holder, ergonomic axial handle with disengageable ratchet, ratchet release on the right side, left curved jaws, with tungsten carbide insert ø 5 mm, length 33 cm X1
12.4 | Needle Holder, ergonomic axial handle with disengageable ratchet, ratchet release on the left side, right curved jaws, with tungsten carbide insert ø 5 mm, length 33 cm X1
12.41 | Knot Tier, size 5 mm, length 36 cm, for extracorporeal knotting X1
12.42 | Fascial Closure Instrument for subcutaneous ligature of trocar incisions, size 2.8 mm, length 17 cm X1
12.43 | Uterine manipulator X1
12.44 | Unipolar High Frequency Cord, with 8 mm plug, length 300 cm X2
12.45 | Bipolar High Frequency Cord, length 300 cm X2
12.46 | Fiber Optic Light Cable, with straight connector, diameter 4.8 mm, length 300 cm X1
12.47 | Container X1
RT2-14-04f | UROLOGY SET
11.6 | Curved long, Dissecting and Grasping Forceps, rotating, dismantling, insulated, with connector pin for unipolar coagulation, with LUER-Lock connector for cleaning, double action jaws, long, size 5 mm, length 36 cm, consisting of: Plastic Handle, without ratchet Metal Outer Sheath, insulated and Forceps Insert X1
11.7 | Curved short, Dissecting and Grasping Forceps, rotating, dismantling, insulated, with connector pin for unipolar coagulation, with LUER-Lock connector for cleaning, double action jaws, long, size 5 mm, length 36 cm, consisting of: Plastic Handle, without ratchet, Metal Outer Sheath, insulated and Forceps Insert X1
11.8 | Atraumatic Grasping Forceps rotating, dismantling, insulated, with connector pin for unipolar coagulation, single action jaws, with especially fine serration, fenestrated, size 5 mm, length 36 cm consisting of:PlasticHandle, without ratchet, Metal OuterSheath, insulated and Forceps Insert X1
11.9 | Atraumatic Grasping Forceps, rotating, with connector pin for unipolar coagulation, size 5 mm, length 36 cm, atraumatic, fenestrated, single action jaws, consisting of: Plastic Handle with disengageable ratchet, with larger contact area, Outer Tube, insulated and Forceps Insert X1
11.1 | Atraumatic Dissecting and Grasping Forceps, rotating, with connector pin for unipolar coagulation, size 5 mm, length 36 cm, atraumatic, fenestrated, double action jaws, consisting of: Plastic Handle, with larger contact area, Outer Tube, insulated and Forceps Insert X1
11.11 | Atraumatic Dissecting and Grasping Forceps, rotating, with connector pin for unipolar coagulation, size 5 mm, length 36 cm, atraumatic, fenestrated, double action jaws, consisting of: Plastic Handle with disengageable ratchet, with larger contact area, Outer Tube, insulated and Forceps Insert X1
11.12 | Grasping Forceps, rotating, with connector pin for unipolar coagulation, size 5 mm, length 36 cm, "Tiger-jaws", 2x4 teeth, single action jaws, consisting of: Plastic-Handle with disengageable ratchet, with larger contact area, Outer Tube, insulated and Forceps Insert X1
11.13 | Blunt Nose Dissecting and Grasping Forceps, rotating, with connector pin for unipolar coagulation, size 5 mm, length 36 cm, heavy, double action jaws, consisting of: Plastic Handle, without ratchet, Outer Tube, insulated and Forceps Insert X1
11.14 | BABCOCK Grasping Forceps, rotating, with connector pin for unipolar coagulation, size 5 mm, length 36 cm, atraumatic, jaws with multiple teeth, fenestrated, single action jaws, consisting of: Plastic Handle with disengageable ratchet, with larger contact area, Outer Tube, insulated and Forceps Insert X1
11.15 | Right Angled Dissecting Forceps, rotating, with connector pin for unipolar coagulation, size 5 mm, length 36 cm, atraumatic, jaws with multiple teeth, fenestrated, single action jaws, consisting of: Plastic Handle with disengageable ratchet, with larger contact area, Outer Tube, insulated and Forceps Insert X1
11.16 | BABCOCK Clamp, rotating, size 10 mm, length 36 cm, double action jaws, consisting of: Metal Handle, with hemostat style ratchet, Outer Tube, insulated and Forceps Insert X1
11.17 | Claw Forceps, rotating, size 10 mm, length 36 cm, 2 x 3 teeth, single action jaws, consisting of: Metal Handle, with hemostat style ratchet, Outer Tube, insulated and Forceps Insert X1
11.18 | Right Angled Dissecting Forceps, rotating, size 10 mm, length 36 cm, 2 x 3 teeth, single action jaws, consisting of: Metal Handle, Outer Tube, insulated and Forceps Insert X1
11.19 | METZENBAUM Scissors, rotating, dismantling, insulated, with connector pin for unipolar coagulation, with LUER-Lock connector for cleaning, double action jaws, curved, length of blades 15 mm, size 5 mm, length 36 cm consisting of: Plastic Handle, insulated, without ratchet Metal Outer Sheath, insulated and Scissors Insert for use with trocars size 6 mm X1
11.2 | Scissors Insert, METZENBAUM Scissors, curved, length of blades 12 mm, double action jaws, size 5 mm, length 36 cm X1
11.21 | Straight Scissors, rotating, with connector pin for unipolar coagulation, size 5 mm, length 36 cm, straight, double action jaws, consisting of: Plastic Handle, without ratchet, Metal Outer Sheath, Scissors Insert X1
11.22 | Click Line Scissors Insert, straight, double action jaws, size 5 mm, length 36 cm X1
11.23 | Curved Bipolar Grasping Forceps, rotating, dismantling, with connector pin for bipolar coagulation, especially suitable for dissection, double action jaws, size 5 mm, length 36 cm, consisting of:Plastic Handle, without ratchet, Metal Outer Sheath and Forceps Insert X1
11.24 | Fenestrated Bipolar Grasping Forceps, rotating, dismantling, with connector pin for bipolar coagulation, with especially fine atraumatic serration, fenestrated jaws, double action jaws, size 5 mm, length 36 cm, consisting of: Plastic Handle, without ratchet Metal X1Outer Sheath and Forceps Insert X1
11.25 | Coagulating and Dissecting Electrode, with channel, L-shaped size 5 mm, length 36 cm, for use with suction and irrigation handles X1
11.26 | Handle with Trumpet Valve, for suction or irrigation, autoclavable, for use with coagulation suction tubes size 5 mm X1
11.27 | Suction and Irrigation Tube, with lateral holes, with two-way stopcock for single-hand control, size 5 mm, length 36 cm X1
11.28 | Puncture Needle, LUER-lock, diameter 1.6 mm, size 5 mm, length 36 cm X1
11.29 | Knot Tier, size 5 mm, length 36 cm, for extracorporeal knotting X1
11.3 | Fascial Closure Instrument for subcutaneous ligature of trocar incisions, size 2.8 mm, length 17 cm X1
11.31 | Unipolar High Frequency Cord, with 4 mm plug, length 300 cm X2
11.32 | Bipolar Cable X1
11.33 | Plastic Container for Sterilizing and Storage, perforated, with lid, for two-level storage, for use with 30 cm and 36 cm Click Line- insruments, external dimensions ( w x d x h ): 600 mm x 145 mm x 260 mm ( X1)

## Laparascope Instruments
RT2-14-05 | LAPSCOPE INSTRUMENTS
RT2-14-05a | Straight Forward Telescope 0° enlarged view, diameter 10 mm, length 29-31 cm, autoclavable, fiber optic light transmission incorporated
RT2-14-05b | Straight Forward Telescope 0° enlarged view, diameter 5 mm, length 29-31 cm, autoclavable, fiber optic light transmission incorporated
RT2-14-05c | Straight Forward Telescope 0° enlarged view, diameter 2.5 mm, length 29-31 cm, autoclavable, fiber optic light transmission incorporated
RT2-14-05d | Forward-Oblique Telescope 30° enlarged view, diameter 10 mm, length 29-31 cm, autoclavable, fiber optic light transmission incorporated
RT2-14-05e | Forward-Oblique Telescope 30° enlarged view, diameter 5 mm, length 29-31 cm, autoclavable, fiber optic light transmission incorporated
RT2-14-05f | Forward-Oblique Telescope 30° enlarged view, diameter 2.5 mm, length 29-31 cm, autoclavable, fiber optic light transmission incorporated
RT2-14-05g | Plastic Container for Sterilizing and Storage, perforated, with lid, for two-level storage, for use with 30 cm and 36 cm intruments , external dimensions ( w x d x h ): 600 mm x 145 mm x 260 mm
RT2-14-05h | Laparoscopic Needle Holder
RT2-14-05i | VERESS Pneumoperitoneum needle, with spring -loaded blunt stylet, LUER-Lock, length 10cm
RT2-14-05j | Trocar size, 11mm,with insufflation stopcock, Cannula with thread, length 100mm(max 110mm)
RT2-14-05k | Trocar size, 3.5mm,with insufflation stopcock, Cannula with thread, length 100mm(max 110mm)
RT2-14-05l | Trocar size, 6mm,with insufflation stopcock, Cannula with thread, length 100mm(max 110mm)

## Laparascope+Exoscope, Neuro
RT2-14-06 | ENDOSCOPIC & EXOSCOPIC, VIDEO SYSTEM, NEUROSURGERY
1 | DESCRIPTION
1.1 | An endoscopic stack system with appropriate endoscopes that have been designed for use in Neurosurgical procedures
1a | Product Offered:
1a.1 | Make:
1a.2 | Model:
1a.3 | Initial Manufacturing year of Model:
1a.4 | Radiation Control License number:
2 | TECHNICAL SPECIFICATION
2.1 | The stack system should come complete with monitor, video processor, Exoscope, camera console, data management system, rigid scopes, 3D fiber optic light cable and the appropriate trolley.
2.2 | System must have ability to process ICG fluorescent imaging
2.3 | System must have module for 4K endoscopy
2.4 | System must have module for 3D endoscopy
2.5 | 3D exoscope with control unit and illumination source
2.6 | Exoscope holding stand must be provided
3 | Video Processor
3a | Make & Model:
3.1 | An Autoclavable light guide cable of at least 5mm diameter and atleast 2.3meter length.
3.2 | Should have digital zoom facility and automatic white balancing
3.3 | Video control unit must have 1 or more camera head inputs with image display settings available and 3D video endoscopy capability
3.4 | The unit offered should have the following output configurations Composite, S-video,Digital, DVI, HDMI, BNC and AVI
4 | HD camera head
4a | Make & Model:
4.1 | Controlled on camera console and head with programmable buttons for capture, zoom, white balancing.
4.2 | A 4K Camera Head with a "clip-on" scope mounting design shall be provided.
4.3 | Camera head should have a minimum of three programmable switches to allow user remote control
4.4 | Freezing facility (Programmable Button)
4.5 | White balancing function must be available on the camera
4.6 | The length of the camera head cable must be between 3m and 5 m
4.7 | It must be autoclavable or other means sterilization, please state the method
5 | Light Source
5a | Make & Model:
5.1 | Lamp design must be LED.
5.2 | Light source must have adjustable light intensity
5.3 | An autoclavable light guide cable of at least 5mm diameter and 2.3m length should be provided.
5.4 | Fibreoptic light cable for 3D endoscopy must be provided
5.5 | Fibreoptic light cable for rigid endoscopy must be provided
6 | Monitor
6a | Make & Model:
6.1 | It should be at least 32" Flat LCD or LED medical grade monitor with the following specifications:
6.2 | The trolley should have antistatic castors with at least 2 equipped with locking system/brake.
6.3 | The unit offered should have a minimum of 3 inputs which include; Video, Y/C, RGB, DVI, HDMI
6.4 | High contrast and wide viewing angle
7 | Data Management System
7a | Make & Model:
7.1 | The unit offered should record both video and still images in full HD
7.2 | Internal memory should be a minimum of 1TB
7.3 | Two-channel simultaneous recording
7.4 | Scope button to record and capture functions
7.5 | The system offered should be compatible with the printer offered so that the recorded items can be printed out.
8 | Trolley
8.1 | The offered unit should come in a purpose built Endoscopic System trolley with a minimum of one drawer for storage
8.2 | The trolley offered should have a tray for the keyboard and mouse.
8.3 | The trolley should have antistatic castors with at least 2 equipped with locking system/brake.
8.4 | The trolley should have a handle for ease access and transportation and a cable management arm.
8.5 | The trolley should have a cover at the rear for safety of cables.
8.6 | Trolley must have a secure gas cylinder holder
8.7 | The trolley must be supplied with a minimum of 3 additional 220 V power sockets.
9 | Printer
9a | Make & Model:
9.1 | A4 or A6 (minimum resolution of 423dpi) colour printer
10 | STANDARD ACCESSORIES & INSTRUMENTS
10.1 | Neurosurgery endoscopes
10.2 | Tray for cleaning, sterilising and storing endoscopes
10.3 | 3D glasses or clip-ons
11 | Conformity Compliance (Please attach certificate)
11.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
11.2 | The unit must comply with electrical safety and performance standard such as IEC 60601-2-18, attached certification
11.2.1 | Please provide unique ref number of the IEC ceriticate:
11.3 | OEM must comply and certified on ISO 9001 quality standards, attach proof of compliance
11.3.1 | Please provide unique ref number of the ISO 9001 ceriticate:
11.4 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
11.4.1 | Please provide unique ref number of the ISO 13485 ceriticate:
11.5 | Model quoted for must be EC certified. Attach a copy of certification
11.5.1 | Please provide unique ref number of the EC ceriticate:
11.6 | The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade (rubber to with 2 screws) type and it must be a minimum length of three (3) meters. NB the mains cable of the unit tendered for must be SABS colour coded.
11.7 | The equipment quoted for must be protected against electromagnetic interference IEC 60601-1-2
11.7.1 | Please provide unique ref number of the IEC ceriticate:
12 | COMPREHENSIVE MAINTENANCE
12.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) must be included in the bid price
12.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
12.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
12.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
12.3 | A 5-year service exchange for the laparoscopic rigid scopes should be quoted for separately
12.4 | The bidder to provide the OEM recommended service interval of the unit:
12.5 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
13 | General
13.1 | Has the product on offer or any of its components being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
13.2 | The unit on offer must be the latest technology.

## Endoscopic System
RT2-14-07 | VIDEO ENDOSCOPIC SYSTEM
1 | DESCRIPTION
1.1 | To help diagnose and treat gastroeasophageal pathology in patients.
1a | Product Offered:
1a.1 | Make:
1a.2 | Model:
1a.3 | Initial Manufacturing year of Model:
2 | TECHNICAL SPECIFICATION
2.1 | Light Source and Video Processor:
2.2 | A light source and processor, either as a combined or independent devices, to suite both adult and paediatric and flexible Gastroscopes.
3 | Video Processor
3a | Make & Model:
3.1 | The processor must be able to interpret a video signal from an endoscope by using a colour chip technology.
3.2 | White balance adjustment with storage function
3.3 | Pre-Freeze image function to select the clearest still image
3.4 | HD Output for compatibility to HD monitor. Compatible with analogy, HDMI, HD-SDI, and DVI output.
3.5 | USB port is required on the unit
3.6 | Image Enhancement Setting: Structural and Edge Enhancement
3.7 | Image Size Selection: Can be selected from keyboard or remote switch
3.8 | Image Contrast Settings: 3 (Normal, High, Low)
4 | Light Source
4a | Make & Model:
4.1 | Lamp design must be LED or Xenon
4.2 | If Xenon, 300W with spare lamp must be provided
4.3 | The unit offered should allow for brightness adjustment
5 | Monitor
5a | Make & Model:
5.1 | It should be at least 26" Flat LCD or LED medical grade monitor with the following specifications:
5.2 | Resolution of at least 1920 x 1200 pixels
5.3 | The unit offered should have a minimum of 3 inputs which include; Video, Y/C, RGB, DVI, HDMI
5.4 | High contrast and wide viewing angle
6 | Printer
6a | Make & Model:
6.1 | A4 - A6 colour printer
6.2 | Resolution of 423dpi
7 | Image Management System
7a | Make & Model:
7.1 | The unit offered should record both video and still images in full HD
7.2 | Supported formats: Internal HDD, Blue Ray, DVD, USB HDD
7.3 | Internal memory should be a minimum of 350GB. External memory to augment internal memory will not be accepted
7.4 | Two-channel simultaneous recording. Bidder to demonstrate during evaluation
7.5 | Scope button to record and capture functions
7.6 | The trolley should have antistatic castors with at least 2 equipped with locking system/brake.
8 | Trolley
8.1 | The offered unit should come in a purpose built Endoscopic System trolley with a minimum of one drawer for storage
8.2 | The trolley offered should have a tray for the keyboard and mouse.
8.3 | The trolley should have antistatic castors with at least 2 equipped with locking system/brake.
8.4 | The trolley should have a handle for ease access and transportation and a cable management arm.
8.5 | The trolley should have a cover at the rear for safety of cables.
8.6 | The trolley must be supplied with a minimum of 3 additional 220 V power sockets.
9 | STANDARD ACCESSORIES
9.1 | Video Connecting Cable
9.2 | Irrigation tube
9.3 | Suction valve
9.4 | Suction bottle
9.5 | Scope Sterilization/Cleaning trolley
9.6 | Endoscopy Cart/trolley
9.7 | Mouth guard
9.8 | Leak tester
10 | IT NETWORKING, CONNECTIVITY AND POWER REQUIREMENTS
10.1 | An RS232/RJ45 data output port with a freely available data format for connection to another monitor or hospital information system shall be provided
10.2 | The unit must operate from a 220 V, 50 Hz. supply
110.31.2 | The mains cable at least three metres long with SABS approved 15A three prong plug.
11 | Conformity Compliance (Please attach certificate)
11.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
11.2 | The unit must comply with electrical safety and performance standard such as IEC 60601-2-18, attached certification
11.2.1 | Please provide unique ref number of the IEC ceriticate:
11.3 | OEM must comply and certified on ISO 9001 quality standards, attach proof of compliance
11.3.1 | Please provide unique ref number of the ISO 9001 ceriticate:
11.4 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
11.4.1 | Please provide unique ref number of the ISO 13485 ceriticate:
11.5 | Model quoted for must be EC certified. Attach a copy of certification
11.5.1 | Please provide unique ref number of the EC ceriticate:
11.6 | The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade (rubber to with 2 screws) type and it must be a minimum length of three (3) meters. NB the mains cable of the unit tendered for must be SABS colour coded.
11.7 | The equipment quoted for must be protected against electromagnetic interference IEC 60601-1-2
11.7.1 | Please provide unique ref number of the IEC ceriticate:
12 | COMPREHENSIVE MAINTENANCE
12.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) must be included in the bid price
12.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
12.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
12.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
12.3 | A 5-year service exchange for the laparoscopic rigid scopes should be quoted for separately
12.4 | The bidder to provide the OEM recommended service interval of the unit:
12.5 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
13 | General
13.1 | Has the product on offer or any of its components being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
13.2 | The unit on offer must be the latest technology.
RT2-14-07a | Flexible Gastroscope, Adult (Compatible & Series award with RT2-14-07)
Make & Model:
1.1 | Instrument Channel: > 2.8mm
1.2 | Insertion tube diameter: < 10mm
1.3 | Working length: 1020mm - 1200mm
1.4 | Field or angle of view: 1400
1.5 | Angulation: minimum of up: 210, Down: 90, Right: 100, Left 100
1.6 | Depth of View: 5 - 100mm
1.7 | Auxillary water jet channel as optional
1.8 | Connector must be completelly water proof
2 | STANDARD ACCESSORIES
2.1 | Scope case
RT2-14-07b | Flexible Gastroscope, Paediatric (Compatible & Series award with RT2-14-07)
Make & Model:
1.1 | Instrument Channel: Greater than 2mm
1.2 | Insertion tube diameter: < 6mm
1.3 | Working length: 1650mm or more
1.4 | Field or angle of view: 1400
1.5 | Angulation: minimum of up: 180, Down: 160, Right: 160, Left 160
1.6 | Depth of View: 3mm - 100mm
1.7 | Auxillary water jet channel as optional
1.8 | Connector must be completelly water proof
2 | STANDARD ACCESSORIES
2.1 | Scope case
RT2-14-07c | Flexible Bronchoscope, Adult (Compatible & Series award with RT2-14-07)
Make & Model:
1.1 | Instrument Channel: ≥2.2mm
1.2 | Insertion tube diameter: ≥6.2mm
1.3 | Working length: 600mm minimum
1.4 | Field or angle of view: 1200
1.5 | Angulation of at least: 100 0 up and 100 0 down
1.6 | Depth of View: 3mm - 100mm
2 | STANDARD ACCESSORIES
2.1 | Scope case
RT2-14-07d | Flexible Bronchoscope, Paediatric (Compatible & Series award with RT2-14-07)
Make & Model:
1.1 | Instrument Channel: ≥2mm
1.2 | Insertion tube diameter: < 4.2mm
1.3 | Working length: 600mm maximum
1.4 | Field or angle of view: 90o - 1400
1.5 | Angulation of at least: 100 0 up and 100 0 down
1.6 | Depth of View: 3mm - 100mm
2 | STANDARD ACCESSORIES
2.1 | Scope case
RT2-14-07e | Flexible Bronchoscope, Neonatal (Compatible & Series award with RT2-14-07)
Make & Model:
1.1 | Instrument Channel: < 1.2mm
1.2 | Insertion tube diameter: < 2.8 mm
1.3 | Working length: 600mm maximum
1.4 | Field or angle of view: 1100
1.5 | Angulation of at least: 100 0 up and 100 0 down
1.6 | Depth of View: 2mm- 50mm
2 | STANDARD ACCESSORIES
2.1 | Scope case
RT2-14-07f | Flexible Colonoscope, Adult (Compatible & Series award with RT2-14-07)
Make & Model:
1.1 | Instrument Channel: > 3.7mm
1.2 | Insertion tube diameter: < 14mm
1.3 | Working length: < 1700mm
1.4 | Field or angle of view: > 140o
1.5 | Angulation: Up & Down: >180o ; Left & Right: >130o
1.6 | Depth of View: 5 - 100mm
1.7 | Must have auxillary water jet channel
1.8 | Connector must be completelly water proof with no need for additional waterproof caps
2 | STANDARD ACCESSORIES
2.1 | Scope case
RT2-14-07g | Flexible Colonoscope, Paediatric (Compatible & Series award with RT2-14-07)
Make & Model:
1.1 | Instrument Channel: > 2.8mm
1.2 | Insertion tube diameter: < 9.8mm
1.3 | Working length: < 1700mm
1.4 | Field or angle of view: > 140o
1.5 | Angulation: Up & Down: >180o ; Left & Right: >130o
1.6 | Depth of View: 5 - 100mm
1.7 | Must have auxillary water jet channel
1.8 | Connector must be completelly water proof with no need for additional waterproof caps
2 | STANDARD ACCESSORIES
2.1 | Scope case
RT2-14-07h | Endoscope Replacement Accessories (Compatible with RT2-14-07)
RT2-14-07h | Video Connecting Cable
RT2-14-07i | Irrigation tube
RT2-14-07k | Suction valve
RT2-14-07l | Suction bottle
RT2-14-07m | Scope Sterilization/Cleaning trolley
RT2-14-07n | Endoscopy Cart/trolley
RT2-14-07o | Mouth guard
RT2-14-07p | Leak tester
RT2-14-07q | Brushes
RT2-14-07r | Brush handles
RT2-14-07s | Suction catheter adaptor