## Company Information
Company name:
CSD Number:
Pysical Address:
Postal Address:
Contact Person:
Telephone No:
Cell No:
Email Address:

## Profile
7. ORTHOPAEDIC EQUIPMENT
# | Item number | Item Description
1 | RT2-07-01 | Electric Drill, Small bones
2 | RT2-07-02 | Electric Drill, Large bones
3 | RT2-07-03 | Battery-Operated Drill, Small bones
4 | RT2-07-04 | Battery-Operated Drill, Large bones
5 | RT2-07-05 | Tourniquet
6 | RT2-07-06 | Plaster Saw
7 | RT2-07-07 | Plaster Saw with Vacuum Extractor
8 | RT2-07-08 | Dermatome, Cordless
9 | RT2-07-09 | Dermatome, Pneumatic
10 | RT2-07-10 | Skin Graft Mesher
11 | RT2-07-11 | Arthroscopic Shaver
8. OPHTHALMOLOGY EQUIPMENT
# | Item number | Item Description
1 | RT2-08-001 | A&B Scanner Combo Unit
2 | RT2-08-002 | Auto Ref-Keratometer
3 | RT2-08-003 | Laser Unit, Argon
4 | RT2-08-004 | Laser Unit, YAG
5 | RT2-08-005 | Luminometer
6 | RT2-08-006 | Optical Coherence Tomography Unit
7 | RT2-08-007 | Phoropter
8 | RT2-08-008 | Slit-Lamp, Basic
9 | RT2-08-009 | Slit-Lamp, Advanced
10 | RT2-08-010 | Tonometer, Hand-held
11 | RT2-08-011 | Visual Field Analyser
12 | RT2-08-012 | Vitrectomy Unit
9. NEURO EQUIPMENT (NEUROLOGY & NEUROSURGERY)
# | Item number | Item Description
1 | RT2-09-001 | EEG, Portable
2 | RT2-09-002 | EEG, Desktop
3 | RT2-09-003 | Ultrasonic Surgical Aspirator
4 | RT2-09-004 | Neuronavigation System
5 | RT2-09-005 | Neuro-Monitor
6 | RT2-09-006 | Neuro-Drills
7 | RT2-09-007 | Amplitude EEG (Cerebral Function Monitoring)
12. MICROSCOPES AND OTHER ADVANCED TECHNOLOGY
# | Item number | Item Description
1 | RT2-12-001 | Microscope, Ophthalmology, Basic
2 | RT2-12-002 | Microscope, Ophthalmology, Advanced
3 | RT2-12-003 | Microscope, ENT, Basic
4 | RT2-12-004 | Microscope, ENT, Advanced
5 | RT2-12-005 | Microscope, Neurosurgery, Basic
6 | RT2-12-006 | Microscope, Neurosurgery, Advanced
7 | RT2-12-007 | Robotic Surgical System
8 | RT2-12-008 | Robotic Digital Surgical Microscope
14. MULTI-DISCIPLINARY SPECIALIZED EQUIPMENT
# | Item number | Item Description
1 | RT2-14-01 | Laparascope, Mid-range
2 | RT2-14-02 | Laparascope, High-end (3D)
3 | RT2-14-03 | Lapscope Instruments
4 | RT2-14-04 | Endoscopic & Exoscopic, Video System, Neurosurgery
5 | RT2-14-05 | Video Endoscopic System
6 | RT2-14-06 | Shear Wave Elastrograph

## Category-7 Ortho
ORTHOPAEDIC EQUIPMENT
# | Item number | Item Description
1 | RT2-07-001 | Electric Drill, Small bones
2 | RT2-07-002 | Electric Drill, Large bones
3 | RT2-07-003 | Battery-Operated Drill, Small bones
4 | RT2-07-004 | Battery-Operated Drill, Large bones
5 | RT2-07-005 | Tourniquet
6 | RT2-07-006 | Plaster Saw
7 | RT2-07-007 | Plaster Saw with Vacuum Extractor
8 | RT2-07-008 | Dermatome, Cordless
9 | RT2-07-009 | Dermatome, Pneumatic
10 | RT2-07-010 | Skin Graft Mesher
11 | RT2-07-011 | Arthroscopic Shaver

## Drill, Electric, Small bones
RT2-07-001 | Drill, Orthopaedic, Electric, Small Bones | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
1 | DESCRIPTION
1.1 | A specialized medical instrument designed for drilling, tapping, and screwing small bones, typically used in orthopaedic, podiatric and hand surgery
1a | Product Offered:
1a.1 | Make:
1a.2 | Model:
1a.3 | Initial Manufacturing year of Model:
2 | SPECIFICATIONS
2.1 | The Orthopeadic drill must be able to do the following functions:
2.1.1 | Cleaning of bones(Reaming)
2.1.2 | Tapping
2.1.3 | Fastening and Loosening of screws
2.1.4 | To insert intramedulary nails
2.1.5 | Traumatology
2.1.6 | Applying of Kirschner wires in bone surgery.
2.2 | The unit must have a handpiece Cord of a minimum length of 3 metres.
2.3 | The unit must be made out of a durable, robust light weight material for easy and smooth operating
2.4 | The unit must operate on 220Vac at 50Hz with power consumption of 500VA
2.5 | The unit must be used for small bones
2.6 | The unit quoted for must be provided with a modular Handpiece with the following attachment (Drill/Ream Mode): SPECIFY torque, speed and cannel size where applicable
2.6.1 | Drill Attachments
2.6.1a | Angled
2.6.1b | Medium straight
2.6.1c | Long Straight
2.6.1d | Long Angled
2.6.1e | Reaming / High torque attachments
2.6.1f | Reaming / High attachments
2.6.1g | TAP mode with drill attachments only
2.6.1h | Screw mode with drill attachments only
3 | Standard Accessories. SPECIFY torque, speed and cannel size where applicable
3.1 | Trinkle/AO attachments, which must accept trinkle and ASIF/AO, type accessories
3.2 | Reaming attachment, which must accept large HEX. AO Reamers
3.3 | Sagittal blade (Set of 10): Sizes as per Endusers choice
3.4 | 4.0mm Jacobs chuck and a key
3.5 | Oscillating blade (Set of 10): Sizes as per Endusers choice
3.6 | Pin driver
3.7 | Sterilization case
3.8 | Small bones blade (Set of 10): Sizes as per Endusers choice
3.9 | Wire driver
4 | Conformity Compliance (Please attach certificate)
4.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
4.2 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
4.2.1 | Please provide unique ref number of the ISO 13485 ceriticate:
4.3 | Model on offer must comply with EU Medical Device Regulation (EU) 2017/745 or equvalent from SAHPRA recognized jurisdictions . Attach a copy of certification
4.3.1 | Please provide unique ref number of the EC ceriticate:
4.4 | Hazardous Substance Act: If this type of equipment / apparatus appears on the schedule of Hazardous substances issued by the Directorate: Health Technology of the Department of Health, a license in terms of the Act on Hazardous Substances (Act 15/1973) must be submitted with the bid / quotation documents. The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
4.5 | All electrical/electronic medical equipment must be licensed by Radiation Control, where applicable a copy of the license must be submitted The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
4.5.1 | Radiation Control License number:
4.6 | The unit shall operate from 220V AC/50Hz mains supply. The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade (rubber to with 2 screws) type and it must be a minimum length of three (3) meters. NB the mains cable of the unit tendered for must be SABS colour coded.
4.7 | The equipment quoted for must be protected against electromagnetic interference IEC 60601-1-2
4.8 | The unit must comply with electrical safety and performance standard such as IEC 60601-2-2, attach certification
4.8.1 | Please provide unique ref number of the IEC ceriticate:
5 | COMPREHENSIVE MAINTENANCE
5.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) should be quoted for, as per the SCC document
5.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
5.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
5.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
5.3 | The bidder to provide the OEM recommended service interval of the unit:
5.4 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
6 | General
6.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
6.2 | The unit on offer must be the latest technology.
6.2.1 | Initial Manufacturing year of Model:
RT2-07-001 | Drill, Orthopaedic, Electric, Small Bones | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
RT2-07-001a | Trinkle / AO attachments, which must accept trinkle and ASIF / AO, type accessories
RT2-07-001b | Reaming attachment, which must accept large HEX. AO Reamers
RT2-07-001c | Sagittal blade (Set of 10): Sizes as per Endusers choice
RT2-07-001d | 4.0mm Jacobs chuck and a key
RT2-07-001e | Oscillating blade (Set of 10): Sizes as per Endusers choice
RT2-07-001f | Pin driver
RT2-07-001g | Wire driver
RT2-07-001h | Sterilization case
RT2-07-001i | Small bones blade (Set of 10): Sizes as per Endusers choice
RT2-07-001j | Compatible power system oil spray

## Drill, Electric, Large bones
RT2-07-002 | Drill, Orthopaedic, Electric, Large Bones | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
1 | DESCRIPTION
1.1 | A specialized medical instrument designed for drilling, tapping, and screwing large bones, typically used in orthopaedic, podiatric and hand surgery
2 | Product Offered:
2.1 | Make:
2.2 | Model:
2.3 | Initial Manufacturing year of Model:
3 | SPECIFICATIONS
3.1 | The Orthopeadic drill must be able to do the following functions:
3.1.1 | Cleaning of bones(Reaming)
3.1.2 | Tapping
3.1.3 | Fastening and Loosening of screws
3.1.5 | To insert intramedulary nails
3.1.7 | Traumatology
3.1.8 | Applying of Kirschner wires in bone surgery.
3.2 | The unit must have a handpiece Cord of a minimum length of 3 metres.
3.3 | The unit must be made out of a durable robust material light weight and for easy and smooth operating
3.4 | The unit must operate on 220Vac at 50Hz with power consumption of 500VA
3.5 | The unit must be used for large bones
3.6 | The unit quoted for must be provided with a modular Handpiece with the following attachment (Drill / Ream Mode): SPECIFY torque, speed and cannel size where applicable
3.6.1 | Drill Attachments
3.6.1a | Angled
3.6.1b | Medium straight
3.6.1c | Long Straight
3.6.1d | Long Angled
3.6.1e | Reaming / High torque attachments
3.6.1f | Reaming / High attachments
3.6.1g | TAP mode with drill attachments only
3.6.1h | Screw mode with drill attachments only
3 | Standard Accessories. SPECIFY torque, speed and cannel size where applicable
3.1 | Trinkle/AO attachments, which must accept trinkle and ASIF/AO, type accessories
3.2 | Reaming attachment, which must accept large HEX. AO Reamers
3.3 | Sagittal saw attachement with a set of 10 blades: Sizes as per Endusers choice
3.4 | 4.0mm Jacobs chuck and a key
3.5 | Oscillating blade (Set of 10): Sizes as per Endusers choice
3.6 | Pin driver
3.7 | Sterilization case
3.8 | Large bones blade (Set of 10): Sizes as per Endusers choice
3.9 | Wire driver
5 | Conformity Compliance (Please attach certificate)
4.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
4.2 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
4.2.1 | Please provide unique ref number of the ISO 13485 ceriticate:
4.3 | Model on offer must comply with EU Medical Device Regulation (EU) 2017/745 or equvalent from SAHPRA recognized jurisdictions . Attach a copy of certification
4.3.1 | Please provide unique ref number of the EC ceriticate:
4.4 | Hazardous Substance Act: If this type of equipment / apparatus appears on the schedule of Hazardous substances issued by the Directorate: Health Technology of the Department of Health, a license in terms of the Act on Hazardous Substances (Act 15/1973) must be submitted with the bid / quotation documents. The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
4.5 | All electrical/electronic medical equipment must be licensed by Radiation Control, where applicable a copy of the license must be submitted The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
4.5.1 | Radiation Control License number:
4.6 | The unit shall operate from 220V AC/50Hz mains supply. The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade (rubber to with 2 screws) type and it must be a minimum length of three (3) meters. NB the mains cable of the unit tendered for must be SABS colour coded.
4.7 | The equipment quoted for must be protected against electromagnetic interference IEC 60601-1-2
4.8 | The unit must comply with electrical safety and performance standard such as IEC 60601-2-2, attach certification
4.8.1 | Please provide unique ref number of the IEC ceriticate:
5 | COMPREHENSIVE MAINTENANCE
5.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) should be quoted for, as per the SCC document
5.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
5.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
5.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
5.3 | The bidder to provide the OEM recommended service interval of the unit:
5.4 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
6 | General
6.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
6.2 | The unit on offer must be the latest technology.
6.2.1 | Initial Manufacturing year of Model:
RT2-07-002 | Drill, Orthopaedic, Electric, Large Bones | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
RT2-07-002a | Trinkle / AO attachments, which must accept trinkle and ASIF / AO, type accessories
RT2-07-002b | Reaming attachment, which must accept large HEX. AO Reamers
RT2-07-002c | Sagittal blade (Set of 10): Sizes as per Endusers choice
RT2-07-002d | 4.0mm Jacobs chuck and a key
RT2-07-002e | Oscillating blade (Set of 10): Sizes as per Endusers choice
RT2-07-002f | Pin driver
RT2-07-002g | Wire driver
RT2-07-002h | Sterilization case
RT2-07-002i | Large bones blade (Set of 10): Sizes as per Endusers choice
RT2-07-002j | Compatible power system oil spray

## Drill, Battery-Op, Small bones
RT2-07-003 | Drill, Orthopaedic, Battery Operated, Small bones | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
1 | DESCRIPTION
1.1 | A specialized medical instrument designed for drilling, tapping, and screwing small bones, typically used in orthopaedic, podiatric and hand surgery
1a | Product Offered:
1a.1 | Make:
1a.2 | Model:
1a.3 | Initial Manufacturing year of Model:
3 | SPECIFICATIONS
3.1 | The Orthopeadic drill must be able to do the following functions:
3.1.1 | Cleaning of bones(Reaming)
3.1.2 | Tapping
3.1.3 | Fastening and Loosening of screws
3.1.5 | To insert intramedulary nails
3.1.7 | Traumatology
3.1.8 | Applying of Kirschner wires in bone surgery.
3.2 | The equipment must be made out of a durable robust material light weight and for easy and smooth operating.
3.3 | The unit must have two (2) triggers, it must allow for variable speed in both forward and reverse directions.
3.4 | The unit must have a pistol grip hand piece for single handed operation.
3.5 | All the attachments for the unit must be quick – coupling.
3.6 | The unit must be supplied with reciprocating saw.
3.7 | The unit must be provided with an Aseptic battery which must be easily attached and detached from the base of the hand-piece. If alternative batteries and methods of sterilization are available, PLEASE SPECIFY.
3.8 | The Orthopeadic drill must have the following:
3.8.1 | The unit must comprise of a built-in motor and rechargeable battery.
3.8.2 | Electronic circuitry for controlling its rotational speed adjustable continuously.
3.8.3 | The battery compartment must be attached to the handpiece.
3.8.4 | The battery compartment should have watertight cap that keeps germs out.
3.8.5 | The battery charger must be a modular type and able to allow minimum of 4 batteries to charge at the same time.
3.8.6 | The battery charger must have a screen/indicator to allow users to view the status of the battery
3.8.7 | The unit must be supplied with a minimum of two (2) Li-Ion rechargeable batteries.
3.8.8 | NiCAd batteries will not be accepted.
3.8.9 | The unit must be supplied with battery transfer shield for transferring of unsterile battery to sterile drill.
4 | The drill must be supplied with the following accessories, SPECIFY torque, speed and cannel size where applicable
4.1 | Basic hand piece with attachment coupling that accept variety of attachements
4.2 | Jacob Chuck and a key
4.3 | AO/ASIF quick coupling
4.4 | Quick coupling for Kirschner wires
4.5 | Quick coupling for drill bits
4.6 | Large quick coupling
4.7 | Reduction drive unit
4.8 | Drill chuck keyless
4.9 | Sagittal saw attachement with a set of 10 blades: Sizes as per Endusers choice
4.10 | Oscillating saw attachement with a set of 10 blades: Sizes as per Endusers choice
4.11 | Sterilization case
4.12 | Hudson Modified Trinkle
4.13 | Li-Ion batteries (x2)
4.14 | Battery charger
4.15 | Reciprocating saw
4.16 | Battery transfer shield
5 | Conformity Compliance (Please attach certificate)
5.1 | The unit must comply with an acceptable international standard such as ISO 9714-1 for Orthopaedic surgical equipment, attached certification
5.1.1 | Please provide unique ref number of the ISO 9714-1 ceriticate:
5.2 | OEM must comply and certified on ISO 9001 quality standards, attach proof of compliance
5.2.1 | Please provide unique ref number of the ISO 9001 ceriticate:
5.3 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
5.3.1 | Please provide unique ref number of the ISO 13485 ceriticate:
5.4 | Model quoted for must be EC certified. Attach a copy of certification
5.4.1 | Please provide unique ref number of the EC ceriticate:
5.5 | The unit must comply with ISO 5835 standard, attached certification
5.5.1 | Please provide unique ref number of the ISO 5835 ceriticate:
6 | COMPREHENSIVE MAINTENANCE
6.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) should be quoted for, as per the SCC document
6.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
6.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
6.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
6.3 | The bidder to provide the OEM recommended service interval of the unit:
6.4 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
7 | General
7.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
7.2 | The unit on offer must be the latest technology.
7.2.1 | Initial Manufacturing year of Model:
RT2-07-003 | Drill, Orthopaedic, Battery Operated, Small bones | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
RT2-07-003a | Basic hand piece with attachment coupling that accept variety of attachements
RT2-07-003b | Jacob Chuck and a key
RT2-07-003c | AO/ASIF quick coupling
RT2-07-003d | Quick coupling for Kirschner wires
RT2-07-003e | Quick coupling for drill bits
RT2-07-003f | Large quick coupling
RT2-07-003g | Reduction drive unit
RT2-07-003h | Drill chuck keyless
RT2-07-003i | Instrument tray where all accessories can be placed for cleaning
RT2-07-003j | Sagittal saw attachement with a set of 10 blades: Sizes as per Endusers choice
RT2-07-003k | Oscillating saw attachement with a set of 10 blades: Sizes as per Endusers choice
RT2-07-003l | Sterilization case
RT2-07-003m | Hudson Modified Trinkle
RT2-07-003n | Li-Ion battery x2
RT2-07-003o | Battery charger
RT2-07-003p | Battery trabsfer shield
RT2-07-003q | High Speed bur attachments uses R60-5XX series rotary burs

## Drill, Battery-Op, Large bones
RT2-07-004 | Drill, Orthopaedic, Battery Operated, Large Bones | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
1 | DESCRIPTION
1.1 | A specialized medical instrument designed for drilling, tapping, and screwing large bones, typically used in orthopaedic, podiatric and hand surgery
1a | Product Offered:
1a.1 | Make:
1a.2 | Model:
1a.3 | Initial Manufacturing year of Model:
3 | SPECIFICATIONS
3.1 | The Orthopeadic drill must be able to do the following functions:
3.1.1 | Cleaning of bones(Reaming)
3.1.2 | Tapping
3.1.3 | Fastening and Loosening of screws
3.1.5 | To insert intramedulary nails
3.1.7 | Traumatology
3.1.8 | Applying of Kirschner wires in bone surgery.
3.2 | The equipment must be made out of a durable robust material light weight and for easy and smooth operating.
3.3 | The unit must have two (2) triggers, it must allow for variable speed in both forward and reverse directions.
3.4 | The unit must have a pistol grip hand piece for single handed operation.
3.5 | All the attachments for the unit must be quick – coupling.
3.6 | The unit must be supplied with reciprocating saw.
3.7 | The unit must be provided with an Aseptic battery which must be easily attached and detached from the base of the hand-piece. If alternative batteries and methods of sterilization are available, PLEASE SPECIFY.
3.8 | The Orthopeadic drill must have the following:
3.8.1 | The unit must comprise of a built-in motor and rechargeable battery.
3.8.2 | Electronic circuitry for controlling its rotational speed adjustable continuously.
3.8.3 | The battery compartment must be attached to the handpiece.
3.8.4 | The battery compartment should have watertight cap that keeps germs out.
3.8.5 | The battery charger must be a modular type and able to allow minimum of 4 batteries to charge at the same time.
3.8.6 | The battery charger must have a screen/indicator to allow users to view the status of the battery
3.8.7 | The unit must be supplied with a minimum of two (2) Li-Ion rechargeable batteries.
3.8.8 | NiCAd batteries will not be accepted.
3.8.9 | The unit must be supplied with battery transfer shield for transferring of unsterile battery to sterile drill.
4 | The drill must be supplied with the following accessories, SPECIFY torque, speed and cannel size where applicable
4.1 | Basic hand piece with attachment coupling that accept variety of attachements
4.2 | Jacob Chuck and a key
4.3 | AO/ASIF quick coupling
4.4 | Quick coupling for Kirschner wires
4.5 | Quick coupling for drill bits
4.6 | Large quick coupling
4.7 | Reduction drive unit
4.8 | Drill chuck keyless
4.9 | Sagittal saw attachement with a set of 10 blades: Sizes as per Endusers choice
4.10 | Oscillating saw attachement with a set of 10 blades: Sizes as per Endusers choice
4.11 | Sterilization case
4.12 | Hudson Modified Trinkle
4.13 | Li-Ion batteries (x2)
4.14 | Battery charger
4.15 | Reciprocating saw
4.16 | Battery transfer shield
5 | Conformity Compliance (Please attach certificate)
5.1 | The unit must comply with an acceptable international standard such as ISO 9714-1 for Orthopaedic surgical equipment, attached certification
5.1.1 | Please provide unique ref number of the ISO 9714-1 ceriticate:
5.2 | OEM must comply and certified on ISO 9001 quality standards, attach proof of compliance
5.2.1 | Please provide unique ref number of the ISO 9001 ceriticate:
5.3 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
5.3.1 | Please provide unique ref number of the ISO 13485 ceriticate:
5.4 | Model quoted for must be EC certified. Attach a copy of certification
5.4.1 | Please provide unique ref number of the EC ceriticate:
5.5 | The unit must comply with ISO 5835 standard, attached certification
5.5.1 | Please provide unique ref number of the ISO 5835 ceriticate:
6 | COMPREHENSIVE MAINTENANCE
6.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) should be quoted for, as per the SCC document
6.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
6.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
6.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
6.3 | The bidder to provide the OEM recommended service interval of the unit:
6.4 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
7 | General
7.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
7.2 | The unit on offer must be the latest technology.
7.2.1 | Initial Manufacturing year of Model:
RT2-07-004 | Drill, Orthopaedic, Battery Operated, large bones | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
RT2-07-004a | Basic hand piece with attachment coupling that accept variety of attachements
RT2-07-004b | Jacob Chuck and a key
RT2-07-004c | AO/ASIF quick coupling
RT2-07-004d | Quick coupling for Kirschner wires
RT2-07-004e | Quick coupling for drill bits
RT2-07-004f | Large quick coupling
RT2-07-004g | Reduction drive unit
RT2-07-004h | Drill chuck keyless
RT2-07-004i | Instrument tray where all accessories can be placed for cleaning
RT2-07-004j | Sagittal saw attachement with a set of 10 blades: Sizes as per Endusers choice
RT2-07-004k | Oscillating saw attachement with a set of 10 blades: Sizes as per Endusers choice
RT2-07-004l | Sterilization case
RT2-07-004m | Hudson Modified Trinkle
RT2-07-004n | Li-Ion battery x2
RT2-07-004o | Battery charger
RT2-07-004p | Battery trabsfer shield
RT2-07-004q | High Speed bur attachments uses R60-5XX series rotary burs

## Tourniquet
RT2-07-005 | Tourniquet System | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
1 | DESCRIPTION
1.1 | Is a medical device used to control bleeding and provide a bloodless surgical field during orthopedic vascular, and other surgical procedures
2 | Product Offered:
2.1 | Make:
2.2 | Model:
2.3 | Initial Manufacturing year of Model:
2.4 | Radiation Control License number:
3 | TECHNICAL SPECIFICATION
3.1 | The unit must comprise of a cuff pressure ranging from 0 - 650 mmHg
3.2 | The cuffs must be washable, autoclavable with Velcro, single and double inflation cuff for the dual port system.
3.3 | The device shall be a dual port system
3.4 | Deflate cuff without disconnecting cuff from machine
3.5 | Displayed Parameters:
3.5.1 | Set Pressure
3.5.2 | Measured Pressure
3.5.3 | Time Lapsed
3.5.4 | Count Down
3.6 | Safety Features
3.6.1 | Self-check upon switching the unit On
3.6.2 | Retain cuff pressure (not released) when power fails
3.6.3 | The unit shall have a quick release/disconnecting facility of the cuff.
3.6.4 | The offered unit shall have Automatic Compensation in case of pressure loss.
3.7 | The unit should include visual and audible alarms for:
3.7.1 | Exceeded time
3.7.2 | Internal failure
3.7.3 | Low battery
3.7.4 | Leakage
3.7.5 | High and low pressure
3.8 | Battery
3.8.1 | Intergrated rechargeable battery back-up.
3.8.2 | Back-up battery must last at least 6 hours
3.9 | Mobile Stand
3.9.1 | The offered unit shall be complete with roll stand and an accessory basket.
3.9.2 | The trolley shall have a minimum of 4 casters with central locking mechanism.
4 | STANDARD ACCESSORIES
4.1 | The trolley should have antistatic castors with at least 2 equipped with locking system/brake.
4.2 | Thigh cuffs: Large, medium, small, extra small and Extra Large
4.3 | Arm cuff: Large, medium, small, extra small and Extra Large
4.4 | Reusable, Esmarch Bandages
4.5 | Hoses for all inflation ports
5 | Conformity Compliance (Please attach certificate)
5.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
5.2 | OEM must comply and certified on ISO 9001 quality standards, attach proof of compliance
5.2.1 | Please provide unique ref number of the ISO 9001 ceriticate:
5.3 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
5.3.1 | Please provide unique ref number of the ISO 13485 ceriticate:
5.4 | Model quoted for must be EC certified. Attach a copy of certification
5.4.1 | Please provide unique ref number of the EC ceriticate:
5.5 | Hazardous Substance Act: If this type of equipment / apparatus appears on the schedule of Hazardous substances issued by the Directorate: Health Technology of the Department of Health, a license in terms of the Act on Hazardous Substances (Act 15/1973) must be submitted with the bid / quotation documents. The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.6 | All electrical/electronic medical equipment must be licensed by Radiation Control, where applicable a copy of the license must be submitted The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.6.1 | Radiation Control License number:
5.7 | The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade (rubber to with 2 screws) type and it must be a minimum length of three (3) meters. NB the mains cable of the unit tendered for must be SABS colour coded.
5.8 | The equipment quoted for must be protected against electromagnetic interference IEC 60601-1-2
6 | COMPREHENSIVE MAINTENANCE
6.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) should be quoted for, as per the SCC document
6.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
6.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
6.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
6.3 | The bidder to provide the OEM recommended service interval of the unit:
6.4 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
7 | General
7.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
7.2 | The unit on offer must be the latest technology.
7.2.1 | Initial Manufacturing year of Model:
RT2-07-005 | Tourniquet System -Accessories and Consumables | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
RT2-07-005a | Thigh cuff, Large
RT2-07-005b | Thigh cuff, Medium
RT2-07-005c | Thigh cuff, Small
RT2-07-005d | Thigh cuff, Extra Small
RT2-07-005e | Thigh cuff, Extra Large
RT2-07-005f | Arm cuff, Large
RT2-07-005g | Arm cuff, Medium
RT2-07-005h | Arm cuff, Small
RT2-07-005i | Arm cuff, Extra Small
RT2-07-005j | Arm cuff, Extra Large
RT2-07-005k | Reusable, Esmarch Bandages
RT2-07-005l | Hose
RT2-07-005m | Battery

## Plaster saw
RT2-07-006 | Electrical Plaster of Paris Cutter with Accessories | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
1 | DESCRIPTION
1.1 | is a specialized surgical tool used to remove casts made of plaster or fiberglass. It plays a critical role in orthopedic practices for safely cutting through rigid immobilization materials without harming the underlying skin or tissues.
1a | Product Offered:
1a.1 | Make:
1a.2 | Model:
1a.3 | Initial Manufacturing year of Model:
2 | SPECIFICATION
2.1 | Each unit shall consist of the following components:
2.1.1 | Electrical plaster cutter as specified
2.1.2 | Stainless Steel cast scissors (manual cast cutter)
2.1.3 | Small cast spreader (stainless steel)
2.1.4 | Large cast spreader (stainless steel)
2.2 | The unit shall have a dual speed control
2.3 | The unit shall be versatile and may be used with battery or plugged directly into the wall. Bidder to state if the unit offered is capable to be used as required.
2.4 | The unit must have a dedicated mobile stand which provides stability and easy transportation.
2.5 | The unit must have an impact resistant and durable housing to protect and ensure longer operational lifespan
2.6 | The unit offered shall be of ergonomic design to ensure comfort and control while cutting.
2.7 | The unit offered shall have a noise reducing mechanism.
2.9 | Technical Details
2.9.1 | SIZE: ±10.0 in. (254mm) in length, 3.0 in. (76mm) in height.
2.9.2 | WEIGHT: ±1.3Kg
2.9.3 | SPEED: High ±15,900 CPM, Low 10,700 CPM
2.9.4 | NOISE: ±76dB (low speed), and 82dB (high speed)
2.9.5 | POWER: 220V 50-60Hz 0.75A
3 | ACCESSORIES
3.1 | POP blades (Pk of 5)
3.2 | Fibreglass material blades (Pk of 5)
3.3 | Stainless Steel cast scissors (manual cast cutter)
3.4 | Small cast spreader (stainless steel)
3.5 | Large cast spreader (stainless steel)
4 | Conformity Compliance (Please attach certificate)
4.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
4.2 | OEM must comply and certified on ISO 9001 quality standards, attach proof of compliance
4.2.1 | Please provide unique ref number of the ISO 9001 ceriticate:
4.3 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
4.3.1 | Please provide unique ref number of the ISO 13485 ceriticate:
4.4 | Model quoted for must be EC certified. Attach a copy of certification
4.4.1 | Please provide unique ref number of the EC ceriticate:
4.5 | The mains cable of the unit being bid for must be 15 amp 3 prong hospital grade (rubber to with 2 screws) type and it must be a minimum length of three (3) meters. NB the mains cable of the unit tendered for must be SABS colour coded.
4.6 | The equipment quoted for must be protected against electromagnetic interference IEC 60601-1-2
5 | COMPREHENSIVE MAINTENANCE
5.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) should be quoted for, as per the SCC document
6.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
7 | General
7.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
7.2 | The unit on offer must be the latest technology.
7.2.1 | Initial Manufacturing year of Model:
RT2-07-006 | Electrical Plaster of Paris Cutter with Accessories | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
RT2-07-006a | POP blades (Pk of 5)
RT2-07-006b | Fibreglass material blades (Pk of 5)
RT2-07-006c | Stainless Steel cast scissors (manual cast cutter)
RT2-07-006d | Small cast spreader (stainless steel)
RT2-07-006e | Large cast spreader (stainless steel)

## Plaster saw with extractor
RT2-07-007 | Electrical Plaster of Paris Cutter With Vacuum and Accessories Set | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
1 | DESCRIPTION
1.1 | is a specialized surgical tool used to remove casts made of plaster or fiberglass. It plays a critical role in orthopedic practices for safely cutting through rigid immobilization materials without harming the underlying skin or tissues.
1a | Product Offered:
1a.1 | Make:
1a.2 | Model:
1a.3 | Initial Manufacturing year of Model:
2 | SPECIFICATION
2.1 | Each set shall consist of the following components:
2.1.1 | Electrical plaster cutter as specified
2.1.2 | Mobile Vacuum Extractor to collect cast dust
2.1.3 | Stainless Steel cast scissors (manual cast cutter)
2.1.4 | Small cast spreader (stainless steel)
2.1.5 | Large cast spreader (stainless steel)
2.2 | The system shall have a dual speed control
2.3 | The offer shall include a vacuum to collect and filter any cast dust while cutting.
2.4 | The unit shall be versatile and may be used with vacuum or plugged directly into the wall. Bidder must state if the unit offered is capable to be used as required.
2.5 | A dedicated mobile stand which provides stability and easy transportation.
2.6 | The unit must have an impact resistant and durable housing to protect and ensure longer operational lifespan
2.7 | The unit offered shall be of ergonomic design to ensure comfort and control while cutting.
2.8 | The unit offered shall have a noise reducing mechanism.
2.9 | The unit shall have an integral vacuum housing of at least 2m length. Bidder to state the length of the vacuum housing of the product offered.
2.10 | The cast vacuum canister must have a quick and easy method of emptying and cleaning of the dust collection filter.
2.11 | The hose mount must allow for easy movement of the vacuum hose.
2.12 | The vacuum must turn on automatically when the cast cutter is activated.
2.13 | Vacuum Extractor must have a Cast Cutter holder or flat surface to place it safely
2.14 | Technical Details
2.14.1 | SIZE: ±10.0 in. (254mm) in length, 3.0 in. (76mm) in height.
2.14.2 | WEIGHT: ±1.3Kg
2.14.3 | SPEED: High ±15,900 CPM, Low 10,700 CPM
2.14.4 | NOISE: ±76dB (low speed), and 82dB (high speed)
2.14.5 | POWER: 220V 50-60Hz 0.75A
2.14.6 | VACUUM: 220-240V Vacuum
4 | STANDARD ACCESSORIES
4.1 | POP blades (Pk of 5)
4.2 | Fibreglass material blades (Pk of 5)
4.3 | Re-usable vacuum filters x10
4.4 | Vacuum bags and liners (10 of each)
4.5 | Stainless Steel cast scissors (manual cast cutter)
4.6 | Small cast spreader (stainless steel)
4.7 | Large cast spreader (stainless steel)
5 | Conformity Compliance (Please attach certificate)
5.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
5.2 | OEM must comply and certified on ISO 9001 quality standards, attach proof of compliance
5.2.1 | Please provide unique ref number of the ISO 9001 ceriticate:
5.3 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
5.3.1 | Please provide unique ref number of the ISO 13485 ceriticate:
5.4 | Model quoted for must be EC certified. Attach a copy of certification
5.4.1 | Please provide unique ref number of the EC ceriticate:
5.5 | The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade (rubber to with 2 screws) type and it must be a minimum length of three (3) meters. NB the mains cable of the unit tendered for must be SABS colour coded.
5.6 | The equipment quoted for must be protected against electromagnetic interference IEC 60601-1-2
6 | COMPREHENSIVE MAINTENANCE
6.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) should be quoted for, as per the SCC document
6.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
7 | General
7.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
7.2 | The unit on offer must be the latest technology.
7.2.1 | Initial Manufacturing year of Model:
RT2-07-007 | Electrical Plaster of Paris Cutter With Vacuum and Accessories Set | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
RT2-07-007a | POP blades (Pk of 5)
RT2-07-007b | Fibreglass material blades (Pk of 5)
RT2-07-007c | Re-usable vacuum filters x10
RT2-07-007d | Vacuum bags and liners (10 of each)
RT2-07-007e | Stainless Steel cast scissors (manual cast cutter)
RT2-07-007f | Small cast spreader (stainless steel)
RT2-07-007g | Large cast spreader (stainless steel)

## Dermatome, Cordless
RT2-07-008 | Dermatome, Cordless | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
1 | DESCRIPTION
1.1 | A specialized surgical instrument used primarily in dermatology and plastic surgery to harvest thin layers of skin for skin grafting.
1a | Product Offered:
1a.1 | Make:
1a.2 | Model:
1a.3 | Initial Manufacturing year of Model:
2 | SPECIFICATION
2.1 | The unit supplied shall be a Dermatome with a cutting width of approximately 70mm.
2.2 | It shall be possible to select cutting widths of up to the above mentioned maximum width.
2.3 | It shall be possible to select different cutting depths (Thickness). The following thickness settings shall be possible:
2.3.1 | 0.08mm
2.3.2 | 0.13mm
2.3.3 | 0.18mm
2.3.4 | 0.23mm
2.3.5 | 0.28mm
2.4 | The unit must be compact and lightweight, and easy to handle.
2.5 | It must be easy to adjust different selections i.e. thickness and width.
2.6 | The unit must be a proven design, and be robust for use in a theatre setting.
2.7 | Bidders must prove that the unit is very accurate when different settings are selected.
2.8 | It must be possible to autoclave the full head and handle of the offered unit.
2.9 | The offer must include a sterilisation case to enable sterilising of the unit.
3 | Standard Accessories and Consumables
3.1 | 25mm (Pk of 10)
3.2 | 50mm (Pk of 10)
3.3 | 70mm (Pk of 10)
3.4 | Sterilizing Case
3.5 | Battery Charger (2-slots)
4 | Conformity Compliance (Please attach certificate)
4.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
4.2 | OEM must comply and certified on ISO 9001 quality standards, attach proof of compliance
4.2.1 | Please provide unique ref number of the ISO 9001 ceriticate:
4.3 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
4.3.1 | Please provide unique ref number of the ISO 13485 ceriticate:
4.4 | Model quoted for must be EC certified. Attach a copy of certification
5.4.1 | Please provide unique ref number of the EC ceriticate:
6 | COMPREHENSIVE MAINTENANCE
6.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) should be quoted for, as per the SCC document
6.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
7 | General
7.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
7.2 | The unit on offer must be the latest technology.
7.2.1 | Initial Manufacturing year of Model:
RT2-07-008 | Dermatome, Cordless | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
RT2-07-008a | 25mm (Pk of 10)
RT2-07-008b | 50mm (Pk of 10)
RT2-07-008c | 70mm (Pk of 10)
RT2-07-008d | Sterilizing Case
RT2-07-008e | Battery Charger (2-slots)

## Dermatome, Pneumatic
RT2-07-009 | Dermatome, Pneumatic | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
1 | DESCRIPTION
1.1 | a specialized surgical instrument used primarily in dermatology and plastic surgery to harvest thin layers of skin for skin grafting.
1a | Product Offered:
1a.1 | Make:
1a.2 | Model:
1a.3 | Initial Manufacturing year of Model:
2 | SPECIFICATION
2.1 | The unit must be powered by compressed air from a pneumatic source.
2.2 | The unit supplied shall be a Dermatome with a cutting width of approximately 70mm.
2.3 | It must be possible to select cutting widths of up to the abovementioned maximum width.
2.4 | It shall be possible to select different cutting depths (Thickness). The following thickness settings shall be possible:
2.4.1 | 0.08mm
2.4.2 | 0.13mm
2.4.3 | 0.18mm
2.4.4 | 0.23mm
2.4.5 | 0.28mm
2.5 | The unit must be compact and lightweight, and easy to handle.
2.6 | It must be easy to adjust different selections i.e. thickness and width.
2.7 | The unit must be a proven design, and be robust for use in a theatre setting.
2.8 | Bidders must prove that the unit is very accurate when different settings are selected.
2.9 | It must be possible to sterilise the full head and handle of the offered unit.
2.10 | The offer must include a sterilisation case to enable sterilising of the unit.
3 | Standard Accessories and Consumables
3.1 | 25mm (Pk of 10)
3.2 | 50mm (Pk of 10)
3.3 | 70mm (Pk of 10)
3.4 | Sterilizing Case
3.5 | 3m Hose
4 | Conformity Compliance (Please attach certificate)
4.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
4.2 | OEM must comply and certified on ISO 9001 quality standards, attach proof of compliance
4.2.1 | Please provide unique ref number of the ISO 9001 ceriticate:
4.3 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
4.3.1 | Please provide unique ref number of the ISO 13485 ceriticate:
4.4 | Model quoted for must be EC certified. Attach a copy of certification
5.4.1 | Please provide unique ref number of the EC ceriticate:
6 | COMPREHENSIVE MAINTENANCE
6.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) should be quoted for, as per the SCC document
6.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
7 | General
7.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
7.2 | The unit on offer must be the latest technology.
7.2.1 | Initial Manufacturing year of Model:
RT2-07-009 | Dermatome, Pneumatic | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
RT2-07-009a | 25mm (Pk of 10)
RT2-07-009b | 50mm (Pk of 10)
RT2-07-009c | 70mm (Pk of 10)
RT2-07-009d | Sterilizing Case
RT2-07-009e | 3m Hose

## Skin Mesher
RT2-07-010 | Skin Mesher | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
1 | SCOPE
1.1 | A medical device used to create uniform, meshed skin grafts for optimal wound closure and healing
1a | Product Offered:
1a.1 | Make:
1a.2 | Model:
1a.3 | Initial Manufacturing year of Model:
2 | SPECIFICATION FOR MESHER
2.1 | The unit must be a skin mesher used for the expansion of harvested skin.
2.2 | It must be a mechanical unit designed for simple operation, utilising disposable carriers to carry the skin through the mesher operation.
2.3 | The unit must be manufactured from high quality, durable stainless steel.
2.4 | It must be possible to change to at least four cutting sizes. Four (4) different cutters and four (4) sets of meshers must be included in the offer
2.5 | The unit must have the following expansion ratios:
2.5.1 | 1.5:1
2.5.2 | 2:1
2.5.3 | 3:1
2.5.4 | 4:1
2.5 | The skin shall be guided through the mesher by a ratchet system, ensuring smooth transport of the skin graft carrier.
2.6 | The mesher instrument shall consist of a ratchet handle, feed roller, cutting roller and a guidance platform. The harvested skin shall be guided along the platform by means of the carriers and the ratchet handle.
3 | Standard Accessories
3.1 | Small carriers (Pack of 10)
3.2 | Large carriers (Pack of 10)
4 | Conformity Compliance (Please attach certificate)
4.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
4.2 | OEM must comply and certified on ISO 9001 quality standards, attach proof of compliance
4.2.1 | Please provide unique ref number of the ISO 9001 ceriticate:
4.3 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
4.3.1 | Please provide unique ref number of the ISO 13485 ceriticate:
4.4 | Model quoted for must be EC certified. Attach a copy of certification
4.4.1 | Please provide unique ref number of the EC ceriticate:
5 | COMPREHENSIVE MAINTENANCE
5.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) should be quoted for, as per the SCC document
5.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
6 | General
6.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
6.2 | The unit on offer must be the latest technology.
6.2.1 | Initial Manufacturing year of Model:
RT2-07-10 | Skin Mesher | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
RT2-07-010a | Small carriers (Pack of 10)
RT2-07-010b | Large carriers (Pack of 10)

## Arthroscopic shaver
RT2-07-011 | Arthroscopic Shaver System | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
1 | DESCRIPTION
1.1 | A medical device used in minimally invasive orthopaedic procedures to remove damaged or inflamed tissue, bone and cartilage joints
1a | Product Offered:
1a.1 | Make:
1a.2 | Model:
2 | Technical Specification
2.1 | The arthroscopic shaver system offered must include a shaver console and Motor Drive Unit (MDU).
2.2 | The arthroscopic shaver system offered must employ reliable electronic circuitry for accurate control and a high degree of patient safety.
2.3 | The arthroscopic shaver system must be controlled by a control unit panel and footswitch
2.4 | The footswitch must be included in the bid price
2.5 | The arthroscopic shaver system offered must be able to automaticaly identify connected handpiece and set of previously used parameters
2.6 | The arthroscopic shaver system must have two handpiece outputs with automatic and manual lock
2.7 | The hand piece must be ergonomic and autoclavable.
2.8 | There must a protection against the overloading of a handpiece and non-volatile memory of previous settings
2.9 | A box of disposable shaver blades must be included in the bid price.
2.10 | The arthroscopic shaver system must be able to clean the joint space by removing soft tissue such as the synovium, and even some of the denser tissues such as degenerated cartilage fragments and dissociated cartilage, as well as the trimming of the meniscus and bone.
2.11 | The arthroscopic shaver system offered must allow for forward, backward and oscillating modes via programmable control buttons on the handpiece.
2.12 | The arthroscopic shaver system offered must have the ability to connect to suction which can be switched between high, low and no suction.
2.13 | The arthroscopic shaver system offered must have variable speed modes
2.14 | The control panel must be on the front panel of the unit offered. It should be easy to read and use
2.16 | FLUID PUMP
2.16.1 | State the make and model of the pump
2.16.2 | The ﬂuid pump must be designed to deliver ﬂuid under a given pressure to an operating area controlled by a control panel .
2.16.3 | Under constant pressure the fluid pump must allow the rinse solution to be introduced to ensure a clear vision and surgical space for the operation by continuously replenishing the washing fluid
2.16.4 | The unit must be user friendly and dedicated to arthroscopic surgeries.
2.16.5 | A "FLUSH" function feature for an efficient rinsing of an operating area must be available.
2.16.6 | The unit offered must employ reliable electronic circuitry for accurate control and it must also offer a high degree of patient safety.
2.16.7 | The unit offered must be provided with clear visual display for monitoring flow and pressure.
2.16.8 | Visible and audible warnings must be provided on the unit offered in the event of patient overpressure and incorrect tubing installation
2.16.9 | The unit offered must be provided with outflow control.
2.16.10 | All necessary accessories and consumables, e.g. patient tubing, must be supplied as part and parcel of this unit at no extra cost to the final bid price in order to enable immediate use of the unit without requiring to order further accessories. Bidder must list all these accessories.
2.16.11 | The unit offered must require no prior calibration before use.
2.16.12 | Remote controller must be included as second option to control the pump.
2.16.13 | The unit must have a mechanism to automaticaly recognise the inserted instrument.
2.16.14 | The unit offered must be able to self-calibrate continuously to ensure settings are reproducible and reliable.
2.16.15 | The unit must able to Improve the flow performance through automatic detection of any scope cannula.
2.17 | The shaver be supplied with following Standard Accessories:
2.17.1 | Control unit
2.17.2 | Handpiece with cable
2.17.3 | Foot pedal
2.17.4 | Irrigation stand
2.17.5 | Set of irrigation lines
2.17.6 | Set of irrigation clips
2.17.7 | Shaver handpiece cleaning brush
2.17.8 | Shaver blade cleaning brush
2.17.9 | Different sizes of disposable blades
3 | Conformity Compliance (Please attach certificate)
3.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
3.2 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
3.2.1 | Please provide unique ref number of the ISO 13485 ceriticate:
3.3 | Model quoted for must be EC certified. Attach a copy of certification
3.3.1 | Please provide unique ref number of the EC ceriticate:
3.4 | Hazardous Substance Act: If this type of equipment / apparatus appears on the schedule of Hazardous substances issued by the Directorate: Health Technology of the Department of Health, a license in terms of the Act on Hazardous Substances (Act 15/1973) must be submitted with the bid / quotation documents. The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
3.5 | All electrical/electronic medical equipment must be licensed by Radiation Control, where applicable a copy of the license must be submitted The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
3.5.1 | Radiation Control License number:
3.7 | The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade (rubber to with 2 screws) type and it must be a minimum length of three (3) meters. NB the mains cable of the unit tendered for must be SABS colour coded.
3.8 | The equipment quoted for must be protected against electromagnetic interference IEC 60601-1-2
4 | COMPREHENSIVE MAINTENANCE
4.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) must be included in the bid price
4.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
4.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
4.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
4.3 | The bidder to provide the OEM recommended service interval of the unit:
4.4 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
5 | General
5.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
5.2 | The unit on offer must be the latest technology.
RT2-07-11 | Arthroscopic Shaver System | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
RT2-07-11a | Control unit
RT2-07-11b | Handpiece with cable
RT2-07-11c | Foot pedal
RT2-07-11d | Irrigation stand
RT2-07-11e | Set of irrigation lines
RT2-07-11f | Set of irrigation clips
RT2-07-11g | Shaver handpiece cleaning brush
RT2-07-11h | Shaver blade cleaning brush
RT2-07-11i | Disposable blades: 3.5mm Dual Edge shaver blade
RT2-07-11j | Disposable blades: 4.0mm Dual Edge shaver blade
RT2-07-11k | Disposable blades: 4.5mm Dual Edge shaver blade
RT2-07-11l | Disposable blades: 3.5mm Angled shaver blade
RT2-07-11m | Disposable blades: 4.0mm Angled shaver blade
RT2-07-11n | Disposable blades: 4.5mm Angled shaver blade

## Category-8 Ophthalmo
OPHTHALMOLOGY EQUIPMENT
# | Item number | Item Description
1 | RT2-08-001 | A&B Scanner Combo Unit
2 | RT2-08-002 | Auto Ref-Keratometer
3 | RT2-08-003 | Laser Unit, Argon
4 | RT2-08-004 | Laser Unit, YAG
5 | RT2-08-005 | Luminometer
6 | RT2-08-006 | Optical Coherence Tomography Unit
7 | RT2-08-007 | Phoropter
8 | RT2-08-008 | Slit-Lamp, Basic
9 | RT2-08-009 | Slit-Lamp, Advanced
10 | RT2-08-010 | Tonometer, Hand-held
11 | RT2-08-011 | Visual Field Analyser
12 | RT2-08-012 | Vitrectomy Unit (Anterior and Posterior)
13 | RT2-08-013 | Vitroctoy Unit (Anterior)

## A&B Scanner Combo
RT2-08-001 | A& B Scanner with Pachymetry | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
1 | Description
1.1 | A versatile medical imaging device combining A-scan and B-scan ultrasound technologies.
2 | Product Offered:
2.1 | Make:
2.2 | Model:
2.3 | Initial Manufacturing year of Model:
2.4 | Radiation Control License number:
3 | Technical Specification
3.1 | The unit must be able to perform the following functions:
3.1.1 | Biometry
3.1.2 | IOL calculation
3.1.3 | Standardised A-diagnostic with S-mode
3.1.4 | Pachymetry with central and map function
3.1.5 | Internal data base
3.2 | The unit must include the following :
3.2.1 | New generation of annular array probes
3.2.2 | Probe holders(preferable on the back of the unit)
3.2.3 | High resolution touch-screen operation: minimum of 10 inches, colour touch screen
3.2.4 | Data communication must be possible via USB or LAN
3.2.5 | Modular configurable with Pachymetry
3.2.6 | Foot switch
3.3 | Include the following probes
3.3.1 | Internal database UBM 40MHz B-probe
3.3.2 | Biometry A-scan 10 MHz A-diagnostic probe
3.3.3 | 10 MHz 2-ring array B-probe
3.3.4 | Pachymetry 20 MHz solid state probe tip 1.5 mm with an angle of at least 45°
3.3.5 | Biometry 10 MHz solid state probe with built in red fixation LED
3.3.6 | A-scan diagnosis 10 MHz solid state probe
3.4 | 10 MHz B-probe , 2-Ring sector scanning
3.4.1 | State the range
3.4.2 | Frame rate basic mode
3.4.3 | Maximum number of pages in a movie
3.4.4 | Image display range Standard
3.4.5 | Wide
3.5 | Axial & lateral resolution / accuracy
3.6 | Colour scale
3.7 | The B-scan must be able to do zooming,report/export image,sampling rate,scanning angle and adjustable gamma
3.8 | Biometry 10 Mhz solid state probe with built in red fixation LED
3.8.1 | State the range
3.8.2 | Measurement range
3.8.3 | Measurement resolution / accuracy
3.8.4 | IOL power calculation
3.9 | A-scan diagnosis 10 MHz solid state probe
3.9.1 | Measurement range
3.9.2 | Measurement resolution / accuracy
3.9.3 | Capture:automatic and manual
3.9.4 | Custom eye type settings
3.10 | Pachymetry 20 MHz solid state probe tip 1.5 mm with an angle of at least 45°
3.10.1 | State the range
3.10.2 | Measurement range
3.10.3 | Measurement resolution / accuracy
3.10.4 | Percent bias
3.10.5 | Plus minus bias
3.10.6 | Velocity range / standard
3.10.7 | IOP adjustment as well as sampling rate
3.11 | The following connectors/ports must be available
3.11.1 | B-scan for 10 or 40 MHz,
3.11.2 | A-scan/pachymeter
3.11.3 | Foot switch
3.11.4 | LAN; 2 x standard USB and 2 x OA-2000/PicBrigde
3.12 | It must be possible on the unit to save the data on the internal database or on the PC. Intuitive(Spontaneous) software guides all measurement results, which can be printed out through built-in printer or external printer.
3.13 | If the unit does not have a built-in printer , the bidder must provide an external printer which must be included in the bid price.
4 | Standard Accessories
4.1 | Internal database UBM 40MHz B-probe
4.2 | Biometry A-scan 10 MHz A-diagnostic probe
4.3 | 10 MHz 2-ring array B-probe
4.4 | Pachymetry 20 MHz solid state probe tip 1.5 mm with an angle of at least 45°
4.5 | Biometry 10 MHz solid state probe with built in red fixation LED
4.6 | A-scan diagnosis 10 MHz solid state probe
4.7 | Printing paper (Pk of 10)
5 | Conformity Compliance (Please attach certificate)
5.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
5.2 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
5.2.1 | Please provide unique ref number of the ISO 13485 ceriticate:
5.3 | Model on offer must comply with EU Medical Device Regulation (EU) 2017/745 or equvalent from SAHPRA recognized jurisdictions . Attach a copy of certification
5.3.1 | Please provide unique ref number of the ceriticate:
5.4 | Hazardous Substance Act: If this type of equipment / apparatus appears on the schedule of Hazardous substances issued by the Directorate: Health Technology of the Department of Health, a license in terms of the Act on Hazardous Substances (Act 15/1973) must be submitted with the bid / quotation documents. The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.5 | All electrical/electronic medical equipment must be licensed by Radiation Control, where applicable a copy of the license must be submitted The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.5.1 | Radiation Control License number:
5.6 | The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade (rubber to with 2 screws) type and it must be a minimum length of three (3) meters. NB the mains cable of the unit tendered for must be SABS colour coded.
5.7 | The equipment quoted for must be protected against electromagnetic interference IEC 60601-1-2
5.8 | The unit must comply with electrical safety and performance standard such as IEC 60601-2-21, attached certification
5.8.1 | Please provide unique ref number of the IEC ceriticate:
6 | COMPREHENSIVE MAINTENANCE
6.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) should be included in the unit price as per the SCC document
6.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
6.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
6.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
6.3 | The bidder to provide the OEM recommended service interval of the unit:
6.4 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
7 | General
7.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
7.2 | The unit on offer must be the latest technology.
7.3 | Initial Manufacturing year of Model:
RT2-08-001 | A& B Scanner with Pachymetry- Accessories and Consumables | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
RT2-08-001a | Internal database UBM 40MHz B-probe
RT2-08-001b | Biometry A-scan 10 MHz A-diagnostic probe
RT2-08-001c | 10 MHz 2-ring array B-probe
RT2-08-001d | Pachymetry 20 MHz solid state probe tip 1.5 mm with an angle of at least 45°
RT2-08-001d | Biometry 10 MHz solid state probe with built in red fixation LED
RT2-08-001e | A-scan diagnosis 10 MHz solid state probe
RT2-08-001f | Printing paper (Pk of 10)

## Auto Ref-Keratometer
RT2-08-002 | Ophthalmic Auto Ref-Keratometer | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
1 | DESCRIPTION
1.1 | An instrument for measuring the curvature of the anterior corneal surface
2 | Product Offered:
2.1 | Make:
2.2 | Model:
2.3 | Initial Manufacturing year of Model:
2.4 | Radiation Control License number:
3 | SPECIFICATIONS
3.1 | The system offered must have the following test modes:
3.1.1 | Refraction
3.1.2 | Keratometry.
3.1.3 | Refraction and Keratometry
3.2 | The measuring range (Diopter-D) for HYPEROPIA must be at least 0D to + 25D in minimum increments of 0,01D steps.
3.3 | The measuring range (Diopter-D) for MYOPIA must be at least 0D to -30D in minimum increments of 0,01D steps.
3.4 | The measuring range (Diopter-D) for ASTIGMATISM must be at least 0D to ±12D in minimum increments of 0,01D steps.
3.5 | The AXIS measuring range must be at least 1° to 180° in steps of one deg (1°) .
3.6 | Testing of the naked eye, the pseudophakic eye and over refraction with contact lenses must be available and selectable.
3.7 | The vertex distance for the refraction must be displayed.
3.8 | The pupillary distance (PD) must be calculated and shown on the printout.
3.9 | PD measurement adjustment with 1mm step must be available. Bidder to specify the range
3.10 | KERATOMETRY must meet the following specifications:
3.10.1 | The CORNEAL CURVATURE RADIUS measuring range must be adjustable with minimum of 0,01mm steps. Bidder to specify the range
3.10.2 | The CORNEAL REFRACTION measuring range must be adjustable with minimum 0,01D steps. Bidder to specify the range
3.10.3 | The CORNEAL ASTIGMATISM ANGLE measuring range must be at 1° to 180° in steps of 1°.
3.10.4 | The readout display must be selectable between MILLIMETER and DIOPTER.
3.10.5 | An TFT colour display of at least 16cm must be incorporated in the system to provide ease of operation and accurate alignment.
3.10.6 | Reading must be displayed on the video monitor for all three modes as follows: REFRACTION; KERATOMETRY; REFRACTION and KERATOMETRY.
3.10.7 | Approximately 10 (ten) or more readings of both REFRACTION and KERATOMETRY must be retained in memory for averaging.
3.11 | Built-in Printer
3.11.1 | PRINTOUT must be obtained at the end of the procedure and the printout must provide the following information:
3.11.1.1 | Patient information
3.11.1.2 | Refractive error measurements
3.11.1.3 | Keratometry Measurements
3.11.1.4 | Additional measurements: Pachymetry (Corneal thickness), Refractive index, Corneal curvature and IOL(Intraocular Lens) Power calculation
3.11.1.5 | Device information: Device model & serial number, Software version and Measurement Date/Time
3.11.1.6 | Other: Print date/time, Operator/Technician ID (Optional) and Hospital name (Optional)
3.12 | A printout for refraction, when selected, must provide the following:
3.12.1 | Vertex distance.
3.12.2 | PD.
3.12.3 | The readings taken for each eye.
3.12.4 | The average of the readings for each eye.
3.13 | A printout for keratometry, when selected, must provide the following
3.13.1 | Corneal curvature radius
3.13.2 | Corneal refraction.
3.13.3 | The average of the keratometry values.
3.13.4 | Corneal astigmatism value in D and axis.
3.14 | A following modes and printout, when selected, must provide the following:
3.14.1 | Pupil size from 1mm – 10mm in 0,1mm increments.
3.14.2 | Corneal size from 1mm – 14mm in 0,1mm increments.
3.14.3 | The system offered must have an automatic shut down/off facility to conserve energy and the life of the system components when not in use.
3.14.4 | The system must be controlled by means of a joystick or any other reliable latest technology.
3.14.5 | The system must include a motorised chin support device capable of vertical adjustment for the patients head.
3.14.6 | The system must produce quality results under normal room lighting conditions.
3.14.7 | The system must include a mobile sturdy table with lockable castors and with an easily adjustable electrical height mechanism, at no extra cost to the final bid price.
4 | STANDARD ACCESSORIES (Any other cables or connectors required for the functionality of the unit must be provided)
4.1 | Printing paper (pack of 10)
4.2 | Dust cover
5 | Conformity Compliance (Please attach certificate)
5.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
5.2 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
5.2.1 | Please provide unique ref number of the ISO 13485 ceriticate:
5.3 | Model on offer must comply with EU Medical Device Regulation (EU) 2017/745 or equvalent from SAHPRA recognized jurisdictions . Attach a copy of certification
5.3.1 | Please provide unique ref number of the ceriticate:
5.4 | Hazardous Substance Act: If this type of equipment / apparatus appears on the schedule of Hazardous substances issued by the Directorate: Health Technology of the Department of Health, a license in terms of the Act on Hazardous Substances (Act 15/1973) must be submitted with the bid / quotation documents. The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.5 | All electrical/electronic medical equipment must be licensed by Radiation Control, where applicable a copy of the license must be submitted The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.5.1 | Radiation Control License number:
5.6 | The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade (rubber to with 2 screws) type and it must be a minimum length of three (3) meters. NB the mains cable of the unit tendered for must be SABS colour coded.
5.7 | The equipment quoted for must be protected against electromagnetic interference IEC 60601-1-2
5.8 | The unit must comply with electrical safety and performance standard such as IEC 60601-2-21, attached certification
5.8.1 | Please provide unique ref number of the IEC ceriticate:
6 | COMPREHENSIVE MAINTENANCE
6.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) should be quoted for, as per the SCC document
6.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
6.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
6.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
6.3 | The bidder to provide the OEM recommended service interval of the unit:
6.4 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
7 | General
7.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
7.2 | The unit on offer must be the latest technology.
7.2.1 | Initial Manufacturing year of Model:
RT2-08-002 | Ophthalmic Auto Ref-Keratometer -Consumables | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
RT2-08-002a | Printing paper (pack of 10)
RT2-08-002b | Dust cover

## Laser Unit, Argon
RT2-08-003 | Argon Laser-Multispot with Slit Lamp and Indirect Opthalmoscope | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
1 | Description
1.1 | A versatile, advanced ophthalmic diagnostic and treatment system. This device combines three essential functions for comprehensive eye examinations and treatments.
2 | Product Offered:
2.1 | Make:
2.2 | Model:
2.3 | Initial Manufacturing year of Model:
2.4 | Radiation Control License number:
3 | Technical Specification
3.1 | The specification establishes the requirement of supply, delivery, end user training, demonstration, commissioning and installation of an Argon Laser-Multisport with Slit lamp and Indirect ophthalmoscope which must be supplied as a system as per the under mentioned Clauses.
3.2 | The unit is intended to be used for the following:
3.2.1 | In the Posterior segment to perform retinal photocoagulation, Panretinal photocoagulation, focal photocoagulation, and grid photocoagulation for vascular and structural abnormalities of the retina and choroid.
3.2.2 | In the treatment of ocular pathology in the anterior segment including the Iridotomy and Trabeculoplasty.
3.3 | The design must be such that the multi-spot approach uses shorter pulse duration of individual pulses to allow for manageable time for multi-spot delivery.
3.4 | The unit must be designed such that the Pan retinal photocoagulation (PRP) give rise to lower levels of discomfort and distress to the patient compared to the single spot PRP treatment.
3.5 | For safety of the patient, the unit must be designed such that the laser is safe, effective and other several advantages associated with shorter exposure including reduced pain, reduced inner retinal damage and reduced scarring.
3.6 | The device must perform single session treatment which help to induce the lower levels of anxiety, headache and photophobia.
3.7 | Unit that offers conventional laser will not be accepted as these require multiple sessions for treatment.
3.8 | The device must provide complete control over all customary settings such as pulse duration, laser power, pattern selection and pattern size.
3.9 | All parameters must be clearly marked with exceptionally easy access to the standby/ready-modes.
3.10 | It must be possible to control aspects such as pattern size, pattern orientation and laser power without removing your view from the patient's fundus.
3.11 | It must be possible to display different types of patterns such as:
3.11.1 | Single spot
3.11.2 | Array
3.11.3 | Triple ring
3.11.4 | Arc
3.11.5 | Line circle
3.11.6 | Circle
3.11.7 | Macular grid
3.11.8 | Grid
3.11.9 | Hexagon
3.12 | The controls and displays must be touchscreen (that allows easy selection of arrays, lines and rectangles) user interface and integrated slit lamp.
3.13 | The system must be equipped with a three position key switch that enables laser and controls, as well as an emergency OFF switch.
3.14 | A system must have a 3D remote control that allows physician to control over lasting positioning, power settings, adjust and rotate laser patterns.
3.15 | A footswitch activator must be provided with a shrouded housing
3.16 | For patients that are unable to sit on a slit lamp, a laser Indirect Ophthalmoscope must be provided. This must be included in the total bid price.
3.17 | For the performance of the argon laser the following features must be possible:
3.17.1 | Power: 0-2000mW
3.17.2 | Wavelength: 577nm
3.17.3 | Pulse duration (single spot): 5-1000ms
3.17.4 | Aim beam: 620nm diode 650nm adjustable
3.17.5 | Delivery spot size: 50, 100, 200, 400um, etc.
3.17.6 | Exposure time: 2milli-seconds-1000milli-seconds
3.17.7 | Cooling system: Thermoelectric or Air cooled
3.17.8 | Pulse counter: 0-99999 with automatic reset available
3.17.9 | Repetition rate: Available with single spot pattern
3.17.10 | 5 inch or above touchscreen monitor display
4 | Standard Accessories and Cosumables
4.1 | Dust cover
4.2 | Ritch-trabeculoplasty
4.3 | Latina SLT Gonio
4.4 | Abraham capsul YAG
4.5 | Abraham iridotomy
4.6 | Peyman Wide Field
4.7 | Goldmann 3 mirror
4.8 | SLT license with Gonio Lens
4.9 | Mainster PRP 165
4.10 | Mainster Standard 90
4.11 | Mainster Wide field 118
5 | Conformity Compliance (Please attach certificate)
5.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
5.2 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
5.2.1 | Please provide unique ref number of the ISO 13485 ceriticate:
5.3 | Model on offer must comply with EU Medical Device Regulation (EU) 2017/745 or equvalent from SAHPRA recognized jurisdictions . Attach a copy of certification
5.3.1 | Please provide unique ref number of the ceriticate:
5.4 | Hazardous Substance Act: If this type of equipment / apparatus appears on the schedule of Hazardous substances issued by the Directorate: Health Technology of the Department of Health, a license in terms of the Act on Hazardous Substances (Act 15/1973) must be submitted with the bid / quotation documents. The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.5 | All electrical/electronic medical equipment must be licensed by Radiation Control, where applicable a copy of the license must be submitted The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.5.1 | Radiation Control License number:
5.6 | The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade (rubber to with 2 screws) type and it must be a minimum length of three (3) meters. NB the mains cable of the unit tendered for must be SABS colour coded.
5.7 | The equipment quoted for must be protected against electromagnetic interference IEC 60601-1-2
5.8 | The unit must comply with electrical safety and performance standard such as IEC 60601-2-21, attached certification
5.8.1 | Please provide unique ref number of the IEC ceriticate:
6 | COMPREHENSIVE MAINTENANCE
6.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) should be quoted for, as per the SCC document
6.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
6.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
6.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
6.3 | The bidder to provide the OEM recommended service interval of the unit:
6.4 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
7 | General
7.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years, If yes, please provide further details of the recall and how was it addressed
7.2 | The unit on offer must be the latest technology.
7.2.1 | Initial Manufacturing year of Model:
RT2-08-003 | Argon Laser-Multispot with Slit Lamp and Indirect Opthalmoscope | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
RT2-08-003a | Dust cover
RT2-08-003b | Ritch-trabeculoplasty
RT2-08-003c | Latina SLT Gonio
RT2-08-003d | Abraham capsul YAG
RT2-08-003e | Abraham iridotomy
RT2-08-003f | Peyman Wide Field
RT2-08-003g | Goldmann 3 mirror
RT2-08-003h | SLT license with Gonio Lens
RT2-08-003i | Mainster PRP 165
RT2-08-003j | Mainster Standard 90
RT2-08-003k | Mainster Wide field 118

## Laser Unit, YAG
RT2-08-004 | Yttrium Aluminum Garnet and Selective Laser Trabeculoplasty (Yag and Slt) | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
1 | DESCRIPTION
1.1 | Ophthalmic laser solution for the ophthalmic practice, providing glaucoma, cataract and retinal therapy in a single platform
2 | Product Offered:
2.1 | Make:
2.2 | Model:
2.3 | Initial Manufacturing year of Model:
2.4 | Radiation Control License number:
3 | SPECIFICATIONS
3.1 | The unit must be operated through an intergrated laser fire Q-switch
3.2 | The unit must allow the user to perform the full range of anterior and posterior YAG laser procedures, as well as SLT, with greater precision and clinical efficacy.
3.3 | In SLT mode, the laser energy must be delivered in short fixed pulses over a large homogeneous spot size without causing permanent destruction to the trabecular meshwork.
3.4 | The YAG mode must deliver high performance photo disruption for capsulotomy and peripheral iridotomy.
3.5 | Optics and Laser delivery technology must feature uniform energy distribution across the entire laser spot area for more precise and safer photocoagulation treatment.
3.6 | The unit offered must automatically sense the delivery devices connected to it and internally adjust Laser parameters to compensate for the chosen delivery device to maintain selected power density on the target tissue.
3.7 | The precise guidance of the aiming and therapy beams on the laser system offered must guarantee optimum illumination of site of treatment at all times.
3.8 | The system must have a repeat function to ensure that at least three laser shots can be delivered in rapid succession.
3.9 | The system must have a dial to offset the aiming beam in steps of maximum 125µm to 500 µm anterior and posterior to the active laser.
3.10 | The ophthalmic laser aiming spot must have adjustable brightness.
3.11 | The ophthalmic laser must have a built in Physician Eye Safety Filter, that protects the user from both the YAG and Photocoagulation laser wavelengths.
3.12 | The ophthalmic laser system offered must ensure that the doctors safety filter is kept in the beam path during treatment.
3.13 | The unit must have air convention cooling system or alternative effective cooling system
3.14 | The following features must be available on the system offered:
3.14.1 | Posterior segment viewing
3.14.2 | Anterior segment viewing
3.14.3 | Dual-color aiming beam
3.14.4 | Precise SLT and YAG treatment
3.14.5 | Easy switch between YAG and SLT
3.14.6 | IOL-friendly photodisruption
3.14.7 | Accurately position optical breakdown
3.14.8 | Blue free illumination for improved visibility
3.14.9 | Integrated laser and slit lamp for greater efficiency, reliability and ease of use
3.14.10 | Optimized to produce two wavelengths from one cavity
3.14.11 | Large diameter oculars for a wider field of view of the target area
3.14.12 | Three posterior defocus positions with an adjustable focal point with a range of at least 350um from anterior to posterior to the aiming beam.
3.14.13 | Mechanism for accurate and safe laser beam guidance for fine-tuned laser application on target tissue.
3.14.14 | The unit must have an option of incorporating an additional module/device able to deliver argon laser retinal photocoagulation in the future.
3.15 | The following must be supplied as part of the ophthalmic laser and must be included in the final bid price:
3.15.1 | A table that is equipped with built-in compact storage space for lenses or other accessories. The table height must be designed such that it can easily be adjusted from 450 - 700 mm using electronic up/down switch.
3.15.2 | The table must allow access for wheelchairs.
3.15.3 | A joystick or equivalent that includes a convenient integrated laser fire-switch.
3.15.4 | The full slit lamp function that includes at least 6-step magnification changer.
3.15.5 | Appropriate lenses to allow safe treatment of the posterior capsule and iris.
4 | Standard Accessories and Consumables
4.1 | Dust Cover
4.2 | Two pairs of safety goggles.
4.3 | Mainster PRP 165
4.4 | Mainster Standard 90
4.5 | Mainster Wide field 118
4.6 | Ritch-trabeculoplasty
4.7 | Latina SLT Gonio
4.8 | Abraham capsul YAG
4.9 | Abraham iridotomy
4.10 | Peyman Wide Field
4.11 | Goldmann 3 mirror
5 | Conformity Compliance (Please attach certificate)
5.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
5.2 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
5.2.1 | Please provide unique ref number of the ISO 13485 ceriticate:
5.3 | Model on offer must comply with EU Medical Device Regulation (EU) 2017/745 or equvalent from SAHPRA recognized jurisdictions . Attach a copy of certification
5.3.1 | Please provide unique ref number of the ceriticate:
5.4 | Hazardous Substance Act: If this type of equipment / apparatus appears on the schedule of Hazardous substances issued by the Directorate: Health Technology of the Department of Health, a license in terms of the Act on Hazardous Substances (Act 15/1973) must be submitted with the bid / quotation documents. The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.5 | All electrical/electronic medical equipment must be licensed by Radiation Control, where applicable a copy of the license must be submitted The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.5.1 | Radiation Control License number:
5.6 | The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade (rubber to with 2 screws) type and it must be a minimum length of three (3) meters. NB the mains cable of the unit tendered for must be SABS colour coded.
5.7 | The equipment quoted for must be protected against electromagnetic interference IEC 60601-1-2
5.8 | The unit must comply with electrical safety and performance standard such as IEC 60601-2-21, attached certification
5.8.1 | Please provide unique ref number of the IEC ceriticate:
6 | COMPREHENSIVE MAINTENANCE
6.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) should be quoted for, as per the SCC document
6.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
6.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
6.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
6.3 | The bidder to provide the OEM recommended service interval of the unit:
6.4 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
7 | General
7.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
7.2 | The unit on offer must be the latest technology.
7.2.1 | Initial Manufacturing year of Model:
RT2-08-004 | Yttrium Aluminum Garnet and Selective Laser Trabeculoplasty (Yag and Slt)- Accessories and Consumables | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
RT2-08-004a | Dust Cover
RT2-08-004b | Two pairs of safety goggles.
RT2-08-004c | Mainster PRP 165
RT2-08-004d | Mainster Standard 90
RT2-08-004e | Mainster Wide field 118
RT2-08-004f | Ritch-trabeculoplasty
RT2-08-004g | Latina SLT Gonio
RT2-08-004h | Abraham capsul YAG
RT2-08-004i | Abraham iridotomy
RT2-08-004j | Peyman Wide Field
RT2-08-004k | Goldmann 3 mirror

## Luminometer
RT2-08-005 | Luminometer | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
1 | DESCRIPTION
1.1 | a device used to measure the intensity of light, typically in the visible or ultraviolet spectrum. It quantifies the amount of light emitted by a sample, substance or reaction.
2 | Product Offered:
2.1 | Make:
2.2 | Model:
2.3 | Initial Manufacturing year of Model:
3 | SPECIFICATIONS
3.1 | The unit must be easy to use
3.2 | The unit must be battery operated instrument
3.3 | High sensitivity and linearity:
3.3.1 | Detection of less than 50amolATP or less 1000 luciferase molecules in one reporter gene assay
3.3.2 | Dynamic range of 5.5 orders of magnitude
3.4 | Two Basic measurement Modes
3.4.1 | Tailor-made mode for running Mycoplasma Detection Assay
3.4.2 | Quick Single Read mode for unprocessed luminescence readings
3.5 | High Data storage capacity & PC Upload Option
3.5.1 | Internal memory with maximum storage of up to 1500 single readings
3.5.2 | Optionally sample results may be uploaded to PC and be viewd and analysed with Microsoft excel
3.6 | Broad range of applications
3.6.1 | Mycoplasma Detection Assay
3.6.2 | ATP Assays
3.6.3 | Adenylate kinase assays
3.6.4 | Luciferase reporter gene analysis
4 | Standard Accessories and Conusmables
4.1 | Battery
5 | Conformity Compliance (Please attach certificate)
5.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
5.2 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
5.2.1 | Please provide unique ref number of the ISO 13485 ceriticate:
5.3 | Model on offer must comply with EU Medical Device Regulation (EU) 2017/745 or equvalent from SAHPRA recognized jurisdictions . Attach a copy of certification
5.3.1 | Please provide unique ref number of the ceriticate:
5.4 | Hazardous Substance Act: If this type of equipment / apparatus appears on the schedule of Hazardous substances issued by the Directorate: Health Technology of the Department of Health, a license in terms of the Act on Hazardous Substances (Act 15/1973) must be submitted with the bid / quotation documents. The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.5 | All electrical/electronic medical equipment must be licensed by Radiation Control, where applicable a copy of the license must be submitted The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.5.1 | Radiation Control License number:
5.6 | The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade (rubber to with 2 screws) type and it must be a minimum length of three (3) meters. NB the mains cable of the unit tendered for must be SABS colour coded.
5.7 | The equipment quoted for must be protected against electromagnetic interference IEC 60601-1-2
5.8 | The unit must comply with electrical safety and performance standard such as IEC 60601-2-21, attached certification
5.8.1 | Please provide unique ref number of the IEC ceriticate:
6 | COMPREHENSIVE MAINTENANCE
6.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) should be quoted for, as per the SCC document
6.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
6.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
6.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
6.3 | The bidder to provide the OEM recommended service interval of the unit:
6.4 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
RT2-08-005 | Luminometer -Accessories and Consumables | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
RT2-08-005a | Battery

## OCT Unit
RT2-08-006 | Optical Coherence Tomography (Oct) with Angiography | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
1 | Description
1.1 | is a non-invasive imaging technique used primarily in ophthalmology to visualize blood flow in the retina and choroid. It combines the high-resolution structural imaging of standard OCT with detailed mapping of blood vessels, enabling clinicians to assess both anatomy and blood flow in the eye.
2 | Product Offered:
2.1 | Make:
2.2 | Model:
2.3 | Initial Manufacturing year of Model:
2.4 | Radiation Control License number:
3 | Specification
3.1 | The specification establishes the requirements, supply, delivery, end user training, demonstration, commissioning and installation of a fully automated Optical Coherence Tomography (OCT) with Angiography or equivalent software that provide detailed retinal and vascular imaging as well as reliable glucoma change analysis and anterior segment population for the busy advanced care practise.
3.2 | The design of the OCT must be such that it is useful for imaging modality, evaluation of common ophthalmologic diseases such age -related macular degeneration(AMD), diabetic retinopathy, artery and vein occlusions, and glaucoma.
3.3 | The OCT must be a non-invasive imaging technique that is capable of generating volumetric angiography images in a matter of seconds.
3.4 | The OCT must be a non-invasive technique that acquires volumetric angiographic information without the use of dye.
3.5 | The OCT must be quick and non-invasive, and provides volumetric data with the clinical capability of specifically localizing and delineating pathology along with the ability to show both structural and blood flow information in tandem.
3.6 | The OCT obtains volumetric scans of 304 × 304 A-scans at 70,000 A-scans per second in approximately 3.0 seconds. The software offers the option of 3 × 3 mm, 6 × 6 mm, and 8 × 8 mm OCT angiograms and automated segmentation of these full-thickness retinal scans into the “superficial” and “deep” inner retinal vascular plexuses, outer retina, and choriocapillaris.
3.7 | The OCT must provides a highly detailed view of the retinal vasculature, which allows for accurate delineation of the foveal avascular zone (FAZ) in diabetic eyes and detection of subtle microvascular abnormalities in diabetic and vascular occlusive eyes.
3.8 | The OCT must provides both structural and functional (i.e. blood flow) information in tandem.
3.9 | The OCT must feature the following
3.9.1 | Angioplex or equivalent OCT Angiography
3.9.2 | Angioplex or equivalent Metrix and change analysis
3.9.3 | Minimum Intensity Projection
3.9.4 | Smart HD scans
3.9.5 | FastTrac Retinal eye tracking
3.9.6 | Macular thickness analysis
3.9.7 | Macular change analysis
3.9.8 | ONH and RNFL OU analysis
3.9.9 | Enhanced guided progression analysis
3.9.10 | Advanced RPE analysis
3.9.11 | Ganglion cell analysis
3.9.12 | PnoMap
3.9.13 | Ganglion cell progression analysis
3.9.14 | EN Face analysis
3.9.15 | Single eye Summary
3.9.16 | Anteria segment basic
3.9.17 | The Anterior Segment Premier module which includes external lens kit and additional software with new scans and analysis for anterior OCT imaging and analysis must be provided with the unit.
3.9.18 | Asian normative database
3.9.19 | Rich analysis and report functions
3.9.20 | True colour fundus photography(panaramic or mosaic overview or stereo images must be possible)
3.9.21 | The OCT must also be desgned for high cross-sectional and three dimensional scans, as well as resolution ocular fundus imaging
3.9.22 | It must be possible to detect early stages retinal disorders and advanced studies of retinal pathology that includes factors that compromise photoreceptor function and choroidal vasculature.
3.9.23 | The OCT must be desgned such that it has an innovative technology for examining and imaging the retina and other eye structures. State the mechanism used.
3.9.24 | The technology used must not only offers documentation of clinical findings but it also often highlights critical diagnostic detail not visible on traditional clinical ophthalmoscopy.
3.10 | The design must include:
3.10.1 | PC with processor
3.10.2 | PC rack
3.10.3 | An intregrated or mounted 32cm flat screen monitor
3.10.4 | Network archiving capability and DICOM
3.10.5 | Motorized optical table
3.10.6 | Horizontal and vertical movement of the camera(up &down;right and left)
3.10.7 | Chinerest movement(up and down, motorized)
3.10.8 | Auto/Man alignment
3.10.9 | Auto/Man focus
3.10.10 | Auto/Man shoot
3.10.11 | Auto tracking(up and down;right and left;back and forth)
3.10.12 | Dioptric compensation for patients eyes
3.10.13 | The working distance of the camera should be a least 46mm
3.10.14 | A built-in camera of about 12-megapixel CCD camera or above
3.10.15 | OCT light source
3.11 | The Anterior Segment Premier module with includes external lens kit and additional software with new scans and analysis for anterior OCT imaging and analysis
3.12 | It must also be possible to capture patient information, and:
3.12.1 | Selecting and capture an icon
3.12.2 | Adjusting the chin position
3.13 | Though the OCT is fully automated it must be possible to do semi-auto capturing i.e to allow an operator to start capturing at any conenience. This enables to easily find the best timing to capture with communicating with patient even in difficult cases.
3.14 | For capturing small pupils with minimum diameter of 2.5mm live FUNDUS view must possible.
3.15 | Where there is opacity in the eye due to cataract, it must be possible for the operator to switch on the cataract mode with one finger touch.
3.16 | For flexible alignment, operation with joystick must be possible. This must be included in the bidders price
3.17 | Standard mode for general screening, analysis and professional mode for advanced screening and analysis must be possible.
3.18 | The unit must be upgradable
3.19 | For protection of sensitive parts, a dust cover must be supplied with the unit.
3.20 | The OCT must have auto fluorescence and ICG Functions.
3.21 | Network Specification
3.21.1 | Operating System:Windows 7, 8 and 10; 64-bit OS compatible. The operating system must be upgradable starting with the most rescent.
3.21.2 | Processor Speed: 3.0 GHz; Intel Quad core (desktop)
3.21.3 | Network Bandwidth: 1 Gbps or higher
3.21.4 | Computer RAM: 8 GB or higher
3.21.5 | Monitor Resolution: 1920 x 1080 at 32-bit color
3.21.6 | Printer for both colour and black & white must be provided.
3.22 | Table Specifications(mm)
3.22.1 | Width (W): 950 (37.4) min
3.22.2 | Length (L): 700 (27.6) min
3.22.3 | Height (H): 690~890 (27.2~35.2) min
3.22.4 | Weight: 65 KG (143 LB) min
4 | Standard Accessories and Consumables
4.1 | Dust cover
4.2 | Printer for both colour and black & white must be provided.
5 | Conformity Compliance (Please attach certificate)
5.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
5.2 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
5.2.1 | Please provide unique ref number of the ISO 13485 ceriticate:
5.3 | Model on offer must comply with EU Medical Device Regulation (EU) 2017/745 or equvalent from SAHPRA recognized jurisdictions . Attach a copy of certification
5.3.1 | Please provide unique ref number of the ceriticate:
5.4 | Hazardous Substance Act: If this type of equipment / apparatus appears on the schedule of Hazardous substances issued by the Directorate: Health Technology of the Department of Health, a license in terms of the Act on Hazardous Substances (Act 15/1973) must be submitted with the bid / quotation documents. The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.5 | All electrical/electronic medical equipment must be licensed by Radiation Control, where applicable a copy of the license must be submitted The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.5.1 | Radiation Control License number:
5.6 | The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade (rubber to with 2 screws) type and it must be a minimum length of three (3) meters. NB the mains cable of the unit tendered for must be SABS colour coded.
5.7 | The equipment quoted for must be protected against electromagnetic interference IEC 60601-1-2
5.8 | The unit must comply with electrical safety and performance standard such as IEC 60601-2-21, attached certification
5.8.1 | Please provide unique ref number of the IEC ceriticate:
6 | COMPREHENSIVE MAINTENANCE
6.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) should be quoted for, as per the SCC document
6.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
6.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
6.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
6.3 | The bidder to provide the OEM recommended service interval of the unit:
6.4 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
7 | General
7.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
7.2 | The unit on offer must be the latest technology.
7.2.1 | Initial Manufacturing year of Model:
RT2-08-006 | Optical Coherence Tomography (Oct) with Angiography , Replacemengt accessories | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
RT2-08-006a | Dust cover
RT2-08-006b | Printer for both colour and black & white must be provided.

## Phoropter
RT2-08-007 | Phoropter | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
1 | DESCRIPTION
1.1 | An instrument for measuring the ametropias, phorias and the amplitude of accommodation of the eyes
2 | Product Offered:
2.1 | Make:
2.2 | Model:
2.3 | Initial Manufacturing year of Model:
2.4 | Radiation Control License number:
3 | SPECIFICATIONS
3.1 | The system offered must have the following test modes.
3.1.1 | Precise astigmatism test
3.1.2 | Accurate presbiopia test
3.1.3 | Precision in biocular test (Fixation disparity)
3.2 | REFRACTION must meet the following specifications:
3.2.1 | The measuring range (Diopter-D) for sphere power +16, 5 D to -19(0, 25 D steps)
3.2.2 | The measuring range cylinder power at 0D to -9D steps 0,25 D
3.2.3 | The AXIS measuring range must be at least 0° to 180° in a maximum of 5 degree steps.
3.3 | The following facilities/features must be available and provided:
3.3.1 | Auxilary lens diat,pd adjustment,standard accessarry lens
3.3.2 | The papillary distance (PD) must be calculated and shown on the print out.-48-80 mm in 1mm steps
3.3.3 | Standard accessing –One near point card with rod holder
3.4 | Polycarbon Cover must meet the following specifications:
3.4.1 | The CORNEAL CURVATURE RADIUS measuring range must be at least 5mm to 10mm in a maximum of 0,01mm steps.
3.4.2 | The CORNEAL REFRACTION measuring range must be at least 67D to 34 D IN 0,12D steps
3.4.3 | The CORNEAL ASTIGMATISM ANGLE measuring range must be at least 0° to 180° in maximum of 1° step.
3.4.4 | The read out display must be selectable in either MILLMETER or DIOPTER.
4 | STANDARD ACCESSORIES
4.1 | Dust cover
5 | Conformity Compliance (Please attach certificate)
5.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
5.2 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
5.2.1 | Please provide unique ref number of the ISO 13485 ceriticate:
5.3 | Model on offer must comply with EU Medical Device Regulation (EU) 2017/745 or equvalent from SAHPRA recognized jurisdictions . Attach a copy of certification
5.3.1 | Please provide unique ref number of the ceriticate:
5.4 | Hazardous Substance Act: If this type of equipment / apparatus appears on the schedule of Hazardous substances issued by the Directorate: Health Technology of the Department of Health, a license in terms of the Act on Hazardous Substances (Act 15/1973) must be submitted with the bid / quotation documents. The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.5 | All electrical/electronic medical equipment must be licensed by Radiation Control, where applicable a copy of the license must be submitted The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.5.1 | Radiation Control License number:
5.6 | The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade (rubber to with 2 screws) type and it must be a minimum length of three (3) meters. NB the mains cable of the unit tendered for must be SABS colour coded.
5.7 | The equipment quoted for must be protected against electromagnetic interference IEC 60601-1-2
5.8 | The unit must comply with electrical safety and performance standard such as IEC 60601-2-21, attached certification
5.8.1 | Please provide unique ref number of the IEC ceriticate:
6 | COMPREHENSIVE MAINTENANCE
6.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) should be included in the unit price as per the SCC document
6.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
6.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
6.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
6.3 | The bidder to provide the OEM recommended service interval of the unit:
6.4 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
7 | General
7.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
7.2 | The unit on offer must be the latest technology.
7.2.1 | Initial Manufacturing year of Model:
RT2-08-007 | Phoropter, Replacement accessories and consumables | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
RT2-08-007a | Dust cover

## Slit-Lamp, Basic
RT2-08-008 | Ophthalmic Slit Lamp | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
1 | DESCRIPTION
1.1 | Medical instrument used for eye examinations, providing detailed view of the anterior segmant of the eye.
2 | Product Offered:
2.1 | Make:
2.2 | Model:
2.3 | Initial Manufacturing year of Model:
3 | SPECIFICATIONS
3.1 | The specification requires a complete slit lamp installed on a stable table with Halogen/LED illumination
3.2 | The slit lamp must have the following features:
3.2.1 | Range for Magnification: from 6x to 40x or above
3.2.2 | Field of view:5mm to 40mm
3.2.3 | Rotation of slit lamp: 180 degrees with reference scale
3.2.4 | Swivel range of still projector +-180 degrees
3.2.5 | Angle of incidence: various steps
3.2.6 | Filters:Blue ; Green and White
3.2.7 | Light source: Halogen/LED. Bidder to state the light source
3.2.8 | Travel of instrument base: Vertical; lateral and axial
3.2.9 | Brighteness: continous adjustable
3.2.10 | Provision for mounting the applanation Tonometer must be provided
3.2.11 | The slit lamp must be provided with chinrest
4 | STANDARD ACCESSORIES
4.1 | dust cover
4.2 | If the unit offered is Halogen illumination, the spare bulb is required.
5 | Conformity Compliance (Please attach certificate)
5.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
5.2 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
5.2.1 | Please provide unique ref number of the ISO 13485 ceriticate:
5.3 | Model on offer must comply with EU Medical Device Regulation (EU) 2017/745 or equvalent from SAHPRA recognized jurisdictions . Attach a copy of certification
5.3.1 | Please provide unique ref number of the ceriticate:
5.4 | Hazardous Substance Act: If this type of equipment / apparatus appears on the schedule of Hazardous substances issued by the Directorate: Health Technology of the Department of Health, a license in terms of the Act on Hazardous Substances (Act 15/1973) must be submitted with the bid / quotation documents. The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.5 | All electrical/electronic medical equipment must be licensed by Radiation Control, where applicable a copy of the license must be submitted The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.5.1 | Radiation Control License number:
5.6 | The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade (rubber to with 2 screws) type and it must be a minimum length of three (3) meters. NB the mains cable of the unit tendered for must be SABS colour coded.
5.7 | The equipment quoted for must be protected against electromagnetic interference IEC 60601-1-2
5.8 | The unit must comply with electrical safety and performance standard such as IEC 60601-2-21, attached certification
5.8.1 | Please provide unique ref number of the IEC ceriticate:
6 | COMPREHENSIVE MAINTENANCE
6.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) should be included in the unit price as per the SCC document
6.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
7 | General
7.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
7.2 | The unit on offer must be the latest technology.
7.2.1 | Initial Manufacturing year of Model:
RT2-08-008 | Ophthalmic Slit Lamp, Replacement accessories and consumables | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
RT2-08-008a | LED replacement bulb
RT2-08-008b | Halogen replacement bulb

## Slit-Lamp, Advanced
RT2-08-009 | Ophthalmic Slit Lamp, Integrated Imaging Module | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
1 | DESCRIPTION
1.1 | Medical instrument used for eye examinations, providing detailed view of the anterior segmant of the eye.
2 | Product Offered:
2.1 | Make:
2.2 | Model:
2.3 | Initial Manufacturing year of Model:
3 | SPECIFICATIONS
3.1 | The specification requires a complete slit lamp installed on electronic controlled table with LED illumination and a monitor
3.2 | The slit lamp must have the following features:
3.2.1 | Range for Magnification: from 6x to 40x or above
3.2.2 | Field of view:5mm to 40mm
3.2.3 | Rotation of slit lamp: 180 degrees with reference scale
3.2.4 | Swivel range of still projector +-180 degrees
3.2.5 | Angle of incidence: various steps
3.2.6 | Filters:Blue ; Green and White
3.2.7 | Light source: LED
3.2.8 | Travel of instrument base: Vertical; lateral and axial
3.2.9 | Brighteness: continous adjustable
3.2.10 | The Slit lamp must be supllied with applanation Tonometer.
3.2.11 | Provision for mounting the applanation Tonometer must be provided
3.2.12 | The slit lamp must be provided with chinrest
3.3 | Imaging module
3.3a | Make:
3.3b | Model:
3.3.1 | The slit lamp must have an intergrated digital camera system (imaging software) and video recording of all examination
3.3.2 | Preferance will be given to units with 8 -to 10 megapixel camera or above.
3.3.3 | The unit offered must allow easy capturing of images. The bidder must specify technology for capturing .
3.3.4 | For teaching and viewing purposes the unit must be provided with monitor of at least 21" with minimum 1280 x 1024 resolution
3.3.5 | All setting must be software driven to make capturing and manipulation of images easier and faster
3.3.6 | The software must allow the user to be able to store and retrieve images.
3.3.7 | The monitor must be offered with mounting swivel bracket and mounted as per end-user requirement.
3.3.8 | The unit must be supplied with storage space of at least 120GB
3.3.9 | The unit must have a dedicated hieght adjustable chair.
4 | STANDARD ACCESSORIES
4.1 | dust cover
4.2 | Hieght adjustable chair
5 | Conformity Compliance (Please attach certificate)
5.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
5.2 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
5.2.1 | Please provide unique ref number of the ISO 13485 ceriticate:
5.3 | Model on offer must comply with EU Medical Device Regulation (EU) 2017/745 or equvalent from SAHPRA recognized jurisdictions . Attach a copy of certification
5.3.1 | Please provide unique ref number of the ceriticate:
5.4 | Hazardous Substance Act: If this type of equipment / apparatus appears on the schedule of Hazardous substances issued by the Directorate: Health Technology of the Department of Health, a license in terms of the Act on Hazardous Substances (Act 15/1973) must be submitted with the bid / quotation documents. The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.5 | All electrical/electronic medical equipment must be licensed by Radiation Control, where applicable a copy of the license must be submitted The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.5.1 | Radiation Control License number:
5.6 | The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade (rubber to with 2 screws) type and it must be a minimum length of three (3) meters. NB the mains cable of the unit tendered for must be SABS colour coded.
5.7 | The equipment quoted for must be protected against electromagnetic interference IEC 60601-1-2
5.8 | The unit must comply with electrical safety and performance standard such as IEC 60601-2-21, attached certification
5.8.1 | Please provide unique ref number of the IEC ceriticate:
6 | COMPREHENSIVE MAINTENANCE
6.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) should be included in the unit price as per the SCC document
6.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
7 | General
7.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
7.2 | The unit on offer must be the latest technology.
7.2.1 | Initial Manufacturing year of Model:
RT2-08-009 | OPHTHALMIC SLIT LAMP, intergrated imaging module Replacement accessories and consumables | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
RT2-08-009a | LED replacement
RT2-08-009b | Monitor
RT2-08-009c | Mounting swivel bracket
RT2-08-009d | Dedicated digital camera system

## Tonometer, Hand-held
RT2-08-010 | Tonometer, Hand-Held | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
1 | DESCRIPTION
1.1 | A portable lightweight medical device used to measusre intraocular pressure in the eye.
2 | Product Offered:
2.1 | Make:
2.2 | Model:
2.3 | Initial Manufacturing year of Model:
3 | SPECIFICATIONS
3.1 | The unit must use LED illumination
3.2 | The unit offered must have a measurement range of 0 - 60mmHg
3.3 | The unit must be of light weight not exceeding 1kg
3.4 | The unit should be powered from a rechargeable battery
3.5 | The unit must have an indicator of the battery status
3.6 | The battery must be able to last up to at least 10 000 readings
3.7 | The unit must be supplied with a battery charger and hard case for storage.
4 | STANDARD ACCESSORIES
4.1 | Battery charger
4.2 | Tip covers (pk of 10)
4.3 | Hard case for storage
5 | Conformity Compliance (Please attach certificate)
5.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
5.2 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
5.2.1 | Please provide unique ref number of the ISO 13485 ceriticate:
5.3 | Model on offer must comply with EU Medical Device Regulation (EU) 2017/745 or equvalent from SAHPRA recognized jurisdictions . Attach a copy of certification
5.3.1 | Please provide unique ref number of the ceriticate:
5.4 | Hazardous Substance Act: If this type of equipment / apparatus appears on the schedule of Hazardous substances issued by the Directorate: Health Technology of the Department of Health, a license in terms of the Act on Hazardous Substances (Act 15/1973) must be submitted with the bid / quotation documents. The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.5 | All electrical/electronic medical equipment must be licensed by Radiation Control, where applicable a copy of the license must be submitted The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.5.1 | Radiation Control License number:
5.6 | The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade (rubber to with 2 screws) type and it must be a minimum length of three (3) meters. NB the mains cable of the unit tendered for must be SABS colour coded.
5.7 | The equipment quoted for must be protected against electromagnetic interference IEC 60601-1-2
5.8 | The unit must comply with electrical safety and performance standard such as IEC 60601-2-21, attached certification
5.8.1 | Please provide unique ref number of the IEC ceriticate:
6 | COMPREHENSIVE MAINTENANCE
6.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) should be included in the unit price as per the SCC document
6.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
7 | General
7.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
7.2 | The unit on offer must be the latest technology.
7.2.1 | Initial Manufacturing year of Model:
RT2-08-010 | Tonometer - Accessories and Consumables | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
RT2-08-010a | Battery charger
RT2-08-010b | Tip covers (pk of 10)
RT2-08-010c | Rechargeable battery
RT2-08-010d | Hard case for storage

## Visual Field Analyser
RT2-08-011 | Visual Field Analyzer | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
1 | Description
Medical device used to measure the field of vision, detecting defects and abnormalities in the visual pathway.
2 | Product Offered:
2.1 | Make:
2.2 | Model:
2.3 | Initial Manufacturing year of Model:
2.4 | Radiation Control License number:
3 | Specification
3.1 | The specification requires a complete Visual Field Analyzer installed on hieght adjustable table
3.2 | The unit must be used for measuring the visual field by detecting monocular visual field and also for screening, monitoring and assiting in the diagnosis of certain conditions in a variety of opthalmic and neurologic diseases
3.3 | The unit must provide information regarding the location of any disease processes or lesion throughout the visual pathway
3.4 | The design of the unit should be such that the patient is presented with a series of light stimuli of varying intensities in different visual field locations and the threshold sensitivity of the retina is assessed based on the patient ability to consciously detect and respond the these stimuli.
3.5 | The unit must able to distinguish focal from diffuse vision loss.
3.6 | The unit tendered must be designed such that it can be used on patient who are wheelchair users, hearing impaired, have postural and fixation problems and low visual acuity
3.7 | Touch-screen and menu-driven interfaces must simplify and expedite instrument operation
3.8 | The unit must offer ergonomic design for maximum comfort, access and versatility.
3.9 | Complete Archiving system. State the method(s) used.
3.10 | The unit must be provided with a software that can be used to reduce blue-yellow threshold test time and also provide a clinically practical tool for early detection of glucoma.
3.11 | The unit must able to operate in two modes:
3.11.1 | Manual:With patient response registered with a mouse click
3.11.2 | Visual grasp: where the eye tracker senses change in gaze direction as evidence of target acquisition.
3.12 | The sensitivity range is from 0-40dB
3.13 | The following hardware components must be provided with the unit.
3.13.1 | Organic light-emiting diode(OLED) as opposed to LCD's
3.13.2 | Binocular Viewpoint Eye Tracker system which uses infrared video tracking with temporal resolution of 60 Hz.
3.13.3 | For patient requiring far field correction, standard trial lenses must be accomondated by lense holders.
3.13.4 | Intraocular distance may be adjusted with push pull mechanism that permits independent movement of each display assembly in the horizontal direction.
3.14 | The device must be operated through a portable windows computer that provides access to the four essential functionalities:
3.14.1 | Test setup, input and display of patient information
3.14.2 | Color-coded display of visual field measurement
3.14.3 | VF Expert automatic interpretation of visual field measurement
3.14.4 | Access to archived data
3.15. | The Monitor capability should include but not limited to the following
3.15.1 | Social Security Disability testing
3.15.2 | Monocular testing
3.15.3 | Binocular testing
3.15.4 | Kinetic Testing
3.15.5 | Custom Testing
3.16 | The unit must display rate of vision loss relative to patient age for individual patient care.
3.17 | Streamlines clinical interpretation and simplifies patient education
3.18 | A mechanism to limit the need for trial lense correction in patients
3.19 | The visual field testing distance must be infinity
3.20 | The minimum stimulus duration should be 200-300ms
3.21 | The unit offered must have the following Fixation control features:
3.21.1 | Blind spot monitoring
3.21.2 | Video eye monitoring
3.21.3 | Gaze tracking
3.21.4 | Head tracking
3.21.5 | Vertex tracking |  | 2
3.21.6 | Remote eye monitor capability
3.22 | The operator interface must be an touch screen with option to input changes.
3.23 | Integrated data storage, retrieval and analysis the following must be available
3.23.1 | PC-based
3.23.2 | Integrated hard drive with storage capacity of 100GB or above
3.23.3 | USB
3.24 | A colour printer with printing papers must be provided with the unit.
3.25 | The software features must include the following :
3.25.1 | Single field analysis
3.25.2 | Visual field index
3.25.3 | Guided Progression Analysis
3.25.4 | Serial field overview
3.25.5 | Networking
3.25.6 | Glacoma Hemifield Test
3.25.7 | DICOM Connectivity
3.26 | When the patient's head is placed on the machine, it must be possible to adjust the patient's head using directional key buttons (i.e vertical and horizontal movements) or any mechanism of the latest technology. This allows the user not to manualy adjust the patient head for the focal point(s) as some patients may be critically ill and minimise time for the whole process.
4 | Standard Accessories and Consumables
4.1 | Dust Cover
5 | Conformity Compliance (Please attach certificate)
5.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
5.2 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
5.2.1 | Please provide unique ref number of the ISO 13485 ceriticate:
5.3 | Model on offer must comply with EU Medical Device Regulation (EU) 2017/745 or equvalent from SAHPRA recognized jurisdictions . Attach a copy of certification
5.3.1 | Please provide unique ref number of the ceriticate:
5.4 | Hazardous Substance Act: If this type of equipment / apparatus appears on the schedule of Hazardous substances issued by the Directorate: Health Technology of the Department of Health, a license in terms of the Act on Hazardous Substances (Act 15/1973) must be submitted with the bid / quotation documents. The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.5 | All electrical/electronic medical equipment must be licensed by Radiation Control, where applicable a copy of the license must be submitted The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.5.1 | Radiation Control License number:
5.6 | The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade (rubber to with 2 screws) type and it must be a minimum length of three (3) meters. NB the mains cable of the unit tendered for must be SABS colour coded.
5.7 | The equipment quoted for must be protected against electromagnetic interference IEC 60601-1-2
5.8 | The unit must comply with electrical safety and performance standard such as IEC 60601-2-21, attached certification
5.8.1 | Please provide unique ref number of the IEC ceriticate:
6 | COMPREHENSIVE MAINTENANCE
6.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) should be included in the unit price as per the SCC document
6.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
6.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
6.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
6.3 | The bidder to provide the OEM recommended service interval of the unit:
6.4 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
7 | General
7.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
7.2 | The unit on offer must be the latest technology.
7.2.1 | Initial Manufacturing year of Model:

## Vitrectomy Unit P&A
RT2-08-012 | Vitrectomy Unit (Anterior and Posterior) | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
1 | DESCRIPTION
1.1 | Ophthalmic surgical units designed to remove vitreous gel and any blood or debris from the eye (vitrectomy).
1.2 | This facility on the unit offered must include the following selectable features:
1.2.1 | Phacoemulsification
1.2.2 | Irrigation
1.2.3 | Irrigation / Aspiration
1.2.4 | Diathermy
1.2.5 | Anterior Vitrectomy
1.2.6 | Posterior Vitrectomy
2 | Product Offered:
2.1 | Make:
2.2 | Model:
2.3 | Initial Manufacturing year of Model:
2.4 | Radiation Control License number:
3 | PHACOEMULSIFICATION:
3.1 | This feature on the unit offered must include the following selectable modes and specifications:
3.1.1 | A fixed mode in which phacoemulsification power and aspiration level are preselected on the console.
3.1.2 | A linear mode in which phacoemulsification power is fixed while aspiration level is linearly controlled via the footswitch.
3.1.3 | A linear mode in which phacoemulsification power is footswitch controlled while aspiration level is preselected on the console.
3.1.4 | A pulsed mode must be available and cover a range of 1 (one) pulse to at least 12 (twelve) pulses per second and it must be adjustable in increments of 1 pulse per second.
3.2 | Phacoemulsification power must be displayed on the console as a percentage of maximum available power.
3.3 | Both the average phaco power and the current power when phaco is in use must be displayed on the console.
3.4 | The elapsed phaco time must be displayed on the console in minutes and seconds.
3.5 | The unit must perform phacoemulsification via a fully autoclavable piezo-electric hand piece.
3.6 | The phaco frequency must be between 20 and 40 kilohertz. Bidder must state the exact frequency of the unit offered.
3.7 | Self-calibration in order to ensure that hand piece and the unit are both in tune must be provided on the unit offered.
3.8 | The unit offered must be able to provide self-diagnosis and trouble shooting with user-friendly display indication on the console.
3.9 | The unit offered must permit programable surgeon setting via independent access keys.
3.10 | The unit must be capable of performing phacoemulsification through an incision in the range of 1.8-3mm. Bidder to specify
3.11 | The tubing packs supplied with the system must be disposable and the cost of which must be included in the final bid price.
3.12 | The fluid collection container supplied with the system must be disposable and the cost of which must be included in the final bid price.
3.13 | Irrigation and Irrigation / Aspiration on the unit offered must meet the following specifications:
3.13.1 | This feature must be controlled by means of a footswitch.
3.13.2 | The footswitch on the unit offered must have at least 3 (three) positions permitting switching as follows:
3.13.2a | Irrigation only.
3.13.2b | Irrigation and aspiration
3.13.2c | Irrigation and aspiration with phacoemulsification
3.13.3 | Aspiration on the unit offered must be available both in a preset fixed mode and a linearly variable mode.
3.13.4 | Aspiration must be able to generate suction pressure in a minimum range of 0 to 500mmHg on the unit offered.
3.13.5 | The unit offered must utilise a venturi pump aspiration system.
3.13.6 | The fluid collection cassette must be measureable and capable of holding a minimum of 150cc on the unit offered.
3.13.7 | The unit offered must deliver a clear audio warning signal when the aspiration cassette is nearly full.
3.13.8 | A footswitch-controlled facility must be available on the unit offered to initiate reflux of aspirated contents.
3.13.9 | The audible tone indicating the level of aspiration power being used must be selectable on the unit offered.
3.14 | The diathermy on the system offered must meet the following specifications:
3.14.1 | A fully isolated, footswitch controlled bipolar diathermy pencil must be included.
3.14.2 | Diathermy power must be adjustable from 0 to 100% in reasonable increments. Bidder must state the selectable increments available on the offered unit.
3.14.3 | A volume adjustable audible tone must indicate when diathermy is active.
3.15 | Anterior Vitrectomy specification must be as follows on the system offered:
3.15.1 | This facility must include vitreous cutting, irrigation and aspiration capability under footswitch control.
3.15.2 | Vitreous cutting rate must be adjustable from 0 to at least 2500 cuts per minute.
3.15.3 | The anterior vitrectomy probe must be of the guillotine variety.
3.16 | Posterior segment facility specifications must be as follows on the system offered:
3.16.1 | The unit must perform posterior vitrectomy with a guillotine type cutting probe.
3.16.2 | Cutting rate must be variable from 0 to 2500 cuts per minute.
3.16.3 | Fluid aspiration must be via linear footswitch control with precise maintenance of low end aspiration.
3.16.4 | The unit must perform intra-ocular illumination via an autoclavable, infrared filtered, fibre optic cable capable of producing a variable intensity while lighting.
3.16.5 | There must be an automatic switch over facility to a spare bulb in the event of bulb failure
3.16.6 | The unit offered must include an intra-ocular scissors drive with single cut, multi cut variable rate and proportional cut capabilities. These must be preset and footswitch operated.
3.16.7 | The unit offered must include an intra-ocular fully isolated bipolar diathermy facility.
3.16.8 | The unit offered must include an air infusion facility which can operate constantly or variably to maintain intra-ocular pressure (IOP) to preset levels.
3.16.9 | The unit offered must include a phaco fragmentation facility capable of performing fixed or pulsed fragmentation with linear aspiration control.
4 | STANDARD ACCESSORIES
4.1 | Phacoemulsification hand pieces x2.
4.2 | 30-degree phaco needles (10/box).
4.3 | Irrigation / aspiration hand pieces 23 gauge with 0.3mm aspiration port and one must be with a straight tip and the other must be with a curved tip (x2)
4.4 | Irrigation sleeves (10/box).
4.5 | Bipolar eraser cautery with connecting cable (10/box)
4.6 | Anterior vitrectomy probe.
4.7 | Posterior vitrectomy probe .
4.8 | Disposable tubing and cassette (10/box)
4.9 | Irrigation and phaco pack (10/box)
4.10 | Microscissors driver tubing and cable
4.11 | Fibre optic cable and probe
4.12 | Intra-ocular eraser type diathermy probe with cable
4.13 | Three infusion cannulas 2.5; 4 and 6mm.
4.14 | Pheumatic vitreous scissors and membrane peeler
4.15 | Machmyer or equivalent vitreous lens.
4.16 | Posterior intra-ocular pressure control tubing sets (10/box)
4.17 | Dust cover
5 | Conformity Compliance (Please attach certificate)
5.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
5.2 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
5.2.1 | Please provide unique ref number of the ISO 13485 ceriticate:
5.3 | Model on offer must comply with EU Medical Device Regulation (EU) 2017/745 or equvalent from SAHPRA recognized jurisdictions . Attach a copy of certification
5.3.1 | Please provide unique ref number of the ceriticate:
5.4 | Hazardous Substance Act: If this type of equipment / apparatus appears on the schedule of Hazardous substances issued by the Directorate: Health Technology of the Department of Health, a license in terms of the Act on Hazardous Substances (Act 15/1973) must be submitted with the bid / quotation documents. The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.5 | All electrical/electronic medical equipment must be licensed by Radiation Control, where applicable a copy of the license must be submitted The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.5.1 | Radiation Control License number:
5.6 | The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade (rubber to with 2 screws) type and it must be a minimum length of three (3) meters. NB the mains cable of the unit tendered for must be SABS colour coded.
5.7 | The equipment quoted for must be protected against electromagnetic interference IEC 60601-1-2
5.8 | The unit must comply with electrical safety and performance standard such as IEC 60601-2-21, attached certification
5.8.1 | Please provide unique ref number of the IEC ceriticate:
6 | COMPREHENSIVE MAINTENANCE
6.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) should be included in the unit price as per the SCC document
6.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
6.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
6.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
6.3 | The bidder to provide the OEM recommended service interval of the unit:
6.4 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
7 | General
7.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
7.2 | The unit on offer must be the latest technology.
7.2.1 | Initial Manufacturing year of Model:
RT2-08-012 | Vitrectomy Unit (Anterior and Posterior) -Accessories and Consumables | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
RT2-08-012a | Phacoemulsification hand pieces
RT2-08-012b | 30-degree phaco needles (10/box).
RT2-08-012c | Irrigation / aspiration hand pieces 23 gauge with 0.3mm aspiration port and one must be with a straight tip and the other must be with a curved tip (x2)
RT2-08-012d | Irrigation sleeves (10/box).
RT2-08-012e | Bipolar eraser cautery with connecting cable (10/box)
RT2-08-012f | Anterior vitrectomy probe.
RT2-08-012g | Irrigation and phaco pack (10/box)
RT2-08-012h | Posterior vitrectomy probe with re-usable and disposable tubing and cassette (10/box)
RT2-08-012i | Microscissors driver tubing and cable
RT2-08-012j | Fibre optic cable and probe
RT2-08-012k | Intra-ocular eraser type diathermy probe with cable
RT2-08-012l | Three infusion cannulas 2.5; 4 and 6mm.
RT2-08-012m | Pheumatic vitreous scissors and membrane peeler
RT2-08-012n | Machmyer vitreous lens.
RT2-08-012o | Posterior intra-ocular pressure control tubing sets (10/box)
RT2-08-012p | Dust cover

## Vitrectomy Unit-A
RT2-08-013 | VITRECTOMY UNIT (Anterior) | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
1 | DESCRIPTION
1.1 | Ophthalmic surgical units designed to remove vitreous gel and any blood or debris from the eye (vitrectomy).
1.2 | This facility on the unit offered must include the following selectable features:
1.2.1 | Phacoemulsification
1.2.2 | Irrigation
1.2.3 | Irrigation / Aspiration
1.2.4 | Diathermy
1.2.5 | Anterior Vitrectomy
2 | Product Offered:
2.1 | Make:
2.2 | Model:
2.3 | Initial Manufacturing year of Model:
2.4 | Radiation Control License number:
3 | PHACOEMULSIFICATION:
3.1 | This feature on the unit offered must include the following selectable modes and specifications:
3.1.1 | A fixed mode in which phacoemulsification power and aspiration level are preselected on the console.
3.1.2 | A linear mode in which phacoemulsification power is fixed while aspiration level is linearly controlled via the footswitch.
3.1.3 | A linear mode in which phacoemulsification power is footswitch controlled while aspiration level is preselected on the console.
3.1.4 | A pulsed mode must be available and cover a range of 1 (one) pulse to at least 12 (twelve) pulses per second and it must be adjustable in increments of 1 pulse per second.
3.2 | Phacoemulsification power must be displayed on the console as a percentage of maximum available power.
3.3 | Both the average phaco power and the current power when phaco is in use must be displayed on the console.
3.4 | The elapsed phaco time must be displayed on the console in minutes and seconds.
3.5 | The unit must perform phacoemulsification via a fully autoclavable piezo-electric hand piece.
3.6 | The phaco frequency must be between 20 and 40 kilohertz. Bidder must state the exact frequency of the unit offered.
3.7 | Self-calibration in order to ensure that hand piece and the unit are both in tune must be provided on the unit offered.
3.8 | The unit offered must be able to provide self-diagnosis and trouble shooting with user-friendly display indication on the console.
3.9 | The unit offered must permit programable surgeon setting via independent access keys.
3.10 | The unit must be capable of performing phacoemulsification through an incision in the range of 1.8-3mm. Bidder to specify
3.11 | The tubing packs supplied with the system must be disposable and the cost of which must be included in the final bid price.
3.12 | The fluid collection container supplied with the system must be disposable and the cost of which must be included in the final bid price.
3.13 | Irrigation and Irrigation / Aspiration on the unit offered must meet the following specifications:
3.13.1 | This feature must be controlled by means of a footswitch.
3.13.2 | The footswitch on the unit offered must have at least 3 (three) positions permitting switching as follows:
3.13.2a | Irrigation only.
3.13.2b | Irrigation and aspiration
3.13.2c | Irrigation and aspiration with phacoemulsification
3.13.3 | Aspiration on the unit offered must be available both in a preset fixed mode and a linearly variable mode.
3.13.4 | Aspiration must be able to generate suction pressure in a minimum range of 0 to 500mmHg on the unit offered.
3.13.5 | The unit offered must utilise a venturi pump aspiration system.
3.13.6 | The fluid collection cassette must be measureable and capable of holding a minimum of 150cc on the unit offered.
3.13.7 | The unit offered must deliver a clear audio warning signal when the aspiration cassette is nearly full.
3.13.8 | A footswitch-controlled facility must be available on the unit offered to initiate reflux of aspirated contents.
3.13.9 | The audible tone indicating the level of aspiration power being used must be selectable on the unit offered.
3.14 | The diathermy on the system offered must meet the following specifications:
3.14.1 | A fully isolated, footswitch controlled bipolar diathermy pencil must be included.
3.14.2 | Diathermy power must be adjustable from 0 to 100% in reasonable increments. Bidder must state the selectable increments available on the offered unit.
3.14.3 | A volume adjustable audible tone must indicate when diathermy is active.
3.15 | Anterior Vitrectomy specification must be as follows on the system offered:
3.15.1 | This facility must include vitreous cutting, irrigation and aspiration capability under footswitch control.
3.15.2 | Vitreous cutting rate must be adjustable from 0 to at least 2500 cuts per minute.
3.15.3 | The anterior vitrectomy probe must be of the guillotine variety.
4 | STANDARD ACCESSORIES
4.1 | Phacoemulsification hand pieces x2.
4.2 | 30-degree phaco needles (10/box).
4.3 | Irrigation / aspiration hand pieces 23 gauge with 0.3mm aspiration port and one must be with a straight tip and the other must be with a curved tip (x2)
4.4 | Irrigation sleeves (10/box).
4.5 | Bipolar eraser cautery with connecting cable (10/box)
4.6 | Anterior vitrectomy probe.
4.7 | Disposable tubing and cassette (10/box)
4.8 | Irrigation and phaco pack (10/box)
4.9 | Microscissors driver tubing and cable
4.10 | Fibre optic cable and probe
4.11 | Intra-ocular eraser type diathermy probe with cable
4.12 | Three infusion cannulas 2.5; 4 and 6mm.
4.13 | Pheumatic vitreous scissors and membrane peeler
4.14 | Machmyer or equivalent vitreous lens.
4.15 | Dust cover
5 | Conformity Compliance (Please attach certificate)
5.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
5.2 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
5.2.1 | Please provide unique ref number of the ISO 13485 ceriticate:
5.3 | Model on offer must comply with EU Medical Device Regulation (EU) 2017/745 or equvalent from SAHPRA recognized jurisdictions . Attach a copy of certification
5.3.1 | Please provide unique ref number of the ceriticate:
5.4 | Hazardous Substance Act: If this type of equipment / apparatus appears on the schedule of Hazardous substances issued by the Directorate: Health Technology of the Department of Health, a license in terms of the Act on Hazardous Substances (Act 15/1973) must be submitted with the bid / quotation documents. The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.5 | All electrical/electronic medical equipment must be licensed by Radiation Control, where applicable a copy of the license must be submitted The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.5.1 | Radiation Control License number:
5.6 | The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade (rubber to with 2 screws) type and it must be a minimum length of three (3) meters. NB the mains cable of the unit tendered for must be SABS colour coded.
5.7 | The equipment quoted for must be protected against electromagnetic interference IEC 60601-1-2
5.8 | The unit must comply with electrical safety and performance standard such as IEC 60601-2-21, attached certification
5.8.1 | Please provide unique ref number of the IEC ceriticate:
6 | COMPREHENSIVE MAINTENANCE
6.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) should be included in the unit price as per the SCC document
6.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
6.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
6.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
6.3 | The bidder to provide the OEM recommended service interval of the unit:
6.4 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
7 | General
7.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
7.2 | The unit on offer must be the latest technology.
7.2.1 | Initial Manufacturing year of Model:
RT2-08-013 | Vitrectomy Unit -Accessories and Consumables | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
RT2-08-013a | Phacoemulsification hand pieces
RT2-08-013b | 30-degree phaco needles (10/box).
RT2-08-013c | Irrigation / aspiration hand pieces 23 gauge with 0.3mm aspiration port and one must be with a straight tip and the other must be with a curved tip (x2)
RT2-08-013d | Irrigation sleeves (10/box).
RT2-08-013e | Bipolar eraser cautery with connecting cable (10/box)
RT2-08-013f | Anterior vitrectomy probe.
RT2-08-013g | Irrigation and phaco pack (10/box)
RT2-08-013h | Microscissors driver tubing and cable
RT2-08-013i | Fibre optic cable and probe
RT2-08-013j | Intra-ocular eraser type diathermy probe with cable
RT2-08-013k | Three infusion cannulas 2.5; 4 and 6mm.
RT2-08-013l | Pheumatic vitreous scissors and membrane peeler
RT2-08-013m | Machmyer vitreous lens.
RT2-08-013n | Dust cover

## Category-9 Neuro
NEURO EQUIPMENT (Neurlogy & Neurosurgery)
# | Item number | Item Description
1 | RT2-09-001 | EEG, Portable
2 | RT2-09-002 | EEG, Desktop
3 | RT2-09-003 | Ultrasonic Surgical Aspirator
4 | RT2-09-004 | Neuronavigation System
5 | RT2-09-005 | Neuro-Monitor
6 | RT2-09-006 | Neuro-Drills
7 | RT2-09-007 | Amplitude EEG (Cerebral Function Monitoring)

## EEG, Portable
RT2-09-001 | EEG Portable | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
1 | DESCRIPTION
1.1 | An electronic monitoring device to measure and record electrical activity in the brain.
2 | Product Offered:
2.1 | Make:
2.2 | Model:
2.3 | Initial Manufacturing year of Model:
2.4 | Radiation Control License number:
3 | TECHNICAL SPECIFICATION
3.1 | The EEG unit quoted for must be user friendly, it must comprise the latest technology and it must be a 32 Channel computerized Portable EEG – Pentium based 17" Notebook with the following minimum spec: Core i7, 8GB of RAM, 256GB SSD or better, CD/DVD RW.
3.2 | The Portable EEG Unit must be upgradeable to Polysomnographic facilities for respiratory assessments and sleep staging.
3.3 | The Portable EEG Unit must be upgradeable to Sleep analysis software, including detailed analysis for respiratory disturbances and sleep stages.
3.4 | The EEG unit must be provided with the following software requirements:
3.4.1 | Spike and seizure detection
3.4.2 | EEG mapping
3.4.3 | Polysomnogram analysis
3.4.4 | Report software
3.5 | The system offered must be specifically based on the easy to use Operating System. The bidder must state the operating system of the unit offered.
3.6 | The system offered must be able to display a minimum of 8 user definable channels of DSA trendgraphs (Density Spectral Array) during recording and Review.
3.7 | The system offered must be able to display 64 channels plus one mark channel.
3.8 | An interfacing between the electrode junction box and computer is essential on the unit offered. (USB or latest equivalent)
3.9 | The system offered must provide for a standard networking capability. The bidder must state the network capability of the unit offered.
3.10 | Automatic photic stimulation with up to 30 steps must be provided on the unit offered.
3.11 | A 3D Voltage mapping feature with customizable views must be provided as standard on the unit offered.
3.12 | Data review during acquisition must be available on the unit offered.
3.13 | Simultaneous acquisition and review of the current recording in a configurable split screen for comparison purpose with freeze function must be provided on the unit offered.
3.14 | A file append feature that allows separate EEG’s to be combined into one file must be provided on the unit offered.
3.15 | A facility to copy any part of the waveform window by just dragging and dropping for comparison functionality must be provided on the unit offered.
3.16 | A facility to save EEG data to CD-R with integrated embedded review software for review on any PC without the need of additional software must be provided on the unit offered.
3.17 | A montage map displaying the electrode and montage positioning with print function must be provided on the unit offered.
3.18 | A high resolution 1200Dpi laser printer or more advanced must be provided on the unit offered.
3.19 | The headbox offered must be mountable and provide for a minimum of 38 simultaneously recorded inputs.
3.20 | The system offered must provide simultaneous inputs for at least: 25 EEG Inputs, 14 Bipolar Inputs, 3 Respiratory Inputs, 4 DC Inputs, 1 x SpO2 and 1 CO2 Inputs.
3.21 | The Input impedance must be 100 MΩ or greater on the unit offered.
3.22 | The Input leakage current must be less than 5nA on the unit offered.
3.23 | The Internal noise level must be less than 1.5 µVp-p (0.53 Hz to 60 Hz) on the unit offered.
3.24 | The CMRR must be 105 dB or greater at 50Hz on the unit offered.
3.25 | The low-cut hardware filter for non DC inputs must be 0.08 Hz or less on the unit offered.
3.26 | The High-cut hardware filter must be 300 Hz with a (-18 dB/oct) roll off or better on the unit offered.
3.27 | The A/D conversion must be a simultaneous sampling of all inputs at 16bits per channel on the unit offered.
3.28 | There must be a selectable sampling rate per input of 100Hz to 1000Hz on the unit offered.
3.29 | The Display Resolution must range from 1024 X 768 to 1600 x 1200 on the unit offered.
3.30 | Overwrite and page by page display modes must be provided on the unit offered.
3.31 | Waveform display position feature must be provided on the unit offered
3.32 | The waveform sweep speed must be from 5seconds to 5 minutes per page on the unit offered.
3.33 | Timing marks, timing scale and event marks must be provided on screen on the unit offered.
3.34 | An EEG input sensitivity of: 1,2,3,(2.5), 5,7,10,15,20,30,50,75,100,150,200 µVmm must be provided on the unit offered.
3.35 | A DC input sensitivity of: off, 10, 15, 20, 30, 50, 70, 100, 150, 200 mV/mm must be provided on the unit offered.
3.36 | The time constant must be variable in steps of; 0.001, 0.003, 0.03, 0.1, 0.3, 0.6, 1.0, 2.0 on the unit offered.
3.37 | The low-cut filter on the unit offered must be variable in steps of: 0.08, 0.16, 0.27, .053, 1.6, 5.3, 53, 159 Hz with a roll off of (-6 dB/ oct).
3.38 | The High-cut filter on the unit offered must be variable with the following standard options: 15, 30, 35, 60, 70, 120 Hz (-12 dB/oct) 50 Hz (RAPID), 3000 Hz (-18 dB/oct).
3.39 | The AC interference filter: at 50 or 60 Hz must have a rejection ratio of 1/25 or more on the unit offered.
3.40 | A calibration waveform on the unit offered must have the option for a Step of sine wave and variable voltages of: 2, 5, 10, 20, 50, 100, 200, 500, 1000 µV (x 1000 for DC input signals).
3.41 | An EEG artifact elimination filter must be available in acquisition and review modes on the unit offered.
3.42 | An impedance check function, on screen with values and at the headbox with simultaneous display using threshold method must be provided on the unit offered.
3.43 | The ability to program at least 36 sets of programmable montages combined with programmable individual amplifier settings must be provided on the unit offered.
3.44 | The reference electrode on the unit offered must be selectable between A1→A2, A1←A2, A1↔A2, A1+A2, VX, AV (with unsuitable electrode deletion function), Aav, Org, SD and OFF. The bidder to state the selecting method if not similar to the above.
3.45 | The following events on the unit offered must be automatically marked: Start stop Recording, montage change with montage name, patient, button press, photic stimulation pulse and frequency, hyperventilation with time stamp.
3.46 | The display resolution on the unit offered must support up to 1600 x 1280.
3.47 | The ability to display up to 64 simultaneous waveform channels and 1 mark channels must be provided on the unit offered.
3.48 | Functions to adjust the waveform colour, position, width, number of channels must be provided on the unit offered.
3.49 | Multiple paging modes such as page overwrite and page-by-page must be provided on the unit offered.
3.50 | A waveform Freeze function must be provided on the unit offered.
3.51 | The waveform sweep speed should be adjustable; 5, 10, 20, 30, 60 s/ page or 5 min/page on the unit offered.
3.52 | The displayed waveforms on the unit offered should include user selectable time markings of 0.1 and 1 second or both.
3.53 | Pre-programmable and real-time user event markings must be provided on the unit offered.
3.54 | EEG scales and cursor measurement scales for amplitude, time and frequency must be provided on the unit offered.
3.55 | The review program on the unit offered must allow adjustment of the following parameters montage, sensitivity, high–cut filter, time constant, reference electrode and display speed.
3.56 | Jump functions, to allow quick access to: Specified event, page by page, and specified time must be provided on the unit offered.
3.57 | Review display modes on the unit offered must include the following for productive EEG scoring: continuous, high speed, high speed with pause, manually page by page, manually second by seconds, and manually waveform centering.
4 | The following sensors and accessories must be included
4.1 | Oral and nasal airflow
4.2 | Respiratory effort sensors for chest and abdomen.
4.3 | Pulse oximetery sensor (finger type) Bidder to specify the technology
4.4 | Snoring sensor.
4.5 | Body Position sensor.
4.6 | Automatic CPAP Titration System with manual remote override ability and analogue signal out for connection to the recording system to monitor flow, CPAP pressure leak etc.
4.7 | Carrying case
5 | Conformity Compliance (Please attach certificate)
5.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
5.2 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
5.2.1 | Please provide unique ref number of the ISO 13485 ceriticate:
5.3 | Model on offer must comply with EU Medical Device Regulation (EU) 2017/745 or equvalent from SAHPRA recognized jurisdictions . Attach a copy of certification
5.3.1 | Please provide unique ref number of the ceriticate:
5.4 | Hazardous Substance Act: If this type of equipment / apparatus appears on the schedule of Hazardous substances issued by the Directorate: Health Technology of the Department of Health, a license in terms of the Act on Hazardous Substances (Act 15/1973) must be submitted with the bid / quotation documents. The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.5 | All electrical/electronic medical equipment must be licensed by Radiation Control, where applicable a copy of the license must be submitted The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.5.1 | Radiation Control License number:
5.6 | The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade (rubber to with 2 screws) type and it must be a minimum length of three (3) meters. NB the mains cable of the unit tendered for must be SABS colour coded.
5.7 | The equipment quoted for must be protected against electromagnetic interference IEC 60601-1-2
5.8 | The unit must comply with electrical safety and performance standard such as IEC 60601-2-21, attached certification
5.8.1 | Please provide unique ref number of the IEC ceriticate:
6 | COMPREHENSIVE MAINTENANCE
6.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) should be included in the unit price as per the SCC document
6.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
6.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
6.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
6.3 | The bidder to provide the OEM recommended service interval of the unit:
6.4 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
7 | General
7.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
7.2 | The unit on offer must be the latest technology.
7.2.1 | Initial Manufacturing year of Model:
RT2-09-001 | EEG Portable- Accessories, Consumables and Upgrades | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
RT2-09-001a | Oral and nasal airflow
RT2-09-001b | Respiratory effort sensors for chest and abdomen.
RT2-09-001c | Pulse oximetery sensor (finger type)
RT2-09-001d | Snoring sensor.
RT2-09-001e | Body Position sensor.
RT2-09-001f | Automatic CPAP Titration System with manual remote override ability and analogue signal out for connection to the recording system to monitor flow, CPAP pressure leak etc.
RT2-09-001g | Carrying case
RT2-09-001h | SLEEP ANALYSIS SOFTWARE (UPGRADE) -specifications below
RT2-09-001ha | Dedicated software for the analysis of sleep disordered breathing & periodic limb movements.
RT2-09-001hb | Display and analysis of sleep stages and respiratory data
RT2-09-001hc | Generation of comprehensive reports.
RT2-09-001hd | Re-formatting of individual channel filters and sensitivity.
RT2-09-001he | Re-formatting of montages.
RT2-09-001hf | Time linked trend graphs superimposed on waveform data
RT2-09-001hg | Automatic analysis data types must include the following: EEG1, EEG2, EOG1, EOG2, ECG, Chin, EMG, LEG, EMG, Airflow, Chest Effort, Abdomen Effort, SpO2, CPAP, BIPAP, Snoring, Body Position, Pulse Rate, Airflow2, ETCO2, R-R, pH.
RT2-09-001hh | Support a dual monitor and enhanced multitasking interface.
RT2-09-001hi | Must allow easy, fast, and accurate scoring and editing.
RT2-09-001hj | Graphical trending of all scanned and scored parameters
RT2-09-001hk | Monitor multiple recordings from a single station.
RT2-09-001hl | On-line scanning, scoring and editing
RT2-09-001hm | Separate adult and infant data analysis settings.
RT2-09-001hn | Single click scoring and editing.
RT2-09-001ho | Multiple synchronized time bases.
RT2-09-001hp | Navigate data by graphical interface, comment query, event query and scroll bar.
RT2-09-001hq | Customized event display.
RT2-09-001hr | View AHI any time in the study or by treatment level.
RT2-09-001hs | Automatic event association.
RT2-09-001ht | Customized workspaces – manual scored changes are stored separately for each scorer allowing QA comparisons and preserving manual changes even if the test is re-analyzed.
RT2-09-001hu | Automatic frame insertion into report from digital video.
RT2-09-001hv | Complete post-study remontaging and refiltering.
RT2-09-001hw | Score data over LAN, WAN or Internet.
RT2-09-001hx | Export and Import of European Data Format (EDF).
RT2-09-001hy | Dedicated Infant data analysis must support.
RT2-09-001hz | Active, Quiet, Transitional and Wake scoring
RT2-09-001haa | Infant staging, trending and graphics
RT2-09-001hab | Periodic breathing analysis
RT2-09-001hac | Ph analysis and reporting (requires PH probe).
RT2-09-001had | Complete infant database with correlations of all types of clinical events with infant staging.
RT2-09-001hae | Flexible and easy-to-use DC channel set up and calibration
RT2-09-001i | REPORTING FUNCTIONALITY (UPGRADE):
RT2-09-001ia | Auto text complete field searches that allow you to find the values or fields you want to insert to into your reports.
RT2-09-001ib | Auto pop-up field descriptions that give you easy to understand field descriptions
RT2-09-001ic | Customizable graphs
RT2-09-001id | Comparison reports give your lab epoch-to-epoch comparisons of the different scores in the lab
RT2-09-001ie | Create customized impression and recommendation templates.
RT2-09-001if | Seamless interface over the LAN, WAN, or Internet for remote interpretation from anywhere.
RT2-09-001ig | HTML report format compatible with most word processing software.
RT2-09-001ih | Direct compatibility with MicrosoftTM Internet Explore
RT2-09-001ii | Auto e-mail directly from the report Generator.
RT2-09-001j | Three re-usable EEG paediatric head caps of the following sizes: 1x 42cm to 46cm; 1x 38cm to 42cm; 1x 34cm to 38cm must be included and supplied with the unit offered at no extra cost to the final bid price. These head caps must include a complete set of re-usable electrodes and electrode board adaptor and these must also be supplied at no extra cost to the final bid pricel

## EEG, Desktop
RT2-09-002 | EEG Machine, Desktop | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
1 | DESCRIPTION
1.1 | An electronic monitoring device to measure and record electrical activity in the brain.
2 | Product Offered:
2.1 | Make:
2.2 | Model:
2.3 | Initial Manufacturing year of Model:
2.4 | Radiation Control License number:
3 | TECHNICAL SPECIFICATION
3.1 | The EEG unit complete with software for acquisition and review and compatible computer with necessary interface and printer as required.
3.2 | The EEG unit quoted for must be user friendly, it must comprise the latest technology and it must be a 32 Channel computerized EEG – PC based system with the following minimum spec: Core i7 CPU, 8GB of RAM, 256GB SSD or better, CD/DVD RW, 19" LCD TFT display, Keyboard, Mouse and UPS. All placed neatly on a trolley.
3.3 | The EEG Unit must be upgradeable to Polysomnographic facilities for respiratory assessments and sleep staging.
3.4 | The EEG Unit must upgradeable to Sleep analysis software, including detailed analysis for respiratory disturbances and sleep stages.
3.5 | The EEG unit must be provided with the following software requirements:
3.5.1 | Spike and seizure detection
3.5.2 | EEG mapping
3.5.3 | Polysomnogram analysis
3.5.4 | Report software
3.6 | The system offered must be specifically based on the easy to use Operating System. The bidder must state the operating system of the unit offered.
3.7 | The system offered must be able to display a minimum of 8 user definable channels of DSA trendgraphs (Density Spectral Array) during recording and Review.
3.8 | The system offered must be able to display 64 channels plus one mark channel.
3.9 | An interfacing between the electrode junction box and computer is essential on the unit offered. (USB or latest equivalent)
3.10 | The electrode junction box must be mounted on a EEG trolley or provided on a separate stand but both option must offer accessory basket.
3.11 | The system offered must provide for a standard networking capability. The bidder must state the network capability of the unit offered.
3.12 | Automatic photic stimulation with up to 30 steps must be provided on the unit offered.
3.13 | A 3D Voltage mapping feature with customizable views must be provided as standard on the unit offered.
3.14 | Data review during acquisition must be available on the unit offered.
3.15 | Simultaneous acquisition and review of the current recording in a configurable split screen for comparison purpose with freeze function must be provided on the unit offered.
3.16 | A file append feature that allows separate EEG’s to be combined into one file must be provided on the unit offered.
3.17 | A facility to copy any part of the waveform window by just dragging and dropping for comparison functionally must be provided on the unit offered.
3.18 | A facility to save EEG data to CD-R with integrated embedded review software for review on any PC without the need of additional software must be provided on the unit offered.
3.19 | A montage map displaying the electrode and montage positioning with print function must be provided on the unit offered.
3.20 | A high resolution 1200Dpi laser printer or more advanced must be provided on the unit offered.
3.21 | The headbox offered must provide for a minimum of 38 simultaneously recorded inputs.
3.22 | The system offered must provide simultaneous inputs for at least: 25 EEG Inputs, 14 Bipolar Inputs, 3 Respiratory Inputs, 4 DC Inputs, 1 x SpO2 and 1 CO2 Inputs.
3.23 | The Input impedance must be 100 MΩ or greater on the unit offered.
3.24 | The Input leakage current must be less than 5nA on the unit offered.
3.25 | The Internal noise level must be less than 1.5 µVp-p (0.53 Hz to 60 Hz) on the unit offered.
3.26 | The CMRR must be 105 dB or greater at 50Hz on the unit offered.
3.27 | The low-cut hardware filter for non DC inputs must be 0.08 Hz or less on the unit offered.
3.28 | The High-cut hardware filter must be 300 Hz with a (-18 dB/oct) roll off or better on the unit offered.
3.29 | The A/D conversion must be a simultaneous sampling of all inputs at 16bits per channel on the unit offered.
3.30 | There must be a selectable sampling rate per input of 100Hz to 1000Hz on the unit offered.
3.31 | The Display Resolution must range from 1024 X 768 to 1600 x 1200 on the unit offered.
3.32 | Overwrite and page by page display modes must be provided on the unit offered.
3.33 | Waveform display position feature must be provided on the unit offered
3.34 | The waveform sweep speed must be from 5seconds to 5 minutes per page on the unit offered.
3.35 | Timing marks. Timing scale and event marks must be provided on screen on the unit offered.
3.36 | An EEG input sensitivity of: 1,2,3,(2.5), 5,7,10,15,20,30,50,75,100,150,200 µVmm must be provided on the unit offered.
3.37 | A DC input sensitivity of: off, 10, 15, 20, 30, 50, 70, 100, 150, 200 mV/mm must be provided on the unit offered.
3.38 | The time constant must be variable in steps of; 0.001, 0.003, 0.03, 0.1, 0.3, 0.6, 1.0, 2.0 on the unit offered.
3.39 | The low-cut filter on the unit offered must be variable in steps of: 0.08, 0.16, 0.27, .053, 1.6, 5.3, 53, 159 Hz with a roll off of (-6 dB/ oct).
3.40 | The High-cut filter on the unit offered must be variable with the following standard options: 15, 30, 35, 60, 70, 120 Hz (-12 dB/oct) 50 Hz (RAPID), 300 Hz (-18 dB/oct).
3.41 | The AC interference filter: at 50 or 60 Hz must have a rejection ratio of 1/25 or more on the unit offered.
3.42 | A calibration waveform on the unit offered must have the option for a Step of sine wave and variable voltages of: 2, 5, 10, 20, 50, 100, 200, 500, 1000 µV (x 1000 for DC input signals).
3.43 | An ECG artifact elimination filter must be available in acquisition and review modes on the unit offered.
3.44 | An impedance check function, on screen with values and at the headbox with simultaneous display using threshold method must be provided on the unit offered.
3.45 | The ability to program at least 36 sets of programmable montages combined with programmable individual amplifier settings must be provided on the unit offered.
3.46 | The reference electrode on the unit offered must be selectable between A1→A2, A1←A2, A1↔A2, A1+A2, VX, AV (with unsuitable electrode deletion function), Aav, Org, SD and OFF.
3.47 | The following events on the unit offered must be automatically marked: Start stop Recording, montage change with montage name, patient, button press, photic stimulation pulse and frequency, hyperventilation with time stamp.
3.48 | The display resolution on the unit offered must support up to 1600 x 1280.
3.49 | The ability to display up to 64 simultaneous waveform channels and 1 mark channels must be provided on the unit offered.
3.50 | Functions to adjust the waveform colour, position, width, number of channels must be provided on the unit offered.
3.51 | Multiple paging modes such as page overwrite and page-by-page must be provided on the unit offered.
3.52 | A waveform Freeze function must be provided on the unit offered.
3.53 | The waveform sweep speed should be adjustable; 5, 10, 20, 30, 60 s/ page or 5 min/page on the unit offered.
3.54 | The displayed waveforms on the unit offered should include user selectable time markings of 0.1 and 1 seconds or both.
3.55 | Pre-programmable and real-time user event markings must be provided on the unit offered.
3.56 | EEG scales and cursor measurement scales for amplitude, time and frequency must be provided on the unit offered.
3.57 | The review program on the unit offered must allow adjustment of the following parameters montage, sensitivity, high–cut filter, time constant, reference electrode and display speed.
3.58 | Jump functions, to allow quick access to: Specified event, page by page, and specified time must be provided on the unit offered.
3.59 | Review display modes on the unit offered must include the following for productive EEG scoring: continuous, high speed, high speed with pause, manually page by page, manually second by seconds, and manually waveform centering.
4 | The following Sensors and Accessories must be included
4.1 | Oral and nasal airflow
4.2 | Respiratory effort sensors for chest and abdomen.
4.3 | Pulse oximetery sensor (finger type) Bidder to specify the technology.
4.4 | Snoring sensor.
4.5 | Body Position sensor.
4.6 | Automatic CPAP Titration System with manual remote override ability and analogue signal out for connection to the recording system to monitor flow, CPAP pressure leak etc.
4.7 | Mobile trolley
4.8 | At least 2KVA UPS
4.9 | Laser printer
4.10 | Printing paper for atleast 1000 EEG tests
5 | Conformity Compliance (Please attach certificate)
5.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
5.2 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
5.2.1 | Please provide unique ref number of the ISO 13485 ceriticate:
5.3 | Model on offer must comply with EU Medical Device Regulation (EU) 2017/745 or equvalent from SAHPRA recognized jurisdictions . Attach a copy of certification
5.3.1 | Please provide unique ref number of the ceriticate:
5.4 | Hazardous Substance Act: If this type of equipment / apparatus appears on the schedule of Hazardous substances issued by the Directorate: Health Technology of the Department of Health, a license in terms of the Act on Hazardous Substances (Act 15/1973) must be submitted with the bid / quotation documents. The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.5 | All electrical/electronic medical equipment must be licensed by Radiation Control, where applicable a copy of the license must be submitted The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.5.1 | Radiation Control License number:
5.6 | The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade (rubber to with 2 screws) type and it must be a minimum length of three (3) meters. NB the mains cable of the unit tendered for must be SABS colour coded.
5.7 | The equipment quoted for must be protected against electromagnetic interference IEC 60601-1-2
5.8 | The unit must comply with electrical safety and performance standard such as IEC 60601-2-21, attached certification
5.8.1 | Please provide unique ref number of the IEC ceriticate:
6 | COMPREHENSIVE MAINTENANCE
6.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) should be included in the unit price as per the SCC document
6.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
6.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
6.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
6.3 | The bidder to provide the OEM recommended service interval of the unit:
6.4 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
7 | General
7.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
7.2 | The unit on offer must be the latest technology.
7.2.1 | Initial Manufacturing year of Model:
RT2-09-002 | EEG Machine, Desktop - Accessories, Consumables and Upgrades | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
RT2-09-002a | Oral and nasal airflow
RT2-09-002b | Respiratory effort sensors for chest and abdomen.
RT2-09-002c | Pulse oximetery sensor (finger type) Bidder to specify the technology.
RT2-09-002d | Snoring sensor.
RT2-09-002e | Body Position sensor.
RT2-09-002f | Automatic CPAP Titration System with manual remote override ability and analogue signal out for connection to the recording system to monitor flow, CPAP pressure leak etc.
RT2-09-002g | Mobile trolley
RT2-09-002h | At least 2KVA UPS
RT2-09-002i | Laser printer
RT2-09-002j | Printing paper for atleast 1000 EEG tests
RT2-09-002k | SLEEP ANALYSIS SOFTWARE (UPGRADE)– Specification below
RT2-09-002ka | Dedicated software for the analysis of sleep disordered breathing & periodic limb movements.
RT2-09-002kb | Display and analysis of sleep stages and respiratory data
RT2-09-002kc | Generation of comprehensive reports.
RT2-09-002kd | Re-formatting of individual channel filters and sensitivity.
RT2-09-002ke | Re-formatting of montages.
RT2-09-002kf | Time linked trend graphs superimposed on waveform data
RT2-09-002kg | Automatic analysis data types must include the following: EEG1, EEG2, EOG1, EOG2, ECG, Chin, EMG, LEG, EMG, Airflow, Chest Effort, Abdomen Effort, SpO2, CPAP, BIPAP, Snoring, Body Position, Pulse Rate, Airflow2, ETCO2, R-R, pH.
RT2-09-002kh | Support a dual monitor and enhanced multitasking interface.
RT2-09-002ki | Must allow easy, fast, and accurate scoring and editing.
RT2-09-002kj | Graphical trending of all scanned and scored parameters
RT2-09-002kk | Monitor multiple recordings from a single station.
RT2-09-002kl | On-line scanning, scoring and editing
RT2-09-002km | Separate adult and infant data analysis settings.
RT2-09-002kn | Single click scoring and editing.
RT2-09-002ko | Multiple synchronized time bases.
RT2-09-002kp | Navigate data by graphical interface, comment query, event query and scroll bar.
RT2-09-002kq | Customized event display.
RT2-09-002kr | View AHI any time in the study or by treatment level.
RT2-09-002ks | Automatic event association.
RT2-09-002kt | Customized workspaces – manual scored changes are stored separately for each scorer allowing QA comparisons and preserving manual changes even if the test is re-analyzed.
RT2-09-002ku | Automatic frame insertion into report from digital video.
RT2-09-002kv | Complete post-study remontaging and refiltering.
RT2-09-002kw | Score data over LAN, WAN or Internet.
RT2-09-002kx | Export and Import of European Data Format (EDF).
RT2-09-002ky | Dedicated Infant data analysis must support.
RT2-09-002kz | Active, Quiet, Transitional and Wake scoring
RT2-09-002kaa | Infant staging, trending and graphics
RT2-09-002kab | Periodic breathing analysis
RT2-09-002kac | Ph analysis and reporting (requires PH probe).
RT2-09-002kad | Complete infant database with correlations of all types of clinical events with infant staging.
RT2-09-002kae | Flexible and easy-to-use DC channel set up and calibration
RT2-09-002l | REPORTING FUNCTIONALITY (UPGRADE):
RT2-09-002la | Auto text complete field searches that allow you to find the values or fields you want to insert to into your reports.
RT2-09-002lb | Auto pop-up field descriptions that give you easy to understand field descriptions
RT2-09-002lc | Customizable graphs
RT2-09-002ld | Comparison reports give your lab epoch-to-epoch comparisons of the different scores in the lab
RT2-09-002le | Create customized impression and recommendation templates.
RT2-09-002lf | Seamless interface over the LAN, WAN, or Internet for remote interpretation from anywhere.
RT2-09-002lg | HTML report format compatible with most word processing software.
RT2-09-002lh | Direct compatibility with MicrosoftTM Internet Explore
RT2-09-002li | Auto e-mail directly from the report Generator.
RT2-09-002lj | Three re-usable EEG paediatric head caps of the following sizes: 1x 42cm to 46cm; 1x 38cm to 42cm; 1x 34cm to 38cm must be included and supplied with the unit offered at no extra cost to the final bid price. These head caps must include a complete set of re-usable electrodes and electrode board adaptor and these must also be supplied at no extra cost to the final bid pricel

## Ultrasonic Surgical Aspirator
RT2-09-003 | Ultrasonic Surgical Aspirator | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
1 | DESCRIPTION
1.1 | Medical device that utilizes high frequency ultrasonic vibrations to fragment, emulsify and aspirate tissues, blood and fluids during surgical procedures.
2 | Product Offered:
2.1 | Make:
2.2 | Model:
2.3 | Initial Manufacturing year of Model:
2.4 | Radiation Control License number:
3 | TECHNICAL SPECIFICATION
3.1 | The Surgical Aspirator must use the latest technology (magnetostrictive or better), incorporating the CEM Nose cone or equivalent and suitable for use in various surgical procedures.
3.2 | The unit must be supplied with a trolley and it should have at least 4 antistatic castors with at least 2 equipped with locking system/brake.
3.3 | It must be suitable for the following applications on soft tissues:
3.3.1 | Fragmentation
3.3.2 | Emulsification
3.3.3 | Aspiration
3.4 | It must be suitable for the following applications on hard tissues:
3.4.1 | Fragmentation
3.4.2 | Emulsification
3.4.3 | Aspiration
3.5 | The Surgical Aspirator must be suitable for the following theatre procedures:
3.5.1 | Neurosurgery
3.5.2 | Gastro-intestinal and affiliated organs surgery
3.5.3 | Urology
3.5.4 | Plastic & reconstructive surgery
3.5.5 | General surgery
3.5.6 | Orthopaedic surgery
3.5.7 | Gynaecological surgery
3.5.8 | Thoracic surgery
3.5.9 | Laparoscopic surgery
3.5.10 | Thorascopic surgery
3.6 | Both audible and visual alarms must be available on the following:
3.6.1 | Occlusion
3.6.2 | Hand piece disconnections
3.6.3 | Power failure
3.6.4 | Safety features
3.7 | The Surgical Aspirator must have a built-in cooling fan.
4 | STANDARD ACCESSORIES
4.1 | It must be supplied with a total of 12 tips comprising of:
4.1.1 | Shear tip
4.1.2 | Micro tip
4.1.3 | Micro tip plus
4.2 | Suction cannister 2L x2
4.3 | Suction tubing
4.4 | Internal vacuum filter
4.5 | External vaccum filter
4.6 | Cleaning brush
4.7 | Foot switch
4.8 | Straight Handpiece
4.9 | Angle Handpiece
5 | Conformity Compliance (Please attach certificate)
5.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
5.2 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
5.2.1 | Please provide unique ref number of the ISO 13485 ceriticate:
5.3 | Model on offer must comply with EU Medical Device Regulation (EU) 2017/745 or equvalent from SAHPRA recognized jurisdictions . Attach a copy of certification
5.3.1 | Please provide unique ref number of the ceriticate:
5.4 | Hazardous Substance Act: If this type of equipment / apparatus appears on the schedule of Hazardous substances issued by the Directorate: Health Technology of the Department of Health, a license in terms of the Act on Hazardous Substances (Act 15/1973) must be submitted with the bid / quotation documents. The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.5 | All electrical/electronic medical equipment must be licensed by Radiation Control, where applicable a copy of the license must be submitted The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.5.1 | Radiation Control License number:
5.6 | The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade (rubber to with 2 screws) type and it must be a minimum length of three (3) meters. NB the mains cable of the unit tendered for must be SABS colour coded.
5.7 | The equipment quoted for must be protected against electromagnetic interference IEC 60601-1-2
5.8 | The unit must comply with electrical safety and performance standard such as IEC 60601-2-21, attached certification
5.8.1 | Please provide unique ref number of the IEC ceriticate:
6 | COMPREHENSIVE MAINTENANCE
6.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) should be included in the unit price as per the SCC document
6.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
6.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
6.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
6.3 | The bidder to provide the OEM recommended service interval of the unit:
6.4 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
7 | General
7.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years If yes, please provide further details of the recall and how was it addressed
7.2 | The unit on offer must be the latest technology.
7.2.1 | Initial Manufacturing year of Model:
RT2-09-003 | Ultrasonic Surgical Aspirator- Accessories and Conusmables | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
RT2-09-003a | Shear tip
RT2-09-003b | Micro tip
RT2-09-003c | Micro tip plus
RT2-09-003d | Suction cannister 2L x2
RT2-09-003e | Suction tubing
RT2-09-003f | Intenal vacuum filter
RT2-09-003g | External vaccum filter
RT2-09-003h | Cleaning brush
RT2-09-003i | Foot switch
RT2-09-003j | Straight Handpiece
RT2-09-003k | Angle Handpiece

## Neuronavigation system
RT2-09-004 | Neuronavigation System | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
1 | DESCRIPTION
1.1 | An intra-operative brain mapping navigation system specifically designed for neurosurgery and not a general navigation system
2 | Product Offered:
2.1 | Make:
2.2 | Model:
2.3 | Initial Manufacturing year of Model:
2.4 | Radiation Control License number:
3 | TECHNICAL SPECIFICATION
3.1 | The system must offer the following:
3.1.1 | High performance Image-Guided-Surgery station.
3.1.2 | Must be integrated in a mobile display cart with separate mobile camera cart for flexible positioning of the camera
3.1.3 | Must feature advanced optical, wireless passive marker tracking technology for pre-operative planning and surgical visualization.
3.1.4 | Must offer optical and Electromagnetic (EM) tracking capabilities, and integrates with external devices like microscopes, ultrasound, and a broad array of instruments.
3.2 | Must be supplied with a UPS that can power the unit for at least 2 hours.
3.3 | The neuronavigation system must be used for all types of emergency and scheduled cranial surgeries, Deep Brain Stimulation (DBS) surgeries, microscopic procedures and scheduled spinal surgeries.
3.4 | The neuronavigation system must have the ability to relate the position of a real surgical instrument in the surgeon's hand or the microscope's focal point to the location of the imaged pathology, updated in real time in an integrated operating room, highlights the modern version of this set of technologies.
3.5 | The unit must be able to visualize the scenario for surgery in a 3-D model of manipulable computer data. In this way the physician can "practice and check" the surgery, try alternative approaches, assess possible difficulties, etc., before the real surgery takes place.
3.6 | The System must include the following:
3.6.1 | Cranial Surgery software and hardware.
3.6.2 | Complete package of DBS software and hardware (enhanced target localization by audiovisual integration of physiological signal, simultaneous display of recording position, advanced 3-D modeling capabilities and display of three orthogonal views).
3.6.3 | Image merging software to merge CT, MRI and Microscope interface.
3.6.4 | 3-D Navigation and volume rendering software for brain and spine
3.6.5 | Spine surgery software and hardware allowing seamless interface to a variety of C-arm Images.
3.7 | The System must be supplied with the following features:
3.7.1 | Neuronavigation systems must offer accurate planning, targeting, and monitoring solutions for Transcranial Magnetic Stimulation (TMS) and/or Transcranial Electrical Stimulation (tES) studies, in addition to 3-D localisation of Neuroimaging systems, such as Electroencephalogram (EEG) and Near-Infrared Spectroscopy (NIRS).
3.7.2 | Neuroimaging devices to offer researchers an ability to view the structure, function and pharmacology of the brain.
3.7.3 | Neurosensory devices for use in neuroscience research.
3.7.4 | Neuromodulation devices to enhance or suppress activities of the brain or nervous system in the research of brain function and disease.
3.7.5 | Monitor with the following minimum specifications:
3.7.5.1 | Built in touch screen monitor 24 inch HD with minimum 1920 x 1200 pixels resolution.
3.7.5.2 | Must accept DICOM format Images from intra-operative MRI, intra-operative CT, and C-Arm.
3.7.5.3 | Must accept standard DICOM format Images from MRI and CT Scanning Departments via CD/USB enabling surgeon to plan surgery in advance on main navigation system.
3.7.5.4 | Surgeon must have control of the navigation system through touch screen and patient frame without compromising the sterility
3.8 | Mandatory accessories to be supplied with the System:
3.8.1 | Small, medium and large passive trackers
3.8.2 | Small, medium and large mounts
3.8.3 | Mount driver
3.8.4 | Sterilization tray
3.8.5 | Cranial biopsy instrument to be supplied with the following:
3.8.5.1 | Articulated arm
3.8.5.2 | Mount, double star adapter mayfield
3.8.5.3 | Biopsy guide
3.8.5.4 | Probe
3.8 | Microscope integration interface only with the following features:
3.8.1 | Display of the microscope focal point on anatomical images
3.8.2 | Display of Region of interest (ROI) and Navigation symbols in oculars
3.8.3 | Display of navigation monitor in oculars 1x tracking hardware (cables & bracket)
3.8.4 | Tracking interface hardware
3.9 | The system must be capable of networking from modular operation theatre to CT Room and Neurosurgery ward
3.10 | Planning station must have a separate, latest and advanced computer for planning of navigated surgical procedures with the following minimum specifications:
3.10.1 | ≥8GB Ram, ≥256 GB SSD, ≥17 inch LED screen, the LED screen and Core i7 CPU must be integrated in one casing, software for processing all captured images, a cordless keyboard and mouse, with latest version of the operating system.
3.10.2 | Preloaded planning software for brain and spine.
4 | Standard Accessories and Consumables
4.1 | Small, medium and large passive trackers
4.2 | Small, medium and large mounts
4.3 | Mount driver
4.4 | Sterilization tray
4.5 | Articulated arm
4.6 | Mount, double star adapter mayfield
4.7 | Biopsy guide
4.8 | Probe
5 | Conformity Compliance (Please attach certificate)
5.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
5.2 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
5.2.1 | Please provide unique ref number of the ISO 13485 ceriticate:
5.3 | Model on offer must comply with EU Medical Device Regulation (EU) 2017/745 or equvalent from SAHPRA recognized jurisdictions . Attach a copy of certification
5.3.1 | Please provide unique ref number of the ceriticate:
5.4 | Hazardous Substance Act: If this type of equipment / apparatus appears on the schedule of Hazardous substances issued by the Directorate: Health Technology of the Department of Health, a license in terms of the Act on Hazardous Substances (Act 15/1973) must be submitted with the bid / quotation documents. The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.5 | All electrical/electronic medical equipment must be licensed by Radiation Control, where applicable a copy of the license must be submitted The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.5.1 | Radiation Control License number:
5.6 | The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade (rubber to with 2 screws) type and it must be a minimum length of three (3) meters. NB the mains cable of the unit tendered for must be SABS colour coded.
5.7 | The equipment quoted for must be protected against electromagnetic interference IEC 60601-1-2
5.8 | The unit must comply with electrical safety and performance standard such as IEC 60601-2-21, attached certification
5.8.1 | Please provide unique ref number of the IEC ceriticate:
6 | COMPREHENSIVE MAINTENANCE
6.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) should be included in the unit price as per the SCC document
6.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
6.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
6.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
6.3 | The bidder to provide the OEM recommended service interval of the unit:
6.4 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
7 | General
7.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
7.2 | The unit on offer must be the latest technology.
7.2.1 | Initial Manufacturing year of Model:
RT2-09-004 | Neuronavigation System- Accessories and Conusmables | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
RT2-09-004a | Small, medium and large passive trackers
RT2-09-004b | Small, medium and large mounts
RT2-09-004c | Mount driver
RT2-09-004d | Sterilization tray
RT2-09-004e | Articulated arm
RT2-09-004f | Mount, double star adapter mayfield
RT2-09-004h | Biopsy guide
RT2-09-004i | Probe

## Intraoperative Neuro Monitor
RT2-09-005 | Intraoperative Neurophysiological Monitor | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
1. | DESCRIPTION
1.1 | Medical device used during surgery to monitor the nervous system's eletrical activity in real-time.
2. | Product Offered
2.1 | Make:
2.2 | Model:
2.3 | Initial Manufacturing year of Model:
2.4 | Radiation Control License number:
3. | Technical Specifications
3.1 | The specifications requires a complete Intraoperative Neurophysiological monitor for monitoring nervous system during surgery.
3.2 | The Intraoperative Neurophysiological monitor must be usable for the following surgeries and the accessories and consumable for these surgeries must be supplied and included in the unit price :
3.2.1 | Spinal Surgery (all spinal cords cases, Corpectomy, Vertebrectomy etc.)
3.2.2 | Brain Surgery (not limited to):
3.2.2.1 | Sub-sensory evoked potentials
3.2.2.2 | Mapping for motor-cortex
3.2.2.3 | Motor-sensory evoked potentials (Brainstem auditory evoked potentials (BAEP),
3.2.2.4 | Visual evoked potentials (VEP)
3.2.2.5 | Somatosensory-evoked potentials (SSEP) etc
3.2.3 | Ear surgery and other surgeries involving arteries: (method in cerebellopontine angle surgeries and acoustic neuroma surgeries etc.)
3.3 | D wave monitoring use direct spinal cord stimulation modality
3.4 | The unit must be designed such that it provides the surgeon with immediate feedback and warning before permanent nerve injury occurs.
3.5 | The monitor must a have mechanism to measure the conduction of sensation above and below the area of surgery.
3.6 | The monitor must also have mechanisms to detect injuries to individual nerve roots such as:
3.6.1 | Electromyography (EMG)
3.6.2 | Transcranial electrical stimulation (TES)
3.6.3 | Brain auditory evoked potential (BAEP) etc.
3.7 | The monitor must be equipped with a loudspeaker that allows immediate audio feedback if the spinal cord or nerve roots are damaged while drilling into the bone, placing hardware in the spine, or decompressing the spine.
3.8 | The system should be equipped with a reliable low-noise universal neurophysiologic amplifier allowing high-quality acquisition of EMG, EEG, EP and other modalities.
3.9 | The system must be equipped with acquisition and stimulator pods to attach recording and stimulating electrodes with at least 5m cables.
3.10 | The Neuro-IOM system must be delivered with the preset templates ensuring all main modalities include the following:
3.10.1 | Direct nerve stimulation
3.10.2 | Free-run EMG
3.10.3 | Somatosensory evoked potentials (SEP)
3.10.4 | Motor evoked potentials (MEP)
3.10.5 | Auditory evoked potentials (AEP)
3.10.6 | Visual evoked potentials (VEP)
3.10.7 | Electro Encephalogram (EEG)
3.10.8 | Direct cortical stimulation
3.10.9 | Train-of-four stimulation (TOF)
3.10.10 | Flexible configuration to ensure any monitoring needs
3.10.11 | Handheld surgeon-controlled probe for direct stimulation of nerves and roots etc.
3.11 | The unit must be supplied with a trolley. The trolley must have the following features:
3.11.1 | Designed to house the complete monitor.
3.11.2 | Must have an power supply adapter which connects components of the system.
3.11.3 | The castors of the trolley should not be less than 150mm in diameter.
3.11.4 | All four castors must swivel and the minimum of two must be lockable.
3.11.5 | Must be manufactured from Stainless Steel material that is suitable for use in the operating theatre environment.
4 | The monitor must be supplied with the following minimum accessories and consumables:
4.1 | Subdermal needle electrodes x 200 packets of 3 needles
4.2 | Disposable corkscrew electrodes x 100
4.3 | Disposable adhesive electrodes x 50
4.4 | EMG endotracheal tubes x 10
4.5 | Monopolar probes/electrodes for stimulation x 10
4.6 | Bipolar probes/electrodes for stimulation x 10
4.7 | Stimulus dissecting instruments x 5
4.8 | Action potentials electrodes x 10
4.9 | Cortical and subdural electrodes x 10
5 | Conformity Compliance (Please attach certificate)
5.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
5.2 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
5.2.1 | Please provide unique ref number of the ISO 13485 ceriticate:
5.3 | Model on offer must comply with EU Medical Device Regulation (EU) 2017/745 or equvalent from SAHPRA recognized jurisdictions . Attach a copy of certification
5.3.1 | Please provide unique ref number of the ceriticate:
5.4 | Hazardous Substance Act: If this type of equipment / apparatus appears on the schedule of Hazardous substances issued by the Directorate: Health Technology of the Department of Health, a license in terms of the Act on Hazardous Substances (Act 15/1973) must be submitted with the bid / quotation documents. The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.5 | All electrical/electronic medical equipment must be licensed by Radiation Control, where applicable a copy of the license must be submitted The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.5.1 | Radiation Control License number:
5.6 | The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade (rubber to with 2 screws) type and it must be a minimum length of three (3) meters. NB the mains cable of the unit tendered for must be SABS colour coded.
5.7 | The equipment quoted for must be protected against electromagnetic interference IEC 60601-1-2
5.8 | The unit must comply with electrical safety and performance standard such as IEC 60601-2-21, attached certification
5.8.1 | Please provide unique ref number of the IEC ceriticate:
6 | COMPREHENSIVE MAINTENANCE
6.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) should be included in the unit price as per the SCC document
6.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
6.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
6.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
6.3 | The bidder to provide the OEM recommended service interval of the unit:
6.4 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
7 | General
7.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
7.2 | The unit on offer must be the latest technology.
7.2.1 | Initial Manufacturing year of Model:
RT2-09-005 | Intraoperative Neurophysiological Monitor- Accessories, Consumables and Upgrades | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
RT2-09-005a | Subdermal needle electrodes x 200 packets of 3 needles
RT2-09-005b | Disposable corkscrew electrodes x 100
RT2-09-005c | Disposable adhesive electrodes x 50
RT2-09-005d | EMG endotracheal tubes x 10
RT2-09-005e | Monopolar probes/electrodes for stimulation x 10
RT2-09-005f | Bipolar probes/electrodes for stimulation x 10
RT2-09-005g | Stimulus dissecting instruments x 5
RT2-09-005h | Action potentials electrodes x 10
RT2-09-005i | Cortical and subdural electrodes x 10

## Neuro Drill
RT2-09-006 | Neuro Drill | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
1 | Description
1.1 | A specialized medical device used in neurosurgery to create precise openings in the skull or spine, facilitating access to the underlying neutral structures.
2 | Product Offered
2.1 | Make:
2.2 | Model:
2.3 | Initial Manufacturing year of Model:
2.4 | Radiation Control License number:
3 | Technical Specification
3.1 | The high power drill must be used for both Nuero (transsphenoidal surgery, cranial surgeries etc.) and spine operations.
3.2 | The drill must offer reliable performance and unique surgical features for the surgeon.
3.3 | The unit offered must be able to perform the following procedures :
3.3.1 | Incision/cutting, removal, drilling, and sawing of soft and hard tissue, bone, and biomaterials in Neurosurgical (Cranial and Craniofacial including Craniotomy).
3.3.2 | Ear, Nose and Throat (ENT), Maxillofacial, Orthopaedic, Arthroscopic, Spinal, Sternotomy, and general surgical procedures.
3.4 | The craniotomy attachments must consist of a combination of a solid heavy-duty craniotomy base which can be coupled with different dura guards.
3.5 | The following dura guards or equivalent must be supplied together with the unit and cost be included in the bid price.
3.5.1 | Fixed dura guard (pediatrics and neonates)
3.5.2 | Standard dura guard
3.5.3 | Long dura guard
3.5.4 | Turnable dura guard
3.5.5 | Craniotomy attachment for multifunctional piece
3.6 | The attachments must be designed to align the surgical tools with the motor to provide support and stability during various surgical procedures.
3.7 | The high speed motor must offer reduced dimensions for improved efficiency in the working environment with the following features:
3.7.1 | An increase in power and torque but also reduced thermal warming
3.7.2 | A tougher, more sturdy cable connection to better withstand the daily routine in operating rooms and sterile processing departments.
3.8 | The Electro control unit must be supplied with integrated irrigation pump
4 | The Neuro drill must be supplied with the standard accessories following:
4.1 | Console/Electro control unit
4.2 | Electro Foot pedal
4.3 | High speed motor with motor cable (minimum 80.000 rpm) and at least 2.5 or above motor cable that is autoclavable)
4.4 | Electro Perforator with coupling
4.5 | Electro Craniotome and multifunction hand piece (2 ring)
4.6 | Disposable irrigation tube set for Irrigation pump and adapters
4.7 | Disposable spray Nozzles
4.8 | Holder for ring burs
4.9 | Holder for craniotome attachments
4.10 | Electro low speed with Intra coupling
4.11 | Holder for saw blades
4.12 | Trolleys with drawers, IV pole for irrigation pump and Electro control unit.
4.13 | The drill must be supplied with the following accessories:
4.13.1 | Standard attachment for neuro and with nozzles
4.13.2 | High power attachment for spine must also come with the nozzle
4.13.3 | Both the micro saws i.e. electro micro surgical saw and electro micro reciprocating saw. Each saw must come standard with the nozzles
4.13.4 | Standard Electro surgical saw and Electro reciprocating saw. Each saw must come standard with the nozzles
4.13.5 | Sterilization Trays and Cases tray where all accessories can be placed for cleaning
4.14.4 | The following High speed hand pieces including the spray nozzle for spine must supplied with the unit
4.14.1 | Length (cm) ¦¦ Quantity:
4.14.2 | 7 cm ¦¦ 3 units
4.14.3 | 10/12/13cm ¦¦ 3 units
4.14.4 | 17 cm ¦¦ 3 units
4.14.5 | Burs for Micro hand pieces
4.14.5.1 | Description ¦¦ Diameter ¦¦ Quantity:
4.14.5.2 | Rosen Burrs ¦¦ 1.0 mm ¦¦ 3 units
4.14.5.3 | Rosen Burrs ¦¦ 3.0 mm ¦¦ 3 units
4.14.5.4 | Rosen Burrs ¦¦ 5.0 mm ¦¦ 3 units
4.14.5.5 | Rosen Burrs(Soft Cut) ¦¦ 4.0 mm ¦¦ 3 units
4.14.5.6 | Diamond Burrs ¦¦ 0.6 mm ¦¦ 3 units
4.14.5.7 | Diamond Burrs ¦¦ 1.0 mm ¦¦ 3 units
4.14.5.8 | Diamond Burrs, coarse ¦¦ 3.0 mm ¦¦ 3 units
4.14.5.9 | Diamond Burrs, coarse ¦¦ 5.0 mm ¦¦ 3 units
4.14.5.10 | Neuro cutter ¦¦ 2.0 mm ¦¦ 3 units
4.14.5.11 | Pin Cutter ¦¦ 1.0 mm ¦¦ 3 units
4.14.5.12 | Twist drills ¦¦ 1.5 mm ¦¦ 3 units
4.14.5.13 | Twist drills ¦¦ 2.0 mm ¦¦ 3 units
5 | Conformity Compliance (Please attach certificate)
5.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
5.2 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
5.2.1 | Please provide unique ref number of the ISO 13485 ceriticate:
5.3 | Model on offer must comply with EU Medical Device Regulation (EU) 2017/745 or equvalent from SAHPRA recognized jurisdictions . Attach a copy of certification
5.3.1 | Please provide unique ref number of the ceriticate:
5.4 | Hazardous Substance Act: If this type of equipment / apparatus appears on the schedule of Hazardous substances issued by the Directorate: Health Technology of the Department of Health, a license in terms of the Act on Hazardous Substances (Act 15/1973) must be submitted with the bid / quotation documents. The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.5 | All electrical/electronic medical equipment must be licensed by Radiation Control, where applicable a copy of the license must be submitted The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.5.1 | Radiation Control License number:
5.6 | The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade (rubber to with 2 screws) type and it must be a minimum length of three (3) meters. NB the mains cable of the unit tendered for must be SABS colour coded.
5.7 | The equipment quoted for must be protected against electromagnetic interference IEC 60601-1-2
5.8 | The unit must comply with electrical safety and performance standard such as IEC 60601-2-21, attached certification
5.8.1 | Please provide unique ref number of the IEC ceriticate:
6 | COMPREHENSIVE MAINTENANCE
6.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) should be included in the unit price as per the SCC document
6.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
6.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
6.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
6.3 | The bidder to provide the OEM recommended service interval of the unit:
6.4 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
7 | General
7.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
7.2 | The unit on offer must be the latest technology.
7.2.1 | Initial Manufacturing year of Model:
RT2-09-006 | Neuro Drill- Accessories, Consumables and Upgrades | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
RT2-09-006a | Console/Electro control unit
RT2-09-006b | Electro Foot pedal
RT2-09-006c | High speed motor with motor cable (minimum 80.000 rpm) and at least 2.5 or above motor cable that is autoclavable)
RT2-09-006d | Electro Perforator with coupling
RT2-09-006e | Electro Craniotome and multifunction hand piece (2 ring)
RT2-09-006f | Disposable irrigation tube set for Irrigation pump and adapters
RT2-09-006g | Disposable spray Nozzles
RT2-09-006h | Holder for ring burs
RT2-09-006i | Holder for craniotome attachments
RT2-09-006j | Electro low speed with Intra coupling
RT2-09-006k | Holder for saw blades
RT2-09-006l | Trolleys with drawers, IV pole for irrigation pump and Electro control unit.
RT2-09-006m | The drill must be supplied with the following accessories:
RT2-09-006ma | Standard attachment for neuro and with nozzles
RT2-09-006mb | High power attachment for spine must also come with the nozzle
RT2-09-006mc | Both the micro saws i.e. electro micro surgical saw and electro micro reciprocating saw. Each saw must come standard with the nozzles
RT2-09-006md | Standard Electro surgical saw and Electro reciprocating saw. Each saw must come standard with the nozzles
RT2-09-006me | Sterilization Trays and Cases tray where all accessories can be placed for cleaning
RT2-09-006n | The following High speed hand pieces including the spray nozzle for spine must supplied with the unit
RT2-09-006n | Length (cm) ¦¦ Quantity:
RT2-09-006na | 7 cm ¦¦ 3 units
RT2-09-006nb | 10/12/13cm ¦¦ 3 units
RT2-09-006nc | 17 cm ¦¦ 3 units
RT2-09-006o | Burs for Micro hand pieces
RT2-09-006o | Description ¦¦ Diameter ¦¦ Quantity:
RT2-09-006oa | Rosen Burrs ¦¦ 1.0 mm ¦¦ 3 units
RT2-09-006ob | Rosen Burrs ¦¦ 3.0 mm ¦¦ 3 units
RT2-09-006oc | Rosen Burrs ¦¦ 5.0 mm ¦¦ 3 units
RT2-09-006od | Rosen Burrs(Soft Cut) ¦¦ 4.0 mm ¦¦ 3 units
RT2-09-006oe | Diamond Burrs ¦¦ 0.6 mm ¦¦ 3 units
RT2-09-006of | Diamond Burrs ¦¦ 1.0 mm ¦¦ 3 units
RT2-09-006og | Diamond Burrs, coarse ¦¦ 3.0 mm ¦¦ 3 units
RT2-09-006oh | Diamond Burrs, coarse ¦¦ 5.0 mm ¦¦ 3 units
RT2-09-006oi | Neuro cutter ¦¦ 2.0 mm ¦¦ 3 units
RT2-09-006oj | Pin Cutter ¦¦ 1.0 mm ¦¦ 3 units
RT2-09-006ok | Twist drills ¦¦ 1.5 mm ¦¦ 3 units
RT2-09-006ol | Twist drills ¦¦ 2.0 mm ¦¦ 3 units

## AEEG
RT2-09-007 | Amplitude Electroencephlograph | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
1 | Description
1.1 | An electronic monitoring device to measure and record electrical activity of the brain through electrodes placed on the scalp recording voltage fluctuation.
2 | Product Offered
2.1 | Make:
2.2 | Model:
2.3 | Initial Manufacturing year of Model:
2.4 | Radiation Control License number:
3 | Technical Specification
3.1 | The monitor must have the ability to provide real-time EEG from both sides of the brain.
3.2 | The monitor must provide the following:
3.2.2 | Real time EEG
3.2.3 | Continuous measurement of impedance in 1, 2, and up to a minimum of 3 channel configurations
3.3 | Software applications must provide data and alerts at the bedside, indicating to clinicians a change in cerebral function
3.4 | The monitor must provide enhanced signal display for easier trace interpretation
3.5 | Cerebral Function Monitoring (CFM) must be possible
3.6 | An additional software feature for CFM sight that provides enhanced signal display for easier trace interpretation must be made available
3.7 | The AEEG must track and record brain wave patterns.
3.8 | The files stored in the monitor must be able to be retrieved, reviewed and shared.
3.9 | The electrodes which must be included in the bid price must be able to analyze the electrical impulses in the brain and send signals to a computer that records the results
3.10 | The unit must be able to detect and investigate epilepsy.
3.11 | The unit must have features that are ideal for teaching and research.
3.12 | Automatic editing of data must be easily done on the system.
3.13 | The software must be able to place all stimulation parameters and electrodes directly into the AEEG recording for archiving.
3.14 | It must be possible to review and check the trend in the past 24 hours.
3.15 | Must be integrated in a mobile display cart with separate mobile camera cart for flexible positioning of the camera
3.16 | The monitor must use only electrodes and be able to pick up a seizure
3.17 | It must be possible on the AEEG to edit/add information on an AEEG strip
3.18 | Personal details must be easily stored with the baby's AEEG strip
4 | STANDARD ACCESSORIES
4.1 | AEEG Electrodes
4.2 | Mobile cart
4.3 | Paper Rolls
5 | Conformity Compliance (Please attach certificate)
5.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
5.2 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
5.2.1 | Please provide unique ref number of the ISO 13485 ceriticate:
5.3 | Model on offer must comply with EU Medical Device Regulation (EU) 2017/745 or equvalent from SAHPRA recognized jurisdictions . Attach a copy of certification
5.3.1 | Please provide unique ref number of the ceriticate:
5.4 | Hazardous Substance Act: If this type of equipment / apparatus appears on the schedule of Hazardous substances issued by the Directorate: Health Technology of the Department of Health, a license in terms of the Act on Hazardous Substances (Act 15/1973) must be submitted with the bid / quotation documents. The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.5 | All electrical/electronic medical equipment must be licensed by Radiation Control, where applicable a copy of the license must be submitted The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.5.1 | Radiation Control License number:
5.6 | The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade (rubber to with 2 screws) type and it must be a minimum length of three (3) meters. NB the mains cable of the unit tendered for must be SABS colour coded.
5.7 | The equipment quoted for must be protected against electromagnetic interference IEC 60601-1-2
5.8 | The unit must comply with electrical safety and performance standard such as IEC 60601-2-21, attached certification
5.8.1 | Please provide unique ref number of the IEC ceriticate:
6 | COMPREHENSIVE MAINTENANCE
6.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) should be included in the unit price as per the SCC document
6.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
6.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
6.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
6.3 | The bidder to provide the OEM recommended service interval of the unit:
6.4 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
7 | General
7.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
7.2 | The unit on offer must be the latest technology.
7.2.1 | Initial Manufacturing year of Model:
RT2-09-007 | Amplitude Electroencephlograph- Accessories, Consumables and Upgrades | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
RT2-09-007a | AEEG Electrodes
RT2-09-007b | Mobile cart
RT2-09-007c | Paper Rolls

## Category 12- Microscopes
Microscopes and other Advanced Technolgy
# | Item number | Item Description
1 | RT2-12-001 | Microscope, Ophthalmology, Basic
2 | RT2-12-002 | Microscope, Ophthalmology, Advanced
3 | RT2-12-003 | Microscope, ENT, Basic
4 | RT2-12-004 | Microscope, ENT, Advanced
5 | RT2-12-005 | Microscope, Neurosurgery, Basic
6 | RT2-12-006 | Microscope, Neurosurgery, Advanced
7 | RT2-12-007 | Robotic Surgical System
8 | RT2-12-008 | Robotic Digital Surgical Microscope

## Microscope, Ophthal, Basic
RT2-12-001 | Microscope, Ophthalmology, Basic | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
1 | DESCRIPTION
1.1 | A surgical microscope suitable for opthalmology surgery with at least 180 degrees variable angle binocular tube configuration with foot pedal.
2 | Product Offered:
2.1 | Make:
2.2 | Model:
2.3 | Initial Manufacturing year of Model:
2.4 | Radiation Control License number:
3 | TECHNICAL SPECIFICATION
3.1 | The unit offered should include focusing, zooming and magnification controls.
3.2 | The unit must have a motorized zoom system with apochromatic lens, zoom ratio 1:6
3.3 | The unit offered should have either an LED light source and/or Xenon (180w) illumination system via a fibre optic cable.
3.4 | A microscope stand must have atleast 4 castors to ensure base stability and a minimum of 2 brakes on castors for smooth operation of the microscope.
3.5 | The microscope stand must have a mounted grip for easy movement of the unit
3.6 | The maximum load of the accessories to the microscope should not affect stability of the microscope
3.7 | The base of the stand shall be compact and have 360 degrees rotatable wheels to allow easy positioning of the microscope.
3.8 | The foot switch should be wireless or wired and waterproof
3.9 | The microscope offered should have pedal and/or hand-switch controls for ease of operation
3.10 | The pedal must be waterproof and inclinable.
3.11 | The surface of the microscope shall be anti-bacterial and have anti-mycotic protection.
3.12 | The optic carrier shall have an electric system to balance the three gravitational angles.
3.13 | Handles covers must be autoclave-able
3.14 | The microscope should have binocular tubes for the main surgeon and second observer tubes for the assistant.
3.15 | Surgeons shall be able to store their personal preferences for magnification, focus and illumination.
3.16 | System Functionality
3.16.1 | Magnification must be atleast 3 step with 1.6 to 20 x range or more
3.16.2 | Optics: Apochromatic with magnification changer
3.16.3 | Stereo base: 22mm - 25mm
3.16.4 | The trolley should have antistatic castors with at least 2 equipped with locking system/brake.
3.16.5 | Eyepiece:
3.16.5a | · 10x wide-field eyepieces
3.16.5b | · 12.5x wide-field eyepieces
3.16.5c | · Dioptric setting ±5 with adjustable eye cup
3.17.6 | Lateral tilt: 0 degrees - 360 degrees
3.17.7 | Inclination angle: 200 degrees
3.17.8 | Binocular tubes: 0 - 180 degrees
3.17.9 | Field of view diameter: 8 mm to 55 mm
3.17.10 | Focal length ranging from 150 - 400mm
3.16 | Illumination
3.16.1 | Illumination: two LED or Xenon lamps source with automatic lamp exchange and emergency back-up lamp.
3.16.2 | Protection filter: UV & IR filter
3.17 | X-Y Coupling
3.17.1 | Should have motorised X-Y movement
3.17.2 | Motorised focusing and zooming
3.17.3 | Should have a key for initial position of X-Y coupling and focus
4 | STANDARD ACCESSORIES
4.1 | Binocular tubes
4.2 | Binocular tubes (Second observer)
4.3 | Autoclavable handle covers
4.4 | Eyepiece cups
4.5 | Dustcover
4.6 | Objective lense: 200mm - 400mm
5 | Conformity Compliance (Please attach certificate)
5.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
5.2 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
5.2.1 | Please provide unique ref number of the ISO 13485 ceriticate:
5.3 | Model on offer must comply with EU Medical Device Regulation (EU) 2017/745 or equvalent from SAHPRA recognized jurisdictions . Attach a copy of certification
5.3.1 | Please provide unique ref number of the ceriticate:
5.4 | Hazardous Substance Act: If this type of equipment / apparatus appears on the schedule of Hazardous substances issued by the Directorate: Health Technology of the Department of Health, a license in terms of the Act on Hazardous Substances (Act 15/1973) must be submitted with the bid / quotation documents. The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.5 | All electrical/electronic medical equipment must be licensed by Radiation Control, where applicable a copy of the license must be submitted The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.5.1 | Radiation Control License number:
5.6 | The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade (rubber to with 2 screws) type and it must be a minimum length of three (3) meters. NB the mains cable of the unit tendered for must be SABS colour coded.
5.7 | The equipment quoted for must be protected against electromagnetic interference IEC 60601-1-2
5.8 | The unit must comply with electrical safety and performance standard such as IEC 60601-2-21, attached certification
5.8.1 | Please provide unique ref number of the IEC ceriticate:
6 | COMPREHENSIVE MAINTENANCE
6.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) should be included in the unit price as per the SCC document
6.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
6.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
6.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
6.3 | The bidder to provide the OEM recommended service interval of the unit:
6.4 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
7 | General
7.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
7.2 | The unit on offer must be the latest technology.
7.2.1 | Initial Manufacturing year of Model:
RT2-12-001 | Microscope, Ophthalmology, Basic- Accessories and Consumables | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
RT2-12-001a | Binocular tubes
RT2-12-001b | Binocular tubes (Second observer)
RT2-12-001c | Autoclavable handle covers
RT2-12-001d | Eyepiece cups
RT2-12-001e | Dustcover
RT2-12-001f | Xenon illumination system
RT2-12-001g | LED illumination system
RT2-12-001h | 10x wide-field eyepieces
RT2-12-001i | 12.5x wide-field eyepieces
RT2-12-001j | Sterile lens holder f = 200
RT2-12-001k | Sterile lens holder f = 175
RT2-12-001l | Stereo co-observation tube/Teaching piece

## Microscope, Ophthal, Advanced
RT2-12-002 | Microscope, Ophthalmology, Advanced | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
1 | DESCRIPTION
1.1 | A eletronic surgical microscope suitable for opthalmology surgery with at least180 degrees variable angle binocular tube configuration with foot pedal and an external monitor. The microscope must have image storage and a high definition 4K camera.
2 | Product Offered:
2.1 | Make:
2.2 | Model:
2.3 | Initial Manufacturing year of Model:
2.4 | Radiation Control License number:
3 | TECHNICAL SPECIFICATION
3.1 | The unit offered should include focusing, zooming and magnification controls.
3.2 | The unit must have a motorized zoom system with apochromatic lens, zoom ratio 1:6
3.3 | The unit must be integrated with both slit illuminator and keratoscope
3.4 | The unit offered should have either an LED light source and/or Xenon (180w) illumination system via a fibre optic cable.
3.5 | A microscope stand must have atleast 4 castors to ensure base stability and all 4 brakes on castors for smooth operation of the microscope.
3.6 | The microscope stand must have a mounted grip for easy movement of the unit
3.7 | The maximum load of the accessories to the microscope should not affect stability of the microscope
3.8 | The base of the stand shall be compact and have 360 degrees rotatable wheels to allow easy positioning of the microscope.
3.9 | The foot switch should be wireless or wired and waterproof
3.10 | The microscope offered should have pedal and/or hand-switch controls for ease of operation
3.11 | The pedal must be waterproof and inclinable, with a minimum of 15 programmable functions to control all the microscope's functions.
4.9 | The surface of the microscope shall be anti-bacterial and have anti-mycotic protection.
3.12 | The optic carrier shall have an electric system to balance the three gravitational angles.
3.13 | The microscope shall have a function to reduce the light intensity at a short distance during examination
5.9 | Handles covers must be autoclave-able
3.14 | The microscope offered shall have an integrated monitor, of at least 21" with HD 4K resolution.
3.15 | The microscope should have binocular tubes for the main surgeon and second observer tubes for the assistant.
6.9 | The control module shall have a screen with touch pad to select different user profiles and settings.
3.16 | Surgeons shall be able to store their personal preferences for magnification, focus and illumination.
3.17 | System Functionality
3.17.1 | Motorized magnification must be atleast 3 step with 1.6 to 20 x range or more
3.17.2 | Optics: Apochromatic with magnification changer
3.17.3 | Stereo base: 22mm - 25mm
3.17.4 | The trolley should have antistatic castors with at least 4 equipped with locking system/brake.
3.17.5 | Eyepiece:
3.17.5a | 10x wide-field eyepieces
3.17.5b | 12.5x wide-field eyepieces
3.17.5c | Dioptric setting ±5 with adjustable eye cup
3.17.6 | Lateral tilt: 0 degrees - 360 degrees
3.17.7 | Inclination angle: 200 degrees
3.17.8 | Binocular tubes: 0 - 180 degrees
3.17.9 | Field of view diameter: 8 mm to 55 mm
3.17.10 | Focal length ranging from 150 - 400mm
3.18 | Illumination
3.18.1 | Illumination: two LED or Xenon lamps source with automatic lamp exchange and emergency back-up lamp.
3.18.2 | Protection filter: UV & IR filter
3.19 | X-Y Coupling
3.19.1 | Should have motorised X-Y movement
3.19.2 | Motorised focusing and zooming
3.19.3 | Should have a key for initial position of X-Y coupling and focus
3.20 | The unit offered must have a data storage capacity of at least 2TB
4 | STANDARD ACCESSORIES
4.1 | Binocular tubes
4.2 | Binocular tubes (Second observer)
4.3 | Autoclavable handle covers
4.5 | Eyepiece cups
4.6 | Dustcover
4.7 | Video adapter
4.8 | Objective lense: 200mm - 400mm
5 | Conformity Compliance (Please attach certificate)
5.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
5.2 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
5.2.1 | Please provide unique ref number of the ISO 13485 ceriticate:
5.3 | Model on offer must comply with EU Medical Device Regulation (EU) 2017/745 or equvalent from SAHPRA recognized jurisdictions . Attach a copy of certification
5.3.1 | Please provide unique ref number of the ceriticate:
5.4 | Hazardous Substance Act: If this type of equipment / apparatus appears on the schedule of Hazardous substances issued by the Directorate: Health Technology of the Department of Health, a license in terms of the Act on Hazardous Substances (Act 15/1973) must be submitted with the bid / quotation documents. The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.5 | All electrical/electronic medical equipment must be licensed by Radiation Control, where applicable a copy of the license must be submitted The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.5.1 | Radiation Control License number:
5.6 | The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade (rubber to with 2 screws) type and it must be a minimum length of three (3) meters. NB the mains cable of the unit tendered for must be SABS colour coded.
5.7 | The equipment quoted for must be protected against electromagnetic interference IEC 60601-1-2
5.8 | The unit must comply with electrical safety and performance standard such as IEC 60601-2-21, attached certification
5.8.1 | Please provide unique ref number of the IEC ceriticate:
6 | COMPREHENSIVE MAINTENANCE
6.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) should be included in the unit price as per the SCC document
6.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
6.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
6.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
6.3 | The bidder to provide the OEM recommended service interval of the unit:
6.4 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
7 | General
7.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
7.2 | The unit on offer must be the latest technology.
7.2.1 | Initial Manufacturing year of Model:
RT2-12-002 | Ophthalmic Surgical Microscope, Advanced Accessories and Consumables | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
RT2-12-002a | Binocular tubes
RT2-12-002b | Binocular tubes (Second observer)
RT2-12-002c | Autoclavable handle covers
RT2-12-002d | Eyepiece cups
RT2-12-002e | Dustcover
RT2-12-002f | Xenon illumination system
RT2-12-002g | LED illumination system
RT2-12-002h | 10x wide-field eyepieces
RT2-12-002i | 12.5x wide-field eyepieces
RT2-12-002j | Sterile lens holder f = 200
RT2-12-002k | Sterile lens holder f = 175
RT2-12-002l | Stereo co-observation tube/Teaching piece

## ENT Microscope, Basic
RT2-12-003 | Microscope, ENT, Basic | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
1 | Product Offered:
1.1 | Make:
1.2 | Model:
1.3 | Initial Manufacturing year of Model:
1.4 | Radiation Control License number:
2 | DESCRIPTION
2.2 | A surgical microscope suitable for Ear, Nose and Throat surgery with at least 180 degrees variable angle binocular tube configuration with foot pedal. The microscope must have image storage and a 3CCD high definition camera.
3 | TECHNICAL SPECIFICATION
3.1 | The unit offered should include focusing, zooming and magnification controls.
3.2 | Magnification must be atleast 3 step with 1.6 to 20 x range
3.3 | The unit offered should have either an LED light source and/or Xenon (180w) illumination system via a fibre optic cable.
3.4 | A microscope stand shall have atleast 4 casters and a minimum of 2 electro-magnetic brakes for smooth operation of the microscope.
3.5 | The weight of the microscope stand without microscope must not be less than 170kg
3.6 | The microscope stand must have a mounted grip for easy movement of the unit
3.7 | The maximum load of the accessories to the microscope should be atleast 5kg
3.8 | The base of the stand must be compact and have 360 degrees rotatable wheels to allow easy positioning of the microscope.
3.9 | The foot switch must be waterproof
3.10 | The microscope offered must have pedal and/or hand-switch controls for ease of operation
3.11 | The pedal must be waterproof and inclinable, with a minimum of 15 programmable functions to control all the microscope's functions.
3.12 | The surface of the microscope must be anti-bacterial and have anti-mycotic protection.
3.13 | The optic carrier must have an electric system to balance the three gravitational angels.
3.14 | The microscope must have a function to automatically reduce the light intensity at a short distance during examination
3.15 | Handles covers must be autoclave-able
3.16 | The power consumption of the microscope offered must not exceed 800VA.
3.17 | The microscope offered must have an integrated monitor, ranging from 21inch to 30inch and HD resolution.
3.18 | The microscope should have binocular tubes for the main surgeon and second observer tubes for the assistant.
3.19 | The control module must have a screen with touch pad to select different user profiles and settings.
3.20 | Surgeons must be able to store their personal preferences for magnification, focus and illumination.
3.21 | The unit offered must have a built-in camera and a storage facility of at least 1TB.
3.22 | The HD monitor must be able to be viewed at different angles to accommodate teaching purposes
3.23 | System Functionality
3.23.1 | Magnification must be atleast 3 step with 1.6 to 20 x range
3.23.2 | Optics: Apochromatic with magnification changer
3.23.3 | Stereo base: 25mm
3.23.4 | The trolley must have antistatic castors with at least 2 equipped with locking system/brake.
3.23.5 | Eyepiece:
3.23.5.1 | 10x wide-field eyepieces
3.23.5.2 | 12.5x wide-field eyepieces
3.23.5.3 | Dioptric setting ±5 with adjustable eye cup
3.24 | Lateral tilt: 0 degrees - 180 degrees
3.25 | Inclination angle: -30 degrees - 120 degrees
3.26 | Illumination: two LED or Xenon lamps source with automatic lamp exchange and emergency back-up lamp.
3.27 | Field of view: 16.5 – 180 mm
3.28 | Focal length ranging from 150 - 400mm
3.29 | Protection filter: UV & IR filter
4 | STANDARD ACCESSORIES
4.1 | HD 21 inch Monitor
4.2 | 3-Chip Camera
4.3 | Binocular tubes
4.4 | Autoclavable handle covers
4.5 | Eyepiece cups
4.6 | Foot pedal
4.7 | Dustcover
4.9 | Video adapter
4.1 | Objective lense: 200mm - 400mm
5 | Conformity Compliance (Please attach certificate)
5.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
5.2 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
5.2.1 | Please provide unique ref number of the ISO 13485 ceriticate:
5.3 | Model on offer must comply with EU Medical Device Regulation (EU) 2017/745 or equvalent from SAHPRA recognized jurisdictions . Attach a copy of certification
5.3.1 | Please provide unique ref number of the ceriticate:
5.4 | Hazardous Substance Act: If this type of equipment / apparatus appears on the schedule of Hazardous substances issued by the Directorate: Health Technology of the Department of Health, a license in terms of the Act on Hazardous Substances (Act 15/1973) must be submitted with the bid / quotation documents. The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.5 | All electrical/electronic medical equipment must be licensed by Radiation Control, where applicable a copy of the license must be submitted The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.5.1 | Radiation Control License number:
5.6 | The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade (rubber to with 2 screws) type and it must be a minimum length of three (3) meters. NB the mains cable of the unit tendered for must be SABS colour coded.
5.7 | The equipment quoted for must be protected against electromagnetic interference IEC 60601-1-2
5.8 | The unit must comply with electrical safety and performance standard such as IEC 60601-2-21, attached certification
5.8.1 | Please provide unique ref number of the IEC ceriticate:
6 | COMPREHENSIVE MAINTENANCE
6.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) should be included in the unit price as per the SCC document
6.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
6.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
6.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
6.3 | The bidder to provide the OEM recommended service interval of the unit:
6.4 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
7 | General
7.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
7.2 | The unit on offer must be the latest technology.
7.2.1 | Initial Manufacturing year of Model:
RT2-12-003 | Ent Operating Microscope, Basic Accessories and Consumables | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
RT2-12-003a | Binocular tubes
RT2-12-003b | Binocular tubes (Second observer)
RT2-12-003c | Autoclavable handle covers
RT2-12-003d | Eyepiece cups
RT2-12-003e | Dustcover
RT2-12-003f | Xenon illumination system
RT2-12-003g | LED illumination system
RT2-12-03h | 10x wide-field eyepieces
RT2-12-003i | 12.5x wide-field eyepieces
RT2-12-003j | Sterile lens holder f = 200
RT2-12-003k | Sterile lens holder f = 175
RT2-12-003l | Stereo co-observation tube/Teaching piece

## ENT Microscope, Advanced
RT2-12-004 | Microscope, ENT, Advanced | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
1 | DESCRIPTION
1.1 | A surgical microscope suitable for Ear, Nose and Throat surgery with at least 180 degrees variable angle binocular tube configuration with foot pedal and an external 3D Monitor. The microscope must have integrated monitor, image storage and a 4K camera.
2 | Product Offered:
2.1 | Make:
2.2 | Model:
2.3 | Initial Manufacturing year of Model:
2.4 | Radiation Control License number:
3 | TECHNICAL SPECIFICATION
3.1 | The unit offered should include focusing, zooming and magnification controls.
3.2 | Magnification must be atleast 3 step with 1.6 to 20 x range
3.3 | The unit offered should have either an LED light source and/or Xenon (300w) illumination system via a fibre optic cable.
3.4 | A microscope stand must have atleast 4 casters and a minimum of 2 electro-magnetic brakes for smooth operation of the microscope.
3.5 | The weight of the microscope stand without microscope must not be less than 170kg
3.6 | The microscope stand must have a mounted grip for easy movement of the unit
3.7 | The maximum load of the accessories to the microscope should be atleast 5kg
3.8 | The base of the stand shall be compact and have 360 degrees rotatable wheels to allow easy positioning of the microscope.
3.9 | The foot switch should be waterproof
3.10 | The microscope offered should have pedal and/or hand-switch controls for ease of operation
3.11 | The pedal must be waterproof and inclinable, with a minimum of 15 programmable functions to control all the microscope's functions.
3.12 | The surface of the microscope shall be anti-bacterial and have anti-mycotic protection.
3.13 | The optic carrier shall have an electric system to balance the three gravitational angels.
3.14 | The microscope shall have a function to automatically reduce the light intensity at a short distance during examination
3.15 | Handles covers must be autoclave-able
3.16 | The power consumption of the microscope offered should not exceed 800VA.
3.17 | The microscope offered shall have an integrated monitor, ranging from 24inch and 4K HD resolution.
3.18 | The microscope should have binocular tubes for the main surgeon and second observer tubes for the assistant.
3.19 | An electronic surgical microscope for ENT with full display monitor, 1TB image storage and a 4K HD camera.
3.20 | The control module shall have a screen with touch pad to select different user profiles and settings.
3.21 | Surgeons shall be able to store their personal preferences for magnification, focus and illumination.
3.22 | The unit offered shall have a built-in camera and a storage facility.
3.23 | At least 40" external monitor must be supplied and should be able to be viewed at different angles to accommodate teaching purposes
3.24 | The bidder must supply mobile stand or swivel mounting bracket for external monitor .
3.25 | System Functionality
3.25.1 | Magnification must be atleast 3 step with 1.6 to 20 x range
3.25.2 | Optics: Apochromatic with magnification changer
3.25.3 | Stereo base: 25mm
3.25.4 | The trolley should have antistatic castors with at least 2 equipped with locking system/brake.
3.25.5 | Eyepiece:
3.25.5.1 | 10x wide-field eyepieces
3.25.5.2 | 12.5x wide-field eyepieces
3.25.5.3 | Dioptric setting ±5 with adjustable eye cup
3.26 | Lateral tilt: 0 degrees - 180 degrees
3.27 | Inclination angle: -30 degrees - 120 degrees
3.28 | Illumination: two LED or Xenon lamps source with automatic lamp exchange and emergency back-up lamp.
3.29 | Field of view: 16.5 – 180 mm
3.30 | Focal length ranging from 200 - 500mm
3.31 | Protection filter: UV & IR filter
4 | STANDARD ACCESSORIES
4.1 | 4K HD Monitor
4.2 | 4K HD Camera
4.3 | Binocular tubes
4.4 | Autoclavable handle
4.5 | Eyepiece cups
4.6 | Wireless foot pedal
4.7 | Dustcover
4.8 | Binocular Tubes (Second observer)
4.9 | Video adapter
4.10 | Objective lense: 200mm - 400mm
4.11 | Xenon illumination system
4.12 | Stereo co-observation tube/Teaching piece
5 | Conformity Compliance (Please attach certificate)
5.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
5.2 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
5.2.1 | Please provide unique ref number of the ISO 13485 ceriticate:
5.3 | Model on offer must comply with EU Medical Device Regulation (EU) 2017/745 or equvalent from SAHPRA recognized jurisdictions . Attach a copy of certification
5.3.1 | Please provide unique ref number of the ceriticate:
5.4 | Hazardous Substance Act: If this type of equipment / apparatus appears on the schedule of Hazardous substances issued by the Directorate: Health Technology of the Department of Health, a license in terms of the Act on Hazardous Substances (Act 15/1973) must be submitted with the bid / quotation documents. The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.5 | All electrical/electronic medical equipment must be licensed by Radiation Control, where applicable a copy of the license must be submitted The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.5.1 | Radiation Control License number:
5.6 | The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade (rubber to with 2 screws) type and it must be a minimum length of three (3) meters. NB the mains cable of the unit tendered for must be SABS colour coded.
5.7 | The equipment quoted for must be protected against electromagnetic interference IEC 60601-1-2
5.8 | The unit must comply with electrical safety and performance standard such as IEC 60601-2-21, attached certification
5.8.1 | Please provide unique ref number of the IEC ceriticate:
6 | COMPREHENSIVE MAINTENANCE
6.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) should be included in the unit price as per the SCC document
6.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
6.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
6.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
6.3 | The bidder to provide the OEM recommended service interval of the unit:
6.4 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
7 | General
7.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
7.2 | The unit on offer must be the latest technology.
7.2.1 | Initial Manufacturing year of Model:
RT2-12-004 | Microscope ENT- Accessories and Consumables | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
RT2-12-004a | Binocular tubes
RT2-12-004b | Binocular tubes (Second observer)
RT2-12-004c | Autoclavable handle covers
RT2-12-004d | Eyepiece cups
RT2-12-004e | Dustcover
RT2-12-004f | Xenon illumination system
RT2-12-004g | LED illumination system
RT2-12-004i | 10x wide-field eyepieces
RT2-12-004j | 12.5x wide-field eyepieces
RT2-12-004k | Sterile lens holder f = 200
RT2-12-004l | Sterile lens holder f = 175
RT2-12-004m | Stereo co-observation tube/Teaching piece
RT2-12-004n | 4K HD Monitor
RT2-12-004o | 4K HD Camera
RT2-12-004p | Video adapter

## Microscope, Neurosurgery
RT2-12-005 | Microscope, Neurosurgery, Basic | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
1 | DESCRIPTION
1.1 | A specialized optical instrument providing magnified visualization for the nuerosurgoens during procedures.
2 | Product Offered:
2.1 | Make:
2.2 | Model:
2.3 | Initial Manufacturing year of Model:
2.4 | Radiation Control License number:
3 | TECHNICAL SPECIFICATION
3.1 | An electronic surgical microscope with full display monitor, image storage and a 4K high definition camera.
3.2 | The unit offered should include focusing, zooming and magnification controls.
3.3 | Magnification must be atleast 6 steps with 1.6 to 20 x range
3.4 | Optics: Apochromatic with magnification changer
3.5 | Stereo base: 22 - 25mm
3.6 | The following Eyepiece options must be available:
3.6.1 | 10x wide-field eyepieces
3.6.2 | 12.5x wide-field eyepieces
3.6.3 | Dioptric setting ±5 with adjustable eye cup
3.7 | The unit offered must support the following movements:
3.7.1 | Lateral tilt: -30 degrees - 50 degrees
3.7.2 | Inclination angle: -30 degrees - 120 degrees
3.8 | Illumination: two LED or Xenon lamps source with automatic lamp exchange and emergency back-up lamp.
3.9 | Field of view: 16.5 – 210 mm
3.10 | Focal length ranging from 150 - 400mm
3.11 | Protection filter: UV & IR filter
3.12 | A microscope stand must have 4 antistatic castors with at least 2 equipped with locking system/brake to ensure base stability for smooth operation.
3.13 | The microscope stand must have a mounted grip for easy movement of the unit
3.14 | The maximum load of the accessories to the microscope should be at least 5kg
3.15 | The base of the stand must be compact and have 360 degrees rotatable wheels to allow easy positioning of the microscope.
3.16 | The foot switch should be wireless or wired and waterproof
3.17 | The microscope offered should have pedal and/or hand-switch controls for ease of operation
3.18 | The pedal must be waterproof and inclinable, with a minimum of 10 programmable functions to control all the microscope's functions.
3.19 | The surface of the microscope must be anti-bacterial and have anti-mycotic protection.
3.20 | The optic carrier must have an electric system to balance the three gravitational angles.
3.21 | The microscope must have a function to automatically reduce the light intensity at a short distance during examination
3.22 | Handles covers must be autoclave-able
3.23 | The microscope offered must have an integrated monitor, of at least 21" with 4K HD resolution.
3.24 | The microscope should have binocular tubes for the main surgeon and second observer tubes for the assistant.
3.25 | The control module must have a screen with touch pad to select different user profiles and settings.
3.26 | Surgeons should be able to programme and store their personal preferred settings for magnification, focus and illumination.
3.27 | The unit offered must be supplied with UPS that can power the unit for at least 2 hours and must be mounted on the stand.
4 | Standard Accessories and Consumables
4.1 | 4K HD Monitor
4.2 | 4K HD Camera
4.3 | Binocular tubes
4.4 | Autoclavable handle
4.5 | Eyepiece cups
4.6 | Wireless foot pedal
4.7 | Dustcover
4.8 | Binocular Tubes (Second observer)
4.9 | Video adapter
4.10 | Objective lense: 200mm - 400mm
4.11 | Xenon illumination system
4.12 | Stereo co-observation tube/Teaching piece
4.13 | Uniterrupted Power Supply (UPS)
5 | Conformity Compliance (Please attach certificate)
5.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
5.2 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
5.2.1 | Please provide unique ref number of the ISO 13485 ceriticate:
5.3 | Model on offer must comply with EU Medical Device Regulation (EU) 2017/745 or equvalent from SAHPRA recognized jurisdictions . Attach a copy of certification
5.3.1 | Please provide unique ref number of the ceriticate:
5.4 | Hazardous Substance Act: If this type of equipment / apparatus appears on the schedule of Hazardous substances issued by the Directorate: Health Technology of the Department of Health, a license in terms of the Act on Hazardous Substances (Act 15/1973) must be submitted with the bid / quotation documents. The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.5 | All electrical/electronic medical equipment must be licensed by Radiation Control, where applicable a copy of the license must be submitted The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.5.1 | Radiation Control License number:
5.6 | The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade (rubber to with 2 screws) type and it must be a minimum length of three (3) meters. NB the mains cable of the unit tendered for must be SABS colour coded.
5.7 | The equipment quoted for must be protected against electromagnetic interference IEC 60601-1-2
5.8 | The unit must comply with electrical safety and performance standard such as IEC 60601-2-21, attached certification
5.8.1 | Please provide unique ref number of the IEC ceriticate:
6 | COMPREHENSIVE MAINTENANCE
6.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) should be included in the unit price as per the SCC document
6.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
6.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
6.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
6.3 | The bidder to provide the OEM recommended service interval of the unit:
6.4 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
7 | General
7.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
7.2 | The unit on offer must be the latest technology.
7.2.1 | Initial Manufacturing year of Model:
RT2-12-005 | Microscope, Neurosurgery, Mobile - Accessories and Consumables | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
RT2-12-005a | Binocular tubes
RT2-12-005b | Binocular tubes (Second observer)
RT2-12-005c | Autoclavable handle covers
RT2-12-005d | Eyepiece cups
RT2-12-005e | Dustcover
RT2-12-005f | Xenon illumination system
RT2-12-005g | LED illumination system
RT2-12-005i | 10x wide-field eyepieces
RT2-12-005j | 12.5x wide-field eyepieces
RT2-12-005k | Sterile lens holder f = 200
RT2-12-005l | Sterile lens holder f = 175
RT2-12-005m | Stereo co-observation tube/Teaching piece
RT2-12-005n | 4K HD Monitor at least 21''(Bidder to state make and model)
RT2-12-005o | 4K HD Camera (Bidder to state make and model)
RT2-12-005p | Video adapter
RT2-12-005q | Uniterrupted Power Supply (UPS)
RT2-12-005r | Micro Inspection Tool
RT2-12-005s | 4K HD Monitor at least 40'' for training purpose (Bidder to state make and model)

## Microscope, Neuro, Advanced
RT2-12-006 | Microscope, Neurosurgery, Advanced | COMPLIES YES/NO | Provide your answer inthis column. You are advised to be straight to the point.
1 | DESCRIPTION
1.1 | A specialized optical instrument providing magnified visualization for the neurosurgeons during procedures.
2 | Product Offered:
2.1 | Make:
2.2 | Model:
2.3 | Initial Manufacturing year of Model:
2.4 | Radiation Control License number:
3 | TECHNICAL SPECIFICATION
3.1 | An electronic surgical microscope with full display monitor, image storage and a 4K high definition camera.
3.2 | The surgical microscope must be specifically designed for neurosurgery and not a general microscope
3.3 | Microscope degree of freedom must allow work in difficult angles to reach areas of interest
3.4 | Microscope robotics must have at least 6 axis of movement
3.5 | MAGNIFICATION SYSTEM
3.5.1 | The system must have a motorized zoom; 1:6 zoom ratio
3.5.2 | The magnification with adjustable speed must be displayed.
3.5.3 | The system must have a user specific start position
3.5.4 | 10x to 12.5x magnetic widefield eyepieces with intergrated eyecups must be available
3.5.5 | Tubes that allow rapid switch between large magnification and maximum overview of the region of interest must be available
3.6 | FOCUSING SYSTEM
3.6.1 | Autofocus with 2 visible laser spots must be possible
3.6.2 | The focusing system must have a motorised working range of at least 225mm to 575mm
3.6.3 | The system must have a internal motorised continuos adjustment
3.6.4 | The magnification must be linked with adjustment of the focus speed
3.6.5 | The system must have a high-speed auto focus, with an accuracy of ± 0.5 mm
3.6.6 | The system must be able to adjust the depth of field electronically.
3.6.7 | the microscope must be able to offer integrated tube rotation for better ergonomic observation
3.7 | ILLUMINATION
3.7.1 | At least 2x300W xenon or LED lamps with auto switch over
3.7.2 | Illumination field diameter able to vary automatically
3.7.3 | Must be able to display the remaining lamp life on the screen
3.7.4 | The setting to adjust light threshold must be available
3.8 | SYSTEM OPERATION
3.8.1 | Multifunctional programmable hand grips must be available
3.8.2 | Must have a central user interface
3.8.3 | The Microscope system must maintain automatic balance across entire tilt and swivel range, without requiring manual axis adjustment.
3.9 | TUBES AND CO-OBSERVATION
3.9.1 | The unit offered must include stereo-co-observation tube
3.9.2 | Stereo co-observation tube must remain fixed when tilting microscope head
3.9.3 | The main tube must be 0° - 360° rotatable
3.9.4 | Face to face surgery must be possible with both surgeon and assistant in a standing position
3.10 | SYSTEM SETUP
3.10.1 | Autobalance of microscope or entire system must be possible by means of a push button.
3.10.2 | Offered system must be draped by means of a integrated vacuum system to remove air from sterile drape for fast and easy draping
3.10.3 | Mouth switch for fine balance must be available
3.10.4 | System must have a central touch screen user interface that programs and allows for operation of the entire system that includes the microscope, patient data management, video camera and recording
3.10.5 | Surgeon must be able to see parameters such as magnification, working distance and light intesity within the field of view of the occulars
3.11 | VIDEO
3.11.1 | At least 4K HD Video Camera Unit with recording and storing capability in both 2D and 3D must be available
3.11.2 | At least 22 inch 4K HD video touchscreen on arm extension to allow visualisation of microscope image
3.11.3 | The system must include integrated video and still image capture capability, stored in the internal 1TB HDD and external USB.
3.11.4 | It must be possible to record still images from the inegrated HD video camera without requiring additional recording devices
3.12 | CONNECTIVITY
3.12.1 | Video-in for external SD video sources
3.12.2 | Microscope interface integration with endoscope and neuronavigation
3.12.3 | DICOM module for data transfer from and to PACS system
3.12.4 | At least 2 x USB port
3.12.5 | HDMI port
3.13 | INTRAOPERATIVE FLUORESCENCE AND INDOCYANINE GREEN (ICG)
3.13.1 | The system must offer an option for fluorescence guided glioma resection
3.13.2 | The system must offer fluorescence based video angiography with automatic analysis tool
3.13.3 | The system must offer automated setup for focus, zoom and light for optimum fluorescene visualisation
3.13.4 | The system must allow parallel automatic recording of fluorescence video and standard white light video
3.13.5 | Fluorescence camera system must have manual and automatic gain control for adjusting fluorescence image quality
3.14 | ELECTRICAL DATA
3.14.1 | The Voltage supply must be 220V,50Hz
3.14.2 | The plug must be SABS approved and colour coded.
4 | Standard Accessories and Consumables
4.1 | 4K HD Monitor
4.2 | 4K HD Camera
4.3 | Binocular tubes
4.4 | Autoclavable handle
4.5 | Eyepiece cups
4.6 | Wireless foot pedal
4.7 | Dustcover
4.8 | Binocular Tubes (Second observer)
4.9 | Video adapter
4.10 | Objective lense: 200mm - 400mm
4.11 | Xenon illumination system
4.12 | Stereo co-observation tube/Teaching piece
4.13 | Uniterrupted Power Supply (UPS)
5 | Conformity Compliance (Please attach certificate)
5.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
5.2 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
5.2.1 | Please provide unique ref number of the ISO 13485 ceriticate:
5.3 | Model on offer must comply with EU Medical Device Regulation (EU) 2017/745 or equvalent from SAHPRA recognized jurisdictions . Attach a copy of certification
5.3.1 | Please provide unique ref number of the ceriticate:
5.4 | Hazardous Substance Act: If this type of equipment / apparatus appears on the schedule of Hazardous substances issued by the Directorate: Health Technology of the Department of Health, a license in terms of the Act on Hazardous Substances (Act 15/1973) must be submitted with the bid / quotation documents. The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.5 | All electrical/electronic medical equipment must be licensed by Radiation Control, where applicable a copy of the license must be submitted The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.5.1 | Radiation Control License number:
5.6 | The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade (rubber to with 2 screws) type and it must be a minimum length of three (3) meters. NB the mains cable of the unit tendered for must be SABS colour coded.
5.7 | The equipment quoted for must be protected against electromagnetic interference IEC 60601-1-2
5.8 | The unit must comply with electrical safety and performance standard such as IEC 60601-2-21, attached certification
5.8.1 | Please provide unique ref number of the IEC ceriticate:
6 | COMPREHENSIVE MAINTENANCE
6.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) should be included in the unit price as per the SCC document
6.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
6.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
6.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
6.3 | The bidder to provide the OEM recommended service interval of the unit:
6.4 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
7 | General
7.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
7.2 | The unit on offer must be the latest technology.
7.2.1 | Initial Manufacturing year of Model:
RT2-12-006 | MICROSCOPE, NEUROSURGERY ADVANCED - Accessories and Consumables | Complies Yes/No | Provide your answer inthis column. You are advised to be straight to the point.
RT2-12-006a | Binocular tubes
RT2-12-006b | Binocular tubes (Second observer)
RT2-12-006c | Autoclavable handle covers
RT2-12-006d | Eyepiece cups
RT2-12-006e | Dustcover
RT2-12-006f | Xenon illumination system
RT2-12-006g | LED illumination system
RT2-12-006i | 10x wide-field eyepieces
RT2-12-006j | 12.5x wide-field eyepieces
RT2-12-006k | Sterile lens holder f = 200
RT2-12-006l | Sterile lens holder f = 175
RT2-12-006m | Stereo co-observation tube/Teaching piece
RT2-12-006n | 4K HD Monitor at least 21''(Bidder to state make and model)
RT2-12-006o | 4K HD Camera (Bidder to state make and model)
RT2-12-006p | Video adapter
RT2-12-006q | Uniterrupted Power Supply (UPS)
RT2-12-006r | Micro Inspection Tool
RT2-12-006s | 4K HD Monitor at least 40'' for training purpose (Bidder to state make and model)

## Robotic Surgery system
RT2-12-007 | Robotic Surgery System, Multi-discipline | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
1 | DESCRIPTION
1.1 | A cutting-edge medical platform enabling minimally invasive procedures across various specialities using robotics technology.
2 | Product Offered:
2.1 | Make:
2.2 | Model:
2.3 | Initial Manufacturing year of Model:
3 | TECHNICAL SPECIFICATION
3.1 | The type of robotic surgery equipment must be able to be used in urology, general surgery, colorectal surgery, upper GI surgery, gynecology, pediatric surgery, ENT, thoracic surgery and endocrine surgery.
3.2 | Customer portal with reporting tools for operating parameters must be possible.
3.3 | The unit must be suitable for childern and adults patients.
3.4 | The robot with its components must have a 3D capability.
3.5 | The system must be compatible to be performed by two operators in robot console simultaneously(Optional).
3.6 | The unit must be able to perform simulated treatments for educational purposes without the use surgical robot.
3.6.1 | All devices/components required for simulation must be supplied.
3.7 | The robot must have at least four (4) arms which can be selected based on surgery requirement.
3.8 | The system must enable the transfer of instruments and optics between ports, with corresponding control functionality automatically relocating to the new port to ensure continuous and precise control of instruments. Bidder to specify other methods if not similar.
3.9 | The instruments must have haptic memory as well as being self-instructing to enable quick removal and insertion of any instrument into the patient.
3.10 | It must be possible for surgery in all four quadrants of the abdomen to be carried out with docking towards the patient.
3.11 | Vessel sealing must be possible from the console
3.11.1 | It must be possible to control the electric vessel sealer for permanent sealing
3.12 | It must be possible to connect monopolar and bipolar instrument at the same time.
3.13 | Multiple screen viewing must be possible.
3.14 | Synchronization between operating table and operating robot is desireable to enable position changes with at least 2 robot arms and the camera arm docked with the instruments inside the patient(Optional).
3.16 | The unit must be ergonomically designed to programme to users' working positions.
3.17 | Memory function for personal settings must be possible (Optional).
3.18 | The unit must be easy to move out of the way during patient emergency.
3.19 | The unit must have tremor or vibration filtration
3.20 | Suction and irrigation must be possible.
3.21 | Camera unit and camera head.
3.21.1 | Image quality must be at least 3D and High Definition.
3.21.2 | Magnification of 10x must be available
3.22 | Light source
3.22.1 | Make & Model:
3.22.2 | Lamp design must be LED.
3.22.3 | Light source must have adjustable light intensity
3.22.4 | Automatic brightness control must be included
3.23 | Monitor
3.23.1 | Make & Model:
3.23.2 | It must be possible to connect more than one monitor to duplicate the operator's view.
3.23.3 | Offered system must contain at least one monitor.
3.23.4 | Monitor position must be adjustable horizontally and vertically.
3.24 | Indocyanine Green (ICG) function or similar
3.24.1 | The offered equipment must provide ICG function or similar as a tracer.
3.24.2 | Activation of NIR to see ICG should be done by the operator without having to release the joystick.
3.24.3 | Image by NIR must be displayed on all screens.
3.25 | Data Management System
3.25.1 | Make & Model:
3.25.2 | It must be possible to store image information in the form of still images and videos.
3.25.3 | Storage capacity of at least 1TB
3.26 | IT requirements
2.26.1 | The system must be connected to network.
2.26.2 | The system documentation must contain a description of which parts of the IT support handle the storage of information permanently or temporarily.
2.26.3 | IT support must be possible to configure in terms of user interfaces, work processes (workflows), reports, forms etc. without having to make customer-unique adaptations.
2.26.4 | Any software upgrade or licensing must be provided at no costs to end-users.
3.27 | Endoscopes
2.27.1 | Endoscopes must be available with at least 0- and 30- degree lens.
2.27.2 | The equipment must have a function that limits fogging on the optics. The bidder to state how it is achieved.
2.27.3 | The bidder must state optics dimensions.
3.28 | ELECTRICAL DATA
3.28.1 | The Voltage supply must be 220V/50Hz
4 | Standard Accessories and Consumables
4.1 | Single-use sterile drapes for 20 operations
4.2 | Trocars
4.3 | Monopolar instruments (End-user choice)
4.4 | Bipolar instruments (End-use choice)
4.5 | Forcep instruments (End-use choice)
4.6 | Sterilization tray
4.7 | Vessel sealing Instruments
4.8 | Suction Instruments
4.9 | Stapler Instruments (End-user choice)
4.10 | 3 D Glasses x 2
5 | Conformity Compliance (Please attach certificate)
5.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
5.2 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
5.2.1 | Please provide unique ref number of the ISO 13485 ceriticate:
5.3 | Model on offer must comply with EU Medical Device Regulation (EU) 2017/745 or equvalent from SAHPRA recognized jurisdictions . Attach a copy of certification
5.3.1 | Please provide unique ref number of the ceriticate:
5.4 | Hazardous Substance Act: If this type of equipment / apparatus appears on the schedule of Hazardous substances issued by the Directorate: Health Technology of the Department of Health, a license in terms of the Act on Hazardous Substances (Act 15/1973) must be submitted with the bid / quotation documents. The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.5 | All electrical/electronic medical equipment must be licensed by Radiation Control, where applicable a copy of the license must be submitted The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.5.1 | Radiation Control License number:
5.6 | The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade (rubber to with 2 screws) type and it must be a minimum length of three (3) meters. NB the mains cable of the unit tendered for must be SABS colour coded.
5.7 | The equipment quoted for must be protected against electromagnetic interference IEC 60601-1-2
5.8 | The unit must comply with electrical safety and performance standard such as IEC 60601-2-21, attached certification
5.8.1 | Please provide unique ref number of the IEC ceriticate:
6 | COMPREHENSIVE MAINTENANCE
6.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) should be addressed as per the SCC document
6.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty must commence after formal Commissioning and handover of the equipment.
6.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
6.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract must commence after the 2-year warranty period has expired.
6.3 | Bidder to state how cleaning and disinfection of robotic instruments & optics is conducted.
6.4 | The bidder to provide the OEM recommended service interval of the unit:
6.5 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
6.6 | Training
6.6.1 | At least 4 surgeons must obtain the amount of training that is deemed necessary for proper handling of the unit at no extra cost to the state including course material.
6.6.2 | An application specialist must be present for a sufficient period of time for approximately three months of operation and as-and-when required.
7 | General
7.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
7.2 | The unit on offer must be the latest technology.
7.3 | Initial Manufacturing year of Model:
RT2-12-007 | Robotic Surgery System, Replacement Accessories and consumables | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
RT2-12-007a | Single-use sterile drapes for 20 operations
RT2-12-007b | Trocars
RT2-12-007c | Monopolar instruments (End-user choice)
RT2-12-007d | Bipolar instruments (End-use choice)
RT2-12-007e | Forcep/Clip Applier instruments (End-use choice)
RT2-12-007f | Sterilization tray
RT2-12-007h | Vessel sealing Instruments
RT2-12-007i | Suction-Irrigation Instruments
RT2-12-007j | Stapler Instruments (End-user choice)
RT2-12-007k | 3 D Glasses x 2

## Digital Surgical Microscope
RT2-12-008 | Robotic Digital Microscope System, Multi-Disciplines | COMPLIES YES/NO | Provide your answer inthis column. You are advised to be straight to the point.
1 | DESCRIPTION
1.1 | An advanced system for precise and magnified visualization combining robotics, optics, and digital technology for various medical specialties.
2 | Product Offered:
2.1 | Make:
2.2 | Model:
2.3 | Initial Manufacturing year of Model:
2.4 | Radiation Control License number:
3 | TECHNICAL SPECIFICATION
3.1 | The surgical microscope must be designed for multidisciplines surgery with special focus on neurosurgery
3.2 | A microscope stand must have 4 antistatic castors with at least 2 equipped with locking system/brake to ensure base stability for smooth operation.
3.3 | Microscope degree of freedom must allow work in difficult angles to reach areas of interest
3.5 | MAGNIFICATION SYSTEM
3.5.1 | The system must have a motorized zoom; 1:6 zoom ratio
3.5.2 | The magnification with adjustable speed must be displayed.
3.5.3 | The system must have a user specific start position
3.5.5 | Tubes that allow rapid switch between large magnification and maximum overview of the region of interest must be available
3.5.6 | Adjustable magnification speed must be possible
3.6 | FOCUSING SYSTEM
3.6.1 | Autofocus with 2 visible laser spots must be available
3.6.2 | The focusing system must have a motorised working range from 200mm to at least 450mm
3.6.3 | The system must have an internal motorised continuos adjustment
3.6.4 | The magnification must be linked with adjustment of the focus speed
3.6.5 | The system must have a high-speed auto focus, with an accuracy of ± 0.5 mm
3.6.6 | The depth of field must be adjustable electronically.
3.7 | ILLUMINATION
3.7.1 | At least 2x300W xenon or LED lamps with auto switch over
3.7.2 | illumination field diameter must be able to vary automatically
3.7.3 | Must be able to display the remaining lamp life on the screen
3.7.4 | Adjustable light threshold settings must be available
3.8 | ROBOTIC ARM OPERATION
3.8.1 | The system must incorporate multifunctional programmable hand grips
3.4 | Microscope robotics must have at least 6 axis of movement
3.8.2 | The unit must have tremor or vibration filtration for stability
It must be possible to pinpoint and lock the camera on target
The unit must allow surgeon to programme positions and save for quick return when needed.
3.8.3 | Automatic microscope balancing must happen in any position of the microscopes'tilt and swivel range and entire microscope must balance all system axes without any manual input or axis adjustment
3.10 | SYSTEM SETUP
3.10.1 | Autobalance of microscope or entire system must be by means of push button
3.10.2 | Offered system must be draped by means of a integrated vacuum system to remove air from sterile drape for fast and easy draping
3.10.3 | Mouth switch for fine balance must be available
3.10.4 | System must have a central touch screen user interface that programs and allows for operation of the entire system that includes the microscope, patient data management, video camera and recording
3.10.5 | Surgeon must be able to see parameters such as magnification, working distance and light intesity on the screen
3.11 | VIDEO
3.11.1 | At least 4k 3D Video Camera Unit with recording and storing capability must be provided.
3.11.2 | The unit must incorporate two (2) surgical monitors mounted on arm extension.
3.11.2.1 | The surgical monitors above, must be at least 26 inch 4k 3D touchscreen medical grade monitors
3.11.3 | Integrated video still image capturing on HDD and USB must be possible
3.11.4 | It must be possible to record still images from the inegrated video camera without requiring additional recording devices
3.11.5 | External monitor of at least 50" 4k 3D video quality and mounted on a trolley or wall bracket must be supplied (As per user requirement).
3.12 | CONNECTIVITY
3.12.1 | Video-in for external SD video sources
3.12.2 | Microscope interface integration with other endoscopic system must be possible.
3.12.3 | DICOM module for data transfer from or to PACS or any Health Information System (HIS) must be available.
3.13 | INTRAOPERATIVE FLUORESCENCE AND ICG
3.13.1 | The system must offer an option for fluorescence guided glioma resection
3.13.2 | The system must offer fluorescence based video angiography with automatic analysis tool
3.13.3 | The system must offer automated setup for focus, zoom and light for optimum fluorescene visualisation
3.13.4 | The system must allow parallel automatic recording of fluorescence video and standard white light video
3.13.5 | Fluorescence camera system must have manual and automatic gain control for adjusting fluorescence image quality
3.14 | ELECTRICAL DATA
3.14.1 | The Voltage supply must be 220V/50Hz
4 | Standard Accessories and Consumables
4.1 | Single-use sterile drapes for 20 operations
3D glasses or clip-on's (Pk of 6)
5 | Conformity Compliance (Please attach certificate)
5.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
5.2 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
5.2.1 | Please provide unique ref number of the ISO 13485 ceriticate:
5.3 | Model on offer must comply with EU Medical Device Regulation (EU) 2017/745 or equvalent from SAHPRA recognized jurisdictions . Attach a copy of certification
5.3.1 | Please provide unique ref number of the ceriticate:
5.4 | Hazardous Substance Act: If this type of equipment / apparatus appears on the schedule of Hazardous substances issued by the Directorate: Health Technology of the Department of Health, a license in terms of the Act on Hazardous Substances (Act 15/1973) must be submitted with the bid / quotation documents. The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.5 | All electrical/electronic medical equipment must be licensed by Radiation Control, where applicable a copy of the license must be submitted The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.5.1 | Radiation Control License number:
5.6 | The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade (rubber to with 2 screws) type and it must be a minimum length of three (3) meters. NB the mains cable of the unit tendered for must be SABS colour coded.
5.7 | The equipment quoted for must be protected against electromagnetic interference IEC 60601-1-2
5.8 | The unit must comply with electrical safety and performance standard such as IEC 60601-2-21, attached certification
5.8.1 | Please provide unique ref number of the IEC ceriticate:
6 | COMPREHENSIVE MAINTENANCE
6.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) should be included in the unit price as per the SCC document
6.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
6.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
6.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
6.3 | The bidder to provide the OEM recommended service interval of the unit:
6.4 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
7 | General
7.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
7.2 | The unit on offer must be the latest technology.
7.2.1 | Initial Manufacturing year of Model:
RT2-12-008 | Robotic Digital Microscope System, Multi-Disciplines, Replacement Accessories and consumables | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
RT2-12-008a | Single-use sterile drapes for 20 operations
RT2-12-008b | 3 D Glasses x 2

## Category 14- Summary
Multi-disciplinary Specialized Equipment
# | Item number | Item Description
1 | RT2-14-001 | Laparascope, Mid-range
2 | RT2-14-002 | Laparascope, High-end (3D)
3 | RT2-14-003 | Lapscope Instruments
4 | RT2-14-004 | Endoscopic & Exoscopic, Video System, Neurosurgery
5 | RT2-14-005 | Video Endoscopic System
6 | RT2-14-006 | Shear Wave Elastrograph

## Laparascope, Mid-range
RT2-14-001 | Video System, Laparoscopic (Mid-Range) | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
1 | DESCRIPTION
1.1 | A medical equipment that provides visualization during minimally invasive surgical procedures.
2 | Product Offered:
2.1 | Make:
2.2 | Model:
2.3 | Initial Manufacturing year of Model:
3 | TECHNICAL SPECIFICATION
3.1 | The stack system must come complete with monitor, video processor, insufflator, camera console, data management system, laparoscopic rigid scopes, fiber optic light cable and the appropriate trolley.
3.2 | Video Processor
3.2.1 | Make:
3.2.2 | Model:
3.2.3 | Must have digital zoom facility and white balancing
3.2.4 | Video control unit must have 1 or more camera head inputs with image display settings available
3.2.5 | The unit offered must have visual output configurations of the latest technology. Bidder to specify
3.3 | Camera Head
3.3.1 | Make:
3.3.2 | Model:
3.3.3 | It must be controlled on camera console and head, with programmable buttons for capture, zoom and white balancing.
3.3.4 | Latest available HD Camera Head with an easy-to-fit scope mounting design must be provided.
3.3.5 | Camera head must have a minimum of three programmable switches to allow user remote control
3.3.6 | It must have capturing facility (Programmable Button)
3.3.7 | White balancing function must be available on the camera
3.3.8 | The length of the camera head cable must be between 3m and 5m
3.3.9 | It must be autoclavable or other means of sterilization, please state the method.
3.4 | CO2 Insufflator
3.4.1 | Make:
3.4.2 | Model:
3.4.3 | Insufflator must have both low and high flow capacity
3.4.4 | The unit must have display to indicate pressure in abdomen
3.4.5 | Visual and audible alarms are required to indicate overpressure.
3.4.6 | There must be an indication of the gas level in the cylinder.
3.4.7 | Automatic pressure relief mechanism when abdominal pressure exceeds set pressure
3.4.8 | Insufflator with smoke evacuation is desirable
3.5 | Light Source
3.5.1 | Make:
3.5.2 | Model:
3.5.3 | Lamp design must be LED.
3.5.4 | Light source must have adjustable light intensity
3.5.5 | An autoclavable fibre optic cable of at least 5mm diameter and 2.3m length must be provided.
3.6 | Monitor
3.6.1 | Make:
3.6.2 | Model:
3.6.3 | It must be at least 26" LED medical grade monitor with high image quality.
3.6.4 | Must be a touch screen monitor or operated with a separate keypad
3.6.5 | The trolley must have antistatic castors with at least 2 equipped with locking system/brake.
3.6.6 | The unit offered must have a minimum of 3 inputs configurations of the latest technology. Bidder to specify
3.6.7 | High contrast and wide viewing angle must be available
3.7 | Data Management System
3.7.1 | Make:
3.7.2 | Model:
3.7.3 | The unit offered must provide both video and still images recording in Full High Definition
3.7.4 | Internal memory must be a minimum of 1TB
3.7.5 | It must be a two-channel simultaneous recording
3.7.6 | Scope button to record and capture functions must be available
3.7.7 | The system offered must be provided with a compatible printer.
3.8 | Printer
3.8.1 | Make:
3.8.2 | Model:
3.8.3 | A4 or A6 colour printer with a minimum resolution of 423dpi must be provided.
3.9 | Trolley
3.9.1 | The offered unit must come in a purpose built Endoscopic System trolley with a minimum of one drawer for storage
3.9.2 | The trolley offered must have a tray for the keyboard and mouse.
3.9.3 | The trolley must have antistatic castors with at least 2 equipped with locking system/brake.
3.9.4 | The trolley must have a handle for ease access and transportation and a cable management arm.
3.9.5 | The trolley must have a cover at the rear for safety of cables.
3.9.6 | Trolley must have a secure gas cylinder holder
3.9.7 | The trolley must be supplied with a minimum of 3 additional 220 V power sockets.
4 | Standard accessories and consumable
4.1 | 3 x Rigid scopes (End-user's choice of size)
4.2 | Printing paper (1x Box)
4.3 | Disposable Insufflation Tubing Set (x10)
4.4 | Fiber Optical Light Guide Cable
4.5 | Universal wrench
4.6 | Urology Camera Head or as per user requirement.
5 | Conformity Compliance (Please attach certificate)
5.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
5.2 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
5.2.1 | Please provide unique ref number of the ISO 13485 ceriticate:
5.3 | Model on offer must comply with EU Medical Device Regulation (EU) 2017/745 or equvalent from SAHPRA recognized jurisdictions . Attach a copy of certification
5.3.1 | Please provide unique ref number of the ceriticate:
5.4 | Hazardous Substance Act: If this type of equipment / apparatus appears on the schedule of Hazardous substances issued by the Directorate: Health Technology of the Department of Health, a license in terms of the Act on Hazardous Substances (Act 15/1973) must be submitted with the bid / quotation documents. The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.5 | All electrical/electronic medical equipment must be licensed by Radiation Control, where applicable a copy of the license must be submitted The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.5.1 | Radiation Control License number:
5.6 | The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade (rubber to with 2 screws) type and it must be a minimum length of three (3) meters. NB the mains cable of the unit tendered for must be SABS colour coded.
5.7 | The equipment quoted for must be protected against electromagnetic interference IEC 60601-1-2
5.8 | The unit must comply with electrical safety and performance standard such as IEC 60601-2-21, attached certification
5.8.1 | Please provide unique ref number of the IEC ceriticate:
6 | COMPREHENSIVE MAINTENANCE
6.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) should be included in the unit price as per the SCC document
6.2 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
6.3 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
6.4 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
6.5 | A 5-year service exchange for the laparoscopic rigid scopes must be quoted for separately
6.6 | The bidder to provide the OEM recommended service interval of the unit:
6.7 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
7 | General
7.1 | Has the product on offer or any of its components being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
7.2 | The unit on offer must be the latest technology.
7.3 | Initial Manufacturing year of Model:
RT2-14-001 | Video System, Laparoscopic Replacement Accessories and Consumables | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
RT2-14-001a | Universal wrench
RT2-14-001b | 3 x Rigid scopes (End-user's choice of size)
RT2-14-001c | Printing paper (1x Box)
RT2-14-001d | Disposable Insufflation Tubing Set (x10)
RT2-14-001e | Fiber Optical Light Guide Cable
RT2-14-001f | Monitor
RT2-14-001g | Monitor Holding Arm (For adding an additional Monitor

## Laparascope, High-End
RT2-14-02 | Video System, Laparoscopic (3d High-End) | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
1 | DESCRIPTION
1.1 | A medical equipment that provides visualization during minimally invasive surgical procedures.
2 | Product Offered:
2.1 | Make:
2.2 | Model:
2.3 | Initial Manufacturing year of Model:
3 | TECHNICAL SPECIFICATION
3.1 | The stack system must come complete with monitor, video processor, insufflator, camera console, data management system, laparoscopic rigid scopes, fiber optic light cable and the appropriate trolley.
3.2 | Video Processor
3.2.1 | Make & Model:
3.2.2 | It must have digital zoom facility and automatic white balancing
3.2.3 | Video control unit must have 1 or more camera head inputs with image display settings available
3.2.4 | The unit offered must have visual output configurations of the latest technology. Bidder to specify
3.3 | Camera Head
3.3.1 | Make & Model:
3.3.2 | It must be controlled on camera console and head, with programmable buttons for capture, zoom and white balancing.
3.3.3 | A 4K 3D Camera Head with a easy-to-fit scope mounting design must be provided.
3.3.4 | Camera head must have a minimum of three programmable switches to allow user remote control
3.3.5 | Freezing facility (Programmable Button)must be available
3.3.6 | White balancing function must be available on the camera
3.3.7 | The length of the camera head cable must be between 3m and 5 m
3.3.8 | It must be autoclavable or other means of sterilization, please state the method
3.4 | CO2 Insufflator
3.4.1 | Make & Model:
3.4.2 | Insufflator must have both low and high flow capacity
3.4.3 | The unit must have display to indicate pressure in abdomen
3.4.4 | Visual and audible alarms are required to indicate overpressure.
3.4.5 | There must be an indication of the gas level in the cylinder.
3.4.6 | Automatic pressure relief mechanism when abdominal pressure exceeds set pressure
3.4.7 | Insufflator with smoke evacuation is desirable
3.5 | Light Source
3.5.1 | Make & Model:
3.5.2 | Lamp design must be LED.
3.5.3 | Light source must have adjustable light intensity
3.5.4 | An autoclavable fibre optic light guide cable of at least 5mm diameter and 2.3m length must be provided.
3.6 | Monitor
3.6.1 | Make & Model:
3.6.2 | It must be at least 30" 4K 3D LED medical grade monitor with high image quality.
3.6.3 | Must be a touch screen monitor or operated with a separate keypad
3.6.4 | The trolley must have antistatic castors with at least 2 equipped with locking system/brake.
3.6.5 | The unit offered must have a minimum of 3 inputs configurations of the latest technology. Bidder to specify
3.6.6 | High contrast and wide viewing angle
3.7 | Data Management System
3.7.1 | Make & Model:
3.7.2 | The unit offered must provide both video and still images recording in Full High Definition
3.7.3 | Internal memory must be a minimum of 1TB
3.7.4 | It must be a two-channel simultaneous recording
3.7.5 | Scope button to record and capture functions must be available
3.7.6 | The system offered must be provided with a compatible printer.
3.8 | Printer
3.8.1 | Make & Model:
3.8.2 | A4 or A6 colour printer with a minimum resolution of 423dpi must be provided.
3.9 | Trolley
3.9.1 | The offered unit must come in a purpose built Endoscopic System trolley with a minimum of one drawer for storage
3.9.2 | The trolley offered must have a tray for the keyboard and mouse.
3.9.3 | The trolley must have antistatic castors with at least 2 equipped with locking system/brake.
3.9.4 | The trolley must have a handle for ease access and transportation and a cable management arm.
3.9.5 | The trolley must have a cover at the rear for safety of cables.
3.9.6 | Trolley must have a secure gas cylinder holder
3.9.7 | The trolley must be supplied with a minimum of 3 additional 220 V power sockets apart from connected listed components.
4 | Standard accessories and consumable
4.1 | 3 x Rigid scopes (End-user's choice of size)
4.2 | Printing paper (1x Box)
4.3 | Disposable Insufflation Tubing Set (x10)
4.4 | Fiber Optical Light Guide Cable
4.5 | Universal wrench
5 | Conformity Compliance (Please attach certificate)
5.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
5.2 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
5.2.1 | Please provide unique ref number of the ISO 13485 ceriticate:
5.3 | Model on offer must comply with EU Medical Device Regulation (EU) 2017/745 or equvalent from SAHPRA recognized jurisdictions . Attach a copy of certification
5.3.1 | Please provide unique ref number of the ceriticate:
5.4 | Hazardous Substance Act: If this type of equipment / apparatus appears on the schedule of Hazardous substances issued by the Directorate: Health Technology of the Department of Health, a license in terms of the Act on Hazardous Substances (Act 15/1973) must be submitted with the bid / quotation documents. The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.5 | All electrical/electronic medical equipment must be licensed by Radiation Control, where applicable a copy of the license must be submitted The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.5.1 | Radiation Control License number:
5.6 | The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade (rubber to with 2 screws) type and it must be a minimum length of three (3) meters. NB the mains cable of the unit tendered for must be SABS colour coded.
5.7 | The equipment quoted for must be protected against electromagnetic interference IEC 60601-1-2
5.8 | The unit must comply with electrical safety and performance standard such as IEC 60601-2-21, attached certification
5.8.1 | Please provide unique ref number of the IEC ceriticate:
6 | COMPREHENSIVE MAINTENANCE
6.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) should be included in the unit price as per the SCC document
6.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
6.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
6.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
6.3 | A 5-year service exchange for the laparoscopic rigid scopes must be quoted for separately
6.4 | The bidder to provide the OEM recommended service interval of the unit:
6.5 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
7 | General
7.1 | Has the product on offer or any of its components being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
7.2 | The unit on offer must be the latest technology.
7.2.1 | Initial Manufacturing year of Model:
RT2-14-002 | Video System, Laparoscopic (3d High-End) Replacement accessories and consumables | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
RT2-14-002a | 3 x Rigid scopes (End-user's choice of size)
RT2-14-002b | Printing paper (1x Box)
RT2-14-002c | Disposable Insufflation Tubing Set (x10)
RT2-14-002d | Fiber Optical Light Guide Cable
RT2-14-002e | Universal wrench
RT2-14-002f | LAP CHOLE SET | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
12.1.1 | Atraumatic Dissecting and Grasping Forceps, rotating, with connector pin for unipolar coagulation, size 5 mm, length 36 cm, atraumatic, fenestrated, double action jaws, consisting of: Plastic Handle with disengageable ratchet, with larger contact area, Outer Tube insulated and Forceps Insert X2
12.1.2 | Coagulating and Dissecting Electrode, L-shaped, with connector pin for unipolar coagulation, size 5 mm, working length 36 cm X1
12.1.3 | Claw Forceps, rotating, size 5mm, length 36 cm, 2 x 3 teeth, single action jaws, consisting of handle with ratchet, Outer Tube, insulated, and Forceps Insert X1
12.1.4 | Suction and Irrigation Tube,with lateral holes, size 5 mm, length 36 cm, for use with suction and irrigation handles X1
12.1.5 | Suction irrigation handle X1
12.1.6 | Fascial Closure Instrument for subcutaneous ligature of trocar incisions, size 2.8 mm, length 17 cm X1
12.1.7 | Unipolar High Frequency Cord, with 8 mm plug, length 300 cm X1
12.1.8 | Curved long, Dissecting and Grasping Forceps, rotating, dismantling, insulated, with connector pin for unipolar coagulation, with LUER-Lock connector for cleaning, double action jaws, long, size 5 mm, length 36 cm, consisting of: Plastic Handle, without ratchet Metal Outer Sheath, insulated and Forceps Insert X1 OR
12.1.9 | Metzenbaum Scissors, rotating, dismantling, with connector pin for unipolar coagulation, with LUER-Lock irrigation connector for cleaning, double action jaws, spoon-shaped jaws, serrated, curved, length of jaws 20 mm, size 5 mm, length 36 cm consisting of: Plastic Handle, without ratchet, Metal Outer Sheath, insulated. and Scissors Insert, for use with trocars size 6 mm X1
RT2-14-002g | LAP NISSEN SET | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
12.2.1 | Atraumatic Dissecting and Grasping Forceps, rotating, with connector pin for unipolar coagulation, size 5 mm, length 36 cm, atraumatic, fenestrated, double action jaws, consisting of: Plastic Handle with disengageable ratchet, with larger contact area, Outer Tube insulated and Forceps Insert X2
12.2.2 | Curved long, Dissecting and Grasping Forceps, rotating, dismantling, insulated, with connector pin for unipolar coagulation, with LUER-Lock connector for cleaning, double action jaws, long, size 5 mm, length 36 cm, consisting of: Plastic Handle, without ratchet Metal Outer Sheath, insulated and Forceps Insert X1
12.2.3 | Metzenbaum Scissors, rotating, dismantling, with connector pin for unipolar coagulation, with LUER-Lock irrigation connector for cleaning, double action jaws, spoon-shaped jaws, serrated, curved, length of jaws 20 mm, size 5 mm, length 36 cm consisting of: Plastic Handle, without ratchet, Metal Outer Sheath, insulated. and Scissors Insert, for use with trocars size 6 mm X1
12.2.4 | Bowel Grasper, fenestrated, double action jaws, size 5mm,length 36 cm X1
12.2.5 | Liver retractor X1
12.2.6 | Suction and Irrigation Tube,with lateral holes, size 5 mm, length 36 cm, for use with suction and irrigation handles X1
12.2.7 | Suction irrigation handle X1
12.2.8 | Fascial Closure Instrument for subcutaneous ligature of trocar incisions, size 2.8 mm, length 17 cm X1
12.2.9 | Unipolar High Frequency Cord, with 8 mm plug, length 300 cm X1
RT2-14-002h | LAP HERNIA SET | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
12.3.1 | Atraumatic Grasping Forceps rotating, dismantling, insulated, with connector pin for unipolar coagulation, single action jaws, with especially fine serration, fenestrated, size 5 mm, length 36 cm consisting of:PlasticHandle,one with and one without ratchet, Metal OuterSheath, insulated and Forceps Insert X2
12.3.2 | Blunt Nose Dissecting and Grasping Forceps, rotating, with connector pin for unipolar coagulation, size 5 mm, length 36 cm, heavy, double action jaws, consisting of: Plastic Handle, without ratchet, Outer Tube, insulated and Forceps Insert X1
12.3.3 | BABCOCK Clamp, rotating, size 10 mm, length 36 cm, double action jaws, consisting of: Metal Handle, with hemostat style ratchet, Outer Tube, insulated and Forceps Insert X1
12.3.4 | Right Angled Dissecting Forceps, rotating, size 5mm , length 36 cm, 2 x 3 teeth, single action jaws, consisting of: Metal Handle, Outer Tube, insulated and Forceps Insert X1
12.3.5 | METZENBAUM Scissors, rotating, dismantling, insulated, with connector pin for unipolar coagulation, with LUER-Lock connector for cleaning, double action jaws, curved, length of blades 15 mm, size 5 mm, length 36 cm consisting of: Plastic Handle, insulated, without ratchet Metal Outer Sheath, insulated and Scissors Insert for use with trocars size 6 mm X1
12.3.6 | Curved Bipolar Grasping Forceps, rotating, dismantling, with connector pin for bipolar coagulation, especially suitable for dissection, double action jaws, size 5 mm, length 36 cm, consisting of:Plastic Handle, without ratchet, Outer Sheath and Forceps Insert X1
12.3.7 | Coagulating and Dissecting Electrode, with channel, L-shaped size 5 mm, length 36 cm X1
12.3.8 | Suction and Irrigation Tube, with lateral holes, with two-way stopcock for single-hand control, size 5 mm, length 36 cm X1
12.3.9 | Suction and Irrigation Handle X1
12.3.10 | Knot Tier, size 5 mm, length 36 cm, for extracorporeal knotting X1
12.3.11 | Fascial Closure Instrument for subcutaneous ligature of trocar incisions, size 2.8 mm, length 17 cm X1
12.3.12 | Unipolar High Frequency Cord, with 4 mm plug, length 300 cm X1
12.3.13 | Bipolar Cable X1
RT2-14-002i | COLORECTAL SET | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
12.4.1 | Atraumatic Grasping Forceps rotating, dismantling, insulated, with connector pin for unipolar coagulation, single action jaws, with especially fine serration, fenestrated, size 5 mm, length 36 cm consisting of:PlasticHandle, without ratchet, Metal OuterSheath, insulated and Forceps Insert X2
12.4.2 | Atraumatic Grasping Forceps, rotating, with connector pin for unipolar coagulation, size 5 mm, length 36 cm, atraumatic, fenestrated, single action jaws, consisting of: Plastic Handle with disengageable ratchet, with larger contact area, Outer Tube, insulated and Forceps Insert X1
12.4.3 | Atraumatic Dissecting and Grasping Forceps, rotating, with connector pin for unipolar coagulation, size 5 mm, length 36 cm, atraumatic, fenestrated, double action jaws, consisting of: Plastic Handle, with larger contact area, Outer Tube, insulated and Forceps Insert X1
12.4.4 | Atraumatic Grasping Forceps, rotating, size 5 mm, length 36 cm, double action,consisting of: Plastic-Handle with disengageable ratchet, with larger contact area, Outer Tube, insulated and Forceps Insert X1
12.4.5 | BABCOCK Grasping Forceps, rotating, with connector pin for unipolar coagulation, size 10mm, length 36 cm, atraumatic, jaws with multiple teeth, fenestrated, single action jaws, consisting of: Plastic Handle with disengageable ratchet, with larger contact area, Outer Tube, insulated and Forceps Insert X1
12.4.6 | Right Angled Dissecting Forceps, rotating, with connector pin for unipolar coagulation, size 5 mm, length 36 cm, atraumatic, jaws with multiple teeth, fenestrated, single action jaws, consisting of: Plastic Handle with disengageable ratchet, with larger contact area, Outer Tube, insulated and Forceps Insert X1
12.4.7 | BABCOCK Clamp, rotating, size 10 mm, length 36 cm, double action jaws, consisting of: Metal Handle, with hemostat style ratchet, Outer Tube, insulated and Forceps Insert X1
12.4.8 | Claw Forceps, rotating, size 10 mm, length 36 cm, 2 x 3 teeth, single action jaws, consisting of: Metal Handle, with hemostat style ratchet, Outer Tube, insulated and Forceps Insert X1
12.4.9 | METZENBAUM Scissors, rotating, dismantling, insulated, with connector pin for unipolar coagulation, with LUER-Lock connector for cleaning, double action jaws, curved, length of blades 15 mm, size 5 mm, length 36 cm consisting of: Plastic Handle, insulated, without ratchet Metal Outer Sheath, insulated and Scissors Insert for use with trocars size 6 mm X1
12.4.10 | METZENBAUM Scissors, curved, length of blades, double action jaws, size 5 mm, length 36 cm X1
12.4.11 | Bowel Grasper, fenestrated, double action jaws, length of jaws 18mm, size 5mm, length 36 cm X2
12.4.12 | MANHES Grasping Forceps,rotating, dismantling, insulated, with connector pin for unipolar coagulation, with LUER-Lock irrigation connector for cleaning, single action jaws, with multiple teeth, width of jaws 4.8 mm, for atraumatic and accurate grasping, size 5 mm,length 36 cm,consisting of Plastic Handle, with ratchet, with larger contact area,Metal Outer Sheath,Forceps Insert X1
12.4.12 | Knot Tier, size 5 mm, length 36 cm, for extracorporeal knotting X1
12.4.13 | Fascial Closure Instrument for subcutaneous ligature of trocar incisions, size 2.8 mm, length 17 cm X1
12.4.14 | Unipolar High Frequency Cord, with 4 mm plug, length 300 cm, X1
12.4.15 | Bipolar Cable X2
RT2-14-002j | GYNAE SET | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
12.18 | VERESS Pneumoperitoneum needle, with spring -loaded blunt stylet, LUER-Lock, length 10cm X1
12.19 | Trocar and Cannula 6mm X3
12.2 | Trocar and Cannula 11mm X2
12.21 | Reducer 11mm to 5mm X1
12.22 | Curved long, Dissecting and Grasping Forceps, rotating, dismantling, insulated, with connector pin for unipolar coagulation, with LUER-Lock connector for cleaning, double action jaws, long, size 5 mm, length 36 cm, consisting of: Plastic Handle, without ratchet Metal Outer Sheath, insulated and Forceps Insert X1
12.23 | Curved short, Dissecting and Grasping Forceps, rotating, dismantling, insulated, with connector pin for unipolar coagulation, with LUER-Lock connector for cleaning, double action jaws, long, size 5 mm, length 36 cm, consisting of: Plastic Handle, without ratchet, Metal Outer Sheath, insulated and Forceps Insert X1
12.24 | Atraumatic Grasping Forceps rotating, dismantling, insulated, with connector pin for unipolar coagulation, single action jaws, with especially fine serration, fenestrated, size 5 mm, length 36 cm consisting of:PlasticHandle, without ratchet, Metal OuterSheath, insulated and Forceps Insert X1
12.25 | Atraumatic Dissecting and Grasping Forceps, rotating, with connector pin for unipolar coagulation, size 5 mm, length 36 cm, atraumatic, fenestrated, double action jaws, consisting of: Plastic Handle with disengageable ratchet, with larger contact area, Outer Tube insulated and Forceps Insert X1
12.26 | Lymph node Grasping Forceps, rotating, with connector pin for unipolar coagulation, size 5 mm, length 36 cm, atraumatic, single action jaws, consisting of: Plastic Handle with disengageable ratchet, with larger contact area Outer Tube, insulated, and Forceps Insert X1
12.27 | Grasping Forceps, rotating, with connector pin for unipolar coagulation, size 5 mm, length 36 cm, jaws with multiple teeth, width of jaws 4,8 mm, single action jaws, consisting of: 33156 Plastic-Handle with disengageable ratchet, with larger contact area 33300 Outer Tube, insulated 33310 ME Forceps Insert X1
12.28 | BABCOCK Grasping Forceps, rotating, with connector pin for unipolar coagulation, size 5 mm, length 36 cm, atraumatic, jaws with multiple teeth, fenestrated, single action jaws, consisting of: Plastic Handle with disengageable ratchet, with larger contact area, Outer Tube, insulated and Forceps Insert X1
12.29 | Claw Forceps, rotating, size 5mm, length 36 cm, 2 x 3 teeth, single action jaws, consisting of: Metal Handle, with hemostat style ratchet, Outer Tube, insulated, and Forceps Insert X1
12.3 | Metzenbaum Scissors, rotating, dismantling, with connector pin for unipolar coagulation, with LUER-Lock irrigation connector for cleaning, double action jaws, spoon-shaped jaws, serrated, curved, length of jaws 20 mm, size 5 mm, length 36 cm consisting of: Plastic Handle, without ratchet, Metal Outer Sheath, insulated. and Scissors Insert for use with trocars size 6 mm X1
12.31 | Biopsy Forceps, rotating, dismantling, insulated, with connector pin for unipolar coagulation, with LUER lock adaptor for cleaning, large jaws, single action jaws, size 5 mm, length 36 cm consisting of: Plastic Handle, without ratchet, Outer Sheath, insulated, and Forceps Insert X1
12.32 | Curved Bipolar Grasping Forceps, rotating, dismantling, with connector pin for bipolar coagulation, especially suitable for dissection, double action jaws, size 5 mm, length 36 cm, consisting of:Plastic Handle, without ratchet, Metal Outer Sheath and Forceps Insert X1
12.33 | Fenestrated Bipolar Grasping Forceps, rotating, dismantling, with connector pin for bipolar coagulation, with especially fine atraumatic serration, fenestrated jaws, double action jaws, size 5 mm, length 36 cm, consisting of: Plastic Handle, without ratchet Metal Outer Sheath and Forceps Insert X1
12.34 | Coagulating and Dissecting Electrode, spatula-sheaped, blunt with connector pin for unipolar coagulation, size 5 mm, working length 36 cm X1
12.35 | Coagulating and Dissecting Electrode, L-shaped, with connector pin for unipolar coagulation, size 5 mm, working length 36 cm X1
12.36 | Palpation Probe X1
12.37 | Myoma Fixation Instrument, screw shaped, ø 5 mm X1
12.38 | Suction and Irrigation Tube, with lateral holes, with two-way stopcock for single-hand control, size 5 mm, length 36 cm X1
12.39 | Needle Holder, ergonomic axial handle with disengageable ratchet, ratchet release on the right side, left curved jaws, with tungsten carbide insert ø 5 mm, length 33 cm X1
12.4 | Needle Holder, ergonomic axial handle with disengageable ratchet, ratchet release on the left side, right curved jaws, with tungsten carbide insert ø 5 mm, length 33 cm X1
12.41 | Knot Tier, size 5 mm, length 36 cm, for extracorporeal knotting X1
12.42 | Fascial Closure Instrument for subcutaneous ligature of trocar incisions, size 2.8 mm, length 17 cm X1
12.43 | Uterine manipulator X1
12.44 | Unipolar High Frequency Cord, with 8 mm plug, length 300 cm X2
12.45 | Bipolar High Frequency Cord, length 300 cm X2
12.46 | Fiber Optic Light Cable, with straight connector, diameter 4.8 mm, length 300 cm X1
12.47 | Container X1
RT2-14-002k | UROLOGY SET | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
11.6 | Curved long, Dissecting and Grasping Forceps, rotating, dismantling, insulated, with connector pin for unipolar coagulation, with LUER-Lock connector for cleaning, double action jaws, long, size 5 mm, length 36 cm, consisting of: Plastic Handle, without ratchet Metal Outer Sheath, insulated and Forceps Insert X1
11.7 | Curved short, Dissecting and Grasping Forceps, rotating, dismantling, insulated, with connector pin for unipolar coagulation, with LUER-Lock connector for cleaning, double action jaws, long, size 5 mm, length 36 cm, consisting of: Plastic Handle, without ratchet, Metal Outer Sheath, insulated and Forceps Insert X1
11.8 | Atraumatic Grasping Forceps rotating, dismantling, insulated, with connector pin for unipolar coagulation, single action jaws, with especially fine serration, fenestrated, size 5 mm, length 36 cm consisting of:PlasticHandle, without ratchet, Metal OuterSheath, insulated and Forceps Insert X1
11.9 | Atraumatic Grasping Forceps, rotating, with connector pin for unipolar coagulation, size 5 mm, length 36 cm, atraumatic, fenestrated, single action jaws, consisting of: Plastic Handle with disengageable ratchet, with larger contact area, Outer Tube, insulated and Forceps Insert X1
11.1 | Atraumatic Dissecting and Grasping Forceps, rotating, with connector pin for unipolar coagulation, size 5 mm, length 36 cm, atraumatic, fenestrated, double action jaws, consisting of: Plastic Handle, with larger contact area, Outer Tube, insulated and Forceps Insert X1
11.11 | Atraumatic Dissecting and Grasping Forceps, rotating, with connector pin for unipolar coagulation, size 5 mm, length 36 cm, atraumatic, fenestrated, double action jaws, consisting of: Plastic Handle with disengageable ratchet, with larger contact area, Outer Tube, insulated and Forceps Insert X1
11.12 | Grasping Forceps, rotating, with connector pin for unipolar coagulation, size 5 mm, length 36 cm, "Tiger-jaws", 2x4 teeth, single action jaws, consisting of: Plastic-Handle with disengageable ratchet, with larger contact area, Outer Tube, insulated and Forceps Insert X1
11.13 | Blunt Nose Dissecting and Grasping Forceps, rotating, with connector pin for unipolar coagulation, size 5 mm, length 36 cm, heavy, double action jaws, consisting of: Plastic Handle, without ratchet, Outer Tube, insulated and Forceps Insert X1
11.14 | BABCOCK Grasping Forceps, rotating, with connector pin for unipolar coagulation, size 5 mm, length 36 cm, atraumatic, jaws with multiple teeth, fenestrated, single action jaws, consisting of: Plastic Handle with disengageable ratchet, with larger contact area, Outer Tube, insulated and Forceps Insert X1
11.15 | Right Angled Dissecting Forceps, rotating, with connector pin for unipolar coagulation, size 5 mm, length 36 cm, atraumatic, jaws with multiple teeth, fenestrated, single action jaws, consisting of: Plastic Handle with disengageable ratchet, with larger contact area, Outer Tube, insulated and Forceps Insert X1
11.16 | BABCOCK Clamp, rotating, size 10 mm, length 36 cm, double action jaws, consisting of: Metal Handle, with hemostat style ratchet, Outer Tube, insulated and Forceps Insert X1
11.17 | Claw Forceps, rotating, size 10 mm, length 36 cm, 2 x 3 teeth, single action jaws, consisting of: Metal Handle, with hemostat style ratchet, Outer Tube, insulated and Forceps Insert X1
11.18 | Right Angled Dissecting Forceps, rotating, size 10 mm, length 36 cm, 2 x 3 teeth, single action jaws, consisting of: Metal Handle, Outer Tube, insulated and Forceps Insert X1
11.19 | METZENBAUM Scissors, rotating, dismantling, insulated, with connector pin for unipolar coagulation, with LUER-Lock connector for cleaning, double action jaws, curved, length of blades 15 mm, size 5 mm, length 36 cm consisting of: Plastic Handle, insulated, without ratchet Metal Outer Sheath, insulated and Scissors Insert for use with trocars size 6 mm X1
11.2 | Scissors Insert, METZENBAUM Scissors, curved, length of blades 12 mm, double action jaws, size 5 mm, length 36 cm X1
11.21 | Straight Scissors, rotating, with connector pin for unipolar coagulation, size 5 mm, length 36 cm, straight, double action jaws, consisting of: Plastic Handle, without ratchet, Metal Outer Sheath, Scissors Insert X1
11.22 | Click Line Scissors Insert, straight, double action jaws, size 5 mm, length 36 cm X1
11.23 | Curved Bipolar Grasping Forceps, rotating, dismantling, with connector pin for bipolar coagulation, especially suitable for dissection, double action jaws, size 5 mm, length 36 cm, consisting of:Plastic Handle, without ratchet, Metal Outer Sheath and Forceps Insert X1
11.24 | Fenestrated Bipolar Grasping Forceps, rotating, dismantling, with connector pin for bipolar coagulation, with especially fine atraumatic serration, fenestrated jaws, double action jaws, size 5 mm, length 36 cm, consisting of: Plastic Handle, without ratchet Metal X1Outer Sheath and Forceps Insert X1
11.25 | Coagulating and Dissecting Electrode, with channel, L-shaped size 5 mm, length 36 cm, for use with suction and irrigation handles X1
11.26 | Handle with Trumpet Valve, for suction or irrigation, autoclavable, for use with coagulation suction tubes size 5 mm X1
11.27 | Suction and Irrigation Tube, with lateral holes, with two-way stopcock for single-hand control, size 5 mm, length 36 cm X1
11.28 | Puncture Needle, LUER-lock, diameter 1.6 mm, size 5 mm, length 36 cm X1
11.29 | Knot Tier, size 5 mm, length 36 cm, for extracorporeal knotting X1
11.3 | Fascial Closure Instrument for subcutaneous ligature of trocar incisions, size 2.8 mm, length 17 cm X1
11.31 | Unipolar High Frequency Cord, with 4 mm plug, length 300 cm X2
11.32 | Bipolar Cable X1
11.33 | Plastic Container for Sterilizing and Storage, perforated, with lid, for two-level storage, for use with 30 cm and 36 cm Click Line- insruments, external dimensions ( w x d x h ): 600 mm x 145 mm x 260 mm ( X1)

## Laparascope Instruments
RT2-14-003 | Lapscope Instruments | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
RT2-14-003a | Straight Forward Telescope 0° enlarged view, diameter 10 mm, length 29-31 cm, autoclavable, fiber optic light transmission incorporated
RT2-14-003b | Straight Forward Telescope 0° enlarged view, diameter 5 mm, length 29-31 cm, autoclavable, fiber optic light transmission incorporated
RT2-14-003c | Straight Forward Telescope 0° enlarged view, diameter 2.5 mm, length 29-31 cm, autoclavable, fiber optic light transmission incorporated
RT2-14-003d | Forward-Oblique Telescope 30° enlarged view, diameter 10 mm, length 29-31 cm, autoclavable, fiber optic light transmission incorporated
RT2-14-003e | Forward-Oblique Telescope 30° enlarged view, diameter 5 mm, length 29-31 cm, autoclavable, fiber optic light transmission incorporated
RT2-14-03f | Forward-Oblique Telescope 30° enlarged view, diameter 2.5 mm, length 29-31 cm, autoclavable, fiber optic light transmission incorporated
RT2-14-003g | Plastic Container for Sterilizing and Storage, perforated, with lid, for two-level storage, for use with 30 cm and 36 cm intruments , external dimensions ( w x d x h ): 600 mm x 145 mm x 260 mm
RT2-14-003h | Laparoscopic Needle Holder
RT2-14-003i | VERESS Pneumoperitoneum needle, with spring -loaded blunt stylet, LUER-Lock, length 10cm
RT2-14-003j | Trocar size, 11mm,with insufflation stopcock, Cannula with thread, length 100mm(max 110mm)
RT2-14-003k | Trocar size, 3.5mm,with insufflation stopcock, Cannula with thread, length 100mm(max 110mm)
RT2-14-003l | Trocar size, 6mm,with insufflation stopcock, Cannula with thread, length 100mm(max 110mm)
RT2-14-003m | LAP CHOLE SET | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
12.1.1 | Atraumatic Dissecting and Grasping Forceps, rotating, with connector pin for unipolar coagulation, size 5 mm, length 36 cm, atraumatic, fenestrated, double action jaws, consisting of: Plastic Handle with disengageable ratchet, with larger contact area, Outer Tube insulated and Forceps Insert X2
12.1.2 | Coagulating and Dissecting Electrode, L-shaped, with connector pin for unipolar coagulation, size 5 mm, working length 36 cm X1
12.1.3 | Claw Forceps, rotating, size 5mm, length 36 cm, 2 x 3 teeth, single action jaws, consisting of handle with ratchet, Outer Tube, insulated, and Forceps Insert X1
12.1.4 | Suction and Irrigation Tube,with lateral holes, size 5 mm, length 36 cm, for use with suction and irrigation handles X1
12.1.5 | Suction irrigation handle X1
12.1.6 | Fascial Closure Instrument for subcutaneous ligature of trocar incisions, size 2.8 mm, length 17 cm X1
12.1.7 | Unipolar High Frequency Cord, with 8 mm plug, length 300 cm X1
12.1.8 | Curved long, Dissecting and Grasping Forceps, rotating, dismantling, insulated, with connector pin for unipolar coagulation, with LUER-Lock connector for cleaning, double action jaws, long, size 5 mm, length 36 cm, consisting of: Plastic Handle, without ratchet Metal Outer Sheath, insulated and Forceps Insert X1 OR
12.1.9 | Metzenbaum Scissors, rotating, dismantling, with connector pin for unipolar coagulation, with LUER-Lock irrigation connector for cleaning, double action jaws, spoon-shaped jaws, serrated, curved, length of jaws 20 mm, size 5 mm, length 36 cm consisting of: Plastic Handle, without ratchet, Metal Outer Sheath, insulated. and Scissors Insert, for use with trocars size 6 mm X1
RT2-14-003n | LAP NISSEN SET | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
12.2.1 | Atraumatic Dissecting and Grasping Forceps, rotating, with connector pin for unipolar coagulation, size 5 mm, length 36 cm, atraumatic, fenestrated, double action jaws, consisting of: Plastic Handle with disengageable ratchet, with larger contact area, Outer Tube insulated and Forceps Insert X2
12.2.2 | Curved long, Dissecting and Grasping Forceps, rotating, dismantling, insulated, with connector pin for unipolar coagulation, with LUER-Lock connector for cleaning, double action jaws, long, size 5 mm, length 36 cm, consisting of: Plastic Handle, without ratchet Metal Outer Sheath, insulated and Forceps Insert X1
12.2.3 | Metzenbaum Scissors, rotating, dismantling, with connector pin for unipolar coagulation, with LUER-Lock irrigation connector for cleaning, double action jaws, spoon-shaped jaws, serrated, curved, length of jaws 20 mm, size 5 mm, length 36 cm consisting of: Plastic Handle, without ratchet, Metal Outer Sheath, insulated. and Scissors Insert, for use with trocars size 6 mm X1
12.2.4 | Bowel Grasper, fenestrated, double action jaws, size 5mm,length 36 cm X1
12.2.5 | Liver retractor X1
12.2.6 | Suction and Irrigation Tube,with lateral holes, size 5 mm, length 36 cm, for use with suction and irrigation handles X1
12.2.7 | Suction irrigation handle X1
12.2.8 | Fascial Closure Instrument for subcutaneous ligature of trocar incisions, size 2.8 mm, length 17 cm X1
12.2.9 | Unipolar High Frequency Cord, with 8 mm plug, length 300 cm X1
RT2-14-003o | LAP HERNIA SET | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
12.3.1 | Atraumatic Grasping Forceps rotating, dismantling, insulated, with connector pin for unipolar coagulation, single action jaws, with especially fine serration, fenestrated, size 5 mm, length 36 cm consisting of:PlasticHandle,one with and one without ratchet, Metal OuterSheath, insulated and Forceps Insert X2
12.3.2 | Blunt Nose Dissecting and Grasping Forceps, rotating, with connector pin for unipolar coagulation, size 5 mm, length 36 cm, heavy, double action jaws, consisting of: Plastic Handle, without ratchet, Outer Tube, insulated and Forceps Insert X1
12.3.3 | BABCOCK Clamp, rotating, size 10 mm, length 36 cm, double action jaws, consisting of: Metal Handle, with hemostat style ratchet, Outer Tube, insulated and Forceps Insert X1
12.3.4 | Right Angled Dissecting Forceps, rotating, size 5mm , length 36 cm, 2 x 3 teeth, single action jaws, consisting of: Metal Handle, Outer Tube, insulated and Forceps Insert X1
12.3.5 | METZENBAUM Scissors, rotating, dismantling, insulated, with connector pin for unipolar coagulation, with LUER-Lock connector for cleaning, double action jaws, curved, length of blades 15 mm, size 5 mm, length 36 cm consisting of: Plastic Handle, insulated, without ratchet Metal Outer Sheath, insulated and Scissors Insert for use with trocars size 6 mm X1
12.3.6 | Curved Bipolar Grasping Forceps, rotating, dismantling, with connector pin for bipolar coagulation, especially suitable for dissection, double action jaws, size 5 mm, length 36 cm, consisting of:Plastic Handle, without ratchet, Outer Sheath and Forceps Insert X1
12.3.7 | Coagulating and Dissecting Electrode, with channel, L-shaped size 5 mm, length 36 cm X1
12.3.8 | Suction and Irrigation Tube, with lateral holes, with two-way stopcock for single-hand control, size 5 mm, length 36 cm X1
12.3.9 | Suction and Irrigation Handle X1
12.3.10 | Knot Tier, size 5 mm, length 36 cm, for extracorporeal knotting X1
12.3.11 | Fascial Closure Instrument for subcutaneous ligature of trocar incisions, size 2.8 mm, length 17 cm X1
12.3.12 | Unipolar High Frequency Cord, with 4 mm plug, length 300 cm X1
12.3.13 | Bipolar Cable X1
RT2-14-003p | COLORECTAL SET | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
12.4.1 | Atraumatic Grasping Forceps rotating, dismantling, insulated, with connector pin for unipolar coagulation, single action jaws, with especially fine serration, fenestrated, size 5 mm, length 36 cm consisting of:PlasticHandle, without ratchet, Metal OuterSheath, insulated and Forceps Insert X2
12.4.2 | Atraumatic Grasping Forceps, rotating, with connector pin for unipolar coagulation, size 5 mm, length 36 cm, atraumatic, fenestrated, single action jaws, consisting of: Plastic Handle with disengageable ratchet, with larger contact area, Outer Tube, insulated and Forceps Insert X1
12.4.3 | Atraumatic Dissecting and Grasping Forceps, rotating, with connector pin for unipolar coagulation, size 5 mm, length 36 cm, atraumatic, fenestrated, double action jaws, consisting of: Plastic Handle, with larger contact area, Outer Tube, insulated and Forceps Insert X1
12.4.4 | Atraumatic Grasping Forceps, rotating, size 5 mm, length 36 cm, double action,consisting of: Plastic-Handle with disengageable ratchet, with larger contact area, Outer Tube, insulated and Forceps Insert X1
12.4.5 | BABCOCK Grasping Forceps, rotating, with connector pin for unipolar coagulation, size 10mm, length 36 cm, atraumatic, jaws with multiple teeth, fenestrated, single action jaws, consisting of: Plastic Handle with disengageable ratchet, with larger contact area, Outer Tube, insulated and Forceps Insert X1
12.4.6 | Right Angled Dissecting Forceps, rotating, with connector pin for unipolar coagulation, size 5 mm, length 36 cm, atraumatic, jaws with multiple teeth, fenestrated, single action jaws, consisting of: Plastic Handle with disengageable ratchet, with larger contact area, Outer Tube, insulated and Forceps Insert X1
12.4.7 | BABCOCK Clamp, rotating, size 10 mm, length 36 cm, double action jaws, consisting of: Metal Handle, with hemostat style ratchet, Outer Tube, insulated and Forceps Insert X1
12.4.8 | Claw Forceps, rotating, size 10 mm, length 36 cm, 2 x 3 teeth, single action jaws, consisting of: Metal Handle, with hemostat style ratchet, Outer Tube, insulated and Forceps Insert X1
12.4.9 | METZENBAUM Scissors, rotating, dismantling, insulated, with connector pin for unipolar coagulation, with LUER-Lock connector for cleaning, double action jaws, curved, length of blades 15 mm, size 5 mm, length 36 cm consisting of: Plastic Handle, insulated, without ratchet Metal Outer Sheath, insulated and Scissors Insert for use with trocars size 6 mm X1
12.4.10 | METZENBAUM Scissors, curved, length of blades, double action jaws, size 5 mm, length 36 cm X1
12.4.11 | Bowel Grasper, fenestrated, double action jaws, length of jaws 18mm, size 5mm, length 36 cm X2
12.4.12 | MANHES Grasping Forceps,rotating, dismantling, insulated, with connector pin for unipolar coagulation, with LUER-Lock irrigation connector for cleaning, single action jaws, with multiple teeth, width of jaws 4.8 mm, for atraumatic and accurate grasping, size 5 mm,length 36 cm,consisting of Plastic Handle, with ratchet, with larger contact area,Metal Outer Sheath,Forceps Insert X1
12.4.12 | Knot Tier, size 5 mm, length 36 cm, for extracorporeal knotting X1
12.4.13 | Fascial Closure Instrument for subcutaneous ligature of trocar incisions, size 2.8 mm, length 17 cm X1
12.4.14 | Unipolar High Frequency Cord, with 4 mm plug, length 300 cm, X1
12.4.15 | Bipolar Cable X2
RT2-14-003q | GYNAE SET | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
12.18 | VERESS Pneumoperitoneum needle, with spring -loaded blunt stylet, LUER-Lock, length 10cm X1
12.19 | Trocar and Cannula 6mm X3
12.2 | Trocar and Cannula 11mm X2
12.21 | Reducer 11mm to 5mm X1
12.22 | Curved long, Dissecting and Grasping Forceps, rotating, dismantling, insulated, with connector pin for unipolar coagulation, with LUER-Lock connector for cleaning, double action jaws, long, size 5 mm, length 36 cm, consisting of: Plastic Handle, without ratchet Metal Outer Sheath, insulated and Forceps Insert X1
12.23 | Curved short, Dissecting and Grasping Forceps, rotating, dismantling, insulated, with connector pin for unipolar coagulation, with LUER-Lock connector for cleaning, double action jaws, long, size 5 mm, length 36 cm, consisting of: Plastic Handle, without ratchet, Metal Outer Sheath, insulated and Forceps Insert X1
12.24 | Atraumatic Grasping Forceps rotating, dismantling, insulated, with connector pin for unipolar coagulation, single action jaws, with especially fine serration, fenestrated, size 5 mm, length 36 cm consisting of:PlasticHandle, without ratchet, Metal OuterSheath, insulated and Forceps Insert X1
12.25 | Atraumatic Dissecting and Grasping Forceps, rotating, with connector pin for unipolar coagulation, size 5 mm, length 36 cm, atraumatic, fenestrated, double action jaws, consisting of: Plastic Handle with disengageable ratchet, with larger contact area, Outer Tube insulated and Forceps Insert X1
12.26 | Lymph node Grasping Forceps, rotating, with connector pin for unipolar coagulation, size 5 mm, length 36 cm, atraumatic, single action jaws, consisting of: Plastic Handle with disengageable ratchet, with larger contact area Outer Tube, insulated, and Forceps Insert X1
12.27 | Grasping Forceps, rotating, with connector pin for unipolar coagulation, size 5 mm, length 36 cm, jaws with multiple teeth, width of jaws 4,8 mm, single action jaws, consisting of: 33156 Plastic-Handle with disengageable ratchet, with larger contact area 33300 Outer Tube, insulated 33310 ME Forceps Insert X1
12.28 | BABCOCK Grasping Forceps, rotating, with connector pin for unipolar coagulation, size 5 mm, length 36 cm, atraumatic, jaws with multiple teeth, fenestrated, single action jaws, consisting of: Plastic Handle with disengageable ratchet, with larger contact area, Outer Tube, insulated and Forceps Insert X1
12.29 | Claw Forceps, rotating, size 5mm, length 36 cm, 2 x 3 teeth, single action jaws, consisting of: Metal Handle, with hemostat style ratchet, Outer Tube, insulated, and Forceps Insert X1
12.3 | Metzenbaum Scissors, rotating, dismantling, with connector pin for unipolar coagulation, with LUER-Lock irrigation connector for cleaning, double action jaws, spoon-shaped jaws, serrated, curved, length of jaws 20 mm, size 5 mm, length 36 cm consisting of: Plastic Handle, without ratchet, Metal Outer Sheath, insulated. and Scissors Insert for use with trocars size 6 mm X1
12.31 | Biopsy Forceps, rotating, dismantling, insulated, with connector pin for unipolar coagulation, with LUER lock adaptor for cleaning, large jaws, single action jaws, size 5 mm, length 36 cm consisting of: Plastic Handle, without ratchet, Outer Sheath, insulated, and Forceps Insert X1
12.32 | Curved Bipolar Grasping Forceps, rotating, dismantling, with connector pin for bipolar coagulation, especially suitable for dissection, double action jaws, size 5 mm, length 36 cm, consisting of:Plastic Handle, without ratchet, Metal Outer Sheath and Forceps Insert X1
12.33 | Fenestrated Bipolar Grasping Forceps, rotating, dismantling, with connector pin for bipolar coagulation, with especially fine atraumatic serration, fenestrated jaws, double action jaws, size 5 mm, length 36 cm, consisting of: Plastic Handle, without ratchet Metal Outer Sheath and Forceps Insert X1
12.34 | Coagulating and Dissecting Electrode, spatula-sheaped, blunt with connector pin for unipolar coagulation, size 5 mm, working length 36 cm X1
12.35 | Coagulating and Dissecting Electrode, L-shaped, with connector pin for unipolar coagulation, size 5 mm, working length 36 cm X1
12.36 | Palpation Probe X1
12.37 | Myoma Fixation Instrument, screw shaped, ø 5 mm X1
12.38 | Suction and Irrigation Tube, with lateral holes, with two-way stopcock for single-hand control, size 5 mm, length 36 cm X1
12.39 | Needle Holder, ergonomic axial handle with disengageable ratchet, ratchet release on the right side, left curved jaws, with tungsten carbide insert ø 5 mm, length 33 cm X1
12.4 | Needle Holder, ergonomic axial handle with disengageable ratchet, ratchet release on the left side, right curved jaws, with tungsten carbide insert ø 5 mm, length 33 cm X1
12.41 | Knot Tier, size 5 mm, length 36 cm, for extracorporeal knotting X1
12.42 | Fascial Closure Instrument for subcutaneous ligature of trocar incisions, size 2.8 mm, length 17 cm X1
12.43 | Uterine manipulator X1
12.44 | Unipolar High Frequency Cord, with 8 mm plug, length 300 cm X2
12.45 | Bipolar High Frequency Cord, length 300 cm X2
12.46 | Fiber Optic Light Cable, with straight connector, diameter 4.8 mm, length 300 cm X1
12.47 | Container X1
RT2-14-003r | UROLOGY SET | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
11.6 | Curved long, Dissecting and Grasping Forceps, rotating, dismantling, insulated, with connector pin for unipolar coagulation, with LUER-Lock connector for cleaning, double action jaws, long, size 5 mm, length 36 cm, consisting of: Plastic Handle, without ratchet Metal Outer Sheath, insulated and Forceps Insert X1
11.7 | Curved short, Dissecting and Grasping Forceps, rotating, dismantling, insulated, with connector pin for unipolar coagulation, with LUER-Lock connector for cleaning, double action jaws, long, size 5 mm, length 36 cm, consisting of: Plastic Handle, without ratchet, Metal Outer Sheath, insulated and Forceps Insert X1
11.8 | Atraumatic Grasping Forceps rotating, dismantling, insulated, with connector pin for unipolar coagulation, single action jaws, with especially fine serration, fenestrated, size 5 mm, length 36 cm consisting of:PlasticHandle, without ratchet, Metal OuterSheath, insulated and Forceps Insert X1
11.9 | Atraumatic Grasping Forceps, rotating, with connector pin for unipolar coagulation, size 5 mm, length 36 cm, atraumatic, fenestrated, single action jaws, consisting of: Plastic Handle with disengageable ratchet, with larger contact area, Outer Tube, insulated and Forceps Insert X1
11.1 | Atraumatic Dissecting and Grasping Forceps, rotating, with connector pin for unipolar coagulation, size 5 mm, length 36 cm, atraumatic, fenestrated, double action jaws, consisting of: Plastic Handle, with larger contact area, Outer Tube, insulated and Forceps Insert X1
11.11 | Atraumatic Dissecting and Grasping Forceps, rotating, with connector pin for unipolar coagulation, size 5 mm, length 36 cm, atraumatic, fenestrated, double action jaws, consisting of: Plastic Handle with disengageable ratchet, with larger contact area, Outer Tube, insulated and Forceps Insert X1
11.12 | Grasping Forceps, rotating, with connector pin for unipolar coagulation, size 5 mm, length 36 cm, "Tiger-jaws", 2x4 teeth, single action jaws, consisting of: Plastic-Handle with disengageable ratchet, with larger contact area, Outer Tube, insulated and Forceps Insert X1
11.13 | Blunt Nose Dissecting and Grasping Forceps, rotating, with connector pin for unipolar coagulation, size 5 mm, length 36 cm, heavy, double action jaws, consisting of: Plastic Handle, without ratchet, Outer Tube, insulated and Forceps Insert X1
11.14 | BABCOCK Grasping Forceps, rotating, with connector pin for unipolar coagulation, size 5 mm, length 36 cm, atraumatic, jaws with multiple teeth, fenestrated, single action jaws, consisting of: Plastic Handle with disengageable ratchet, with larger contact area, Outer Tube, insulated and Forceps Insert X1
11.15 | Right Angled Dissecting Forceps, rotating, with connector pin for unipolar coagulation, size 5 mm, length 36 cm, atraumatic, jaws with multiple teeth, fenestrated, single action jaws, consisting of: Plastic Handle with disengageable ratchet, with larger contact area, Outer Tube, insulated and Forceps Insert X1
11.16 | BABCOCK Clamp, rotating, size 10 mm, length 36 cm, double action jaws, consisting of: Metal Handle, with hemostat style ratchet, Outer Tube, insulated and Forceps Insert X1
11.17 | Claw Forceps, rotating, size 10 mm, length 36 cm, 2 x 3 teeth, single action jaws, consisting of: Metal Handle, with hemostat style ratchet, Outer Tube, insulated and Forceps Insert X1
11.18 | Right Angled Dissecting Forceps, rotating, size 10 mm, length 36 cm, 2 x 3 teeth, single action jaws, consisting of: Metal Handle, Outer Tube, insulated and Forceps Insert X1
11.19 | METZENBAUM Scissors, rotating, dismantling, insulated, with connector pin for unipolar coagulation, with LUER-Lock connector for cleaning, double action jaws, curved, length of blades 15 mm, size 5 mm, length 36 cm consisting of: Plastic Handle, insulated, without ratchet Metal Outer Sheath, insulated and Scissors Insert for use with trocars size 6 mm X1
11.2 | Scissors Insert, METZENBAUM Scissors, curved, length of blades 12 mm, double action jaws, size 5 mm, length 36 cm X1
11.21 | Straight Scissors, rotating, with connector pin for unipolar coagulation, size 5 mm, length 36 cm, straight, double action jaws, consisting of: Plastic Handle, without ratchet, Metal Outer Sheath, Scissors Insert X1
11.22 | Click Line Scissors Insert, straight, double action jaws, size 5 mm, length 36 cm X1
11.23 | Curved Bipolar Grasping Forceps, rotating, dismantling, with connector pin for bipolar coagulation, especially suitable for dissection, double action jaws, size 5 mm, length 36 cm, consisting of:Plastic Handle, without ratchet, Metal Outer Sheath and Forceps Insert X1
11.24 | Fenestrated Bipolar Grasping Forceps, rotating, dismantling, with connector pin for bipolar coagulation, with especially fine atraumatic serration, fenestrated jaws, double action jaws, size 5 mm, length 36 cm, consisting of: Plastic Handle, without ratchet Metal X1Outer Sheath and Forceps Insert X1
11.25 | Coagulating and Dissecting Electrode, with channel, L-shaped size 5 mm, length 36 cm, for use with suction and irrigation handles X1
11.26 | Handle with Trumpet Valve, for suction or irrigation, autoclavable, for use with coagulation suction tubes size 5 mm X1
11.27 | Suction and Irrigation Tube, with lateral holes, with two-way stopcock for single-hand control, size 5 mm, length 36 cm X1
11.28 | Puncture Needle, LUER-lock, diameter 1.6 mm, size 5 mm, length 36 cm X1
11.29 | Knot Tier, size 5 mm, length 36 cm, for extracorporeal knotting X1
11.3 | Fascial Closure Instrument for subcutaneous ligature of trocar incisions, size 2.8 mm, length 17 cm X1
11.31 | Unipolar High Frequency Cord, with 4 mm plug, length 300 cm X2
11.32 | Bipolar Cable X1
11.33 | Plastic Container for Sterilizing and Storage, perforated, with lid, for two-level storage, for use with 30 cm and 36 cm Click Line- insruments, external dimensions ( w x d x h ): 600 mm x 145 mm x 260 mm ( X1)

## Endoscopic & Exoscopic Video
RT2-14-004 | Endoscopic & Exoscopic, Video System, Neurosurgery | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
1 | DESCRIPTION
1.1 | A neurosurgery video system used for minimally invasive brain and spine procedures
2 | Product Offered:
2.1 | Make:
2.2 | Model:
2.3 | Initial Manufacturing year of Model:
3 | TECHNICAL SPECIFICATION
3.1 | The stack system must come complete with monitor, video processor, Exoscope, camera console, data management system, rigid scopes, 3D fiber optic light cable and the appropriate trolley.
3.2 | System must have ability to process ICG fluorescent imaging
3.3 | System must have module for 4K 3D endoscopy
3.4 | 3D exoscope with control unit and illumination source must be provided with a holding stand.
3.5 | Video Processor
3.5.1 | Make & Model:
3.5.2 | must have digital zoom facility and automatic white balancing
3.5.3 | Video control unit must have 1 or more camera head inputs with image display settings available and 3D video endoscopy capability
3.5.4 | The unit offered must have visual output configurations of the latest technology. Bidder to specify
3.6 | Camera head
3.6.1 | Make & Model:
3.6.2 | It must be controlled on camera console and head, with programmable buttons for capture, zoom and white balancing.
3.6.3 | A 4K 3D Camera Head with a easy-to-fit scope mounting design must be provided.
3.6.4 | Camera head must have a minimum of three programmable switches to allow user remote control
3.6.5 | Freezing facility (Programmable Button)must be available
3.6.6 | White balancing function must be available on the camera
3.6.7 | The length of the camera head cable must be between 3m and 5 m
3.6.8 | It must be autoclavable or other means of sterilization, please state the method
3.7 | Light Source
3.7.1 | Make & Model:
3.7.2 | Lamp design must be LED.
3.7.3 | Light source must have adjustable light intensity
3.7.4 | An autoclavable 3D fibre optic light guide cable of at least 5mm diameter and 2.3m length must be provided.
3.7.5 | Fibre optic light cable for rigid endoscopy must be provided
3.8 | Monitor
3.8.1 | Make & Model:
3.8.2 | It must be at least 32" 4K 3D LED medical grade monitor with high image quality.
3.8.3 | The trolley must have antistatic castors with at least 2 equipped with locking system/brake.
3.8.4 | The unit offered must have a minimum of 3 inputs configurations of the latest technology. Bidder to specify
3.8.5 | High contrast and wide viewing angle must be available
3.9 | Data Management System
3.9.1 | Make & Model:
3.9.2 | The unit offered must provide both video and still images recording in Full High Definition
3.9.3 | Internal memory must be a minimum of 1TB
3.9.4 | It must be a two-channel simultaneous recording
3.9.5 | Scope button to record and capture functions must be available
3.9.6 | The system offered must be provided with a compatible printer.
3.10 | Printer
3.10.1 | Make & Model:
3.10.2 | A4 or A6 colour printer with a minimum resolution of 423dpi must be provided.
3.11 | Trolley
3.11.1 | The offered unit must come in a purpose built Endoscopic System trolley with a minimum of one drawer for storage
3.11.2 | The trolley offered must have a tray for the keyboard and mouse.
3.11.3 | The trolley must have antistatic castors with at least 2 equipped with locking system/brake.
3.11.4 | The trolley must have a handle for ease access and transportation and a cable management arm.
3.11.5 | The trolley must have a cover at the rear for safety of cables.
3.11.6 | Trolley must have a secure gas cylinder holder
3.11.7 | The trolley must be supplied with a minimum of 3 additional 220 V power sockets apart from connected listed components.
4 | STANDARD ACCESSORIES & INSTRUMENTS
4.1 | 3x Neuroscurgery endoscopes (End-user's choice of size)
4.2 | Tray for cleaning, sterilising and storing endoscopes
4.3 | 3D glasses or clip-on's x2
4.4 | Printing paper (1x Box)
4.5 | Disposable Insufflation Tubing Set (x10)
4.6 | Fiber Optical Light Guide Cable
4.7 | Universal wrench
5 | Conformity Compliance (Please attach certificate)
5.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
5.2 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
5.2.1 | Please provide unique ref number of the ISO 13485 ceriticate:
5.3 | Model on offer must comply with EU Medical Device Regulation (EU) 2017/745 or equvalent from SAHPRA recognized jurisdictions . Attach a copy of certification
5.3.1 | Please provide unique ref number of the ceriticate:
5.4 | Hazardous Substance Act: If this type of equipment / apparatus appears on the schedule of Hazardous substances issued by the Directorate: Health Technology of the Department of Health, a license in terms of the Act on Hazardous Substances (Act 15/1973) must be submitted with the bid / quotation documents. The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.5 | All electrical/electronic medical equipment must be licensed by Radiation Control, where applicable a copy of the license must be submitted The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.5.1 | Radiation Control License number:
5.6 | The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade (rubber to with 2 screws) type and it must be a minimum length of three (3) meters. NB the mains cable of the unit tendered for must be SABS colour coded.
5.7 | The equipment quoted for must be protected against electromagnetic interference IEC 60601-1-2
5.8 | The unit must comply with electrical safety and performance standard such as IEC 60601-2-21, attached certification
5.8.1 | Please provide unique ref number of the IEC ceriticate:
6 | COMPREHENSIVE MAINTENANCE
6.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) should be included in the unit price as per the SCC document
6.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
6.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
6.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
6.2 | A 5-year service exchange for the laparoscopic rigid scopes must be quoted for separately
6.3 | The bidder to provide the OEM recommended service interval of the unit:
6.4 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
7 | General
7.1 | Has the product on offer or any of its components being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
7.2 | The unit on offer must be the latest technology.
7.3 | Initial Manufacturing year of Model:
RT2-14-004 | Endoscopic & Exoscopic, Video System, Neurosurger Replacement Accessories and Consumables | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
RT2-14-004a | 3x Neuroscurgery endoscopes (End-user's choice of size)
RT2-14-004b | Tray for cleaning, sterilising and storing endoscopes
RT2-14-004c | 3D glasses or clip-on's x2
RT2-14-004d | Printing paper (1x Box)
RT2-14-004e | Disposable Insufflation Tubing Set (x10)
RT2-14-004f | Fiber Optical Light Guide Cable
RT2-14-004g | Universal wrench

## Endoscopic System
RT2-14-005 | Video Endoscopic System | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
1 | DESCRIPTION
1.1 | A video system used for minimally invasive brain and spine procedures
2. | Product Offered:
2.1 | Make:
2.2 | Model:
2.3 | Initial Manufacturing year of Model:
3 | TECHNICAL SPECIFICATION
3.2 | Video Processor
3.2.1 | Make & Model:
3.2.2 | The processor must be able to interpret a video signal from an endoscope by using a colour chip technology.
3.2.3 | White balance adjustment with storage function
3.2.4 | Pre-Freeze image function to select the clearest still image
3.2.5 | HD Output for compatibility to HD monitor. Compatible with latest technology.
3.2.6 | USB port is required on the unit
3.2.7 | Image Enhancement Setting: Structural and Edge Enhancement must be available
3.2.8 | Image Size Selection: Can be selected from keyboard or remote switch
3.2.9 | Image Contrast Settings: 3 (Normal, High, Low) must be available.
3.3 | Light Source
3.3.1 | Make & Model:
3.3.2 | Lamp design must be LED or Xenon. Bidder to specify.
3.3.3 | If Xenon, 300W with spare lamp must be provided
3.3.4 | The unit offered must allow for brightness adjustment
3.4 | Monitor
3.4.1 | Make:
3.4.2 | Model:
3.4.3 | It must be at least 26" LED medical grade monitor with high image quality.
3.4.4 | Must be a touch screen monitor or operated with a separate keypad
3.4.5 | The trolley must have antistatic castors with at least 2 equipped with locking system/brake.
3.4.6 | The unit offered must have a minimum of 3 inputs configurations of the latest technology. Bidder to specify
3.4.7 | High contrast and wide viewing angle must be available
3.5 | Printer
3.5.1 | Make:
3.5.2 | Model:
3.5.3 | A4 or A6 colour printer with a minimum resolution of 423dpi must be provided.
3.6 | Image Management System
3.6.1 | Make & Model:
3.6.2 | The unit offered must provide both video and still images recording in Full High Definition
3.6.3 | Supported formats: Internal HDD, Blue Ray, DVD, USB and external HDD
3.6.4 | Internal memory must be a minimum of 1TB
3.6.5 | It must be a two-channel simultaneous recording
3.6.6 | Scope button to record and capture functions must be available
3.6.7 | The trolley must have antistatic castors with at least 2 equipped with locking system/brake.
3.7 | Trolley
3.7.1 | The offered unit must come in a purpose built Endoscopic System trolley with a minimum of one drawer for storage
3.7.2 | The trolley offered must have a tray for the keyboard and mouse.
3.7.3 | The trolley must have antistatic castors with at least 2 equipped with locking system/brake.
3.7.4 | The trolley must have a handle for ease access and transportation and a cable management arm.
3.7.5 | The trolley must have a cover at the rear for safety of cables.
3.7.6 | The trolley must be supplied with a minimum of 3 additional 220 V power sockets apart from connected listed components.
3.8 | IT NETWORKING AND CONNECTIVITY REQUIREMENTS
3.8.1 | An RS232/RJ45 data output port with a freely available data format for connection to another monitor or hospital information system shall be provided
4 | STANDARD ACCESSORIES
4.1 | Video Connecting Cable
4.2 | Irrigation tube
4.3 | Suction valve
4.4 | Suction bottle
4.5 | Scope Sterilization/Cleaning trolley
4.6 | Endoscopy Cart/trolley
4.7 | Mouth guard
4.8 | Leak tester
5 | Conformity Compliance (Please attach certificate)
5.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
5.2 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
5.2.1 | Please provide unique ref number of the ISO 13485 ceriticate:
5.3 | Model on offer must comply with EU Medical Device Regulation (EU) 2017/745 or equvalent from SAHPRA recognized jurisdictions . Attach a copy of certification
5.3.1 | Please provide unique ref number of the ceriticate:
5.4 | Hazardous Substance Act: If this type of equipment / apparatus appears on the schedule of Hazardous substances issued by the Directorate: Health Technology of the Department of Health, a license in terms of the Act on Hazardous Substances (Act 15/1973) must be submitted with the bid / quotation documents. The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.5 | All electrical/electronic medical equipment must be licensed by Radiation Control, where applicable a copy of the license must be submitted The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.5.1 | Radiation Control License number:
5.6 | The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade (rubber to with 2 screws) type and it must be a minimum length of three (3) meters. NB the mains cable of the unit tendered for must be SABS colour coded.
5.7 | The equipment quoted for must be protected against electromagnetic interference IEC 60601-1-2
5.8 | The unit must comply with electrical safety and performance standard such as IEC 60601-2-21, attached certification
5.8.1 | Please provide unique ref number of the IEC ceriticate:
6 | COMPREHENSIVE MAINTENANCE
6.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) should be included in the unit price as per the SCC document
6.2 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
6.3 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
6.4 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
6.5 | A 5-year service exchange for the laparoscopic rigid scopes must be quoted for separately
6.6 | The bidder to provide the OEM recommended service interval of the unit:
6.7 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
7 | General
7.1 | Has the product on offer or any of its components being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
7.2 | The unit on offer must be the latest technology.
7.3 | Initial Manufacturing year of Model:
RT2-14-005a | Flexible Gastroscope, Adult | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
Make & Model:
1.1 | Instrument Channel: > 2.8mm
1.2 | Insertion tube diameter: < 10mm
1.3 | Working length: 1020mm - 1200mm
1.4 | Field or angle of view: 1400
1.5 | Angulation: minimum of up: 210, Down: 90, Right: 100, Left 100
1.6 | Depth of View: 5 - 100mm
1.7 | Auxillary water jet channel as optional
1.8 | Connector must be completelly water proof
1,9 | Scope case
RT2-14-005b | Flexible Gastroscope, Paediatric | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
Make & Model:
1.1 | Instrument Channel: Greater than 2mm
1.2 | Insertion tube diameter: < 6mm
1.3 | Working length: 1650mm or more
1.4 | Field or angle of view: 1400
1.5 | Angulation: minimum of up: 180, Down: 160, Right: 160, Left 160
1.6 | Depth of View: 3mm - 100mm
1.7 | Auxillary water jet channel as optional
1.8 | Connector must be completelly water proof
1,9 | Scope case
RT2-14-005c | Flexible Bronchoscope, Adult | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
Make & Model:
1.1 | Instrument Channel: ≥2.2mm
1.2 | Insertion tube diameter: ≥6.2mm
1.3 | Working length: 600mm minimum
1.4 | Field or angle of view: 1200
1.5 | Angulation of at least: 100 0 up and 100 0 down
1.6 | Depth of View: 3mm - 100mm
1,7 | Scope case
RT2-14-005d | Flexible Bronchoscope, Paediatric | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
Make & Model:
1.1 | Instrument Channel: ≥2mm
1.2 | Insertion tube diameter: < 4.2mm
1.3 | Working length: 600mm maximum
1.4 | Field or angle of view: 90o - 1400
1.5 | Angulation of at least: 100 0 up and 100 0 down
1.6 | Depth of View: 3mm - 100mm
1,7 | Scope case
RT2-14-005e | Flexible Bronchoscope, Neonatal | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
Make & Model:
1.1 | Instrument Channel: < 1.2mm
1.2 | Insertion tube diameter: < 2.8 mm
1.3 | Working length: 600mm maximum
1.4 | Field or angle of view: 1100
1.5 | Angulation of at least: 100 0 up and 100 0 down
1.6 | Depth of View: 2mm- 50mm
1,7 | Scope case
RT2-14-005f | Flexible Colonoscope, Adult | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
Make & Model:
1.1 | Instrument Channel: > 3.7mm
1.2 | Insertion tube diameter: < 14mm
1.3 | Working length: < 1700mm
1.4 | Field or angle of view: > 140o
1.5 | Angulation: Up & Down: >180o ; Left & Right: >130o
1.6 | Depth of View: 5 - 100mm
1.7 | Must have auxillary water jet channel
1.8 | Connector must be completelly water proof with no need for additional waterproof caps
1,9 | Scope case
RT2-14-005g | Flexible Colonoscope, Paediatric | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
Make & Model:
1.1 | Instrument Channel: > 2.8mm
1.2 | Insertion tube diameter: < 9.8mm
1.3 | Working length: < 1700mm
1.4 | Field or angle of view: > 140o
1.5 | Angulation: Up & Down: >180o ; Left & Right: >130o
1.6 | Depth of View: 5 - 100mm
1.7 | Must have auxillary water jet channel
1.8 | Connector must be completelly water proof with no need for additional waterproof caps
1,9 | Scope case
RT2-14-005 | Endoscope Replacement Accessories | Complies Yes/No | Price
RT2-14-005h | Video Connecting Cable
RT2-14-005i | Irrigation tube
RT2-14-005k | Suction valve
RT2-14-005l | Suction bottle
RT2-14-005m | Scope Sterilization/Cleaning trolley
RT2-14-005n | Endoscopy Cart/trolley
RT2-14-005o | Mouth guard
RT2-14-005p | Leak tester
RT2-14-005q | Brushes
RT2-14-005r | Brush handles
RT2-14-005s | Suction catheter adaptor

## Shear Wave Elastrograph
RT2-14-006 | Shear-Wave Elastograph | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
1 | Description
1.1 | A medical device to assess tissue stiffness and elasticity, aiding in the diagnosis and monitoring of several conditions.
2 | Product Offered:
2.1 | Make:
2.2 | Model:
2.3 | Initial Manufacturing year of Model:
2.4 | Radiation Control License number:
3 | Technical Specification
3.1 | The unit must offer B-mode, Doppler, and Shear-Wave Elastrography types of ultrasound
3.2 | The unit must be a fully digital unit with a memory beam former and processor to generate sharp defined ultrasound images, reduce artifacts and improve image quality.
3.3 | It must utilise the latest software to perform investigations.
3.4 | High image quality will be one of the main criteria that will be used to evaluate the offers.
3.5 | Although offered units might comply with all the specifications, image quality must be a critical evaluation criterion, specifically with increased BMI cases.
3.6 | The unit must be of the latest technology.
3.7 | The unit will be used for general applications as a multi-purpose unit for Abdominal (liver, kidney) and Superficial (breast, thyroid, musculoskeletal) assessments.
3.8 | IMAGING MODALITIES
3.8.1 | The unit must offer excellent high image quality in B mode and fully digital colour flow Doppler. Colour Doppler should have no delay after activation.
3.8.2 | The system Dynamic Range must not be less than 172 dB. The Dynamic range should be able to be changed by a minimum of 4dB increments.
3.8.3 | The unit must have direct storage on hard drive of not less than 1TB. State all other storage facility that is included in the offered price. (Such as CD, DVD, MOD). Recorded CD roms must be compatible with any standard PC, without requiring extra software. USB download of images from the system must be possible.
3.8.4 | A digital thermal printer must be included in the offer for the machine.
3.8.5 | The system must offer a Digital Continuous Beam Former to generate sharp defined ultrasound beams, to reduce artifacts and to provide more consistent overall image quality.
3.8.6 | The unit must offer multi-frequency on all probes to provide sharp reproducible image quality in both superficial and deep anatomy.
3.8.7 | The unit must be equipped with a non-interlaced colour monitor with a screen size of not less than 21 inch, measured diagonally. The monitor must be of the flat screen LED type.
3.8.8 | The unit should have at least two (2) active probe connectors with electronic switching.
3.8.9 | The unit must be able to offer scanning methods for Electronic Convex and Linear technology.
3.8.10 | The unit must offer depth gain in not less than 8 steps. Unit must be programmable to start up automatically at a depth of operator’s choice and programmable to start up at different depths for specific examinations.
3.8.11 | Image processing must be done via a rotary dial or similar technology with at least 5 combinations for optimising the best image quality for each patient’s physique (user programmable). Image processing should visibly improve image and not just change it.
3.8.12 | Pan-Zoom enlargement must be in real and frozen mode up to 10X zoom.
3.8.13 | A colour pallet must be available for the colourization of the B mode and M mode and PW Doppler trace.
3.8.14 | Independent M mode and PW Doppler gain adjustment must be available.
3.8.15 | Sweep speed must be adjustable in at least 4 steps.Unit must have selectable sweep rates of 1,2,3,4,8 s/screen
3.8.16 | The unit must offer Tissue Harmonics as standard in the system.
3.8.17 | Grey scale level must be at least 256 shades.
3.8.18 | The unit must have not less than 1024 processing channels.
3.8.19 | The device must be built in a sturdy transportable trolley. All of the unit's components must be kept on the trolley, which also has enough room set aside for the probes and any disposables required for the machine. The device will be utilised for mobile ultrasonography in various locations. a) System width in centimetres b) System depth in centimetres.
3.8.20 | The frame rate must not be less than 394 frames per second. Indicate the frame rate that the system offers
3.8.21 | Image depth must be not less than 30cm. Indicate the image depth that the system offers
3.8.22 | The unit should offer quick scan to automatically adjusts STC/TGC and gain through one-touch operation. This must visibly improve the image
3.8.23 | Post processing must be available on frozen images
3.8.24 | Unit must offer user programmable annotation for quick labelling. User must be able to program approximately 40 words of their own choice per selected examination type for annotation purposes
3.8.25 | Proprietary algorithms for wave propagation and elasticity mapping must be available.
3.8.26 | The unit must have advanced tissue motion correction to ensure accurate wave propagation.
3.8.27 | Auto-optimization for image contrast and elastography sensitivity must be available.
3.8.28 | Automatic shear wave velocity calculation and stiffness map generation must be available.
3.8.29 | The unit must have manual and automatic ROI selection.
3.9 | DOPPLER
3.9.1 | PW Doppler must be offered on all imaging probes.
3.9.2 | Maximum detectable velocity for PW must be stated.
3.9.3 | Pulse Repetition Frequency must vary between 0,7 – 19,8 kHz pulsed wave and 2,1 – 40kHz continuous wave
3.9.4 | Independent angle steering up to ±20° for colour and Doppler must be offered.
3.9.5 | Independent gain control for PW must be available.
3.9.6 | Colour pallet offering 7 types independent of B mode and Pulse Doppler pallet to be available.
3.9.7 | Sampling with Pulse Doppler must vary between 1 and 20 mm.
3.9.8 | Gate size must vary at least between 1 and 8mm.
3.9.9 | Base line shift must be possible in at least 8 steps.
3.9.10 | Spectral invert must be possible.
3.9.11 | At least 6 wall filter selections must be available. Bidder to state minimum.
3.9.12 | Accurate autotrace function for both arterial and venous waveforms must be available.
3.10 | COLOUR DOPPLER
3.10.1 | A minimum of 8 colour maps must be available for colour display.
3.10.2 | Colour maps must be available to be selected in real time and freeze mode.
3.10.3 | Independent colour gain adjustment must be from 1 to 16 dB in 0,5 dB increments. Indicate range
3.10.4 | Maximum colour frame rate must not be less than 285 frames per second.
3.10.5 | Colour flow must operate to the full depth of 30 cm.
3.10.6 | Colour maps and processing selections must be able to be selected by the user and stored in the presets.
3.10.7 | Quad signal processing and parallel signal processing must be available. Bidders to supply details.
3.10.8 | Power Doppler on all transducers must be available.
3.10.9 | The unit offered must be able of real time triplex mode.
3.10.10 | Colour mapping must be instantaneous – no time delay /lag.
3.11 | IMAGE MEMORY
3.11.1 | Image memory to operate in both dual and quad screen format.
3.11.2 | ECG synchronised recording must be possible. (optional)
3.11.3 | It must be possible to record B loop up to 500 frames.
3.11.4 | Cine loop and measurements must be active simultaneously.
3.11.5 | M/D review of a minimum of 15 seconds must be possible.
3.11.6 | B Mode Imaging on linear transducer to be electronically steerable.
3.12 | TRANSDUCERS
3.12.1 | The following frequencies to address the relevant applications must be possible
3.12.1.1 | Convex probe (Abdominal applications) 1.0 to 6.0 MHz
3.12.1.2 | Linear probe 4.0 to 18.0 MHz
3.12.1.3 | Endocavitary probe 5.0 to 12.0 MHz
3.12.2 | All imaging transducers must be multi-frequency.
3.12.3 | Tissue Harmonics must be available on all offered transducers.
3.12.4 | The selected frequency of the imaging transducer must be displayed on the screen.
3.12.5 | All transducers must be digitally controlled so continuous beam forming is possible. State number of focal zones.
3.12.6 | A minimum of 2 active probes must be available. It must be possible to electronically switch between these 2 probes.
3.13 | SHEAR WAVE SPEED AND STIFFNESS MEASUREMENT
3.13.1 | Shear Wave Speed Range:
3.13.1.1 | Typical range: 0.5 to 10 m/s.
3.13.2 | Young’s Modulus Measurement:
3.13.2.1 | Range: 5 kPa to 300 kPa, depending on the tissue.
3.13.3 | Depth of Penetration for Shear Waves:
3.13.3.1 | Up to 8 cm in abdominal applications (with convex transducers).
3.13.3.2 | Limited to 3-4 cm in superficial applications (with linear transducers).
3.13.4 | Frame Acquisition Time for Elastography:
3.13.4.1 | 5-20 ms, ensuring real-time feedback during scanning.
3.13.5 | Quantitative Elastography Maps:
3.13.5.1 | Color-coded elastograms (commonly blue-red scales) for tissue stiffness visualization.
3.13.6 | ROI (Region of Interest) Size Adjustability:
3.13.6.1 | User can adjust the size and location of ROI for targeted measurements.
3.14 | DICOM COMPLIANCE
3.14.1 | The offered unit must be fully DICOM compliant.
3.14.2 | It must be possible for the unit to be connected to PACS.
3.15 | CLINICAL APPLICATION IN SOFTWARES TO BE AVAILABLE
3.15.1 | Liver Elastography
3.15.2 | Breast Elastography
3.15.3 | Musculoskeletal Elastography
3.15.4 | Thyroid Elastography
3.15.5 | Prostate Elastography
3.16 | Power and Batteries
3.16.1 | The monitor must operate off both the 220VV , 50Hz single phase AC power supply and an internally fitted rechargeable battery. The unit must be fused in both the LIVE and NEUTRAL.
3.16.2 | The unit offered must be supplied with an internal rechargeable battery.
3.16.3 | The rechargeable battery must be automatically charged when the unit is connected to a live A.C. supply
3.16.4 | In the event of A.C. mains failure the battery must automatically provide continuous operation
3.16.5 | The unit must be provided with visible indication to reflect whether the unit is being operated off the AC power or the rechargeable battery power
3.16.6 | The unit offered must be provided with circuitry, which must ensure that the rechargeable battery will be protected against over-charging and over-discharging
3.16.7 | Units that are quoted on and employ external battery chargers and power supply will not be considered.
3.16.8 | Battery back-up option must be provided (battery allowing unit to be placed in a sleep mode while being moved from one location to the next so as to provide rapid resumption of operation)
3.16.9 | The rechargeable battery must be of a reasonable capacity, such that with battery power the unit must be able to operate continuously for a minimum of four (4)hours
3.16.9.1 | UPS option in the event that offered unit's battery capacity does not meet the 4-hour requirement
3.16.9.2 | Please state the following:
3.16.9.2.1 | Capacity
3.16.9.2.2 | Voltage
3.16.9.2.3 | Type of battery used
3.16.9.2.4 | Length of battery operating time
4 | Standard Accessories and Consumables
4.1 | Convex probe 1.0 to 6.0 MHz
4.2 | Linear probe 4.0 to 18.0 MHz
4.3 | Endocavitary probe 5.0 to 12.0 MHz
4.4 | Printer
4.5 | Ultrasound Gel 2L
4.6 | Uninterrupted Power Supply (UPS)
4.7 | Printing paper (box of 10 rolls)
5 | Conformity Compliance (Please attach certificate)
5.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
5.2 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
5.2.1 | Please provide unique ref number of the ISO 13485 ceriticate:
5.3 | Model on offer must comply with EU Medical Device Regulation (EU) 2017/745 or equvalent from SAHPRA recognized jurisdictions . Attach a copy of certification
5.3.1 | Please provide unique ref number of the EC ceriticate:
5.4 | Hazardous Substance Act: If this type of equipment / apparatus appears on the schedule of Hazardous substances issued by the Directorate: Health Technology of the Department of Health, a license in terms of the Act on Hazardous Substances (Act 15/1973) must be submitted with the bid / quotation documents. The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.5 | All electrical/electronic medical equipment must be licensed by Radiation Control, where applicable a copy of the license must be submitted The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.5.1 | Radiation Control License number:
5.6 | The unit shall operate from 220V AC/50Hz mains supply. The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade (rubber to with 2 screws) type and it must be a minimum length of three (3) meters. NB the mains cable of the unit tendered for must be SABS colour coded.
5.7 | The equipment quoted for must be protected against electromagnetic interference IEC 60601-1-2
5.9 | The unit must comply with electrical safety and performance standard such as IEC 60601-2-37, attached certification
5.9.1 | Please provide unique ref number of the IEC ceriticate:
6 | COMPREHENSIVE MAINTENANCE
6.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) must be quoted for, as per the SCC document
6.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
6.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
6.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
6.3 | The bidder to provide the OEM recommended service interval of the unit:
6.4 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
7 | General
7.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
7.2 | The unit on offer must be the latest technology.
7.2.1 | Initial Manufacturing year of Model:
RT2-14-006 | Shear-Wave Elastograph Unit Accessories | Complies Yes/No
RT2-14-006a | Convex probe 1.0 to 6.0 MHz
RT2-14-006b | Linear probe 4.0 to 18.0 MHz
RT2-14-006c | Endocavitary probe 5.0 to 12.0 MHz
RT2-14-006d | Printer
RT2-14-006e | Ultrasound Gel 2L
RT2-14-006f | Uninterrupted Power Supply (UPS)
RT2-14-006g | Printing paper (box of 10 rolls)