## Company Information
Company name:
CSD Number:
Pysical Address:
Postal Address:
Contact Person:
Telephone No:
Cell No:
Email Address:

## Profile
3. PAEDIATRIC CARE EQUIPMENT
# | Item number | Item Description
1 | RT2-03-01 | Apnea Monitor, Impedance sensor
2 | RT2-03-02 | Apnea Monitor, Pressure sensor
3 | RT2-03-03 | Apnea Monitor, Airflow sensor
4 | RT2-03-04 | Bilirubin Meter, Xenon
5 | RT2-03-05 | Bilirubin Meter, LED
6 | RT2-03-06 | Radiant Warmer-Incubator Combination
7 | RT2-03-07 | ICU Crib
8 | RT2-03-08 | Incubator, ICU
9 | RT2-03-09 | Incubator, Standard
10 | RT2-03-10 | Incubator, Transport
11 | RT2-03-11 | Infant Radiant Warmer
12 | RT2-03-12 | Phototherapy Unit, LED
13 | RT2-03-13 | Phototherapy Unit, Blanket, LED
14 | RT2-03-14 | Cooling Device
15 | RT2-03-15 | Ultrasound, Neonatal
16 | RT2-03-16 | Vein Viewer
4. OBSTETRIC & GYNAECOLOGY EQUIPMENT
# | Item number | Item Description
1 | RT2-04-01 | Cardiotocograph, Single
2 | RT2-04-02 | Cardiotocograph, Single + MHR
3 | RT2-04-03 | Cardiotocograph, Twins
4 | RT2-04-04 | Cardiotocograph, Twins + MHR
5 | RT2-04-05 | Fetal Doppler
6 | RT2-04-06 | Video Colposcope
7 | RT2-04-07 | Vacuum Extractor
8 | RT2-04-08 | Ultrasound Machine, O&G

## Category-3 Paeds
PAEDIATRIC CARE EQUIPMENT
# | Item number | Item Description
1 | RT2-03-001 | Apnea Monitor
2 | RT2-03-002 | Bilirubin Meter
3 | RT2-03-003 | Radiant Warmer-Incubator Combination
4 | RT2-03-004 | ICU Crib
5 | RT2-03-005 | Incubator, ICU
6 | RT2-03-006 | Incubator, Standard
7 | RT2-03-007 | Incubator, Transport
8 | RT2-03-008 | Infant Radiant Warmer
9 | RT2-03-009 | Phototherapy Unit, LED
10 | RT2-03-010 | Phototherapy Unit, Blanket, LED
11 | RT2-03-011 | Cooling Device
12 | RT2-03-012 | Ultrasound, Neonatal
13 | RT2-03-013 | Vein Viewer

## Apnea Monitor, Impedance sensor
RT2-03-001 | APNEA MONITOR | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
1 | Description
1.1 | A non invasive device that tracks and monitors neonates and infants' breathing patterns and detecting apneic episodes.
2 | Product Offered:
2.1 | Make:
2.2 | Model:
2.3 | Initial Manufacturing year of Model:
2.4 | Radiation Control License number:
3 | Technical Specification
3.1 | The unit must continuously monitor respiration in neonates and infants.
3.2 | The bidder to specify sensor technology of the unit offered
3.3 | Bidder must describe the working principle of the breathing sensor/transducer.
3.4 | The monitor must trigger both audible and visible alarms after a preselected delay of 10, 15 or 20 sec following the onset of apnea.
3.5 | The alarm must be audible at a distance of at least 10 metres
3.6 | The monitor must indicate every breathing motion detected.
3.7 | The unit must have a numeric display for breath rate of the patient and it must be visible under all lighting conditions.
3.8 | The unit must operate from an internal rechargeable battery in case of mains failure.
3.9 | The unit must have the following alarm conditions:
3.9.1 | Low battery: audible and visual alarms
3.9.2 | Breathing cessation: audible and visual alarms
3.9.3 | Sensor disconnect: audible and visual alarms
3.10 | The sensor/transducer must be safely isolated from the main supply.
3.11 | The unit must be waterproof
3.12 | The design must be flat with rubber footing to provide stability from any disturbance or movements on flat surface with the option to be mounted on to the IV Pole or overbed rail.
4 | Standard Accessories
4.1 | Disposable sensor x10
4.2 | Rechargeable battery
4.3 | Battery charger
5 | Conformity Compliance (Please attach certificate)
5.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
5.2 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
5.2.1 | Please provide unique ref number of the ISO 13485 ceriticate:
5.3 | Model on offer must comply with EU Medical Device Regulation (EU) 2017/745 or equvalent from SAHPRA recognized jurisdictions . Attach a copy of certification
5.3.1 | Please provide unique ref number of the ceriticate:
5.4 | Hazardous Substance Act: If this type of equipment / apparatus appears on the schedule of Hazardous substances issued by the Directorate: Health Technology of the Department of Health, a license in terms of the Act on Hazardous Substances (Act 15/1973) must be submitted with the bid / quotation documents. The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.5 | All electrical/electronic medical equipment must be licensed by Radiation Control, where applicable a copy of the license must be submitted The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.5.1 | Radiation Control License number:
5.6 | The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade (rubber to with 2 screws) type and it must be a minimum length of three (3) meters. NB the mains cable of the unit tendered for must be SABS colour coded.
5.7 | The equipment quoted for must be protected against electromagnetic interference IEC 60601-1-2
5.8 | The unit must comply with electrical safety and performance standard such as IEC 60601-2-21, attached certification
5.8.1 | Please provide unique ref number of the IEC ceriticate:
6 | COMPREHENSIVE MAINTENANCE
6.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) should be included in the unit price as per the SCC document
6.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
6.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
6.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
6.3 | The bidder to provide the OEM recommended service interval of the unit:
6.4 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
7 | General
7.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
7.2 | The unit on offer must be the latest technology.
7.2.1 | Initial Manufacturing year of Model:
RT2-03-001 | Apnea Monitor ACCESSORIES & CONSUMABLES | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
RT2-03-001a | Disposable sensor pack of 10
RT2-03-001b | Re-usable sensor
RT2-03-001c | Rechargeable battery
RT2-03-001d | Battery charger

## Bilirubin Meter
RT2-03-002 | Bilirubin Meter | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
1 | Description
1.1 | Non-invasive device used to measure bilirubin levels in newborns to diagnose and monitor jaundice.
2 | Product Offered:
2.1 | Make:
2.2 | Model:
2.3 | Initial Manufacturing year of Model:
3 | Technical Specification
3.1 | The unit offered must be effective regardless of skin colour
3.2 | It must be suitable for use on pre-term neonates of gestational age > 24 weeks and post-natal age 0 - 20 days
3.3 | The unit must be light weight, handheld and easy to operate.
3.4 | The measurement must be done non-invasively.
3.5 | Measurements must be done by means of a light measured at specific wavelengths between 300 and 900 nm
3.6 | The measuring method must be based on the determination of the yellowness of the subcutaneous tissue by using two optical paths to measure optical density difference at two wavelengths
3.7 | The bidder must specify light source technology that meets the above wavelength.
3.8 | The unit must take the measurement, process and display the information without need of any peripheral devices.
3.9 | Must not be sensitive to user-applied pressure or angle of device
3.10 | Must be able to display result of a single measurement
3.11 | Must not be sensitive to motion artifacts (or movements) of the baby as well as the ambient light
3.12 | The unit must be able to operate with or without light head cover.
3.13 | The unit must not require routine calibration.
3.14 | The light source life: > 150 000 measurements
3.15 | The unit must be capable of being used before, during and post phototherapy.
3.16 | The unit must incorporate a LCD to display the following:
3.16.1 | The results
3.16.2 | Operating parameter
3.16.3 | Errors
3.16.4 | Date and time
3.16.5 | Battery status
3.17 | The unit must report results in mg/dl or µmol/L
3.18 | The total serum billirubin range must be from 0 to 20 mg/dl or 0 to 340 µmol/L.
3.19 | Repeatability must be ±1.5mg/dl or 25 µmol/L.
3.20 | Accuracy during and post phototherapy must be ±1.75mg/dl or 30µmol/L.
3.21 | The following error signals must be displayed on the unit.
3.21.1 | Unacceptable calibration
3.21.2 | Unacceptable patient measurement
3.21.3 | Too much movement during measurement
3.21.4 | Measurement is out of clinically proven range
3.21.5 | Light bulb faulty
3.21.6 | System error
3.21.7 | Low battery
3.21 | The unit must incorporate a rechargeable battery, and must operate only from the battery power.
3.22 | Battery must be supplied
3.23 | The battery must allow at least 100 measurements on a full charge.
3.24 | Expected battery life: at least 1 year
3.25 | A battery charger, operating from 220 Volt, 50Hz mains power must be supplied
3.26 | The battery charger must have a battery overcharge protection circuit.
3.27 | The unit must indicate that the battery is being charged when the unit is plugged into mains power.
3.28 | It must indicate either the charge or the number of tests left in the battery
3.29 | A protective carry case must be supplied
3.30 | The unit must not weigh more than 250grams
4 | Standard Accessories
4.1 | Protective carry case
4.2 | Rechargeable battery
4.3 | Battery charger
5 | Conformity Compliance (Please attach certificate)
5.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
5.2 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
5.2.1 | Please provide unique ref number of the ISO 13485 ceriticate:
5.3 | Model on offer must comply with EU Medical Device Regulation (EU) 2017/745 or equvalent from SAHPRA recognized jurisdictions . Attach a copy of certification
5.3.1 | Please provide unique ref number of the ceriticate:
5.4 | Hazardous Substance Act: If this type of equipment / apparatus appears on the schedule of Hazardous substances issued by the Directorate: Health Technology of the Department of Health, a license in terms of the Act on Hazardous Substances (Act 15/1973) must be submitted with the bid / quotation documents. The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.5 | All electrical/electronic medical equipment must be licensed by Radiation Control, where applicable a copy of the license must be submitted The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.5.1 | Radiation Control License number:
5.6 | The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade (rubber to with 2 screws) type and it must be a minimum length of three (3) meters. NB the mains cable of the unit tendered for must be SABS colour coded.
5.7 | The equipment quoted for must be protected against electromagnetic interference IEC 60601-1-2
5.8 | The unit must comply with electrical safety and performance standard such as IEC 60601-2-21, attached certification
5.8.1 | Please provide unique ref number of the IEC ceriticate:
6 | COMPREHENSIVE MAINTENANCE
6.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) should be included in the unit price as per the SCC document
6.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
6.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
6.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
6.3 | The bidder to provide the OEM recommended service interval of the unit:
6.4 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
7 | General
7.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
7.2 | The unit on offer must be the latest technology.
7.2.1 | Initial Manufacturing year of Model:
RT2-03-002 | Bilirubin Accessories & Consumables | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
RT2-03-002a | Tip covers pack of 10

## Radiant Warmer-Incubator Combo
RT2-03-003 | Radiant Warmer-Incubator Combination | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
1 | Description
1.1 | A medical device designed to provide a controlled environment for newborn babies by heated air around them to promote healthy growth and development.
2 | Product Offered:
2.1 | Make:
2.2 | Model:
2.3 | Initial Manufacturing year of Model:
2.4 | Radiation Control License number:
3 | Technical Specification
3.1 | The unit must have option to be used as an open incubator (radiant warmer) and as closed incubator with just a simple touch
3.2 | Platform and mattress:
3.2.1 | Adjustable tilt to Trendellenburg and reverse Trendellenburg positions
3.2.2 | Locking mechanism to keep the mattress platform locked in desired positions.
3.2.3 | Must accommodate X-ray film cassettes beneath the mattress.
3.2.4 | The mattress must be waterproof and must be X-ray translucent.
3.2.5 | Mattress size: approximately 500 x 650 mm(±50mm tolerance).
3.2.6 | Mattress thickness: at least 25 mm
3.2.7 | Mattress must support patient weight up to 4,5 kg.
3.2.8 | All four sides of the platform must be fitted with transparent guards/panels.
3.2.9 | The front and two side panels must be hinged.
3.2.10 | Facility for holding patient cables, tubes, drains etc to prevent them from being trapped or dislodged even when sides are opened
3.3 | Heater and Controls
3.3.1 | The unit on offer must provide dual heating components(Overhead canopy and Inside the chassis) with the automatic switch over whenever a specific mode is selected between open and closed incubator.
3.4.1 | Radiant heater housed in an overhead canopy, with internal reflector, mounted on the rear mainframe.
3.4.2 | Indicator to show heater is switched ON
3.4.3 | The overhead canopy must swivel left, right then must tilt and lock in the desired position in order to continuously warm the baby. Bidders to specify the technology/mechanism.
3.4.4 | Microprocessor control with automatic self-test when switched on.
3.4.5 | The heat controller must provide the following modes for temperature control:
3.4.5.1 | Pre-heat
3.4.5.2 | Manual heat control: adjustable heater power from 0 - 100%.
3.4.5.3 | Skin temperature control: servo-control using a skin temperature probe
3.4.6 | A reusable skin temperature probe must be supplied
3.4.7 | Skin temperature probe cable length: at least 1m
3.4.8 | The skin temperature set point must be selectable in the range of 34 °C to 38 °C, in 0,1 °C increments.
3.4.9 | The skin temperature measurement range: 30 °C to 40 °C.
3.4.10 | Skin temperature measurement accuracy: 0,5 °C
3.4.11 | If a malfunction occurs affecting safety or performance the heater must be automatically turned off
3.5 | Alarms
3.5.1 | Audible and visual alarms when the 220 V power is lost during operation.
3.5.2 | Audible and visual alarms if the skin temperature deviates by ± 1 °C from the set temperature.
3.5.3 | Audible and visual alarms if the skin temperature exceeds 39°C
3.5.4 | Audible and visual alarms if skin temperature probe is disconnected or fails during control mode.
3.5.5 | Should the skin temperature probe become disconnected or fail, the unit must alarm (alarm must not be silenced more than 2 munites) and automatically switch to manual mode.
3.5.6 | Alarm test facility
3.5.7 | Temporary mute of audible alarms.
3.5.8 | Apgar timer, with audible alarm, with digital display: 1, 5 and 10 minutes elapsed time.
3.5.9 | Apgar timer with on/off switch and reset button
3.6 | Suction
3.6.1 | The item should be equipped with an integrated suction unit.
3.6.2 | Suction regulator adjustable: 0 to - 20 kPa
3.6.3 | Collection bottle (850 ml minimum) with overflow stop valve
3.6.4 | Adequate tubing must be supplied
3.6.5 | The suction regulator must be able to operate from the hospital’s piped vacuum system.
3.6.6 | Collection bottle must accommodate disposable liners
3.6.7 | The suction unit must have a bacterial filter
3.7 | Other
3.7.1 | The unit must be provided with an examination lamp, with on/off switch.
3.7.2 | The unit must be supplied with an oxygen cylinder holder.
3.7.3 | The unit must have a at least Two(2) drawers
3.7.4 | The minimum acceptable height of the heater canopy must be 1.85 metres above floor
3.7.5 | The mattress platform must be hieght adjustable with a minimum of 900 mm above floor.
3.7.6 | It must have an Instrument tray, approximately 200 x 300 mm and must hold up to 2 kg load.
3.7.7 | It must have a swivel mounted monitor shelf, approximately 300 x 400 mm, and must hold at least 5 kg load.
3.7.8 | IV pole must be provided with the unit
3.7.9 | Accessory rail must be provided on the rear mainframe.
3.7.10 | The unit must have 4 swivel castors of minimum diameter 100 mm and must be stable during movement. At least 2 castors must have locking/ brake mechanism.
3.8 | Resuscitator
3.8.1 | The crib must be equipped with an infant resuscitator integrated into the crib
3.8.2 | Unit must be gas powered via the wall oxygen supply
3.8.3 | Oxygen flow regulator: 5 - 15 l/min
3.8.4 | Adjustable Peak Inspiratory Pressure (PIP): 2 - 75 cm H2O
3.8.5 | Adjustable Positive End Expiratory Pressure (PEEP): 1 - 25 cm H2O
3.8.6 | Resuscitator must provide controlled pressure settings by means of a pressure limiting device
3.8.7 | Pressure gauge: -10 to 80 cm H2O
3.8.8 | Maximum pressure relief: 5 - 70cmH²O at 8lpm.
3.8.9 | The unit must be handed over in full operating order
3.9 | Weighing Scale
3.9.1 | The crib must be equipped with an weighing scale integrated into unit.
3.9.2 | The scale must have total weighing capacity of up to 8Kg
3.9.3 | It must have accuracy of ±5g
3.9.4 | It must have On/Off buttom with auto switch function when in use
3.9.5 | Hold function to retain the measured weight must be available
3.9.6 | Breastmilk intake measuring function for neonates is desireable in the unit to be offered
3.10 | Displays:
3.10.1 | Measured temperature: digital display
3.10.2 | Heater power
3.10.3 | Indicator of the selected mode (manual/control)
3.10.4 | Baby weight - scale
3.10.5 | Alarms
3.10.6 | Errors
4 | The following accessories must be supplied with the unit:
4.1 | Re-Usable Skin Temperature probe
4.2 | Disposable Skin temperature probe
4.3 | Disposable hydrophobic bacterial filters, box of 10
4.4 | IV pole
4.5 | Shelf for monitor
4.6 | Instrument tray
5 | Conformity Compliance (Please attach certificate)
5.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
5.2 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
5.2.1 | Please provide unique ref number of the ISO 13485 ceriticate:
5.3 | Model on offer must comply with EU Medical Device Regulation (EU) 2017/745 or equvalent from SAHPRA recognized jurisdictions . Attach a copy of certification
5.3.1 | Please provide unique ref number of the ceriticate:
5.4 | Hazardous Substance Act: If this type of equipment / apparatus appears on the schedule of Hazardous substances issued by the Directorate: Health Technology of the Department of Health, a license in terms of the Act on Hazardous Substances (Act 15/1973) must be submitted with the bid / quotation documents. The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.5 | All electrical/electronic medical equipment must be licensed by Radiation Control, where applicable a copy of the license must be submitted The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.5.1 | Radiation Control License number:
5.6 | The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade (rubber to with 2 screws) type and it must be a minimum length of three (3) meters. NB the mains cable of the unit tendered for must be SABS colour coded.
5.7 | The equipment quoted for must be protected against electromagnetic interference IEC 60601-1-2
5.8 | The unit must comply with electrical safety and performance standard such as IEC 60601-2-21, attached certification
5.8.1 | Please provide unique ref number of the IEC ceriticate:
6 | COMPREHENSIVE MAINTENANCE
6.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) should be included in the unit price as per the SCC document
6.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries and Oxygen sensor), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
6.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
6.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries and Oxygen sensor), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
6.3 | The bidder to provide the OEM recommended service interval of the unit:
6.4 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
7 | General
7.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
7.2 | The unit on offer must be the latest technology.
7.2.1 | Initial Manufacturing year of Model:
RT2-03-003 | Radiant Warmer-Incubator Combination | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
RT2-03-003a | Re-Usable Skin Temperature probe
RT2-03-003b | Disposable Skin temperature probe
RT2-03-003c | Disposable skin temp probes, box of 10
RT2-03-003d | IV pole
RT2-03-003e | Shelf for monitor
RT2-03-003f | Instrument tray
RT2-03-003g | Battery
RT2-03-003h | Air/oxygen blender (UPGRADE)
RT2-03-003i | Built-in scale (UPGRADE)
RT2-03-003j | LED Phototherapy (UPGRADE)
RT2-03-003k | Pulse Oximeter (UPGRADE)
RT2-03-003l | Suction Unit
RT2-03-003m | Oval sleeves (Pk of 4)
RT2-03-003n | Iris sleeves (Pk of 4)
RT2-03-003jo | Mattress
RT2-03-003p | Grommets as per number of ports

## ICU Crib
RT2-03-004 | ICU Crib | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
1 | Description
1.1 | A specially design hospital bed for critically ill newborns, infants and children requiring intensive monitoring and life-sustataining interventions.
2 | Product Offered:
2.1 | Make:
2.2 | Model:
2.3 | Initial Manufacturing year of Model:
2.4 | Radiation Control License number:
3 | Technical Specification
3.1 | The unit must comprise of the platform and mattress with the following features:
3.1.1 | Adjustable tilt to Trendellenburg and reverse Trendellenburg positions of at least 12 degrees
3.3.2 | Locking mechanism to keep the mattress platform locked in desired positions.
3.3.3 | Must accommodate X-ray film cassettes beneath the mattress.
3.3.4 | The mattress must be waterproof and must be X-ray translucent.
3.3.5 | Mattress size: approximately 500 x 650 mm(±50mm tolerance).
3.3.6 | Mattress thickness: at least 25 mm
3.3.7 | Mattress must support patient weight up to 4,5 kg.
3.3.8 | All four sides of the platform must be fitted with transparent guards/panels.
3.3.9 | The front and two side panels must be hinged and collapsable.
3.3.10 | Facility for holding patient cables, tubes, drains etc to prevent then being trapped or dislodged even when sides are opened
3.4 | Heater and Controls
3.4.1 | Radiant heater housed in an overhead canopy, with internal reflector, mounted on the rear mainframe.
3.4.2 | Indicator to show heater is switched ON
3.4.3 | The overhead canopy must swivel left, right then must tilt and lock in the desired position in order to continuously warm the baby. Bidders to specify the technology/mechanism.
3.4.4 | Microprocessor control with automatic self-test when switched on.
3.4.5 | The heat controller must provide the following modes for temperature control:
3.4.5.1 | Pre-heat
3.4.5.2 | Manual heat control: adjustable heater power from 0 - 100%.
3.4.5.3 | Skin temperature control: servo-control using a skin temperature probe
3.4.6 | A reusable skin temperature probe must be supplied
3.4.7 | Skin temperature probe cable length of at least 1m
3.4.8 | The skin temperature set point must be selectable in the range of 34 °C to 38 °C, in 0,1 °C increments.
3.4.9 | The skin temperature measurement range: 30 °C to 40 °C.
3.4.10 | Skin temperature measurement accuracy: 0,5 °C
3.4.11 | If a malfunction occurs affecting safety or performance the heater must be automatically turned off
3.5 | Displays:
3.5.1 | Heater power
3.5.2 | Set temperature
3.5.3 | Measured skin temperature
3.5.4 | Operation modes
3.5.5 | Apgar Timer
3.6 | Alarms: The following visual and audio alarms must be present for safety
3.6.1 | Power failure
3.6.2 | Apgar timer, with audible alarm, with digital display: 1, 5 and 10 minutes elapsed time.
3.6.3 | Skin temperature probe detachment
3.6.4 | Skin temperature probe failure
3.6.5 | Heater failure
3.6.6 | High and Low temperature if the skin temperature deviate by 1°C
3.7 | It must be possible to temporarily mute of audible alarms.
3.8 | Apgar timer with on/off switch and reset button
3.7 | Air/Oxygen Blender
3.7.1 | The unit must have intergrated Air/Oxygen blender
3.7.2 | Deliver Oxygen Range 21 % to 100 %
3.7.3 | Flow Range 2 to 15 lpm.
3.8 | Other
3.8.1 | The overhead canopy must be provided with an examination lamp, with on/off switch.
3.8.2 | The ICU crib must be supplied with an oxygen cylinder holder.
3.8.3 | The ICU crib must have at least one drawer
3.8.4 | The minimum acceptable height of the heater canopy must be at least 1.85 metres above floor
3.8.5 | The minimum acceptable height of the mattress platform must 900 mm above the floor.
3.8.6 | It must have an Instrument tray, approximately 200 x 300 mm and must hold up to 2 kg load.
3.8.7 | It must have a swivel mounted monitor shelf, approximately 300 x 400 mm, and must hold at least 5 kg load.
3.8.8 | IV pole must be provided
3.8.9 | Accessory rail must be provided on the rear mainframe.
3.8.10 | The unit must have 4 swivel castors of minimum diameter 100 mm and must be stable during movement. At least 2 castors must have locking/ brake mechanism.
3.9 | Resuscitator
3.9.1 | The crib must be equipped with an infant resuscitator integrated into the crib
3.9.2 | Unit must be gas powered via the wall oxygen supply
3.9.3 | Oxygen flow regulator: 5 - 15 l/min
3.9.4 | Adjustable Peak Inspiratory Pressure (PIP): 2 - 75 cm H2O
3.9.5 | Adjustable Positive End Expiratory Pressure (PEEP): 1 - 25 cm H2O
3.9.6 | Resuscitator must provide controlled pressure settings by means of a pressure limiting device
3.9.7 | Pressure gauge: -10 to 80 cm H2O
3.9.8 | Maximum pressure relief: 5 - 70cmH²O at 8lpm.
3.10 | A starter-pack of consumables must be supplied with the unit.
3.11 | Weighing Scale
3.11.1 | The crib must be equipped with an weighing scale integrated into unit.
3.11.2 | The scale must have total weighing capacity of up to 8Kg
3.11.3 | It must have accuracy of ±5g
3.11.4 | It must have On/Off buttom with auto switch function when in use
3.11.5 | Hold function to retain the measured weight must be available
3.11.6 | Breastmilk intake measuring function for neonates is desireable in the unit to be offered
4 | The following accessories must be supplied with the unit:
4.1 | Re-usable Skin Temperature probe
4.2 | Disposable Skin Temperature probe pack of 10
4.3 | Disposable hydrophobic bacterial filters, box of 10
4.4 | IV pole
4.5 | Shelf for monitor
4.6 | Instrument tray
4.7 | Disposable Patient resuscitator circuits pack of 10
5 | Conformity Compliance (Please attach certificate)
5.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
5.2 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
5.2.1 | Please provide unique ref number of the ISO 13485 ceriticate:
5.3 | Model on offer must comply with EU Medical Device Regulation (EU) 2017/745 or equvalent from SAHPRA recognized jurisdictions . Attach a copy of certification
5.3.1 | Please provide unique ref number of the ceriticate:
5.4 | Hazardous Substance Act: If this type of equipment / apparatus appears on the schedule of Hazardous substances issued by the Directorate: Health Technology of the Department of Health, a license in terms of the Act on Hazardous Substances (Act 15/1973) must be submitted with the bid / quotation documents. The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.5 | All electrical/electronic medical equipment must be licensed by Radiation Control, where applicable a copy of the license must be submitted The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.5.1 | Radiation Control License number:
5.6 | The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade (rubber to with 2 screws) type and it must be a minimum length of three (3) meters. NB the mains cable of the unit tendered for must be SABS colour coded.
5.7 | The equipment quoted for must be protected against electromagnetic interference IEC 60601-1-2
5.8 | The unit must comply with electrical safety and performance standard such as IEC 60601-2-21, attached certification
5.8.1 | Please provide unique ref number of the IEC ceriticate:
6 | COMPREHENSIVE MAINTENANCE
6.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) should be included in the unit price as per the SCC document
6.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
6.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
6.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
6.3 | The bidder to provide the OEM recommended service interval of the unit:
6.4 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
7 | General
7.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
7.2 | The unit on offer must be the latest technology.
7.2.1 | Initial Manufacturing year of Model:
RT2-03-004 | ICU Crib | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
RT2-03-004a | Re-usable Skin Temperature probe
RT2-03-004b | Disposable Skin Temperature probe pack of 10
RT2-03-004c | IV pole
RT2-03-004d | Shelf for monitor
RT2-03-004e | Instrument tray
RT2-03-004f | Disposable Patient resuscitator circuits pack of 10
RT2-03-004h | Air/oxygen blender
RT2-03-004i | Built-in scale
RT2-03-004j | LED Phototherapy (UPGRADE)
RT2-03-004l | Pulse Oxymeter (UPGRADE)
RT2-03-004m | Intergrated suction (UPGRADE)

## Incubator, ICU
RT2-03-005 | Incubator ICU | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
1 | Description
1.1 | A medical device designed to provide a controlled environment for newborn babies by circulating heated air around them to promote healthy growth and development.
2 | Product Offered:
2.1 | Make:
2.2 | Model:
2.3 | Initial Manufacturing year of Model:
2.4 | Radiation Control License number:
3 | Technical Specification
3.1 | The control module must provide an even distribution of heat and humidity within the hood
3.2 | The unit must perform an automatic self test when switched on.
3.3 | The incubator frame and chassis must be robust
3.5 | Must be fitted with a robust transparent double hood. Approximate dimensions (LxWxH): 900 x 450 x 430 mm
3.6 | The front of the hood must have a full width transparent access panel, to allow the mattress tray to be extended outside the hood enclosure for examination and treatment of the infant without raising the hood completely .
3.7 | When the mattress tray is extended outside the hood enclosure, it must remain attached to the incubator and support the weight of the baby.
3.8 | The hood must be hinged to allow easy access for bedding the infant.
3.9 | The hood must have a safety latch to prevent accidental opening
3.10 | The hood must remain in a stable position when opened
3.11 | The hood must have at least four arm openings with robust latches that are elbow operated. Two of these openings must be on the front and two at the rear of the hood
3.12 | The hood must also have two iris openings, on the left and right hand sides
3.13 | The hood must have 2 small self-sealing ports, for IV tubes etc.
3.14 | The hood must seal well onto the chassis frame to avoid any form of heat loss.
3.15 | It must be possible to easily manipulate the bed into the Trendelenberg and reverse Trendelenberg positions.
3.16 | The unit must have an IV pole and a monitor stand.
3.17 | The monitor stand must support a mass of at least 5kg
3.18 | The incubator must be easy to clean, disinfect, and/or sterilize
3.19 | Switches and controls must be protected against accidental setting changes
3.20 | TEMPERATURE CONTROL:
3.20.1 | The control unit must be microprocessor-based.
3.20.2 | The control module must have servo heat control mode
3.20.3 | The unit must not take more than 30 minutes from power-on to reach the desired temperature setting.
3.20.4 | User selectable air temperature and skin temperatures must be available.
3.20.5 | Air temperature must be selectable from ± 25°C to ± 37°C with tolerance of ±2°C
3.20.6 | Air temperature selectable above 39°C must have override confirmation
3.20.7 | Skin temperature must be selectable from ± 34°C to ± 37°C.
3.20.8 | Skin temperature selectable above 39°C must have override confirmation
3.20.9 | Temperature variation less than +/- 0.2°C
3.20.10 | A temperature sensor must be available in the hood for temperature control and monitoring.
3.20.11 | In addition to the primary sensor for controlling the temperature, the unit must have a secondary fail-safe sensor which will act as a safety back-up and activate audible and visible alarms and shut off power to the heater if the temperature exceeds 40oC.
3.20.12 | The supply to the heater may only be restored after a manual reset and if the incubator’s secondary temperature monitoring device senses that the internal temperature has decreased to a safe level.
3.20.13 | At switch-on the incubator must automatically heat to 34°C +/- 1°C
3.20.14 | When the hood’s front access panel is opened, it must have a minimal effect on the set temperature within the incubator.
3.20.15 | Reusable skin temperature probe must be supplied
3.21 | DISPLAY
3.21.1 | Display for the following:
3.21.1a | Skin temperature.
3.21.1b | Air temperature.
3.21.1c | Set temperature
3.21.1d | Relative humidity
3.21.1e | Details of all alarms
3.21.2 | Display must be large and clear enough to be viewed from three metres under all lighting conditions.
3.22 | Alarms: The following visual and audio alarms must be present for safety
3.22.1 | Default alarms must be enabled when the incubator is turned on
3.22.2 | Heater power cut-off in the event of:
3.22.2a | Over temperature
3.22.2b | Failure of air circulation
3.23 | Skin temperature differs from set temperature by +/-1°C
3.23.1 | Air temperature differs from set temperature by +/-1°C
3.23.2 | Low humidity
3.23.3 | Power failure
3.23.4 | Heater failure
3.23.5 | Skin or air probe disconnection or failure.
3.23.6 | The audible alarm must have a mute option. Visual alarm may not be reset until the cause of the alarm has been rectified
3.23.7 | Test facility for audible and visual alarms.
3.23.8 | Low water alarm while humidity function is switched on
3.24 | GENERAL
3.24.1 | Exhaled air must continuously be removed to prevent carbon dioxide build-up inside the hood.
3.24.2 | Fresh air drawn to the hood must be filtered by an easily replaceable bacterial filter
3.24.3 | The mattress must be approximately 630 x 350mm(±50mm tolerance) with thickness of 30 mm and must be very durable, comfortable and have a waterproof cover.
3.24.4 | The sound level within the baby compartment must not exceed 45 dB(A) except if an audible alarm is sounding, then the sound level within the baby compartment must not exceed 80dB(A).
3.24.5 | The air velocity above the mattress must not exceed 0.35 m/s.
3.24.8 | All metal parts of the incubator must be well earthed.
3.24.9 | The incubator’s control module must be well designed to guard against entry of liquids
3.24.10 | All parts on the incubator and trolley must be rust resistant
3.24.11 | X-Ray tray incorporated without need to open incubator
3.24.12 | Rail for accessories on head side
3.24.14 | Labels and markings must be durable enough to withstand routine cleaning and normal wear
3.24.15 | Appropriate warning legends must be provided on the unit.
3.24.16 | The unit must be simple to use and maintain.
3.24.17 | The unit must be designed for easy access to serviceable parts
3.24.18 | The incubator must have no sharp edges.
3.24.19 | The unit must be mounted on a robust trolley
3.24.20 | Trolley must have 100 mm diameter (minimum) antistatic swivel castors.
3.24.21 | At least two of the castors must have brakes.
3.24.22 | Handles for pushing the trolley must be available
3.24.23 | Holder for a standard 1 kg oxygen cylinder must be mounted to the trolley.
3.24.24 | Switches, knobs, and other controls must be designed to withstand conditions of heavy use.
3.24.25 | At least Two(2) storage drawer must be provided.
3.26 | A starter-pack of consumables must be supplied with the unit.
4 | The following accessories must be supplied with the unit:
4.1 | Bacterial filter (Pk of 4)
4.2 | Skin temperature probe (reusable)
4.3 | Oval sleeves (Pk of 4)
4.4 | Iris sleeves (Pk of 4)
4.5 | Mattress
4.6 | Grommets as per number of ports
5 | Conformity Compliance (Please attach valid certificate with ref number and expiry date)
5.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
5.2 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
5.2.1 | Please provide unique ref number of the ISO 13485 ceriticate:
5.3 | Model on offer must comply with EU Medical Device Regulation (EU) 2017/745 or equvalent from SAHPRA recognized jurisdictions . Attach a copy of certification
5.3.1 | Please provide unique ref number of the ceriticate:
5.4 | Hazardous Substance Act: If this type of equipment / apparatus appears on the schedule of Hazardous substances issued by the Directorate: Health Technology of the Department of Health, a license in terms of the Act on Hazardous Substances (Act 15/1973) must be submitted with the bid / quotation documents. The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.5 | All electrical/electronic medical equipment must be licensed by Radiation Control, where applicable a copy of the license must be submitted The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.5.1 | Radiation Control License number:
5.6 | The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade (rubber to with 2 screws) type and it must be a minimum length of three (3) meters. NB the mains cable of the unit tendered for must be SABS colour coded.
5.7 | The equipment quoted for must be protected against electromagnetic interference IEC 60601-1-2
5.8 | The unit must comply with electrical safety and performance standard such as IEC 60601-2-21, attached certification
5.8.1 | Please provide unique ref number of the IEC ceriticate:
6 | COMPREHENSIVE MAINTENANCE
6.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) should be included in the unit price as per the SCC document
6.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
6.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
6.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
6.3 | The bidder to provide the OEM recommended service interval of the unit:
6.4 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
7 | General
7.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
7.2 | The unit on offer must be the latest technology.
7.2.1 | Initial Manufacturing year of Model:
RT2-03-005 | Incubator, ICU Accessories | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
RT2-03-005a | Bacterial filter (Pk of 10)
RT2-03-005b | Iris port sleeves (Pk of 10)
RT2-03-005c | Oval port sleeves (Pk of 10)
RT2-03-005d | Skin temperature probe (reusable)(Pk of 5)
RT2-03-005e | Disposable Skin temperature probe (Pk of 10)
RT2-03-005f | Covers for temperature probe (Pk of 25)
RT2-03-005g | Internal battery for alarms
RT2-03-005h | Scale (UPGRADE)
RT2-03-005i | Observation lamp (UPGRADE)
RT2-03-005j | Mattress cover
RT2-03-005k | X-ray tray
RT2-03-005l | Accessory rail
RT2-03-005m | Sensor module
RT2-03-005n | IV Pole

## Incubator, Standard
RT2-03-006 | Infant Incubator (Standard) | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
1 | Description
1.1 | A medical device designed to provide a controlled environment for newborn babies by circulating heated air around them to promote healthy growth and development.
2 | Product Offered:
2.1 | Make:
2.2 | Model:
2.3 | Initial Manufacturing year of Model:
2.4 | Radiation Control License number:
3 | Technical Specification
3.1 | The control module must provide an even distribution of heat and humidity inside the hood
3.2 | The unit must perform an automatic self test when switched on.
3.3 | The incubator frame and chassis must be robust
3.4 | Must be fitted with a robust transparent double hood. Approximate dimensions (LxWxH): 900 x 450 x 430 mm
3.5 | The front of the hood must have a full width transparent access panel, to allow the mattress tray to be extended outside the hood enclosure for examination and treatment of the infant without raising the hood completely .
3.6 | When the mattress tray is extended outside the hood enclosure, it must remain attached to the incubator and support the weight of the baby.
3.7 | The hood must be hinged to allow easy access for bedding the baby.
3.8 | The hood must have a safety latch to prevent accidental opening
3.9 | The hood must remain in a stable position when opened
3.10 | The hood must have at least four arm openings with robust latches that are elbow operated. Two of these openings must be on the front and two at the rear of the hood
3.11 | The hood must also have two iris openings, on the left and right hand sides
3.12 | The hood must have at least 4 small self-sealing ports, for IV tubes etc.
3.13 | The hood must seal well onto the chassis frame to avoid heat loss.
3.14 | Easy manipulation of the bed into Trendelenberg and reverse Trendelenberg positions.
3.15 | The unit must have an IV pole
3.16 | The incubator must be easy to clean, disinfect, and/or sterilize
3.17 | Switches and controls must be protected against accidental setting changes
3.18 | TEMPERATURE CONTROL:
3.18.1 | The control unit must be microprocessor-based.
3.18.2 | The control module must have servo heat control modes
3.18.3 | The unit must not take more than 30 minutes from power-on to reach the desired temperature setting.
3.18.4 | User selectable air temperature and skin temperatures.
3.18.5 | Air temperature must be selectable from ± 25°C to ± 37°C with tolerance of ±2°C
3.18.6 | Air temperature selectable above 39°C must have override confirmation
3.18.7 | Skin temperature must be selectable from 34°C to 37°C.
3.18.8 | Skin temperature selectable above 39°C must have override confirmation
3.18.9 | Temperature variation less than +/- 0.2°C
3.18.10 | A temperature sensor must be available in the hood for temperature control and monitoring.
3.18.11 | In addition to the primary sensor for controlling the temperature, the unit must have a secondary fail-safe sensor which will act as a safety back-up and activate audible and visible alarms and shut off power to the heater if the temperature exceeds 40oC.
3.18.12 | The supply to the heater may only be restored after a manual reset and if the incubator’s secondary temperature monitoring device senses that the internal temperature has decreased to a safe level.
3.18.13 | At switch-on the incubator must automatically heat to 34°C +/- 1°C
3.18.14 | When the hood’s front access panel is opened, it must have a minimal effect on the set temperature within the incubator.
3.18.15 | Reusable skin temperature probe must be supplied
3.19 | DISPLAY
3.19.1 | LCD or LED display for the following:
3.19.1a | Skin temperature.
3.19.1b | Air temperature.
3.19.1c | Set temperature
3.19.1d | Details of all alarms
3.19.1e | Display must be large and clear enough to be viewed from three metres under all lighting conditions.
3.20 | Alarms: The following visual and audio alarms must be present for safety
3.20.1 | Default alarms must be enabled when the incubator is turned on
3.20.2 | Heater power cut-off in the event of:
3.20.3 | Over temperature
3.20.4 | Failure of air circulation
3.20.5 | Skin temperature differs from set temperature by +/-1°C
3.20.6 | Air temperature differs from set temperature by +/-1°C
3.20.7 | Low humidity
3.20.8 | Power failure
3.20.9 | Heater failure
3.20.10 | Skin or air probe disconnection or failure.
3.20.11 | The audible alarm must have a mute option. Visual alarm may not be reset until the cause of the alarm has been rectified
3.20.12 | Test facility for audible and visual alarms.
3.20.13 | Low water alarm while humidity function is switched on
3.21 | GENERAL
3.21.1 | The air velocity above the mattress must not exceed 0.35 m/s.
3.21.2 | Exhaled air must continuously be removed to prevent carbon dioxide build-up inside the hood.
3.21.3 | Fresh air drawn to the hood must be filtered by an easily replaceable bacterial filter
3.21.4 | The mattress must be approximately 630 x 350mm(±50mm tolerance) with thickness of 30 mm and must be very durable, comfortable and have a waterproof cover.
3.21.5 | The sound level within the baby compartment must not exceed 45 dB(A) except if an audible alarm is sounding, then the sound level within the baby compartment must not exceed 80 dB(A).
3.21.6 | All metal parts of the incubator must be well earthed.
3.21.7 | The incubator’s control module must be well designed to guard against entry of liquids
3.21.8 | All parts on the incubator and trolley must be rust resistant
3.21.9 | Labels and markings must be durable enough to withstand routine cleaning and normal wear
3.21.10 | The unit must be simple to use and maintain.
3.21.11 | The unit must be mounted on a robust trolley
3.21.12 | Trolley must have 100 mm diameter (minimum) antistatic swivel castors.
3.21.13 | At least two of the castors must have brakes.
3.21.14 | Handles for pushing the trolley must be available
3.21.15 | The unit must be designed for easy access to serviceable parts
3.21.16 | The incubator must have no sharp edges.
3.21.17 | Switches, knobs, and other controls must be designed to withstand conditions of heavy use.
3.21.18 | At least Two(2) storage drawer must be provided.
3.21.19 | A starter-pack of consumables must be supplied with the unit.
4 | The following accessories must be supplied with the unit:
4.1 | Bacterial filter (Pk of 4)
4.2 | Skin temperature probe (reusable)
4.3 | Oval sleeves (Pk of 4)
4.4 | Iris sleeves (Pk of 4)
4.5 | Mattress
4.6 | Grommets as per number of ports
5 | Conformity Compliance (Please attach valid certificate with ref number and expiry date)
5.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
5.2 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
5.2.1 | Please provide unique ref number of the ISO 13485 ceriticate:
5.3 | Model on offer must comply with EU Medical Device Regulation (EU) 2017/745 or equvalent from SAHPRA recognized jurisdictions . Attach a copy of certification
5.3.1 | Please provide unique ref number of the ceriticate:
5.4 | Hazardous Substance Act: If this type of equipment / apparatus appears on the schedule of Hazardous substances issued by the Directorate: Health Technology of the Department of Health, a license in terms of the Act on Hazardous Substances (Act 15/1973) must be submitted with the bid / quotation documents. The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.5 | All electrical/electronic medical equipment must be licensed by Radiation Control, where applicable a copy of the license must be submitted The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.5.1 | Radiation Control License number:
5.6 | The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade (rubber to with 2 screws) type and it must be a minimum length of three (3) meters. NB the mains cable of the unit tendered for must be SABS colour coded.
5.7 | The equipment quoted for must be protected against electromagnetic interference IEC 60601-1-2
5.8 | The unit must comply with electrical safety and performance standard such as IEC 60601-2-21, attached certification
5.8.1 | Please provide unique ref number of the IEC ceriticate:
6 | COMPREHENSIVE MAINTENANCE
6.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) should be included in the unit price as per the SCC document
6.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
6.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
6.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
6.3 | The bidder to provide the OEM recommended service interval of the unit:
6.4 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
7 | General
7.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
7.2 | The unit on offer must be the latest technology.
7.2.1 | Initial Manufacturing year of Model:
RT2-03-006 | Incubator, Standard Accessories | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
RT2-03-006a | Bacterial filter (Pk of 10)
RT2-03-006b | Iris port sleeves (Pk of 10)
RT2-03-006c | Oval port sleeves (Pk of 10)
RT2-03-006d | Skin temperature probe (reusable)(Pk of 5)
RT2-03-006e | Disposable Skin temperature probe (Pk of 10)
RT2-03-006f | Covers for temperature probe (Pk of 25)
RT2-03-006g | Internal battery for alarms
RT2-03-006h | Scale (UPGRADE)
RT2-03-006i | Observation lamp (UPGRADE)
RT2-03-006j | Mattress cover
RT2-03-006k | X-ray tray
RT2-03-006l | Accessory rail
RT2-03-006m | Sensor module
RT2-03-006n | IV Pole

## Incubator, Transport
RT2-03-007 | Incubator, Transport | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
1 | Description
1.1 | A medical device designed to provide a controlled environment for newborn babies by circulating heated air around them to promote healthy growth and development.
2 | Product Offered:
2.1 | Make:
2.2 | Model:
2.3 | Initial Manufacturing year of Model:
2.4 | Radiation Control License number:
3 | Technical Specification
3.1 | The unit offered must be capable of providing thermal support during transportation even in adverse weather conditions maintaining air temperature for at least 3 hours when operating on battery power.
3.2 | Internal noise levels in the unit offered must be less than 50 dB
3.3 | Control through microprcessed panel with environmental temperature wide view displays
3.4 | Air Controlled mode temperature range 20 - 39⁰ C
3.5 | Servo Controlled mode temperature range 34- 37⁰ C
3.6 | Manual modes should have 15 second interval alarms on overrides.
3.7 | Temperature measurements displayed in increments of 0.1⁰C
3.8 | Required Alarms :
3.8.1 | High Air Temperature
3.8.2 | Low Air Temperature
3.8.3 | High Skin Temperature
3.8.4 | Low Skin Temperature
3.8.5 | Fan Failure
3.8.6 | Sensor Failure
3.8.7 | Power Failure
3.8.8 | Battery power level and when recharge is required.
3.9 | Battery Requirements:
3.9.1 | Battery must not take longer than 16 hours to recharge
3.9.2 | Unit must automatically switch over to battery power ones removed from line power.
3.9.3 | The bidder must indicate type of battery and life expectancy
3.10 | Minimum required ports :
3.10.1 | 2 Hand Ports
3.10.2 | 6 Tubing Ports
3.10.3 | 1 Oxygen inlet port
3.11 | Castors' Requirements
3.11.1 | Castor Size ≥ 12.7 cm
3.11.2 | At least 2 of the castors must have locks/brake mechanism.
3.11.3 | Castors must provide easy manuvrering.
3.11.4 | Space should be provided to store and secure two gas Cylinders
3.11.5 | Aluminum stand with three positions system coupling for in-hospital transport purposes.
4 | Standard Accessories
4.1 | Bacterial filter (Pk of 4)
4.2 | Skin temperature probe (reusable)
4.3 | Covers for temperature probe (Pk of 25)
4.4 | Oval sleeves (Pk of 4)
4.5 | Iris sleeves (Pk of 4)
4.6 | Mattress
4.7 | Grommets as per number of ports
4.8 | Rechargeable Battery
4.9 | Charging Cable Ambulance power point
5 | Conformity Compliance (Please attach certificate)
5.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
5.2 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
5.2.1 | Please provide unique ref number of the ISO 13485 ceriticate:
5.3 | Model on offer must comply with EU Medical Device Regulation (EU) 2017/745 or equvalent from SAHPRA recognized jurisdictions . Attach a copy of certification
5.3.1 | Please provide unique ref number of the ceriticate:
5.4 | Hazardous Substance Act: If this type of equipment / apparatus appears on the schedule of Hazardous substances issued by the Directorate: Health Technology of the Department of Health, a license in terms of the Act on Hazardous Substances (Act 15/1973) must be submitted with the bid / quotation documents. The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.5 | All electrical/electronic medical equipment must be licensed by Radiation Control, where applicable a copy of the license must be submitted The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.5.1 | Radiation Control License number:
5.6 | The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade (rubber to with 2 screws) type and it must be a minimum length of three (3) meters. NB the mains cable of the unit tendered for must be SABS colour coded.
5.7 | The equipment quoted for must be protected against electromagnetic interference IEC 60601-1-2
5.8 | The unit must comply with electrical safety and performance standard such as IEC 60601-2-21, attached certification
5.8.1 | Please provide unique ref number of the IEC ceriticate:
6 | COMPREHENSIVE MAINTENANCE
6.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) should be included in the unit price as per the SCC document
6.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
6.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
6.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
6.3 | The bidder to provide the OEM recommended service interval of the unit:
6.4 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
7 | General
7.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
7.2 | The unit on offer must be the latest technology.
7.2.1 | Initial Manufacturing year of Model:
RT2-03-007 | Incubator, Transport Accessories and Consumables | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
RT2-03-007a | Bacterial filter (Pk of 4)
RT2-03-007b | Iris port sleeves (Pk of 10)
RT2-03-007c | Disposable Iris port sleeves (Pk of 10)
RT2-03-007d | Oval port sleeves (Pk of 10)
RT2-03-007e | Skin temperature probe (reusable)
RT2-03-007f | Covers for temperature probe (Pk of 25)
RT2-03-007g | Disposable skin temperature probe (Pk of 10)
RT2-03-007h | Examination lamp bulb
RT2-03-007i | Ambulance collapsable incubator stand/trolley
RT2-03-007j | Mattress
RT2-03-007k | Grommets as per number of ports
RT2-03-007l | Rechargeable Battery
RT2-03-007m | Charging Cable Ambulance power point
RT2-03-007n | IV Pole

## Infant Radiant Warmer
RT2-03-008 | Infant Radiant warmer | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
1 | Description
1.1 | A warming device to provide uniform heat to the body using an overhead heating source.
2 | Product Offered:
2.1 | Make:
2.2 | Model:
2.3 | Initial Manufacturing year of Model:
2.4 | Radiation Control License number:
3 | Technical Specification
3.1 | The unit must offer long-term care of a neonate
3.2 | Platform and mattress:
3.2.1 | Adjustable tilt to Trendelenburg and reverse Trendelenburg positions at least 12 degrees
3.2.2 | Locking mechanism to keep the mattress platform locked in all positions.
3.2.3 | The mattress must be waterproof and must be X-ray translucent.
3.2.4 | Mattress size: approximately 500 x 650 mm (±50mm tolerance).
3.2.5 | Mattress thickness: at least 25 mm
3.2.6 | Mattress must support patient weight up to 4,5 kg.
3.2.7 | All four sides of the platform must be fitted with transparent guards/panels.
3.2.8 | The front and two side panels must be hinged and easily collapsable.
3.3 | Heater
3.3.1 | Radiant heater must be housed in an overhead canopy, with internal reflector, mounted on the rear mainframe.
3.3.2 | It must have an nndicator to show heater is switched ON
3.3.3 | The overhead canopy must swivel left and right and must lock in the desired position.
3.3.4 | The heater must be able to heat while at a swivel and tilted position.
3.4 | Controls
3.4.1 | Heat control: must be adjustable heater power from 0 - 100%.
3.4.2 | If a malfunction occurs affecting safety or performance the heater must be automatically turned off
3.4.3 | Skin temperature probe to monitor infant temperature.
3.4.4 | Heat output must be controlled manually or through baby mode (feedback mode) for thermoregulation.
3.5 | Displays :
3.5.1 | Heater power
3.5.2 | Set temperature
3.5.3 | Measured skin temperature
3.5.4 | Operation modes
3.5.5 | Apgar Timer
3.6 | Alarms: The following visual and audio alarms must be present for safety
3.6.1 | Power failure
3.6.2 | Apgar timer, with audible alarm, with digital display: 1, 5 and 10 minutes elapsed time.
3.6.3 | Skin temperature probe detachment
3.6.4 | Skin temperature probe failure
3.6.5 | Heater failure
3.6.6 | High and Low temperature
3.7 | Apgar timer with on/off switch and reset button
3.8 | Other
3.8.1 | The unit must be provided with an examination lamp, with on/off switch to provide adequate lighting to the mattress platform.
3.8.2 | The infant radiant warmer must have a minimum of one drawer
3.8.3 | The minimum acceptable height of the heater canopy must be 1.85 metres above the floor
3.8.4 | Adjustable mattress platform height from the floor must be at least 900 mm
3.8.5 | The mains cable of the unit must be at least three metres long and be fitted with an SABS approved plug and wire colour code
3.8.7 | the unit must have 4 swivel castors of minimum diameter 100 mm and must be stable during movement.
3.8.8 | Two of the castors must be lockable.
3.8.9 | The unit must have oxygen cylinder holder
4 | The following accessories must be supplied with the unit:
4.1 | Reusable Temperature probe x1
4.2 | Disposable Tempeture probe pack of 10
4.3 | Mattress
4.4 | Accessory rail
4.5 | Monitor stand
4.6 | IV Pole
5 | Conformity Compliance (Please attach certificate)
5.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
5.2 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
5.2.1 | Please provide unique ref number of the ISO 13485 ceriticate:
5.3 | Model on offer must comply with EU Medical Device Regulation (EU) 2017/745 or equvalent from SAHPRA recognized jurisdictions . Attach a copy of certification
5.3.1 | Please provide unique ref number of the ceriticate:
5.4 | Hazardous Substance Act: If this type of equipment / apparatus appears on the schedule of Hazardous substances issued by the Directorate: Health Technology of the Department of Health, a license in terms of the Act on Hazardous Substances (Act 15/1973) must be submitted with the bid / quotation documents. The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.5 | All electrical/electronic medical equipment must be licensed by Radiation Control, where applicable a copy of the license must be submitted The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.5.1 | Radiation Control License number:
5.6 | The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade (rubber to with 2 screws) type and it must be a minimum length of three (3) meters. NB the mains cable of the unit tendered for must be SABS colour coded.
5.7 | The equipment quoted for must be protected against electromagnetic interference IEC 60601-1-2
5.8 | The unit must comply with electrical safety and performance standard such as IEC 60601-2-21, attached certification
5.8.1 | Please provide unique ref number of the IEC ceriticate:
6 | COMPREHENSIVE MAINTENANCE
6.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) should be included in the unit price as per the SCC document
6.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
6.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
6.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
6.3 | The bidder to provide the OEM recommended service interval of the unit:
6.4 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
7 | General
7.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
7.2 | The unit on offer must be the latest technology.
7.2.1 | Initial Manufacturing year of Model:
RT2-03-008 | Infant radiant warmer Accessories | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
RT2-03-008a | Reusable Temperature probe
RT2-03-008b | Mattress
RT2-03-008c | LED Phototherapy (UPGRADE)
RT2-03-008d | IV pole
RT2-03-008e | Disposable temperature probe (Pk of 10)
RT2-03-008f | Accessory rail
RT2-03-008g | Monitor stand
RT2-03-008h | Collapsable side guards/panels
RT2-03-008i | Collapsable front guards/panels
RT2-03-008j | Rear guards/panels
RT2-03-008k | Heater element
RT2-03-008l | Hinges x4

## Phototherapy Lamp, LED
RT2-03-009 | Phototherapy Light (LED) | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
1 | Description
A medical device emitting specific wavelengths of lights to treat jaundice skin conditions in newborns.
2 | Product Offered:
2.1 | Make:
2.2 | Model:
2.3 | Initial Manufacturing year of Model:
3 | Technical Specification
3.1 | The phototherapy light system must be suitable for the treatment of neonatal hyperbilirubinemia.
3.2 | Microprocessor control with LCD to display set treatment time, lamp usage time and intensity level.
3.3 | The light system must be suitable for babies in a bassinet, incubator, open incubator or radiant warmer.
3.4 | The light source must be LEDs, emiting light in the range of 450 – 475nm
3.5 | Target light to ensure correct placement of lamp over patient
3.6 | The unit must have at least two intensity settings, high and low.
3.7 | The high setting must give an output of 30-45 µW/cm²/nm at a distance of 300 mm from the light enclosure.
3.8 | The low setting must give an output of 12-15 µW/cm²/nm at a distance of 300 mm from the light enclosure.
3.9 | The LED light source must ensure a uniform illumination pattern at the skin surface
3.10 | The light source housing must tilt at least to 30o in either direction to provide effective phototherapy
3.11 | The lamp effective surface area must be at least 400 x 200 mm at 300 mm from the light source
3.12 | The unit must be height adjustable : approximately 1200 - 1600 mm from light source to floor
3.13 | The light enclosure must be protected from the ingress of fluids.
3.14 | The light unit must be mounted on a roll stand.
3.15 | The roll stand must have lockable castors.
3.16 | The base of the roll stand must not tip, even when the light enclosure is tilted.
3.17 | Must have a re-settable treatment timer that will switch off the light when the set time has expired
3.18 | The unit must have a time recorder that operates only when the phototherapy light is switched on. It must be possible to reset the time recorder to zero when the light source is replaced.
3.19 | It must not be possible to accidently reset the time recorder
3.20 | LED array replacement must be a simple operation
3.21 | It must be possible to replace only a section of the LED array without comprimising the intensity of lights
4 | Conformity Compliance (Please attach certificate)
4.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
4.2 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
4.2.1 | Please provide unique ref number of the ISO 13485 ceriticate:
4.3 | Model on offer must comply with EU Medical Device Regulation (EU) 2017/745 or equvalent from SAHPRA recognized jurisdictions . Attach a copy of certification
4.3.1 | Please provide unique ref number of the ceriticate:
4.4 | Hazardous Substance Act: If this type of equipment / apparatus appears on the schedule of Hazardous substances issued by the Directorate: Health Technology of the Department of Health, a license in terms of the Act on Hazardous Substances (Act 15/1973) must be submitted with the bid / quotation documents. The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
4.5 | All electrical/electronic medical equipment must be licensed by Radiation Control, where applicable a copy of the license must be submitted The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
4.5.1 | Radiation Control License number:
4.6 | The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade (rubber to with 2 screws) type and it must be a minimum length of three (3) meters. NB the mains cable of the unit tendered for must be SABS colour coded.
4.7 | The equipment quoted for must be protected against electromagnetic interference IEC 60601-1-2
4.8 | The unit must comply with electrical safety and performance standard such as IEC 60601-2-21, attached certification
4.8.1 | Please provide unique ref number of the IEC ceriticate:
5 | COMPREHENSIVE MAINTENANCE
5.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) should be included in the unit price as per the SCC document
5.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
5.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
5.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
5.3 | The bidder to provide the OEM recommended service interval of the unit:
5.4 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
6 | General
6.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
6.2 | The unit on offer must be the latest technology.
6.2.1 | Initial Manufacturing year of Model:
RT2-03-009 | Phototherapy Light (LED) Accessories and Consumables | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
RT2-03-009a | Monitoring of baby temperature (UPGRADE)
RT2-03-009b | LED examination light/Bar replacement

## Phototherapy Blanket, LED
RT2-03-010 | Phototherapy Blanket LED | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
1 | Description
A flexible blanket-like pad with embedded LED lights that emit specific wavelength of light to treat jaundice.
2 | Product Offered:
2.1 | Make:
2.2 | Model:
2.3 | Initial Manufacturing year of Model:
3 | Technical Specification
3.1 | The phototherapy light system must be suitable for the treatment of neonatal hyperbilirubinemia.
3.2 | Microprocessor control with LCD to display set treatment time, therapy time, lamp usage time and intensity level.
3.3 | The light system must be suitable for babies in a bassinet, incubator, or radiant warmer.
3.4 | The light source must be LEDs, emiting light in the range of 450 – 475nm
3.5 | The unit must have at least two intensity settings, high and low.
3.6 | Intensity: Maximum 40-50uW/cm2/nm and Minimum 30-35uW/cm2/nm
3.7 | Effective treatment area: Small pad of 250-350cm2
3.8 | The unit must not emit light in the Ultraviolet or Infrared radiation range
3.9 | The unit must consist of a soft, flexible, oval design that conforms to the shape of the baby to deliver phototherapy via an LED light source
3.10 | The unit must be able to shut off automatically in the event of the blanket overheating
3.11 | Must have a re-settable treatment timer that will switch off the light when the set time has expired
3.12 | The unit must have a time recorder that operates only when the phototherapy light is switched on. It must be possible to reset the time recorder to zero when the light source is replaced.
3.13 | It must not be possible to accidently reset the time recorder
3.14 | The unit must operate off the 220 Volt, 50Hz power supply.
4 | Conformity Compliance (Please attach certificate)
4.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
4.2 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
4.2.1 | Please provide unique ref number of the ISO 13485 ceriticate:
4.3 | Model on offer must comply with EU Medical Device Regulation (EU) 2017/745 or equvalent from SAHPRA recognized jurisdictions . Attach a copy of certification
4.3.1 | Please provide unique ref number of the ceriticate:
4.4 | Hazardous Substance Act: If this type of equipment / apparatus appears on the schedule of Hazardous substances issued by the Directorate: Health Technology of the Department of Health, a license in terms of the Act on Hazardous Substances (Act 15/1973) must be submitted with the bid / quotation documents. The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
4.5 | All electrical/electronic medical equipment must be licensed by Radiation Control, where applicable a copy of the license must be submitted The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
4.5.1 | Radiation Control License number:
4.6 | The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade (rubber to with 2 screws) type and it must be a minimum length of three (3) meters. NB the mains cable of the unit tendered for must be SABS colour coded.
4.7 | The equipment quoted for must be protected against electromagnetic interference IEC 60601-1-2
4.8 | The unit must comply with electrical safety and performance standard such as IEC 60601-2-21, attached certification
4.8.1 | Please provide unique ref number of the IEC ceriticate:
5 | COMPREHENSIVE MAINTENANCE
5.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) should be included in the unit price as per the SCC document
5.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
5.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
5.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
5.3 | The bidder to provide the OEM recommended service interval of the unit:
5.4 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
6 | General
6.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
6.2 | The unit on offer must be the latest technology.
6.2.1 | Initial Manufacturing year of Model:
RT2-03-010 | Phototherapy Blanket LED | Complies Yes/No
RT2-03-010a | Effective treatment area: Large pad of 450-550cm2

## Cooling Device, Neonatal
RT2-03-011 | Neonatal Cooling Device | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
1 | DESCRIPTION
1.1 | A medical device designed to cool a newborn's body to prevent brain damage.
Product Offered:
Make:
Model:
Initial Manufacturing year of Model:
Radiation Control License number:
2 | TECHNICAL SPECIFICATION
2.1 | The unit is specifically required for the cooling of babies suffering from Hypoxic Ischemic Encephalopathy (HIE) after birth.
2.2 | The unit must be equipped with a servo control to monitor and regulate the set temperature
2.3 | The unit being offered must comprise the latest technology. It must be user friendly and incorporate reliable safety features
2.4 | The unit must be equipped with an alarm that is set off if the temperature varies more than 0.5 degrees from set temperature.
2.5 | The display of the unit must clearly show the current temperature of the patient
2.6 | It must be possible for the user to set the following targets:
2.6.1 | Target Temperature
2.6.2 | Duration to reach temperature
2.6.3 | Rate of re-warming/cooling
2.7 | It must be possible to make changes to settings during operations, and all changes must be recorded
2.8 | The temperature of the patient must be monitored by a rectal probe (at least two probes to be supplied with the unit)
2.9 | The baby must be cooled by passing cooling fluids around the baby through a blanket or wrap
2,10 | Bidders to state the sizes of the blanket or wrap, and how many is included in the offer
2.11 | The unit must only make use of a disposable blankets or wraps.
2.12 | The device must include alarms for at least the following:
2.12.1 | No Power
2.12.2 | No flow
2.12.3 | Low fluid
2.12.4 | Temperature out of range
2.12.5 | System failure
2.13 | The unit must have the following modes of setting for different operations
2.13.1 | Programmable servo control
2.13.1.1 | User can set duration of cooling, rate of rewarming and target temperature. The unit then performs the complete cycle automatically
2.13.2 | Servo controlled rectal temperature
2.13.2.1 | The unit must maintain the rectal temperature that is set by the clinician
2.13.3 | Constant blanket temperature
2.13.3.1 | The unit maintains a constant set temperature for the blanket or wrap
3 | STANDARD ACCESSORIES
3.1 | Dedicated trolley for the unit.
3.2 | Reusable Rectal probe
3.3 | Rectal Probe Covers
3.4 | Reusable Skin probe
3.5 | 2metre connecting hose
3.6 | Disposable Blanket of approximately 30 X 45 cm
3.7 | Disposable Blanket of approximately 50 X 90 cm
3.8 | Wrap around blanket
3.9 | 5 litre coolant fluid
3.10 | Refill/drainage hose
5 | Conformity Compliance (Please attach certificate)
5.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
5.2 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
5.2.1 | Please provide unique ref number of the ISO 13485 ceriticate:
5.3 | Model on offer must comply with EU Medical Device Regulation (EU) 2017/745 or equvalent from SAHPRA recognized jurisdictions . Attach a copy of certification
5.3.1 | Please provide unique ref number of the ceriticate:
5.4 | Hazardous Substance Act: If this type of equipment / apparatus appears on the schedule of Hazardous substances issued by the Directorate: Health Technology of the Department of Health, a license in terms of the Act on Hazardous Substances (Act 15/1973) must be submitted with the bid / quotation documents. The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.5 | All electrical/electronic medical equipment must be licensed by Radiation Control, where applicable a copy of the license must be submitted The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.5.1 | Radiation Control License number:
5.6 | The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade (rubber to with 2 screws) type and it must be a minimum length of three (3) meters. NB the mains cable of the unit tendered for must be SABS colour coded.
5.7 | The equipment quoted for must be protected against electromagnetic interference IEC 60601-1-2
5.8 | The unit must comply with electrical safety and performance standard such as IEC 60601-2-21, attached certification
5.8.1 | Please provide unique ref number of the IEC ceriticate:
6 | COMPREHENSIVE MAINTENANCE
6.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) should be included in the unit price as per the SCC document
6.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries, Flow sensors and Oxygen sensor), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
6.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
6.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries, Flow sensors and Oxygen sensor), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
6.3 | The bidder to provide the OEM recommended service interval of the unit:
6.4 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
7 | General
7.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
7.2 | The unit on offer must be the latest technology.
7.2.1 | Initial Manufacturing year of Model:
RT2-03-011 | Patient Cooling Device | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point
RT2-03-011a | Dedicated trolley for the unit.
RT2-03-011b | Reusable Rectal probe
RT2-03-011c | Rectal Probe Covers
RT2-03-011d | Reusable Skin probe
RT2-03-011e | 2metre connecting hose
RT2-03-011f | Disposable Blanket of approximately 30 X 45 cm
RT2-03-011g | Disposable Blanket of approximately 50 X 90 cm
RT2-03-011h | Wrap around blanket
RT2-03-011i | 5 litre coolant fluid
RT2-03-011j | Refill/drainage hose

## Ultrasound, Neonatal
RT2-03-012 | Ultrasound Unit, Neonatal | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
1 | Description
1.1 | A medical device that uses high frequency sound waves to create real images or videos of the internal organs or other soft tissues.
2 | Product Offered:
2.1 | Make:
2.2 | Model:
2.3 | Initial Manufacturing year of Model:
2.4 | Radiation Control License number:
3 | Technical Specification
3.1 | The unit must be a fully digital unit with a memory beam former and processor to generate sharp defined ultrasound images, reduce artefacts and improve image quality.
3.2 | It must utilise the latest software to perform investigations.
3.3 | High image quality will be one of the main criteria that will be used to evaluate the offers.
3.4 | Although offered units might comply with all the specifications, image quality must be a critical evaluation criterion, specifically with increased BMI cases.
3.5 | The unit must be of the latest technology.
3.6 | It must be a state of the art fully digital phase array ultrasound unit.
3.7 | IMAGING MODALITIES
3.7.1 | The unit must be latest technology; please state year in which this unit was first manufactured.
3.7.2 | The unit must offer excellent high image quality in B mode and fully digital colour flow Doppler. Colour Doppler should have no delay after activation.
3.7.3 | The system Dynamic Range must not be less than 172 dB. The Dynamic range should be able to be changed in 5dB increments.
3.7.4 | The unit must have direct storage on hard drive of not less than 80 GB. State all other storage facility that is included in the offered price. (Such as CD, DVD, MOD). Recorded CD roms must be compatible with any standard PC, without requiring extra software. USB download of images from the system must be possible
3.7.5 | A digital black and white thermal printer must be included in the offer for the machine.
3.7.6 | The system must offer a Digital Continuous Beam Former to generate sharp defined ultrasound beams, to reduce artefacts and to provide more consistent overall image quality.
3.7.7 | The unit must offer multi-frequency on all probes to provide sharp reproducible image quality in both superficial and deep anatomy.
3.7.8 | The unit must be equipped with a non-interlaced colour monitor with a screen size of not less than 15 inch, measured from corner to corner. The monitor must be of the flat screen TFT LCD type.
3.7.9 | The unit should have at least three active probe connectors with electronic switching.
3.7.10 | The unit must be able to offer scanning methods for Electronic Convex and Linear technology.
3.7.11 | The unit must offer depth gain in not less than 8 steps. Unit must be programmable to start up automatically at a depth of operator’s choice and programmable to start up at different depths for specific examinations.
3.7.12 | Image processing must be done via a rotary dial or similar technology with at least 5 combinations for optimising the best image quality for each patient’s physique (user programmable). Image processing should visibly improve image and not just change it.
3.7.13 | Pan-Zoom enlargement must be in real and frozen mode up to 10X zoom.
3.7.14 | A colour pallet must be available for the colourization of the B mode, M mode and PW Doppler trace.
3.7.15 | Independent M mode and PW Doppler gain adjustment must be available.
3.7.16 | Sweep speed must be adjustable in at least 4 steps.Unit must have selectable sweep rates of 1,2,3,4,8 s/screen
3.7.17 | The unit must offer Tissue Harmonics as standard in the system.
3.7.18 | Grey scale level must be at least 256 shades.
3.7.19 | The unit must have not less than 1024 processing channels.
3.7.20 | The unit must be housed in a robust mobile cart. The cart must house all elements of the unit, and must have enough dedicated storage space for the probes and all disposables needed on the machine. The unit to be used for mobile ultrasound in different areas. a) Width of system in centimetres b) Depth of system in centimetres.
3.7.21 | The frame rate must not be less than 394 frames per second. Indicate the frame rate that the system offers
3.7.22 | Image depth must be not less than 30cm.
3.7.23 | The unit should offer quick scan to automatically adjusts STC/TGC and gain through one-touch operation. This must visibly improve the image
3.7.24 | Post processing must be available on frozen images
3.7.25 | Unit must offer user programmable annotation for quick labelling. User must be able to program approximately 40 words of their own choice per selected examination type for annotation purposes
3.8 | DOPPLER
3.8.1 | PW Doppler must be offered on all imaging probes.
3.8.2 | Maximum detectable velocity for PW must be stated.
3.8.3 | Pulse Repetition Frequency must vary between 0,7 – 19,8 kHz pulsed wave and 2,1 – 40kHz continuous wave
3.8.4 | Independent angle steering up to ±20° for colour and Doppler must be offered.
3.8.5 | Independent gain control for PW must be available.
3.8.6 | Colour pallet offering 7 types independent of B mode and Pulse Doppler pallet to be available.
3.8.7 | Sampling with Pulse Doppler must vary between 1 and 20 mm.
3.8.8 | Gate size must vary at least between 1 and 8mm.
3.8.9 | Base line shift must be possible in at least 8 steps.
3.8.10 | Spectral invert must be possible.
3.8.11 | At least 6 wall filter selections must be available. Bidder to state minimum.
3.8.12 | Accurate autotrace function for both arterial and venous waveforms must be available.
3.9 | COLOUR DOPPLER
3.9.1 | A minimum of 8 colour maps must be available for colour display.
3.9.2 | Colour maps must be available to be selected in real time and freeze mode.
3.9.3 | Independent colour gain adjustment must be from 1 to 16 dB in 0,5 dB steps. Indicate range
3.9.4 | Maximum colour frame rate must not be less than 285 frames per second.
3.9.5 | Colour flow must operate to the full depth of 30 cm.
3.9.6 | Colour maps and processing selections must be able to be selected by the user and stored in the presets.
3.9.7 | Quad signal processing and parallel signal processing must be available. Bidders to supply details.
3.9.8 | Power Doppler on all transducers must be available.
3.9.9 | The unit offered must be able to provide a real time triplex mode.
3.9.10 | Colour mapping must be instantaneous – no time delay /lag.
3.10 | IMAGE MEMORY
3.10.1 | Image memory to operate in both dual and quad screen format.
3.10.2 | ECG synchronised recording must be possible.
3.10.3 | It must be possible to record B loop up to 500 frames.
3.10.4 | Cine loop and measurements must be active simultaneously.
3.10.5 | M/D review of a minimum of 15 seconds must be possible.
3.10.6 | B Mode Imaging on linear transducer to be electronically steerable.
3.11 | TRANSDUCERS
3.11.1 | The following probes must be included for the machine:
3.11.1.1 | Phased Array probe - 2.0 to 7.5 Mhz)
3.11.1.2 | Linear probe - 5 to 10 Mhz
3.11.1.3 | Neonatal Cardiac probe at least 3 to 9 Mhz
3.11.2 | All imaging transducers must be multi-frequency.
3.11.3 | Tissue Harmonics must be available on all offered transducers.
3.11.4 | The selected frequency of the imaging transducer must be displayed on the screen.
3.11.5 | All transducers must be digitally controlled so continuous beam forming is possible. State number of focal zones.
3.11.6 | A minimum of 2 active probes must be available. It must be possible to electronically switch between these 2 probes.
3.12 | MEASUREMENTS
3.12.1 | Patient data MUST offer the option to enter previous ultrasound date and measurements.
3.12.2 | Gestational age according to Last Menstrual Period(LMP) must be displayed on working screen, not just on patient information page.
3.12.3 | On-screen measurements must be possible from real time. 16 callipers (8 channel)each individually numbered for measuring distance/volume/angle and area must be offered
3.12.4 | The unit must offer a storage facility for up to 300 patient measurement data. It must be possible to retrieve the data for follow-ups.
3.12.5 | The unit must offer 50th percentile growth graphs of the Doppler measurements RI/RP. It must be possible to retrieve the data for follow-ups.
3.12.6 | The unit must offer user programmable annotation for quick labelling.
3.12.7 | State if the unit comply to the following: 1. Panoramic view. 2. Quad signal processing
3.12.8 | All software upgrades must be offered optionally, with full description what improvement the software upgrade will bring to the offered unit.
3.12.9 | The unit must be supplied with an uninterrupted power supply (UPS), to be able to keep the unit running for at least two(2) hours in case of power failure.
3.13 | DICOM COMPLIANCE
3.13.1 | The offered unit must be fully DICOM compliant.
3.13.2 | It must be possible for the unit to be connected to PACS.
4 | Standard Accessories
4.1 | Phased Array probe 2.0 to 7.5 Mhz)
4.2 | Linear probe 5 to 10 Mhz
4.3 | Neonatal Cardiac probe 3 to 9 Mhz
4.4 | Uninterrupted Power Supply (UPS)
4.5 | Gel
4.6 | Printer
4.7 | Printing paper
5 | Conformity Compliance (Please attach certificate)
5.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
5.2 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
5.2.1 | Please provide unique ref number of the ISO 13485 ceriticate:
5.3 | Model on offer must comply with EU Medical Device Regulation (EU) 2017/745 or equvalent from SAHPRA recognized jurisdictions . Attach a copy of certification
5.3.1 | Please provide unique ref number of the ceriticate:
5.4 | Hazardous Substance Act: If this type of equipment / apparatus appears on the schedule of Hazardous substances issued by the Directorate: Health Technology of the Department of Health, a license in terms of the Act on Hazardous Substances (Act 15/1973) must be submitted with the bid / quotation documents. The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.5 | All electrical/electronic medical equipment must be licensed by Radiation Control, where applicable a copy of the license must be submitted The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.5.1 | Radiation Control License number:
5.6 | The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade (rubber to with 2 screws) type and it must be a minimum length of three (3) meters. NB the mains cable of the unit tendered for must be SABS colour coded.
5.7 | The equipment quoted for must be protected against electromagnetic interference IEC 60601-1-2
5.8 | The unit must comply with electrical safety and performance standard such as IEC 60601-2-21, attached certification
5.8.1 | Please provide unique ref number of the IEC ceriticate:
6 | COMPREHENSIVE MAINTENANCE
6.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) should be included in the unit price as per the SCC document
6.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
6.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
6.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
6.3 | The bidder to provide the OEM recommended service interval of the unit:
6.4 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
7 | General
7.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
7.2 | The unit on offer must be the latest technology.
7.2.1 | Initial Manufacturing year of Model:
RT2-03-012 | Ultrasound Unit, Neonatal | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point
RT2-03-012a | Phased Array probe - 2.0 to 7.5 Mhz)
RT2-03-012b | Linear probe - 5 to 10 Mhz
RT2-03-012c | Neonatal Cardiac probe – 3 to 9 Mhz
RT2-03-012d | Uninterrupted Power Supply (UPS)
RT2-03-012e | Gel
RT2-03-012f | Printer
RT2-03-012g | Printing paper

## Vein Viewer
RT2-03-013 | Vein Viewer/Finder | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
1 | DESCRIPTION
1.1 | A medical device that uses near-infrared light, ultrasound or other technologies to visualize and locate the veins beneath the skin surface.
2 | Product Offered:
2.1 | Make:
2.2 | Model:
2.3 | Initial Manufacturing year of Model:
3 | SPECIFICATIONS
3.1 | The unit offered must be able to be used on the following patients:
3.1.1 | Neonates
3.1.2 | Paediatrics
3.1.3 | Adults
3.1.4 | Obese or Edema
3.1.5 | Patients with anemia, hypotension or excessive loss of blood
3.2 | The unit offered must be portable, non-contact, lightweight, hand-held and ability to place on stand for hands free functionality
3.3 | The unit must use safe infrared light source or any other technology without any additional eye protection for both the user and patient
3.4 | The effective working distance of the unit must not be more than 31cm
3.5 | The device must perform accurately on different skin pigmentation/tone.
3.6 | The unit must display real-time vein even with patient movement
3.7 | The colour image on the display of the unit must be clear and accurate
3.8 | The unit must function with both AC and rechargeable Li-ion battery.
3.9 | The battery must be able to function for more than 2 hours of continuous work
3.10 | The unit must provide vein map projection on different skin tone
3.11 | All shortlisted units will undergo Clinical trial for a BEC recommended duration of which the Clinician report will form part of evaluation
4 | Standard Accessories and Conusmables
4.1 | Battery
4.2 | Carry case
5 | Conformity Compliance (Please attach certificate)
5.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
5.2 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
5.2.1 | Please provide unique ref number of the ISO 13485 ceriticate:
5.3 | Model on offer must comply with EU Medical Device Regulation (EU) 2017/745 or equvalent from SAHPRA recognized jurisdictions . Attach a copy of certification
5.3.1 | Please provide unique ref number of the ceriticate:
5.4 | Hazardous Substance Act: If this type of equipment / apparatus appears on the schedule of Hazardous substances issued by the Directorate: Health Technology of the Department of Health, a license in terms of the Act on Hazardous Substances (Act 15/1973) must be submitted with the bid / quotation documents. The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.5 | All electrical/electronic medical equipment must be licensed by Radiation Control, where applicable a copy of the license must be submitted The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.5.1 | Radiation Control License number:
5.6 | The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade (rubber to with 2 screws) type and it must be a minimum length of three (3) meters. NB the mains cable of the unit tendered for must be SABS colour coded.
5.7 | The equipment quoted for must be protected against electromagnetic interference IEC 60601-1-2
5.8 | The unit must comply with electrical safety and performance standard such as IEC 60601-2-21, attached certification
5.8.1 | Please provide unique ref number of the IEC ceriticate:
6 | COMPREHENSIVE MAINTENANCE
6.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) should be included in the unit price as per the SCC document
6.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
7 | General
7.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
7.2 | The unit on offer must be the latest technology.
RT2-03-013 | Vein Viewer/Finder -Accessories and Consumables | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point
RT2-03-013a | Battery
RT2-03-013b | Carry case.

## Category-4 O&G Equipment
OBSTETRIC & GYNAECOLOGY EQUIPMENT
# | Item number | Item Description
1 | RT2-04-01 | Cardiotocograph, Single
2 | RT2-04-02 | Cardiotocograph, Single + MHR
3 | RT2-04-03 | Cardiotocograph, Twins
4 | RT2-04-04 | Cardiotocograph, Twins + MHR
5 | RT2-04-05 | Fetal Doppler
6 | RT2-04-06 | Video Colposcope
7 | RT2-04-07 | Vacuum Extractor
8 | RT2-04-08 | Ultrasound Machine, O&G

## Cardiotocograph, single
RT2-04-001 | Cardiotocograph, Singleton | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
1 | Description
A medical device that monitors and records the fetal heart rate and uterine contractions of a singleton pregnancy (one fetus) during pregnancy, labor, and delivery
2 | Product Offered:
2.1 | Make:
2.2 | Model:
2.3 | Initial Manufacturing year of Model:
2.4 | Radiation Control License number:
3 | Technical Specification
3.1 | Required for external monitoring of fetal heart rate (FHR), and external monitoring of uterine activity of pregnant patients
3.2 | The unit must be able to monitor one(1) fetal heart rate.
3.3 | It must be able to perform automatic self-check when switched on
3.4 | There must be a visible Fetal Heart Beat indicator which will flash with each detected heart beat
3.5 | The unit must provide an audio signal, with variable volume, for every heart beat detected
3.6 | The fetal heart rate must be measured by means of wide beam pulsed Doppler ultrasound, with auto correlation processing.
3.7 | Must be able to detect and display FHR
3.8 | The oscillator frequency of not < 1,2 MHz.
3.9 | Ultrasound intensity must be less than 10 mW/cm2
3.10 | The transducer must be manufactured of high-impact plastic that is waterproof, robust and be able to pass a 1m drop test.
3.11 | The belts and buckles supplied for positioning of the transducer must allow easy angling for optimum positioning.
3.12 | Adjustable baseline position must be possible .
3.13 | A test button must be provided that produces a series of pulses for calibration which must be recorded and displayed.
3.14 | The unit must have a LCD display of at least 16 cm (6 Inch).
3.15 | Displays must be clearly visible under all lighting conditions.
3.17 | The following parameters must be displayed:
3.17.1 | Fetal heart rate approximately 50 to 200 beats per minute
3.17.2 | Uterine activity: approximately 0 to 100 relative units.
3.18 | Fetal heart rate must update once per second when FHR is greater than 60 and once each beat when FHR is less than 60.
3.19 | The monitor must have the following controls:
3.19.1 | Volume Up/Down
3.19.2 | TOCO/IUP Zero
3.19.3 | Printer On/Off
3.19.4 | Printer fast forward
3.20 | The unit must have audible and visual alarms for low/high FHR
3.21 | Electronic storage capacity of a minimum of 300 patient records or at least 24 hours trends with First In First Out(FIFO) file management method.
3.22 | RECORDER
3.22.1 | Integrated dual channel thermal recorder must be available
3.22.2 | Heart rate recording must be possible.
3.22.3 | All dot print test must be available
3.22.4 | The strip chart paper must be easy to load
3.23 | Fetal heart rate:
3.23.1 | Range: approximately 50 to 210 beats per minute.
3.23.2 | Vertical chart: at least 7 cm.
3.23.3 | Vertical scale sensitivity: approximately 20 bpm per cm.
3.23.4 | Resolution: 1 bpm
3.24 | Uterine activity:
3.24.1 | Range: 0 – 100 units
3.24.2 | Vertical chart: at least 4 cm.
3.24.3 | Vertical scale sensitivity: approximately 25 relative units per cm
3.24.4 | Resolution: 1 relative unit
3.25 | Paper chart speed must be selectable for 1, 2 and 3 cm per minute.
3.26 | The time and date must be printed by the recorder at least every 10 minutes.
3.27 | Recording time must be at least 24 hours at 1 cm per minute.
3.28 | The unit must have a facility for remote event marking
3.29 | TRANSDUCERS
3.29.1 | Ultrasound and toco transducers must be supplied
3.29.2 | All transducers must be of very robust design. Bidders must guarantee shock resistance on the transducers e.g. will withstand a drop from one metre high onto a hard surface
3.29.3 | Bidders must state whether the ultrasound transducer is repairable and which parts are replaceable or whether it is a sealed unit that must be replaced when faulty.
3.29.4 | The transducer cables must be a minimum of 2 m long
3.29.5 | Connecting plug must be keyed and colour coded to avoid it being plugged or forced into the wrong receptacle on the unit.
3.29.6 | Transducer holders must be provided
3.30 | General
3.30.1 | All metal parts must be earthed.
3.30.2 | Must be fitted with rechargeable batteries
3.30.3 | The unit must be fitted with an internal battery charger
3.30.4 | Indicator for battery status must be available
3.30.5 | A trolley or wall mounting bracket must be supplied as per end user requirements.
3.30.6 | The trolley must have four castor wheels, two of them with brakes
3.30.7 | The trolley must have a basket or holder for accessories
3.30.8 | A starter-pack of consumables must be supplied with the unit.
4 | Standard Accessories and Consumables
4.1 | US transducer
4.2 | Toco transducer
4.3 | Thermal printing paper (Pk of 5)
4.4 | Belts x2 (For each transducer)
4.5 | Marker
4.6 | Trolley or wall mounting bracket
5 | Conformity Compliance (Please attach certificate)
5.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
5.2 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
5.2.1 | Please provide unique ref number of the ISO 13485 ceriticate:
5.3 | Model on offer must comply with EU Medical Device Regulation (EU) 2017/745 or equvalent from SAHPRA recognized jurisdictions . Attach a copy of certification
5.3.1 | Please provide unique ref number of the ceriticate:
5.4 | Hazardous Substance Act: If this type of equipment / apparatus appears on the schedule of Hazardous substances issued by the Directorate: Health Technology of the Department of Health, a license in terms of the Act on Hazardous Substances (Act 15/1973) must be submitted with the bid / quotation documents. The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.5 | All electrical/electronic medical equipment must be licensed by Radiation Control, where applicable a copy of the license must be submitted The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.5.1 | Radiation Control License number:
5.6 | The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade (rubber to with 2 screws) type and it must be a minimum length of three (3) meters. NB the mains cable of the unit tendered for must be SABS colour coded.
5.7 | The equipment quoted for must be protected against electromagnetic interference IEC 60601-1-2
5.8 | The unit must comply with electrical safety and performance standard such as IEC 60601-2-21, attached certification
5.8.1 | Please provide unique ref number of the IEC ceriticate:
6 | COMPREHENSIVE MAINTENANCE
6.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) should be included in the unit price as per the SCC document
6.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
6.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
6.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
6.3 | The bidder to provide the OEM recommended service interval of the unit:
6.4 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
7 | General
7.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
7.2 | The unit on offer must be the latest technology.
7.2.1 | Initial Manufacturing year of Model:
RT2-04-001 | Cardiotocograph Replacement Accessories and Consumables | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
RT2-04-001a | US transducer
RT2-04-001b | Toco transducer
RT2-04-001c | Thermal printing paper (Pk of 5)
RT2-04-001d | Belt
RT2-04-001e | Marker
RT2-04-001f | Trolley
RT2-04-001g | Wall mounting bracket

## Cardiotocograph with MHR
RT2-04-002 | Cardiotocograph, singleton with Maternal Heart Rate(MHR) monitoring | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
1 | Description
1.1 | A medical device that monitors and records the fetal and maternal heart rate and uterine contractions of a singleton pregnancy (one fetus) during pregnancy, labor, and delivery
2 | Product Offered:
2.1 | Make:
2.2 | Model:
2.3 | Initial Manufacturing year of Model:
2.4 | Radiation Control License number:
3 | Technical Specification
3.1 | Required for external monitoring of fetal heart rate (FHR), and external monitoring of uterine activity of pregnant patients and must include monitoring and display of maternal heart rate
3.2 | The unit must be able to monitor one(1) fetal heart rate.
3.3 | Must perform automatic self-check when switched on
3.4 | There must a visible Fetal Heart Beat indicator which will flash with each detected heart beat
3.5 | The unit must provide an audio signal, with variable volume, for every heart beat detected
3.6 | The fetal heart rate must be measured by means of wide beam pulsed Doppler ultrasound, with auto correlation processing.
3.8 | The oscillator frequency < 1,2 MHz.
3.9 | Ultrasound intensity must be less than 10 mW/cm2
3.10 | The transducer must be manufactured of high-impact plastic that is waterproof, robust and be able to pass a 1m drop test.
3.11 | The belts and buckles supplied for positioning of the transducer must allow easy angling for optimum positioning.
3.12 | Adjustable baseline position must be possible .
3.13 | A test button must be provided that produces a series of pulses for calibration which must be recorded and displayed.
3.14 | The unit must have a LCD display of at least 16 cm (6 Inch).
3.15 | Displays must be clearly visible under all lighting conditions.
3.16 | Must allow Maternal Heart Rate (MHR) and Fetal Heart Rate(FHR) coincidence detection.
3.17 | The following parameters must be displayed:
3.18 | Fetal heart rate approximately 50 to 200 beats per minute
3.19 | Uterine activity: approximately 0 to 100 relative units.
3.20 | Heart Rate monitoring
3.20.1 | Fetal heart rate must update once per second when FHR is greater than 60 and once each beat when FHR is less than 60.
3.20.2 | Maternal heart rate must be obtained, and displayed, reliably using preferably ECG or plethysmography.
3.21 | The monitor must have the following controls:
3.21.1 | Volume Up/Down
3.21.2 | TOCO/IUP Zero
3.21.3 | Printer On/Off
3.21.4 | Printer fast forward
3.22 | Must have audible and visual alarms for low/high FHR
3.23 | Electronic storage capacity of a minimum of 300 patient records or at least 24 hours trends with First In First Out(FIFO) file management method.
3.24 | RECORDER
3.24.1 | Integrated dual channel thermal recorder
3.24.2 | Heart rate recording must be available.
3.24.3 | All dot print test must be available
3.24.4 | The strip chart paper must be easy to load
3.25 | Fetal heart rate:
3.25.1 | Range: approximately 50 to 210 beats per minute.
3.25.2 | Vertical chart: at least 7 cm.
3.25.3 | Vertical scale sensitivity: approximately 20 bpm per cm.
3.25.4 | Resolution: 1 bpm
3.26 | Uterine activity:
3.26.1 | Range: 0 – 100 units
3.26.2 | Vertical chart: at least 4 cm.
3.26.3 | Vertical scale sensitivity: approximately 25 relative units per cm
3.26.4 | Resolution: 1 relative unit
3.26.5 | Paper chart speed must be selectable for 1, 2 and 3 cm per minute.
3.26.6 | The time and date must be printed by the recorder at least every 10 minutes.
3.26.7 | Recording time must be at least 24 hours at 1 cm per minute.
3.26.8 | The unit must have a facility for remote event marking
3.27 | TRANSDUCERS
3.27.1 | Ultrasound and toco transducers must be supplied
3.27.2 | All transducers must be of very robust design. Bidders must guarantee shock resistance on the transducers e.g. will withstand a drop from one metre high onto a hard surface
3.27.3 | Bidders must state whether the ultrasound transducer is repairable and which parts are replaceable or whether it is a sealed unit that must be replaced when faulty.
3.27.4 | The transducer cables must be a minimum of 2 m long
3.27.5 | Connecting plug must be keyed and colour coded to avoid it being plugged or forced into the wrong receptacle on the unit.
3.27.6 | Transducer holders must be provided
3.28 | General
3.28.1 | All metal parts must be earthed.
3.28.2 | Must be fitted with rechargeable batteries
3.28.3 | The unit must be fitted with an internal battery charger
3.28.4 | Indicator for battery status
3.28.5 | A trolley or wall mounting bracket must be supplied as per end user requirements.
3.28.6 | The trolley must have four castor wheels, two of them with brakes
3.28.7 | The trolley must have a basket/holder for accessories
3.28.8 | A starter-pack of consumables must be supplied with the unit.
4 | Standard Accessories and Consumables
4.1 | US transducer
4.2 | Toco transducer
4.3 | Thermal printing paper (Pk of 5)
4.4 | Belts x2 (For each transducer)
4.5 | Marker
4.6 | Trolley or wall mounting bracket
5 | Conformity Compliance (Please attach certificate)
5.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
5.2 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
5.2.1 | Please provide unique ref number of the ISO 13485 ceriticate:
5.3 | Model on offer must comply with EU Medical Device Regulation (EU) 2017/745 or equvalent from SAHPRA recognized jurisdictions . Attach a copy of certification
5.3.1 | Please provide unique ref number of the ceriticate:
5.4 | Hazardous Substance Act: If this type of equipment / apparatus appears on the schedule of Hazardous substances issued by the Directorate: Health Technology of the Department of Health, a license in terms of the Act on Hazardous Substances (Act 15/1973) must be submitted with the bid / quotation documents. The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.5 | All electrical/electronic medical equipment must be licensed by Radiation Control, where applicable a copy of the license must be submitted The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.5.1 | Radiation Control License number:
5.6 | The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade (rubber to with 2 screws) type and it must be a minimum length of three (3) meters. NB the mains cable of the unit tendered for must be SABS colour coded.
5.7 | The equipment quoted for must be protected against electromagnetic interference IEC 60601-1-2
5.8 | The unit must comply with electrical safety and performance standard such as IEC 60601-2-21, attached certification
5.8.1 | Please provide unique ref number of the IEC ceriticate:
6 | COMPREHENSIVE MAINTENANCE
6.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) should be included in the unit price as per the SCC document
6.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
6.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
6.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
6.3 | The bidder to provide the OEM recommended service interval of the unit:
6.4 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
7 | General
7.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years If yes, please provide further details of the recall and how was it addressed
7.2 | The unit on offer must be the latest technology.
7.2.1 | Initial Manufacturing year of Model:
RT2-04-002 | Cardiotocograph Replacement Accessories and Consumables | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
RT2-04-002a | US transducer
RT2-04-002b | Toco transducer
RT2-04-002c | Thermal printing paper (Pk of 5)
RT2-04-002d | Belt
RT2-04-002e | Marker
RT2-04-002f | Trolley
RT2-04-002g | Wall mounting bracket

## Cardiotocograph, Twins
RT2-04-003 | Cardiotocograph, twins | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
1 | Description
A medical device that monitors and records the fetal heart rate and uterine contractions of a twins pregnancy (two fetus) during pregnancy, labor, and delivery
2 | Product Offered:
2.1 | Make:
2.2 | Model:
2.3 | Initial Manufacturing year of Model:
2.4 | Radiation Control License number:
3 | Technical Specification
3.1 | Required for external monitoring of fetal heart rate (FHR), and external monitoring of uterine activity of pregnant patients
3.2 | It must be possible to monitor Fetal heart rate of two fetus.
3.3 | Must perform automatic self-check when switched on
3.4 | There must be a visible Fetal Heart Beat indicator which will flash with each detected heart beat
3.5 | The unit must provide an audio signal, with variable volume, for every heart beat detected
3.6 | The fetal heart rate must be measured by means of wide beam pulsed Doppler ultrasound, with auto correlation processing.
3.7 | Must be able to detect and display multiple FHR while discriminating between different foetuses
3.8 | The oscillator frequency < 1,2 MHz.
3.9 | Ultrasound intensity must be less than 10 mW/cm2
3.10 | The transducer must be manufactured of high-impact plastic that is waterproof, robust and be able to pass a 1m drop test.
3.11 | The belts and buckles supplied for positioning of the transducer must allow easy angling for optimum positioning.
3.12 | Adjustable baseline position must be possible .
3.13 | A test button must be provided that produces a series of pulses for calibration which must be recorded and displayed.
3.14 | The unit must have a LCD display of at least 16 cm (6 Inch).
3.15 | Displays must be clearly visible under all lighting conditions.
3.16 | Must allow Maternal Heart Rate (MHR) and Fetal Heart Rate(FHR) coincidence detection.
3.17 | The following parameters must be displayed:
3.17.1 | Fetal heart rate approximately 50 to 200 beats per minute
3.17.2 | Uterine activity: approximately 0 to 100 relative units.
3.18 | Fetal heart rate must update once per second when FHR is greater than 60 and once each beat when FHR is less than 60.
3.19 | The monitor must have the following controls:
3.19.1 | Volume Up/Down
3.19.2 | TOCO/IUP Zero
3.19.3 | Printer On/Off
3.19.4 | Printer fast forward
3.20 | Must have audible and visual alarms for low/high FHR
3.21 | Fetal heart rate must update once per second when FHR is greater than 60 and once each beat when FHR is less than 60.
3.22 | RECORDER
2.22.1 | Integrated dual channel thermal recorder
2.22.2 | Dual heart rate recording for twins must be available.
2.22.3 | All dot print test must be available
2.22.4 | The strip chart paper must be easy to load
3.23 | Fetal heart rate:
3.23.1 | Range: approximately 50 to 210 beats per minute.
3.23.2 | Vertical chart: at least 7 cm.
3.23.3 | Vertical scale sensitivity: approximately 20 bpm per cm.
3.23.4 | Resolution: 1 bpm
3.24 | Uterine activity:
3.24.1 | Range: 0 – 100 units
3.24.2 | Vertical chart: at least 4 cm.
3.24.3 | Vertical scale sensitivity: approximately 25 relative units per cm
3.24.4 | Resolution: 1 relative unit
3.25 | Paper chart speed must be selectable for 1, 2 and 3 cm per minute.
3.26 | The time and date must be printed by the recorder at least every 10 minutes.
3.27 | Recording time must be at least 24 hours at 1 cm per minute.
3.28 | The unit must have a facility for remote event marking
3.29 | TRANSDUCERS
3.29.1 | 2x Ultrasound and toco transducers must be supplied
3.29.2 | All transducers must be of very robust design. Bidders must guarantee shock resistance on the transducers e.g. will withstand a drop from one metre high onto a hard surface
3.29.3 | Bidders must state whether the ultrasound transducer is repairable and which parts are replaceable or whether it is a sealed unit that must be replaced when faulty.
3.29.4 | The transducer cables must be a minimum of 2 m long
3.29.5 | Connecting plug must be keyed and colour coded to avoid it being plugged or forced into the wrong receptacle on the unit.
3.29.6 | Transducer holders must be provided
3.30 | Other system requirements
3.30.1 | All metal parts must be earthed.
3.30.2 | Must be fitted with rechargeable batteries
3.30.3 | The unit must be fitted with an internal battery charger
3.30.4 | Indicator for battery status must be displayed.
3.30.5 | A trolley or wall mounting bracket must be supplied as per end user requirements.
3.30.6 | The trolley must have four castor wheels, two of them with brakes
3.30.7 | The trolley must have a basket/holder for accessories
3.30.8 | A starter-pack of consumables must be supplied with the unit.
4 | Standard Accessories and Consumables
4.1 | US transducers x2
4.2 | Toco transducer
4.3 | Thermal printing paper (Pk of 5)
4.4 | Belts x3 (For each transducer)
4.5 | Marker
4.6 | Trolley or wall mounting bracket
5 | Conformity Compliance (Please attach certificate)
5.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
5.2 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
5.2.1 | Please provide unique ref number of the ISO 13485 ceriticate:
5.3 | Model on offer must comply with EU Medical Device Regulation (EU) 2017/745 or equvalent from SAHPRA recognized jurisdictions . Attach a copy of certification
5.3.1 | Please provide unique ref number of the ceriticate:
5.4 | Hazardous Substance Act: If this type of equipment / apparatus appears on the schedule of Hazardous substances issued by the Directorate: Health Technology of the Department of Health, a license in terms of the Act on Hazardous Substances (Act 15/1973) must be submitted with the bid / quotation documents. The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.5 | All electrical/electronic medical equipment must be licensed by Radiation Control, where applicable a copy of the license must be submitted The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.5.1 | Radiation Control License number:
5.6 | The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade (rubber to with 2 screws) type and it must be a minimum length of three (3) meters. NB the mains cable of the unit tendered for must be SABS colour coded.
5.7 | The equipment quoted for must be protected against electromagnetic interference IEC 60601-1-2
5.8 | The unit must comply with electrical safety and performance standard such as IEC 60601-2-21, attached certification
5.8.1 | Please provide unique ref number of the IEC ceriticate:
6 | COMPREHENSIVE MAINTENANCE
6.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) should be included in the unit price as per the SCC document
6.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
6.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
6.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
6.3 | The bidder to provide the OEM recommended service interval of the unit:
6.4 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
7 | General
7.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
7.2 | The unit on offer must be the latest technology.
7.2.1 | Initial Manufacturing year of Model:
RT2-04-003 | Cardiotocograph Replacement Accessories and Consumables | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
RT2-04-003a | US transducer
RT2-04-003b | Toco transducer
RT2-04-003c | Thermal printing paper (Pk of 5)
RT2-04-003d | Belt
RT2-04-003e | Trolley
RT2-04-003f | Wall mount bracket
RT2-04-003g | Marker

## Cardiotocographs, Twins + MHR
RT2-04-004 | Cardiotocograph Twins with Maternal Heart Rate(MHR) | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
1 | Description
1.1 | A medical device that monitors and records the fetal and maternal heart rate and uterine contractions of a twins pregnancy (two fetus) during pregnancy, labor, and delivery
2 | Product Offered:
2.1 | Make:
2.2 | Model:
2.3 | Initial Manufacturing year of Model:
2.4 | Radiation Control License number:
3 | Technical Specification
3.1 | Required for external monitoring of fetal heart rate (FHR), and external monitoring of uterine activity of pregnant patients and must include monitoring and display of maternal heart rate
3.2 | It must be possible to monitor Fetal heart rate of two fetus.
3.3 | Must perform automatic self-check when switched on
3.4 | There must a visible Fetal Heart Beat indicator which will flash with each detected heart beat
3.5 | The unit must provide an audio signal, with variable volume, for every heart beat detected
3.6 | The fetal heart rate must be measured by means of wide beam pulsed Doppler ultrasound, with auto correlation processing.
3.7 | Must be able to detect and display multiple FHR while discriminating between different foetuses
3.8 | The oscillator frequency < 1,2 MHz.
3.9 | Ultrasound intensity must be less than 10 mW/cm2
3.10 | The transducer must be manufactured of high-impact plastic that is waterproof, robust and be able to pass a 1m drop test.
3.11 | The belts and buckles supplied for positioning of the transducer must allow easy angling for optimum positioning.
3.12 | Adjustable baseline position must be possible .
3.13 | A test button must be provided that produces a series of pulses for calibration which must be recorded and displayed.
3.14 | The unit must have a LCD display of at least 16 cm (6 Inch).
3.15 | Displays must be clearly visible under all lighting conditions.
3.16 | Must allow Maternal Heart Rate (MHR) and Fetal Heart Rate(FHR) coincidence detection.
3.17 | The following parameters must be displayed:
3.18 | Fetal heart rate approximately 50 to 200 beats per minute
3.19 | Uterine activity: approximately 0 to 100 relative units.
3.20 | Heart Rate monitoring
3.20.1 | Fetal heart rate must update once per second when FHR is greater than 60 and once each beat when FHR is less than 60.
3.20.2 | Maternal heart rate must be obtained, and displayed, reliably using preferably ECG or plethysmography
3.21 | The monitor must have the following controls:
3.21.1 | Volume Up/Down
3.21.2 | TOCO/IUP Zero
3.21.3 | Printer On/Off
3.21.4 | Printer fast forward
3.22 | Must have audible and visual alarms for low/high FHR
3.23 | Electronic storage capacity of a minimum of 300 patient records or at least 24 hours trends with First In First Out(FIFO) file management method.
3.24 | RECORDER
3.24.1 | Integrated dual channel thermal recorder
3.24.2 | Heart rate recording must be available.
3.24.3 | All dot print test must be possible
3.24.4 | The strip chart paper must be easy to load
3.25 | Fetal heart rate:
3.25.1 | Range: approximately 50 to 210 beats per minute.
3.25.2 | Vertical chart: at least 7 cm.
3.25.3 | Vertical scale sensitivity: approximately 20 bpm per cm.
3.25.4 | Resolution: 1 bpm
3.26 | Uterine activity:
3.26.1 | Range: 0 – 100 units
3.26.2 | Vertical chart: at least 4 cm.
3.26.3 | Vertical scale sensitivity: approximately 25 relative units per cm
3.26.4 | Resolution: 1 relative unit
3.26.5 | Paper chart speed must be selectable for 1, 2 and 3 cm per minute.
3.26.6 | The time and date must be printed by the recorder at least every 10 minutes.
3.26.7 | Recording time must be at least 24 hours at 1 cm per minute.
3.26.8 | The unit must have a facility for remote event marking
3.27 | TRANSDUCERS
3.27.1 | 2x Ultrasound and toco transducers must be supplied
3.27.2 | All transducers must be of very robust design. Bidders must guarantee shock resistance on the transducers e.g. will withstand a drop from one metre high onto a hard surface
3.27.3 | Bidders must state whether the ultrasound transducer is repairable and which parts are replaceable or whether it is a sealed unit that must be replaced when faulty.
3.27.4 | The transducer cables must be a minimum of 2 m long
3.27.5 | Connecting plug must be keyed and colour coded to avoid it being plugged or forced into the wrong receptacle on the unit.
3.27.6 | Transducer holders must be provided
3.28 | General
3.28.1 | All metal parts must be earthed.
3.28.2 | Must be fitted with rechargeable batteries
3.28.3 | The unit must be fitted with an internal battery charger
3.28.4 | Indicator for battery status must be displayed
3.28.5 | A trolley or wall mounting bracket must be supplied as per end user requirements.
3.28.6 | The trolley must have four castor wheels, two of them with brakes
3.28.7 | The trolley must have a basket/holder for accessories
3.28.8 | A starter-pack of consumables must be supplied with the unit.
4 | Standard Accessories
4.1 | US transducers x2
4.2 | Toco transducer
4.3 | Thermal printing paper (Pk of 5)
4.4 | Belts x3 (For each transducer)
4.5 | Marker
4.6 | Trolley or wall mounting bracket
5 | Conformity Compliance (Please attach certificate)
5.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
5.2 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
5.2.1 | Please provide unique ref number of the ISO 13485 ceriticate:
5.3 | Model on offer must comply with EU Medical Device Regulation (EU) 2017/745 or equvalent from SAHPRA recognized jurisdictions . Attach a copy of certification
5.3.1 | Please provide unique ref number of the ceriticate:
5.4 | Hazardous Substance Act: If this type of equipment / apparatus appears on the schedule of Hazardous substances issued by the Directorate: Health Technology of the Department of Health, a license in terms of the Act on Hazardous Substances (Act 15/1973) must be submitted with the bid / quotation documents. The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.5 | All electrical/electronic medical equipment must be licensed by Radiation Control, where applicable a copy of the license must be submitted The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.5.1 | Radiation Control License number:
5.6 | The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade (rubber to with 2 screws) type and it must be a minimum length of three (3) meters. NB the mains cable of the unit tendered for must be SABS colour coded.
5.7 | The equipment quoted for must be protected against electromagnetic interference IEC 60601-1-2
5.8 | The unit must comply with electrical safety and performance standard such as IEC 60601-2-21, attached certification
5.8.1 | Please provide unique ref number of the IEC ceriticate:
6 | COMPREHENSIVE MAINTENANCE
6.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) should be included in the unit price as per the SCC document
6.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
6.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
6.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
6.3 | The bidder to provide the OEM recommended service interval of the unit:
6.4 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
7 | General
7.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
7.2 | The unit on offer must be the latest technology.
7.2.1 | Initial Manufacturing year of Model:
RT2-04-004 | Cardiotocograph Replacement Accessories and Consumables | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
RT2-04-004a | US transducer
RT2-04-004b | Toco transducer
RT2-04-004c | Thermal printing paper (Pk of 5)
RT2-04-004d | Belt
RT2-04-004e | Marker
RT2-04-004f | Wall mount bracket
RT2-04-004g | Trolley

## Fetal Doppler
RT2-04-05 | Doppler Fetal Heart Rate Monitor | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
1 | Description
1.1 | A non-invasive medical device using Doppler ultrasound technology to measure fetal heart rate (FHR) during pregnancy, labor and postpartum
2 | Product Offered:
2.1 | Make:
2.2 | Model:
2.3 | Initial Manufacturing year of Model:
2.4 | Radiation Control License number:
3 | Technical Specification
3.1 | The unit must be a portable Doppler for measuring fetal heart rate.
3.2 | The unit must be able to accommodate two types of obstetric probes – 2 MHz and 3 MHz
3.3 | The 2 MHz probe must be supplied with the unit
3.4 | The probes must be robust and well constructed and the crystals of the transducer must be well protected against accidental droppage.
3.5 | Fetal heart rate measuring range: approximately 60 - 220 beats per minute.
3.6 | Resolution: 1 bpm
3.7 | Accuracy +- 2 bpm
3.8 | The following must be displayed:
3.8.1 | Fetal Heart rate
3.8.2 | Heartbeat synchronous indicator
3.9 | The unit maust have a low battery indicator
3.10 | Battery charging indicator must be available
3.11 | The power output of the obstetric probes must not exceed 10 mW/cm².
3.12 | The unit must operate with rechargable batteries
3.13 | Battery and charger must be supplied.
3.14 | Battery charger must operate from a 220 V, 50 Hz power supply
3.15 | Batteries must have capacity for more than 250 examinations of 1 minute each
3.16 | The handheld unit must be lightweight, approximately 2 kg
3.17 | The unit must be equipped with a high quality speaker
3.18 | The unit must accommodate an external headset. Headset may be offered as an option
3.19 | If the headset is used for the audio output, the internal loudspeaker of the unit must be disabled.
3.20 | The volume of the audio output must be adjustable.
3.21 | The unit must have a clear LED or LCD display showing fetal heart rate
3.22 | The unit must automatically shut down after one minute of inactivity
3.23 | The unit must have a storage holder for the Doppler probes.
3.24 | A starter-pack of consumables must be supplied with the unit.
4 | Standard Accessories
4.1 | 2 MHz probe
4.2 | Carry bag
4.3 | Battery charger
5 | Conformity Compliance (Please attach certificate)
5.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
5.2 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
5.2.1 | Please provide unique ref number of the ISO 13485 ceriticate:
5.3 | Model on offer must comply with EU Medical Device Regulation (EU) 2017/745 or equvalent from SAHPRA recognized jurisdictions . Attach a copy of certification
5.3.1 | Please provide unique ref number of the ceriticate:
5.4 | Hazardous Substance Act: If this type of equipment / apparatus appears on the schedule of Hazardous substances issued by the Directorate: Health Technology of the Department of Health, a license in terms of the Act on Hazardous Substances (Act 15/1973) must be submitted with the bid / quotation documents. The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.5 | All electrical/electronic medical equipment must be licensed by Radiation Control, where applicable a copy of the license must be submitted The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.5.1 | Radiation Control License number:
5.6 | The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade (rubber to with 2 screws) type and it must be a minimum length of three (3) meters. NB the mains cable of the unit tendered for must be SABS colour coded.
5.7 | The equipment quoted for must be protected against electromagnetic interference IEC 60601-1-2
5.8 | The unit must comply with electrical safety and performance standard such as IEC 60601-2-21, attached certification
5.8.1 | Please provide unique ref number of the IEC ceriticate:
6 | COMPREHENSIVE MAINTENANCE
6.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) should be included in the unit price as per the SCC document
6.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
6.2 | The bidder to provide the OEM recommended service interval of the unit:
6.3 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
7 | General
7.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
7.2 | The unit on offer must be the latest technology.
7.2.1 | Initial Manufacturing year of Model:
RT2-04-005 | Doppler Fetal Heart Rate Monitor Replacement Accessories | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
RT2-04-005a | 2 MHz probe
RT2-04-005b | Carry bag
RT2-04-005c | A 3 MHz probe may be offered as an option
RT2-04-005d | Headset
RT2-04-005e | Battery
RT2-04-005f | Battery charger

## Video Colposcope
RT2-04-006 | Video Colposcope | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
1 | DESCRIPTION
1.1 | A medical device combining colposcopy and video technology to examine the cervix, vagina and vulva for signs of disease.
2 | Product Offered:
2.1 | Make:
2.2 | Model:
2.3 | Initial Manufacturing year of Model:
3 | TECHNICAL SPECIFICATION
3.1 | The unit offered must be supplied complete with camera on a vertical mobile stand, monitor mounted on a trolley, Image storage with a minimum of 120GB, Keyboard, mouse and Printer
3.2 | Camera
3.2.1 | Three (3) level electronic green filter LED cold lighting system to enhance vascular image display without compromise of illumination
3.2.2 | The unit must have a fast auto focusing system to ensure clear images
3.2.3 | The unit must have a remote capture control using a mouse or keyboard or foot paddle
3.2.4 | The unit must have a real time display of magnification
3.2.5 | Must have an angle adjustment function
3.2.6 | Must be equipped with Image freeze control button
3.2.7 | Precise vertical and horizontal positioning with adjustable height
3.2.8 | Adjustable zoom, focus and field view
3.3 | Monitor
3.3.1 | It must be atleast 23" Flat LCD or LED medical grade monitor with the following specifications:
3.3.1a | Resolution of at least 1920 x 1080 pixels
3.3.1b | The unit offered must have a minimum of 3 inputs which include; Video, Y/C, RGB, DVI, HDMI
3.3.1c | High contrast and wide viewing angle
3.4 | Image Management System
3.4.1 | The unit offered must record both video and still images in full HD
3.4.2 | Internal memory must be a minimum of 120GB
3.4.3 | Button to record and capture functions
3.4.4 | The system offered must be compatible with the printer offered so that the recorded items can be printed out.
3.4.5 | The system must offer provision for transfer of data to an external drive (HDD, USB or CD/DVD). The bidder to specify.
3.5 | Trolley
3.5.1 | The offered unit must come in a purpose built Colposcope System trolley with a minimum of one drawer for storage
3.5.2 | The trolley offered must have a tray for the keyboard and mouse.
3.5.3 | The trolley must have antistatic castors with at least 2 equipped with locking system/brake.
3.5.4 | The trolley must have a handle for ease access and transportation and a cable management arm.
3.5.5 | The trolley must be supplied with a minimum of 3 additional 220 V power sockets.
3.6 | IT NETWORKING AND CONNECTIVITY
3.6.1 | An RS232/RJ45 or any other latest data output port with a freely available data format for connection to hospital information system must be provided
4 | STANDARD ACCESSORIES and CONSUMABLES
4.1 | Mouse
4.2 | Printer
4.3 | Keyboard
4.4 | A4 or A6 paper (1000 sheets)
4.5 | PC software
4.6 | Dust cover must be included and supplied for the colposcope system offered
5 | Conformity Compliance (Please attach certificate)
5.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
5.2 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
5.2.1 | Please provide unique ref number of the ISO 13485 ceriticate:
5.3 | Model on offer must comply with EU Medical Device Regulation (EU) 2017/745 or equvalent from SAHPRA recognized jurisdictions . Attach a copy of certification
5.3.1 | Please provide unique ref number of the ceriticate:
5.4 | Hazardous Substance Act: If this type of equipment / apparatus appears on the schedule of Hazardous substances issued by the Directorate: Health Technology of the Department of Health, a license in terms of the Act on Hazardous Substances (Act 15/1973) must be submitted with the bid / quotation documents. The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.5 | All electrical/electronic medical equipment must be licensed by Radiation Control, where applicable a copy of the license must be submitted The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.5.1 | Radiation Control License number:
5.6 | The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade (rubber to with 2 screws) type and it must be a minimum length of three (3) meters. NB the mains cable of the unit tendered for must be SABS colour coded.
5.7 | The equipment quoted for must be protected against electromagnetic interference IEC 60601-1-2
5.8 | The unit must comply with electrical safety and performance standard such as IEC 60601-2-21, attached certification
5.8.1 | Please provide unique ref number of the IEC ceriticate:
6 | COMPREHENSIVE MAINTENANCE
6.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) should be included in the unit price as per the SCC document
6.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
6.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
6.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
6.3 | The bidder to provide the OEM recommended service interval of the unit:
6.4 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
7 | General
7.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
7.2 | The unit on offer must be the latest technology.
7.2.1 | Initial Manufacturing year of Model:
RT2-04-006 | VIDEO COLPOSCOPE Replacement accessories | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
RT2-04-006a | Mouse
RT2-04-006b | Printer
RT2-04-006c | Keyboard
RT2-04-006d | A4 or A6 paper (1000 sheets)
RT2-04-006e | PC software
RT2-04-006f | Dust cover must be included and supplied for the colposcope system offered

## Vacuum Extractor
RT2-04-007 | Vacuum Extractor, Obstetric - Portable Electric | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
1 | DESCRIPTION
1.1 | A device that helps move a baby through the birth canal during a vaginal delivery using gentle suction.
2 | Product Offered:
2.1 | Make:
2.2 | Model:
2.3 | Initial Manufacturing year of Model:
3 | SPECIFICATIONS
3.1 | The unit must be designed for performing vacuum (ventonse) extraction of a baby in order to assist with childbirth.
3.2 | The suction unit being tendered on must be lightweight, easily portable, it must employ a reliable robust electric vacuum pump.
3.3 | The pump must be oilless, maintenance-free piston type.
3.4 | The unit must have a strong durable A.B.S. or equivalent corrosion resistant case with a handle.
3.5 | There must be a Jar Holder or Clip for a Jar to be fixed to the suction unit cage.
3.6 | The unit must have a graduated, polycarbonate collection jar with a minimum capacity of 2L with a float assembly in the jar lid. The jar must be suitable for use with disposable liners
3.7 | The collection jar must be autoclavable or other means of sterilization, please specify if not autoclavable
3.8 | The unit must have a vacuum control to adjust the vacuum over a range of minus 20Kpa to minus 90Kpa.
3.9 | The maximum vacuum must not be less than minus 90Kpa with a flow rate of not less than 14 to 15Lpm. N.B. The bidder must state the maximum vacuum and flow rate on the unit offered.
3.10 | The unit must have anti-bacterial filters fitted at the inlet side and exhaust side of the pump. These filters must be fitted where it is conveniently located such that they are easily accessible to the clinical staff for changing.
3.11 | In addition to the two normal filters specified above, in line shut-off filter must be supplied and fitted to the unit offered inside the housing of the unit and just before the inlet to the pump.
3.12 | The unit must have an ON / OFF switch fitted to the case.
3.13 | The vacuum gauge must be fitted to the pump case near the vacuum control knob.
3.14 | The motor must be 220 Volt + 10%, 50 Hz single phase AC.
3.15 | The motor must be protected by an inline circuit interrupter
3.16 | The noise level of the unit offered must not exceed 60dB.
3.17 | A complete set of re-usable silicone extractor cups with stainless steel handles and fittings such as chains, plates, filters and appropriate length tubing for each with cup sizes must be available.
3.18 | Please state the sterilization method that must be used to sterilize re-usable silicone extractor cups.
3.19 | All parts must be of a durable material construction and they must be corrosion resistant.
4 | STANDARD ACCESSORIES
4.1 | 1 x 40mm extractor cup
4.2 | 2 x 50mm extractor cups
4.3 | 2 x 60mm extractor cups
4.4 | 1 x 70mm extractor cup
4.5 | Polycarbonate Jar
4.6 | Silicone tubing (3m)
5 | Conformity Compliance (Please attach certificate)
5.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
5.2 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
5.2.1 | Please provide unique ref number of the ISO 13485 ceriticate:
5.3 | Model on offer must comply with EU Medical Device Regulation (EU) 2017/745 or equvalent from SAHPRA recognized jurisdictions . Attach a copy of certification
5.3.1 | Please provide unique ref number of the ceriticate:
5.4 | Hazardous Substance Act: If this type of equipment / apparatus appears on the schedule of Hazardous substances issued by the Directorate: Health Technology of the Department of Health, a license in terms of the Act on Hazardous Substances (Act 15/1973) must be submitted with the bid / quotation documents. The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.5 | All electrical/electronic medical equipment must be licensed by Radiation Control, where applicable a copy of the license must be submitted The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.5.1 | Radiation Control License number:
5.6 | The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade (rubber to with 2 screws) type and it must be a minimum length of three (3) meters. NB the mains cable of the unit tendered for must be SABS colour coded.
5.7 | The equipment quoted for must be protected against electromagnetic interference IEC 60601-1-2
5.8 | The unit must comply with electrical safety and performance standard such as IEC 60601-2-21, attached certification
5.8.1 | Please provide unique ref number of the IEC ceriticate:
6 | COMPREHENSIVE MAINTENANCE
6.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) should be included in the unit price as per the SCC document
6.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
6.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
6.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
6.3 | The bidder to provide the OEM recommended service interval of the unit:
6.4 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
7 | General
7.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
7.2 | The unit on offer must be the latest technology.
7.2.1 | Initial Manufacturing year of Model:
RT2-04-007 | Vacuum Extractor, Obstetric - Portable Electric | Complies Yes/No
RT2-04-007a | 1 x 40mm extractor cup
RT2-04-007b | 2 x 50mm extractor cups
RT2-04-007c | 2 x 60mm extractor cups
RT2-04-007d | 1 x 70mm extractor cup
RT2-04-007e | Polycarbonate Jar
RT2-04-007f | Silicone tubing (3m)

## Ultrasound, Obstetric
RT2-04-008 | Ultrasound Unit, Obstetric | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
1 | Description
1.1 | A medical device that uses high frequency sound waves to create real images or videos of the internal organs or other soft tissues for Abdominal, Vascular, Obstetric, Gynaecology, Musculo-Skeletal diagnostic purpose.
2 | Product Offered:
2.1 | Make:
2.2 | Model:
2.3 | Initial Manufacturing year of Model:
2.4 | Radiation Control License number:
3 | Technical Specification
3.1 | The unit must be a fully digital unit with a memory beam former and processor to generate sharp defined ultrasound images, reduce artefacts and improve image quality.
3.2 | It must utilise the latest software to perform investigations.
3.3 | High image quality will be one of the main criteria that will be used to evaluate the offers.
3.4 | Although offered units might comply with all the specifications, image quality must be a critical evaluation criterion, specifically with increased BMI cases.
3.5 | The unit must be of the latest technology.
3.6 | It must be a state of the art fully digital phase array ultrasound unit.
3.7 | The unit will be used for general applications as a multi-purpose unit for Abdominal, Vascular, Obstetric, Gynaecology, Musculo-Skeletal, Small Parts and Paediatric.
3.8 | IMAGING MODALITIES
3.8.1 | The unit must be latest technology; please state year in which this unit was first manufactured.
3.8.2 | The unit must offer excellent high image quality in B mode and fully digital colour flow Doppler. Colour Doppler should have no delay after activation.
3.8.3 | The system Dynamic Range must not be less than 172 dB. The Dynamic range should be able to be changed minimum of 4dB increments.
3.8.4 | The unit must have direct storage on hard drive of not less than 1TB. State all other storage facility that is included in the offered price. (Such as CD, DVD, MOD). Recorded CD roms must be compatible with any standard PC, without requiring extra software. USB download of images from the system must be possible
3.8.5 | A digital black and white thermal printer must be included in the offer for the machine.
3.8.6 | The system must offer a Digital Continuous Beam Former to generate sharp defined ultrasound beams, to reduce artefacts and to provide more consistent overall image quality.
3.8.7 | The unit must offer multi-frequency on all probes to provide sharp reproducible image quality in both superficial and deep anatomy.
3.8.8 | The unit must be equipped with a non-interlaced colour monitor with a screen size of not less than 21 inch, measured diagonally. The monitor must be of the flat screen TFT LCD type.
3.8.9 | The unit should have at least two (2) active probe connectors with electronic switching.
3.8.10 | The unit must be able to offer scanning methods for Electronic Convex and Linear technology.
3.8.11 | The unit must offer depth gain in not less than 8 steps. Unit must be programmable to start up automatically at a depth of operator’s choice and programmable to start up at different depths for specific examinations.
3.8.12 | Image processing must be done via a rotary dial or similar technology with at least 5 combinations for optimising the best image quality for each patient’s physique (user programmable). Image processing should visibly improve image and not just change it.
3.8.13 | Pan-Zoom enlargement must be in real and frozen mode up to 10X zoom.
3.8.14 | A colour pallet must be available for the colourization of the B mode and M mode and PW Doppler trace.
3.8.15 | Independent M mode and PW Doppler gain adjustment must be available.
3.8.16 | Sweep speed must be adjustable in at least 4 steps.Unit must have selectable sweep rates of 1,2,3,4,8 s/screen
3.8.17 | The unit must offer Tissue Harmonics as standard in the system.
3.8.18 | Grey scale level must be at least 256 shades.
3.8.19 | The unit must have not less than 1024 processing channels.
3.8.20 | The device must be built in a sturdy transportable trolley. All of the unit's components must be kept on the trolley, which also has enough room set aside for the probes and any disposables required for the machine. The device will be utilised for mobile ultrasonography in various locations. a) System width in centimetres b) System depth in centimetres.
3.8.21 | The frame rate must not be less than 394 frames per second. Indicate the frame rate that the system offers
3.8.22 | Image depth must be not less than 30cm.
3.8.23 | The unit should offer quick scan to automatically adjusts STC/TGC and gain through one-touch operation. This must visibly improve the image
3.8.24 | Post processing must be available on frozen images
3.8.25 | Unit must offer user programmable annotation for quick labelling. User must be able to program approximately 40 words of their own choice per selected examination type for annotation purposes
3.8.26 | Gestational age according to Last Menstrual Period(LMP) must be displayed on the working screen (not just on the patient information page)
3.9 | DOPPLER
3.9.1 | PW Doppler must be offered on all imaging probes.
3.9.2 | Maximum detectable velocity for PW must be stated.
3.9.3 | Pulse Repetition Frequency must vary between 0,7 – 19,8 kHz pulsed wave and 2,1 – 40kHz continuous wave
3.9.4 | Independent angle steering up to ±20° for colour and Doppler must be offered.
3.9.5 | Independent gain control for PW must be available.
3.9.6 | Colour pallet offering 7 types independent of B mode and Pulse Doppler pallet to be available.
3.9.7 | Sampling with Pulse Doppler must vary between 1 and 20 mm.
3.9.8 | Gate size must vary at least between 1 and 8mm.
3.9.9 | Base line shift must be possible in at least 8 steps.
3.9.10 | Spectral invert must be possible.
3.9.11 | At least 6 wall filter selections must be available. Bidder to state minimum.
3.9.12 | Accurate autotrace function for both arterial and venous waveforms must be available.
3.10 | COLOUR DOPPLER
3.10.1 | A minimum of 8 colour maps must be available for colour display.
3.10.2 | Colour maps must be available to be selected in real time and freeze mode.
3.10.3 | Independent colour gain adjustment must be from 1 to 16 dB in 0,5 dB increments. Indicate range
3.10.4 | Maximum colour frame rate must not be less than 285 frames per second.
3.10.5 | Colour flow must operate to the full depth of 30 cm.
3.10.6 | Colour maps and processing selections must be able to be selected by the user and stored in the presets.
3.10.7 | Quad signal processing and parallel signal processing must be available. Bidders to supply details.
3.10.8 | Power Doppler on all transducers must be available.
3.10.9 | The unit offered must be able of real time triplex mode.
3.10.10 | Colour mapping must be instantaneous – no time delay /lag.
3.11 | IMAGE MEMORY
3.11.1 | Image memory to operate in both dual and quad screen format.
3.11.2 | ECG synchronised recording must be possible. (optional)
3.11.3 | It must be possible to record B loop up to 500 frames.
3.11.4 | Cine loop and measurements must be active simultaneously.
3.11.5 | M/D review of a minimum of 15 seconds must be possible.
3.11.6 | B Mode Imaging on linear transducer to be electronically steerable.
3.12 | TRANSDUCERS
3.12.1 | The following frequencies to address the relevant applications must be possible
3.12.2 | Curvi-linear probe (Abdominal, vascular, O&G applications- 2.0 to 7.0 MHz
3.12.3 | Linear probe 4.5 to 16 MHz
3.12.4 | Transvaginal probe 3.9 to 15.0 MHz
3.12.5 | All imaging transducers must be multi-frequency.
3.12.6 | Tissue Harmonics must be available on all offered transducers.
3.12.7 | The selected frequency of the imaging transducer must be displayed on the screen.
3.12.8 | All transducers must be digitally controlled so continuous beam forming is possible. State number of focal zones.
3.12.9 | A minimum of 2 active probes must be available. It must be possible to electronically switch between these 2 probes.
3.13 | MEASUREMENTS
3.13.1 | Patient data MUST offer the option to enter previous ultrasound date and measurements, (not just LMP from previous examination).
3.13.2 | Gestational age according to LMP must be displayed on working screen, not just on patient information page.
3.13.3 | On-screen measurements must be possible from real time. 16 callipers (8 channel)each individually numbered for measuring distance/volume/angle and area must be offered
3.13.4 | Full range of obstetrics measurements and weight charts which must include Hadlock, Chitty and others.
3.13.5 | The unit must offer a storage facility for foetal growth measurement data. It must be possible to retrieve the data for follow-ups.
3.13.6 | The unit must offer 50th percentile growth graphs of the Doppler measurements RI/RP. It must be possible to retrieve the data for follow-ups.
3.13.7 | 50th Percentile growth graphs must be offered on all measurements for example BPD,HC, AC, FL Fetal weight . It must be possible to retrieve the data for follow-ups.
3.13.8 | The unit must offer user programmable annotation for quick labelling.
3.13.9 | State if the unit comply to the following: 1. Panoramic view. 2. Quad signal processing
3.13.10 | All software upgrades must be offered optionally, with full description what improvement the software upgrade will bring to the offered unit.
3.13.11 | The unit must be supplied with an uninterrupted power supply (UPS), to be able to keep the unit running for at least two(2) hours in case of power failure.
3.14 | DICOM COMPLIANCE
3.14.1 | The offered unit must be fully DICOM compliant.
3.14.2 | It must be possible for the unit to be connected to PACS.
4 | Standard Accessories and Consumables
4.1 | Curvi-linear probe
4.2 | Linear probe
4.3 | Transvaginal probe
4.4 | Printer
4.5 | Ultrasound Gel 2L
4.6 | Uninterrupted Power Supply (UPS)
4.7 | Printing paper (box of 10 rolls)
5 | Conformity Compliance (Please attach certificate)
5.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
5.2 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
5.2.1 | Please provide unique ref number of the ISO 13485 ceriticate:
5.3 | Model on offer must comply with EU Medical Device Regulation (EU) 2017/745 or equvalent from SAHPRA recognized jurisdictions . Attach a copy of certification
5.3.1 | Please provide unique ref number of the ceriticate:
5.4 | Hazardous Substance Act: If this type of equipment / apparatus appears on the schedule of Hazardous substances issued by the Directorate: Health Technology of the Department of Health, a license in terms of the Act on Hazardous Substances (Act 15/1973) must be submitted with the bid / quotation documents. The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.5 | All electrical/electronic medical equipment must be licensed by Radiation Control, where applicable a copy of the license must be submitted The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.5.1 | Radiation Control License number:
5.6 | The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade (rubber to with 2 screws) type and it must be a minimum length of three (3) meters. NB the mains cable of the unit tendered for must be SABS colour coded.
5.7 | The equipment quoted for must be protected against electromagnetic interference IEC 60601-1-2
5.8 | The unit must comply with electrical safety and performance standard such as IEC 60601-2-21, attached certification
5.8.1 | Please provide unique ref number of the IEC ceriticate:
6 | COMPREHENSIVE MAINTENANCE
6.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) should be included in the unit price as per the SCC document
6.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
6.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
6.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
6.3 | The bidder to provide the OEM recommended service interval of the unit:
6.4 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
7 | General
7.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
7.2 | The unit on offer must be the latest technology.
7.2.1 | Initial Manufacturing year of Model:
RT2-04-008 | Ultrasound Unit Accessories | Complies Yes/No
RT2-04-008a | Curvi-linear probe
RT2-04-008b | Linear probe
RT2-04-008c | Transvaginal probe
RT2-04-008d | Printer
RT2-04-008e | Ultrasound Gel 2L
RT2-04-008f | Uninterrupted Power Supply (UPS)
RT2-04-008g | Printing paper (box of 10 rolls)