## Caterogy-1 Anaesthesia
ANAESTHESIA EQUIPMENT
# | Item number | Item Description
1 | RT2-01-01 | Anaesthetic Machine, Basic
2 | RT2-01-02 | Anaesthetic Machine, Basic with Monitor
3 | RT2-01-03 | Anaesthetic Machine, Standard
4 | RT2-01-04 | Anaesthetic Machine, Standard with Monitor
5 | RT2-01-05 | Anaesthetic Machine, Advanced
6 | RT2-01-06 | Anaesthetic Machine, Advanced with Monitor
7 | RT2-01-07 | Monitor, Anaesthetic
8 | RT2-01-08 | Anaesthetic Machine, MRI Compatible
9 | RT2-01-09 | Anaesthetic Machine, MRI Compatible with Monitor
10 | RT2-01-10 | Monitor, MRI Compatible
11 | RT2-01-11 | Video Laryngoscope
12 | RT2-01-12 | Video Laryngoscope with Flexible Bronchoscope
13 | RT2-01-13 | Video Disposable Flexible Bronchoscope
14 | RT2-01-14 | Infusion Pump, MRI compatible
15 | RT2-01-15 | Syringe Driver, TCI
16 | RT2-01-16 | Syringe Driver, PCA
17 | RT2-01-17 | Blood-Fluid Warmer
18 | RT2-01-18 | Patient Warmer, Forced-Air
19 | RT2-01-19 | Patient Warmer
20 | RT2-01-20 | Nerve Stimulator
21 | RT2-01-21 | Thromboelastogram
22 | RT2-01-22 | Cerebral Oximeter, Advanced
23 | RT2-01-23 | Ultrasound Machine, Point-of-care

## Anaesthetic Machine, Basic
RT2-01-01 | ANAESTHESIA DELIVERY SYSTEM: BASIC WITHOUT PHYSIOLOGICAL MONITOR | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
1 | Description
1.1 | A medical device designed to administer anesthetics to patients during surgical procedures.
2 | Product Offered (Anaesthetic Machine)
2.1 | Make:
2.2 | Model:
2.3 | Initial Manufacturing year of Model:
2.4 | Radiation Control License number:
Purpose of Use
3 | Technical Specification
3.1 | The complete anaesthesia workstation shall have (but not be limited to) the following aspects, features and components, as described in the specifications:
3.2 | Anaesthesia delivery system including 2x vaporisers
3.3 | Ventilator suitable for ventilating paediatric to adult patients
3.4 | The unit shall be easily moved around without the possibility of toppling over
3.5 | The frame of the unit shall be manufactured from corrosion free material, yet be strong and durable
3.6 | The unit shall be provided with at least four castors that comply with SANS 621 detailing the castor's material type, size and movement. The castors shall be equipped with a braking system
3.7 | The unit shall be supplied with a swivel monitoring shelf to mount the physiological monitor. The monitor (described later in the specifications) shall be mounted on the frame of the machine with a dedicated monitor mounting system
3.8 | The unit shall have at least two drawers for secure storage of airway management equipment and other critical requirements
3.9 | The unit shall be supplied with at least three (3) extra 220 Volt 15A SABS approved power sockets for additional monitoring/syringe pumps
3.10 | The unit shall have a working area, as well as a pull-out or flip-up writing table for writing notes
3.11 | The unit shall preferably be equipped with a working light to enable the anaesthetist to work in a darkened theatre
3.12 | The unit shall incorporate a suction bottle mounting rail (Gabler) on the side of the machine, with a mounted suction control
3.13 | All critical controls shall be within reach of an anaesthetist sitting in front of the machine
3.14 | A anaesthetic machine master switch is essential for start-up of both pneumatic and electrical supplies. Both must switch on simultaneously with a single action
3.15 | The unit shall have a routine electronic checklist for diagnostic testing of flow delivery and preferably circuit integrity. It shall be possible to bypass this in the event of urgent machine requirement
3.16 | Pipeline gas inlets shall be non-interchangeable. Hoses shall be colour-coded, and shall be connectable to the medical service outlets in ……… Hospital. It shall be possible to connect the machine to the outlets for Nitrous Oxide, Oxygen, Vacuum and Medical Air (Standard colours according National Standards)
3.17 | At least one cylinder regulator shall be provided for Oxygen. This must be pin-indexed as per ISO standard. The trolley must have an attachment to support the cylinder
3.18 | Gas regulators must preferably have metal diaphragms.
3.19 | A safety pressure relief valve must protect the entire pneumatic system from excessive input pressure from pipeline and cylinder supply. Bidders to state the pressure relief setting on this system
3.20 | Back flow of gases into pipeline or cylinders shall be prevented by a non-return valve system. With a cylinder detached from the pin-indexed yoke and the pipeline hoses connected, no gas leak shall occur at the yoke. Preferential flow shall be facilitated from the pipeline inlets. Bidders to describe how this is achieved
3.21 | Gauges must be supplied for each individual gas. Pipeline and cylinder gauges shall be differentially positioned, and clearly indicated for purpose
3.22 | The anaesthetic machine or ventilator must have a built-in manometer gauge, calibrated between the range of -20 to +100cm water. A digital readout is acceptable
3.23 | All internal pneumatic circuitry must be contained within the enclosed work surface, to avoid pulling and accidental disconnection of gas pipes. All gas tubing must be diameter indexed with non-interchangeable connections
3.24 | The unit shall be equipped to administer low flow anaesthesia. For this reason, the unit must be equipped with rotameters that can display(show) the flow rates for high and low flow anaesthesia, typically by using cascading rotameters in the following range:
3.24.1 | Oxygen low flow tube 200ml/minute – 1 litre/min At higher flow rates the gas must cascade into a second flow tube (1 – 10 l/min)
3.24.2 | Nitrous Oxide low flow tube 200ml/minute – 1 litre/min At higher flow rates the gas must cascade into a second flow tube (1 – 10 l/min)
3.24.3 | Air tube 200ml – 10litres/min Cascading tubes as above are acceptable, but not essential
3.25 | Flowmeter tubes shall preferably be diameter indexed to prevent damage and interchangeability during transportation service and maintenance. Secondary regulators must be fitted between incoming pipeline gas supply and the flowmeter control valves to ensure stability of the bobbin at low end of scale (This applies if manual or mechanical flowmeters are offered, this point does not apply if electronic flowmeters are offered.)
3.26 | Machines supplied with electronic flowmeters must be provided with a manual flow meter for oxygen delivery independent of electronic control
3.27 | Flowmeter sequence must be in accordance with South African Society of Anaesthesiologists (SASA) Practice Guidelines – (Oxygen on the right)
3.28 | The unit shall include a safety mechanism that ensures a minimum concentration of 25% oxygen in a oxygen/nitrous oxide gas mixture across all flow rates above 1 litre. It is desired that at flow rates of below 1 litre, the Oxygen flow rate cannot be decreased below 250ml/min
3.29 | The Anaesthetic machine shall be equipped with an oxygen monitor – Even if the monitor has an oxygen monitoring capability, it shall also be built into the machine
3.30 | The oxygen monitor shall have a low and high alarm – visual and audible. The low alarm shall automatically default to no less than 18%.
3.31 | It shall be possible to calibrate the low oxygen control to 21%.
3.32 | The unit shall include a high-flow Oxygen flush mechanism. This system shall bypass the rotameters and vaporisers, and shall be non-lockable. The activation mechanism shall be a recessed, self closing push button, and shall provide an Oxygen flush in excess of 30 litres/min
3.33 | The unit must incorporate a continuous alarm should the Oxygen pressure fall below 193 KPa. Should pressure further fall below 138 KPa, then all Nitrous Oxide must be automatically cut off from the patient circuit
3.34 | The unit must have a vaporiser manifold which is of the lock-on type, allowing the vaporiser to be removed from the machine, and replaced with an alternative. No tools shall be required to perform this action
3.35 | The unit must have 2 slots/stations for at least 2 vaporisers, to be mounted side by side
3.36 | The machine/vaporiser must automatically prevent more than one vaporiser from being turned on simultaneously. This must be achieved by an interlock mechanism, and not by a separate vaporiser selector switch
3.37 | Two vaporisers shall be supplied with each machine
3.37.1 | 1 X Isoflurane
3.37.2 | 1 X Sevoflurane
3.38 | Each vaporiser shall be colour-coded, and provided with a safety device to avoid accidental filling with incorrect product according to International Standards
3.39 | The unit must have an open-end gas scavenging reservoir fitted. Active scavenging from the reservoir must be approximately 30 litres/minute. Scavenging outlet must be compatible with the scavenging system being installed at the users's Hospital
3.40 | The unit must be equipped with a carbon dioxide circle absorber, APL valve, antipollution hood and ventilator-manual switch. This switch when directed to the ventilator shall isolate the APL valve from the circle. All patient connectors shall conform to ISO standards (male with outer diameter of 22mm and inner diameter of 15mm)
3.41 | Two non-disposable canisters must be included in the offer. The bidder must state the type of material of the canisters and the temperature which the canister will reach the stirility point.
3.42 | The soda lime must be isolated while the circle is in use.
3.43 | The absorber system shall be either integrated into the frame or be positioned under the shelf for transport purposes
3.44 | All materials in contact with the exhaled patient gases shall be autoclavable, (excluding flow sensors and Oxygen cell)
3.45 | A water trap to accommodate excess water generated is essential. This shall preferably be integrated into the soda lime canister
3.46 | The frame shall be equipped with a power back-up system to enable full operation of the anaesthetic machine and ventilator for at least 60 minutes with batteries fully charged, under typical operating conditions
3.47 | VENTILATOR
3.47.1 | The ventilator shall allow a “closed-loop” monitoring configuration to ensure that the selected volume or pressure is achieved
3.47.2 | Any change to fresh gas flow shall be compensated
3.47.3 | Any flow sensors used shall be robust and easily removed for cleaning
3.47.4 | Flow sensors shall be remote from patient, or integrated within the system
3.47.5 | Bidders to clearly state the system used, as well as cost and frequency of replacement required for the flow sensors
3.47.6 | The ventilator shall be able to ventilate patients in Volume Controlled and Pressure Controlled modes
3.47.7 | It shall be possible to ventilate patients with a Tidal Volume from 30ml to 1 500ml
3.47.8 | The following (approximate) ventilator parameter ranges shall be applicable:
3.47.8.1 | Respiration Rate: 4 to 80 breaths/min
3.47.8.2 | Minute Volume: 0 to 60 l/min
3.47.8.3 | Tidal Volume: 30 to 1 500ml. (State different values for Volume and Pressure Controlled modes)
3.47.8.4 | Be able to provide PEEP in the range: 0 to 25cm Water
3.47.9 | Bidders to state all monitoring capabilities offered as part of the ventilator
3.47.10 | The following alarms shall form part of the capabilities of the offered ventilator:
3.47.10.1 | Tidal Volume
3.47.10.2 | Minute Volume
3.47.10.3 | Inspired Oxygen (Fi02)
3.47.10.4 | Apnoea
3.47.10.5 | Low Airway Pressure
3.47.10.6 | High Airway pressure
3.47.10.7 | Bidders to clearly indicate that these alarms are available, as well as the ranges and settings possible
3.47.11 | The ventilator shall have a mute alarm function, but it shall not be possible to cancel the alarm until the problem has been rectified. Mute function shall only be applicable for two minutes time period
3.49.12 | The ventilator shall have power back-up to allow full function of the ventilator and anaesthetic machine for not less than 60 minutes
4 | A starter-pack of consumables must be supplied with the unit comprising at least the following for both Paediatric and Adult patients:
4.1 | Adult Patients starter-pack:
4.1 | Anaesthesia circuit (Adult, single/double limb design) Disposable. 1.8m (Complete with Y-piece, Elbow, Catheter mount, Rebreathing bag) x10
4.2 | Anaesthesia circuit (Paediatric, single/double limb design) Disposable. 1.8m (Complete with Y-piece, Elbow, Catheter mount, Rebreathing bag) x10
4.3 | Re-breathing circuit complete x10
4.4 | Non-re-breathing circuit complete x10
4.5 | Flow sensor (pack of 2) x5
4.6 | Spirometry line x10
4.7 | Sample line (pack of 10)
4.8 | Face masks: (sizes 00, 0, 1, 2, 3, 4, 5) x10
4.9 | Bacterial/viral filter x10
4.11 | Manual ventilation hose x10
4.12 | Fresh gas tubing
4.13 | Seal for back-up cylinder connection on the unit
4.14 | Scavenger tubing (as per institution's scavenging outlet)
4.15 | Inspiratory and expiratory unidirectional valves
4.16 | Side stream CO2 kit (adult/paed)
4.17 | Disposable Adult Spirometry Set complete with spirometry tube, sample line and flow sensor, 3 m, box of 20 pieces
4.18 | Disposable Paediatric Spirometry Set complete with spirometry tube, sample line and flow sensor, 3 m, box of 20 pieces
4.19 | Gas sample tube with luer lock connection, 3m, disposable (10 pcs)
4.20 | Suction Tubing Set inc connector
4.21 | Resuscitator Adult, 1500mL, autoclavable, silicone
4.22 | Resuscitator Paediatric, 550mL, autoclavable, silicone
4.23 | Tubing set Paediatric for circle systems
5 | Conformity Compliance (Please attach certificate)
5.1 | The unit must comply with electrical safety and performance standard such as IEC 60601-2-13 for Anaesthetic systems, attached certification
5.1.1 | Please provide unique ref number of the IEC ceriticate:
5.2 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
5.2.1 | Please provide unique ref number of the ISO 13485 ceriticate:
5.3 | Model on offer must comply with EU Medical Device Regulation (EU) 2017/745 or equvalent from SAHPRA recognized jurisdictions . Attach a copy of certification
5.3.1 | Please provide unique ref number of the ceriticate:
5.4 | Hazardous Substance Act: If this type of equipment / apparatus appears on the schedule of Hazardous substances issued by the Directorate: Health Technology of the Department of Health, a license in terms of the Act on Hazardous Substances (Act 15/1973) must be submitted with the bid / quotation documents. The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.5 | All electrical/electronic medical equipment must be licensed by Radiation Control, where applicable a copy of the license must be submitted The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.5.1 | Radiation Control License number:
5.6 | The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade (rubber to with 2 screws) type and it must be a minimum length of three (3) meters. NB the mains cable of the unit tendered for must be SABS colour coded and approved.
6 | COMPREHENSIVE MAINTENANCE
6.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) should be included in the bid price as per the SCC document
6.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries, Flow sensors and Oxygen sensor), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
6.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
6.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries, Flow sensors and Oxygen sensor), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
6.3 | The bidder to provide the OEM recommended service interval of the unit:
6.4 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
7 | General
7.1 | Has the product on offer or any of its components been recalled by the Regulatory Authority or OEM initiated recall in the last 5 years? If yes, please provide details of the type of recall and how was it addressed
7.2 | The unit on offer must be the latest technology.
Initial Manufacturing year of Model:
7.2.1 | Latest year update of Model:
| ANAESTHESIA DELIVERY SYSTEM: BASIC WITHOUT PHYSIOLOGICAL MONITOR, REPLACEMENT ACCESSORIES | Complies Yes/No
RT2-01-01a | Anaesthesia circuit (Adult, single/double limb design) Disposable. 1.8m (Complete with Y-piece, Elbow, Catheter mount, Rebreathing bag)
RT2-01-01b | Anaesthesia circuit (Paediatric, single/double limb design) Disposable. 1.8m (Complete with Y-piece, Elbow, Catheter mount, Rebreathing bag)
RT2-01-01c | Anaesthesia circuit (neonatal, single/double limb design) Disposable. 1.8m (Complete with Y-piece, Elbow, Catheter mount, Rebreathing bag)
RT2-01-01d | Re-breathing circuit complete
RT2-01-01e | Non-re-breathing circuit complete
RT2-01-01f | Flow sensor (pack of 2)
RT2-01-01g | Spirometry line
RT2-01-01h | Sample line (pack of 10)
RT2-01-01i-1 | Face masks: size 00
RT2-01-01i-2 | Face masks: size 0
RT2-01-01i-3 | Face masks: size 1
RT2-01-01i-4 | Face masks: size 2
RT2-01-01i-5 | Face masks: size 3
RT2-01-01i-5 | Face masks: size 4
RT2-01-01i-6 | Face masks: size 5
RT2-01-01j | Bacterial/viral filter
RT2-01-01k | Y connector
RT2-01-01l | Manual ventilation hose
RT2-01-01m | Fresh gas tubing
RT2-01-01n | Seal for back-up cylinder connection on the unit
RT2-01-01o | Scavenger tubing (as per institution's scavenging outlet)
RT2-01-01p | Inspiratory and expiratory unidirectional valves
RT2-01-01q | Side stream CO2 kit (adult/paed)
RT2-01-01r | Disposable Adult Spirometry Set complete with spirometry tube, sample line and flow sensor, 3 m, box of 20 pieces
RT2-01-01s | Disposable Paediatric Spirometry Set complete with spirometry tube, sample line and flow sensor, 3 m, box of 20 pieces
RT2-01-01t | Gas sample tube with luer lock connection, 3m, disposable (10 pcs)
RT2-01-01u | Suction Tubing Set inc connector
RT2-01-01v | Resuscitator Adult, 1500mL, autoclavable, silicone
RT2-01-01w | Resuscitator Paediatric, 550mL, autoclavable, silicone
RT2-01-01x | Tubing set Paediatric for circle systems
RT2-01-01y | sevoflurane vaporizer
RT2-01-01z | Isoflurane vaporizer
RT2-01-01zz | Desflurane vaporizer

## Anaesthetic, Basic + monitor
RT2-01-02 | ANAESTHESIA DELIVERY SYSTEM: Basic with Integrated Physiological Monitor | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
1 | Description
1.1 | A medical device designed to administer anesthetics to patients during surgical procedures.
1 | Product Offered (Anaesthetic Machine)
1.1a | Make:
1.2b | Model:
1.3c | Initial Manufacturing year of Model:
1.4d | Radiation Control License number:
2 | Product Offered (Monitor)
2.1a | Make:
2.2b | Model:
2.3c | Initial Manufacturing year of Model:
2.4d | Radiation Control License number:
3 | Technical Specification
3.1 | The complete anaesthesia workstation shall have (but not be limited to) the following aspects, features and components, as described in the specifications:
3.2 | Anaesthesia delivery system including 2x vaporisers
3.3 | Ventilator suitable for ventilating paediatric to adult patients
3.4 | The unit shall be easily moved around without the possibility of toppling over
3.5 | The frame of the unit shall be manufactured from corrosion free material, yet be strong and durable
3.6 | The unit shall be provided with at least four castors that comply with SANS 621 detailing the castor's material type, size and movement. The castors shall be equipped with a braking system
3.7 | The unit shall be supplied with a swivel monitoring shelf to mount the physiological monitor. The monitor (described later in the specifications) shall be mounted on the frame of the machine with a dedicated monitor mounting system
3.8 | The unit shall have at least two drawers for secure storage of airway management equipment and other critical requirements
3.9 | The unit shall be supplied with at least three (3) extra 220 Volt 15A SABS approved power sockets for additional monitoring/syringe pumps
3.10 | The unit shall have a working area, as well as a pull-out or flip-up writing table for writing notes
3.11 | The unit shall preferably be equipped with a working light to enable the anaesthetist to work in a darkened theatre
3.12 | The unit shall incorporate a suction bottle mounting rail (Gabler) on the side of the machine, with a mounted suction control
3.13 | All critical controls shall be within reach of an anaesthetist sitting in front of the machine
3.14 | A anaesthetic machine master switch is essential for start-up of both pneumatic and electrical supplies. Both must switch on simultaneously with a single action
3.15 | The unit shall have a routine electronic checklist for diagnostic testing of flow delivery and preferably circuit integrity. It shall be possible to bypass this in the event of urgent machine requirement
3.16 | Pipeline gas inlets shall be non-interchangeable. Hoses shall be colour-coded, and shall be connectable to the medical service outlets in ……… Hospital. It shall be possible to connect the machine to the outlets for Nitrous Oxide, Oxygen, Vacuum and Medical Air (Standard colours according National Standards)
3.17 | At least one cylinder regulator shall be provided for Oxygen. This must be pin-indexed as per ISO standard. The trolley must have an attachment to support the cylinder
3.18 | Gas regulators must preferably have metal diaphragms.
3.19 | A safety pressure relief valve must protect the entire pneumatic system from excessive input pressure from pipeline and cylinder supply. Bidders to state the pressure relief setting on this system
3.20 | Back flow of gases into pipeline or cylinders shall be prevented by a non-return valve system. With a cylinder detached from the pin-indexed yoke and the pipeline hoses connected, no gas leak shall occur at the yoke. Preferential flow shall be facilitated from the pipeline inlets. Bidders to describe how this is achieved
3.21 | Gauges must be supplied for each individual gas. Pipeline and cylinder gauges shall be differentially positioned, and clearly indicated for purpose
3.22 | The anaesthetic machine or ventilator must have a built-in manometer gauge, calibrated between the range of -20 to +100cm water. A digital readout is acceptable
3.23 | All internal pneumatic circuitry must be contained within the enclosed work surface, to avoid pulling and accidental disconnection of gas pipes. All gas tubing must be diameter indexed with non-interchangeable connections
3.24 | The unit shall be equipped to administer low flow anaesthesia. For this reason, the unit must be equipped with rotameters that can display(show) the flow rates for high and low flow anaesthesia, typically by using cascading rotameters in the following range:
3.24.1 | Oxygen low flow tube 200ml/minute – 1 litre/min At higher flow rates the gas must cascade into a second flow tube (1 – 10 l/min)
3.24.2 | Nitrous Oxide low flow tube 200ml/minute – 1 litre/min At higher flow rates the gas must cascade into a second flow tube (1 – 10 l/min)
3.24.3 | Air tube 200ml – 10litres/min Cascading tubes as above are acceptable, but not essential
3.25 | Flowmeter tubes shall preferably be diameter indexed to prevent damage during transportation service and maintenance. Secondary regulators must be fitted between incoming pipeline gas supply and the flowmeter control valves to ensure stability of the bobbin at low end of scale (This applies if manual or mechanical flowmeters are offered, this point does not apply if electronic flowmeters are offered.)
3.26 | Machines supplied with electronic flowmeters must be provided with a manual flow meter for oxygen delivery independent of electronic control
3.27 | Flowmeter sequence must be in accordance with South African Society of Anaesthesiologists (SASA) Practice Guidelines – (Oxygen on the right)
3.28 | The unit shall include a safety mechanism that ensures a minimum concentration of 25% oxygen in a oxygen/nitrous oxide gas mixture across all flow rates above 1 litre. It is desired that at flow rates of below 1 litre, the Oxygen flow rate cannot be decreased below 250ml/min
3.29 | The Anaesthetic machine shall be equipped with an oxygen monitor – Even if the monitor has an oxygen monitoring capability, it shall also be built into the machine
3.30 | The oxygen monitor shall have a low and high alarm – visual and audible. The low alarm shall automatically default to no less than 18%.
3.31 | It shall be possible to calibrate the low oxygen control to 21%.
3.32 | The unit shall include a high-flow Oxygen flush mechanism. This system shall bypass the rotameters and vaporisers, and shall be non-lockable. The activation mechanism shall be a recessed, self closing push button, and shall provide an Oxygen flush in excess of 30 litres/min
3.33 | The unit must incorporate a continuous alarm should the Oxygen pressure fall below 193 KPa. Should pressure further fall below 138 KPa, then all Nitrous Oxide must be automatically cut off from the patient circuit
3.34 | The unit must have a vaporiser manifold which is of the lock-on type, allowing the vaporiser to be removed from the machine, and replaced with an alternative. No tools shall be required to perform this action
3.35 | The unit must have 2 slots/stations for at least 2 vaporisers, to be mounted side by side
3.36 | The machine/vaporiser must automatically prevent more than one vaporiser from being turned on simultaneously. This must be achieved by an interlock mechanism, and not by a separate vaporiser selector switch
3.37 | Two vaporisers shall be supplied with each machine
3.37.1 | 1 X Isoflurane
3.37.2 | 1 X Sevoflurane
3.38 | Each vaporiser shall be colour-coded, and provided with a safety device to avoid accidental filling with incorrect product according to International Standards
3.39 | The unit must have an open-end gas scavenging reservoir fitted. Active scavenging from the reservoir must be approximately 30 litres/minute. Scavenging outlet must be compatible with the scavenging system being installed at the users's Hospital
3.40 | The unit must be equipped with a carbon dioxide circle absorber, APL valve, antipollution hood and ventilator-manual switch. This switch when directed to the ventilator shall isolate the APL valve from the circle. All patient connectors shall conform to ISO standards (male with outer diameter of 22mm and inner diameter of 15mm)
3.41 | Two non-disposable canisters must be included in the offer. The bidder must state the type of material of the canisters and the temperature which the canister will reach the stirility point.
3.42 | The soda lime must be isolated while the circle is in use.
3.43 | The absorber system shall be either integrated into the frame or be positioned under the shelf for transport purposes
3.44 | All materials in contact with the exhaled patient gases shall be autoclavable, (excluding flow sensors and Oxygen cell)
3.45 | A water trap to accommodate excess water generated is essential. This shall preferably be integrated into the soda lime canister
3.46 | The frame shall be equipped with a power back-up system to enable full operation of the anaesthetic machine and ventilator for at least 60 minutes with batteries fully charged, under typical operating conditions
3.47 | VENTILATOR
3.47.1 | The ventilator shall allow a “closed-loop” monitoring configuration to ensure that the selected volume or pressure is achieved
3.47.2 | Any change to fresh gas flow shall be compensated
3.47.3 | Any flow sensors used shall be robust and easily removed for cleaning
3.47.4 | Flow sensors shall be remote from patient, or integrated within the system
3.47.5 | Bidders to clearly state the system used, as well as cost and frequency of replacement required for the flow sensors
3.47.6 | The ventilator shall be able to ventilate patients in Volume Controlled and Pressure Controlled modes
3.47.7 | It shall be possible to ventilate patients with a Tidal Volume from 30ml to 1 500ml
3.47.8 | The following (approximate) ventilator parameter ranges shall be applicable:
3.47.8.1 | Respiration Rate: 4 to 80 breaths/min
3.47.8.2 | Minute Volume: 0 to 60 l/min
3.47.8.3 | Tidal Volume: 30 to 1 500ml. (State different values for Volume and Pressure Controlled modes)
3.47.8.4 | Be able to provide PEEP in the range: 0 to 25cm Water
3.47.9 | Bidders to state all monitoring capabilities offered as part of the ventilator
3.47.10 | The following alarms shall form part of the capabilities of the offered ventilator:
3.47.10.1 | Tidal Volume
3.47.10.2 | Minute Volume
3.47.10.3 | Inspired Oxygen (Fi02)
3.47.10.4 | Apnoea
3.47.10.5 | Low Airway Pressure
3.47.10.6 | High Airway pressure
3.47.10.7 | Bidders to clearly indicate that these alarms are available, as well as the ranges and settings possible
3.47.11 | The ventilator shall have a mute alarm function, but it shall not be possible to cancel the alarm until the problem has been rectified. Mute function shall only be applicable for two minutes time period
3.49.12 | The ventilator shall have power back-up to allow full function of the ventilator and anaesthetic machine for not less than 60 minutes
3.51 | MONITOR
3.51.1 | The anaesthetic machine shall be equipped with a monitor specifically designed for use in operating theatres
3.51.2 | The monitor functions can be incorporated in the design of the anaesthetic machine, or can be a separate monitor mounted on the anaesthetic machine on a suitable swivel arm
3.51.3 | The monitor shall be able to monitor and display the following parameters:
3.51.3.1 | ECG
3.51.3.2 | Respiration
3.51.3.3 | Pulse Oximetry (Oxygen Saturation)
3.51.3.4 | Non Invasive Blood Pressure
3.51.3.5 | Side stream capnography
3.51.3.6 | Automatic Agent Monitoring
3.51.3.7 | 2 channels of Invasive Blood Pressure
3.51.3.8 | Temperature
3.51.4 | The monitor shall be of modern design based on ergonomic principles
3.51.5 | The front panel shall be designed to facilitate easy cleaning and disinfecting
3.51.6 | Preference will be given to units incorporating touch screen as well as “navigation point” controls for easy information input and onscreen navigation
3.51.7 | The display of the monitor shall have a display area of not less than 26 cm in size (diagonally measured: from corner to corner)
3.51.8 | The display shall be flicker free and easy to read. Bidders to state if brightness control is possible
3.51.9 | Preference will be given to units with a colour display, and different parameters displayed in different colours
3.51.10 | In addition to the waveform display section, the monitors shall have an area dedicated for alphanumeric display of parameter identification and measured values, alarm and status messages
3.51.11 | Bidders to state if the screen can be configured by the user to display parameters in different set-ups
3.51.12 | Bidders to state if trending is possible on the monitor, and state trending options
3.51.13 | Audible and visual alarms shall be available for all measured parameters. State the method employed to avoid “nuisance alarms”
3.51.14 | Bidders to state if alarms are graded into levels of urgency, and explain the method and grading
3.51.15 | The monitor shall provide clinically reasonable default alarms when switched on. These alarms shall be programmable, and shall be capable of being overridden manually
3.51.16 | It shall be possible to override the audible alarm for a short period, but no override of the visual alarm shall be possible
3.51.17 | It shall be possible to suspend the alarms of certain parameters during clinical procedures, but the alarm shall reactivate automatically after a suitable time delay. Any silencing or disabling of alarms shall be clearly indicated on the monitor
3.51.18 | Input impedance of the system shall be of such a nature as to protect the patient at all times
3.51.19 | Maximum allowable leakage current to the patient shall be less than 10 µA
3.51.20 | The monitor shall be protected against defibrillator and electro surgical interference
3.51.21 | The monitor shall be of low power consumption, and shall preferably not require fan or other forced ventilation
3.51.22 | All cables and sockets shall be colour coded for easy recognition, and shall conform to AAMI quality standards or equivalents
3.51.23 | The monitor shall be equipped with a battery to keep the monitor functional for at least 15 minutes after power failure
3.51.24 | ECG/RESPIRATION
3.51.24.1 | A full set of 3 lead ECG cables shall be included in the offer
3.51.24.2 | The ECG amplifier shall employ a differential input with high noise suppression
3.51.24.3 | It shall be possible to change the sweep speed of the ECG between 25 and 50mm/sec
3.51.24.4 | Bidders to state if a freeze facility for ECG is available
3.51.24.5 | The input impedance shall be higher than 10 MegOhms
3.51.24.6 | The heart rate measuring range shall be between 0 and 250 beats per minute, with an accuracy of not more than 5 bpm
3.51.24.7 | An AC filter shall be provided
3.51.24.8 | Alarm limits shall be adjustable in the following range
3.51.24.8.1 | Low – 40 to 150
3.51.24.8.2 | High – 70 to 250
3.51.24.9 | The violation of alarm limits shall result in an audible and visual alarm. It shall be possible to temporarily silence the audible alarm. State the time duration of the temporary silence
3.51.24.10 | The respiration rate shall be measured by means of impedance via the ECG leads
3.51.24.11 | Measurement range for respiration shall be approximately 4 to 120 breaths/min
3.51.25 | NON INVASIVE BLOOD PRESSURE
3.51.25.1 | The monitor shall make use of the oscillometric or other reliable measuring method
3.51.25.2 | Systolic, Diastolic and Mean pressure measurements shall be displayed
3.51.25.3 | It shall be possible to program the sampling interval for automatic NIBP measurements
3.51.25.4 | The shortest period of sampling shall not be longer than one minute
3.51.25.5 | The cuff pressure shall be limited to a comfortable level for both paediatric and adult patients
3.51.25.6 | The NIBP cuffs shall be robust, interchangeable and easy to clean
3.51.25.7 | A full set of reusable cuffs shall be supplied with each monitor. This shall include the range from neonatal, paediatric to obese and thigh cuffs. Bidders to state which cuffs are included in the offer
3.51.25.8 | Two sets of NIBP hoses shall be included in the offer per monitor– One set for adult measurement, and one set for paediatric measurement
3.51.25.9 | It shall be possible to set high and low alarm limits for Systolic, Diastolic and Mean Arterial Pressure. Default settings to be stated
3.51.26 | OXYGEN SATURATION
3.51.26.1 | Continuous monitoring of oxygen saturation levels by means of the photometric principle shall be possible
3.51.26.2 | The measuring range shall be from low to 100%. Bidders to state the accuracy levels
3.51.26.3 | It must be the latest technology with clinical research support based on accuracy and real-time monitoring. Please submit the published peer reviewed/researched paper as supporting documents.
3.51.26.4 | A finger probe for adults shall be included in the offer, and a suitable probe for paediatric use. These shall be per monitor. Bidders to state the replacement price for each probe, and state other types of probes available
3.51.26.5 | The monitor shall provide an audible tone indicating fall or change in oxygen saturation
3.51.26.6 | The monitor shall have a continuous oxygen saturation trace, and units displaying oxygen saturation with only a bar graph will not be acceptable
3.51.27 | CAPNOGRAPHY / ANAESTHETIC AGENT MONITORING
3.51.27.1 | The capnography shall be an integral part of the monitor, and no loose attachments shall be acceptable
3.51.27.2 | The capnography measurement shall be in kPa and mmHg
3.51.27.3 | Bidders to state the method employed to get rid of water build-up in the measuring line
3.51.27.4 | Bidders to state the cost of the measuring line. A set of at least 10 lines shall be included per monitor
3.51.27.5 | The monitor shall also include automatic agent identification of at least the following anaesthetic agents
3.51.27.5.1 | Halothane
3.51.27.5.2 | Isoflurane
3.51.27.5.3 | Sevoflurane
3.51.27.6 | The unit shall automatically alert when an anesthetic agent mixture is detected
3.51.27.7 | Inspiration and expiration values of all gases shall be displayed
3.51.27.8 | The unit shall have individual alarm limits for each parameter
3.51.27.9 | The unit shall have automatic zero calibration
3.51.28 | INVASIVE BLOOD PRESSURE
3.51.28.1 | The monitor shall have the ability to monitor 2 channels of Invasive Pressure
3.51.28.2 | Measurement range shall be approximately -40 to 320mm Hg
3.51.29 | TEMPERATURE
3.51.29.1 | The monitor shall be able to measure at least one temperature
3.51.29.2 | Temperature range shall be approximately 10º to 45º Celsius
3.51.29.3 | Two temperature probes shall be included per monitor – one oral and one rectal probe
4 | The following Accessories and Consumables must be supplied with the unit:
4.1 | Re-breathing bag, reusable, rubber
4.2 | Reservoir bag
4.3 | Pressure limiting valve, adjustable
4.4 | Oxygen fuel cell analyser with fuel cell
4.5 | Adult & Paediatric volume sensor re-usable
4.6 | Water trap for multigas monitor, re-usable (10 pcs.)
4.7 | A full set paediatric 3/5 lead ECG cables
4.8 | A full set adult 3/5 lead ECG cables
4.9 | Reusable paediatric cuffs
4.10 | Reusable adult cuffs
4.11 | Reusable neonate cuffs
4.12 | Reusable obese cuffs
4.13 | Air hoses
4.14 | Finger sensor for paeds
4.15 | Finger sensor for adults
4.16 | Earlobe sensor for paeds
4.17 | Earlobe sensor for adults
4.18 | Finger wrap-around probe for neonates
4.19 | Sensor lead
4.20 | Two reusable temperature probes
4.21 | State the cost and frequency of replacement for the flow sensors.
4.22 | State the cost and frequency of replacement for the fuel cell.
4.23 | Vaporizer key-fillers (Sevo & Iso)
4.24 | Sevoflurane vaporizer
4.25 | Isoflurane vaporizer
4.26 | Anaesthesia circuit (Adult, single/double limb design) Disposable. 1.8m (Complete with Y-piece, Elbow, Catheter mount, Rebreathing bag)
4.27 | Anaesthesia circuit (Paediatric, single/double limb design) Disposable. 1.8m (Complete with Y-piece, Elbow, Catheter mount, Rebreathing bag)
4.28 | Anaesthesia circuit (neonatal, single/double limb design) Disposable. 1.8m (Complete with Y-piece, Elbow, Catheter mount, Rebreathing bag)
4.29 | Re-breathing circuit complete
4.30 | Non-re-breathing circuit complete
4.31 | Flow sensor (pack of 2)
4.32 | Spirometry line
4.33 | Sample line (pack of 10)
4.34 | Face masks: sizes 00, 0, 1, 2, 3, 4, 5.
4.35 | Bacterial/viral filter
4.36 | Y connector
4.37 | Manual ventilation hose
4.38 | Fresh gas tubing
4.39 | Seal for back-up cylinder connection on the unit
4.40 | Scavenger tubing (as per institution's scavenging outlet)
4.41 | Inspiratory and expiratory unidirectional valves
4.42 | Side stream CO2 kit (adult/paed)
4.43 | Disposable Adult Spirometry Set complete with spirometry tube, sample line and flow sensor, 3 m, box of 20 pieces
4.44 | Disposable Paediatric Spirometry Set complete with spirometry tube, sample line and flow sensor, 3 m, box of 20 pieces
4.45 | Gas sample tube with luer lock connection, 3m, disposable (10 pcs)
4.46 | Suction Tubing Set inc connector
4.47 | Resuscitator Adult, 1500mL, autoclavable, silicone
4.48 | Resuscitator Paediatric, 550mL, autoclavable, silicone
4.49 | Tubing set Paediatric for circle systems
5 | Conformity Compliance (Please attach certificate)
5.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 and 601-1-2 for medical equipment, attached certification
5.2 | The unit must comply with electrical safety and performance standard such as IEC 60601-2-13 for Anaesthetic systems, attached certification
5.2.1 | Please provide unique ref number of the IEC ceriticate:
5.4 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
5.4.1 | Please provide unique ref number of the ISO 13485 ceriticate:
5.5 | Model on offer must comply with EU Medical Device Regulation (EU) 2017/745 or equvalent from SAHPRA recognized jurisdictions . Attach a copy of certification
5.5.1 | Please provide unique ref number of the ceriticate:
5.6 | Hazardous Substance Act: If this type of equipment / apparatus appears on the schedule of Hazardous substances issued by the Directorate: Health Technology of the Department of Health, a license in terms of the Act on Hazardous Substances (Act 15/1973) must be submitted with the bid / quotation documents. The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.7 | All electrical/electronic medical equipment must be licensed by Radiation Control, where applicable a copy of the license must be submitted The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.7.1 | Radiation Control License number:
5.8 | The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade (rubber to with 2 screws) type and it must be a minimum length of three (3) meters. NB the mains cable of the unit tendered for must be SABS colour coded.
5.9 | The equipment quoted for must be protected against electromagnetic interference, IEC 60601-1-2
5.10 | The unit must comply with electrical safety and performance standard such as IEC 60601-2-49 for Patient Monitoring, attached certification
5.10.1 | Please provide unique ref number of the IEC ceriticate:
6 | COMPREHENSIVE MAINTENANCE
6.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) should be quoted for, as per the SCC document
6.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries, Flow sensors and Oxygen sensor), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
6.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
6.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries, Flow sensors and Oxygen sensor), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
6.3 | The bidder to provide the OEM recommended service interval of the unit:
6.4 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
7 | General
7.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
7.2 | The unit on offer must be the latest technology.
7.2.1 | Initial Manufacturing year of Model: 2022
ANAESTHESIA DELIVERY SYSTEM: Basic with integrated physiological monitor, Replacement Accessories
RT2-01-02a | A full set paediatric 3/5 lead ECG cable
RT2-01-02b | A full set adult 3/5 lead ECG cable
RT2-01-02c | Reusable paediatric cuff
RT2-01-02d | Reusable adult cuff
RT2-01-02e | Reusable neonate cuff
RT2-01-02f | Reusable obese cuff
RT2-01-02g | Air hose
RT2-01-02h | Finger sensor for paeds
RT2-01-02i | Finger sensor for adults
RT2-01-02j | Earlobe sensor for paeds
RT2-01-02k | Earlobe sensor for adults
RT2-01-02l | Finger wrap-around probe for neonates
RT2-01-02m | Sensor lead
RT2-01-02n | Reusable temperature probe

## Anaesthetic, Standard
RT2-01-03 | ANAESTHESIA DELIVERY SYSTEM: Standard without physiological monitor | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
1 | Description
1.1 | A medical device designed to administer anesthetics to patients during surgical procedures.
2 | Product Offered (Anaesthetic Machine)
2.1 | Make:
2.2 | Model:
2.3 | Initial Manufacturing year of Model:
2.4 | Radiation Control License number:
3 | Technical Specification
3.1 | Anaesthesia delivery system suitable for neonatal, paediatric and adult anaesthesia
3.2 | Ventilator suitable for ventilating neonatal, paediatrics to adult patients
3.3 | Multigas analyser:
3.3.1 | Oxygen
3.3.2 | Nitrous oxide
3.3.3 | Carbon dioxide
3.4 | Anaesthetic agent monitoring:
3.4.1 | Isofluorane
3.4.2 | Sevofluorane
3.4.3 | Desfluorane
3.5 | The monitor must be able to simultaneously display the inspiratory and expiratory O2, CO2, N2O and Anaesthetic agent concentrations.
3.6 | The monitor must be able to identify the Anaesthetic agent automatically. (During a change between agents it is very desirable that the monitor must be able to measure the agents simultaneously with an alert message. i.e.: wash in and wash out values of both agents)
3.7 | Ketone, methane & Alcohol concentrations should not influence the measurements of the anaesthetic gasses analyzed
3.8 | Gas analyzer Mac point should be age corrected. The monitor should be able to display measured MAC values
3.9 | Self test requiring limited user intervention.
3.10 | The self test must provide the measurements and values for automatic leak detection and compliance.
3.11 | Automatic calibration of sensors for gases and anaesthetic vapours must form part of the self test and require limited user intervention.
3.12 | The unit must provide conventional, low-flow (down to 200 ml/min) anaesthesia.
3.13 | It must be possible to override the self test facility. The system must display the override status
3.14 | The unit must return analysed gas back to the circle system
3.15 | The unit must be supplied with a minimum of 2 storage drawer
3.16 | The unit must be supplied with a minimum of 4 additional 220 V power sockets. These sockets must be excluded from the battery backup circuit in case of a power failure
3.17 | The unit must have a working area or a retractable shelf. Width: not less than 300 mm. Depth: not less than 200 mm
3.18 | The unit must have a light to enable the anaesthetist to work in a darkened theatre.
3.19 | Connectivity with automated record system should be possible (Optional)
3.20 | Trolley
3.20.1 | The unit must be mounted on a strong , robust and compact trolley manufactured from corrosion free material.
3.20.2 | The trolley must have no sharp corners and be easy to clean.
3.20.3 | The unit shall be provided with at least four castors that comply with SANS 621 detailing the castor's material type, size and movement. The castors shall be equipped with a braking system
3.20.4 | Trolley must have two lockable castors or a central locking brake.
3.20.5 | The unit must accommodate two(2) gas cylinders of the appropriate size for use. The size should be compatible with current state vendor of oxygen starting from 0.94 to 1.84kg.
3.20.6 | Trolley must have accessible handles for ease of movement
3.20.7 | The trolley, with all components mounted, must be stable without toppling over when placed on a 30° slope.
3.20.8 | The unit shall be supplied with a castor guard system, preventing the entanglement and trapping of cables within the wheels.
3.20.9 | The trolley must have a mounting system suitable for a physiologial monitor.
3.20.10 | The unit should be ergonomically designed to ensure maximum ease and safety when used.
3.20.11 | All elements of the system shall be software driven allowing easy future upgrades. All parts of the aesthesia machine: gas mixer, vaporizer and ventilator shall preferably be software driven
3.21 | Pipeline and Cylinder Gas Supplies
3.21.1 | The anaesthesia system must be designed to use external piped medical gasses (O2, N2O and medical air) and vacuum from medical service outlet sockets.
3.21.2 | Connection hoses for above must be provided
3.21.3 | Hoses must be fitted with Heyer type non-interchangeable probes for the existing medical gas outlets as specified by the purchaser.
3.21.4 | The medical air hose connection or CO2 circle system must include a water trap
3.21.5 | A safety pressure relief valve of 60 - 80 cm H2O must be provided to protect the system from excessive pressure .
3.21.6 | The unit must function safely and correctly within the pipeline supply pressure range of: 270 - 600 kPa
3.21.7 | The system must preferentially utilise the piped oxygen supply over the cylinder supply, if pipeline pressure is > 270 kPa
3.21.8 | Connection hoses must be:
3.21.8.1 | Flexible and at least 6 metres long
3.21.8.2 | Antistatic
3.21.9 | Colour coded as follows:
3.21.9.1 | O2: white
3.21.9.2 | N2O: blue
3.21.9.3 | Vacuum: yellow
3.21.9.4 | Medical air: black
3.21.9.5 | Scavenge low vacuum: orange or yellow
3.21.9 | Where relevant, fitted with outlet probes complying with SANS 1409:2014
3.21.10 | Permanantly bonded to the outlet probe and the machine connector
3.21.11 | Hoses for gases above atmospheric pressure, must designed for pressures of at least 1 000 kPa.
3.21.12 | Vacuum hoses may not collapse or kink at 90 kPa below atmospheric pressure, when subjected to 180° bend of inside radius equal to the outside diameter of the hose.
3.21.13 | All connection hoses must be fed to a central manifold
3.21.14 | All high pressure gas and vacuum pipe connections (other than hoses) and all pipes connected to pressure regulators and pressure gauges must be stainless steel (18/8 grade) pipes that are oil-free and colour coded as for hoses (see above) over their entire lengths
3.21.15 | An in-line non-return valve must be fitted to each gas pipeline on the high pressure (inlet) line
3.21.16 | An in-line non-return valve must be fitted to each gas cylinder yoke
3.21.17 | A replaceable gas filter, to block particles greater than 65 µm, must be fitted on the upstream side of the non-return valve
3.21.18 | Each gas pipeline inlet must have a pressure gauge or electronic display of pressure. The pressure range (415 kPa ± 75 kPa) must be clearly indicated in colour
3.21.19 | Must be fitted with a gas cylinder yoke to accommodate one oxygen gas cylinder with a pin index system
3.21.20 | Each gas cylinder yoke must have a pressure gauge or electronic display of pressure
3.21.21 | The pressure gauges or electronic displays of pressure for both pipeline & cylinder supplies, calibrated in kPa for each line, must be labelled and visible from the front of the unit.
3.21.22 | O2 and N2O regulators must be provided complying with CKS 605:1987.
3.21.23 | A pressure regulator complying with CKS 605 must be supplied for each gas inlet and must reduce the gas pressure to 415 kPa without freezing up, creeping, chattering or whistling
3.21.24 | Outlet sockets and probes must comply with SANS 1409:2014.
3.21.25 | Sockets with hoses must be colour-coded as per SANS 1409:2014:
3.21.26 | Oxygen: White background with black letters
3.21.27 | Nitrous oxide: Blue background with white letters
3.21.28 | Medical air: One half black with white letters, the other half white with black letters
3.21.29 | Vacuum: Yellow background with black letters
3.21.30 | Continuous display, visible from the front of the machine, of both pipeline and cylinder pressures even during power failure
3.21.31 | Pipeline and cylinder pressures must be differentially displayed, and clearly indicated
3.21.32 | Safety pressure relief valves must protect the entire pneumatic system from excessive input pressures from pipeline and cylinder supplies.
3.21.33 | When a cylinder is detached from the pin-indexed yoke no gas leak must occur at the yoke.
3.21.34 | Airway pressure must be monitored and displayed
3.21.35 | Safety relief valve for all gases must be supplied and set at 600kPa
3.22 | Scavenging
3.22.1 | The unit must have a scavenging system with a hose at least 5 metres long with the suitable connector as per the institutions scavenging installation.
3.22.2 | Air-break device to prevent negative pressure on the breathing system
3.22.3 | If an active sytem is supplied a pressure indicator must be present
3.22.4 | The scavenging outlet must be a 30mm ISO connector.
3.23 | Patient Gas Delivery System (Flowmeters/Rotameters)
3.23.1 | The unit must have cascading flowmeters (mechanical or electronic) that show flow rates for high and low flow anaesthesia in the following ranges:
3.23.1.1 | Oxygen low flow: 200 - 1 000 ml/min
3.23.1.2 | At higher flow rates the gas must cascade into a second flowmeter (1 – 10 l/min)
3.23.1.3 | Nitrous oxide low flow: 200 - 1 000 ml/min
3.23.1.4 | At higher flow rates the gas must cascade into a second flowmeter (1 – 10 l/min)
3.23.1.5 | Air tube 200ml/min – 10 l/min
3.23.1.6 | Flowmeters, if not electronic, must be diameter indexed to prevent incorrect replacement during maintenance.
3.23.2 | Gas supply must conform to:
3.23.2.1 | O2 minimum flow range of 200 ml/min to 10 l/min. Accuracy ± 5%
3.23.2.2 | medical air minimum flow range 0 to 10 l/min. Accuracy ± 5%
3.23.2.3 | nitrous oxide minimum flow range 0 to 10 l/min. Accuracy ± 5%
3.23.3 | Flowmeters must be accurate to at least 5%
3.23.4 | Flowmeter valve control
3.23.5 | Flowmeter control knobs
3.23.6 | Flowmeter sequence to be according to standards approved by the South Africal Society of Anaesthesiologists - Oxygen on the right
3.23.7 | The delivery system must prevent simultaneous delivery of two anaesthetic vapours.
3.23.8 | A separate common gas outlet to accept a standard 22 mm outside diameter and 15 mm internal diameter.
3.23.9 | A separate oxygen gas outlet for connection to and oxygen mask/Ambu bag tubing
3.24 | Hypoxic guard
3.24.1 | The unit must include a safety mechanism that ensures a minimum concentration of 25% oxygen in any oxygen/nitrous oxide gas mixture across all flow rates.
3.24.2 | The unit must include a safety mechanism that ensures a minimum concentration of 21% oxygen in any oxygen/air gas mixture across all flow rates.
3.24.3 | The unit must include a high-flow oxygen flush mechanism with a flow rate of at least 35 l/min
3.24.4 | Oxygen flush must be non-lockable, recessed and self closing
3.24.5 | This system must bypass the flowmeters and vaporisers
3.25 | Breathing System
3.25.1 | Total circuit volume, including absorber, must not exceed 2,7 litres
3.25.2 | With APL valve 5 - 75 cm H2O
3.25.3 | With ventilator-manual switch. This switch when directed to the ventilator must isolate the APL valve from the circle
3.25.4 | Breathing system must be fully autoclavable to 134°C
3.25.5 | There must be a mechanism to remove water from the breathing system
3.25.6 | Breathing system must be latex free
3.25.7 | If unit is using bellows, they must be rising and visible
3.25.8 | The gas machine must have 15/22 fresh gas outlet and thus accept all commonly used breathing systems. It shall be possible to use Bain and Jackson Rees systems to allow easy spontaneous breathing and fast response times in paediatric anaesthesia
3.25.9 | The absorber canister should be exchangeable during machine use without loss of pressure and circuit gas
3.26 | Integrated Ventilator
3.26.1 | Ventilator suitable for ventilating adults, paediatric and neonatal patients without changing ventilator components (e.g. bellows)
3.26.2 | The ventilator must be electronically controlled
3.26.3 | The ventilation technology offered may be turbine, volume reflector, piston or ascending bellows
3.26.4 | The minimum basic modes to be supplied:
3.26.4.1 | Manual/Spontaneous Ventilation
3.26.4.2 | Volume Control Ventilation
3.26.4.3 | Pressure Control Ventilation
3.26.4.4 | SIMV
3.26.5 | Additional modes of ventilation may be offered and quoted for separately
3.26.6 | The ventilator must be integrated with the workstation and user friendly
3.26.7 | Frequency: minimum range of 5 – 60 bpm.
3.26.8 | Inspiratory time: minimum range of 0.2 – 10.0 seconds.
3.26.9 | Minimum Inspiration/Expiration Ratio: 4:1 – 1:4.
3.26.10 | Adjustable inspiratory pause/plateau time (volume control ventilation): 0-30% of inspiratory time
3.26.11 | Peak Inspiratory Pressure: minimum range of: PEEP + 5 to 70 cm H2O.
3.26.12 | Positive End Expiratory Pressure (PEEP): Integrated, electronically controlled. Minimum range: 0 – 30 cm H2O.
3.26.13 | Tidal Volume: 20ml – 1500ml per breath. Minimum accuracy: 15%
3.26.14 | Peak Inspiratory Flow: At least 120 lpm
3.26.15 | The ventilator must be capable of flow or pressure triggering.
3.26.16 | Adjustable Inspiratory Rise time/Ramp with a range of 0.1 - 2 sec
3.26.17 | Compliance compensation/Tidal volume compensation/Compressible volume compensation or equivalent.
3.26.18 | Patient gas flow compensation must be automatic so that alterations in patient gas flow do not affect delivered tidal volumes
3.26.19 | In event of power or technical failure, manual ventilation must be possible.
3.26.20 | The ventilator must be compatible with a circle system and open system
3.27 | Display:
3.27.1 | All displays must be in acceptable English
3.27.2 | Colour display must include numerical and graphic display of parameters measured and set ventilator settings
3.27.3 | Graphic display to feature at least 3 waveforms: - pressure, flow and volume.
3.27.4 | Must also display flow-volume and pressure-volume loops with a reference loop
3.27.5 | The flow sensor has to be service free
3.27.6 | Alarms : The ventilator must have the following user adjustable alarms
3.27.6.1 | Apnoea alarm
3.27.6.2 | High airway pressure alarm
3.27.6.3 | Continuous pressure alarm
3.27.6.4 | High and low expiratory tidal volume alarm
3.27.6.5 | Minute volume, high/low alarm
3.27.6.6 | Inspired oxygen (Fi02), high/low alarm
3.27.6.7 | PEEP, low alarm
3.27.6.8 | Leakage / Check breathing circuit alarm or equivalent
3.27.6.9 | Alarm profile must be user definable or default to patient appropriate settings
3.27.6.10 | The ventilator must have a mute alarm function, active only for a limited time(no more than 2 minutes).
3.27.6.11 | It must not be possible to cancel the visual alarm until the problem has been rectified.
3.27.7 | All components of the ventilator and breathing system in contact with the patient’s airway must be latex free.
3.27.8 | All components of the ventilator and breathing system in contact with the patient’s airway must be steam autoclavable up to 134ᴼ C
3.28 | Integrated Gas Monitoring
3.28.1 | Oxygen Monitoring
3.28.1.1 | The system must display both inspired and end-tidal oxygen concentration.
3.28.1.2 | The oxygen analyser must have user adjustable low and high alarms – visual and audible. The low alarm must automatically default to no less than 21%.
3.28.1.3 | The system must be accurate (+/- 2% in 21-100%) range, and have a rapid response time (≤ 500ms).
3.28.1.4 | The unit must sound a continuous alarm if the oxygen pressure falls below 190 kPa, then all nitrous oxide must be automatically cut from the patient circuit.
3.28.1.5 | Analyser preferrably of paramagnetic type. If fuel cell offered, provide full information on cost of disposables and frequency of replacement
3.28.2 | Circle System
3.28.2.1 | The circle absorber system must be integrated as part of the ventilation system.
3.28.2.2 | The circle absorber system must have an effective system to eliminate excess water created by the chemical reactions either by heating or water traps.
3.28.2.3 | The circle absorber system must be autoclavable at 134ºC.
3.28.2.4 | The circle absorber system must be supplied with a carbon dioxide absorber canister.
3.28.2.5 | The absorber canister should be reusable/refillable. Disposable cannisters should be quoted separately
3.28.2.6 | It must be easy to remove and replace the absorber canister while the workstation is in operation.
3.28.2.7 | Misalignment of the canister must be readily apparent to the user.
3.28.2.8 | The system must have a circle bypass facility while changing/refilling the absorber canister.
3.28.2.9 | The breathing circuit must contain an integrated adjustable pressure limiting (APL) valve with a minimum range of 0-70 cm H2O.
3.28.2.10 | The APL valve must have a tactile or electronic indicator when a value of between 20-30 cm H2O is exceeded.
3.28.2.11 | The circle absorber system must be latex free.
3.28.2.12 | The bag-to-ventilator switching mechanism must be clearly labeled and user-friendly. Mechanical or electronic
3.29 | Patient circuit pressure release valve
3.29.1 | Adjustable up to 70 cm water
3.29.2 | Audible alarm when the valve opens
3.29.3 | Manual override, non-lockable, to keep the release valve closed when the set pressure is exceeded.
3.30 | Flowmeter pressure release valve
3.30.1 | Situated between the flowmeters and the vaporisers
3.30.2 | Must open at 105 kPa +- 5 kPa
3.31 | Other System requirements
3.31.1 | The unit must operate from both the 220 V, 50 Hz mains supply and internal rechargeable batteries
3.31.2 | The batteries must be charged while the unit is operating from the mains supply.
3.31.3 | The batteries must automatically take over in the event of a mains power failure.
3.31.4 | The batteries must have the capacity for full operation of the anaesthetic machine and ventilator for at least 90 minutes with batteries fully charged, under typical operating conditions.
3.31.5 | Graphic battery power indicator
3.31.6 | Recharge time < 8 hrs
3.31.7 | Mains cable at least four (4) metres long with SABS approved 15A three prong plug.
3.31.8 | The unit must have a rail for attaching a suction regulator.
3.31.9 | The unit must be handed over in full operating order
4 | The following Accessories and Consumables must be supplied with the unit:
4.1 | Re-breathing bag, reusable, rubber
4.2 | Reservoir bag
4.3 | Pressure limiting valve, adjustable
4.4 | Oxygen fuel cell analyser
4.5 | Flow sensors.
4.6 | Fuel cell.
4.7 | Vaporizer key-fillers (Sevo, & Iso)
4.8 | The unit must be supplied with two vaporizers (one must sevoflurane, the other must be of the choice of the buying institution)
4.9 | Anaesthesia circuit (Adult, single/double limb design) Disposable. 1.8m (Complete with Y-piece, Elbow, Catheter mount, Rebreathing bag)
4.10 | Anaesthesia circuit (Paediatric, single/double limb design) Disposable. 1.8m (Complete with Y-piece, Elbow, Catheter mount, Rebreathing bag)
4.11 | Anaesthesia circuit (neonatal, single/double limb design) Disposable. 1.8m (Complete with Y-piece, Elbow, Catheter mount, Rebreathing bag)
4.12 | Re-breathing circuit complete
4.13 | Non-re-breathing circuit complete
4.14 | Maintenance-free Flow sensor (pack of 2)
4.15 | Spirometry line
4.16 | Sample line (pack of 10)
4.17 | Face masks: sizes 00, 0, 1, 2, 3, 4, 5.
4.18 | Bacterial/viral filter
4.19 | Y connector
4.20 | Manual ventilation hose
4.21 | Fresh gas tubing
4.22 | Seal for back-up cylinder connection on the unit
4.23 | Scavenger tubing
4.24 | Inspiratory and expiratory unidirectional valves
4.25 | Side stream CO2 kit (adult/paed)
4.26 | Disposable Adult Spirometry Set complete with spirometry tube, sample line and flow sensor, 3 m, box of 20 pieces
4.27 | Disposable Paediatric Spirometry Set complete with spirometry tube, sample line and flow sensor, 3 m, box of 20 pieces
4.28 | Gas sample tube with luer lock connection, 3m, disposable (10 pcs)
4.29 | Suction Tubing Set inc connector
4.30 | Resuscitator Adult, 1500mL, autoclavable, silicone
4.31 | Resuscitator Paediatric, 550mL, autoclavable, silicone
4.32 | Tubing set Paediatric for circle systems
5 | OPTIONAL ITEMS
5.1 | Prices incl VAT must be given for all optional items
5.2 | Brochures with full technical specifications must be supplied for each optional item offered
6 | Conformity Compliance (Please attach certificate)
6.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
6.2 | The unit must comply with electrical safety and performance standard such as IEC 60601-2-13 for Anaesthetic systems, attached certification
6.2.1 | Please provide unique ref number of the IEC ceriticate:
6.4 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
6.4.1 | Please provide unique ref number of the ISO 13485 ceriticate:
6.5 | Model on offer must comply with EU Medical Device Regulation (EU) 2017/745 or equvalent from SAHPRA recognized jurisdictions . Attach a copy of certification
6.5.1 | Please provide unique ref number of the ceriticate:
6.6 | Hazardous Substance Act: If this type of equipment / apparatus appears on the schedule of Hazardous substances issued by the Directorate: Health Technology of the Department of Health, a license in terms of the Act on Hazardous Substances (Act 15/1973) must be submitted with the bid / quotation documents. The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
6.7 | All electrical/electronic medical equipment must be licensed by Radiation Control, where applicable a copy of the license must be submitted The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
6.7.1 | Radiation Control License number:
6.8 | The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade (rubber to with 2 screws) type and it must be a minimum length of three (3) meters. NB the mains cable of the unit tendered for must be SABS colour coded.
6.9 | The equipment quoted for must be protected against electromagnetic interference IEC 60601-1-2
7 | COMPREHENSIVE MAINTENANCE
7.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) should be quoted for, as per the SCC document
7.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries, Flow sensors and Oxygen sensor), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
7.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
7.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries, Flow sensors and Oxygen sensor), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
7.3 | The bidder to provide the OEM recommended service interval of the unit:
7,4 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
8 | General
8.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
8.2 | The unit on offer must be the latest technology.
8.2.1 | Initial Manufacturing year of Model:
| ANAESTHESIA DELIVERY SYSTEM: STANDARD WITHOUT PHYSIOLOGICAL MONITOR, REPLACEMENT ACCESSORIES | Complies Yes/No | RT13 items to be removed
RT2-01-03a | Anaesthesia circuit (Adult, single/double limb design) Disposable. 1.8m (Complete with Y-piece, Elbow, Catheter mount, Rebreathing bag)
RT2-01-03b | Anaesthesia circuit (Paediatric, single/double limb design) Disposable. 1.8m (Complete with Y-piece, Elbow, Catheter mount, Rebreathing bag)
RT2-01-03c | Anaesthesia circuit (neonatal, single/double limb design) Disposable. 1.8m (Complete with Y-piece, Elbow, Catheter mount, Rebreathing bag)
RT2-01-03d | Re-breathing circuit complete
RT2-01-03e | Non-re-breathing circuit complete
RT2-01-03f | Flow sensor (pack of 2)
RT2-01-03g | Spirometry line
RT2-01-03h | Sample line (pack of 10)
RT2-01-03i-1 | Face masks: sizes 00
RT2-01-03i-2 | Face masks: sizes 0
RT2-01-03i-3 | Face masks: sizes 1
RT2-01-03i-4 | Face masks: sizes 2
RT2-01-03i-5 | Face masks: sizes 3
RT2-01-03i-6 | Face masks: sizes 4
RT2-01-03i-7 | Face masks: sizes 5
RT2-01-03j | Bacterial/viral filter
RT2-01-03k | Y connector
RT2-01-03l | Manual ventilation hose
RT2-01-03m | Fresh gas tubing
RT2-01-03n | Seal for back-up cylinder connection on the unit
RT2-01-03o | Scavenger tubing (as per institution's scavenging outlet)
RT2-01-03p | Inspiratory and expiratory unidirectional valves
RT2-01-03q | Side stream CO2 kit (adult/paed)
RT2-01-03r | Disposable Adult Spirometry Set complete with spirometry tube, sample line and flow sensor, 3 m, box of 20 pieces
RT2-01-03s | Disposable Paediatric Spirometry Set complete with spirometry tube, sample line and flow sensor, 3 m, box of 20 pieces
RT2-01-03t | Gas sample tube with luer lock connection, 3m, disposable (10 pcs)
RT2-01-03u | Suction Tubing Set inc connector
RT2-01-03v | Resuscitator Adult, 1500mL, autoclavable, silicone
RT2-01-03w | Resuscitator Paediatric, 550mL, autoclavable, silicone
RT2-01-03x | Tubing set Paediatric for circle systems
RT2-01-03y | sevoflurane vaporizer
RT2-01-03z | Isoflurane vaporizer
RT2-01-03zz | Desflurane vaporizer

## Anaesthetic, Standard + Monitor
RT2-01-04 | ANAESTHESIA DELIVERY SYSTEM: Standard with Integrated Physiological Monitor | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
1 | Description
1.1 | A medical device designed to administer anesthetics to patients during surgical procedures.
2a | Product Offered (Anaesthetic Machine)
2a.1 | Make:
2a.2 | Model:
2a.3 | Initial Manufacturing year of Model:
2a.4 | Radiation Control License number:
2b | Product Offered (Monitor)
2b.1 | Make:
2b.2 | Model:
2b.3 | Initial Manufacturing year of Model:
2b.4 | Radiation Control License number:
3 | Technical Specification
3.1 | Anaesthesia delivery system suitable for neonatal, paediatric and adult anaesthesia
3.2 | Ventilator suitable for ventilating neonatal to adult patients
3.3 | Multigas analyser:
3.3.1 | Oxygen
3.3.2 | Nitrous oxide
3.3.3 | Carbon dioxide
3.4 | Anaesthetic agent monitoring:
3.4.1 | Isofluorane
3.4.2 | Sevofluorane
3.4.3 | Desfluorane
3.5 | The monitor must be able to simultaneously display the inspiratory and expiratory O2, CO2, N2O and Anaesthetic agent concentrations.
3.6 | The monitor must be able to identify the Anaesthetic agent automatically. (During a change between agents it is very desirable that the monitor must be able to measure the agents simultaneously with an alert message. i.e.: wash in and wash out values of both agents)
3.7 | Ketone, methane & Alcohol concentrations should not influence the measurements of the anaesthetic gasses analyzed
3.8 | Gas analyzer Mac point should be age corrected. The monitor should be able to display measured MAC values
3.9 | Self test requiring limited user intervention.
3.10 | The self test must provide the measurements and values for automatic leak detection and compliance.
3.11 | Automatic calibration of sensors for gases and anaesthetic vapours must form part of the self test and require very limited user intervention.
3.12 | The unit must provide conventional, low-flow (down to 200 ml/min) anaesthesia.
3.13 | It must be possible to override the self test facility. The system must display the override status
3.14 | The unit must return analysed gas back to the circle system
3.15 | The unit must be supplied with a minimum of 2 storage drawers
3.16 | The unit must be supplied with a minimum of 4 additional 220 V power sockets. These sockets must be excluded from the battery backup circuit in case of a power failure
3.17 | The unit must have a working area or a retractable shelf. Width: not less than 300 mm. Depth: not less than 200 mm
3.18 | The unit must have a light to enable the anaesthetist to work in a darkened theatre.
3.20 | Trolley
3.20.1 | The unit must be mounted on a strong , robust and compact trolley manufactured from corrosion free material.
3.20.2 | The trolley must have no sharp corners and be easy to clean.
3.20.3 | The trolley must have at least four antistatic castors of diameter 100 mm or greater.
3.20.4 | Trolley must have two lockable castors or a central locking brake.
3.20.5 | The unit must accommodate an oxygen cylinder. Cylinder size according to user requirement. Size to be compatible with current state vendor of oxygen.
3.20.6 | Trolley must have accessible handles for ease of movement
3.20.7 | The trolley, with all components mounted, must be stable without toppling over when placed on a 30° slope.
3.20.8 | The unit must be supplied with a castor guard system, preventing the entaglement and trapping of cables within the wheels.
3.20.9 | The trolley must have a mounting system suitable for a physiologial monitor.
3.20.10 | The unit should be ergonomically designed to ensure maximum ease and safety when used.
3.20.11 | All elements of the system shall be software driven allowing easy future upgrades. All parts of the aesthesia machine: gas mixer, vaporizer and ventilator shall preferably be software driven
3.21 | Pipeline and Cylinder Gas Supplies
3.21.1 | The anaesthesia system must be designed to use external piped medical gasses (O2, N2O and medical air) and vacuum from medical service outlet sockets.
3.21.2 | Connection hoses for above must be provided
3.21.3 | Hoses must be fitted with Heyer type non-interchangeable probes for the existing medical gas outlets as specified by the purchaser.
3.21.4 | The medical air hose connection or CO2 circle system must include a water trap
3.21.5 | A safety pressure relief valve of 60 - 80 cm H2O must be provided to protect the system from excessive pressure .
3.21.6 | The unit must function safely and correctly within the pipeline supply pressure range of: 270 - 600 kPa
3.21.7 | The system must preferentially utilise the piped oxygen supply over the cylinder supply, if pipeline pressure is > 270 kPa
3.21.8 | Connection hoses must be:
3.21.8.1 | Flexible and at least 6 metres long
3.21.8.2 | Antistatic
3.21.9 | Colour coded as follows:
3.21.9.1 | O2: white
3.21.9.2 | N2O: blue
3.21.9.3 | Vacuum: yellow
3.21.9.4 | Medical air: black
3.21.9.5 | Scavenge low vacuum: orange or yellow
3.21.9 | Where relevant, fitted with outlet probes complying with SANS 1409:2014
3.21.10 | Permanantly bonded to the outlet probe and the machine connector
3.21.11 | Hoses for gases above atmospheric pressure, must designed for pressures of at least 1 000 kPa.
3.21.12 | Vacuum hoses may not collapse or kink at 90 kPa below atmospheric pressure, when subjected to 180° bend of inside radius equal to the outside diameter of the hose.
3.21.13 | All connection hoses must be fed to a central manifold
3.21.14 | All high pressure gas and vacuum pipe connections (other than hoses) and all pipes connected to pressure regulators and pressure gauges must be stainless steel (18/8 grade) pipes that are oil-free and colour coded as for hoses (see above) over their entire lengths
3.21.15 | An in-line non-return valve must be fitted to each gas pipeline on the high pressure (inlet) line
3.21.16 | An in-line non-return valve must be fitted to each gas cylinder yoke
3.21.17 | A replaceable gas filter, to block particles greater than 65 µm, must be fitted on the upstream side of the non-return valve
3.21.18 | Each gas pipeline inlet must have a pressure gauge or electronic display of pressure. The pressure range (415 kPa ± 75 kPa) must be clearly indicated in colour
3.21.19 | Must be fitted with a gas cylinder yoke to accommodate one oxygen gas cylinder with a pin index system
3.21.20 | Each gas cylinder yoke must have a pressure gauge or electronic display of pressure
3.21.21 | The pressure gauges or electronic displays of pressure for both pipeline & cylinder supplies, calibrated in kPa for each line, must be labelled and visible from the front of the unit.
3.21.22 | O2 and N2O regulators must be provided complying with CKS 605:1987.
3.21.23 | A pressure regulator complying with CKS 605 must be supplied for each gas inlet and must reduce the gas pressure to 415 kPa without freezing up, creeping, chattering or whistling
3.21.24 | Outlet sockets and probes must comply with SANS 1409:2014.
3.21.25 | Sockets with hoses must be colour-coded as per SANS 1409:2014:
3.21.26 | Oxygen: White background with black letters
3.21.27 | Nitrous oxide: Blue background with white letters
3.21.28 | Medical air: One half black with white letters, the other half white with black letters
3.21.29 | Vacuum: Yellow background with black letters
3.21.30 | Continuous display, visible from the front of the machine, of both pipeline and cylinder pressures even during power failure
3.21.31 | Pipeline and cylinder pressures must be differentially displayed, and clearly indicated
3.21.32 | Safety pressure relief valves must protect the entire pneumatic system from excessive input pressures from pipeline and cylinder supplies.
3.21.33 | When a cylinder is detached from the pin-indexed yoke no gas leak must occur at the yoke.
3.21.34 | Airway pressure must be monitored and displayed
3.21.35 | Safety relief valve for all gases must be supplied and set at 600kPa
3.22 | Scavenging
3.22.1 | The unit must have a scavenging system with a hose at least 5 metres long with the suitable connector as per the institutions scavenging installation.
3.22.2 | Air-break device to prevent negative pressure on the breathing system
3.22.3 | If an active sytem is supplied a pressure indicator must be present
3.22.4 | The scavenging outlet must be a 30mm ISO connector.
3.23 | Patient Gas Delivery System (Flowmeters/Rotameters)
3.23.1 | The unit must have cascading flowmeters (mechanical or electronic) that show flow rates for high and low flow anaesthesia in the following ranges:
3.23.1.1 | Oxygen low flow: 200 - 1 000 ml/min
3.23.1.2 | At higher flow rates the gas must cascade into a second flowmeter (1 – 10 l/min)
3.23.1.3 | Nitrous oxide low flow: 200 - 1 000 ml/min
3.23.1.4 | At higher flow rates the gas must cascade into a second flowmeter (1 – 10 l/min)
3.23.1.5 | Air tube 200ml/min – 10 l/min
3.23.1.6 | Flowmeters, if not electronic, must be diameter indexed to prevent incorrect replacement during maintenance.
3.23.2 | Gas supply must conform to:
3.23.2.1 | O2 minimum flow range of 200 ml/min to 10 l/min. Accuracy ± 5%
3.23.2.2 | medical air minimum flow range 0 to 10 l/min. Accuracy ± 5%
3.23.2.3 | nitrous oxide minimum flow range 0 to 10 l/min. Accuracy ± 5%
3.23.3 | Flowmeters must be accurate to at least 5%
3.23.4 | Flowmeter valve control
3.23.5 | Flowmeter control knobs
3.23.6 | Flowmeter sequence to be according to standards approved by the South Africal Society of Anaesthesiologists - Oxygen on the right
3.23.7 | The delivery system must prevent simultaneous delivery of two anaesthetic vapours.
3.23.8 | A separate common gas outlet to accept a standard 22 mm outside diameter and 15 mm internal diameter.
3.23.9 | A separate oxygen gas outlet for connection to and oxygen mask/Ambu bag tubing
3.24 | Hypoxic guard
3.24.1 | The unit must include a safety mechanism that ensures a minimum concentration of 25% oxygen in any oxygen/nitrous oxide gas mixture across all flow rates.
3.24.2 | The unit must include a safety mechanism that ensures a minimum concentration of 21% oxygen in any oxygen/air gas mixture across all flow rates.
3.24.3 | The unit must include a high-flow oxygen flush mechanism with a flow rate of at least 35 l/min
3.24.4 | Oxygen flush must be non-lockable, recessed and self closing
3.24.5 | This system must bypass the flowmeters and vaporisers
3.25 | Breathing System
3.25.1 | Total circuit volume, including absorber, must not exceed 2,7 litres
3.25.2 | With APL valve 5 - 75 cm H2O
3.25.3 | With ventilator to manual switch. This switch when directed to the ventilator must isolate the APL valve from the circle
3.25.4 | Breathing system must be fully autoclavable to 134°C
3.25.5 | There must be a mechanism to remove water from the breathing system
3.25.6 | Breathing system must be latex free
3.25.7 | Reusable corrugated breathing circuit adequate for both adult and paediatric use to be supplied
3.25.8 | Bidders to supply the breathing circuit as required by the user
3.25.9 | If unit is using bellows, they must be rising and visible
3.25.10 | The gas machine must have 15/22 fresh gas outlet and thus accept all commonly used breathing systems. It shall be possible to use Bain and Jackson Rees systems to allow easy spontaneous breathing and fast response times in paediatric anaesthesia
3.25.11 | The absorber canister should be exchangeable during machine use without loss of pressure and circuit gas
3.25.12 | The absorber canister should be exchangeable during machine use without loss of pressure and circuit gas
3.26 | Integrated Ventilator
3.26.1 | Ventilator suitable for ventilating adults, paediatric and neonatal patients without changing ventilator components (e.g. bellows)
3.26.2 | The ventilator must be electronically controlled
3.26.3 | The ventilation technology offered may be turbine, volume reflector, piston or ascending bellows
3.26.4 | The minimum basic modes to be supplied:
3.26.4.1 | Manual/Spontaneous Ventilation
3.26.4.2 | Volume Control Ventilation
3.26.4.3 | Pressure Control Ventilation
3.26.4.4 | SIMV
3.26.5 | Additional modes of ventilation may be offered and quoted for separately
3.26.6 | The ventilator must be integrated with the workstation and user friendly
3.26.7 | Frequency: minimum range of 5 – 60 bpm.
3.26.8 | Inspiratory time: minimum range of 0.2 – 10.0 seconds.
3.26.9 | Minimum Inspiration/Expiration Ratio: 4:1 – 1:4.
3.26.10 | Adjustable inspiratory pause/plateau time (volume control ventilation): 0-30% of inspiratory time
3.26.11 | Peak Inspiratory Pressure: minimum range of: PEEP + 5 to 70 cm H2O.
3.26.12 | Positive End Expiratory Pressure (PEEP): Integrated, electronically controlled. Minimum range: 0 – 30 cm H2O.
3.26.13 | Tidal Volume: 20ml – 1500ml per breath. Minimum accuracy: 15%
3.26.14 | Peak Inspiratory Flow: At least 120 lpm
3.26.15 | The ventilator must be capable of flow or pressure triggering.
3.26.16 | Adjustable Inspiratory Rise time/Ramp with a range of 0.1 - 2 sec
3.26.17 | Compliance compensation/Tidal volume compensation/Compressible volume compensation or equivalent.
3.26.18 | Patient gas flow compensation must be automatic so that alterations in patient gas flow do not affect delivered tidal volumes
3.26.19 | In event of power or technical failure, manual ventilation must be possible.
3.27 | Display:
3.27.1 | All displays must be in acceptable English
3.27.2 | Colour display must include numerical and graphic display of parameters measured
3.27.3 | Graphic display to feature at least 3 waveforms: - pressure, flow and volume.
3.27.4 | Must also display flow-volume and pressure-volume loops with a reference loop
3.27.5 | Alarms : The ventilator must have the following user adjustable alarms
3.27.6 | Apnoea alarm
3.27.6.1 | High airway pressure alarm
3.27.6.2 | Continuous pressure alarm
3.27.6.3 | High and low expiratory tidal volume alarm
3.27.6.4 | Minute volume, high/low alarm
3.27.6.5 | Inspired oxygen (Fi02), high/low alarm
3.27.6.6 | PEEP, low alarm
3.27.6.7 | Leakage / Check breathing circuit alarm or equivalent
3.27.6.8 | Alarm profile must be user definable or default to patient appropriate settings
3.27.6.9 | The ventilator must have a mute alarm function, active only for a limited time (no more than 2 minutes).
3.27.6.10 | It must not be possible to cancel the visual alarm until the problem has been rectified.
3.27.6.11 | All components of the ventilator and breathing system in contact with the patient’s airway must be latex free.
3.27.7 | All components of the ventilator and breathing system in contact with the patient’s airway must be steam autoclavable up to 134ᴼ C
3.28 | Integrated Gas Monitoring
3.28.1 | Oxygen Monitoring
3.28.1.1 | The system must display both inspired and end-tidal oxygen concentration.
3.28.1.2 | The oxygen analyser must have user adjustable low and high alarms – visual and audible. The low alarm must automatically default to no less than 21%.
3.28.1.3 | The system must be accurate (+/- 2% in 21-100%) range, and have a rapid response time (≤ 500ms).
3.28.1.4 | The unit must sound a continuous alarm if the oxygen pressure falls below 190 kPa, then all nitrous oxide must be automatically cut from the patient circuit.
3.28.1.5 | Analyser preferrably of paramagnetic type. If fuel cell offered, provide full information on cost of disposables and frequency of replacement
4 | Patient Physiological Monitoring
4.1 | The patient monitor must be modular and suitable for use for neonatal, paediatric and adult patients.
4.2 | Must interface with the gas monitoring system if required
4.3 | The following paramaters must be monitored:
4.3.1 | - ECG: 3-lead and 5-lead. Must be capable of displaying 2 channels simultaneously.
4.3.2 | - ST-segment and arrhythmia analysis.
4.3.3 | - NIBP
4.3.4 | - Oxygen saturation (SpO2)
4.3.5 | - Plethysmography/respiratory rate
4.3.6 | - Temperature x 2 (minimum)
4.3.7 | - Invasive pressure x 2 (minimum)
4.4 | The patient monitor must be fully integrated with the workstation.
4.5 | Display must be no less than 400 mm diagonally.
4.6 | Minimum 6 waveforms must be able to be displayed simultaneously.
4.7 | Full colour display.
4.8 | The monitor must be capable of displaying trends for at least 24 hours in tabular and graphical format.
4.9 | The trend resolution must be user-adjustable and have a minimum resolution of 1 min.
4.10 | The front panel must be easy to clean and disinfect.
4.11 | The display must be flicker free and easy to read.
4.12 | Brightness control must be possible.
4.13 | In addition to the waveform display section, the monitor must have alphanumeric display of parameter identification, measured values, alarm and status messages.
4.13.1 | Audible and visual alarms must be available for all measured parameters.
4.13.2 | Default alarm limits if the limits are not set by the user
4.13.3 | Alarm indication: Visual and audible with mute
4.13.4 | Adjustable audio alarm volume
4.13.5 | Low battery: Alarm at least 10 minutes before battery power is exhausted
4.13.6 | It must be possible to mute the audible alarm for a short period (no longer than 2 minutes), but no override of the visual alarm must be possible.
4.13.7 | It must be possible to suspend the alarms of certain parameters during clinical procedures, but the alarm must reactivate automatically after a suitable time delay. Any silencing or disabling of alarms must be clearly indicated on the monitor.
4.14 | Input impedance of the system > 10 MΩ.
4.15 | Maximum allowable leakage current to the patient must be less than 10 µA.
4.16 | The monitor must be protected against defibrillator and electrosurgical interference.
4.17 | The monitor must be of low power consumption, and must not require fan or other forced ventilation.
4.18 | All cables and sockets must be colour coded for easy recognition, and must conform to international standards
4.19 | ECG
4.19.1 | A full set of 3/5 lead ECG cables must be included in the offer.
4.19.2 | The ECG amplifier must employ a differential input with high noise suppression.
4.19.3 | Sweep speed: at least 25 to 50 mm/sec.
4.19.4 | Freeze facility for ECG.
4.19.5 | The input impedance must be > 10 MΩ.
4.19.6 | The heart rate measuring range: 20 to 300 bpm, with an accuracy < 5 bpm.
4.19.7 | An 50 Hz filter must be provided.
4.19.8 | Alarm limits must be adjustable in the following range:
4.19.9 | Low: 40 to 150
4.19.10 | High: 70 to 250.
4.20 | Non Invasive Blood Pressure
4.20.1 | Must be suitable for adults, paeds and neonates
4.20.2 | The unit must use the oscillometric method of measurement.
4.20.3 | Modes: Manual and automatic
4.20.4 | Reusable neonate, paediatric, adult and obese cuffs must be supplied
4.20.5 | The cuffs must be robust, interchangeable and easy to clean.
4.20.6 | Must display systolic, diastolic and mean pressures for each measurement
4.20.7 | Adult: Systolic: 30 - 270 mm Hg, Diastolic: 20 - 240 mm Hg, MAP: 20 - 250 mm Hg
4.20.8 | Paeds: Systolic: 30 - 160 mm Hg, Diastolic: 20 - 120, MAP: 20 - 120 mm Hg
4.20.9 | Values must be displayed numerically
4.20.10 | Auto zero when unit is switched on
4.20.11 | The interval between measurements must be selectable from 1 – 120 minutes.
4.20.12 | Adjustable high and low alarm limits for systolic, diastolic, mean arterial pressure and heart rate.
4.20.13 | Over pressure protection: Automatic deflation at 300 mm Hg adult, 150 mm Hg neonate
4.20.14 | Average measurement time: approximately 30 s
4.20.15 | If measurement exceeds two minutes the cuff must deflate, an alarm must sound and likely faults indicated.
4.20.16 | The air hose must be at least two metres long.
4.21 | Oxygen Saturation.
4.21.1 | Continuous monitoring of oxygen saturation levels by means of the photometric principle.
4.21.2 | Must be suitable for adults, paeds and neonates
4.21.3 | The unit must continuously display the SaO2 % waveform
4.21.4 | Oxygen saturation (SaO2 %): Range of measurement: 0 - 100%
4.21.5 | The accuracy must be better than 2,5 % (70 - 100%) and 5% (50 - 70%).
4.21.6 | Pulse rate (beats/min): Range of measurement: 20 - 300 bpm
4.21.7 | The pulse rate accuracy must be better than ± 2 %.
4.21.8 | An audible bleep should sound for each heartbeat, the pitch of which should vary with the oxygen saturation.
4.21.9 | Automatic rejection of motion artefacts
4.22 | SpO2 Sensor:
4.22.1 | It must be the latest technology with clinical research support based on accuracy and real-time monitoring. Please submit the published peer reviewed/researched paper as supporting documents.
4.22.2 | The following reusable sensors must be provided:
4.22.2.1 | Finger sensor for adults and paeds.
4.22.2.2 | Earlobe sensor for adults and paeds.
4.22.2.3 | Finger wrap-around probe for neonates
4.22.3 | The sensor lead must be approximately 3 metres long.
4.23 | Temperature
4.23.1 | Temperature range must be approximately 15º to 45º C.
4.23.2 | Two reusable temperature probes must be included
4.24 | Circle System
4.24.1 | The circle absorber system must be integrated as part of the ventilation system.
4.24.2 | The circle absorber system must have an effective system to eliminate excess water created by the chemical reactions either by heating or water traps.
4.24.3 | The circle absorber system must be autoclavable at 134ºC.
4.24.4 | The circle absorber system must be supplied with a carbon dioxide absorber canister.
4.24.5 | The absorber canister should be reusable/refillable. Disposable cannisters should be quoted separately
4.24.6 | It must be easy to remove and replace the absorber canister while the workstation is in operation.
4.24.7 | Misalignment of the canister must be readily apparent to the user.
4.24.8 | The system must have a circle bypass facility while changing/refilling the absorber canister.
4.24.9 | The breathing circuit must contain an integrated adjustable pressure limiting (APL) valve with a minimum range of 0-70 cm H2O.
4.24.10 | The APL valve must have a tactile or electronic indicator when a value of between 20-30 cm H2O is exceeded.
4.24.11 | The circle absorber system must be latex free.
4.24.12 | The bag-to-ventilator switching mechanism must be clearly labeled and user-friendly. Mechanical or electronic
4.25 | Patient circuit pressure release valve
4.25.1 | Adjustable up to 70 cm water
4.25.2 | Audible alarm when the valve opens
4.25.3 | Manual override, non-lockable, to keep the release valve closed when the set pressure is exceeded.
4.26 | Flowmeter pressure release valve
4.26.1 | Situated between the flowmeters and the vaporisers
4.26.2 | Must open at 105 kPa +- 5 kPa
4.27 | Other System Requirements
4.27.1 | The unit must operate from both the 220 V, 50 Hz mains supply and internal rechargeable batteries
4.27.2 | The batteries must be charged while the unit is operating from the mains supply.
4.27.3 | The batteries must automatically take over in the event of a mains power failure.
4.27.4 | The batteries must have the capacity for full operation of the anaesthetic machine and ventilator for at least 90 minutes with batteries fully charged, under typical operating conditions.
4.27.5 | Graphic battery power indicator
4.27.6 | Recharge time < 8 hrs
4.27.7 | Mains cable at least four (4) metres long with SABS approved 15A three prong plug.
4.27.8 | The unit must have a rail for attaching a suction regulator.
4.28 | Data management:
4.28.1 | The workstation must be capable of downloading data to a USB device or memory card.
4.28.2 | The monitor must be able to communicate with a central station
4.28.3 | A printer must be incorporated in the unit
4.29 | The unit must be handed over in full operating order
5 | The following Accessories and Consumablesmust be supplied with the unit:
5.1 | Re-breathing bag, reusable, rubber
5.2 | Reservoir bag
5.3 | Pressure limiting valve, adjustable
5.4 | Oxygen fuel cell analyser with fuel cell
5.5 | Adult & Paediatric volume sensor re-usable
5.6 | Water trap for multigas monitor, re-usable (10 pcs.)
5.7 | A full set paediatric 3/5 lead ECG cables
5.8 | A full set adult 3/5 lead ECG cables
5.9 | Reusable paediatric cuffs
5.10 | Reusable adult cuffs
5.11 | Reusable neonate cuffs
5.12 | Reusable obese cuffs
5.13 | Air hoses
5.14 | Finger sensor for paeds
5.15 | Finger sensor for adults
5.16 | Earlobe sensor for paeds
5.17 | Earlobe sensor for adults
5.18 | Finger wrap-around probe for neonates
5.19 | Sensor lead
5.20 | Two reusable temperature probes
5.21 | State the cost and frequency of replacement for the flow sensors.
5.22 | State the cost and frequency of replacement for the fuel cell.
5.23 | Vaporizer key-fillers (Sevo & Iso)
5.24 | The unit must be supplied with two vaporizers (1x sevoflurane and 1x Iso)
5.25 | Anaesthesia circuit (Adult, single/double limb design) Disposable. 1.8m (Complete with Y-piece, Elbow, Catheter mount, Rebreathing bag)
5.26 | Anaesthesia circuit (Paediatric, single/double limb design) Disposable. 1.8m (Complete with Y-piece, Elbow, Catheter mount, Rebreathing bag)
5.27 | Anaesthesia circuit (neonatal, single/double limb design) Disposable. 1.8m (Complete with Y-piece, Elbow, Catheter mount, Rebreathing bag)
5.28 | Re-breathing circuit complete
5.29 | Non-re-breathing circuit complete
5.30 | Flow sensor (pack of 2)
5.31 | Spirometry line
5.32 | Sample line (pack of 10)
5.33 | Face masks: sizes 00, 0, 1, 2, 3, 4, 5.
5.34 | Bacterial/viral filter
5.35 | Y connector
5.36 | Manual ventilation hose
5.37 | Fresh gas tubing
5.38 | Seal for back-up cylinder connection on the unit
5.39 | Scavenger tubing (as per institution's scavenging outlet)
5.40 | Inspiratory and expiratory unidirectional valves
5.41 | Side stream CO2 kit (adult/paed)
5.42 | Disposable Adult Spirometry Set complete with spirometry tube, sample line and flow sensor, 3 m, box of 20 pieces
5.43 | Disposable Paediatric Spirometry Set complete with spirometry tube, sample line and flow sensor, 3 m, box of 20 pieces
5.44 | Gas sample tube with luer lock connection, 3m, disposable (10 pcs)
5.45 | Suction Tubing Set inc connector
5.46 | Resuscitator Adult, 1500mL, autoclavable, silicone
5.47 | Resuscitator Paediatric, 550mL, autoclavable, silicone
5.48 | Tubing set Paediatric for circle systems
6 | OPTIONAL ITEMS
6.1 | Prices incl VAT must be given for all optional items
6.2 | Brochures with full technical specifications must be supplied for each optional item offered
7 | Conformity Compliance (Please attach certificate)
7.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
7.2 | The unit must comply with electrical safety and performance standard such as IEC 60601-2-13 for Anaesthetic systems, attached certification
7.2.1 | Please provide unique ref number of the IEC ceriticate:
7.4 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
7.4.1 | Please provide unique ref number of the ISO 13485 ceriticate:
7.5 | Model on offer must comply with EU Medical Device Regulation (EU) 2017/745 or equvalent from SAHPRA recognized jurisdictions . Attach a copy of certification
7.5.1 | Please provide unique ref number of the ceriticate:
7.6 | Hazardous Substance Act: If this type of equipment / apparatus appears on the schedule of Hazardous substances issued by the Directorate: Health Technology of the Department of Health, a license in terms of the Act on Hazardous Substances (Act 15/1973) must be submitted with the bid / quotation documents. The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
7.7 | All electrical/electronic medical equipment must be licensed by Radiation Control, where applicable a copy of the license must be submitted The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
7.7.1 | Radiation Control License number:
7.8 | The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade (rubber to with 2 screws) type and it must be a minimum length of three (3) meters. NB the mains cable of the unit tendered for must be SABS colour coded.
7.9 | The equipment quoted for must be protected against electromagnetic interference IEC 60601-1-2
7.10 | The unit must comply with electrical safety and performance standard such as IEC 60601-2-49 for Patient Monitoring, attached certification
7.10.1 | Please provide unique ref number of the IEC ceriticate:
8 | COMPREHENSIVE MAINTENANCE
8.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) should be quoted for, as per the SCC document
8.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries, Flow sensors and Oxygen sensor), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
8.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
8.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries, Flow sensors and Oxygen sensor), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
8.3 | The bidder to provide the OEM recommended service interval of the unit:
8.4 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
9 | General
9.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
9.2 | The unit on offer must be the latest technology.
9.2.1 | Initial Manufacturing year of Model:
| ANAESTHESIA DELIVERY SYSTEM: STANDARD with integrated physiological monitor, Replacement Accessories | Complies Yes/No | Comments
RT2-01-04a | A full set paediatric 3/5 lead ECG cable
RT2-01-04b | A full set adult 3/5 lead ECG cable
RT2-01-04c | Reusable paediatric cuff
RT2-01-04d | Reusable adult cuff
RT2-01-04e | Reusable neonate cuff
RT2-01-04f | Reusable obese cuff
RT2-01-04g | Air hose
RT2-01-04h | Finger sensor for paeds
RT2-01-04i | Finger sensor for adults
RT2-01-04j | Earlobe sensor for paeds
RT2-01-04k | Earlobe sensor for adults
RT2-01-04l | Finger wrap-around probe for neonates
RT2-01-04m | Sensor lead
RT2-01-04n | Reusable temperature probe
Ventilator modes- UPGRADES
RT2-01-04o | Pressure Support with Mandatory Apnoea back-up ventilation
RT2-01-04p | Volume Controlled Ventilation + SIMV + Pressure Support
RT2-01-04q | Pressure Control Ventilation + SIMV + Presssure Support
RT2-01-04r | CPAP

## Anaesthetic, Advanced
RT2-01-05 | ANAESTHESIA DELIVERY SYSTEM: Advanced without Physiological Monitor | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
1 | Description
A medical device designed to administer anesthetics to patients during surgical procedures.
2 | Product Offered (Anaesthetic Machine)
2.1 | Make:
2.2 | Model:
2.3 | Initial Manufacturing year of Model:
2.4 | Radiation Control License number:
3 | Technical Specification
3.1 | An advanced anaesthesia delivery system comprising a completely integrated electronic anaesthesia workstation.
3.2 | The unit must include electronicaly integrated oxygen, nitrous oxide and air delivery and monitoring.
3.3 | Ventilation and anaesthesia vapour monitoring must be fully integrated within the unit.
3.4 | The anaesthesia unit must be able to communicate/interface with a physiological monitor
3.5 | The anaesthesia system must comprise:
3.5.1 | Anaesthesia system suitable for neonatal (down to 500g), paediatric and adult anaesthesia.
3.5.2 | Ventilator suitable for ventilating neonatal (down to 500g), paediatric, adult and morbid/obese patients
3.6 | Multigas analyser:
3.6.1 | Oxygen
3.6.2 | Nitrous oxide
3.6.3 | Carbon dioxide
3.7 | Anaesthetic vapour monitoring:
3.7.1 | Isoflurane
3.7.2 | Sevoflurane
3.7.3 | Desflurane
3.8 | Complete and automated self test requiring very limited user intervention.
3.9 | The self test must provide the measurements and values for automatic leakage and compliance.
3.10 | Automatic calibration of sensors for anaesthetic gases and vapours must form part of the self test and require very limited user intervention.
3.11 | It must be possible to override the self test facility. The system must display the override status
3.12 | The unit must provide conventional, low-flow (down to 200 ml/min) anaesthesia.
3.13 | The unit must return analysed gas back to the circle system
3.14 | A safety oxygen function is required for the delivery of oxygen through the anaesthesia unit in the event of an electronic failure of the gas delivery system.
3.15 | The unit must be supplied with a minimum of 2 storage drawer
3.16 | The unit must be supplied with a minimum of 4 additional 220 V power sockets. These sockets must be excluded from the battery backup circuit in case of a power failure
3.17 | The unit must have a working area or a retractable shelf. Width: not less than 300 mm. Depth: not less than 200 mm
3.18 | The unit must have a light to enable the anaesthetist to work in a darkened theatre.
3.20 | Trolley
3.20.1 | The unit must be mounted on a strong , robust and compact trolley manufactured from corrosion free material.
3.20.2 | The trolley must have no sharp corners and be easy to clean.
3.20.3 | The trolley must have at least four antistatic castors of diameter 100 mm or greater.
3.20.4 | Trolley must have two lockable castors or a central locking brake.
3.20.5 | The unit must accommodate an oxygen cylinder. Cylinder size according to user requirement. Size to be compatible with current state vendor of oxygen.
3.20.6 | Trolley must have accessible handles for ease of movement
3.20.7 | The trolley, with all components mounted, must be stable without toppling over when placed on a 30° slope.
3.20.8 | The unit must be supplied with a castor guard system, preventing the entaglement and trapping of cables within the wheels.
3.20.9 | The trolley must have a mounting system suitable for a physiologial monitor.
3.21 | Pipeline and Cylinder Gas Supplies
3.21.1 | The anaesthesia system must be designed to use external piped medical gasses (O2, N2O and medical air) and vacuum from medical service outlet sockets.
3.21.2 | Connection hoses for above must be provided
3.21.3 | Hoses must be fitted with Heyer type non-interchangeable probes for the existing medical gas outlets as specified by the purchaser.
3.21.4 | The medical air hose connection or CO2 circle system must include a water trap
3.21.5 | A safety pressure relief valve of 60 - 80 cm H2O must be provided to protect the system from excessive pressure .
3.21.6 | The unit must function safely and correctly within the pipeline supply pressure range of: 270 - 600 kPa
3.21.7 | The system must preferentially utilise the piped oxygen supply over the cylinder supply, if pipeline pressure is > 270 kPa
3.21.8 | Connection hoses must be:
3.21.8.1 | Flexible and at least 6 metres long
3.21.8.2 | Antistatic
3.21.9 | Colour coded as follows:
3.21.9.1 | O2: white
3.21.9.2 | N2O: blue
3.21.9.3 | Vacuum: yellow
3.21.9.4 | Medical air: black
3.21.9.5 | Scavenge low vacuum: orange or yellow
3.21.10 | Where relevant, fitted with outlet probes complying with SANS 1409:2014
3.21.11 | Permanently bonded to the outlet probe and the machine connector
3.21.12 | Hoses for gases above atmospheric pressure, must designed for pressures of at least 1 000 kPa.
3.21.13 | Vacuum hoses may not collapse or kink at 90 kPa below atmospheric pressure, when subjected to 180° bend of inside radius equal to the outside diameter of the hose.
3.21.14 | All connection hoses must be fed to a central manifold
3.21.15 | All high pressure gas and vacuum pipe connections (other than hoses) and all pipes connected to pressure regulators and pressure gauges must be stainless steel (18/8 grade) pipes that are oil-free and colour coded as for hoses (see above) over their entire lengths
3.21.16 | An in-line non-return valve must be fitted to each gas pipeline on the high pressure (inlet) line
3.21.17 | An in-line non-return valve must be fitted to each gas cylinder yoke
3.21.18 | A replaceable gas filter, to block particles greater than 65 µm, must be fitted on the upstream side of the non-return valve
3.21.19 | Each gas pipeline inlet must have a pressure gauge or electronic display of pressure. The pressure range (415 kPa ± 75 kPa) must be clearly indicated in colour
3.21.20 | Must be fitted with a gas cylinder yoke to accommodate one oxygen gas cylinder with a pin index system
3.21.21 | Each gas cylinder yoke must have a pressure gauge or electronic display of pressure
3.21.22 | The pressure gauges or electronic displays of pressure for both pipeline & cylinder supplies, calibrated in kPa for each line, must be labelled and visible from the front of the unit.
3.21.23 | O2 and N2O regulators must be provided complying with CKS 605:1987.
3.21.24 | A pressure regulator complying with CKS 605 must be supplied for each gas inlet and must reduce the gas pressure to 415 kPa without freezing up, creeping, chattering or whistling
3.21.25 | Outlet sockets and probes must comply with SANS 1409:2014.
3.21.26 | Sockets with hoses must be colour-coded as per SANS 1409:2014:
3.21.26.1 | Oxygen: White background with black letters
3.21.26.2 | Nitrous oxide: Blue background with white letters
3.21.26.3 | Medical air: One half black with white letters, the other half white with black letters
3.21.26.4 | Vacuum: Yellow background with black letters
3.21.27 | Continuous display, visible from the front of the machine, of both pipeline and cylinder pressures even during power failure
3.21.28 | Pipeline and cylinder pressures must be differentially displayed, and clearly indicated
3.21.29 | Safety pressure relief valves must protect the entire pneumatic system from excessive input pressures from pipeline and cylinder supplies.
3.21.30 | When a cylinder is detached from the pin-indexed yoke no gas leak must occur at the yoke.
3.21.31 | Airway pressure must be monitored and displayed
3.21.32 | Safety relief valve for all gases must be supplied and set at 600kPa
3.22 | Scavenging
3.22.1 | The unit must have a scavenging system with a hose at least 5 metres long with the suitable connector as per the institutions scavenging installation.
3.22.2 | Air-break device to prevent negative pressure on the breathing system
3.22.3 | If an active sytem is supplied a pressure indicator must be present
3.22.4 | The scavenging outlet must be a 30mm ISO connector.
3.23 | Patient Gas Delivery System
3.23.1 | Patient gas delivery must be electronically controlled
3.23.2 | Must display the delivered patient gas flow using an ISO colour standard for each gas.
3.23.3 | The unit must be able to deliver a total patient gas flow of 200 ml/min to 18 l/min (Minimum). The unit must have the ability to deliver at least 10 l/min of oxygen. These must be incremental adjustments
3.23.4 | Gas supply must conform to:
3.23.5 | O2 minimum flow range of 200 ml/min to 10 l/min. Accuracy ± 5%
3.23.6 | medical air minimum flow range 0 to 10 l/min. Accuracy ± 5%
3.23.7 | nitrous oxide minimum flow range 0 to 10 l/min. Accuracy ± 5%
3.23.8 | Flowmeter sequence to be according to standards approved by the South African Society of Anaesthesiologists - Oxygen on the right
3.23.9 | The delivery system must prevent simultaneous delivery of two anaesthetic vapours.
3.23.10 | A separate common gas outlet to accept a standard 22 mm outside diameter and 15 mm internal diameter.
3.24 | Hypoxic guard
3.24.1 | The unit must include a safety mechanism that ensures a minimum concentration of 25% oxygen in any oxygen/nitrous oxide gas mixture across all flow rates.
3.24.2 | The unit must include a safety mechanism that ensures a minimum concentration of 21% oxygen in any oxygen/air gas mixture across all flow rates.
3.24.3 | The unit must include a high-flow oxygen flush mechanism with a flow rate of at least 35 l/min
3.24.4 | Oxygen flush must be non-lockable, recessed and self closing
3.24.5 | This system must bypass the flowmeters and vaporisers
3.25 | Breathing System
3.25.1 | Total circuit volume, including absorber, must not exceed 2,7 litres
3.25.2 | With APL valve 5 - 75 cm H2O
3.25.3 | With ventilator to manual switch. This switch when directed to the ventilator must isolate the APL valve from the circle
3.25.4 | Breathing system must be fully autoclavable to 134°C
3.25.5 | There must be a mechanism to remove water from the breathing system
3.25.6 | Breathing system must be latex free
3.25.7 | Bag-to-ventilator switch automatically begins in the ventilator position
3.25.8 | Bidders to supply the breathing circuit as required by the user
3.26 | Integrated Ventilator
3.26.1 | Ventilator suitable for ventilating adults, paediatric and neonates (> 500g) patients without changing ventilator components (e.g. bellows)
3.26.2 | The ventilator must be electronically controlled
3.26.3 | The ventilation technology offered may be turbine, volume reflector, piston or ascending bellows
3.26.4 | The minimum basic modes to be supplied:
3.26.4.1 | Manual/Spontaneous Ventilation
3.26.4.2 | Volume Control Ventilation
3.26.4.3 | Pressure Control Ventilation
3.26.4.4 | Pressure Support with Mandatory Apnoea back-up ventilation
3.26.4.5 | Volume Controlled Ventilation + SIMV + Pressure Support
3.26.4.6 | Pressure Control Ventilation + SIMV + Presssure Support
3.26.4.7 | CPAP
3.26.5 | Additional modes of ventilation may be offered and quoted for separately
3.26.6 | The ventilator must be integrated with the workstation and be user friendly
3.26.7 | Frequency: minimum range of 5 – 60 bpm.
3.26.8 | Inspiratory time: minimum range of 0.2 – 10.0 seconds.
3.26.9 | Minimum Inspiration/Expiration Ratio: 4:1 – 1:4.
3.26.10 | Adjustable inspiratory pause/plateau time (volume control ventilation): 0-30% of inspiratory time
3.26.11 | Peak Inspiratory Pressure: minimum range of: PEEP + 5 to 70 cm H2O.
3.26.12 | Positive End Expiratory Pressure (PEEP): Integrated, electronically controlled. Minimum range: 0 – 30 cm H2O.
3.26.13 | Tidal Volume: 20ml – 1500ml per breath. Minimum accuracy: 15%
3.26.14 | Peak Inspiratory Flow: At least 90 lpm
3.26.15 | The ventilator must be capable of flow or pressure triggering.
3.26.16 | Adjustable Inspiratory Rise time/Ramp with a range of 0.1 - 2 sec
3.26.17 | Compliance compensation/Tidal volume compensation/Compressible volume compensation or equivalent.
3.26.18 | Patient gas flow compensation must be automatic so that alterations in patient gas flow does not affect delivered tidal volumes
3.26.19 | It must be possible to set CPAP in manual/spontaneous mode within a range of: 1 to 10 cm H2O
3.26.20 | In event of power or technical failure, manual ventilation must be possible.
3.27 | Display:
3.27.1 | All displays must be in acceptable English
3.27.2 | Colour display must include numerical and graphic display of parameters measured
3.27.3 | Graphic display to feature at least 3 waveforms: - pressure, flow and volume.
3.27.4 | Must also display flow-volume and pressure-volume loops with a reference loop
3.27.5 | Alarms : The ventilator must have the following user adjustable alarms
3.27.5.1 | Apnoea alarm
3.27.5.1 | High airway pressure alarm
3.27.5.1 | Continuous pressure alarm
3.27.5.1 | High and low expiratory tidal volume alarm
3.27.5.1 | Minute volume, high/low alarm
3.27.5.1 | Inspired oxygen (Fi02), high/low alarm
3.27.5.1 | PEEP, low alarm
3.27.5.1 | Leakage / Check breathing circuit alarm or equivalent
3.27.5.1 | Alarm profile must be user definable or default to patient appropriate settings
3.27.5.1 | The ventilator must have a mute alarm function, active only for a limited time.
3.27.5.1 | It must not be possible to cancel the visual alarm until the problem has been rectified.
3.27.6 | All components of the ventilator and breathing system in contact with the patient’s airway must be latex free.
3.27.7 | All components of the ventilator and breathing system in contact with the patient’s airway must be steam autoclavable up to 134ᴼ C
3.28 | Integrated Gas Monitoring
3.28.1 | Oxygen Monitoring
3.28.1.1 | The system must display both inspired and end-tidal oxygen concentration.
3.28.1.2 | The oxygen analyser must have user adjustable low and high alarms – visual and audible. The low alarm must automatically default to no less than 21%.
3.28.1.3 | The system must be accurate (+/- 2% in 21-100%) range, and have a rapid response time (≤ 500ms).
3.28.1.4 | The unit must sound a continuous alarm if the oxygen pressure falls below 190 kPa, then all nitrous oxide must be automatically cut from the patient circuit.
3.28.1.5 | Bidder to give details of the oxygen analyser
3.28.1.6 | The lifespan of the oxygen analyser must be guaranteed for at least 10 years.
3.28.1.7 | A system using replaceable oxygen fuel cells will not be accepted
3.28.2 | Anaesthetic vapour, carbon dioxide and nitrous oxide monitoring:
3.28.2.1 | Anaesthetic vapour, N2O monitoring and capnography must be integrated into the anaesthetic workstation.
3.28.2.2 | The monitoring must be of the side-stream type.
3.28.2.3 | Sampling: 100 - 200 ml/min
3.28.2.4 | Sampled gases must be internally returned to the breathing system.
3.28.2.5 | CO2 must be presented as a numerical display of inspired and end-tidal concentration and a capnography waveform.The displayed units must be user-selectable: kPa or mmHg.
3.28.2.6 | The monitoring system must have automatic identification of anaesthetic vapours.
3.28.2.7 | The system must have a graphical display indicating the type of anaesthetic vapour/s currently in use and the concentration thereof. The colour coding must be as per international colour coding for anaesthetic vapours.
3.28.2.8 | The anaestheric vapour must be displayed as vol % and must include inspiratory and expiratory concentrations.
3.28.2.9 | The monitoring system must be able to simultaneously detect, identify and display mixed anaesthetic vapours.
3.28.2.10 | High and low anaesthetic vapour alarms must be present and these must be user-definable.
3.28.2.11 | There must be automatic calculation of minimum alveolar concentration.
3.28.2.12 | The system must be able to calculate anaesthetic gas utilization in ml per case
3.28.2.13 | The unit must have an indication of the level of vapour in the vaporiser
3.28.2.14 | Integrated vapour analysis must deliver an audible "low vapour " alarm and a critical alarm for an empty tank
3.29 | Circle System
3.29.1 | The circle absorber system must be integrated as part of the ventilation system.
3.29.2 | The circle absorber system must have an effective system to eliminate excess water created by the chemical reactions either by heating or water traps.
3.29.3 | The circle absorber system must be autoclavable at 134ºC.
3.29.4 | The circle absorber system must be supplied with a carbon dioxide absorber canister.
3.29.5 | The absorber canister may be reusable/refillable
3.29.6 | It must be easy to remove and replace the absorber canister while the workstation is in operation.
3.29.7 | Misalignment of the canister must be readily apparent to the user.
3.29.8 | The system must have a circle bypass facility while changing/refilling the absorber canister.
3.29.9 | The breathing circuit must contain an integrated adjustable pressure limiting (APL) valve with a minimum range of 5-75 cm H2O.
3.29.10 | The APL valve must have a tactile or electronic indicator when a value of between 20-30 cm H2O is exceeded.
3.29.11 | The circle absorber system must be latex free.
3.29.12 | The bag-to-ventilator switching mechanism must be clearly labeled and user-friendly. Mechanical or electronic
3.30 | Other system requirements
3.30.1 | The unit must operate from both the 220 V, 50 Hz mains supply and internal rechargeable batteries
3.30.2 | The batteries must be charged while the unit is operating from the mains supply.
3.30.3 | The batteries must automatically take over in the event of a mains power failure.
3.30.4 | The batteries must have the capacity for full operation of the anaesthetic machine and ventilator for at least 90 minutes with batteries fully charged, under typical operating conditions.
3.30.5 | Graphic battery power indicator
3.30.6 | Recharge time < 8 hrs
3.30.7 | Mains cable at least four metres long with SABS approved 15A three prong plug.
3.30.8 | The unit must have a rail for attaching a suction regulator.
3.31 | Data management: The workstation must be capable of downloading data to a USB device or memory card.
3.31.1 | The unit must be handed over in full operating order
4 | The following Accessories and Consumables must be supplied with the unit:
4.1 | Re-breathing bag, reusable, rubber
4.2 | Reservoir bag
4.3 | Pressure limiting valve, adjustable
4.4 | Adult & Paediatric volume sensor re-usable
4.5 | Water trap for multigas monitor, re-usable (10 pcs.)
4.6 | Vaporizer key-fillers (Sevo, & other choice below)
4.7 | The unit must be supplied with two vaporizers (one must sevoflurane, the other must be of the choice of the buying institution)
4.8 | Anaesthesia circuit (Adult, single/double limb design) Disposable. 1.8m (Complete with Y-piece, Elbow, Catheter mount, Rebreathing bag)
4.9 | Anaesthesia circuit (Paediatric, single/double limb design) Disposable. 1.8m (Complete with Y-piece, Elbow, Catheter mount, Rebreathing bag)
4.10 | Anaesthesia circuit (neonatal, single/double limb design) Disposable. 1.8m (Complete with Y-piece, Elbow, Catheter mount, Rebreathing bag)
4.11 | Re-breathing circuit complete
4.12 | Non-re-breathing circuit complete
4.13 | Flow sensor (pack of 2)
4.14 | Spirometry line
4.15 | Sample line (pack of 10)
4.16 | Face masks: sizes 00, 0, 1, 2, 3, 4, 5.
4.17 | Bacterial/viral filter
4.18 | Y connector
4.19 | Manual ventilation hose
4.20 | Fresh gas tubing
4.21 | Seal for back-up cylinder connection on the unit
4.22 | Scavenger tubing
4.23 | Inspiratory and expiratory unidirectional valves
4.24 | Side stream CO2 kit (adult/paed)
4.25 | Disposable Adult Spirometry Set complete with spirometry tube, sample line and flow sensor, 3 m, box of 20 pieces
4.26 | Disposable Paediatric Spirometry Set complete with spirometry tube, sample line and flow sensor, 3 m, box of 20 pieces
4.27 | Gas sample tube with luer lock connection, 3m, disposable (10 pcs)
4.28 | Suction Tubing Set inc connector
4.29 | Resuscitator Adult, 1500mL, autoclavable, silicone
4.30 | Resuscitator Paediatric, 550mL, autoclavable, silicone
4.31 | Tubing set Paediatric for circle systems
5 | OPTIONAL ITEMS: Any accessories and modules may be offered as optional items
5.1 | Prices incl VAT must be given for all optional items
5.2 | Brochures with full technical specifications must be supplied for each optional item offered
6 | Conformity Compliance (Please attach certificate)
6.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
6.2 | The unit must comply with electrical safety and performance standard such as IEC 60601-2-13 for Anaesthetic systems, attached certification
6.2.1 | Please provide unique ref number of the IEC ceriticate:
6.4 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
6.4.1 | Please provide unique ref number of the ISO 13485 ceriticate:
6.5 | Model on offer must comply with EU Medical Device Regulation (EU) 2017/745 or equvalent from SAHPRA recognized jurisdictions . Attach a copy of certification
6.5.1 | Please provide unique ref number of the ceriticate:
6.6 | Hazardous Substance Act: If this type of equipment / apparatus appears on the schedule of Hazardous substances issued by the Directorate: Health Technology of the Department of Health, a license in terms of the Act on Hazardous Substances (Act 15/1973) must be submitted with the bid / quotation documents. The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
6.7 | All electrical/electronic medical equipment must be licensed by Radiation Control, where applicable a copy of the license must be submitted The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
6.7.1 | Radiation Control License number:
6.8 | The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade (rubber to with 2 screws) type and it must be a minimum length of three (3) meters. NB the mains cable of the unit tendered for must be SABS colour coded.
6.9 | The equipment quoted for must be protected against electromagnetic interference IEC 60601-1-2
7 | COMPREHENSIVE MAINTENANCE
7.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) should be quoted for, as per the SCC document
7.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries, Flow sensors and Oxygen sensor), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
7.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
7.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries, Flow sensors and Oxygen sensor), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
7.3 | The bidder to provide the OEM recommended service interval of the unit:
7.4 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
8 | General
8.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
8.2 | The unit on offer must be the latest technology.
8.2.1 | Initial Manufacturing year of Model:
| ANAESTHESIA DELIVERY SYSTEM: ADVANCED WITHOUT PHYSIOLOGICAL MONITOR, REPLACEMENT ACCESSORIES | Complies Yes/No
RT2-01-05a | Anaesthesia circuit (Adult, single/double limb design) Disposable. 1.8m (Complete with Y-piece, Elbow, Catheter mount, Rebreathing bag)
RT2-01-05b | Anaesthesia circuit (Paediatric, single/double limb design) Disposable. 1.8m (Complete with Y-piece, Elbow, Catheter mount, Rebreathing bag)
RT2-01-05c | Anaesthesia circuit (neonatal, single/double limb design) Disposable. 1.8m (Complete with Y-piece, Elbow, Catheter mount, Rebreathing bag)
RT2-01-05d | Re-breathing circuit complete
RT2-01-05e | Non-re-breathing circuit complete
RT2-01-05f | Flow sensor (pack of 2)
RT2-01-05g | Spirometry line
RT2-01-05h | Sample line (pack of 10)
RT2-01-05i-1 | Face masks: sizes 00
RT2-01-05i-2 | Face masks: sizes 0
RT2-01-05i-3 | Face masks: sizes 1
RT2-01-05i-4 | Face masks: sizes 2
RT2-01-05i-5 | Face masks: sizes 3
RT2-01-05i-6 | Face masks: sizes 4
RT2-01-05i-7 | Face masks: sizes 5
RT2-01-05j | Bacterial/viral filter
RT2-01-05k | Y connector
RT2-01-05l | Manual ventilation hose
RT2-01-05m | Fresh gas tubing
RT2-01-05n | Seal for back-up cylinder connection on the unit
RT2-01-05o | Scavenger tubing (as per institution's scavenging outlet)
RT2-01-05p | Inspiratory and expiratory unidirectional valves
RT2-01-05q | Side stream CO2 kit (adult/paed)
RT2-01-05r | Disposable Adult Spirometry Set complete with spirometry tube, sample line and flow sensor, 3 m, box of 20 pieces
RT2-01-05s | Disposable Paediatric Spirometry Set complete with spirometry tube, sample line and flow sensor, 3 m, box of 20 pieces
RT2-01-05t | Gas sample tube with luer lock connection, 3m, disposable (10 pcs)
RT2-01-05u | Suction Tubing Set inc connector
RT2-01-05v | Resuscitator Adult, 1500mL, autoclavable, silicone
RT2-01-05w | Resuscitator Paediatric, 550mL, autoclavable, silicone
RT2-01-05x | Tubing set Paediatric for circle systems
RT2-01-05y | sevoflurane vaporizer
RT2-01-05z | Isoflurane vaporizer
RT2-01-05zz | Desflurane vaporizer

## Anaesthetic, Advanced + Monitor
RT2-01-06 | ANAESTHESIA DELIVERY SYSTEM: Advanced with integrated Physiological Monitor | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
1 | Description
1.1 | A medical device designed to administer anesthetics to patients during surgical procedures.
1a | Product Offered (Anaesthetic Machine)
1a.1 | Make:
1a.2 | Model:
1a.3 | Initial Manufacturing year of Model:
1a.4 | Radiation Control License number:
1b | Product Offered (Monitor)
1b.1 | Make:
1b.2 | Model:
1b.3 | Initial Manufacturing year of Model:
1b.4 | Radiation Control License number:
2 | Technical Specification
2.1 | An advanced anaesthesia delivery system comprising a completely integrated electronic anaesthesia workstation is required.
2.2 | The unit must include electronicaly integrated oxygen, nitrous oxide and air delivery and monitoring.
2.3 | Ventilation and anaesthesia vapour monitoring must be fully integrated within the unit.
2.4 | The patient physiological monitoring must communicate/interface with the anaesthesia unit.
2.5 | The anaesthesia system must comprise:
2.5.1 | Anaesthesia system suitable for neonatal (down to 500g), paediatric and adult anaesthesia.
2.5.2 | Ventilator suitable for ventilating neonatal (down to 500g), paediatric, adult and morbid/obese patients
2.6 | Physiological monitor
2.7 | Multigas analyser:
2.7.1 | Oxygen
2.7.2 | Nitrous oxide
2.7.3 | Carbon dioxide
2.8 | Anaesthetic vapour monitoring:
2.8.1 | Isoflurane
2.8.2 | Sevoflurane
2.8.3 | Desflurane
2.9 | Complete and automated self test requiring very limited user intervention.
2.10 | The self test must provide the measurements and values for automatic leakage and compliance.
2.11 | Automatic calibration of sensors for anaesthetic gases and vapours must form part of the self test and require very limited user intervention.
2.12 | It must be possible to override the self test facility. The system must display the override status
2.13 | The unit must provide conventional, low-flow (down to 200 ml/min) anaesthesia.
2.14 | The unit must return analysed gas back to the circle system
2.15 | A safety oxygen function is required for the delivery of oxygen through the anaesthesia unit in the event of an electronic failure of the gas delivery system.
2.16 | The patient monitor must be fixed to the anaesthesia device with a suitable mounting system or integrated into the device.
2.17 | The unit must be supplied with a minimum of 2 storage drawer
2.18 | The unit must be supplied with a minimum of 4 additional 220 V power sockets. These sockets must be excluded from the battery backup circuit in case of a power failure
2.19 | The unit must have a working area or a retractable shelf. Width: not less than 300 mm. Depth: not less than 200 mm
2.20 | The unit must have a light to enable the anaesthetist to work in a darkened theatre.
2.21 | Trolley
2.21.1 | The unit must be mounted on a strong , robust and compact trolley manufactured from corrosion free material.
2.21.2 | The trolley must have no sharp corners and be easy to clean.
2.21.3 | The trolley must have at least four antistatic castors of diameter 100 mm or greater.
2.21.4 | Trolley must have two lockable castors or a central locking brake.
2.21.5 | The unit must accommodate an oxygen cylinder. Cylinder size according to user requirement. Size to be compatible with current state vendor of oxygen.
2.21.6 | Trolley must have accessible handles for ease of movement
2.21.7 | The trolley, with all components mounted, must be stable without toppling over when placed on a 30° slope.
2.21.8 | The unit must be supplied with a castor guard system, preventing the entaglement and trapping of cables within the wheels.
2.22 | Pipeline and Cylinder Gas Supplies
2.22.1 | The anaesthesia system must be designed to use external piped medical gasses (O2, N2O and medical air) and vacuum from medical service outlet sockets.
2.22.2 | Connection hoses for above must be provided
2.22.3 | Hoses must be fitted with Heyer type non-interchangeable probes for the existing medical gas outlets as specified by the purchaser.
2.22.4 | The medical air hose connection or CO2 circle system must include a water trap
2.22.5 | A safety pressure relief valve of 60 - 80 cm H2O must be provided to protect the system from excessive pressure .
2.22.6 | The unit must function safely and correctly within the pipeline supply pressure range of: 270 - 600 kPa
2.22.7 | The system must preferentially utilise the piped oxygen supply over the cylinder supply, if pipeline pressure is > 270 kPa
2.22.8 | Connection hoses must be:
2.22.8.1 | Flexible and at least 6 metres long
2.22.8.2 | Antistatic
2.22.9 | Colour coded as follows:
2.22.9.1 | O2: white
2.22.9.2 | N2O: blue
2.22.9.3 | Vacuum: yellow
2.22.9.4 | Medical air: black/white
2.22.9.5 | Scavenge low vacuum: orange or yellow
2.22.10 | Where relevant, fitted with outlet probes complying with SANS 1409:2014
2.22.11 | Permanently bonded to the outlet probe and the machine connector
2.22.12 | Hoses for gases above atmospheric pressure, must designed for pressures of at least 1 000 kPa.
2.22.13 | Vacuum hoses may not collapse or kink at 90 kPa below atmospheric pressure, when subjected to 180° bend of inside radius equal to the outside diameter of the hose.
2.22.14 | All connection hoses must be fed to a central manifold
2.22.15 | All high pressure gas and vacuum pipe connections (other than hoses) and all pipes connected to pressure regulators and pressure gauges must be stainless steel (18/8 grade) pipes that are oil-free and colour coded as for hoses (see above) over their entire lengths
2.22.16 | An in-line non-return valve must be fitted to each gas pipeline on the high pressure (inlet) line
2.22.17 | An in-line non-return valve must be fitted to each gas cylinder yoke
2.22.18 | A replaceable gas filter, to block particles greater than 65 µm, must be fitted on the upstream side of the non-return valve
2.22.19 | Each gas pipeline inlet must have a pressure gauge or electronic display of pressure. The pressure range (415 kPa ± 75 kPa) must be clearly indicated in colour
2.22.20 | Must be fitted with a gas cylinder yoke to accommodate one oxygen gas cylinder with a pin index system
2.22.21 | Each gas cylinder yoke must have a pressure gauge or electronic display of pressure
2.22.22 | The pressure gauges or electronic displays of pressure for both pipeline & cylinder supplies, calibrated in kPa for each line, must be labelled and visible from the front of the unit.
2.22.23 | O2 and N2O regulators must be provided complying with CKS 605:1987.
2.22.24 | A pressure regulator complying with CKS 605 must be supplied for each gas inlet and must reduce the gas pressure to 415 kPa without freezing up, creeping, chattering or whistling
2.22.25 | Outlet sockets and probes must comply with SANS 1409:2014.
2.22.26 | Sockets with hoses must be colour-coded as per SANS 1409:2014:
2.22.26.1 | Oxygen: White background with black letters
2.22.26.2 | Nitrous oxide: Blue background with white letters
2.22.26.3 | Medical air: One half black with white letters, the other half white with black letters
2.22.26.4 | Vacuum: Yellow background with black letters
2.22.27 | Continuous display, visible from the front of the machine, of both pipeline and cylinder pressures even during power failure
2.22.28 | Pipeline and cylinder pressures must be differentially displayed, and clearly indicated
2.22.29 | Safety pressure relief valves must protect the entire pneumatic system from excessive input pressures from pipeline and cylinder supplies.
2.22.30 | When a cylinder is detached from the pin-indexed yoke no gas leak must occur at the yoke.
2.22.31 | Airway pressure must be monitored and displayed
2.22.32 | Safety relief valve for all gases must be supplied and set at 600kPa
2.23 | Scavenging
2.23.1 | The unit must have a scavenging system with a hose at least 5 metres long with the suitable connector as per the institutions scavenging installation.
2.23.2 | Air-break device to prevent negative pressure on the breathing system
2.23.3 | If an active sytem is supplied a pressure indicator must be present
2.23.4 | The scavenging outlet must be a 30mm ISO connector.
2.24 | Patient Gas Delivery System
2.24.1 | Patient gas delivery must be electronically controlled
2.24.2 | and display the delivered patient gas flow using an ISO colour standard for each gas.
2.24.3 | The unit must be able to deliver a total patient gas flow of 200 ml/min to 18 l/min (Minimum). The unit must have the ability to deliver at least 10 l/min of oxygen. These must be incremental adjustments
2.24.4 | Gas supply must conform to:
2.24.4.1 | O2 minimum flow range of 200 ml/min to 10 l/min. Accuracy ± 5%
2.24.4.2 | medical air minimum flow range 0 to 10 l/min. Accuracy ± 5%
2.24.4.3 | nitrous oxide minimum flow range 0 to 10 l/min. Accuracy ± 5%
2.24.5 | Flowmeter sequence to be according to standards approved by the South Africal Society of Anaesthesiologists - Oxygen on the right
2.24.6 | The delivery system must prevent simultaneous delivery of two anaesthetic vapours.
2.24.7 | A separate common gas outlet to accept a standard 22 mm outside diameter and 15 mm internal diameter.
2.25 | Hypoxic guard
2.25.1 | The unit must include a safety mechanism that ensures a minimum concentration of 25% oxygen in any oxygen/nitrous oxide gas mixture across all flow rates.
2.25.2 | The unit must include a safety mechanism that ensures a minimum concentration of 21% oxygen in any oxygen/air gas mixture across all flow rates.
2.25.3 | The unit must include a high-flow oxygen flush mechanism with a flow rate of at least 35 l/min
2.25.4 | Oxygen flush must be non-lockable, recessed and self closing
2.25.5 | This system must bypass the flowmeters and vaporisers
2.26 | Breathing System
2.26.1 | Total circuit volume, including absorber, must not exceed 2,7 litres
2.26.2 | With APL valve 5 - 75 cm H2O
2.26.3 | With ventilator to manual switch. This switch when directed to the ventilator must isolate the APL valve from the circle
2.26.4 | There must be a mechanism to remove water from the breathing system
2.26.5 | Breathing system must be fully autoclavable to 134°C
2.26.6 | Breathing system must be latex free
2.26.7 | Bag-to-ventilator switch automatically begins in the ventilator position
2.26.8 | Reusable corrugated breathing circuit adequate for both adult and paediatric use to be supplied
2.26.9 | Bidders to supply the breathing circuit as required by the user
2.27 | Integrated Ventilator
2.27.1 | Ventilator suitable for ventilating adults, paediatric and neonates (> 500g) patients without changing ventilator components (e.g. bellows)
2.27.2 | The ventilator must be electronically controlled
2.27.3 | The ventilation technology offered may be turbine, volume reflector, piston or ascending bellows
2.27.4 | The minimum basic modes to be supplied:
2.27.4.1 | Manual/Spontaneous Ventilation
2.27.4.2 | Volume Control Ventilation
2.27.4.3 | Pressure Control Ventilation
2.27.4.4 | Pressure Support with Mandatory Apnoea back-up ventilation
2.27.4.5 | Volume Controlled Ventilation + SIMV + Pressure Support
2.27.4.6 | Pressure Control Ventilation + SIMV + Presssure Support
2.27.4.7 | CPAP
2.27.5 | Additional modes of ventilation may be offered and quoted for separately
2.27.6 | The ventilator must be integrated with the workstation and user friendly
2.27.7 | Frequency: minimum range of 5 – 60 bpm.
2.27.8 | Inspiratory time: minimum range of 0.2 – 10.0 seconds.
2.27.9 | Minimum Inspiration/Expiration Ratio: 4:1 – 1:4.
2.27.10 | Adjustable inspiratory pause/plateau time (volume control ventilation): 0-30% of inspiratory time
2.27.11 | Peak Inspiratory Pressure: minimum range of: PEEP + 5 to 70 cm H2O.
2.27.12 | Positive End Expiratory Pressure (PEEP): Integrated, electronically controlled. Minimum range: 0 – 30 cm H2O.
2.27.13 | Tidal Volume: 20ml – 1500ml per breath. Minimum accuracy: 15%
2.27.14 | Peak Inspiratory Flow: At least 90 lpm
2.27.15 | The ventilator must be capable of flow or pressure triggering.
2.27.16 | Adjustable Inspiratory Rise time/Ramp with a range of 0.1 - 2 sec
2.27.17 | Compliance compensation/Tidal volume compensation/Compressible volume compensation or equivalent.
2.27.18 | Patient gas flow compensation must be automatic so that alterations in patient gas flow does not affect delivered tidal volumes
2.27.19 | It must be possible to set CPAP in manual/spontaneous mode within a range of: 1 to 10 cm H2O
2.27.20 | In event of power or technical failure, manual ventilation must be possible.
2.28 | Display:
2.28.1 | All displays must be in acceptable English
2.28.2 | Colour display must include numerical and graphic display of parameters measured
2.28.3 | Graphic display to feature at least 3 waveforms: - pressure, flow and volume.
2.28.4 | Must also display flow-volume and pressure-volume loops with a reference loop
2.28.5 | Alarms : The ventilator must have the following user adjustable alarms:
2.28.5.1 | Apnoea alarm
2.28.5.2 | High airway pressure alarm
2.28.5.3 | Continuous pressure alarm
2.28.5.4 | High and low expiratory tidal volume alarm
2.28.5.5 | Minute volume, high/low alarm
2.28.5.6 | Inspired oxygen (Fi02), high/low alarm
2.28.5.7 | PEEP, low alarm
2.28.5.8 | Leakage / Check breathing circuit alarm or equivalent
2.28.5.9 | Alarm profile must be user definable or default to patient appropriate settings
2.28.5.10 | The ventilator must have a mute alarm function, active only for a limited time. No more than 2 minutes
2.28.5.11 | It must not be possible to cancel the visual alarm until the problem has been rectified.
2.28.6 | All components of the ventilator and breathing system in contact with the patient’s airway must be latex free.
2.28.7 | All components of the ventilator and breathing system in contact with the patient’s airway must be steam autoclavable up to 134ᴼ C
2.29 | Integrated Gas Monitoring
2.29.1 | Oxygen Monitoring
2.29.1.1 | The system must display both inspired and end-tidal oxygen concentration.
2.29.1.2 | The oxygen analyser must have user adjustable low and high alarms – visual and audible. The low alarm must automatically default to no less than 21%.
2.29.1.3 | The system must be accurate (+/- 2% in 21-100%) range, and have a rapid response time (≤ 500ms).
2.29.1.4 | The unit must sound a continuous alarm if the oxygen pressure falls below 190 kPa, then all nitrous oxide must be automatically cut from the patient circuit.
2.29.1.5 | Bidder to give details of the oxygen analyser
2.29.1.6 | The lifespan of the oxygen analyser must be guaranteed for at least 10 years.
2.29.1.7 | A system using replaceable oxygen fuel cells will not be accepted
2.29.2 | Anaesthetic vapour, carbon dioxide and nitrous oxide monitoring:
2.29.2.1 | Anaesthetic vapour, N2O monitoring and capnography must be integrated into the anaesthetic workstation.
2.29.2.2 | The monitoring must be of the side-stream type.
2.29.2.3 | Sampling: 100 - 200 ml/min
2.29.2.4 | Sampled gases must be internally returned to the breathing system.
2.29.2.5 | There must be no water build-up in the measuring line.
2.29.2.6 | CO2 must be presented as a numerical display of inspired and end-tidal concentration and a capnography waveform.The displayed units must be user-selectable: kPa or mmHg.
2.29.2.7 | The monitoring system must have automatic identification of anaesthetic vapours.
2.29.2.8 | The system must have a graphical display indicating the type of anaesthetic vapour/s currently in use and the concentration thereof. The colour coding must be as per international colour coding for anaesthetic vapours.
2.29.2.9 | The anaestheric vapour must be displayed as vol % and must include inspiratory and expiratory concentrations.
2.29.2.10 | The monitoring system must be able to simultaneously detect, identify and display mixed anaesthetic vapours.
2.29.2.11 | High and low anaesthetic vapour alarms must be present and these must be user-definable.
2.29.2.12 | There must be automatic calculation of minimum alveolar concentration.
2.29.2.13 | The system must be able to calculate anaesthetic gas utilization in ml per case
2.29.2.14 | The unit must have an indication of the level of vapour in the vaporiser
2.29.2.15 | Integrated vapour analysis must deliver an audible "low vapour " alarm and a critical alarm for an empty tank
2.30 | Circle System
2.30.1 | The circle absorber system must be integrated as part of the ventilation system.
2.30.2 | The circle absorber system must have an effective system to eliminate excess water created by the chemical reactions either by heating or water traps.
2.30.3 | The circle absorber system must be autoclavable at 134ºC.
2.30.4 | The circle absorber system must be supplied with a non-disposable CO2 absorber canister
2.30.5 | The absorber canister may be reusable/refillable
2.30.6 | It must be easy to remove and replace the absorber canister while the workstation is in operation.
2.30.7 | Misalignment of the canister must be readily apparent to the user.
2.30.8 | The system must have a circle bypass facility while changing/refilling the absorber canister.
2.30.9 | The breathing circuit must contain an integrated adjustable pressure limiting (APL) valve with a minimum range of 0-70 cm H2O.
2.30.10 | The APL valve must have a tactile or electronic indicator when a value of between 20-30 cm H2O is exceeded.
2.30.11 | The circle absorber system must be latex free.
2.30.12 | The bag-to-ventilator switching mechanism must be clearly labeled and user-friendly. Mechanical or electronic
2.31 | Patient Physiological Monitoring
2.31.1 | The patient monitor must be modular and suitable for use for neonatal, paediatric and adult patients.
2.31.2 | The following paramaters must be monitored:
2.31.2.1 | - ECG: 3-lead and 5-lead. Must be capable of displaying 2 channels simultaneously.
2.31.2.2 | - ST-segment and arrhythmia analysis.
2.31.2.3 | - NIBP
2.31.2.4 | - Oxygen saturation (SpO2)
2.31.2.5 | - Plethysmography/respiratory rate
2.31.2.6 | - Temperature x 2 (minimum)
2.31.2.7 | - Invasive pressure x 4 (minimum)
2.31.3 | The patient monitor must be fully integrated with the workstation.
2.31.4 | Display must be no less than 400 mm diagonally.
2.31.5 | Minimum 6 waveforms must be able to be displayed simultaneously.
2.31.6 | Full colour display.
2.31.7 | The monitor must be capable of displaying trends for at least 24 hours in tabular and graphical format.
2.31.8 | The trend resolution must be user-adjustable and have a minimum resolution of 1 min.
2.31.9 | The front panel must be easy to clean and disinfect.
2.31.10 | The display must be flicker free and easy to read.
2.31.11 | Brightness control must be possible.
2.31.12 | In addition to the waveform display section, the monitor must have alphanumeric display of parameter identification, measured values, alarm and status messages.
2.31.13 | Audible and visual alarms must be available for all measured parameters.
2.31.14 | Default alarm limits if the limits are not set by the user
2.31.15 | Alarm indication: Visual and audible with mute
2.31.16 | Adjustable audio alarm volume
2.31.17 | Low battery: Alarm at least 10 minutes before battery power is exhausted
2.31.18 | It must be possible to mute the audible alarm for a short period (no longer than 2 minutes), but no override of the visual alarm must be possible.
2.31.19 | It must be possible to suspend the alarms of certain parameters during clinical procedures, but the alarm must reactivate automatically after a suitable time delay. Any silencing or disabling of alarms must be clearly indicated on the monitor.
2.31.20 | Input impedance of the system > 10 MΩ.
2.31.21 | Maximum allowable leakage current to the patient must be less than 10 µA.
2.31.22 | The monitor must be protected against defibrillator and electrosurgical interference.
2.31.23 | The monitor must be of low power consumption, and must not require fan or other forced ventilation.
2.31.24 | All cables and sockets must be colour coded for easy recognition, and must conform to international standards
2.32 | ECG
2.32.1 | A full set (adult and paediatric) of 3/5 lead ECG cables must be included in the offer.
2.32.2 | The ECG amplifier must employ a differential input with high noise suppression.
2.32.3 | Sweep speed: at least 25 to 50 mm/sec.
2.32.4 | Freeze facility for ECG.
2.32.5 | The input impedance must be > 10 MΩ.
2.32.6 | The heart rate measuring range: 20 to 300 bpm, with an accuracy < 5 bpm.
2.32.7 | An 50 Hz filter must be provided.
2.32.8 | Alarm limits must be adjustable in the following range:
2.32.9 | Low: 40 to 150
2.32.10 | High: 70 to 250.
2.33 | Non Invasive Blood Pressure
2.33.1 | Must be suitable for adults, paeds and neonates
2.33.2 | The unit must use the oscillometric method of measurement.
2.33.3 | Modes: Manual and automatic
2.33.4 | Reusable neonate, paediatric, adult and obese cuffs must be supplied
2.33.5 | The cuffs must be robust, interchangeable and easy to clean.
2.33.6 | Must display systolic, diastolic and mean pressures for each measurement
2.33.7 | Adult: Systolic: 30 - 270 mm Hg, Diastolic: 20 - 240 mm Hg, MAP: 20 - 250 mm Hg
2.33.8 | Paeds: Systolic: 30 - 160 mm Hg, Diastolic: 20 - 120, MAP: 20 - 120 mm Hg
2.33.9 | Values must be displayed numerically
2.33.10 | Auto zero when unit is switched on
2.33.11 | The interval between measurements must be selectable from 1 – 120 minutes.
2.33.12 | Adjustable high and low alarm limits for systolic, diastolic, mean arterial pressure and heart rate.
2.33.13 | Over pressure protection: Automatic deflation at 300 mm Hg adult, 150 mm Hg neonate
2.33.14 | Average measurement time: approximately 30 s
2.33.15 | If measurement exceeds two minutes the cuff must deflate, an alarm must sound and likely faults indicated.
2.33.16 | The air hose must be at least two metres long.
2.34 | Oxygen Saturation.
2.34.1 | Continuous monitoring of oxygen saturation levels by means of the photometric principle.
2.34.2 | Must be suitable for adults, paeds and neonates
2.34.3 | The unit must continuously display the SaO2 % waveform
2.34.4 | Oxygen saturation (SaO2 %): Range of measurement: 0 - 100%
2.34.5 | The accuracy must be better than 2,5 % (70 - 100%) and 5% (50 - 70%).
2.34.6 | Pulse rate (beats/min): Range of measurement: 20 - 300 bpm
2.34.7 | The pulse rate accuracy must be better than ± 2 %.
2.34.8 | An audible bleep should sound for each heartbeat, the pitch of which should vary with the oxygen saturation.
2.34.9 | Automatic rejection of motion artefacts
2.35 | SpO2 Sensor:
2.35.1 | It must be the latest technology with clinical research support based on accuracy and real-time monitoring. Please submit the published peer reviewed/researched paper as supporting documents.
2.35.2 | The following reusable sensors must be provided:
2.35.3 | Finger sensor for adults and paeds.
2.35.4 | Earlobe sensor for adults and paeds.
2.35.5 | Finger wrap-around probe for neonates
2.35.6 | The sensor lead must be approximately 3 metres long.
2.36 | Temperature
2.36.1 | Temperature range must be approximately 15º to 45º C.
2.36.2 | Two reusable temperature probes must be included
2.37 | Other system requirements
2.37.3 | The unit must operate from both the 220 V, 50 Hz mains supply and internal rechargeable batteries
2.37.4 | The batteries must be charged while the unit is operating from the mains supply.
2.37.5 | The batteries must automatically take over in the event of a mains power failure.
2.37.6 | The batteries must have the capacity for full operation of the anaesthetic machine, ventilator and monitor for at least 90 minutes with batteries fully charged, under typical operating conditions.
2.37.7 | Graphic battery power indicator
2.37.8 | Recharge time < 8 hrs
2.37.9 | Mains cable at least four metres long with SABS approved 15A three prong plug.
2.37.10 | The unit must have a rail for attaching a suction regulator.
2.38 | Data management:
2.38.1 | The workstation must be capable of downloading data to a USB device or memory card.
2.38.2 | The monitor must be able to communicate with a central station
2.38.3 | A printer must be incorporated in the unit
2.39 | The unit must be handed over in full operating order
3 | The following Accessories and Consumables must be supplied with the unit:
3.1 | Re-breathing bag, reusable, rubber
3.2 | Reservoir bag
3.3 | Pressure limiting valve, adjustable
3.4 | Adult & Paediatric volume sensor re-usable
3.5 | Water trap for multigas monitor, re-usable (10 pcs.)
3.6 | A full set paediatric 3/5 lead ECG cables
3.7 | A full set adult 3/5 lead ECG cables
3.8 | Reusable paediatric cuffs
3.9 | Reusable adult cuffs
3.10 | Reusable neonate cuffs
3.11 | Reusable obese cuffs
3.12 | Air hose
3.13 | Finger sensor for paeds
3.14 | Finger sensor for adults
3.15 | Earlobe sensor for paeds
3.16 | Earlobe sensor for adults
3.17 | Finger wrap-around probe for neonates
3.18 | Sensor lead
3.19 | Two reusable temperature probes
3.20 | Vaporizer key-fillers (Sevo, & Iso)
3.21 | The unit must be supplied with two vaporizers (1x sevoflurane, & 1x Iso)
3.22 | Anaesthesia circuit (Adult, single/double limb design) Disposable. 1.8m (Complete with Y-piece, Elbow, Catheter mount, Rebreathing bag)
3.23 | Anaesthesia circuit (Paediatric, single/double limb design) Disposable. 1.8m (Complete with Y-piece, Elbow, Catheter mount, Rebreathing bag)
3.24 | Anaesthesia circuit (neonatal, single/double limb design) Disposable. 1.8m (Complete with Y-piece, Elbow, Catheter mount, Rebreathing bag)
3.25 | Re-breathing circuit complete
3.26 | Non-re-breathing circuit complete
3.27 | Flow sensor (pack of 2)
3.28 | Spirometry line
3.29 | Sample line (pack of 10)
3.30 | Face masks: sizes 00, 0, 1, 2, 3, 4, 5.
3.31 | Bacterial/viral filter
3.32 | Y connector
3.33 | Manual ventilation hose
3.34 | Fresh gas tubing
3.35 | Seal for back-up cylinder connection on the unit
3.36 | Scavenger tubing (as per institution's scavenging outlet)
3.37 | Inspiratory and expiratory unidirectional valves
3.38 | Side stream CO2 kit (adult/paed)
3.39 | Disposable Adult Spirometry Set complete with spirometry tube, sample line and flow sensor, 3 m, box of 20 pieces
3.40 | Disposable Paediatric Spirometry Set complete with spirometry tube, sample line and flow sensor, 3 m, box of 20 pieces
3.41 | Gas sample tube with luer lock connection, 3m, disposable (10 pcs)
3.42 | Suction Tubing Set inc connector
3.43 | Resuscitator Adult, 1500mL, autoclavable, silicone
3.44 | Resuscitator Paediatric, 550mL, autoclavable, silicone
3.45 | Tubing set Paediatric for circle systems
4 | Conformity Compliance (Please attach certificate)
4.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
4.2 | The unit must comply with electrical safety and performance standard such as IEC 60601-2-13 for Anaesthetic systems, attached certification
4.2.1 | Please provide unique ref number of the IEC ceriticate:
4.3 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
4.3.1 | Please provide unique ref number of the ISO 13485 ceriticate:
4.4 | Model on offer must comply with EU Medical Device Regulation (EU) 2017/745 or equvalent from SAHPRA recognized jurisdictions . Attach a copy of certification
4.4.1 | Please provide unique ref number of the ceriticate:
4.5 | Hazardous Substance Act: If this type of equipment / apparatus appears on the schedule of Hazardous substances issued by the Directorate: Health Technology of the Department of Health, a license in terms of the Act on Hazardous Substances (Act 15/1973) must be submitted with the bid / quotation documents. The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
4.6 | All electrical/electronic medical equipment must be licensed by Radiation Control, where applicable a copy of the license must be submitted The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
4.6.1 | Radiation Control License number:
4.7 | The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade (rubber to with 2 screws) type and it must be a minimum length of three (3) meters. NB the mains cable of the unit tendered for must be SABS colour coded.
4.8 | The equipment quoted for must be protected against electromagnetic interference IEC 60601-1-2
4.9 | The unit must comply with electrical safety and performance standard such as IEC 60601-2-49 for Patient Monitoring, attached certification
4.9.1 | Please provide unique ref number of the IEC ceriticate:
5 | COMPREHENSIVE MAINTENANCE
5.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) should be quoted for, as per the SCC document
5.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries, Flow sensors and Oxygen sensor), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
5.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
5.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries, Flow sensors and Oxygen sensor), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
5.3 | The bidder to provide the OEM recommended service interval of the unit:
5.4 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
6 | General
6.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
6.2 | The unit on offer must be the latest technology.
6.2.1 | Initial Manufacturing year of Model:
| ANAESTHESIA DELIVERY SYSTEM: ADVANCED with integrated physiological monitor, Replacement Accessories | Complies Yes/No
RT2-01-06a | A full set paediatric 3/5 lead ECG cable
RT2-01-06b | A full set adult 3/5 lead ECG cable
RT2-01-06c | Reusable paediatric cuff
RT2-01-06d | Reusable adult cuff
RT2-01-06e | Reusable neonate cuff
RT2-01-06f | Reusable obese cuff
RT2-01-06g | Air hose
RT2-01-06h | Finger sensor for paeds
RT2-01-06i | Finger sensor for adults
RT2-01-06j | Earlobe sensor for paeds
RT2-01-06k | Earlobe sensor for adults
RT2-01-06l | Finger wrap-around probe for neonates
RT2-01-06m | Sensor lead
RT2-01-06n | Reusable temperature probe
UPGRADES
RT2-01-06o | AIMS (Anaesthesia information management system)
RT2-01-06p | Auxilliary Delivery System
RT2-01-06q | Neuromuscular transmission monitoring
RT2-01-06r | Metabolic monitoring
RT2-01-06s | Depth of anaesthesia monitoring system
RT2-01-06t | Invasive cardiac output monitoring
RT2-01-06u | Non-invasive cardiac output monitoring
RT2-01-06v | Cerebral oximetry

## Anaesthetic Monitor
RT2-01-07 | ANAESTHESIA PHYSIOLOGICAL MONITOR | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
1 | Description
1.1 | A medical device that tracks a patient's vital signs and physiological parameters during anesthesia care.
1a | Product Offered (Monitor)
1a.1 | Make:
1a.2 | Model:
1a.3 | Initial Manufacturing year of Model:
1a.4 | Radiation Control License number:
2 | Technical Specification
2.1 | The patient monitor must be modular and suitable for use for neonatal, paediatric and adult patients.
2.2 | The following paramaters must be monitored:
2.2.1 | - ECG: 3-lead and 5-lead. Must be capable of displaying 2 channels simultaneously.
2.2.2 | - ST-segment and arrhythmia analysis.
2.2.3 | - NIBP
2.2.4 | - Oxygen saturation (SpO2)
2.2.5 | - Plethysmography/respiratory rate
2.2.6 | - Temperature x 2 (minimum)
2.2.7 | - Invasive pressure x 4 (minimum)
2.3 | The patient monitor must be fully integrated with the workstation.
2.4 | Display must be no less than 400 mm diagonally.
2.5 | Minimum 6 waveforms must be able to be displayed simultaneously.
2.6 | Full colour display.
2.7 | The monitor must be capable of displaying trends for at least 24 hours in tabular and graphical format.
2.8 | The trend resolution must be user-adjustable and have a minimum resolution of 1 min.
2.9 | The front panel must be easy to clean and disinfect.
2.10 | The display must be flicker free and easy to read.
2.11 | Brightness control must be possible.
2.12 | In addition to the waveform display section, the monitor must have alphanumeric display of parameter identification, measured values, alarm and status messages.
2.13 | Audible and visual alarms must be available for all measured parameters.
2.14 | Default alarm limits if the limits are not set by the user
2.15 | Alarm indication: Visual and audible with mute
2.16 | Adjustable audio alarm volume
2.17 | Low battery: Alarm at least 10 minutes before battery power is exhausted
2.18 | It must be possible to mute the audible alarm for a short period (no longer than 2 minutes), but no override of the visual alarm must be possible.
2.19 | It must be possible to suspend the alarms of certain parameters during clinical procedures, but the alarm must reactivate automatically after a suitable time delay. Any silencing or disabling of alarms must be clearly indicated on the monitor.
2.20 | Input impedance of the system > 10 MΩ.
2.21 | Maximum allowable leakage current to the patient must be less than 10 µA.
2.22 | The monitor must be protected against defibrillator and electrosurgical interference.
2.23 | The monitor must be of low power consumption, and must not require fan or other forced ventilation.
2.24 | All cables and sockets must be colour coded for easy recognition, and must conform to international standards
2.25 | ECG
2.25.1 | A full set (adult and paediatric) of 3/5 lead ECG cables must be included in the offer.
2.25.2 | The ECG amplifier must employ a differential input with high noise suppression.
2.25.3 | Sweep speed: at least 25 to 50 mm/sec.
2.25.4 | Freeze facility for ECG.
2.25.5 | The input impedance must be > 10 MΩ.
2.25.6 | The heart rate measuring range: 20 to 300 bpm, with an accuracy < 5 bpm.
2.25.7 | An 50 Hz filter must be provided.
2.25.8 | Alarm limits must be adjustable in the following range:
2.25.8.1 | Low: 40 to 150
2.25.8.2 | High: 70 to 250.
2.26 | Non Invasive Blood Pressure
2.26.1 | Must be suitable for adults, paeds and neonates
2.26.2 | The unit must use the oscillometric method of measurement.
2.26.3 | Modes: Manual and automatic
2.26.4 | Reusable neonate, paediatric, adult and obese cuffs must be supplied
2.26.5 | The cuffs must be robust, interchangeable and easy to clean.
2.26.6 | Must display systolic, diastolic and mean pressures for each measurement
2.26.6.1 | Adult: Systolic: 30 - 270 mm Hg, Diastolic: 20 - 240 mm Hg, MAP: 20 - 250 mm Hg
2.26.6.2 | Paeds: Systolic: 30 - 160 mm Hg, Diastolic: 20 - 120, MAP: 20 - 120 mm Hg
2.26.7 | Values must be displayed numerically
2.26.8 | Auto zero when unit is switched on
2.26.9 | The interval between measurements must be selectable from 1 – 120 minutes.
2.26.10 | Adjustable high and low alarm limits for systolic, diastolic, mean arterial pressure and heart rate.
2.26.11 | Over pressure protection: Automatic deflation at 300 mm Hg adult, 150 mm Hg neonate
2.26.12 | Average measurement time: approximately 30 s
2.26.13 | If measurement exceeds two minutes the cuff must deflate, an alarm must sound and likely faults indicated.
2.26.14 | The air hose must be at least two metres long.
2.27 | Oxygen Saturation.
2.27.1 | Continuous monitoring of oxygen saturation levels by means of the photometric principle.
2.27.2 | Must be suitable for adults, paeds and neonates
2.27.3 | The unit must continuously display the SaO2 % waveform
2.27.4 | Oxygen saturation (SaO2 %): Range of measurement: 0 - 100%
2.27.5 | The accuracy must be better than 2,5 % (70 - 100%) and 5% (50 - 70%).
2.27.6 | Pulse rate (beats/min): Range of measurement: 20 - 300 bpm
2.27.7 | The pulse rate accuracy must be better than ± 2 %.
2.27.8 | An audible bleep should sound for each heartbeat, the pitch of which should vary with the oxygen saturation.
2.27.9 | Automatic rejection of motion artefacts
2.28 | SpO2 Sensor:
2.28.1 | It must be the latest technology with clinical research support based on accuracy and real-time monitoring. Please submit the published peer reviewed/researched paper as supporting documents.
2.28.2 | The following reusable sensors must be provided:
2.28.2.1 | Finger sensor for adults and paeds.
2.28.2.2 | Earlobe sensor for adults and paeds.
2.28.2.3 | Finger wrap-around probe for neonates
2.28.3 | The sensor lead must be approximately 3 metres long.
2.29 | Temperature
2.29.1 | Temperature range must be approximately 15º to 45º C.
2.29.2 | Two reusable temperature probes must be included
3 | The following accessories must be supplied with the unit:
3.1 | A full set paediatric 3/5 lead ECG cables
3.2 | A full set adult 3/5 lead ECG cables
3.3 | Reusable paediatric cuffs
3.4 | Reusable adult cuffs
3.5 | Reusable neonate cuffs
3.6 | Reusable obese cuffs
3.7 | Air hose
3.8 | Finger sensor for paeds
3.9 | Finger sensor for adults
3.10 | Earlobe sensor for paeds
3.11 | Earlobe sensor for adults
3.12 | Finger wrap-around probe for neonates
3.13 | Sensor lead
3.14 | Two reusable temperature probes
4 | Conformity Compliance (Please attach certificate)
4.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
4.2 | The unit must comply with electrical safety and performance standard such as IEC 60601-2-49 for Patient Monitoring, attached certification
4.2.1 | Please provide unique ref number of the IEC ceriticate:
4.3 | OEM must comply and certified on ISO 9001 quality standards, attach proof of compliance
4.3.1 | Please provide unique ref number of the ISO 9001 ceriticate:
4.4 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
4.4.1 | Please provide unique ref number of the ISO 13485 ceriticate:
4.5 | Model on offer must comply with EU Medical Device Regulation (EU) 2017/745 or equvalent from SAHPRA recognized jurisdictions . Attach a copy of certification
4.5.1 | Please provide unique ref number of the ceriticate:
4.6 | Hazardous Substance Act: If this type of equipment / apparatus appears on the schedule of Hazardous substances issued by the Directorate: Health Technology of the Department of Health, a license in terms of the Act on Hazardous Substances (Act 15/1973) must be submitted with the bid / quotation documents. The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
4.7 | All electrical/electronic medical equipment must be licensed by Radiation Control, where applicable a copy of the license must be submitted The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
4.7.1 | Radiation Control License number:
4.8 | The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade (rubber to with 2 screws) type and it must be a minimum length of three (3) meters. NB the mains cable of the unit tendered for must be SABS colour coded.
4.9 | The equipment quoted for must be protected against electromagnetic interference IEC 60601-1-2
5 | COMPREHENSIVE MAINTENANCE
5.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) should be quoted for, as per the SCC document
5.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries, Flow sensors and Oxygen sensor), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
5.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
5.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries, Flow sensors and Oxygen sensor), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
5.3 | The bidder to provide the OEM recommended service interval of the unit:
5.4 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
6 | General
6.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
6.2 | The unit on offer must be the latest technology.
6.2.1 | Initial Manufacturing year of Model:
| ANAESTHESIA DELIVERY SYSTEM: ADVANCED with integrated physiological monitor, Replacement Accessories | Complies Yes/No
RT2-01-07a | A full set paediatric 3/5 lead ECG cable
RT2-01-07b | A full set adult 3/5 lead ECG cable
RT2-01-07c | Reusable paediatric cuff
RT2-01-07d | Reusable adult cuff
RT2-01-07e | Reusable neonate cuff
RT2-01-07f | Reusable obese cuff
RT2-01-07g | Air hose
RT2-01-07h | Finger sensor for paeds
RT2-01-07i | Finger sensor for adults
RT2-01-07j | Earlobe sensor for paeds
RT2-01-07k | Earlobe sensor for adults
RT2-01-07l | Finger wrap-around probe for neonates
RT2-01-07m | Sensor lead
RT2-01-07n | Reusable temperature probe

## Anaesthetic, MRI compatible
RT2-01-08 | ANAESTHETIC DELIVERY SYSTEM MRI Compatible, WITHOUT Physiological Monitor | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
1 | DESCRIPTION
1.1 | An advanced anaesthesia delivery system comprising a completely integrated electronic anaesthesia workstation and can be used in MRI scanner rooms with magnets of 1.5 and 3 Tesla with a field strength between 300-400 gauss.
1a | Product Offered (Anaesthetic Machine)
1a.1 | Make:
1a.2 | Model:
1a.3 | Initial Manufacturing year of Model:
1a.4 | Radiation Control License number:
2 | TECHNICAL SPECIFICATION
2.1 | The unit must include electronicaly integrated oxygen, nitrous oxide and air delivery and monitoring. The Unit must have a physically integrated magnetic field strength monitor (Gauss Alarm)
2.2 | Ventilation and anaesthesia vapour monitoring must be fully integrated within the unit.
2.3 | The anaesthesia unit must be able to communicate/interface with a physiological monitor
2.4 | The anaesthesia system must comprise:
2.5 | Anaesthesia system suitable for neonatal (down to 500g), paediatric and adult anaesthesia.
2.6 | Ventilator suitable for ventilating neonatal (down to 500g), paediatric, adult and morbid/obese patients
3 | Multigas analyser: Must be available on Anaesthetic machine
3.1 | Oxygen
3.2 | Nitrous oxide
3.3 | Carbon dioxide
3.4 | Anaesthetic vapour monitoring:
3.4.1 | Halothane
3.4.2 | Isoflurane
3.4.3 | Sevoflurane
3.4.4 | Desflurane
3.5 | Complete and automated funtional test/check requiring minimal user intervention or disassembly.
3.6 | The functional test/check must provide the measurements and values for automatic leakage and compliance.
3.7 | Automatic calibration of sensors for anaesthetic gases and vapours must form part of the self test and require no user intervention.
3.8 | It must be possible to override the functional test/check facility. The system must display the override status
3.9 | The unit must provide conventional, low-flow (down to 200 ml/min) anaesthesia.
3.10 | The unit must return analysed gas back to the circle system
3.11 | A safety oxygen function is required for the delivery of oxygen through the anaesthesia unit in the event of an electronic failure of the gas delivery system.
3.12 | The unit must be supplied with a minimum of 1 storage drawer
3.13 | The unit must be supplied with a minimum of 3 additional 220 V power sockets. These sockets must be excluded from the battery backup circuit in case of a power failure
3.14 | The unit must have a working area or a retractable shelf. Width: not less than 300 mm. Depth: not less than 200 mm
3.15 | The unit must have a light to enable the anaesthetist to work in a darkened theatre.
4 | Trolley Made out of MRI compatable material
4.1 | The unit must be mounted on a strong , robust and compact trolley manufactured from corrosion free material. The material must be MRI compatible.
4.2 | The trolley must have no sharp corners and be easy to clean.
4.3 | The trolley must have at least four antistatic castors of diameter 125 mm or greater.
4.4 | Trolley must have two lockable castors or a central locking brake.
4.5 | The unit must accommodate an oxygen cylinder. Cylinder size according to user requirement. Size to be compatible with current state vendor of oxygen and cylinder must be MRI compatible (made out of alluminium).
4.6 | Trolley must have accessible handles for ease of movement
4.7 | The trolley, with all components mounted, must be stable without toppling over when placed on a 30° slope.
4.8 | The unit must be supplied with a castor guard system, preventing the entaglement and trapping of cables within the wheels.
4.9 | The trolley must have a mounting system suitable for a physiologial monitor.
4.10 | Pipeline and Cylinder Gas Supplies All made out of MRI compatible material. The anaesthesia system must be designed to use external piped medical gasses (O2, N2O, medical air), vacuum (scavenging) from medical service outlet sockets.
4.11 | Connection hoses for above must be provided and must be fitted with Heyer type non-interchangeable probes for the existing medical gas outlets.
4.12 | Hoses must be fitted with Heyer type non-interchangeable probes for the existing medical gas outlets.
4.13 | The medical air hose connection must include a water trap
4.14 | A safety pressure relief valve of 60 - 80 cm H2O must be provided to protect the system from excessive pressure .
4.15 | The unit must function safely and correctly within the pipeline supply pressure range of: 270 - 600 kPa
4.16 | The system must preferentially utilise the piped oxygen supply over the cylinder supply, if pipeline pressure is > 270 kPa
5 | Connection hoses must be Flexible and at least 6 metres long:
5.1 | Antistatic
5.2 | Colour coded as follows:
5.3 | O2: white
5.4 | N2O: blue
5.5 | Vacuum: yellow
5.6 | Medical air: black
5.7 | Scavenge low vacuum: orange or yellow
5.8 | Where relevant, fitted with outlet probes complying with SANS 1409:2014
5.9 | Permanently bonded to the outlet probe and the machine connector
5.10 | Hoses for gases above atmospheric pressure, must be designed for pressures of at least 1 000 kPa.
5.11 | Vacuum hoses may not collapse or kink at 90 kPa below atmospheric pressure, when subjected to 180° bend of inside radius equal to the outside diameter of the hose.
5.12 | All connection hoses must be fed to a central manifold.
5.13 | All high pressure gas and vacuum pipe connections (other than hoses) and all pipes connected to pressure regulators and pressure gauges must be stainless steel (18/8 grade) pipes that are oil-free and colour coded as for hoses (see above) over their entire lengths
5.14 | An in-line non-return valve must be fitted to each gas pipeline on the high pressure (inlet) line
5.15 | An in-line non-return valve must be fitted to each gas cylinder yoke
5.16 | A replaceable gas filter, to block particles greater than 65 µm, must be fitted on the upstream side of the non-return valve
5.17 | Each gas pipeline inlet must have a pressure gauge or electronic display of pressure. The pressure range (415 kPa ± 75 kPa) must be clearly indicated in colour
5.18 | Must be fitted with a gas cylinder yoke to accommodate one oxygen gas cylinder with a pin index system, MRI compatible
5.19 | Each gas cylinder yoke must have a pressure gauge or electronic display of pressure
5.20 | The pressure gauges or electronic displays of pressure for both pipeline & cylinder supplies, calibrated in kPa for each line, must be labelled and visible from the front of the unit.
5.21 | O2 and N2O regulators must be provided complying with CKS 605:1987.
5.22 | A pressure regulator complying with CKS 605 must be supplied for each gas inlet and must reduce the gas pressure to 415 kPa without freezing up, creeping, chattering or whistling
5.23 | Outlet sockets and probes must comply with SANS 1409:2014.
5.24 | Sockets with hoses must be colour-coded as per SANS 1409:2014:
5.25 | Oxygen: White background with black letters
5.26 | Nitrous oxide: Blue background with white letters
5.27 | Medical air: One half black with white letters, the other half white with black letters
5.28 | Vacuum: Yellow background with black letters
5.29 | Continuous display, visible from the front of the machine, of both pipeline and cylinder pressures even during power failure
5.30 | Pipeline and cylinder pressures must be differentially displayed, and clearly indicated
5.31 | Safety pressure relief valves must protect the entire pneumatic system from excessive input pressures from pipeline and cylinder supplies. Safety relief valve for all gases must be supplied and set at 600kPa
5.32 | When a cylinder is detached from the pin-indexed yoke no gas leak must occur at the yoke.
5.33 | Airway pressure must be monitored and displayed
5.34 | Scavenging-MRI compatible material
5.35 | Scavenging-MRI compatible material
5.36 | The unit must have a scavenging system with a hose at least 5 metres long with the suitable connector as per the institutions scavenging installation.
5.37 | Air-break device to prevent negative pressure on the breathing system
5.38 | If an active sytem is supplied a pressure indicator must be present
5.39 | The scavenging outlet must be a 30mm ISO connector.
6 | Patient Gas Delivery System
6.1 | Patient gas delivery must be electronically controlled\Rotameters
6.2 | Must display the delivered patient gas flow using an ISO colour standard for each gas.
6.3 | The unit must be able to deliver a total patient gas flow of 200 ml/min to 18 l/min (Minimum). The unit must have the ability to deliver at least 10 l/min of oxygen. These must be incremental adjustments
6.4 | Gas supply must conform to:
6.5 | O2 minimum flow range of 200 ml/min to 10 l/min. Accuracy ± 5%
6.6 | medical air minimum flow range 0 to 10 l/min. Accuracy ± 5%
6.7 | nitrous oxide minimum flow range 0 to 10 l/min. Accuracy ± 5%
6.8 | Flowmeter sequence to be according to standards approved by the South Africal Society of Anaesthesiologists - Oxygen on the right
6.9 | The delivery system must prevent simultaneous delivery of two anaesthetic vapours.
6.10 | A separate common gas outlet to accept a standard 22 mm outside diameter and 15 mm internal diameter.
7 | Hypoxic guard
7.1 | The unit must include a safety mechanism that ensures a minimum concentration of 25% oxygen in any oxygen/nitrous oxide gas mixture across all flow rates.
7.2 | The unit must include a safety mechanism that ensures a minimum concentration of 21% oxygen in any oxygen/air gas mixture across all flow rates.
7.3 | Accidental simultaneous delivery of N2O and Air into the breathing system, must either be prevented by means of an interlock switch or clearly indicated so that the anaesthetist can rectify the error. Delivery indicated purely by mechanical rotameters will not achieve the requirements of this clause.
7.4 | The unit must include a high-flow oxygen flush mechanism with a flow rate of at least 35 l/min
7.5 | Oxygen flush must be non-lockable, recessed and self closing
7.6 | This system must bypass the flowmeters and vaporisers
8 | Breathing System
8.1 | Total circuit volume, including absorber, must not exceed 2,7 litres
8.2 | With APL valve 5 - 75 cm H2O
8.3 | With ventilator to manual switch. This switch when directed to the ventilator must isolate the APL valve from the circle.
8.4 | Breathing system must be fully autoclavable to 134°C
8.5 | There must be a mechanism to remove water from the breathing system
8.6 | Breathing system must be latex free
8.7 | Bag-to-ventilator switch automatically begins in the ventilator position
8.8 | Reusable corrugated breathing circuit to be supplied
8.9 | Bidders to supply the breathing circuit as required by the user
9 | Integrated Ventilator
9.1 | Ventilator suitable for ventilating adults, paediatric and neonates (> 500g) patients without changing ventilator components (e.g. bellows)
9.2 | The ventilator must be electronically controlled
9.3 | The ventilation technology offered may be turbine, volume reflector, piston or ascending bellows
9.4 | The minimum basic modes to be supplied:
9.5 | Manual/Spontaneous Ventilation
9.6 | Volume Control Ventilation
9.7 | Pressure Control Ventilation
9.8 | Pressure Support with Mandatory Apnoea back-up ventilation
9.9 | Volume Controlled Ventilation + SIMV + Pressure Support
9.10 | Pressure Control Ventilation + SIMV + Presssure Support
9.11 | CPAP
9.12 | Additional modes of ventilation may be offered and quoted for separately
9.13 | The ventilator must be integrated with the workstation and user friendly
10 | The ventilator must meet the following minimum criteria:
10.1 | Frequency: minimum range of 5 – 30 bpm.
10.2 | Inspiratory time: minimum range of 0.2 – 10.0 seconds.
10.3 | Minimum Inspiration/Expiration Ratio: 4:1 – 1:4.
10.4 | Adjustable inspiratory pause/plateau time (volume control ventilation): 0-30% of inspiratory time
10.5 | Peak Inspiratory Pressure: minimum range of: PEEP + 5 to 30 cm H2O.
10.6 | Positive End Expiratory Pressure (PEEP): Integrated, electronically controlled. Minimum range: 0 – 30 cm H2O.
10.7 | Tidal Volume: 20ml – 1500ml per breath. Minimum accuracy: 15%
10.8 | Peak Inspiratory Flow: At least 120 /pm
10.9 | The ventilator must be capable of flow or pressure triggering.
10.10 | Adjustable Inspiratory Rise time/Ramp with a range of 0.1 - 2 sec
10.11 | Compliance compensation/Tidal volume compensation/Compressible volume compensation or equivalent.
10.12 | Patient gas flow compensation must be automatic so that alterations in patient gas flow does not affect delivered tidal volumes
10.13 | It must be possible to set CPAP in manual/spontaneous mode within a range of: 1 to 10 cm H2O
10.14 | In event of power or technical failure, manual ventilation must be possible.
11 | Display:
11.1 | All displays must be in English
11.2 | Colour display must include numerical and graphic display of parameters measured
11.3 | Graphic display to feature at least 3 waveforms: - pressure, flow and volume.
11.4 | Must also display flow-volume and pressure-volume loops with a reference loop
12 | Alarms : The ventilator must have the following user adjustable alarms
12.1 | Apnoea alarm
12.2 | High airway pressure alarm
12.3 | Continuous pressure alarm
12.4 | High and low expiratory tidal volume alarm
12.5 | Minute volume, high/low alarm
12.6 | Inspired oxygen (Fi02), high/low alarm
12.7 | PEEP, low alarm
12.8 | Leakage / Check breathing circuit alarm or equivalent
12.9 | Alarm profile must be user definable or default to patient appropriate settings
12.10 | The ventilator must have a mute alarm function, active only for a limited time (maximum 2 minutes)
12.11 | It must not be possible to cancel the visual alarm until the problem has been rectified.
12.12 | All components of the ventilator and breathing system in contact with the patient’s airway must be latex free.
12.13 | All components of the ventilator and breathing system in contact with the patient’s airway must be steam autoclavable up to 134ᴼ C
Integrated Gas Monitoring
13 | Oxygen Monitoring
13.1 | The system must display both inspired and end-tidal oxygen concentration.
13.2 | The oxygen analyser must have user adjustable low and high alarms – visual and audible. The low alarm must automatically default to no less than 21%.
13.3 | The system must be accurate (+/- 2% in 21-100%) range, and have a rapid response time (≤ 500ms).
13.4 | The unit must sound a continuous alarm if the oxygen pressure falls below 190 kPa, then all nitrous oxide must be automatically cut from the patient circuit.
14 | Anaesthetic vapour, carbon dioxide and nitrous oxide monitoring:
14.1 | Anaesthetic vapour, N2O monitoring and capnography must be electronically integrated into the anaesthetic workstation.
14.2 | The monitoring must be of the side-stream type.
14.3 | Sampling: 100 - 200 ml/min
14.4 | Sampled gases must be internally returned to the breathing system.
14.5 | CO2 must be presented as a numerical display of inspired and end-tidal concentration and a capnography waveform.The displayed units must be user-selectable: kPa or mmHg.
14.6 | The monitoring system must have automatic identification of anaesthetic vapours.
14.7 | The system must have a graphical display indicating the type of anaesthetic vapour/s currently in use and the concentration thereof. The colour coding must be as per international colour coding for anaesthetic vapours.
14.8 | The anaestheric vapour must be displayed as vol % and must include inspiratory and expiratory concentrations.
14.9 | The monitoring system must be able to simultaneously detect, identify and display mixed anaesthetic vapours.
14.10 | High and low anaesthetic vapour alarms must be present and these must be user-definable.
14.11 | There must be automatic calculation of minimum alveolar concentration.
14.12 | The system must be able to calculate anaesthetic gas utilization in ml per case
14.13 | The unit must have an indication of the level of vapour in the vaporiser
14.14 | Integrated vapour analysis must deliver an audible "low vapour " alarm and a critical alarm for an empty tank
15 | Circle System
15.1 | The circle absorber system must be integrated as part of the ventilation system.
15.2 | The circle absorber system must have an effective system to eliminate excess water created by the chemical reactions either by heating or water traps.
15.3 | The circle absorber system must be autoclavable at 134ºC.
15.4 | The circle absorber system must be supplied with a carbon dioxide absorber canister.
15.5 | The absorber canister may be reusable/refillable
15.6 | It must be easy to remove and replace the absorber canister while the workstation is in operation.
15.7 | Misalignment of the canister must be readily apparent to the user.
15.8 | The system must have a circle bypass facility while changing/refilling the absorber canister.
15.9 | The breathing circuit must contain an integrated adjustable pressure limiting (APL) valve with a minimum range of 0-70 cm H2O.
15.10 | The APL valve must have a tactile or electronic indicator when a value of between 20-30 cm H2O is exceeded.
15.11 | The circle absorber system must be latex free.
15.12 | The bag-to-ventilator switching mechanism must be clearly labeled and user-friendly. Mechanical or electronic
16 | General
16.1 | The unit must operate from both the 220 V, 50 Hz mains supply and internal rechargeable batteries
16.2 | The batteries must be charged while the unit is operating from the mains supply.
16.3 | The batteries must automatically take over in the event of a mains power failure.
16.4 | The batteries must have the capacity for full operation of the anaesthetic machine and ventilator for at least 90 minutes with batteries fully charged, under typical operating conditions.
16.5 | Graphic battery power indicator
16.6 | Recharge time < 8 hrs
16.7 | Mains cable at least six metres long with SABS approved 15A three prong plug.
16.8 | The unit must have a rail for attaching a suction regulator.
17 | Data management:
17.1 | The workstation must be capable of downloading data to a USB device or memory card.
17.2 | The Unit must have an intergated field strength monitor with acoustic, audible and visual alarms for Advisory,Caution and Warning.
17.3 | A laser printer must be incorporated in the unit and be MRI compatible
18 | Accessories and Consumables
18.1 | A full set of basic accessories for adult/paediatric/neonate must be supplied with the system,appropriate for use in an MRI suite
18.2 | List all the consumables required for full operation of the equipment, with prices separately
18.3 | All cables to be supplied as per user requirement
19 | STANDARD ACCESSORIES
19.1 | The unit must be supplied with a starter pack for full functionality upon commissioning.
19.2 | The unit must be delivered with a set of adult (A), paediatric (P) and neonatal (N) consumables:
19.2.1 | Patient circuit (A, P, N) - box of 10 each
19.2.2 | Filters (box of 10)
19.2.3 | Inflatable reusable face mask (A, P, N) - 3 full set of each
19.2.4 | Spirometery (box of 10 at least)
19.2.5 | Sampling line (box of 10 atleast)
19.2.6 | Flow sensors (box of 10 taleast)
19.2.7 | CO2 absorber canister disposable reusable x1 and disposable (box of 10 atleast)
19.2.8 | Vaporizer key-fillers (Sevo, & other choice below)
19.3 | The unit must be supplied with 2x MRI compatible vaporizers (one must sevoflurane, the other must be of the choice of the buying institution)
19.4 | Water traps (box of 10 at least)
21 | Vaporisers\MRI compatible(Tec 5 or 7)
21.1 | Vaporisers must be tilt, pressure and temperature compensated
21.2 | Mechanical, electronic or injection technology vaporisers are acceptable
21.3 | Filling mechanism:
21.3.1 | Vaporiser filling mechanism must be compatible with the current state approved vendor of inhalational volatile vapours.
21.3.2 | Each vaporiser must be supplied with 2 compatible filling adapters as per user requirement
21.3.3 | Minimum accuracy of the supplied volatile vapour delivery system must be at least 20% of set value
21.3.4 | Visible level indicator
22 | Auxilliary Delivery System
22.1 | An external auxilliary flowmeter suitable for delivery of oxygen via a face mask, supplied with a 6mm tapered connection. Flow range 0-10 l/min and above
23 | Trending:
23.1 | Trend information and event recording for:
23.1.1 | peak pressure
23.1.2 | plateau pressure
23.1.3 | PEEP
23.1.4 | expiratory tidal volume
23.1.5 | respiratory rate
23.2 | This must be provided in graphical and tabular formats.
23.3 | The trend and event information must cover at least the last 24 hours of operation.
23.4 | Resolution must be adjustable, with a minimum resolution of 1 minute intervals.
23.5 | The adjustment of the resolution must not affect the display of parameters which were sampled at different times to the resolution setting in such a way as to display a null reading, when one was in fact obtained.
24 | Software
24.1 | Any software for upgrades may be offered
25 | Suction Regulator
25.1 | A suction regulator capable of -50 kPa (minimum)
25.2 | The suction regulator must have a built-in diaphragm with a 0 – 100 kPa vacuum gauge and an “on-off” switch.
25.3 | It must aspirate 1 litre water in 45 seconds at a pressure of at least - 50 kPa
25.4 | A collection jar calibrated to at least 1 500 ml must be included.
25.5 | The collection jar must accommodate a disposable liner, as per user requirement
25 | Conformity Compliance (Please attach certificate)
25.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
25.2 | The unit must comply with electrical safety and performance standard such as IEC 60601-2-13 for Anaesthetic systems, attached certification
25.2.1 | Please provide unique ref number of the IEC ceriticate:
25.4 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
25.4.1 | Please provide unique ref number of the ISO 13485 ceriticate:
25.5 | Model on offer must comply with EU Medical Device Regulation (EU) 2017/745 or equvalent from SAHPRA recognized jurisdictions . Attach a copy of certification
25.5.1 | Please provide unique ref number of the ceriticate:
25.6 | Hazardous Substance Act: If this type of equipment / apparatus appears on the schedule of Hazardous substances issued by the Directorate: Health Technology of the Department of Health, a license in terms of the Act on Hazardous Substances (Act 15/1973) must be submitted with the bid / quotation documents. The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
25.7 | All electrical/electronic medical equipment must be licensed by Radiation Control, where applicable a copy of the license must be submitted The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
25.7.1 | Radiation Control License number:
25.8 | The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade (rubber to with 2 screws) type and it must be a minimum length of three (3) meters. NB the mains cable of the unit tendered for must be SABS colour coded.
25.9 | The equipment quoted for must be protected against electromagnetic interference IEC 60601-1-2
26 | COMPREHENSIVE MAINTENANCE
26.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) should be quoted for, as per the SCC document
26.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries, Flow sensors and Oxygen sensor), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
26.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
26.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries, Flow sensors and Oxygen sensor), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
26.3 | The bidder to provide the OEM recommended service interval of the unit:
26.4 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
27 | General
27.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
27.2 | The unit on offer must be the latest technology.
27.2.1 | Initial Manufacturing year of Model:
| ANAESTHETIC DELIVERY SYSTEM MRI Compatible, WITHOUT Physiological monitor- Accessories and Consumables | Complies Yes/No
RT2-01-08a | Anaesthesia circuit (Adult, single/double limb design) Disposable. 1.8m (Complete with Y-piece, Elbow, Catheter mount, Rebreathing bag)
RT2-01-08b | Anaesthesia circuit (Paediatric, single/double limb design) Disposable. 1.8m (Complete with Y-piece, Elbow, Catheter mount, Rebreathing bag)
RT2-01-08c | Anaesthesia circuit (neonatal, single/double limb design) Disposable. 1.8m (Complete with Y-piece, Elbow, Catheter mount, Rebreathing bag)
RT2-01-08d | Re-breathing circuit complete
RT2-01-08e | Non-re-breathing circuit complete
RT2-01-08f | Flow sensor (pack of 2)
RT2-01-08g | Spirometry line
RT2-01-08h | Sample line (pack of 10)
RT2-01-08i | Face masks: sizes (00,0,1,2,3,4, 5)- End-user to specify size
RT2-01-08j | Bacterial/viral filter
RT2-01-08k | Y connector
RT2-01-08l | Manual ventilation hose
RT2-01-08m | Fresh gas tubing
RT2-01-08n | Seal for back-up cylinder connection on the unit
RT2-01-08o | Scavenger tubing (as per institution's scavenging outlet)
RT2-01-08p | Inspiratory and expiratory unidirectional valves
RT2-01-08q | Side stream CO2 kit (adult/paed)
RT2-01-08r | Disposable Adult Spirometry Set complete with spirometry tube, sample line and flow sensor, 3 m, box of 20 pieces
RT2-01-08s | Disposable Paediatric Spirometry Set complete with spirometry tube, sample line and flow sensor, 3 m, box of 20 pieces
RT2-01-08t | Gas sample tube with luer lock connection, 3m, disposable (10 pcs)
RT2-01-08u | Suction Tubing Set inc connector
RT2-01-08v | Resuscitator Adult, 1500mL, autoclavable, silicone
RT2-01-08w | Resuscitator Paediatric, 550mL, autoclavable, silicone
RT2-01-08x | Tubing set Paediatric for circle systems
RT2-01-08y | MRI compatible Sevoflurane vaporizer
RT2-01-08z | MRI compatible Isoflurane vaporizer
RT2-01-08zz | MRI compatible Desflurane vaporizer

## Anaesthetic, MRI comp + Monitor
RT2-01-09 | ANAESTHETIC DELIVERY SYSTEM WITH PHYSIOLOGICAL MONITOR MRI COMPATABLE | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
1 | DESCRIPTION
1.1 | An advanced anaesthesia delivery system comprising a completely integrated electronic anaesthesia workstation and can be used in MRI scanner rooms with magnets of 1.5 and 3 Tesla with a field strength between 300-400 gauss with compatible wireless patient monitor.
1a | Product Offered (Anaesthetic Machine)
1a.1 | Make:
1a.2 | Model:
1a.3 | Initial Manufacturing year of Model:
1a.4 | Radiation Control License number:
1b | Product Offered (Monitor)
1b.1 | Make:
1b.2 | Model:
1b.3 | Initial Manufacturing year of Model:
1b.4 | Radiation Control License number:
2 | TECHNICAL SPECIFICATION
2.1 | The unit must include electronicaly integrated oxygen, nitrous oxide and air delivery and monitoring.The Unit must have a physically integrated magnetic field strength monitor(Gauss Alarm)
2.2 | Ventilation and anaesthesia vapour monitoring must be fully integrated within the unit.
2.3 | The anaesthesia unit must be able to communicate/interface with a physiological monitor
2.4 | The anaesthesia system must comprise:
2.5 | Anaesthesia system suitable for neonatal (down to 500g), paediatric and adult anaesthesia.
2.6 | Ventilator suitable for ventilating neonatal (down to 500g), paediatric, adult and morbid/obese patients
3 | Multigas analyser: Must be available either on Anaesthetic machine or Physiological monitor
3.1 | Oxygen
3.2 | Nitrous oxide
3.3 | Carbon dioxide
3.4 | Anaesthetic vapour monitoring:
3.4.1 | Halothane
3.4.2 | Isoflurane
3.4.3 | Sevoflurane
3.4.4 | Desflurane
3.5 | Complete and automated functional test/check requiring no user intervention or disassembly.
3.6 | The functional test/check must provide the measurements and values for automatic leakage and compliance.
3.7 | Automatic calibration of sensors for anaesthetic gases and vapours must form part of the self test and require no user intervention.
3.8 | It must be possible to override the functional test/check facility. The system must display the override status
3.9 | The unit must provide conventional, low-flow (down to 200 ml/min) anaesthesia.
3.10 | The unit must return analysed gas back to the circle system
3.11 | A safety oxygen function is required for the delivery of oxygen through the anaesthesia unit in the event of an electronic failure of the gas delivery system.
3.12 | The unit must be supplied with a minimum of 1 storage drawer
3.13 | The unit must be supplied with a minimum of 3 additional 220 V power sockets. These sockets must be excluded from the battery backup circuit in case of a power failure
3.14 | The unit must have a working area or a retractable shelf. Width: not less than 300 mm. Depth: not less than 200 mm
3.15 | The unit must have a light to enable the anaesthetist to work in a darkened theatre.
4 | Trolley Made out of MRI compatable material
4.1 | The unit must be mounted on a strong , robust and compact trolley manufactured from corrosion free material. The material must be MRI compatible.
4.2 | The trolley must have no sharp corners and be easy to clean.
4.3 | The trolley must have at least four antistatic castors of diameter 125 mm or greater.
4.4 | Trolley must have two lockable castors or a central locking brake.
4.5 | The unit must accommodate an oxygen cylinder. Cylinder size according to user requirement. Size to be compatible with current state vendor of oxygen and cylinder must be MRI compatible (made out of alluminium).
4.6 | Trolley must have accessible handles for ease of movement
4.7 | The trolley, with all components mounted, must be stable without toppling over when placed on a 30° slope.
4.8 | The unit must be supplied with a castor guard system, preventing the entaglement and trapping of cables within the wheels.
4.9 | The trolley must have a mounting system suitable for a physiologial monitor.
4.10 | Pipeline and Cylinder Gas Supplies All made out of MRI compatible material. The anaesthesia system must be designed to use external piped medical gasses (O2, N2O, medical air), vacuum (scavenging) from medical service outlet sockets.
4.11 | Connection hoses for above must be provided and must be fitted with Heyer type non-interchangeable probes for the existing medical gas outlets.
4.12 | Hoses must be fitted with Heyer type non-interchangeable probes for the existing medical gas outlets.
4.13 | The medical air hose connection must include a water trap
4.14 | A safety pressure relief valve of 60 - 80 cm H2O must be provided to protect the system from excessive pressure .
4.15 | The unit must function safely and correctly within the pipeline supply pressure range of: 270 - 600 kPa
4.16 | The system must preferentially utilise the piped oxygen supply over the cylinder supply, if pipeline pressure is > 270 kPa
5 | Connection hoses must be Flexible and at least 6 metres long:
5.1 | Antistatic
5.2 | Colour coded as follows:
5.2.1 | O2: white
5.2.2 | N2O: blue
5.2.3 | Vacuum: yellow
5.2.4 | Medical air: black
5.2.5 | Scavenge low vacuum: orange or yellow
5.3 | Where relevant, fitted with outlet probes complying with SANS 1409:2014
5.4 | Permanently bonded to the outlet probe and the machine connector
5.5 | Hoses for gases above atmospheric pressure, must be designed for pressures of at least 1 000 kPa.
5.6 | Vacuum hoses may not collapse or kink at 90 kPa below atmospheric pressure, when subjected to 180° bend of inside radius equal to the outside diameter of the hose.
5.7 | All connection hoses must be fed to a central manifold.
5.8 | All high pressure gas and vacuum pipe connections (other than hoses) and all pipes connected to pressure regulators and pressure gauges must be stainless steel (18/8 grade) pipes that are oil-free and colour coded as for hoses (see above) over their entire lengths
5.9 | An in-line non-return valve must be fitted to each gas pipeline on the high pressure (inlet) line
5.10 | An in-line non-return valve must be fitted to each gas cylinder yoke
5.11 | A replaceable gas filter, to block particles greater than 65 µm, must be fitted on the upstream side of the non-return valve
5.12 | Each gas pipeline inlet must have a pressure gauge or electronic display of pressure. The pressure range (415 kPa ± 75 kPa) must be clearly indicated in colour
5.13 | Must be fitted with a gas cylinder yoke to accommodate one oxygen gas cylinder with a pin index system, MRI compatible
5.14 | Each gas cylinder yoke must have a pressure gauge or electronic display of pressure
5.15 | The pressure gauges or electronic displays of pressure for both pipeline & cylinder supplies, calibrated in kPa for each line, must be labelled and visible from the front of the unit.
5.16 | O2 and N2O regulators must be provided complying with CKS 605:1987.
5.17 | A pressure regulator complying with CKS 605 must be supplied for each gas inlet and must reduce the gas pressure to 415 kPa without freezing up, creeping, chattering or whistling
5.18 | Outlet sockets and probes must comply with SANS 1409:2014.
5.19 | Sockets with hoses must be colour-coded as per SANS 1409:2014:
5.20 | Oxygen: White background with black letters
5.21 | Nitrous oxide: Blue background with white letters
5.22 | Medical air: One half black with white letters, the other half white with black letters
5.23 | Vacuum: Yellow background with black letters
5.24 | Continuous display, visible from the front of the machine, of both pipeline and cylinder pressures even during power failure
5.25 | Pipeline and cylinder pressures must be differentially displayed, and clearly indicated
5.26 | Safety pressure relief valves must protect the entire pneumatic system from excessive input pressures from pipeline and cylinder supplies. Safety relief valve for all gases must be supplied and set at 600kPa
5.27 | When a cylinder is detached from the pin-indexed yoke no gas leak must occur at the yoke.
5.28 | Airway pressure must be monitored and displayed
5.29 | Scavenging-MRI compatible material
5.30 | Scavenging-MRI compatible material
5.31 | The unit must have a scavenging system with a hose at least 5 metres long with the suitable connector as per the institutions scavenging installation.
5.32 | Air-break device to prevent negative pressure on the breathing system
5.33 | If an active sytem is supplied a pressure indicator must be present
5.34 | The scavenging outlet must be a 30mm ISO connector.
6 | Patient Gas Delivery System
6.1 | Patient gas delivery must be electronically controlled\Rotameters
6.2 | Must display the delivered patient gas flow using an ISO colour standard for each gas.
6.3 | The unit must be able to deliver a total patient gas flow of 200 ml/min to 18 l/min (Minimum). The unit must have the ability to deliver at least 10 l/min of oxygen. These must be incremental adjustments
6.4 | Gas supply must conform to:
6.5 | O2 minimum flow range of 200 ml/min to 10 l/min. Accuracy ± 5%
6.6 | medical air minimum flow range 0 to 10 l/min. Accuracy ± 5%
6.7 | nitrous oxide minimum flow range 0 to 10 l/min. Accuracy ± 5%
6.8 | Flowmeter sequence to be according to standards approved by the South Africal Society of Anaesthesiologists - Oxygen on the right
6.9 | The delivery system must prevent simultaneous delivery of two anaesthetic vapours.
6.10 | A separate common gas outlet to accept a standard 22 mm outside diameter and 15 mm internal diameter.
7 | Hypoxic guard
7.1 | The unit must include a safety mechanism that ensures a minimum concentration of 25% oxygen in any oxygen/nitrous oxide gas mixture across all flow rates.
7.2 | The unit must include a safety mechanism that ensures a minimum concentration of 21% oxygen in any oxygen/air gas mixture across all flow rates.
7.3 | Accidental simultaneous delivery of N2O and Air into the breathing system, must either be prevented by means of an interlock switch or clearly indicated so that the anaesthetist can rectify the error. Delivery indicated purely by mechanical rotameters will not achieve the requirements of this clause.
7.4 | The unit must include a high-flow oxygen flush mechanism with a flow rate of at least 35 l/min
7.5 | Oxygen flush must be non-lockable, recessed and self closing
7.6 | This system must bypass the flowmeters and vaporisers
8 | Breathing System
8.1 | Total circuit volume, including absorber, must not exceed 2,7 litres
8.2 | With APL valve 5 - 75 cm H2O
8.3 | With ventilator to manual switch. This switch when directed to the ventilator must isolate the APL valve from the circle.
8.4 | Breathing system must be fully autoclavable to 134°C
8.5 | There must be a mechanism to remove water from the breathing system
8.6 | Breathing system must be latex free
8.7 | Bag-to-ventilator switch automatically begins in the ventilator position
8.8 | Reusable corrugated breathing circuit to be supplied
8.9 | Bidders to supply the breathing circuit as required by the user
9 | Integrated Ventilator
9.1 | Ventilator suitable for ventilating adults, paediatric and neonates (> 500g) patients without changing ventilator components (e.g. bellows)
9.2 | The ventilator must be electronically controlled
9.3 | The ventilation technology offered may be turbine, volume reflector, piston or ascending bellows
9.4 | The minimum basic modes to be supplied:
9.4.1 | Manual/Spontaneous Ventilation
9.4.2 | Volume Control Ventilation
9.4.3 | Pressure Control Ventilation
9.4.4 | Pressure Support with Mandatory Apnoea back-up ventilation
9.4.5 | Volume Controlled Ventilation + SIMV + Pressure Support
9.4.6 | Pressure Control Ventilation + SIMV + Presssure Support
9.4.7 | CPAP
9.5 | Additional modes of ventilation may be offered and quoted for separately
9.6 | The ventilator must be integrated with the workstation and user friendly
10 | The ventilator must meet the following minimum criteria:
10.1 | Frequency: minimum range of 5 – 30 bpm.
10.2 | Inspiratory time: minimum range of 0.2 – 10.0 seconds.
10.3 | Minimum Inspiration/Expiration Ratio: 4:1 – 1:4.
10.4 | Adjustable inspiratory pause/plateau time (volume control ventilation): 0-30% of inspiratory time
10.5 | Peak Inspiratory Pressure: minimum range of: PEEP + 5 to 30 cm H2O.
10.6 | Positive End Expiratory Pressure (PEEP): Integrated, electronically controlled. Minimum range: 0 – 30 cm H2O.
10.7 | Tidal Volume: 20ml – 1500ml per breath. Minimum accuracy: 15%
10.8 | Peak Inspiratory Flow: At least 120 /pm
10.9 | The ventilator must be capable of flow or pressure triggering.
10.10 | Adjustable Inspiratory Rise time/Ramp with a range of 0.1 - 2 sec
10.11 | Compliance compensation/Tidal volume compensation/Compressible volume compensation or equivalent.
10.12 | Patient gas flow compensation must be automatic so that alterations in patient gas flow does not affect delivered tidal volumes
10.13 | It must be possible to set CPAP in manual/spontaneous mode within a range of: 1 to 10 cm H2O
10.14 | In event of power or technical failure, manual ventilation must be possible.
11 | Display:
11.1 | All displays must be in English
11.2 | Colour display must include numerical and graphic display of parameters measured
11.3 | Graphic display to feature at least 3 waveforms: - pressure, flow and volume.
11.4 | Must also display flow-volume and pressure-volume loops with a reference loop
12 | Alarms : The ventilator must have the following user adjustable alarms
12.1 | Apnoea alarm
12.2 | High airway pressure alarm
12.3 | Continuous pressure alarm
12.4 | High and low expiratory tidal volume alarm
12.5 | Minute volume, high/low alarm
12.6 | Inspired oxygen (Fi02), high/low alarm
12.7 | PEEP, low alarm
12.8 | Leakage / Check breathing circuit alarm or equivalent
12.9 | Alarm profile must be user definable or default to patient appropriate settings
12.10 | The ventilator must have a mute alarm function, active only for a limited time (Maximum 2minutes)
12.11 | It must not be possible to cancel the visual alarm until the problem has been rectified.
12.12 | All components of the ventilator and breathing system in contact with the patient’s airway must be latex free.
12.13 | All components of the ventilator and breathing system in contact with the patient’s airway must be steam autoclavable up to 134ᴼ C
Integrated Gas Monitoring
13 | Oxygen Monitoring
13.1 | The system must display both inspired and end-tidal oxygen concentration.
13.2 | The oxygen analyser must have user adjustable low and high alarms – visual and audible. The low alarm must automatically default to no less than 21%.
13.3 | The system must be accurate (+/- 2% in 21-100%) range, and have a rapid response time (≤ 500ms).
13.4 | The unit must sound a continuous alarm if the oxygen pressure falls below 190 kPa, then all nitrous oxide must be automatically cut from the patient circuit.
14 | Anaesthetic vapour, carbon dioxide and nitrous oxide monitoring:
14.1 | Anaesthetic vapour, N2O monitoring and capnography must be electronically integrated into the anaesthetic workstation.
14.2 | The monitoring must be of the side-stream type.
14.3 | Sampling: 100 - 200 ml/min
14.4 | Sampled gases must be internally returned to the breathing system.
14.5 | CO2 must be presented as a numerical display of inspired and end-tidal concentration and a capnography waveform.The displayed units must be user-selectable: kPa or mmHg.
14.6 | The monitoring system must have automatic identification of anaesthetic vapours.
14.7 | The system must have a graphical display indicating the type of anaesthetic vapour/s currently in use and the concentration thereof. The colour coding must be as per international colour coding for anaesthetic vapours.
14.8 | The anaestheric vapour must be displayed as vol % and must include inspiratory and expiratory concentrations.
14.9 | The monitoring system must be able to simultaneously detect, identify and display mixed anaesthetic vapours.
14.10 | High and low anaesthetic vapour alarms must be present and these must be user-definable.
14.11 | There must be automatic calculation of minimum alveolar concentration.
14.12 | The system must be able to calculate anaesthetic gas utilization in ml per case
14.13 | The unit must have an indication of the level of vapour in the vaporiser
14.14 | Integrated vapour analysis must deliver an audible "low vapour " alarm and a critical alarm for an empty tank
15 | Circle System
15.1 | The circle absorber system must be integrated as part of the ventilation system.
15.2 | The circle absorber system must have an effective system to eliminate excess water created by the chemical reactions either by heating or water traps.
15.3 | The circle absorber system must be autoclavable at 134ºC.
15.4 | The circle absorber system must be supplied with a carbon dioxide absorber canister.
15.5 | The absorber canister may be reusable/refillable
15.6 | It must be easy to remove and replace the absorber canister while the workstation is in operation.
15.7 | Misalignment of the canister must be readily apparent to the user.
15.8 | The system must have a circle bypass facility while changing/refilling the absorber canister.
15.9 | The breathing circuit must contain an integrated adjustable pressure limiting (APL) valve with a minimum range of 0-70 cm H2O.
15.10 | The APL valve must have a tactile or electronic indicator when a value of between 20-30 cm H2O is exceeded.
15.11 | The circle absorber system must be latex free.
15.12 | The bag-to-ventilator switching mechanism must be clearly labeled and user-friendly. Mechanical or electronic
16 | General
16.1 | The unit must operate from both the 220 V, 50 Hz mains supply and internal rechargeable batteries
16.2 | The batteries must be charged while the unit is operating from the mains supply.
16.3 | The batteries must automatically take over in the event of a mains power failure.
16.4 | The batteries must have the capacity for full operation of the anaesthetic machine and ventilator for at least 90 minutes with batteries fully charged, under typical operating conditions.
16.5 | Graphic battery power indicator
16.6 | Recharge time < 8 hrs
16.7 | Mains cable at least six metres long with SABS approved 15A three prong plug.
16.8 | The unit must have a rail for attaching a suction regulator.
17 | Data management:
17.1 | The workstation must be capable of downloading data to a USB device or memory card.
17.2 | The Unit must have an intergated field strength monitor with acoustic, audible and visual alarms for Advisory,Caution and Warning.
17.3 | A laser printer must be incorporated in the unit and be MRI compatible
18 | Patient Physiological Monitor, MRI and Anaesthetic workstation compatible
18.1 | The patient monitor must be compatible with the Anaesthesia work station and the MRI equipment, wireless components where possible.
18.2 | The patient monitor must be modular or pre-configured and suitable for use for neonatal, paediatric and adult patients.
18.3 | The following paramaters must be monitored:
18.3.1 | - ECG: 3-lead and 5-lead. Must be capable of displaying 2 channels simultaneously.
18.3.2 | - ST-segment and arrhythmia analysis.
18.3.3 | - NIBP
18.3.4 | - Oxygen saturation (SpO2)
18.3.5 | - Plethysmography/respiratory rate
18.3.6 | - 2 x Temperature
18.3.7 | - At least 2 x Invasive pressure
18.4 | Display must be no less than 15" diagonally.
18.5 | Minimum 6 waveforms must be able to be displayed simultaneously.
18.6 | Full colour display.
18.7 | The front panel must be easy to clean and disinfect.
18.8 | The display must be flicker free and easy to read.
18.9 | Brightness control must be possible.
18.10 | In addition to the waveform display section, the monitor must have alpha-numeric display of parameter identification, measured values, alarm and status messages.
18.11 | Audible and visual alarms must be available for all measured parameters.
18.12 | Default alarm limits if the limits are not set by the user
18.13 | Alarm indication: Visual and audible with mute
18.14 | Adjustable audio alarm volume
18.15 | Low battery: Alarm at least 10 minutes before battery power is exhausted
18.16 | It must be possible to mute the audible alarm for a short period, but no override of the visual alarm must be possible.
18.17 | It must be possible to suspend the alarms of certain parameters during clinical procedures, but the alarm must reactivate automatically after a suitable time delay. Any silencing or disabling of alarms must be clearly indicated on the monitor.
18.18 | Input impedance of the system > 10 MΩ.
18.19 | Maximum allowable leakage current to the patient must be less than 10 µA.
18.20 | The monitor must be protected against defibrillator and electrosurgical interference.
18.21 | The monitor must be of low power consumption, and must not require fan or other forced ventilation.
18.22 | All cables and sockets must be colour coded for easy recognition, and must conform to international standards
18.23 | Display of trends for at least 8 hours in tabular and graphical format.
18.24 | The trend resolution must be user-adjustable and have a minimum resolution of 1 min.
19 | ECG
19.1 | A full set of 3/5 lead ECG cables must be included in the offer.
19.2 | The ECG amplifier must employ a differential input with high noise suppression.
19.3 | Sweep speed: 25 to 50 mm/sec.
19.4 | Freeze facility for ECG.
19.5 | The input impedance must be > 10 MΩ.
19.6 | The heart rate measuring range: 20 to 300 bpm, with an accuracy < 5 bpm.
19.7 | A 50 Hz filter must be provided.
19.8 | Alarm limits must be adjustable in the following range:
19.8.1 | Low: 40 to 150
19.8.2 | High: 70 to 250.
20 | Non Invasive Blood Pressure
20.1 | Must be suitable for adults, paeds and neonates
20.2 | The unit must use the oscillometric method of measurement.
20.3 | Modes: Manual and automatic
20.4 | Reusable neonate, paediatric, and adult cuffs must be supplied
20.5 | The cuffs must be robust, interchangeable and easy to clean.
20.6 | Must display systolic, diastolic and mean pressures for each measurement
20.7 | Adult: Systolic: 30 - 270 mm Hg, Diastolic: 20 - 240 mm Hg, MAP: 20 - 250 mm Hg
20.8 | Paeds: Systolic: 30 - 160 mm Hg, Diastolic: 20 - 120, MAP: 20 - 120 mm Hg
20.9 | Values must be displayed numerically
20.1 | Auto zero when unit is switched on
20.11 | The interval between measurements must be selectable from 1 – 120 minutes.
20.12 | Adjustable high and low alarm limits for systolic, diastolic, mean arterial pressure and heart rate.
20.13 | Over pressure protection: Automatic deflation at 300 mm Hg adult, 150 mm Hg neonate
20.14 | Average measurement time: approximately 30 s
20.15 | If measurement exceeds two minutes the cuff must deflate, an alarm must sound and likely faults indicated.
20.16 | The NIBP hose must be at least two metres long.
21 | Oxygen Saturation.
21.1 | Continuous monitoring of oxygen saturation levels by means of the photometric principle and wireless technology
21.2 | Must be suitable for adults, paeds and neonates
21.3 | The unit must continuously display the SpO2 % waveform
21.4 | Oxygen saturation (SpO2 %): Range of measurement: 0 - 100%
21.5 | The accuracy must be better than 2,5 % (70 - 100%) and 5% (50 - 70%).
21.6 | Pulse rate (beats/min): Range of measurement: 20 - 300 bpm
21.7 | The pulse rate accuracy must be better than ± 2 %.
21.8 | An audible bleep should sound for each heartbeat, the pitch of which should vary with the oxygen saturation.
21.9 | Automatic rejection of variuos artefacts
22 | SpO2 Sensor:
22.1 | Wireless Telemetry
22.2 | It must be the latest technology with clinical research support based on accuracy and real-time monitoring. Please submit the published peer reviewed/researched paper as supporting documents.
22.3 | The following reusable sensors must be provided:
22.4 | Finger sensor for adults and paeds.
22.5 | Earlobe sensor for adults and paeds.
22.6 | Finger wrap-around probe for neonates
22.7 | The sensor lead must be approximately 3 metres long if not wireless
23 | Temperature
23.1 | Temperature range must be approximately 15º C to 45º C.
23.2 | Two reusable temperature probes must be included
24 | Accessories and Consumables
24.1 | A full set of basic accessories for adult/paediatric/neonate must be supplied with the system,appropriate for use in an MRI suite
25 | STANDARD ACCESSORIES
25.1 | The unit must be supplied with a starter pack for full functionality upon commissioning.
25.2 | The unit must be delivered with a set of adult (A), paediatric (P) and neonatal (N) accessories:
25.2.0 | Invasive pressure cables compatible with the transducer system available at the institution
25.2.1 | SpO2 probe complete (A, P, N) x2
25.2.2 | ECG cable complete (A, P, N) x2
25.2.3 | Temperature probe x2
25.2.4 | NIBP hose and cuff (A, P, N) x2
25.2.5 | Patient circuit (A, P, N) - box of 10 each
25.2.6 | Filters (box of 10)
25.2.7 | Inflatable reusable face mask (A, P, N) - 3 full set of each
25.2.8 | Spirometery (box of 10 at least)
25.2.9 | Sampling line (box of 10 atleast)
25.2.10 | Flow sensors (box of 10 taleast)
25.2.11 | CO2 absorber canister disposable reusable x1 and disposable (box of 10 atleast)
25.2.12 | Vaporizer key-fillers (Sevo, & other choice below)
25.3 | The unit must be supplied with two vaporizers (one must sevoflurane, the other must be of the choice of the buying institution)
25.4 | Water traps (box of 10 at least)
27 | Vaporisers\MRI compatible(Tec 5 or 7)
27.1 | Vaporisers must be tilt, pressure and temperature compensated
27.2 | Mechanical, electronic or injection technology vaporisers are acceptable
27.3 | Filling mechanism:
27.3.1 | Vaporiser filling mechanism must be compatible with the current state approved vendor of inhalational volatile vapours.
27.3.2 | Each vaporiser must be supplied with 2 compatible filling adapters as per user requirement
27.3.3 | Minimum accuracy of the supplied volatile vapour delivery system must be at least 20% of set value
27.3.4 | Visible level indicator
28 | Auxilliary Delivery System
28.1 | An external auxilliary flowmeter suitable for delivery of oxygen via a face mask, supplied with a 6mm tapered connection. Flow range 0-10 l/min and above
29 | Trending:
29.1 | Trend information and event recording for:
29.1.1 | peak pressure
29.1.2 | plateau pressure
29.1.3 | PEEP
29.1.4 | expiratory tidal volume
29.1.5 | respiratory rate
29.2 | This must be provided in graphical and tabular formats.
29.3 | The trend and event information must cover at least the last 24 hours of operation.
29.4 | Resolution must be adjustable, with a minimum resolution of 1 minute intervals.
29.5 | The adjustment of the resolution must not affect the display of parameters which were sampled at different times to the resolution setting in such a way as to display a null reading, when one was in fact obtained.
30 | Software
30.1 | Any software for upgrades may be offered
31 | Suction Regulator
31.1 | A suction regulator capable of -50 kPa (minimum)
31.2 | The suction regulator must have a built-in diaphragm with a 0 – 100 kPa vacuum gauge and an “on-off” switch.
31.3 | It must aspirate 1 litre water in 45 seconds at a pressure of at least - 50 kPa
31.4 | A collection jar calibrated to at least 1 500 ml must be included.
31.5 | The collection jar must accommodate a disposable liner, as per user requirement
32 | Conformity Compliance (Please attach certificate)
32.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
32.2 | The unit must comply with electrical safety and performance standard such as IEC 60601-2-13 for Anaesthetic systems, attached certification
32.2.1 | Please provide unique ref number of the IEC ceriticate:
32.3 | The unit must comply with electrical safety and performance standard such as IEC 60601-2-49 for Patient Monitoring, attached certification
32.3.1 | Please provide unique ref number of the IEC ceriticate:
32.5 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
32.5.1 | Please provide unique ref number of the ISO 13485 ceriticate:
32.6 | Model on offer must comply with EU Medical Device Regulation (EU) 2017/745 or equvalent from SAHPRA recognized jurisdictions . Attach a copy of certification
32.6.1 | Please provide unique ref number of the ceriticate:
32.7 | Hazardous Substance Act: If this type of equipment / apparatus appears on the schedule of Hazardous substances issued by the Directorate: Health Technology of the Department of Health, a license in terms of the Act on Hazardous Substances (Act 15/1973) must be submitted with the bid / quotation documents. The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
32.8 | All electrical/electronic medical equipment must be licensed by Radiation Control, where applicable a copy of the license must be submitted The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
32.8.1 | Radiation Control License number:
32.9 | The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade (rubber to with 2 screws) type and it must be a minimum length of three (3) meters. NB the mains cable of the unit tendered for must be SABS colour coded.
32.10 | The equipment quoted for must be protected against electromagnetic interference IEC 60601-1-2
33 | COMPREHENSIVE MAINTENANCE
33.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) should be quoted for, as per the SCC document
33.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries, Flow sensors and Oxygen sensor), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
33.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
33.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries, Flow sensors and Oxygen sensor), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
33.3 | The bidder to provide the OEM recommended service interval of the unit:
33.4 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
34 | General
34.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
34.2 | The unit on offer must be the latest technology.
34.2.1 | Initial Manufacturing year of Model:
| ACCESSORIES FOR ANAESTHETIC DELIVERY SYSTEM WITH PHYSIOLOGICAL MONITOR MRI COMPATABLE- Accessories and Consumables | Complies Yes/No
RT2-01-09a | A full set paediatric 3/5 lead ECG cable
RT2-01-09b | A full set adult 3/5 lead ECG cable
RT2-01-09c | Reusable paediatric cuff
RT2-01-09d | Reusable adult cuff
RT2-01-09e | Reusable neonate cuff
RT2-01-09f | Reusable obese cuff
RT2-01-09g | Air hose
RT2-01-09h | Finger sensor for paeds
RT2-01-09i | Finger sensor for adults
RT2-01-09j | Earlobe sensor for paeds
RT2-01-09k | Earlobe sensor for adults
RT2-01-09l | Finger wrap-around probe for neonates
RT2-01-09m | Sensor lead
RT2-01-09n | Reusable temperature probe

## Monitor, MRI Compatible
RT2-01-10 | PATIENT MONITOR COMPATIBLE TO MRI, with Multigas Analysis | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
1 | DESCRIPTION
1.1 | A specialised medical device designed for use in Magnetic Resonance Imaging (MRI) environments to safely monitor patients during MRI procedures, even when using anaesthetisia or other medical gasses.
1a | Product Offered:
1a.1 | Make:
1a.2 | Model:
1a.3 | Initial Manufacturing year of Model:
1a.4 | Radiation Control License number:
2 | TECHNICAL SPECIFICATION
2.1 | The offered patient physiological monitor must be compatible with various MRI-Compatible Anaesthetisia workstations.
2.2 | wireless components where possible. The monitor must comprise of a detouchable module for continuous monitoring of the patient during the MRI procedure.
2.3 | The patient monitor must be modular or pre-configured and suitable for use for neonatal, paediatric and adult patients.
2.4 | The following paramaters must be monitored:
2.4.1 | - ECG: 3-lead and 5-lead. Must be capable of displaying 2 channels simultaneously.
2.4.2 | - ST-segment and arrhythmia analysis.
2.4.3 | - NIBP
2.4.4 | - Oxygen saturation (SpO2)
2.4.5 | - Plethysmography/respiratory rate
2.4.6 | - 2 x Temperature
2.4.7 | - At least 2 x Invasive pressure
3 | Multigas analyser
3.1 | Oxygen
3.2 | Nitrous oxide
3.3 | Carbon dioxide
3.4 | Anaesthetic vapour monitoring:
3.4.1 | Halothane
3.4.2 | Isoflurane
3.4.3 | Sevoflurane
3.4.4 | Desflurane
3.5 | Display must be no less than 15" diagonally.
3.6 | Minimum 6 waveforms must be able to be displayed simultaneously.
3.7 | Full colour display.
3.8 | The front panel must be easy to clean and disinfect.
3.9 | The display must be flicker free and easy to read.
3.10 | Brightness control must be possible.
3.11 | In addition to the waveform display section, the monitor must have alpha-numeric display of parameter identification, measured values, alarm and status messages.
3.12 | Audible and visual alarms must be available for all measured parameters.
3.13 | Default alarm limits if the limits are not set by the user
3.14 | Alarm indication: Visual and audible with mute
3.15 | Adjustable audio alarm volume
3.16 | Low battery: Alarm at least 10 minutes before battery power is exhausted
3.17 | It must be possible to mute the audible alarm for a short period, but no override of the visual alarm must be possible.
3.17.1 | It must be possible to suspend the alarms of certain parameters during clinical procedures, but the alarm must reactivate automatically after a suitable time delay. Any silencing or disabling of alarms must be clearly indicated on the monitor.
3.18 | Input impedance of the system > 10 MΩ.
3.19 | Maximum allowable leakage current to the patient must be less than 10 µA.
3.20 | The monitor must be protected against defibrillator and electrosurgical interference.
3.21 | The monitor must be of low power consumption, and must not require fan or other forced ventilation.
3.22 | All cables and sockets must be colour coded for easy recognition, and must conform to international standards
3.23 | Display of trends for at least 72 hours in tabular and graphical format.
3.24 | The trend resolution must be user-adjustable and have a minimum resolution of 1 min.
4 | ECG
4.1 | A full set of 3/5 lead ECG cables must be included in the offer.
4.2 | The ECG amplifier must employ a differential input with high noise suppression.
4.3 | Sweep speed: 25 to 50 mm/sec.
4.4 | Freeze facility for ECG.
4.5 | The input impedance must be > 10 MΩ.
4.6 | The heart rate measuring range: 20 to 300 bpm, with an accuracy < 5 bpm.
4.7 | A 50 Hz filter must be provided.
4.8 | Alarm limits must be adjustable in the following range:
4.8.1 | Low: 40 to 150
4.8.2 | High: 70 to 250.
5 | Non Invasive Blood Pressure
5.1 | Must be suitable for adults, paeds and neonates
5.2 | The unit must use the oscillometric method of measurement.
5.3 | Modes: Manual and automatic
5.4 | Reusable neonate, paediatric, adult and obese cuffs must be supplied
5.5 | The cuffs must be robust, interchangeable and easy to clean.
5.6 | Must display systolic, diastolic and mean pressures for each measurement
5.7 | Adult: Systolic: 30 - 270 mm Hg, Diastolic: 20 - 240 mm Hg, MAP: 20 - 250 mm Hg
5.8 | Paeds: Systolic: 30 - 160 mm Hg, Diastolic: 20 - 120, MAP: 20 - 120 mm Hg
5.9 | Values must be displayed numerically
5.10 | Auto zero when unit is switched on
5.11 | The interval between measurements must be selectable from 1 – 120 minutes.
5.12 | Adjustable high and low alarm limits for systolic, diastolic, mean arterial pressure and heart rate.
5.13 | Over pressure protection: Automatic deflation at 300 mm Hg adult, 150 mm Hg neonate
5.14 | Average measurement time: approximately 30 s
5.15 | If measurement exceeds two minutes the cuff must deflate, an alarm must sound and likely faults indicated.
5.16 | The NIBP hose must be at least two metres long.
6 | Oxygen Saturation.
6.1 | Continuous monitoring of oxygen saturation levels by means of the photometric principle and wireless technology
6.2 | Must be suitable for adults, paeds and neonates
6.3 | The unit must continuously display the SpO2 % waveform
6.4 | Oxygen saturation (SpO2 %): Range of measurement: 0 - 100%
6.5 | The accuracy must be better than 2,5 % (70 - 100%) and 5% (50 - 70%).
6.6 | Pulse rate (beats/min): Range of measurement: 20 - 300 bpm
6.7 | The pulse rate accuracy must be better than ± 2 %.
6.8 | An audible bleep should sound for each heartbeat, the pitch of which should vary with the oxygen saturation.
6.9 | Automatic rejection of variuos artefacts
7 | SpO2 Sensor:
7.1 | Wireless Telemetry
7.2 | It must be the latest technology with clinical research support based on accuracy and real-time monitoring. Please submit the published peer reviewed/researched paper as supporting documents.
7.3 | The following reusable sensors must be provided:
7.4 | Finger sensor for adults and paeds.
7.5 | Earlobe sensor for adults and paeds.
7.6 | Finger wrap-around probe for neonates
7.7 | The sensor lead must be approximately 3 metres long if not wireless
8 | Temperature
8.1 | Temperature range must be approximately 15º to 45º C.
8.2 | Two reusable temperature probes must be included
9 | STANDARD ACCESSORIES
9.1 | The unit must be supplied with a starter pack for full functionality upon commissioning.
9.2 | The unit must be delivered with a set of adult (A), paediatric (P) and neonatal (N) consumables:
9.2.1 | SpO2 probe complete (A, P, N) x2
9.2.2 | ECG cable complete (A, P, N) x2
9.2.3 | Temperature probe x2
9.2.4 | NIBP hose and cuff (A, P, N) x2
11 | Conformity Compliance (Please attach certificate)
11.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
11.2 | The unit must comply with electrical safety and performance standard such as IEC 60601-2-49 for Patient Monitoring, attached certification
11.2.1 | Please provide unique ref number of the IEC ceriticate:
11.3 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
11.3.1 | Please provide unique ref number of the ISO 13485 ceriticate:
11.4 | Model on offer must comply with EU Medical Device Regulation (EU) 2017/745 or equvalent from SAHPRA recognized jurisdictions . Attach a copy of certification
11.4.1 | Please provide unique ref number of the ceriticate:
11.5 | Hazardous Substance Act: If this type of equipment / apparatus appears on the schedule of Hazardous substances issued by the Directorate: Health Technology of the Department of Health, a license in terms of the Act on Hazardous Substances (Act 15/1973) must be submitted with the bid / quotation documents. The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
11.6 | All electrical/electronic medical equipment must be licensed by Radiation Control, where applicable a copy of the license must be submitted The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
11.6.1 | Radiation Control License number:
11.7 | The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade (rubber to with 2 screws) type and it must be a minimum length of three (3) meters. NB the mains cable of the unit tendered for must be SABS colour coded.
11.8 | The equipment quoted for must be protected against electromagnetic interference IEC 60601-1-2
12 | COMPREHENSIVE MAINTENANCE
12.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) should be quoted for, as per the SCC document
12.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
12.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
12.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
12.3 | The bidder to provide the OEM recommended service interval of the unit:
12.4 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
13 | General
13.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
13.2 | The unit on offer must be the latest technology.
13.2.1 | Initial Manufacturing year of Model:
| ACCESSORIES FOR PATIENT MONITOR COMPATABLE TO MRI, with Multigas analysis | Complies Yes/No
RT2-01-10a | A full set paediatric 3/5 lead ECG cable
RT2-01-10b | A full set adult 3/5 lead ECG cable
RT2-01-10c | Reusable paediatric cuff
RT2-01-10d | Reusable adult cuff
RT2-01-10e | Reusable neonate cuff
RT2-01-10f | Reusable obese cuff
RT2-01-10g | Air hose
RT2-01-10h | Finger sensor for paeds
RT2-01-10i | Finger sensor for adults
RT2-01-10j | Earlobe sensor for paeds
RT2-01-10k | Earlobe sensor for adults
RT2-01-10l | Finger wrap-around probe for neonates
RT2-01-10m | Sensor lead
RT2-01-10n | Reusable temperature probe
UPGRADES
RT2-01-10o | Neuromuscular transmission monitoring
RT2-01-10p | Metabolic monitoring
RT2-01-10q | Electroencephalography (EEG) and depth of anaesthesia monitoring system
RT2-01-10r | Invasive cardiac output monitoring
RT2-01-10s | Non-invasive cardiac output monitoring
RT2-01-10t | Cerebral oximetry

## Video Laryngoscope
RT2-01-11 | INTUBATION VIDEO LARYNGOSCOPE | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
1 | DESCRIPTION
1.1 | A medical device used for visualizing the larynx and trachea during intubation procedures.
1a | Product Offered:
1a.1 | Make:
1a.2 | Model:
1a.3 | Initial Manufacturing year of Model:
1a.4 | Radiation Control License number:
2 | TECHNICAL SPECIFICATION
2.1 | The unit must be mounted on a mobile stand.
2.2 | The unit must have facility to be used on all patients ranging from preterm neonatal (from 600g) to morbidly obese adults
2.3 | The handheld unit will be attached to an electronic display to facilitate visualization of the larynx
2.4 | The device should incorporate NTSC colour video or latest technology.
2.5 | Slimline durable handheld intubation scope similar in design to a laryngoscope handle and blade system
3 | The Blade should have:
3.1 | The acutely angled blade must have an angulation of at least 60⁰ and be available for adult and paediatrics sizes
3.2 | An automatic anti fogging mechanism and anti glare technology in order to ensure a clear image is displayed on the display unit during use. This anti fogging mechanism will not require any action on the part of the user to function.
3.3 | An automatic autofocus mechanism in order to ensure a clear image is displayed on the display unit during use.
3.4 | Fully reusable and disposable blades will be accepted.
3.5 | The device should be supplied with a set of at least three(3) blades for neonatals, paediatrics or adults, the end user shall specify the set of preference.
3.6 | The full video laryngoscope blade, meaning both the length of the blade and the blade handle, must be able to be sterilised. The Bidder must clearly state the different sterilization methods that may be performed on the offered laryngoscope type blade and other components. Suitable wall cleaning charts must be provided to ensure appropriate cleaning and unit maintenance of the device.
4 | The offered video intubation system shall consist of the following components which are fully intergrated into the system (not limitted to):
4.1 | Slimline durable handheld intubation scope similar in design to a laryngoscope handle and blade system.
4.2 | Display and Control unit monitor must comprise of at least 7" LED.
4.3 | The video screen monitor must be separate from the laryngoscope blade, must offer user friendly and allow for immediate use when switched on
4.4 | Mobile stand on castor wheels (at least two of them with brakes)
4.5 | IV pole mount for semi-permanent fixation
4.6 | Mounting adaptor to secure the monitor onto the mobile stand or IV pole mount
4.7 | Suitable blade holder capable of storing a minimum of 2 blades. The holder must be compatible for attachment to the mobile stand and IV pole mount
4.8 | The offered system shall be fully operational on delivery and will not require any additional accessories or consumable items (at least one month of consumable items must be provided) in order to function completely.
4.9 | The control panel and all control switches must be well sealed preventing entry of fluids to the internal workings and circuits of the unit.
4.10 | All menus and key indicators shall be in English and/or internationally recognised symbols.
4.11 | The display shall be separate from the laryngoscope blade and shall be of a reasonable size and have a clear image display.
4.12 | The display must have the capability to connect via a video cable for monitor slaving option
4.13 | The unit must have the ability to record the intubation process either directly or with an attachment
4.14 | The Control panel must indicate whether the unit is operating off the 220 Volt, 50Hz A.C. line power or from its own internal battery power supply. The bidder should state the capacity and type of internal battery employed
4.15 | The battery power should last for 2 hrs or more.
5 | STANDARD ACCESSORIES
5.1 | Mcgill forceps x1 (size as per End-user's choice)
5.2 | 3 x blades (As per clause 3.5 above)
5.3 | Guide
5.4 | Intubation stylet/introducer
5.5 | Acutely bent blade
6 | Conformity Compliance (Please attach certificate)
6.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
6.2 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
6.3.1 | Please provide unique ref number of the ISO 13485 ceriticate:
6.4 | Model on offer must comply with EU Medical Device Regulation (EU) 2017/745 or equvalent from SAHPRA recognized jurisdictions . Attach a copy of certification
6.4.1 | Please provide unique ref number of the ceriticate:
6.5 | Hazardous Substance Act: If this type of equipment / apparatus appears on the schedule of Hazardous substances issued by the Directorate: Health Technology of the Department of Health, a license in terms of the Act on Hazardous Substances (Act 15/1973) must be submitted with the bid / quotation documents. The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
6.7 | All electrical/electronic medical equipment must be licensed by Radiation Control, where applicable a copy of the license must be submitted The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
6.7.1 | Radiation Control License number:
6.8 | The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade (rubber to with 2 screws) type and it must be a minimum length of three (3) meters. NB the mains cable of the unit tendered for must be SABS colour coded.
6.9 | The equipment quoted for must be protected against electromagnetic interference IEC 60601-1-2
7 | COMPREHENSIVE WARRANTY COVER
7.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) should be quoted for, as per the SCC document
7.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
8 | General
8.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
8.2 | The unit on offer must be the latest technology.
8.2.1 | Initial Manufacturing year of Model:
| INTUBATION VIDEO LARYNGOSCOPE ACCESSORIES & CONSUMABLES | Complies Yes/No
RT2-01-11a | Mcgill forceps x1 (size as per End-user's choice)
RT2-01-11b | Guide
RT2-01-11c | Intubation stylet/introducer
RT2-01-11d | Acutely bent blade
RT2-01-11e | Reusable Bariatric blade
RT2-01-11f | Reusable Miller blade size 00
RT2-01-11g | Reusable Miller blade size 0
RT2-01-11h | Reusable Miller blade size 1
RT2-01-11i | Reusable Miller blade size 2
RT2-01-11j | Reusable McIntosh / Macintosh blade size 0
RT2-01-11k | Reusable McIntosh / Macintosh blade size 1
RT2-01-11l | Reusable McIntosh / Macintosh blade size 2
RT2-01-11m | Reusable McIntosh / Macintosh blade size 3
RT2-01-11n | Reusable McIntosh / Macintosh blade size 4
RT2-01-11o | Reusable Laryngoscope with 60 degrees angulation size 1
RT2-01-11p | Reusable Laryngoscope with 60 degrees angulation size 2
RT2-01-11q | Reusable Laryngoscope with 60 degrees angulation size 3
RT2-01-11r | Reusable Laryngoscope with 60 degrees angulation size 4
RT2-01-11s | A connecting cable (monitor and the blades).
RT2-01-11t | External power supply (where applicable)
RT2-01-11u | Disposable Laryngoscope with 60 degrees angulation size 1
RT2-01-11v | Disposable Laryngoscope with 60 degrees angulation size 2
RT2-01-11w | Disposable Laryngoscope with 60 degrees angulation size 3
RT2-01-11x | Disposable Laryngoscope with 60 degrees angulation size 4

## Video Laryngoscope + Broncho
RT2-01-12 | INTUBATION VIDEO LARYNGOSCOPE with Flexible bronchscope | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
1 | DESCRIPTION
1.1 | A medical device used for visualizing the larynx and trachea during intubation procedures or video flexible bronchoscope.
1a | Product Offered:
1a.1 | Make:
1a.2 | Model:
1a.3 | Initial Manufacturing year of Model:
1a.4 | Radiation Control License number:
2 | TECHNICAL SPECIFICATION
2.1 | The unit must be mounted on a mobile stand.
2.2 | The unit must have facility to be used on all patients ranging from preterm neonatal (from 600g) to morbidly obese adults
2.3 | The handheld unit will be attached to an electronic display to facilitate visualization of the larynx
2.4 | The device should incorporate NTSC colour video or latest technology.
2.5 | Slimline durable handheld intubation scope similar in design to a laryngoscope handle and blade system
3 | The offered video intubation system shall consist of the following components which are fully intergrated into the system (not limitted to):
3.1 | Slimline durable handheld intubation scope similar in design to a laryngoscope handle and blade system.
3.2 | Display and Control unit monitor must comprise of at least 7" LED.
3.3 | Power supply (internal and external);
3.4 | Mobile stand on castor wheels (at least two of them with brakes)
3.5 | IV pole mount for semi-permanent fixation
3.6 | Mounting adaptor to secure the monitor onto the mobile stand or IV pole mount
3.7 | Suitable blade holder / basket capable of storing a minimum of two blades. The holder/ basket must be compatible for attachment to the mobile stand and IV pole mount
3.8 | The offered system shall be fully operational on delivery and will not require any additional accessories or consumable items (at least one month of consumable items) in order to function completely.
3.9 | The control panel and all control switches must be well sealed preventing entry of fluids to the internal workings and circuits of the unit.
3.10 | All menus and key indicators shall be in English and/or internationally recognised symbols.
3.11 | The display shall be separate from the laryngoscope blade and shall be of a reasonable size and have a clear image display.
3.12 | The display must have the capability to connect via a video cable for monitor slaving option
3.13 | Device must have the ability to record and capture the intubation process either directly or with an attachment
3.14 | The Control panel must indicate whether the unit is operating off the 220 Volt, 50Hz A.C. line power or from its own internal battery power supply. The bidder should state the capacity and type of internal battery employed
3.15 | The battery power should last for 2 hrs or more.
3.16 | The unit should be able to accommodate a flexible reusable bronchoscope (size as per end user request).
3.17 | Unit should be able to accommodate the video laryngoscope blades
4 | The Blade should have:
4.1 | The acutely angled blade must have an angulation of at least 60⁰ and be available for adult and paediatrics sizes
4.2 | An automatic anti fogging mechanism and anti glare technology in order to ensure a clear image is displayed on the display unit during use. This anti fogging mechanism will not require any action on the part of the user to function.
4.3 | An automatic autofocus mechanism in order to ensure a clear image is displayed on the display unit during use.
4.4 | Fully reusable and disposable blades will be accepted.
4.5 | The device should be supplied with one full set of three blades (neonatal or paediatric or adult) of the choice of the end user.
4.6 | The full video laryngoscope blade, meaning both the length of the blade and the blade handle, must be able to be sterilised. The Bidder to clearly state the different sterilization methods that may be performed on the offered laryngoscope type blade and other components. Suitable wall cleaning charts are to be provided to ensure appropriate cleaning and unit maintenance of the device.
5 | BRONCHOSCOPE
5.1 | Must have a connector cable to allow view to be projected to the monitor
5.2 | Minimum tip deflection should be 100º degrees up and 100º degrees down
5.3 | The image produced should be clearly defined without pixelation
5.4 | Must have a viewing angle of > 90º
5.5 | Must have a direction of view of 0 º
5.6 | Must have a minimum working length of ± 60 cm
5.7 | Must have an instrument channel with an ID of ± 1.2 mm
5.8 | A single channel should be available for suction with an appropriate suction controller
5.9 | The suction controller should be removable to allow alternative uses such as insufflation of oxygen or administration of medication
5.10 | Must have a distal tip OD of ± 4.0 mm
5.11 | State if any other size flexible bronchoscopes are available
6 | STANDARD ACCESSORIES
6.1 | mcgill forceps x1
6.2 | guide
6.3 | Acutely bent blade
6.4 | Flexible bronchoscope size 4 (outer diametre 4mm, length 65cm), if there are any other size available please state.
7 | Conformity Compliance (Please attach certificate)
7.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
7.2 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
7.3.1 | Please provide unique ref number of the ISO 13485 ceriticate:
7.4 | Model on offer must comply with EU Medical Device Regulation (EU) 2017/745 or equvalent from SAHPRA recognized jurisdictions . Attach a copy of certification
7.4.1 | Please provide unique ref number of the ceriticate:
7.5 | Hazardous Substance Act: If this type of equipment / apparatus appears on the schedule of Hazardous substances issued by the Directorate: Health Technology of the Department of Health, a license in terms of the Act on Hazardous Substances (Act 15/1973) must be submitted with the bid / quotation documents. The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
7.6 | All electrical/electronic medical equipment must be licensed by Radiation Control, where applicable a copy of the license must be submitted The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
7.6.1 | Radiation Control License number:
7.7 | The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade (rubber to with 2 screws) type and it must be a minimum length of three (3) meters. NB the mains cable of the unit tendered for must be SABS colour coded.
7.8 | The equipment quoted for must be protected against electromagnetic interference IEC 60601-1-2
8 | COMPREHENSIVE WARRANTY COVER
8.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) should be quoted for, as per the SCC document
8.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
9 | General
9.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
9.2 | The unit on offer must be the latest technology.
9.2.1 | Initial Manufacturing year of Model:
| Accessories for INTUBATION VIDEO LARYNGOSCOPE with Flexible bronchscope | Complies Yes/No
RT2-01-12a | Flexible bronchoscope size 4 (outer diametre 4mm, length 65cm), if there are any other size available please state.

## Single use flexible broncho
RT2-01-13 | INTUBATION VIDEO FLEXIBLE BRONCHOSCOPE (with disposable scopes) | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
1 | DESCRIPTION
1.1 | A robust video bronchoscopy system to facilitate visualization of the larynx and the trachea achieving intubation by means of a video flexible, disposable bronchoscope .
1a | Product Offered:
1a.1 | Make:
1a.2 | Model:
1a.3 | Initial Manufacturing year of Model:
1a.4 | Radiation Control License number:
2 | TECHNICAL SPECIFICATION
2.1 | The unit must be mounted on a mobile stand.
2.2 | The unit must have facility to be used on all patients ranging from preterm neonatal (from 600g) to morbidly obese adults
2.3 | The handheld unit will be attached to an electronic display to facilitate visualization of the larynx and trachea
2.4 | The device should incorporate NTSC colour video or latest technology.
2.5 | Slimline durable handheld intubation scope
3 | The offered video intubation system shall consist of the following components which are fully intergrated into the system (not limitted to):
3.1 | Slimline durable handheld intubation scope
3.3 | Display and Control unit monitor; at least 7" - 10" TFT LCD or LED panel monitor will be acceptable.
3.4 | Power supply (internal and external);
3.2 | Mobile stand on castor wheels (at least two of them with brakes)
3.2 | IV pole mount for semi-permanent fixation
3.2 | Mounting adaptor to secure the monitor onto the mobile stand or IV pole mount
3.2 | Suitable scope holder capable of storing the scope. The holder must be compatible for attachment to the mobile stand and IV pole mount
3.2 | The offered system shall be fully operational on delivery and will not require any additional accessories or consumable items in order to function completely.
3.2 | The control panel and all control switches must be well sealed preventing entry of fluids to the internal workings and circuits of the unit.
3.2 | All menus and key indicators shall be in English and/or internationally recognised symbols.
3.2 | The display shall be separate from the laryngoscope blade and shall be of a reasonable size and have a clear image display.
3.2 | The display must have the capability to connect via a video cable for monitor slaving option
3.2 | Device must have the ability to record and capture the intubation process either directly or with an attachment
3.2 | The Control panel must indicate whether the unit is operating off the 220 Volt, 50Hz A.C. line power or from its own internal battery power supply. The bidder should state the capacity and type of internal battery employed
3.2 | The battery power should last for 2 hrs or more.
3.2 | The unit should be able to accommodate a flexible reusable bronchoscope (size as per end user request).
5 | BRONCHOSCOPE
5.1 | Must have a connector cable to allow view to be projected to the monitor
5.2 | Minimum tip deflection should be 100º degrees up and 100º degrees down
5.3 | The image produced should be clearly defined without pixelation
5.4 | Must have a viewing angle of > 90º
5.5 | Must have a direction of view of 0 º
5.6 | Must have a minimum working length of ± 60 cm
5.11 | State the different sizes of flexible bronchoscopes available
6 | STANDARD ACCESSORIES
6.1 | 3 flexible bronchoscopes as per end user request
7 | OPTIONAL ACCESSORIES
7.1 | List all the accessories available with prices
8 | Conformity Compliance (Please attach certificate)
8.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
8.3 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
8.3.1 | Please provide unique ref number of the ISO 13485 ceriticate:
8.4 | Model on offer must comply with EU Medical Device Regulation (EU) 2017/745 or equvalent from SAHPRA recognized jurisdictions . Attach a copy of certification
8.4.1 | Please provide unique ref number of the ceriticate:
8.5 | Hazardous Substance Act: If this type of equipment / apparatus appears on the schedule of Hazardous substances issued by the Directorate: Health Technology of the Department of Health, a license in terms of the Act on Hazardous Substances (Act 15/1973) must be submitted with the bid / quotation documents. The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
8.6 | All electrical/electronic medical equipment must be licensed by Radiation Control, where applicable a copy of the license must be submitted The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
8.6.1 | Radiation Control License number:
8.7 | The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade (rubber to with 2 screws) type and it must be a minimum length of three (3) meters. NB the mains cable of the unit tendered for must be SABS colour coded.
8.7.1 | The equipment quoted for must be protected against electromagnetic interference IEC 60601-1-2
9 | COMPREHENSIVE WARRANTY COVER
9.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) should be quoted for, as per the SCC document
9.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
10 | General
10.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
10.2 | The unit on offer must be the latest technology.
10.2.1 | Initial Manufacturing year of Model:

## Syringe Driver, TCI
RT2-01-15 | SYRINGE DRIVERS TCI | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
1 | DESCRIPTION
1.1 | An infusion pump system used for Target Controlled Infusion (TCI) of medicines, typically in anesthesia and critical care settings.
1a | Product Offered:
1a.1 | Make:
1a.2 | Model:
1a.3 | Initial Manufacturing year of Model:
1a.4 | Radiation Control License number:
2 | TECHNICAL SPECIFICATION
2.1 | A Target Controlled Infusion (TCI) Syringe Driver programmable with computer controlled infusion capabilities that comprises of the latest technology must be supplied with a priced list of the syringes and extension sets that the unit has been calibrated on.
2.2 | The bidder must be prepared to provide testing/certification documentation to prove the compatibility of generic syringes available in the country and must be prepared to provide certification of the compitibility or incompatibility of various generic syringes for the life span of the TCI pump
2.3 | The unit must automatically sense the syringe size and give a continuous visble indication, on the front panel, of the syringe size fitted.
2.4 | The unit offered must:
2.4.1 | Be user friendly
2.4.2 | Be lightweight – State Weight
2.4.3 | Be portable
2.4.4 | Be able to be mounted/ attached to a RAIL or to a mobile drip stand with a universal adaptor
2.5 | Have a carry handle
2.6 | The unit offered must be well constructed with robust impact resistant casing and panels. The control panel, all push buttons and switches must be of a flush membrane type or an equivalent, to provide resistance for the entry of fluids in to the internal workings and circuits of the unit.
3 | Power and Batteries
3.1 | The monitor must operate off both the 220V-240V , 50Hz single phase AC power supply and an internally fitted rechargeable battery. The pump must be fused in both the LIVE and NEUTRAL.
3.2 | The unit offered must be supplied with an internal rechargeable battery.
3.3 | The internal rechargeable battery must be automatically charged when the unit is connected to a live A.C. supply
3.4 | In the event of A.C. mains failure the battery must automatically provide continuous operation
3.5 | The unit must be provided with visible indication to reflect whether the unit is being operated off the AC power or the internal rechargeable battery power
3.6 | The unit offered must be provided with circuitry, which must ensure that the internal rechargeable battery will be protected against over-charging and over-discharging
3.7 | Unit that are quoted on and employ external battery chargers and power supply will not be considered.
3.8 | The internal rechargeable battery must be of a reasonable capacity, such that with battery power the unit must be able to operate continuously for a minimum of four (4)hours at user set infusion rate of 125 ml/h
3.9 | Please state the following:
3.9.1 | Capacity
3.9.2 | Voltage
3.9.3 | Type of battery used
3.9.4 | Length of battery operating time
4 | Modes of operation:
4.1 | The unit offered must be capable of functioning in two different modes.
4.2 | Standard mode:
4.2.1 | Volume or drug delivery in an appropriate per unit volume and per unit weight
4.2.2 | Must be capable of an adjustable bolus administration up to rate of 1200 ml/hour for TIVA
4.2.3 | A drug library must be available that, when selected, allows the name, concentration, and infusion rate of the selected drug to be displayed.
5 | Pharmacokinetic mode:
5.1 | Delivery of a pre-specified drug using standardized pharmacokinetic models.
5.2 | When in the pharmacokinetic mode the unit must clearly display:
5.2.1 | Drug name
5.2.2 | Concentration
5.2.3 | Plasma/effect site concentration
5.2.4 | Plasma/effect site target.
5.2.5 | Initial induction dose
5.2.6 | Initial induction rate
5.2.7 | Initial induction volume
5.2.8 | Time of induction
5.2.9 | Initial maintenance rate
5.2.10 | Patient age
5.2.11 | Patient Length/height
5.2.12 | Patient gender
5.2.13 | Patient weight
5.2.14 | Body Mass Index
5.2.15 | Drug model being used
5.2.16 | Decrement time
5.2.17 | Decrement concentration
5.2.18 | Elapsed time
5.2.19 | Volume infused
5.2.20 | Dose infused
5.3 | The unit must be provided as standard with the following pharmacokinetic models:
5.3.1 | Propofol – Marsh – plasma targeting
5.3.2 | Propofol – Schnider – plasma and effect site targeting
5.3.3 | Propofol – Kataria – paediatric model
5.3.4 | Remifentanil – Minto – plasma targeting
5.3.5 | Remifentanil – effect site targeting
5.3.6 | Sufentanil – Gepts – plasma targeting
5.3.7 | Sufentanil – effect site targeting
5.3.8 | The Eleveld universal modal is highly desirable
5.4 | The supplier must provide performance verification data for the unit that demonstrates the validity of their pharmacokinetic models
5.5 | As a safety mechanism - when functioning in the pharmacokinetic mode the user should not be able to exit the mode without first specifically terminating the programme
5.6 | The pharmacokinetic mode should stay engaged even when the pump is switched off and restarted
5.7 | The unit being tendered on must be equipped with computer interface technology. E.g RS232, USB or Wifi. Please state the interface tech
6 | Rate settings
6.1 | The unit offered must also be capable of delivering a high flow rate and the flow rate must be user selectable through the whole range.
6.2 | The unit must be capable of an adjustable bolus administration up to rate of 1200 ml/hour when required during normal infusion delivery
6.3 | The unit must be able to revert to the previously selected normal infusion delivery after administering a bolus.
6.4 | The unit must have a pump run indicator, which will provide the operator with visual indication that the pump is carrying out infusion delivery
6.5 | There must be separate switch controls and indicators for starting and stopping infusion delivery
6.6 | When the unit is turned off, the last settings such as volume and dose delivery must be retained in memory and displayed at switch on.
6.7 | There must be fast purge mode of operation, which will allow priming of IV line. It must not be possible to operate the unit continuously in this mode
6.8 | It must not be possible to change the flow rate on the pump during delivery. The pump must first be stopped, flow rate changed and then restarted.
6.9 | If no flow rate/zero flow rate is selected, it must not be possible to start the unit
6.1 | The unit must provide an audible and visible alarm within a set period should the pump not be restarted. State this time interval.
6.11 | There must be adequate safety precautions against tampering of settings by unauthorized persons. Provide details.
6.12 | The bolus following release of occlusion at 1 ml/h and all higher rates (at alarm), must not exceed 0,6 ml.
7 | Accuracy
7.1 | The accuracy (long term) measured over 60 min at 1ml/h and all higher rates must be better than ± 5% for the unit being offered
7.2 | The accuracy (short term) at 5 ml/h must be better than ± 5% of mean on 2 minutes observation window for the unit being offered.
7.3 | State the actual accuracy of the unit offered.
7.4 | Bidders must submit a written statement / report on the performance of the offered unit in the presence of a working electro-surgery unit
7.5 | Time duration from start-up to attain 95% of “rate set” at 1ml/h and all higher rate < 10 min for the unit being offered.
8 | Display
8.1 | The following must be available and clearly displayed on the front panel under all lighting conditions:
8.2 | A.C. mains power supply operation
8.3 | Battery power supply operation
8.4 | Drug name
8.5 | Plasma/effect site concentration and / or Plasma/effect site target
8.6 | Infusion RATE selected
8.7 | Total volume infused. Any bolus administered by the unit during normal infusion must be added to the total “volume infused” and displayed.
8.8 | Elapsed time
8.9 | Alarm condition and possible alarm / error messages
8.1 | Alarms
8.10.1 | All alarm conditions must be accompanied by an audible and visual warnings
8.10.2 | It must be possible to silence the audible alarm for a maximum of two minutes after which if the fault condition still exists the unit must automatically provide an audible alarm again
8.11 | Trends of drug administration should be recallable
9 | The unit offered must activate alarms for the following minimum conditions
9.1 | LOW battery. State the duration of battery power operation left when this alarm sounds
9.2 | Infusion set not properly loaded
9.3 | Syringe removed
9.4 | Syringe barrel clamp alarm (or equivalent)
9.5 | Syringe plunger disengaged
9.6 | Internal malfunction.
9.7 | Malfunctions detected during self-test at power up
9.8 | Continuous audible and visual alarm for end of travel of syringe plunger.
9.9 | An audible and visible pre-alarm to indicate that the syringe travel plunger is nearing its end and also when the set “Volume to be infused” is approaching final completion
9.1 | Occlusion during infusion administration
9.11 | Maximum occlusion pressure alarm: < 500 mmHg with the use of a 50 ml syringe at ≥1 ml/h The time to alarm following occlusion at 1 ml/h and all higher rates < 36 min for the unit being tendered on. State time to alarm
10 | Conformity Compliance (Please attach certificate)
10.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
10.2 | The unit must comply with electrical safety and performance standard such as IEC 60601-2-24, attached certification
10.2.1 | Please provide unique ref number of the IEC ceriticate:
10.4 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
10.4.1 | Please provide unique ref number of the ISO 13485 ceriticate:
10.5 | Model on offer must comply with EU Medical Device Regulation (EU) 2017/745 or equvalent from SAHPRA recognized jurisdictions . Attach a copy of certification
10.5.1 | Please provide unique ref number of the ceriticate:
10.6 | Hazardous Substance Act: If this type of equipment / apparatus appears on the schedule of Hazardous substances issued by the Directorate: Health Technology of the Department of Health, a license in terms of the Act on Hazardous Substances (Act 15/1973) must be submitted with the bid / quotation documents. The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
10.7 | All electrical/electronic medical equipment must be licensed by Radiation Control, where applicable a copy of the license must be submitted The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
10.7.1 | Radiation Control License number:
10.8 | The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade (rubber to with 2 screws) type and it must be a minimum length of three (3) meters. NB the mains cable of the unit tendered for must be SABS colour coded.
10.9 | The equipment quoted for must be protected against electromagnetic interference IEC 60601-1-2
11 | COMPREHENSIVE MAINTENANCE
11.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) should be quoted for, as per the SCC document
11.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
11.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
11.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
11.3 | The bidder to provide the OEM recommended service interval of the unit:
11.4 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
12 | General
12.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
12.2 | The unit on offer must be the latest technology.
12.2.1 | Initial Manufacturing year of Model:
| SYRINGE DRIVER TCI -ACCESSORIES and CONSUMABLES | Complies Yes/No
RT2-01-15a | Extension sets Clear Fluid (Straight) - 100 per box
RT2-01-15b | Extension sets (Spiral)- 100/box
RT2-01-15c | Extension sets (Light resistant/sensitive)- 100 per box
RT2-01-15d | Extension sets Low Sorb- 100 per box
RT2-01-15e | Extension sets Blood- 100 per box
RT2-01-15f | Neonatal Closed blood set with Filter - 100 per box
RT2-01-15g | Compatible Syringes LUER LOCK Size: 2ml (Pk of 10)
RT2-01-15h | Compatible Syringes LUER LOCK Size: 3ml (Pk of 10)
RT2-01-15i | Compatible Syringes LUER LOCK Size: 5ml (Pk of 10)
RT2-01-15j | Compatible Syringes LUER LOCK Size: 10ml (Pk of 10)
RT2-01-15k | Compatible Syringes LUER LOCK Size: 20ml (Pk of 10)
RT2-01-15l | Compatible Syringes LUER LOCK Size: 30ml (Pk of 10)
RT2-01-15m | Compatible Syringes LUER LOCK Size: 50ml (Pk of 10)
RT2-01-15n | Compatible Syringes LUER LOCK Size: 60ml (Pk of 10)

## Syringe Driver, PCA
RT2-01-16 | SYRINGE DRIVERS PCA | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
1 | DESCRIPTION
1.1 | A patient-controlled analgesia (PCA) pump is a programmable device that releases medicine for pain that you control with a button.
1a | Product Offered:
1a.1 | Make:
1a.2 | Model:
1a.3 | Initial Manufacturing year of Model:
1a.4 | Radiation Control License number:
2 | TECHNICAL SPECIFICATION
2.1 | A Syringe Driver (including patient hand set and locking facility) that are re-usable and programmable to provide patient controlled intravenous and epidural analgesia as well as regional analgesia infusions.
2.2 | Must be usable for intravenous, epidural, regional analgesia and wound infiltration infusions.
2.3 | It should pe progammable to provide continuous infusion only, bolus only, continuous pus bolus and auto bolus infusion
2.4 | Weight based dosing software with ceiling ptotection must be available.
2.5 | A searchable drug library should be avaialble.
2.6 | It should be possible to store programmed information.
2.7 | At least 2 levels of security acces to prevent unauthorised programme changes.
2.8 | Flow rate range:
2.8.1 | Basal rate 0 to 100 ml/hr in 0.1 ml increments
2.8.2 | KVO Rate 0 (OFF) - 5 ml/hr
2.8.3 | Flow rate accuracy - ± 6 %
2.8.4 | Volume to be infused: 0.1 – 1000 ml
2.9 | Infusion safety limitations
2.9.1 | ultrasonic Air Sensor Air bubble size 0.04 – 1 ml
2.9.2 | Occlusion Pressure adjustable 250-1500 mmHg (adjustable depending on infusion mode)
2.10 | The unit should be light weight and small for easy mobility of patients. State weight.
2.11 | It should have a carry handle
2.12 | The unit offered must be well constructed with robust impact resistant casing and panels. The control panel, all push buttons and switches must be of a flush membrane type or an equivalent, to provide resistance for the entry of fluids in to the internal workings and circuits of the unit.
2.13 | The unit must be user friendly for patients and doctors.
2.14 | Disposables/accesories must be clearly described and compatibility (or not) with generic disposables should be confirmed with documentary evidence.
2.15 | If syringes is part of the equipment, the size should be automatically sensed and displayed.
2.16 | Non-freeflow or backflow disposable administration sets should be offered (colour coded for different applications e.g. IV and epidural).
2.18 | The unit, including reservoir of administered drug should be lockable to prevent tampering. State if additional accessories are necessary for this.
1.18 | An electronic safety programme is desireable.
3 | Display
3.1 | The following must be available and clearly displayed on the front panel under all lighting conditions, preferably full colour:
3.2 | AC mains power supply operation
3.3 | Battery power supply operation
3.4 | Drug name & concentration
3.5 | Infusion rate selected
3.6 | Bolus volume
3.7 | Lock out time
3.8 | A history must be logged for at least 72 hours
3.9 | The history of bolus administration and total dose must be able to be called to view on screen
3.10 | I should be possible to export the history to computer
3.11 | Alarm condition and possible alarm / error messages
4 | Alarms
4.1 | All alarm conditions must be accompanied by an audible and visual warnings
4.2 | It must be possible to silence the audible alarm for a maximum of two minutes after which if the fault condition still exists the unit must automatically provide an audible alarm again
4.3 | Minimum alarms to be diplayed
4.3.1 | Air in line
4.3.2 | Up occlusion, Down occlusion
4.3.3 | Pump unattended
4.3.4 | End of infusion
4.3.5 | Near end of Infusion
4.3.6 | Low Battery, End Battery
4.3.7 | Pump Unattended
4.3.8 | Door open
4.3.9 | Missing set key
4.3.10 | System error
5 | Power and Batteries
5.1 | The monitor must operate off both the 220V-240V , 50Hz single phase AC power supply and an internally fitted rechargeable battery. The pump must be fused in both the LIVE and NEUTRAL.
5.2 | The unit offered must be supplied with an internal rechargeable battery.
5.3 | The internal rechargeable battery must be automatically charged when the unit is connected to a live AC supply.
5.4 | In the event of A.C. mains failure the battery must automatically provide continuous operation
5.5 | The unit must be provided with visible indication to reflect whether the unit is being operated off the AC power or the internal rechargeable battery power.
5.6 | The unit offered must be provided with circuitry, which must ensure that the internal rechargeable battery will be protected against over-charging and over-discharging
5.7 | Unit that are quoted on and employ external battery chargers and power supply will not be considered.
5.8 | The internal rechargeable battery must be of a reasonable capacity, such that with battery power the unit must be able to operate continuously for a minimum of twelve (12) hours at al settings.
5.9 | Please state the following:
5.9.1 | Capacity
5.9.2 | Voltage
5.9.3 | Type of battery used
5.9.4 | Length of battery operating time
6 | Operating conditions
6.1 | Must function between 18 and 40⁰C
6.2 | Must function between 15 and 90 % humidity
6.3 | Must function at all SA altitudes
7 | Conformity Compliance (Please attach certificate)
7.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification.
7.2 | The unit must comply with electrical safety and performance standard such as IEC 60601-2-24, attach certification.
7.2.1 | Please provide unique ref number of the IEC ceriticate:
7.3 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
7.3.1 | Please provide unique ref number of the ISO 13485 ceriticate:
74 | Model on offer must comply with EU Medical Device Regulation (EU) 2017/745 or equvalent from SAHPRA recognized jurisdictions . Attach a copy of certification
7.4.1 | Please provide unique ref number of the ceriticate:
7.5 | Hazardous Substance Act: If this type of equipment / apparatus appears on the schedule of Hazardous substances issued by the Directorate: Health Technology of the Department of Health, a license in terms of the Act on Hazardous Substances (Act 15/1973) must be submitted with the bid / quotation documents. The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
7.6 | All electrical/electronic medical equipment must be licensed by Radiation Control, where applicable a copy of the license must be submitted The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
7.6.1 | Radiation Control License number:
7.7 | The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade (rubber to with 2 screws) type and it must be a minimum length of three (3) meters. NB the mains cable of the unit tendered for must be SABS colour coded.
7.8 | The equipment quoted for must be protected against electromagnetic interference IEC 60601-1-2
8 | COMPREHENSIVE MAINTENANCE
8.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) should be quoted for, as per the SCC document
8.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
8.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
8.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
8.3 | The bidder to provide the OEM recommended service interval of the unit:
8.4 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
9 | General
9.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
9.2 | The unit on offer must be the latest technology.
9.2.1 | Initial Manufacturing year of Model:
| SYRINGE DRIVER PCA -ACCESSORIES and CONSUMABLES | Complies Yes/No
RT2-01-16a | Extension sets Clear Fluid (Straight) - 100/box
RT2-01-16b | Extension sets (Spiral)- 100/box
RT2-01-16c | Extension sets (Light resistant/sensitive)- 100/box
RT2-01-16d | Extension sets Low Sorb- 100/box
RT2-01-16e | Extension sets Blood- 100/box
RT2-01-16f | Neonatal Closed blood set with Filter - 100/box
RT2-01-16g | Compatible Syringes LUER LOCK Size: 2ml (Pk of 10)
RT2-01-16h | Compatible Syringes LUER LOCK Size: 3ml (Pk of 10)
RT2-01-16i | Compatible Syringes LUER LOCK Size: 5ml (Pk of 10)
RT2-01-16j | Compatible Syringes LUER LOCK Size: 10ml (Pk of 10)
RT2-01-16k | Compatible Syringes LUER LOCK Size: 20ml (Pk of 10)
RT2-01-16l | Compatible Syringes LUER LOCK Size: 30ml (Pk of 10)
RT2-01-16m | Compatible Syringes LUER LOCK Size: 50ml (Pk of 10)
RT2-01-16n | Compatible Syringes LUER LOCK Size: 60ml (Pk of 10)

## Blood-Fluid Warmer
RT2-01-17 | BLOOD/FLUID WARMER | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
1 | DESCRIPTION
1.1 | A device that is used to warm blood, blood products, and other intravenous fluids to a safe temperature before transfusion or administration to patients.
1a | Product Offered:
1a.1 | Make:
1a.2 | Model:
1a.3 | Initial Manufacturing year of Model:
1a.4 | Radiation Control License number:
2 | TECHNICAL SPECIFICATION
2.1 | A blood solution warmer must be lightweight, robustly constructed and its design must be such that it must be possible for it to be attached to a normal infusion stand and a rail or equivalent.
2.2 | The unit on offer must be fitted and supplied with a universal attachment for enabling attachment of the unit to both an infusion stand and a Gabler rail, and the cost of which must be included in the final bid price.
2.3 | If the unit on offer uses any other dedicated disposables or consumable items, the bidder must list them with the prices including V.A.T on the schedule of optional accessories.
2.4 | The heat exchanger control circuit of the unit offered must provide well-controlled uniform heat distribution, preventing any hot spots developing that may damage blood.
2.5 | The design of the unit on offer must be such that there will be minimal cooling down of blood and fluid that is being administered to the patient after warming.
2.6 | The unit must have reliable safety protection against over temperature. Both an audible and visual alarm must be provided on the quoted unit should the temperature rise to a preset temperature, which is over and above the maximum user selectable temperature. The over temperature protection must also additionally turn off the power to the heater should the temperature exceed this preset temperature. Bidder must state the preset temperature at which this safety protection against over temperature will turn of the power to the heater.
2.7 | Unit must warm 150ml/min of blood from 4 degrees Celcius to 38 degrees Celcius
2.8 | The unit on offer must operate off the 220V to 240V 50Hz single phase a.c. supply and be fused in both the LIVE and NEUTRAL.
2.9 | The unit offered must be provided with a suitable display of the temperature control set point and must be clear viewing under all lighting conditions.
2.10 | State what measures are used to prevent or eliminate air bubbles from the blood and fluid circuit.
3 | STANDARD ACCESSORIES
3.1 | Mounting bracket
4 | Conformity Compliance (Please attach certificate)
4.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
4.2 | The unit must comply with electrical safety and performance standard such as IEC 60601-2-35, attached certification
4.2.1 | Please provide unique ref number of the IEC ceriticate:
4.4 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
4.4.1 | Please provide unique ref number of the ISO 13485 ceriticate:
4.5 | Model on offer must comply with EU Medical Device Regulation (EU) 2017/745 or equvalent from SAHPRA recognized jurisdictions . Attach a copy of certification
5.5.1 | Please provide unique ref number of the ceriticate:
4.6 | Hazardous Substance Act: If this type of equipment / apparatus appears on the schedule of Hazardous substances issued by the Directorate: Health Technology of the Department of Health, a license in terms of the Act on Hazardous Substances (Act 15/1973) must be submitted with the bid / quotation documents. The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
4.7 | All electrical/electronic medical equipment must be licensed by Radiation Control, where applicable a copy of the license must be submitted The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.7.1 | Radiation Control License number:
4.8 | The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade (rubber to with 2 screws) type and it must be a minimum length of three (3) meters. NB the mains cable of the unit tendered for must be SABS colour coded.
4.9 | The equipment quoted for must be protected against electromagnetic interference IEC 60601-1-2
5 | COMPREHENSIVE MAINTENANCE
5.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) should be quoted for, as per the SCC document
5.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
5.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
5.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
5.3 | The bidder to provide the OEM recommended service interval of the unit:
5.4 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
6 | General
6.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
6.2 | The unit on offer must be the latest technology.
6.2.1 | Initial Manufacturing year of Model:
| BLOOD/FLUID WARMER -ACCESSORIES and CONSUMABLES | Complies Yes/No
RT2-01-017a | Cassette/cartridge (pk of 5)
RT2-01-017b | Flow warming set, Adult (pk of 10)
RT2-01-017c | Flow warming set, Peads (pk of 10)

## Patient Warmer, Forced-Air
RT2-01-18 | PATIENT WARMER (FORCED AIR) | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
1 | DESCRIPTION
1.1 | A medical device designed to keep patients warm and confortable during medical procedures, recovery and hospital stays using a steam of warm air.
1a | Product Offered:
1a.1 | Make:
1a.2 | Model:
1a.3 | Initial Manufacturing year of Model:
1a.4 | Radiation Control License number:
2 | TECHNICAL SPECIFICATION
2.1 | The unit must be able to operate off the 220V – 240V, 50Hz single phase A.C. power supply.
3 | The physical properties of the unit should fulfill the following:
3.1 | Weight should not exceed 8kg
3.2 | The unit must come with sturdy mounting capability
3.3 | It should be no larger than 38cm (height) X 41cm (width) X 33cm (depth)
3.4 | The operating noise level should be less than 60db
3.5 | The Control Panel must have LED or equivalent display and all switches must preferably be touch activated.
4 | The Control panel must offer a choice of at least 3 selectable temperature settings as well as ambient temperature setting. The temperature settings must be typically as follows:
4.1 | Low: set value between 30-34 degrees Celsius with an accuracy of ± 1.5 degrees Celsius
4.2 | Medium: set value between 35-39 degrees Celsius with an accuracy of ± 1.5 degrees Celsius
4.3 | High: set value between 39-43 degrees Celsius with an accuracy of ± 1.5 degrees Celsius
4.4 | The unit must be able to circulate room ambient temperature air
4.5 | The unit must indicate when the selected set temperature has been reached and thereafter maintain the selected set temperature
4.6 | The unit must reach the selected temperature within 90 seconds of the selection of that temperature.
4.7 | The unit must incorporate visual and audible temperature discrepancy alarms. These should be activated when the temperature measured at the hose-end exceeds or falls sort of the set temperature by more than 1 degree Celsius. Bidder to demonstrate these alarms during evaluation
4.8 | An over-temperature state must interrupt the heating function of the unit. The over-temperature alarm must activate when the measure hose-end temperature exceeds the set temperature by more than 3 degrees Celsius. Bidder to demostrate these alarms during evaluation
4.9 | The air supply inlet must incorporate a 0,3 micron (or smaller) HEPA filter.
4.10 | The unit must generate airflow of at least 19 litres per second.
5 | THE OUTLET HOSE : The outlet hose must be of the re-usable type and it must be easily removable for cleaning or replacement.
5.1 | The hose must have the following features:
5.1.1 | It must be able to swivel freely
5.1.2 | It must have securing devices for to prevent unintended disconnection
5.1.3 | Its minimum length must be 1.5 metre and suitable for both paediatrics and adults.
6 | THE BLANKETS
6.1 | These must be disposable, lightweight, soft, radiolucent and latex-free.
6.2 | The unit must utilize blankets that allow even heat distribution to the patient without the creation of hotspots.
7 | STANDARD ACCESSORIES
7.1 | A starter pack of the following blankets to be provided: 2 x Boxes (10 ea) of full body . (i.e 1 box paeds and 1box adults).
7.2 | Mobile stand with at least 4 castors and 2 fitted with brakes. (Complete for mounting the unit securely)
8 | Conformity Compliance (Please attach certificate)
8.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
8.2 | The unit must comply with electrical safety and performance standard such as IEC 60601-2-35, attached certification
8.2.1 | Please provide unique ref number of the IEC ceriticate:
8.4 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
8.4.1 | Please provide unique ref number of the ISO 13485 ceriticate:
8.5 | Model on offer must comply with EU Medical Device Regulation (EU) 2017/745 or equvalent from SAHPRA recognized jurisdictions . Attach a copy of certification
8.5.1 | Please provide unique ref number of the ceriticate:
8.6 | Hazardous Substance Act: If this type of equipment / apparatus appears on the schedule of Hazardous substances issued by the Directorate: Health Technology of the Department of Health, a license in terms of the Act on Hazardous Substances (Act 15/1973) must be submitted with the bid / quotation documents. The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
8.7 | All electrical/electronic medical equipment must be licensed by Radiation Control, where applicable a copy of the license must be submitted The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
8.7.1 | Radiation Control License number:
8.8 | The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade (rubber to with 2 screws) type and it must be a minimum length of three (3) meters. NB the mains cable of the unit tendered for must be SABS colour coded.
8.9 | The equipment quoted for must be protected against electromagnetic interference IEC 60601-1-2
9 | COMPREHENSIVE MAINTENANCE
9.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) should be quoted for, as per the SCC document
9.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
9.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
9.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
9.3 | The bidder to provide the OEM recommended service interval of the unit:
9.4 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
10 | General
10.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
10.2 | The unit on offer must be the latest technology.
10.2.1 | Initial Manufacturing year of Model:
| PATIENT WARMER (FORCED AIR), REPLACEMENT ACCESSORIES & CONSUMABLES | Complies Yes/No
RT2-01-18a | Upper body blanket
RT2-01-18b | Lower body blanket
RT2-01-18c | Surgical Access blanket
RT2-01-18d | Underbody Peadiatric blanket
RT2-01-18e | Underbody Adult blanket
RT2-01-18f | Full body Peaditric blanket
RT2-01-18g | Full body Adult blanket
RT2-01-18h | Air Hose
RT2-01-18i | Mobile stand
RT2-01-18j | HEPA filter (pk of 5)

## Patient Warmer
RT2-01-19 | PATIENT WARMER | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
1 | DESCRIPTION
1.1 | A medical device designed to keep patients warm and confortable during medical procedures, recovery and hospital stays.
1a | Product Offered:
1a.1 | Make:
1a.2 | Model:
1a.3 | Initial Manufacturing year of Model:
1a.4 | Radiation Control License number:
2 | TECHNICAL SPECIFICATION
2.1 | The unit being offered must be suitable for use in the operating theatre, ICU, Accident and Emergency department.
2.2 | It must be suitable for both paediatric and adult use and it must be supplied with a mattress / heating pad suitable for both adult and paediatric applications.
2.3 | The unit being offered must comprise the latest technology. It must be user friendly and incorporate reliable safety features
2.4 | The unit being quoted for must operate from both the 220V-240V, 50Hz A.C. mains supply. The unit must be fused in both LIVE and NEUTRAL.
2.5 | The operating voltage required by the mattress / heating pad must be restricted to a maximum of 30V and this voltage must be provided to the mattress / heating pad by the control unit.
2.6 | The cable connecting the mattress / heating pad from the control unit must be robust, extremely flexible and it must be a minimum length of 2.5 m.
2.7 | The control unit must be a microprocessor-based system that will provide accurate and uniform distribution and regulation of the temperature within the mattress / heating pad and it must also ensure that the unit has a fail safe mechanism.
2.8 | In the event where the mattress / heating pad is made up of more than one heating element, the temperature within the mattress / heating pad must be monitored by the control unit and the average value must be displayed. The individual heating units should not differ in temperature by more than 1 degree celcius.
2.9 | The control unit must provide a user selectable set temperature in a minimum range of 33oC up to 39oC and the set temperature selector control must be located on the front panel.
2.10 | The offered unit must be capable of monitoring the actual temperature of the heating pad.
2.11 | The heating element of the offered unit must be integrated into the heating pad.
2.12 | The user selected set temperature must be clearly displayed on the front panel.
2.13 | In addition to displaying the user selected set temperature, the control unit must also display the actual heating pad temperature on a suitable display located on the front panel.
2.14 | Sensors within the heating pad must be so arranged that “hot spots” within the heating pad will be prevented from occurring.
2.15 | The control unit must be guarded against external RF interference for the unit offered.
2.16 | The temperature accuracy of the unit must be within a tolerance of 0,5oC.
2.17 | The control unit must be provided with both audible and visible alarms should the temperature exceed 41oC and at which stage heater power must also be automatically shut off on the unit offered.
2.18 | The cable must be securely attached to the heating pad and must be waterproof.
2.19 | The heating pad must be watertight and easy to clean.
2.20 | The heating pad must have good electrical insulation and isolated from direct patient contact.
2.21 | The heating pad must be manufacturered from high density material that does not create pressure points on patient's skin.
2.22 | The control unit must not weigh more than 7.5kg and be supplied with a mounting bracket suitable for both IV pole and bedrail.
3 | The control unit offered must also be provided with the following audible and visible alarms:
3.1 | For heating pad fault.
3.2 | Control unit fault.
3.3 | The control unit offered must also be provided with an audible mains failure alarm, which must be able to function for a minimum period of 10 minutes (1 or 3 minutes) in the event of a mains failure.
3.4 | The control unit offered must also provide an audible and visible alarm should the actual temperature of the mattress / heating pad deviate from the user selected set temperature. A maximum of 1 degree celcius will be acceptable.
4 | STANDARD ACCESSORIES
4.1 | The unit to be supplied with under body heating pad for adult x1 and paeds x1
4.2 | Mobile stand with at least 4 castors and 2 fitted with brakes. (Complete for mounting the unit securely)
5 | Conformity Compliance (Please attach certificate)
5.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
5.2 | The unit must comply with electrical safety and performance standard such as IEC 60601-2-35, attached certification
5.2.1 | Please provide unique ref number of the IEC ceriticate:
5.3 | OEM must comply and certified on ISO 9001 quality standards, attach proof of compliance
5.3.1 | Please provide unique ref number of the ISO 9001 ceriticate:
5.4 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
5.4.1 | Please provide unique ref number of the ISO 13485 ceriticate:
5.5 | Model on offer must comply with EU Medical Device Regulation (EU) 2017/745 or equvalent from SAHPRA recognized jurisdictions . Attach a copy of certification
5.5.1 | Please provide unique ref number of the ceriticate:
5.6 | Hazardous Substance Act: If this type of equipment / apparatus appears on the schedule of Hazardous substances issued by the Directorate: Health Technology of the Department of Health, a license in terms of the Act on Hazardous Substances (Act 15/1973) must be submitted with the bid / quotation documents. The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.7 | All electrical/electronic medical equipment must be licensed by Radiation Control, where applicable a copy of the license must be submitted The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.7.1 | Radiation Control License number:
5.8 | The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade (rubber to with 2 screws) type and it must be a minimum length of three (3) meters. NB the mains cable of the unit tendered for must be SABS colour coded.
5.9 | The equipment quoted for must be protected against electromagnetic interference IEC 60601-1-2
6 | COMPREHENSIVE MAINTENANCE
6.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) should be quoted for, as per the SCC document
6.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
6.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
6.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
6.3 | The bidder to provide the OEM recommended service interval of the unit:
6.4 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
7 | General
7.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
7.2 | The unit on offer must be the latest technology.
7.2.1 | Initial Manufacturing year of Model:
| PATIENT WARMER ACCESSORIES | Complies Yes/No
RT2-01-19a | Paediatrics Heating pad
RT2-01-19b | Adults Heating pad
RT2-01-19c | Mobile Stand
RT2-01-19d | Upper body blanket
RT2-01-19e | Lower body blanket
RT2-01-19f | Surgical Access blanket
RT2-01-19g | Underbody Peadiatric blanket
RT2-01-19h | Underbody Adult blanket

## Nerve Stimulator
RT2-01-20 | NERVE STIMULATOR | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
1 | DESCRIPTION
1.1 | A device which utilises electrical current to stimulate the nerves for therapeutic/diagnostic purposes
1a | Product Offered:
1a.1 | Make:
1a.2 | Model:
1a.3 | Initial Manufacturing year of Model:
1a.4 | Radiation Control License number:
2 | TECHNICAL SPECIFICATION
2.1 | This specification establishes the requirements, supply, delivery, end user training, demonstration, commission and installation of a HAND HELD NERVE STIMULATOR, that is:
2.2 | Compact and light
2.3 | Robust
2.4 | User friendly
2.5 | Microprocessor controlled with the latest up to date technology.
2.6 | The unit offered must offer a full range of BOTH conventional and recently developed nerve stimulation patterns for monitoring of neuromuscular blockade during surgery or in the intensive care setting.
2.7 | The unit must have a constant current variable voltage feature, which will allow delivery of comparable stimuli throughout the whole operation, despite changes in impedance.
2.8 | The pulse stimulus must be of a monophasic, square type to ensure a constant current is maintained throughout the stimulus.
2.9 | The unit must offer the user all the conventional patterns of stimulation which must be easy to identify and select on the unit.
2.10 | The patterns of stimulation must include:
2.11
2.12 | Train of four: Four successive 2Hz pulses per 2 seconds (one pulse every 0.5 seconds) with a 10 second interval between trains,on demand or twelve-second repeat
2.13 | Tetanic stimulation: on demand at 50Hz or 100Hz for up to a minimum of five seconds
2.14 | Post -Tetanic count.
2.15 | Double burst stimulation: The burst repeats must be separated by 10 seconds. Selectable 3,3 or 3,2 patterns
2.16 | The unit must be provided with a digital alphanumeric display, which must provide clear viewing of all parameters under all lighting conditions.
2.17 | The unit offered must provide display of the current output
2.18 | The unit must be able to deliver a constant current output which must be selectable in the range of 0-80mA
2.19 | The output pulse must be:
2.20 | A monophasic square wave form
2.21 | Longer than 0.1milliseconds and less than 0.5milliseconds
2.22 | Longer duration pulses and other waveforms if available must be clearly indicated when selected.
2.23 | The bidder must clearly state the output pulse forms available on the unit offered.
3 | BATTERIES
3.1 | The unit must be battery operated.
3.2 | The battery cover must be easy to open and must require no tools to open and close the cover.
3.3 | It is important that the unit offered is provided with audible and visual indication of pulse delivery
3.4 | The unit must be provided with a Low Battery Indicator.
3.5 | The unit must have an Automatic Power Shutdown feature in order to conserve battery power.
3.6 | The units should have a programming mode whereby users may define and store any combination of pulse patterns and time intervals.
4 | STANDARD ACCESSORIES
4.1 | Electrodes set with crocodile clips or equivalent
4.2 | Battery
5 | Conformity Compliance (Please attach certificate)
5.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
5.2 | The unit must comply with electrical safety and performance standard such as IEC 60601-2-10, attached certification
5.2.1 | Please provide unique ref number of the IEC ceriticate:
5.3 | OEM must comply and certified on ISO 9001 quality standards, attach proof of compliance
5.3.1 | Please provide unique ref number of the ISO 9001 ceriticate:
5.4 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
5.4.1 | Please provide unique ref number of the ISO 13485 ceriticate:
5.5 | Model on offer must comply with EU Medical Device Regulation (EU) 2017/745 or equvalent from SAHPRA recognized jurisdictions . Attach a copy of certification
5.5.1 | Please provide unique ref number of the ceriticate:
5.6 | Hazardous Substance Act: If this type of equipment / apparatus appears on the schedule of Hazardous substances issued by the Directorate: Health Technology of the Department of Health, a license in terms of the Act on Hazardous Substances (Act 15/1973) must be submitted with the bid / quotation documents. The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.7 | All electrical/electronic medical equipment must be licensed by Radiation Control, where applicable a copy of the license must be submitted The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.7.1 | Radiation Control License number:
6 | COMPREHENSIVE WARRANTY COVER
6.1 | A 2-YEAR warranty inclusive maintenance and service (Year 1 to Year 2) should be quoted for, as per the SCC document
6.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included.
7 | General
7.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
7.2 | The unit on offer must be the latest technology.
7.2.1 | Initial Manufacturing year of Model:
| NERVE STIMULATOR | Complies Yes/No
RT2-01-20a | Electrodes set with crocodile clips or equivalent
RT2-01-20b | Battery

## Thromboelastogram
RT2-01-21 | Thromboelastography Unit | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
1 | DESCRIPTION
1.1 | A unit used to perform a rapid point-of-care test to assess global haemostasis in whole blood in the setting of trauma and surgery, to guide appropriate usage of blood products
1a | Product Offered:
1a.1 | Make:
1a.2 | Model:
1a.3 | Initial Manufacturing year of Model:
1a.4 | Radiation Control License number:
2 | TECHNICAL SPECIFICATION
2.1 | A viscoelastic haemostatic analyzer or equivalent ,cartridge based.Test must include dynamic assessment of clot development and strength,assess platelet function.The measurements must be in digital and graphic form.
2.2 | The unit must be portable and light in weight the bidder must state dimensions and weight.
2.3 | LCD touch screen
2.4 | Printer
2.5 | Operating temperature: 10 degrees C to 32 degrees C(50 degrees F to 89.6 degrees F)
2.6 | Operating pressure: 0 to 2000 m(0 to 6,560 ft) above see level
2.7 | Operating Humidity: 20% to 80% RH(non-condensing)
2.8 | Voltage: 100V-240V AC
2.9 | Power supply output voltage:12 VOLTS DC
2.10 | Output power/current ratings: 5A,60W (max)
2.11 | Supply frequency:50/60HZ
2.12 | Viscoelastic haemostatic testing system: cartridge based.
2.13 | Test must include dynamic assessment of clot development and strength
2.14 | Should have the capacity to assess platelet function
2.15 | Should be able to assess clot strength via suitable technology using resonanance technology.
2.16 | Point of care test results should be available within 5 to 10 mins
2.17 | The quality control should be done regularly.Period must be specified by the bidding company.
2.18 | The level of the quality checks must be indicated by the bidding company
2.19 | Trolley
2.19.1 | The unit must be mounted on a strong , robust and compact trolley manufactured from corrosion free material.
2.19.2 | The trolley must have no sharp corners and be easy to clean.
2.19.3 | The trolley must have at least four antistatic castors of diameter 100 mm or greater.
2.19.4 | Trolley must have two lockable castors or a central locking brake.
2.19.5 | Trolley must have accessible handles for ease of movement
3 | Other system requirements
3.1 | The unit must operate from both the 220 V, 50 Hz mains supply and internal rechargeable batteries
3.2 | The batteries must be charged while the unit is operating from the mains supply.
3.3 | The batteries must automatically take over in the event of a mains power failure.
3.4 | The batteries must have the capacity for full operation of the machine for at least 45 minutes with batteries fully charged, under typical operating conditions.
3.5 | Graphic battery power indicator
3.6 | Where an uninterrupted power supply (UPS) is offered instead of internal battery, there must not be extra cost for it and replacement cost must be stated and its maintenance cost thereof.
3.7 | Mains cable at least four metres long with SABS approved 15A three prong plug.
3.8 | Data management: The workstation must be capable of downloading data to a USB device or memory card.
3.9 | The unit must be handed over in full operating order
4 | The following accessories/consumables must be supplied with the unit:
4.1 | Disposable Assay Cartirdges
4.2 | Disposable quality control Cartirdges as many as needed(level1,level 2 or equivalent)
4.3 | Vial with lyophilized biological material for the quality control cartridge.
4.4 | Pipettes or equivalent for blood sample.
4.5 | A starter-pack of consumables must be supplied with the unit comprising of all the above accessories/consumables.
5 | Conformity Compliance (Please attach certificate)
5.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
5.2 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
5.2.1 | Please provide unique ref number of the ISO 13485 ceriticate:
5.3 | Model on offer must comply with EU Medical Device Regulation (EU) 2017/745 or equvalent from SAHPRA recognized jurisdictions . Attach a copy of certification
5.3.1 | Please provide unique ref number of the ceriticate:
5.4 | Hazardous Substance Act: If this type of equipment / apparatus appears on the schedule of Hazardous substances issued by the Directorate: Health Technology of the Department of Health, a license in terms of the Act on Hazardous Substances (Act 15/1973) must be submitted with the bid / quotation documents. The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.5 | Must comply with 2.5 Occupational Health and Safety Act, Act no 85 of 1993 and Regulations
5.6 | All electrical/electronic medical equipment must be licensed by Radiation Control, where applicable a copy of the license must be submitted The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.6.1 | Radiation Control License number:
5.7 | The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade (rubber to with 2 screws) type and it must be a minimum length of three (3) meters. NB the mains cable of the unit tendered for must be SABS colour coded.
5.8 | The equipment quoted for must be protected against electromagnetic interference IEC 60601-1-2
6 | COMPREHENSIVE MAINTENANCE
6.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) should be quoted for, as per the SCC document
6.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
6.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
6.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
6.3 | The bidder to provide the OEM recommended service interval of the unit:
6.4 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
7 | General
7.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
7.2 | The unit on offer must be the latest technology.
7.2.1 | Initial Manufacturing year of Model:
| Thromboelastography Unit -Accessories and Consumables | Complies Yes/No
RT2-01-21a | Disposable Assay Cartridge
RT2-01-21b | Disposable quality control Cartridge (level1,level 2 or equivalent)
RT2-01-21c | Vial with lyophilized biological material for the quality control cartridge.
RT2-01-21d | Pipettes or equivalent for blood sample.
RT2-01-21e | Uninterrupted Power Supply (UPS)
RT2-01-21f | UPS replacement Battery

## Cerebral Oximeter, Advanced
RT2-01-22 | CELEBRAL OXIMETER, Advanced | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
1 | Description
1.1 | A non-invasive medical device that measures regional cerebral oxygen saturation (rSO2) and provides additional features for enhanced monitoring and analysis.
1a | Product Offered:
1a.1 | Make:
1a.2 | Model:
1a.3 | Initial Manufacturing year of Model:
1a.4 | Radiation Control License number:
2 | Technical Specification
2.1 | A moblie peripheral pulse oximeter that can measure additional parameters non-invasively
2.2 | The unit must be able to function independently
2.3 | Docking with a central moitor can be offered optional. If so full particulars should be provided by bidder
2.4 | Suitable for adult and paediatric use.
2.5 | The design must enable stability without movement on a flat surface or desktop
2.6 | The case of the unit must be robust and impact resistant.
2.7 | The design must be such that spilled liquids do not gain access to the internal working of the unit.
2.8 | If available a stand or fastening bracket should be offered
2.9 | State of the art modern electronics with Automatic rejection of motion artefacts
2.10 | Self test procedure at switch-on
2.11 | The unit must have trending for at least 72 hours
2.12 | The unit must be handed over in full operating order
2.13 | A starter-pack of consumables must be supplied with the unit.
3 | The following parameters in addition to pulse oximetry must be available in different user selectable configurations (i.e. not all listed parameter should always be part of the monitor). (Clear information must be provided if any additional equipment is necessary for measurement of certain parameters.)
3.1 | Peripheral pulse oximetry (Sp O2)
3.2 | Pulse rate (PR)
3.3 | Perfusion index (Oi)
3.4 | Pethismographic variability ratio (PVi)
3.5 | Total Haemoglobin (SpHb)
3.6 | Carboxyhaemoglobin (SpCO)
3.7 | Methaemoglobin (SpMet)
3.8 | Oxygen content (SpOC)
3.9 | Acoustic respiratory rate (RRa)
3.10 | Capnography (Et CO2)
3.11 | Fractional CO2 (FiCO2)
3.12 | Respiratory rate (RR)
3.13 | Brain function monitoring
3.14 | Near infrared Oximetry (Regional oximetry)
3.15 | All parameter must measure within a 5% accuracy. Heart rate within 1 beat.
4 | DISPLAY
4.1 | A full colour display visible from at least 40 degrees and offer clear viewing under all lighting conditions.
4.2 | If hand-held, automatic display rotation
4.3 | Touch screen controls is desireable
4.4 | Display should be user customisable
4.5 | Trens must be displayable (user selectable) and viewable with different time frames up to 72 hours
4.6 | Pulse oximetry:
4.6.1 | The unit must continuously measure and display SpO2 %
4.6.2 | Oxygen saturation (SpO2 %): Range of measurement: 0 - 100%
4.6.3 | The accuracy must be better than 2,5 % (70 - 100%) and 5% (50 - 70%).
4.6.4 | Plethysmographic waveform should be displayed
4.7 | Pulse rate numerical display
4.7.1 | Pulse rate (beats/min): Range of measurement: 20 - 300 bpm
4.7.2 | The pulse rate accuracy must be better than ± 2 %.
4.7.3 | An audible bleep should sound for each heartbeat, the pitch of which should vary with the oxygen saturation.
4.8 | Perfusion index - numerical display
4.9 | Pethismographic variability ratio (PVi) - numerical display (%)
4.10 | Total Haemoglobin (SpHb) - numerical display (%)
4.11 | Carboxyhaemoglobin (SpCO) - numerical display (%)
4.12 | Methaemoglobin (SpMet) - numerical display (%)
4.13 | Oxygen content (SpOC) - numerical display (%)
4.14 | Acoustic respiratory rate (RRa) - numerical display with waveform option selectable
4.15 | Capnography (Et CO2) numerical % graphical display in %, kPa, mmHg
4.15.1 | FiCO2 range: 0-25 vol%, 0-32 kPa, 0-200 mmHg
4.15.2 | FetCO2 0-25 vol%, 0-32 kPa, 0-200 mmHg
4.15.3 | RR 0-150
4.16 | Brain function monitoring
4.16.1 | Measure and isplay 2 righ + 2 left frontal EEG waveforms
4.16.2 | Processed EEG value related to anaesthetic effect (0-100)
4.16.3 | Right & Left Power/Density spectral array spectrograms
4.17 | Near infrared Oximetry (Regional oximetry)
4.17.1 | Must measure and display 2 channels (usually L and R) numerical and trend (0 to 100%)
4.17.2 | Must diplay difference to baseline
4.17.3 | Must display area unde the curve
4.17.4 | Battery status and connection to mains electricity must displayed
5 | ALARMS
5.1 | Default alarms must be available
5.2 | Al parameters should have user slecetable alarm limits
5.2.1 | SpO2 LOW: 50 % - 99 %
5.2.2 | Pulse rate LOW: 40 – 150 b.p.m.
5.2.3 | Pulse rate HIGH: 70 – 250 b.p.m.
5.2.4 | Probe disconnected to patient
5.3 | Alarm limits should be diplayed.
5.4 | Alarms sould have both an audio and visual element
5.5 | The audio alam should be silenceble for up to 2 minutes but visula alarms should stay on until problem is resolved
5.6 | Adjustable alarm volume
5.7 | Low battery: Alarm at least 10 minutes before battery power is exhausted
6 | SENSORS & LEADS TO BE OFFERED
6.1 | DC-I, Adult Reusable Sensor, 3 ft., Non-Sterile
6.2 | DCI-P, Pediatric Reusable Sensor, 3 ft., Non-Sterile
6.3 | DB-I, Adult Reusable Soft Finger Sensor, 3 ft., Non-Sterile
6.4 | TC-I, Reusable Tip-Clip Ear Sensor, 3 ft., Non-Sterile
6.5 | Y-I, Multisite Reusable sensor, 3 ft., Non-Sterile
6.6 | TF-I, Reusable Transflectance Forehead sensor, 3 ft., Non-Sterile
6.7 | Patient Cable, MD20-12, 20-pin Patient Cable, 12 ft
6.8 | Dual Cable MD20-12
6.9 | EEG Sensor, Adult, Non-Sterile
6.10 | EEG Sensor, Pediatric, Non-Sterile
6.11 | O3 Regional Oximetry Sensors, Adult
6.12 | O3 Regional Oximetry Sensors, Pediatric
6.13 | O3 Regional Oximetry Sensors, Infant/Neonatal
6.14 | Adult/Pediatric Airway Adapter Set 3m
6.15 | Infant/Neonatal, Airway Adapter Set
6.16 | Adult Nasal CO2 Cannula
6.17 | Infant Nasal CO2 Cannula
6.18 | Adult Nasal CO2 Cannula with O2
6.19 | Pediatric Nasal CO2 Cannula with O2
7 | POWER AND BATTERY
7.1 | The unit must operate from 220 V, 50 Hz mains supply
7.2 | Mains cable at least three metres long with SABS approved 15A three prong plug.
7.3 | The unit must include an internal rechargeable battery
7.4 | Battery charger to be incorporated in the unit
7.5 | The unit must be able to function from the mains supply regardless of the battery condition.
7.6 | Battery must have the capacity for the unit to operate fully for at least five hours
7.7 | The battery must be charged while the unit is operating from the mains supply.
7.8 | The battery must automatically take over in the event of a mains power failure.
7.9 | Automatic switch-off after a period of non-use
8 | Operating conditions
8.1 | 10-45 ⁰C
8.2 | All SA Alititudes
8.3 | Humidity: 15-90%
9 | Conformity Compliance (Please attach certificate)
9.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
9.2 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
9.2.1 | Please provide unique ref number of the ISO 13485 ceriticate:
9.3 | Model on offer must comply with EU Medical Device Regulation (EU) 2017/745 or equvalent from SAHPRA recognized jurisdictions . Attach a copy of certification
9.3.1 | Please provide unique ref number of the ceriticate:
9.4 | Hazardous Substance Act: If this type of equipment / apparatus appears on the schedule of Hazardous substances issued by the Directorate: Health Technology of the Department of Health, a license in terms of the Act on Hazardous Substances (Act 15/1973) must be submitted with the bid / quotation documents. The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
9.5 | All electrical/electronic medical equipment must be licensed by Radiation Control, where applicable a copy of the license must be submitted The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
9.5.1 | Radiation Control License number:
9.6 | The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade (rubber to with 2 screws) type and it must be a minimum length of three (3) meters. NB the mains cable of the unit tendered for must be SABS colour coded.
9.7 | The equipment quoted for must be protected against electromagnetic interference IEC 60601-1-2
10 | COMPREHENSIVE MAINTENANCE
10.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) should be quoted for, as per the SCC document
10.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
10.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
10.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
10.3 | The bidder to provide the OEM recommended service interval of the unit:
10.4 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
11 | General
11.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
11.2 | The unit on offer must be the latest technology.
11.2.1 | Initial Manufacturing year of Model:
| CELEBRAL OXIMETER Replacement Accessories & Consumables | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
RT2-01-22a | DC-I, Adult Reusable Sensor, 3 ft., Non-Sterile
RT2-01-22b | DCI-P, Pediatric Reusable Sensor, 3 ft., Non-Sterile
RT2-01-22c | DB-I, Adult Reusable Soft Finger Sensor, 3 ft., Non-Sterile
RT2-01-22d | TC-I, Reusable Tip-Clip Ear Sensor, 3 ft., Non-Sterile
RT2-01-22e | Y-I, Multisite Reusable sensor, 3 ft., Non-Sterile
RT2-01-22f | TF-I, Reusable Transflectance Forehead sensor, 3 ft., Non-Sterile
RT2-01-22g | Patient Cable, MD20-12, 20-pin Patient Cable, 12 ft.
RT2-01-22h | Dual Cable MD20-12
RT2-01-22i | EEG Sensor, Adult, Non-Sterile
RT2-01-22i | EEG Sensor, Pediatric, Non-Sterile
RT2-01-22j | O3 Regional Oximetry Sensors, Adult
RT2-01-22k | O3 Regional Oximetry Sensors, Pediatric
RT2-01-22l | O3 Regional Oximetry Sensors, Infant/Neonatal
RT2-01-22m | Adult/Pediatric Airway Adapter Set 3m
RT2-01-22n | Infant/Neonatal, Airway Adapter Set
RT2-01-22o | Adult Nasal CO2 Cannula
RT2-01-22p | Infant Nasal CO2 Cannula
RT2-01-22q | Adult Nasal CO2 Cannula with O2
RT2-01-22r | Pediatric Nasal CO2 Cannula with O2
RT2-01-22s | Pleth Variability Index (PVI) Parameter (License)
RT2-01-22t | Total Hemoglobin (SpHb) Parameter (License)
RT2-01-22u | Carboxyhemoglobin (SpCO) Parameter (License)
RT2-01-22v | Methemoglobin (SpMet) Parameter (License)
RT2-01-22x | Rainbow® Acoustic Monitoring (RRa™) Parameter (License)
RT2-01-22y | Acoustic Respiration Cloth Sensor, Adult/Pediatric Sensor, (RRa), Single Patient Use, Non-Sterile
RT2-01-22z | Acoustic Respiration Cloth Sensor, Inf/Neo, Acoustic Respiration Infant/Neonatal Sensor
RT2-01-22aa | Acoustic Respiration Cloth Sensor, Adult/Pediatric Adhesive Sensors, (RRa), Single Patient Use, Non-Sterile
RT2-01-22ab | Acoustic Respiration Cloth Sensor, Short Term Monitoring Adult / Pediatric Adhesive Sensors, (RRa), Single Patient Use, Non-Sterile
RT2-01-22ac | Neo, Adult/Neonatal Disposable Sensor, SpHb, SpMet, SpO2

## Ultrasound, PoC
RT2-01-23 | Ultrasound Machine, Point-of-Care | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
1 | DESCRIPTION
1.1 | A portable, user friendly medical imaging tool that enables healthcare professionals to perform ultrasound examinations at the bedside or in the clinic.
Product Offered:
1a.1 | Make:
1a.2 | Model:
1a.3 | Initial Manufacturing year of Model:
1a.4 | Radiation Control License number:
2 | TECHNICAL SPECIFICATION
2.1 | Mid-range echocardiographic machine
2.2 | Digital broadband beam-former
2.3 | Digital channels must be between 300 000 - 500 000.
2.4 | Single crystal technology or equivalent on transthoracic and Transesophageal transducers is needed, state transducer technology
2.5 | 2D Imaging
2.6 | Dynamic range more than 250dB
2.7 | Real-Time Compounding
2.8 | Adaptive Image Processing for noise and artifact reduction
2.9 | Ergonomic transducers designs with lightweight flexible cables
2.10 | High-resolution interactive graphical color touch panel with adjustment for various ambient light conditions
2.13 | At least 3 active transducer connectors and one non Imaging transducer port, preferably 4 or more
2.14 | Intelligent one-button optimization in 2D and Doppler modes
2.15 | The unit must be offered with at least 2-hour back-up battery
2.16 | The unit must have an intergrated ECG module
3 | CLINICAL APPLICATION IN SOFTWARES
3.1 | The echocardiogram machine must have the following software applications:
3.1.1 | Adult Echocardiography
3.1.2 | Pediatric Echocardiography
3.1.3 | Stress Echocardiography
3.1.4 | Transesophageal Echocardiography (optional)
3.1.5 | Full Cardiac Echocardiogram Software Package
4 | IMAGING MODES AND IMAGE DISPLAY
4.1 | The operating modes should include: 2D, M, 2D/M, PWD, High PRF, Color Doppler, Power Doppler, Directional Power Doppler, steerable CWD, triplex and Harmonic with pulse inversion technology.
4.2 | Tissue Doppler with Frame rate ranging from 400 - 600 frames per second.
4.3 | M mode color Doppler and M mode tissue Doppler
4.4 | Innovative Transducer technology with for greater acoustic efficiency and bandwidth
4.5 | Quantification of the mitral valve using the 4D / live 3D transesophageal transducer.
4.6 | Coded beam forming
4.7 | Live (on screen) automatic Doppler Analysis
4.8 | Tissue Specific Imaging
4.9 | Dual Live display with BW reference image and color Doppler image
4.10 | Colorization of volume imaging
4.11 | Trapezoid scanning providing convex view on linear probes.
4.12 | High Definition Write Zoom and Read Zoom with pan feature
4.13 | Scanning depth ranges minimum 32 cm
4.14 | Measurement tools including: distance, depth, area, and circumference
4.15 | Volume Flow Measurements
4.16 | On-board workstation-class data management with thumbnail previews and storage of images, loops, and reports
4.17 | Retrospective and prospective clip capture to internal drive or removable media
4.18 | Integrated DVD/CD burning capability for storage of DICOM images or export in JPEG and avi for PC compatibility
4.19 | DICOM 3.0 Print and Store capability to internal drive or DVD/CD
4.20 | Internal HDD capacity atleast 500GB
4.21 | Onboard quantification for :
4.22 | a. Automated Strain Imaging and Speckle tracking.
4.23 | b. Display of automated borders
4.24 | c. Display of Mitral valve annular motion
4.25 | d. Automated Inter-Medial Thickness Measurement
5 | TRANSDUCERS
5.1 | High Frequency Linear Transducer 3-12MHz
5.2 | High Frequency Linear Transducer 5-18MHz
5.3 | Adult Phased Array Transducer 2-4MHz
6 | PERIPHERALS (standard)
6.1 | Wheelbased cart:
6.1.1 | The cart must have 4 castors, two of which must have brakes
6.1.2 | The cart must have minimum two transducer holders
6.1.3 | The cart must have at least 1 drawer and 1 shelf
7 | STANDARD ACCESSORIES
7.1 | 1. Black and White photo printer
7.2 | A6 100 x sheets of printing paper
7.3 | ECG cable complete
8 | OPTIONAL ACCESSORIES
8.1 | 3. Innovative linear and curved transducer technology
8.2 | Two Full Off-line Work stations with Full Reporting Systems
8.3 | 5.Linear Carotid transducer Probe
8.4 | 6. Non-Imaging Pencil Probe
9 | Conformity Compliance (Please attach certificate)
9.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
9.2 | The unit must comply with electrical safety and performance standard such as IEC 60601-2-37, attached certification
9.2.1 | Please provide unique ref number of the IEC ceriticate:
9.4 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
9.4.1 | Please provide unique ref number of the ISO 13485 ceriticate:
9.5 | Model on offer must comply with EU Medical Device Regulation (EU) 2017/745 or equvalent from SAHPRA recognized jurisdictions . Attach a copy of certification
9.5.1 | Please provide unique ref number of the ceriticate:
9.6 | Hazardous Substance Act: If this type of equipment / apparatus appears on the schedule of Hazardous substances issued by the Directorate: Health Technology of the Department of Health, a license in terms of the Act on Hazardous Substances (Act 15/1973) must be submitted with the bid / quotation documents. The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
9.7 | Must comply with 2.5 Occupational Health and Safety Act, Act no 85 of 1993 and Regulations
9.8 | All electrical/electronic medical equipment must be licensed by Radiation Control, where applicable a copy of the license must be submitted The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
9.8.1 | Radiation Control License number:
9.9 | The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade (rubber to with 2 screws) type and it must be a minimum length of three (3) meters. NB the mains cable of the unit tendered for must be SABS colour coded.
9.10 | The equipment quoted for must be protected against electromagnetic interference IEC 60601-1-2
10 | COMPREHENSIVE MAINTENANCE
10.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) should be quoted for, as per the SCC document
10.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
10.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
10.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
11 | General
11.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
| Ultrasound Machine, PoC Accessories & Consumables | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
RT2-01-23a | High Frequency Linear Transducer 3-12MHz
RT2-01-23b | High Frequency Linear Transducer 5-18MHz
RT2-01-23c | Adult Phased Array Transducer 2-4MHz
RT2-01-23d | Black and White printer
RT2-01-23e | A6 100 x sheets of printing paper (per Box)
RT2-01-23f | ECG cable complete
RT2-01-23g | Convex transducer 3-5MHz
RT2-01-23h | Paediatric Phased array Transducer 3-7MHz
RT2-01-23i | Linear Transducer 4-10MHz
RT2-01-23j | Adult TEE tranducer (In storage case)
RT2-01-23k | Paeds TEE transducer (In storage case)
RT2-01-23l | TEE software (UPGRADE
RT2-01-23m | Cart/trolley

## Category-2 Vent & Respiratory
VENTILATION & RESPIRATORY EQUIPMENT
# | Item number | Item Description
1 | RT2-02-01 | Ventilator, Neonatal-Paeds
2 | RT2-02-02 | Ventlator, Paeds-Adult
3 | RT2-02-03 | Ventilator, Negative Pressure
4 | RT2-02-04 | Oxygen High Flow Therapy
5 | RT2-02-05 | Ventilator, CPAP
6 | RT2-02-06 | Ventilator, Transport, Paeds-Adult
7 | RT2-02-07 | Ventilator, Transport, Neonatal-Paeds
8 | RT2-02-08 | High Frequency Oscilator
9 | RT2-02-09 | Resuscitator, Gas-powered, Infant
10 | RT2-02-10 | Humidifier, Servo control
11 | RT2-02-11 | Air-O2 Blender, Low Flow
12 | RT2-02-12 | Air-O2 Blender, High Flow
13 | RT2-02-13 | Pulmonary Function Unit
14 | RT2-02-14 | Oxygen Flowmeter, Single
15 | RT2-02-15 | Oxygen Flowmeter, Double
16 | RT2-02-16 | Nebulizer, Ultrasonic
17 | RT2-02-17 | Electrical Impedance Tomography, Lung
18 | RT2-02-18 | Non-Invasive Ventilation Monitor
19 | RT2-02-19 | Spirometer, portable

## Ventilator, Neonatal-Paeds
RT2-02-01 | Ventilator: Neonatal/Paediatric | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
1 | Description
1.1 | A device designed to provide respiratory support to newborn, infants, and children who require mechanical ventilation.
1a | Product Offered:
1a.1 | Make:
1a.2 | Model:
1a.3 | Initial Manufacturing year of Model:
1a.4 | Radiation Control License number:
2 | Technical Specification
2.1 | The ventilator must be pneumatically powered
2.2 | The ventilator must be microprocessor controlled
2.3 | The ventilator must be a Possitive Pressure ventilator
2.4 | Multi-mode state of the art touch screen or equivalent graphical interface
2.5 | Ability to display Airway graphics and Respiratory/Lung Function parameters. (This can be done by paging through the screen)
2.6 | Breath triggering: Pressure or Flow
2.7 | Breath termination: Time or Flow
2.8 | Adjustable I:E Ratio Normal 1:2 (Range 1:1-1:4) Inspiratory Time: 0.1 to 3 seconds (Range 0.11-10s)
2.9 | Gas flow: Decelerating and Constant (Range 10 - 200 L/min)
2.10 | The ventilator must provide the following ventilation modes (Invasively):
2.10.1 | Pressure Limited Mode (Pressure Control) + SIMV
2.10.2 | Volume Limited Mode ( Volume Control) + SIMV
2.10.3 | PEEP
2.10.4 | CPAP (0-40)
2.10.5 | Pressure Support Ventilation (PSV) + PEEP
2.10.6 | APRV
2.10.7 | PRVC/Volume Target Ventilation (Optional)
2.11 | The ventilator must provide the following ventilation modes (Non-Invasively):
2.11.1 | Nasal CPAP
2.11.2 | Mask CPAP
2.11.3 | Bi-PAP
2.11.4 | Hi-Flow Oxygen
2.12 | Integrated Nebulizer (Optional)- Ultrasonic
2.13 | Inspiratory Hold capability up to 30 sec
2.14 | Expiratory Hold capability up to 30 sec
2.15 | Back up Mode if Apnea occurs in all modes
2.16 | Back up/ Apnea ventilation in all modes
2.17 | Pre-Oxygenation capability with 100%
2.18 | Respiratory rate/min range 1-160/min
2.19 | Peak Inspired Pressure: 0 to at least 60cm H2O
2.20 | Inspratory Flow: 0,4 to 100 L/min
2.21 | Pause adjustable from 0 to 2 sec
2.22 | Flow Trigger: 0,1 to 5 L/min
2.23 | Proximal trigger option for neonate/infant.
2.24 | Pressure Trigger: 0.5 to 10 mbar below PEEP/CPAP.
2.25 | Pressure High: 0 to 70 mbar.
2.26 | Pressure Low: 0 to 45 mbar.
2.27 | Oxygen Concentration from 21% to 100%.
2.28 | Bias flow.
2.29 | Leak Compensation: Up to 8 L/min.
2.30 | Alarms:
2.30.1 | Both audible and visual alarms.
2.30.2 | Patient disconnect
2.30.3 | Oxygen concentration variation
2.30.4 | Low inspiratory pressure.
2.30.5 | High breath rates
2.30.6 | Incompatible settings.
2.30.7 | Low PEEP.
2.30.8 | Over and under patient pressure
2.30.9 | Over and under patient volume
2.30.10 | High continuous pressure
2.30.11 | Inverse I:E ratio
2.30.12 | Air/Oxygen % deviation
2.30.13 | Power failure
2.30.14 | Air and Oxygen supply failure
2.30.15 | Air and Oxygen low pressure.
2.30.16 | APNEA.
2.30.17 | Alarm mute. (Resets after 60 seconds.)
2.30.18 | Low battery.
2.31 | The unit must have automatic switch over to a built in battery with a capacity to ensure full functionality of the unit for not less than 2 hours.
2.32 | The unit must have automatic switch over that will switch between the piped air and oxygen supply should one of the supplies fail.
2.33 | The unit must have a circuit compliance compensation 0.0 to 7.5 ml/cmH2O and must be automatically measured.
2.34 | Proximal flow and pressure sensing technology for neonates.
2.36 | Graphics Display: The following parameters and measurements must be displayed:
2.36.1 | Inspired tidal volume in ml
2.36.2 | Expired tidal volume in ml
2.36.3 | Peak inspiratory pressure.
2.36.4 | Proximal airway pressure.
2.36.5 | Inspired tidal volume per kg in ml/kg
2.36.6 | Expired tidal volume per kg in ml/kg
2.36.7 | Spontaneous tidal volume in ml
2.36.8 | Total breath rate in bpm
2.36.9 | Spontaneous breath rate in bpm
2.36.10 | Minute volume in litres
2.36.11 | Spontaneous minute volume in litres
2.36.12 | Peak airway pressure in cmH2O
2.36.13 | Mean airways pressure in cmH2O
2.36.14 | Plateau pressure in cmH2O
2.36.15 | I:E ratio
2.36.16 | Set I:E ratio
2.36.17 | Inspired time in seconds
2.36.18 | Expired time in seconds
2.36.19 | Static airways compliance in ml/cmH2O
2.36.20 | Static airways compliance per kg in ml/cmH2O
2.36.21 | Dynamic airways compliance in ml/cmH2O
2.36.22 | Dynamic airways compliance per kg in ml/cmH2O
2.36.23 | Peak inspired flow rate in l/m
2.36.24 | Peak expired flow rate in l/m
2.36.25 | Inspired O2 in %
2.36.26 | Tube leak in %
2.37 | At least 3 selectable waveforms must be displayed at any time on the graphic display screen.
2.38 | The graphics display system must have freeze capabilities.
2.39 | Trending of at least 24 hours on at least 8 parameters of pressure- flow and tidal volumes.
2.40 | The unit must be able to calculate and display pressure- and flow/volume loops.
2.41 | The ventilator must be mounted on a stable trolley with at least two of the castors fitted with brakes that conform to SANS 601.
2.42 | The unit must have circuit holding arm
2.43 | The ventilator trolley must be suitable for mounting of the dual servo humidifier
2.44 | The unit must be supplied a pole or universal clamp unit to mount the exiting humidifiers in use onto the unit.
2.45 | Mains cable at least three metres long with SABS approved 15A three prong plug.
2.46 | The ventilator must come complete with colour-coded air and O2 hoses 3-metre length. (O2 and air Heyer type keys, SABS 1409, must also be supplied.)
2.47 | The unit must have a rail and circuit hanger of a suitable strength on both sides of the ventilator trolley.
2.48 | The unit must have an integrated blender with O2 concentration 21 to 100 % ± 3 %.
2.49 | The unit must be handed over in full operating order
3 | A starter-pack of consumables must be supplied with the unit comprising at least the following:
3.1 | Bacteria filter (pk of 10)
3.2 | Neonatal patient circuit (Pk of 10)
3.3 | Paeds patient circuits (Pk of 10)
3.4 | Test lung for neonates
3.5 | Two reusable flow sensors
3.6 | Circuit support arm
4 | Conformity Compliance (Please attach certificate)
4.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
4.2 | The unit must comply with electrical safety and performance standard such as IEC 60601-2-12, attached certification
4.2.1 | Please provide unique ref number of the IEC ceriticate:
4.3 | OEM must comply and certified on ISO 9001 quality standards, attach proof of compliance
4.3.1 | Please provide unique ref number of the ISO 9001 ceriticate:
4.4 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
4.4.1 | Please provide unique ref number of the ISO 13485 ceriticate:
4.5 | Model on offer must comply with EU Medical Device Regulation (EU) 2017/745 or equvalent from SAHPRA recognized jurisdictions . Attach a copy of certification
4.5.1 | Please provide unique ref number of the ceriticate:
4.6 | Hazardous Substance Act: If this type of equipment / apparatus appears on the schedule of Hazardous substances issued by the Directorate: Health Technology of the Department of Health, a license in terms of the Act on Hazardous Substances (Act 15/1973) must be submitted with the bid / quotation documents. The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
4.7 | All electrical/electronic medical equipment must be licensed by Radiation Control, where applicable a copy of the license must be submitted The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
4.7.1 | Radiation Control License number:
4.8 | The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade (rubber to with 2 screws) type and it must be a minimum length of three (3) meters. NB the mains cable of the unit tendered for must be SABS colour coded.
4.9 | The equipment quoted for must be protected against electromagnetic interference IEC 60601-1-2
5 | COMPREHENSIVE MAINTENANCE
5.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) should be quoted for, as per the SCC document
5.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries, Flow sensors and Oxygen sensor), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
5.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
5.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries, Flow sensors and Oxygen sensor), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
5.3 | The bidder to provide the OEM recommended service interval of the unit:
5.4 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
6 | General
6.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
6.2 | The unit on offer must be the latest technology.
6.2.1 | Initial Manufacturing year of Model:
| Neonatal/Paediatric Ventilator Accessories, Consumables & Additional options | Complies Yes/No
RT2-02-01a | Addition of HFOV Mode (UPGRADE)
RT2-02-01b | Bacteria filter (pk of 10)
RT2-02-01c | Neonatal patient circuit (Pk of 10)
RT2-02-01d | Paeds patient circuits (Pk of 10)
RT2-02-01e | Test lung for neonates
RT2-02-01f | Disposable Flow sensors
RT2-02-01f | Reusable Flow sensors
RT2-02-01g | Oxygen cell
RT2-02-01h | Medical Air compressor (Mountable to ventilator trolley)

## Ventilator, Paeds-Adult
RT2-02-02 | VENTILATOR PAEDIATRIC/ADULT | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
1 | Description
1.1 | A device designed to provide respiratory support to adults and children who require mechanical ventilation.
1 | Product Offered:
1a.1 | Make:
1a.2 | Model:
1a.3 | Initial Manufacturing year of Model:
1a.4 | Radiation Control License number:
2 | Technical Specification
2.1 | The unit must be able to ventilate and monitor paediatric and adult patients
2.2 | The ventilator must be a time-cycled, volume- and pressure-controlled unit.
2.3 | The ventilator must have an integrated, electronically controlled air compressor or turbine
2.4 | The ventilator must have an oxygen port for use with an oxygen cylinder during transport
2.5 | Suitable preset pin indexed regulators must be supplied that can handle flows up to 100 L/pm
2.6 | The unit must have automatic switch-over unit that will switch between the piped air and turbine/oxygen supply should one of the supplies fail.
2.7 | Self diagnosis with each start-up and regular integrity testing of all system parameters
2.8 | Ventilator must have "Standby" mode which suspends alarms, breath cycling and trends.
2.9 | "Standby" mode must retain most current ventilator settings and alarms, resume normal ventilation with one touch
2.10 | The following breath types must be selectable:
2.10.1 | Volume
2.10.2 | Pressure
2.10.3 | Pressure regulated volume control
2.10.4 | Pressure support ventilation
2.10.5 | Spontaneous
2.11 | The ventilator must have at the least the following modes of ventilation:
2.11.1 | Invasive and non-invasive ventilation must be possible
2.11.2 | Volume ventilation with:
2.11.2.1 | Assist/control
2.11.2.2 | SIMV (Synchronous Intermittent Mandatory Ventilation)
2.11.2.3 | PSV (Pressure Support Ventilation)
2.11.2.4 | Option to disable demand flow
2.11.2.5 | Option of square or decelerating flow waveform
2.11.3 | Pressure ventilation (IPPV) with:
2.11.3.1 | Assist/control
2.11.3.2 | SIMV (Synchronous Intermittent Mandatory Ventilation)
2.11.3.3 | PSV (Pressure Support Ventilation)
2.11.3.4 | CPAP
2.11.3.5 | NPPV (Non-invasive positive pressure ventilation) with leak compensation and apnea back-up
2.11.4 | PRVC (Pressure Regulated Volume Control)
2.11.5 | APRV (Airway Pressure Release Ventilation) /BiPAP ( Biphasic , Duo Level, BiLevel)
2.11.6 | Spontaneous ventilation
2.11.7 | Apnea monitoring must be active in A/C, SIMV, CPAP and APRV
2.11.7.1 | Apnea mode: user settable time duration before ventilation onset
2.12 | BREATH CONTROL
2.12.1 | Breath rate: 1 - 150 bpm
2.12.2 | Tidal volume: 50 - 3000 ml
2.12.3 | Peak flow: 1 - 120 litres/min
2.12.4 | Volume limit: 5 ml - 2,5 litres
2.12.5 | PEEP: 0 - 50 cm H2O
2.12.6 | Inspiratory flow trigger: 0.5 - 15 litres/min
2.12.7 | Pressure trigger: 0,5 - 10 mbar below PEEP/CPAP
2.12.8 | Bias flow: 0.4 - 5 litres/min
2.12.9 | Expiratory hold: 0 - 5 sec
2.12.10 | Inspiratory hold: 0 - 5 sec
2.12.11 | Inspiratory pause: 0 - 2 sec
2.12.12 | Inspiratory time: 0,2 - 5 sec
2.12.13 | Manual breath - add single control breath
2.12.14 | Inspiratory waveform: square and decelerating waves
2.12.15 | Inspiratory pressure control level: 1 - 100 cm H2O
2.13 | In ARPV mode:
2.13.1 | Pressure high: 0 - 60 cm H2O
2.13.2 | Time high: 0,2 - 30 sec
2.13.3 | Pressure low: 0 - 40 cm H2O
2.13.4 | Time low: 0,3 - 30 sec
2.14 | Leakage compensation in all non-invasive ventilation modes must be possible.
2.15 | The unit must have a circuit compliance compensation 0 to 7.5 ml/cmH2O and must be automatically measured.
2.16 | I:E ratio: 4:1 to 1:4
2.17 | PEEP/CPAP must be possible from 0 to approximately 50 cm H2O
2.18 | OXYGEN CONTROL
2.18.1 | FiO2: 21 – 100%
2.18.2 | Increase % oxygen: single key activation to deliver up to100% O2 for 2 min. FiO2 to be adjustable by user
2.18.3 | Ventilator must have an integrated nebulizer port capable of delivering a set FiO2, synchronized with the inspiratory phase. Must be timed and volume compensated with a peak flow ≥ 15 L/min
2.19 | MONITOR
2.19.1 | Volume monitoring:
2.19.1.1 | Monitored values to include Tidal and Minute Volumes divided by patent's weight (ml/kg)
2.19.1.2 | Tidal volume
2.19.1.3 | Minute volume
2.19.1.4 | Mandatory expired tidal volume: 0 - 4 litres
2.19.1.5 | Spontaneous expired tidal volume
2.19.1.6 | Peak Inspiratory Flow Rate (PIFR): 0 to 300 litres/min
2.19.1.7 | Peak Expiratory Flow Rate (PEFR): 0 to 300 litres/min
2.19.2 | Pressure monitoring:
2.19.2.1 | Peak inspiratory pressure: 0 - 120 cm H2O
2.19.2.2 | Mean airway pressure: 0 - 120 cm H2O
2.19.2.3 | Plateau pressure: 0 - 100 cm H2O
2.19.2.4 | PEEP: 0 - 50 cm H2O
2.19.2.5 | Oxygen inlet pressure: 40 - 95 psig
2.19.2.6 | Delivered O2: 21 - 100 %
2.19.3 | Time cycle monitoring:
2.19.3.1 | Breath rate: 0 - 150 bpm
2.19.3.2 | Spontaneous breath rate: 0 - 150 bpm
2.19.3.3 | Inspiratory time: 0 - 99 sec
2.19.3.4 | Exhalation time: 0 - 99 sec
2.19.3.5 | I:E ratio: 1:4 - 4:1
2.19.3.6 | Rapid Mustow breathing index 0 - 500 bpm/litre
2.20 | The ventilator must have a colour LCD screen with at least three simultaneous real-time waveforms
2.21 | The following parameters must be available as time-base curves:
2.21.1 | Pressure
2.21.2 | Flow
2.21.3 | Tidal volume
2.22 | Display of the following patient parameters must be user selectable and at least three selected parameters must be displayed as digital values
2.22.1 | Spontaneous minute volume
2.22.2 | Inspiratory flow
2.22.3 | Peak airway pressure
2.22.4 | Mean airway pressure
2.22.5 | Plateau pressure
2.22.6 | PEEP
2.22.7 | Expired tidal volume
2.22.8 | Expired minute volume
2.22.9 | Inspiratory time
2.22.10 | Breath rate
2.22.11 | Spontaneous rate
2.22.12 | Internal oxygen inlet pressure (pipeline inlet pressure)
2.22.13 | Battery charge status
2.22.14 | I:E ratio
2.22.15 | Set I:E ratio
2.22.16 | Static airways compliance in ml/cmH2O
2.22.17 | Static airways compliance per kg in ml/cmH2O
2.22.18 | Dynamic airways compliance in ml/cmH2O
2.22.19 | Dynamic airways compliance per kg in ml/cmH2O
2.22.20 | Peak inspired flow rate in l/m
2.22.21 | Peak expired flow rate in l/m
2.22.22 | Inspired O2 in %
2.22.23 | Tube leak in %
2.23 | The ventilator must measure and display inspiratory/expiratory loops.
2.24 | Waveforms and loops must be colour coded to denote inspiratory and expiratory breath phases
2.25 | The graphics display system must have a freeze capability.
2.26 | Loops: Flow/Volume Loop, Pressure/Volume Loop
2.27 | Reference loops must be able to be stored and displayed in the background in different colours
2.28 | All primary controls must be visible and changeable from all screens without entering a menu
2.29 | Monitored values can be arranged to match sequence of hospital's ventilator check sheet
2.30 | SETUP
2.30.1 | Patient details
2.30.2 | Patient size: Adult / Paediatric
2.30.3 | Patient weight
2.31 | TRENDS
2.31.1 | Trends: at least 8 parameters for at least 24 hours
2.31.2 | Trends must be configurable to display at least 4 values as histograms with ability to review with cursor/time stamp
2.32 | Automatically recorded events:
2.32.1 | New patient
2.32.2 | Settings Change
2.32.3 | Increase 02
2.32.4 | Power On
2.32.5 | Insp Hold
2.32.6 | Power Off
2.32.7 | Exp Hold
2.32.8 | Power Recovery
2.32.9 | Nebulizer
2.32.10 | Enter Standby
2.32.11 | Manual Breath
2.32.12 | Exit Standby
2.33 | ALARMS: The ventilator must have the following audible and visual alarms:
2.33.1 | Low/high airway pressure
2.33.2 | Low/high minute volume
2.33.3 | Low/high tidal volume
2.33.4 | Low/high breath rate:1 to 200 bpm
2.33.5 | Low/high FiO2
2.33.6 | Loss of PEEP
2.33.7 | Low/high PEEP
2.33.8 | Low/high PIP
2.33.9 | Leakage alarm
2.33.10 | Apnea: 10 - 60 sec
2.33.11 | Patient disconnect
2.33.12 | Oxygen concentration variation (± 5%)
2.33.13 | Air/oxygen % deviation
2.33.14 | Power failure
2.33.15 | Air and oxygen supply failure
2.33.16 | Air and oxygen low pressure
2.33.17 | Battery low (When approximately 15 min operation left).
2.33.18 | Apnea
2.34 | Alarm limits must be user-adjustable
2.35 | Default alarm limits if user does not set alarm limits
2.36 | Alarm tone differentiates high, medium and low priority alarms
2.37 | Adjustable alarm volume
2.38 | Temporary mute for audible alarm: 0 - 2 min
2.39 | Alarm Reset. Deactivates visual indicators for resolved alarm conditions
2.40 | Machine Alarms:
2.40.1 | Safety Valve Open
2.40.2 | Ventilator Inoperative
2.40.3 | Fan Failure
2.40.4 | Loss of Air -Air Supply pressure < 18.0 psig
2.40.5 | Loss of O2 – O2 Supply pressure <18.0 psig (when % O2 > 21)
2.41 | GAS SUPPLY
2.41.1 | Oxygen pipeline supply: Pressure: 200 to 600 kPa) Minimum flow: 80 L/min at 1.5 bar
2.41.2 | Air pipeline supply: Pressure: 200 to 600 kPa) Minimum flow: 80 L/min at 1.5 bar
2.41.3 | Compressor air: 0.2 to 0.7 bar
2.41.4 | Medical air and oxygen hoses, 3m long, colour coded, mounted with Heyer oxygen and medical air probes. The Heyer probes must be compatible with the wall points used in the instutution.
2.42 | POWER SUPPLY
2.42.1 | The unit must be powered by 220V, 50 Hz mains supply
2.42.2 | The unit must have automatic switch over to a built in battery with a capacity to ensure full functionality of the unit for not less than 2 hours.
2.42.3 | Automatic switch to battery in case of power failure,
2.42.4 | Automatic battery recharge when connected to mains
2.42.5 | Power cord at least 3 m long with SABS approved 15A three prong plug
2.43 | GENERAL
2.43.1 | The ventilator must be mounted on a trolley with at least 4 antistatic swivel castors of at least 125mm diameter of which 2 must be lockable.
2.43.2 | Trolley must have a push-handle
2.43.3 | The trolley must have an oxygen cylinder holder
2.43.4 | The trolley must have a rail on both sides to fit the circuit arm
2.43.5 | Removeable arm with circuit hanger to fit the side rails
2.43.6 | Expiratory inlet to include integral water trap, valved to allow emptying without interruption of normal patient
2.43.7 | Air filter capacity at inlet: 99 % (for > 0.5 um)
2.43.8 | State lifetime and price of Oxygen sensor
2.43.9 | The unit must be handed over in full operating order
2.44 | The following Accessories and Consumables must be supplied with the unit:
2.44.1 | Two reusable patient breathing circuits
2.44.2 | Built-in nebulizer
2.44.3 | Reusable, sterilizable patient masks
2.44.4 | Reusable, sterilizable patient connectors
2.44.5 | Reusable flow sensor.
2.44.6 | Internal O2-Air mixer.
2.44.7 | FiO2 analyzer, internal.
2.44.8 | Test lung and test circuit
2.44.9 | Bacteria filter (pk of 10)
2.44.10 | Adult disposable patient circuit (Pk of 10)
2.44.11 | Paeds disposable patient circuits (Pk of 10)
2.44.12 | Air filters (pk of 10)
2.44.13 | Flow sensor (pk of 5)
2.44.14 | Oxygen sensor
2.44.15 | Expiratory valve (complete)
2.46 | Optional Items
2.46.1 | MANEUVERS
2.46.1.1 | Ventilator must provide separate manoeuvres for MIP/NIF P100, AutoPEEP airway, Slow Flow Loop Inflection Points
2.46.1.2 | Ventilator must return to normal ventilation at end of manoeuvre without any changes required to primary controls
2.46.1.3 | Ventilator must have "Sensitivity" setting for maneuvre only, separate from primary "Sensitivity" setting. Allowing Patient to abort maneuvre.
2.46.1.4 | Inflection point maneuvre must calculate upper and lower inflection points and display on Pressure/Volume loop
2.46.1.5 | Stored inflection point loops must be able to be superimposed on current loop for comparison
2.46.2 | Humidifier
2.46.2.1 | The humidifier must be dual servo controlled with an alarm system
2.46.2.2 | All accessories for the humidifier must be included
2.46.2.3 | Humidifier must have a hose heating element, dual temperature probe and electrical adapter.
2.46.4 | End Tidal CO2 curve
2.46.5 | Printer
3 | Conformity Compliance (Please attach certificate)
3.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
3.2 | OEM must comply and certified on ISO 9001 quality standards, attach proof of compliance
3.2.1 | Please provide unique ref number of the ISO 9001 ceriticate:
3.3 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
3.3.1 | Please provide unique ref number of the ISO 13485 ceriticate:
3.4 | Model on offer must comply with EU Medical Device Regulation (EU) 2017/745 or equvalent from SAHPRA recognized jurisdictions . Attach a copy of certification
3.4.1 | Please provide unique ref number of the ceriticate:
3.5 | Hazardous Substance Act: If this type of equipment / apparatus appears on the schedule of Hazardous substances issued by the Directorate: Health Technology of the Department of Health, a license in terms of the Act on Hazardous Substances (Act 15/1973) must be submitted with the bid / quotation documents. The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
3.6 | All electrical/electronic medical equipment must be licensed by Radiation Control, where applicable a copy of the license must be submitted The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
3.6.1 | Radiation Control License number:
3.7 | The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade (rubber to with 2 screws) type and it must be a minimum length of three (3) meters. NB the mains cable of the unit tendered for must be SABS colour coded.
3.8 | The equipment quoted for must be protected against electromagnetic interference IEC 60601-1-2
3.9 | The unit must comply with electrical safety and performance standard such as IEC 60601-2-12, attached certification
3.9.1 | Please provide unique ref number of the IEC ceriticate:
4 | COMPREHENSIVE MAINTENANCE
4.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) should be quoted for, as per the SCC document
4.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
4.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
4.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
4.3 | The bidder to provide the OEM recommended service interval of the unit:
4.4 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
5 | General
5.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
5.2 | The unit on offer must be the latest technology.
5.2.1 | Initial Manufacturing year of Model:
| Ventilator, Paeds/Adults Accessories and Consumables | Complies Yes/No
RT2-02-02a | Reusable patient breathing circuits
RT2-02-02b | Reusable, sterilizable patient masks
RT2-02-02c | Reusable, sterilizable patient connectors
RT2-02-02d | Reusable flow sensor.
RT2-02-02e | Test lung
RT2-02-02f | Bacteria filter (pk of 10)
RT2-02-02g | Adult disposable patient circuit (Pk of 10)
RT2-02-02h | Paeds disposable patient circuits (Pk of 10)
RT2-02-02i | Air filters (pk of 10)
RT2-02-02j | Flow sensor (pk of 5)
RT2-02-02k | Oxygen sensor
RT2-02-02l | Expiratory valve (complete)

## Negative Pressure Ventilator
RT2-02-03 | Non-Invasive Negative Pressure Ventilator, Biphasic Cuirass | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
1 | DESCRIPTION
1.1 | A ventilator used to perform an efficient and effective method of non-invasive external ventilation and function as a real alternative to traditional forms of ventilation by utilising a hard clear plastic shell (cuirass) as well as a secretion clearance device with cough assistance
1a | Product Offered:
1a.1 | Make:
1a.2 | Model:
1a.3 | Initial Manufacturing year of Model:
1a.4 | Radiation Control License number:
2 | TECHNICAL SPECIFICATION
2.1 | The unit must provide the following:
2.2 | Lightweight and compact for ease of placement in an ED, ICU or operating theatre
2.3 | Extremely durable to ensure repeated use
2.4 | Power Supply:
2.4.1 | External power supply: 220V model @ 50 Hz
2.4.2 | Power supply cable provided which is compatible with RSA power sockets
2.4.3 | The unit must have internal rechargeable NiMH battery
2.5 | Ventilator Specifications
2.5.1 | Frequency: 6 -1200 cycles per minute
2.5.2 | I:E ratio 1:6 - 6: 1
2.5.3 | Maximum Inspiratory Pressure: -50cmH20
2.5.4 | Maximum Expiratory Pressure: +50cmH20
2.6 | Modes of Operation to include:
2.6.1 | Control (Mandatory Controlled Ventilation)
2.6.2 | Respiratory Synchronised
2.6.3 | Respiratory Triggered
2.6.4 | Secretion Clearance
2.6.5 | Continuous Negative pressure ventilation
3 | STANDARD ACCESSORIES
3.1 | Cuirasses:
3.1.1 | Range of adult and paediatric sizes available (two sizes as per End-User's preference should come standard with every ventilator unit)
3.2 | Ventilator circuits (pack of 10)
4 | Conformity Compliance (Please attach certificate)
4.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
4.3 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
4.3.1 | Please provide unique ref number of the ISO 13485 ceriticate:
4.4 | Model on offer must comply with EU Medical Device Regulation (EU) 2017/745 or equvalent from SAHPRA recognized jurisdictions . Attach a copy of certification
4.4.1 | Please provide unique ref number of the ceriticate:
4.5 | Hazardous Substance Act: If this type of equipment / apparatus appears on the schedule of Hazardous substances issued by the Directorate: Health Technology of the Department of Health, a license in terms of the Act on Hazardous Substances (Act 15/1973) must be submitted with the bid / quotation documents. The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
4.6 | All electrical/electronic medical equipment must be licensed by Radiation Control, where applicable a copy of the license must be submitted The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
4.6.1 | Radiation Control License number:
4.7 | The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade (rubber to with 2 screws) type and it must be a minimum length of three (3) meters. NB the mains cable of the unit tendered for must be SABS colour coded.
4.8 | The equipment quoted for must be protected against electromagnetic interference IEC 60601-1-2
4.9 | The unit must comply with electrical safety and performance standard such as IEC 60601-2-12, attached certification
4.9.1 | Please provide unique ref number of the IEC ceriticate:
5 | COMPREHENSIVE MAINTENANCE
5.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) should be quoted for, as per the SCC document
5.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries, Flow sensors and Oxygen cell), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
5.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
5.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries, Flow sensors and Oxygen cell), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
5.3 | The bidder to provide the OEM recommended service interval of the unit:
5.4 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
6 | General
6.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
6.2 | The unit on offer must be the latest technology.
6.2.1 | Initial Manufacturing year of Model:
| Non-Invasive Negative Pressure Ventilator Accessories and Consumables | Complies Yes/No
RT2-02-03a | Cuirasses (Adult size)
RT2-02-03b | Cuirasses (Paeds size)
RT2-02-03c | Patient circuit (Pk of 5)

## Oxygen High Flow Therapy
RT2-02-04 | Oxygen High Flow Therapy Unit | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
1 | DESCRIPTION
1.1
2 | Product Offered:
2.1 | Make:
2.2 | Model:
2.3 | Initial Manufacturing year of Model:
3 | SPECIFICATIONS
3.1 | The unit must be able to be used for:
3.1.1 | Oxygen therapy
3.1.2 | Humidification
3.1.3 | Endotracheal intubation
3.1.4 | Tracheotomy
3.2 | The unit offered must be portable and ability to place on stand
3.3 | The unit must allow adjustment oxygen flow from minimum 2 L/min and maximum of 80 L/min
3.4 | The unit must have Oxygen concentration from 21% to 100%.
3.5 | While in operation the set temperature on the unit must be able to adjusted between 29 oC to 37 oC
3.6 | Bidders to state maximum temperature deviation from set temperature
3.7 | It shall be easy to refill the fluid, i.e. Autofill chamber with constant compressible volume.
3.8 | The unit must indicate the therapy time on a display
3.9 | The unit must have a maximum output pressure of 25 cmH2O
3.10 | The noise level when operating on maximum flow must not exceed 45 dB
3.11 | The following alarms must be possible both visual and audible:
3.11.1 | Power failure
3.11.2 | Check water
3.11.3 | Low FiO2
3.11.4 | High FiO2
3.11.5 | Low temperature
3.11.6 | High temperature
4 | Standard Accessories and Consumables
4.1 | Auto-fill chamber
4.2 | Heating tube
4.3 | Elbow kit
4.4 | Tracheal inbation interface
4.5 | Air filter
4.6 | Small Nasal Cannula
4.7 | Medium Nasal Cannula
4.8 | Large Nasal Cannula
4.9 | Oxygen inlet hose
5 | Conformity Compliance (Please attach certificate)
5.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
5.2 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
5.3.1 | Please provide unique ref number of the ISO 13485 ceriticate:
5.4 | Model on offer must comply with EU Medical Device Regulation (EU) 2017/745 or equvalent from SAHPRA recognized jurisdictions . Attach a copy of certification
5.4.1 | Please provide unique ref number of the ceriticate:
6 | COMPREHENSIVE MAINTENANCE
6.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) must be included in the bid price
6.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
7 | General
7.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
7.2 | The unit on offer must be the latest technology.
| Oxygen High Flow Therapy -Accessories and Consumables | Provide your answers in this Column. You are advised to be straight to the point | Complies Yes/No
RT2-02-04a | Battery
RT2-02-04b | Auto-fill chamber
RT2-02-04c | Heating tube
RT2-02-04d | Elbow kit
RT2-02-04e | Tracheal inbation interface
RT2-02-04f | Air filter
RT2-02-04g | Small Nasal Cannula
RT2-02-04h | Medium Nasal Cannula
RT2-02-04i | Large Nasal Cannula
RT2-02-04j | Oxygen inlet hose

## Ventilator, CPAP
RT2-02-05 | VENTILATOR CPAP | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
1 | Description
1.1 | A CPAP (Continous Positive Airway Pressure) Ventilator is a medical device that delivers a constant flow of air pressure into the airways to keep them open during sleep.
1a | Product Offered:
1a.1 | Make:
1a.2 | Model:
1a.3 | Initial Manufacturing year of Model:
1a.4 | Radiation Control License number:
2 | Technical Specification
2.1 | An electronically controlled & monitored CPAP flow system to safely and effectively deliver fresh gas via a nasal interface or mask to the airways of neonate and paediatric patients.
2.2 | Electrically powered and microprocessor driven, monitored and controlled.
2.3 | The CPAP system must consist of the following:
2.3.1 | Blender
2.3.2 | O2 analyser
2.3.3 | Flowmeter
2.3.4 | Electronic pressure gauge
2.3.5 | Alarm system
2.4 | Manual breath inspiration/hold feature with ability to ensure sustained lung inflation which is used as a recruitment manoeuvre on infants with atelectasis
2.5 | Manual time for a manual breath adjustable from 2 to 10 sec for this manoeuvre
2.6 | Manual breath delivered in any mode adjustable 5 - 15 cm H2O
2.7 | Oxygen flush to ensure good oxygenation during suctioning etc 23 - 100%
2.8 | Must have leak compensation to allow for various cannulae and prongs to be used
2.9 | Able to deliver Demand flow in nCPAP thereby reducing the work of breathing and not have pre-set limited flow in nCPAP
2.10 | CPAP pressure range: 0 to 35 cm H2O
2.11 | nCPAP 2 - 13 cm H2O
2.12 | O2 flush time adjustable 30 - 120 sec
2.13 | The unit must be able to supply two independent pressure levels via 2 flow meters.
2.14 | Correctly coded high pressure gas supply hoses for oxygen and air of three metre length must be supplied with the unit. Appropriate type oxygen and air probes to be fitted. Probes to be in accordance with SABS 1409.
2.15 | Supply air/oxygen mixture to the patient at a flow rate between 0 – 15 lpm (95% accuracy)
2.16 | The control unit must automatically go through a self diagnosis and checking process.
2.17 | The control unit must have direct input connections for oxygen and medical air
2.18 | The source for oxygen either from a central wall outlet (gas bank) or an oxygen cylinder.
2.19 | The source for medical air from a central wall outlet (gas bank) and medical air cylinder.
2.20 | The input supply pressure for oxygen and medical air: 200 – 500 kPa.
2.21 | The unit must automatically switch between the piped air and oxygen supply should one of the supplies fail. (Alarm condition should be triggered in any event)
2.22 | Must provide compensation for leaks
2.23 | The unit must be user friendly with all circuit connections, front panel controls and displays ergonomically grouped by function and clearly marked.
2.24 | Oxygen concentration setting variable between 21 % and 100 %.
2.25 | Have a manometer with a range of 0 – 12 cm H2O
2.26 | Breath rate: 1 - 120 bpm
2.27 | Timed inspiration: 0.1 to 3.0 sec.
2.28 | Apnea backup ventilation: 1 - 80 bpm.
2.29 | Apnea interval: 10 - 30 seconds.
2.30 | Maximum safety pressure set to ± 3 cm H20 above measured airway pressure.
2.31 | Unit must have safety features to prevent:
2.31.1 | Volutrauma
2.31.2 | Barotrauma
2.31.3 | Over stretching of the lungs
2.32 | The unit must have a proximal pressure monitoring line.
2.33 | Must have a method to prevent it becoming occluded by water vapour.
2.34 | Must monitor the following parameters:
2.34.1 | nCPAP
2.34.2 | MAP
2.34.3 | PIP
2.34.4 | PEEP
2.34.5 | Fi02 (21% to 100%)
2.35 | Graphics Display
2.35.1 | Measured values to be displayed:
2.35.1.1 | CPAP
2.35.1.2 | PIP
2.35.1.3 | MAP
2.35.2 | The graphics display system must have freeze capabilities.
2.36 | The offered unit must display trending of at least 72 hours on at least 20 parameters of pressure-, flow- and tidal volumes.
2.37 | Alarms:
2.37.1 | Patient disconnection
2.37.2 | High airway pressure
2.37.3 | Low airway pressure.
2.37.4 | High oxygen alarm
2.37.5 | Low oxygen alarm
2.37.6 | Low nCPAP alarm: -2 to -20 cm H2O
2.37.7 | High nCPAP alarm: +2 to + 20 cm H2O
2.37.8 | Oxygen supply failure
2.37.9 | Air supply failure
2.37.10 | Apnoea: off; 10 - 20 sec
2.37.11 | Electrical supply failure
2.37.12 | Low battery
2.38 | A safety mechanical pressure release (overpressure) valve must be present
2.39 | The above alarm parameters must auto set themselves within 2 minutes of power up of drive gas flows and oxygen concentrations or by means of a single key setting.
2.40 | Violation of any of the set alarm limits must result in both audible and visual alarms.
2.41 | It must be possible to silence the alarm for a maximum of 120 seconds.
2.42 | The unit must be mounted on a mobile stand/trolley, with 4 to 5 castors (all lockable) with a facility to connect a humidifier.
2.43 | The unit must have a rail and circuit hanger on both sides of the trolley.
2.44 | The unit must operate from a 220 V, 50Hz supply.
2.45 | Mains cable at least three metres long with SABS approved 15A three prong plug.
2.46 | The offered unit shall have automatic switch over to a built in battery unit with a capacity to ensure full functionality of the unit for not less than 2 hour.
2.47 | The unit must be fitted with an internal rechargable battery.
2.48 | Battery must be capable of at least 2 hours operation
2.49 | Battery charger to be incorporated in the unit
2.50 | Must display battery status and give an alarm when battery is low
2.51 | The unit must be handed over in full operating order
3 | Standard Accessories and Consumables
3.1 | 5 sets of silicone nasal prongs (each set must include 3 different neonatal sizes)
3.2 | 5 sets of CPAP bonnets (each set to include 3 different neonatal sizes)
3.3 | 5 neonatal circuits
3.4 | Humidifier
3.5 | Temperature probes for humidifier
3.6 | Circuit holding arm
4 | Conformity Compliance (Please attach certificate)
4.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
4.2 | OEM must comply and certified on ISO 9001 quality standards, attach proof of compliance
4.2.1 | Please provide unique ref number of the ISO 9001 ceriticate:
4.3 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
4.3.1 | Please provide unique ref number of the ISO 13485 ceriticate:
4.4 | Model on offer must comply with EU Medical Device Regulation (EU) 2017/745 or equvalent from SAHPRA recognized jurisdictions . Attach a copy of certification
4.4.1 | Please provide unique ref number of the ceriticate:
4.5 | Hazardous Substance Act: If this type of equipment / apparatus appears on the schedule of Hazardous substances issued by the Directorate: Health Technology of the Department of Health, a license in terms of the Act on Hazardous Substances (Act 15/1973) must be submitted with the bid / quotation documents. The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
4.6 | All electrical/electronic medical equipment must be licensed by Radiation Control, where applicable a copy of the license must be submitted The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
4.6.1 | Radiation Control License number:
4.7 | The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade (rubber to with 2 screws) type and it must be a minimum length of three (3) meters. NB the mains cable of the unit tendered for must be SABS colour coded.
4.8 | The equipment quoted for must be protected against electromagnetic interference IEC 60601-1-2
4.9 | The unit must comply with electrical safety and performance standard such as IEC 60601-2-12, attached certification
4.9.1 | Please provide unique ref number of the IEC ceriticate:
5 | COMPREHENSIVE MAINTENANCE
5.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) should be quoted for, as per the SCC document
5.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries, Flow sensors and oxygen cell), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
5.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
5.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries, Flow sensors and oxygen cell), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
5.3 | The bidder to provide the OEM recommended service interval of the unit:
5.4 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
6 | General
6.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
6.2 | The unit on offer must be the latest technology.
6.2.1 | Initial Manufacturing year of Model:
| VENTILATOR CPAP ACCESSORIES and CONSUMABLES | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
RT2-02-05a | 5 sets of silicone nasal prongs (each set must include 3 different neonatal sizes)
RT2-02-05b | 5 sets of CPAP bonnets (each set to include 3 different neonatal sizes)
RT2-02-05c | 5 neonatal circuits
RT2-02-05d | Temperature probes for humidifier
RT2-02-05e | Circuit holding arm

## Ventilator, Transport, Adult
RT2-02-06 | Transport Adult / Paediatric | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
1 | Description
1.1 | A portable medical device designed to provide mechanical ventilation during transport or in emergency situations.
1a | Product Offered:
1a.1 | Make:
1a.2 | Model:
1a.3 | Initial Manufacturing year of Model:
1a.4 | Radiation Control License number:
2 | Technical Specification
2.1 | High-end Transport ventilator, suitable for adult / paediatric patients, that integrates ventilation with comprehensive monitoring capability,with intergrated Air supply via compressor or scroll pump. Hot swap battery function. An optional colour display of ventilation parameters, graphics and trends must be available.
2.2 | High pressure medical oxygen connection must be interchangeable by the clinician for Oxygen with appropriate CGA gas fitting from wall supply to Oxygen Cylinder.
2.3 | Ventilator must have intergrated Air compressor/ scroll pump able to deliver +/- 120 lpm
2.4 | Ventilator shall have leak compensation, able to compensate for baseline leaks which may occur at the patient/mask interface or around artificial airways
2.5 | Ventilator shall have "Standby" mode, which suspends alarms and breath cycling, suspends trends.
2.6 | "Standby" mode shall retain current ventilator settings and alarms, resume normal ventilation with one touch
2.7 | CONFIGURATION AND SETUP
2.7.1 | User interface should allow for entering or selecting:
2.7.2 | Current patient or new patient
2.7.3 | Patient size range – Infant/Adult / Pediatric
2.8 | VENTILATOR MODE CONTROL and TYPES
2.8.1 | BREATH TYPES
2.8.1a | Volume
2.8.1b | Pressure
2.8.1c | Pressure support ventilation
2.8.1d | Spontaneous
2.8.2 | MODE TYPES
2.8.2.1 | Assist/control (A/C) – pressure and volume:
2.8.2.1a | Volume Controlled mode shall have option to disable demand flow and shall offer option of square or decelerating flow waveform
2.8.2.2 | Synchronous Intermittent Mandatory Ventilation (SIMV) – Pressure, Volume and PRVC
2.8.2.3 | Continuous Positive Airway Pressure / Pressure Support Ventilation (CPAP/PSV)
2.8.2.3a | Variable Rise Time in all modes.
2.8.2.3b | Variable breath Time Termination.
2.8.2.3c | Pressure Control: Variable Flow Termination.
2.8.2.4 | Non-invasive positive pressure ventilation (NPPV)
2.8.2.4a | Leak Compensation.
2.8.2.4b | Apnea backup should be active in A/C, SIMV, CPAP and Non Invasive.
2.8.2.4c | Apnea modes shall include user settable time duration before Apnea ventilation onset
2.9 | BREATH CONTROLS
2.9.1 | Breath rate: 1 – 80bpm
2.9.2 | Inspiratory flow trigger: Off, 1 – 9.0 L/min
2.9.3 | Manual breath – add single control breath
2.9.4 | Expiratory hold: max 6.0 sec
2.9.5 | Inspiratory hold: max 6.0 sec
2.9.6 | PC / PS / Spont. Flow 160 lpm
2.9.7 | Tidal Volume: 50 – 2000 ml
2.9.8 | Inspiratory pressure control level: 0 – 90cmH2O
2.9.9 | Inspiratory pressure Support level: 0 – 60cmH2O
2.9.10 | Inspiratory Time: 0.3 – 9.9 sec
2.9.11 | Calculated I:E ratio: 1:4 and 4:1
2.9.12 | PEEP: 0 – 20 cmH2O
2.10 | OXYGEN CONTROL
2.10.1 | % Oxygen: 21 – 100% ( 1% adjustments)
2.10.2 | Use of Low pressure Oxygen on/off ( Calculated Time left in Cylinder )
2.11 | PATIENT MONITORING (Volume, Pressure, Rate & Time)
2.11.1 | Ventilator shall have a variable orifice flow sensor for all patient ranges
2.11.2 | Monitored values to include Tidal and Minute Volumes.
2.11.3 | Tidal Volume (Vt) -Absolute range: 0 to 4 L (All patients) Normalized range: 0.0 to 30.0 ml / Kg (All patients)
2.11.4 | Exhaled Tidal Volume:
2.11.5 | Inspired Tidal Volume:
2.11.6 | Minute Volume - Absolute range:O to 99.9 L (All patients) Normalized Range: 0 to 999 (All patients)
2.11.7 | Patient Effort Indicator:
2.11.8 | Breath Rate -0 to 80 BPM
2.11.9 | Inspiratory Time (Tj)
2.11.10 | I:E Ratio (I:E) -1 :4 to 4:1
2.11.11 | Rapid Shallow Breathing Index (f / Vt) -0 to 500 b2 / min / L
2.11.12 | Peak Inspiratory Pressure (PPEAK) -0 to 120 cmH20
2.11.13 | Mean Airway Pressure (PMEAN) -0 to 99 cmH20
2.11.14 | PEEP -0 to 50 cmH20
2.11.15 | Oxygen Inlet Pressure (02 Inlet) 140 to 550 kpa (1.4-5.5 bar)
2.11.16 | Delivered O2 (Fi02): 21% to 100%
2.12 | ALARMS
2.12.1 | Alarms shall have distinctive audible tone
2.12.2 | Audible alarm volume
2.12.3 | Alarm Silence -120 sec. max
2.12.4 | Alarm Reset. Deactivates visual indicators for resolved alarm conditions
2.12.5 | Volume Alarms
2.12.5a | Low Exhaled Minute Volume (Low Ve) -Off (indicated by 0), 1 to 99 L
2.12.6 | Pressure Alarms
2.12.6a | Low PPEAK Off to 60 cmH20
2.12.6b | High PPEAK Alarm 5 to 100 cmH20
2.12.6c | High PEEP: Off, 3 to 40 cmH20
2.12.6d | Low PEEP Off -3 to 20 cmH20
2.12.6e | Maximum Circuit Pressure Limit -Maximum pressure at the y-connection: 125 cmH20
2.12.7 | Rate / Time Alarms
2.12.7a | Apnoea Interval. 10 to 60 seconds
2.12.7b | High Rate:
2.12.7c | Auto Apnoea Backup:
2.12.8 | Machine Alarms
2.12.8a | Safety Valve Open (SAFETY VALVE)
2.12.8b | Ventilator Inoperative (VENT INOP)
2.12.8c | Fan Failure
2.13 | Display
2.13.1 | Graphical display screen must be available
2.14 | MANEUVRES
2.14.1 | Spontaneous Breathing Trial (SBT)
2.14.2 | Ventilator shall return to normal ventilation at end of maneuver without any changes required to primary controls
2.15 | POWER SUPPLY
2.15.1 | Input voltage: 115 – 230 V and 11 –15 VDC
2.15.2 | Operating frequency: 50-60 Hz
2.15.3 | Power cord > 3 m fitted with standard 3 point plug
2.15.4 | Internal Battery back-up: Re-chargeable batteries with minimum duration of 1 hour.
2.15.5 | State duration of battery operation in minutes
2.15.6 | Maximum time to full charge 8 to 12 hours
2.16 | GAS SUPPLY
2.16.1 | Oxygen Pipeline Supply: Pressure Range 40 to 80 psig (2.8 to 5.5 bar)
2.16.2 | Low Pressure Source: <80 lpm <10 PSIG
2.17 | DATA INPUT I OUTPUT
2.17.1 | Digital Communication -RS-232 Ch1 (input for software updates/enhancements) RS-232 Ch2 (Undefined)
3 | STANDARD ACCESSORIES
3.1 | Adult breathing circuit
3.2 | External charger
3.3 | Carry bag and bed clamp
4 | Conformity Compliance (Please attach certificate)
4.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
4.2 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
4.2.1 | Please provide unique ref number of the ISO 13485 ceriticate:
4.3 | Model on offer must comply with EU Medical Device Regulation (EU) 2017/745 or equvalent from SAHPRA recognized jurisdictions . Attach a copy of certification
4.3.1 | Please provide unique ref number of the ceriticate:
4.4 | Hazardous Substance Act: If this type of equipment / apparatus appears on the schedule of Hazardous substances issued by the Directorate: Health Technology of the Department of Health, a license in terms of the Act on Hazardous Substances (Act 15/1973) must be submitted with the bid / quotation documents. The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
4.5 | All electrical/electronic medical equipment must be licensed by Radiation Control, where applicable a copy of the license must be submitted The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
4.5.1 | Radiation Control License number:
4.6 | The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade (rubber to with 2 screws) type and it must be a minimum length of three (3) meters. NB the mains cable of the unit tendered for must be SABS colour coded.
4.7 | The equipment quoted for must be protected against electromagnetic interference IEC 60601-1-2
4.8 | The unit must comply with electrical safety and performance standard such as IEC 60601-2-12, attached certification
4.8.1 | Please provide unique ref number of the IEC ceriticate:
5 | COMPREHENSIVE MAINTENANCE
5.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) should be quoted for, as per the SCC document
5.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries, Flow sensors and oxygen cell), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
5.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
5.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries, Flow sensors and oxygen cell), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
5.3 | The bidder to provide the OEM recommended service interval of the unit:
5.4 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
6 | General
6.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
6.2 | The unit on offer must be the latest technology.
6.2.1 | Initial Manufacturing year of Model:
| Transport Adult / Paediatric Accessories and Consumables | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
RT2-02-06a | Adult breathing circuit
RT2-02-06b | External charger
RT2-02-06c | Carry bag and
RT2-02-06d | Bed clamp

## Ventilator, Transport, Neonatal
RT2-02-07 | Ventilator ,Transport , Neonatal | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
1 | Description
1.1 | A portable , compact medical device designed for transporting and ventilating newborn infants.
1a | Product Offered:
1a.1 | Make:
1a.2 | Model:
1a.3 | Initial Manufacturing year of Model:
1a.4 | Radiation Control License number:
2 | Technical Specification
2.1 | A Basic Transport ventilator, suitable for Neonatal patients, that integrates ventilation with basic monitoring capability and a miniminum battery life of 4 hours.
2.2 | High pressure medical oxygen connection must be interchangeable by the clinician for Oxygen with appropriate CGA gas fitting from wall supply to Oxygen Cylinder.
3 | CONFIGURATION AND SETUP
3.1 | User interface should allow for entering or selecting:
3.2 | Oxygen /Air mixture 21% to 100%
3.3 | Mode of Ventilation
3.4 | Must be able to measure mainstream eCo2
4 | VENTILATOR MODE CONTROL and TYPES
4.1 | BREATH TYPES
4.1.1 | Volume
4.1.2 | Pressure
4.1.3 | Spontaneous /CPAP
4.2 | MODE TYPES
4.2.1 | Assist/control (A/C) – pressure and volume
4.2.2 | Continuous Positive Airway Pressure (CPAP)
4.2.3 | Variable breath Time Termination.
4.2.4 | Apnea Alarm should be active in A/C, SIMV, CPAP
4.3 | BREATH CONTROLS
4.3.1 | Breath rate: 1 - 80 bpm
4.3.2 | Tidal Volume: 5 ml
4.3.3 | Inspiratory pressure control level: 0 – 60cmH2O
4.3.4 | Inspiratory Time: 0.25 – 2 sec
4.3.5 | Calculated I:E ratio: 1:4 and 4:1
4.3.6 | PEEP: 0 – 20 cmH2O
4.4 | OXYGEN CONTROL
4.4.1 | % Oxygen: 21% – 100%
4.5 | PATIENT MONITORING
4.5.1 | Rate & Time Monitors
4.5.1.1 | Breath Rate -0 to 150 BPM
4.5.2 | Pressure Monitors
4.5.2.1 | Peak Inspiratory Pressure (PPEAK) -0 to 100 cmH20
4.5.2.2 | Mean Airway Pressure (PMEAN) -0 to 99 cmH20
4.5.2.3 | PEEP -0 to 50 cmH20
5 | ALARMS
5.1 | Alarms shall have distinctive audible tone
5.2 | Alarm Silence -120 sec. max
5.3 | Alarm Reset. Deactivates visual indicators for resolved alarm conditions
5.4 | Low Paw cmH2O
5.5 | High Paw cmH20
5.6 | Maximum Circuit Pressure Limit -Maximum pressure at the y-connection:60 cmH20
5.7 | Apnoea Interval
5.8 | Safety Valve Open (SAFETY VALVE)
5.9 | Ventilator Inoperative (VENT INOP)
5.10 | Fan Failure
5.11 | Loss of O2 – O2 Supply pressure and Duration Calculation:
6 | MANEUVERS
6.1 | Manual Breath
7 | POWER SUPPLY
7.1 | Input voltage: 115 – 230 V and 11 –15 VDC
7.2 | Operating frequency: 50-60 Hz
7.3 | Power cord > 3 m fitted with standard 3 point plug
7.4 | Internal Battery back-up: Re-chargeable batteries with minimum duration of 4 hour.
7.5 | State duration of battery operation in minutes
7.6 | Maximum time to full charge 4 to 6 hours
8 | GAS SUPPLY
8.1 | Oxygen Pipeline Supply: Pressure Range 280 to 550 kpa (2.8 to 5.5 bar)
8.2 | Low Pressure Source: <80 lpm <68 kpa
9 | ACCESSORIES
9.1 | Neonatal breathing circuits
9.2 | External charger
9.3 | Carry bag
9.4 | Mainstream etCo2 line
9.5 | Sidestream etCo2 line
9.6 | Bed clamp
10 | Conformity Compliance (Please attach certificate)
10.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
10.2 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
10.2.1 | Please provide unique ref number of the ISO 13485 ceriticate:
10.3 | Model on offer must comply with EU Medical Device Regulation (EU) 2017/745 or equvalent from SAHPRA recognized jurisdictions . Attach a copy of certification
10.3.1 | Please provide unique ref number of the ceriticate:
10.4 | Hazardous Substance Act: If this type of equipment / apparatus appears on the schedule of Hazardous substances issued by the Directorate: Health Technology of the Department of Health, a license in terms of the Act on Hazardous Substances (Act 15/1973) must be submitted with the bid / quotation documents. The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
10.5 | All electrical/electronic medical equipment must be licensed by Radiation Control, where applicable a copy of the license must be submitted The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
10.5.1 | Radiation Control License number:
10.6 | The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade (rubber to with 2 screws) type and it must be a minimum length of three (3) meters. NB the mains cable of the unit tendered for must be SABS colour coded.
10.7 | The equipment quoted for must be protected against electromagnetic interference IEC 60601-1-2
10.8 | The unit must comply with electrical safety and performance standard such as IEC 60601-2-12, attached certification
10.8.1 | Please provide unique ref number of the IEC ceriticate:
11 | COMPREHENSIVE MAINTENANCE
11.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) should be quoted for, as per the SCC document
10.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries,Flow sensors & Oxygen cell), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
10.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
10.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries, Flow sensors and oxygen cell), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
10.3 | The bidder to provide the OEM recommended service interval of the unit:
10.4 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
11 | General
11.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
11.2 | The unit on offer must be the latest technology.
11.2.1 | Initial Manufacturing year of Model:
| Ventilator ,Transport , Neonatal Accessories and Consumables | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
RT2-02-07a | Neonatal breathing circuit
RT2-02-07b | External charger
RT2-02-07c | Carry bag and
RT2-02-07d | Bed clamp
RT2-02-07e | Neonatal breathing circuits
RT2-02-07f | External charger
RT2-02-07g | Carry bag
RT2-02-07h | Mainstream etCo2 line
RT2-02-07i | Sidestream etCo2 line
RT2-02-07j | Bed clamp

## HF Oscillator
RT2-02-08 | HF Oscillator, Neonatal/Paediatric | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
Description
A medical device that provides respiratory support to both critically ill neonates and paediatric.
1a | Product Offered:
1a.1 | Make:
1a.2 | Model:
1a.3 | Initial Manufacturing year of Model:
1a.4 | Radiation Control License number:
1 | Technical Specification
1.1 | Configuration
1.1.1 | Positive pressure
1.1.2 | Invasive
1.1.3 | High frequency oscillator
1.1.4 | Patient population range: neonatal and pediatric
2 | Characteristics
2.1 | The unit must be a piston-driven oscillation with no need for centering.
2.1.1 | Bidder to provide evidence & demonstrate that their unit can support upto pediatric patients
2.2 | Dump valve opens at Paw > 50cm H2O or Paw < 20% of set Max Paw.
2.3 | Power requirements: 220 V AC, 50Hz
3 | Range of controls
3.1 | Mean airway pressure 3-45 cmH2O
3.2 | Oscillatory pressure 0-90 cmH2O
3.3 | % Inspiratory time 30% to 50%
3.4 | Frequency 3-15 Hz
3.5 | Bias flow 0-40l/min
3.6 | Max mean pressure 10-45 cmH2O
4 | Alarms
4.1 | Oscillator overheated.
4.2 | Battery low.
4.3 | Source gas low.
4.4 | Above “set max Paw”.
4.5 | Below “set min Paw”.
4.6 | Power failure.
5 | Monitored variables
5.1 | Mean airway pressure: +- 130 cmH2O; pressure amplitude: +- 130 cmH2O.
5.2 | Respiratory rate in Hz.
6 | ACCESSORIES
6.1 | Neonatal breathing circuits
6.2 | External charger
6.3 | Carry bag
6.4 | Mainstream etCo2 line
6.5 | Sidestream etCo2 line
6.6 | Bed clamp
7 | Conformity Compliance (Please attach certificate)
7.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 and 601-1-2 for medical equipment, attached certification
7.2 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
7.2.1 | Please provide unique ref number of the ISO 13485 ceriticate:
7.3 | Model on offer must comply with EU Medical Device Regulation (EU) 2017/745 or equvalent from SAHPRA recognized jurisdictions . Attach a copy of certification
7.3.1 | Please provide unique ref number of the ceriticate:
7.4 | Hazardous Substance Act: If this type of equipment / apparatus appears on the schedule of Hazardous substances issued by the Directorate: Health Technology of the Department of Health, a license in terms of the Act on Hazardous Substances (Act 15/1973) must be submitted with the bid / quotation documents. The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
7.5 | All electrical/electronic medical equipment must be licensed by Radiation Control, where applicable a copy of the license must be submitted The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
7.6 | The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade (rubber to with 2 screws) type and it must be a minimum length of three (3) meters. NB the mains cable of the unit tendered for must be SABS colour coded.
7.7 | The equipment quoted for must be protected against electromagnetic interference
8 | COMPREHENSIVE MAINTENANCE
8.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) should be quoted for, as per the SCC document
8.2 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries, Flow sensors and Oxygen sensor), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
8.3 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
8.4 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries, Flow sensors and Oxygen sensor), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
| HF Oscillator Accessories and Consumables | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
RT2-02-08a | Neonatal breathing circuits
RT2-02-08b | External charger
RT2-02-08c | Carry bag
RT2-02-08d | Mainstream etCo2 line
RT2-02-08e | Sidestream etCo2 line
RT2-02-08f | Bed clamp
RT2-02-08g | Disposable neonatal breathing circuit (Pk of 10)
RT2-02-08h | Disposable paeds breathing circuit (Pk of 10)
RT2-02-08i | Test lung for neonates
RT2-02-08j | Circuit Support Arm
RT2-02-08k | Disposable Flow sensors
RT2-02-08l | Reusable Flow sensor
RT2-02-08m | Oxygen sensor
RT2-02-08o | Exhalation diaphragm

## Resus, Gas-powered, Infant
RT2-02-09 | Resuscitator, infant, gas powered | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
Description
A portable medical device used to provide ventilation and resuscitation to newborn infants in distress
1a | Product Offered:
1a.1 | Make:
1a.2 | Model:
1a.3 | Initial Manufacturing year of Model:
1 | Technical Specification
1.1 | Must be able to provide 100% oxygen for infants up to 10 kg
1.2 | The gas which drives the resuscitator must be the oxygen
1.3 | Oxygen flow regulator: 5 - 15 l/min
1.4 | Adjustable Peak Inspiratory Pressure (PIP): 2 - 75 cm H2O
1.5 | Adjustable Positive End Expiratory Pressure (PEEP): 1 - 10 cm H2O
1.6 | Resuscitator must provide controlled pressure settings by means of a pressure limiting device
1.7
1.8 | Pressure gauge: -10 to 80 cm H2O
1.9 | Maximum pressure relief: 20 - 70cmH²O @ 8lpm.
1.10 | The unit must be handed over in full operating order
1.11 | A starter-pack of consumables must be supplied with the unit.
2 | Conformity Compliance (Please attach certificate)
2.1 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
2.1.1 | Please provide unique ref number of the ISO 13485 ceriticate:
2.2 | Model on offer must comply with EU Medical Device Regulation (EU) 2017/745 or equvalent from SAHPRA recognized jurisdictions . Attach a copy of certification
2.2.1 | Please provide unique ref number of the ceriticate:
3 | COMPREHENSIVE MAINTENANCE
3.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) should be quoted for, as per the SCC document
3.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
3.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
3.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
3.3 | The bidder to provide the OEM recommended service interval of the unit:
3.4 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
4 | General
4.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
4.2 | The unit on offer must be the latest technology.
4.2.1 | Initial Manufacturing year of Model:
| Resuscitator, infant, gas powered, Additional Options | Complies Yes/No
RT2-02-09a | Air/O2 blender (Upgrade)
RT2-02-09b | Suction availability (Upgrade)

## Humidifier, Servo control
RT2-02-10 | Heated Humidifier, servo control type | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
1 | DESCRIPTION
1.1 | A medical device that adds moisture and warmth to the air delivered to patients through a ventilator.
1a | Product Offered:
1a.1 | Make:
1a.2 | Model:
1a.3 | Initial Manufacturing year of Model:
2 | Technical Specification
2.1 | The Air Humidification System needs to provide optimal humidity to all intensive and high care patients that is suitable for neonates to adults invansive and non inavsive, including nasal prong O2
2.2 | The unit will be used to heat and humidify passing air, replacing the function of the vital nasal and throat passages in patients
2.3 | The units shall have silent operation
2.4 | Proximal and Distal temperature servo feedback loop must be able to maintain the distal setup temperature constant. While in operation the set temperature on the unit must be able to adjusted
2.5 | The unit shall have no moving parts that require maintenance.
2.6 | The unit must have a mounting capability
2.7 | Bidders to state maximum temperature deviation from set temperature
2.8 | It shall be easy to refill the fluid, i.e. Autofill chamber with constant compressible volume.
2.9 | The unit shall require less than 30 minutes to warm up
2.10 | The unit shall operate from a standard 220V 50Hz electrical supply
2.11 | The unit shall operate with a flow range: Invasive 3 - 60 L/min, Non Invasive 3 -40 L/min
3 | STANDARD ACCESSORIES to be added
3.1 | Mounting bracket
3.2 | Heater wire x 5
3.3 | Temperature probe x 5
3.4 | Humidifier chamber
4 | Conformity Compliance (Please attach certificate)
4.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
4.2 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
4.2.1 | Please provide unique ref number of the ISO 13485 ceriticate:
4.3 | Model on offer must comply with EU Medical Device Regulation (EU) 2017/745 or equvalent from SAHPRA recognized jurisdictions . Attach a copy of certification
4.3.1 | Please provide unique ref number of the ceriticate:
4.4 | The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade (rubber to with 2 screws) type and it must be a minimum length of three (3) meters. NB the mains cable of the unit tendered for must be SABS colour coded.
4.5 | The equipment quoted for must be protected against electromagnetic interference IEC 60601-1-2
5 | COMPREHENSIVE MAINTENANCE
5.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) should be quoted for, as per the SCC document
5.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS, Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
6 | General
6.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
6.2 | The unit on offer must be the latest technology.
6.2.1 | Initial Manufacturing year of Model:
| Heated Humidifier, servo control type Accessories and Consumables | Complies Yes/No
RT2-02-10a | Mounting bracket
RT2-02-10b | Heater wire
RT2-02-10c | Temperature probe
RT2-02-10d | Humidifier chamber

## Air-O2 Blender, Low Flow
RT2-02-11 | BLENDER - AIR/OXYGEN, LOW FLOW | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
1 | DESCRIPTION
1.1 | A medical device used to mix oxygen and air to deliver precise blend of gases to patients at low flow rates.
1a | Product Offered:
1a.1 | Make:
1a.2 | Model:
1a.3 | Initial Manufacturing year of Model:
1a.4 | Original Equipment Manufacturer:
2 | TECHNICAL SPECIFICATION
2.1 | The blender must be able to operate accurately on gas supply of between 2.5 and 5 bar.
2.2 | The Blender must be accurate up to 3%.
2.3 | The blender must have 2 intergrated flow meters:
2.4 | Flow meters with a range of 0 to 15 L.P.M flow and 0 to 1 or 3.5 L.P.M, with the latter calibrated in increments of at least 0.1 LPM.
2.5 | The O2 concentration control adjustment range must be between 21% - 100%.
2.6 | The alarm must be audible from 5 meters and must be high pitched. The alarm must be activated when there is a drop in wall O2 or Air pressure. The alarm must be demonstrated to an evaluation committee.
2.7 | The alarm bypass/ reset must operate when inlet pressure is corrected
2.8 | The pressure drop must be less than 45kpa at 3,5 bar inlet pressures and 40 L.P.M. flow rate.
2.9 | The weight must not exceed 3kg, please state the weight in kg
2.10 | The unit must have 2 outlet ports
2.11 | The blender must come with a bracket suitable to the end user facility e.g. rail Bracket/ wall mounted bracket.
2.12 | The blender must come complete with colour-coded air and oxygen hoses and “S.I.S.S.” type O2 gas outlet connector and “N.I.S.T” type air and O2 connectors and must be supplied with colour- coded air and oxygen hoses with Heyer- type connectors that are compatible with the wall O2 points at end user institution.
4 | STANDARD ACCESSORIES
4.1 | Mounting bracket (Rail or pole bracket suitable for the institution)
4.2 | At least 1.5m Colour-coded hoses complete with heyer connectors
5 | Conformity Compliance (Please attach certificate)
5.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
5.2 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
5.2.1 | Please provide unique ref number of the ISO 13485 ceriticate:
5.3 | Model on offer must comply with EU Medical Device Regulation (EU) 2017/745 or equvalent from SAHPRA recognized jurisdictions . Attach a copy of certification
5.3.1 | Please provide unique ref number of the ceriticate:
6 | COMPREHENSIVE MAINTENANCE
6.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) should be quoted for, as per the SCC document
6.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries, Flow sensors and Oxygen sensor), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
6.3 | The bidder to provide the OEM recommended service interval of the unit:
6.4 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
| BLENDER - AIR/OXYGEN, LOW FLOW ACCESSORIES | Complies Yes/No
RT2-02-11a | Mobile stand
RT2-02-11b | Mounting bracket
RT2-02-11c | Flow meter
RT2-02-11d | Control knob

## Air-O2 Blender, High Flow
RT2-02-12 | BLENDER - AIR/OXYGEN, HIGH FLOW | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
1 | DESCRIPTION
1.1 | A medical device used to mix Medical Air and hospital grade Oxygen (O2) into a gas source ranging from 21% (normal) to 100% (max) oxygen
1a | Product Offered:
1a.1 | Make:
1a.2 | Model:
1a.3 | Initial Manufacturing year of Model:
2 | TECHNICAL SPECIFICATION
2.1 | The blender must be able to operate accurately on gas supply of between 2.5 bars (29 psi) and 5 bars (72.5 psi).
2.2 | The blender must have a primary and an auxiliary outlet:
2.3 | Primary 15 – 20 L.P.M. flow range.
2.4 | Auxiliary 2 – 90 L.P.M. flow range with a bleed of no more than 10 – 12 L.P.M.
2.5 | The oxygen concentration control adjustment range must be between 21% and 100%.
2.6 | The maximum total flow at 60% on the control knob at 3,5 bar inlet pressure settings must be at least 120 L.P.M.
2.7 | Accuracy must be (with inlet gasses within 2 (two) bar and each gas pressure greater than 2 (two) bar but less than 5 (five) bar) + 10% of setting from 21% to 50% or + 5% oxygen which ever is smaller, accuracy maintained up to the audible alarm point.
2.8 | The alarm bypass activation must work when the inlet gas pressures differ between 1,37 bar (20 psi) and 2 bars (29 psi).
2.9 | The alarm must be audible from 5 meters. The alarm must be activated when there is a drop in wall O2 or Air pressure. It must be high pitched.
2.1 | The alarm / bypass reset must operate when inlet pressure differential is 0,5 bar (7.25 psi) or less.
2.11 | The pressure drop must be less than 45kpa at 3,5 bar (50.77 psi) inlet pressures and 40 L.P.M. flow rate.
2.12 | The weight must not exceed 3 kg. (please state the weight in Kg)
2.13 | The blender must come complete with colour-coded air and oxygen hoses and “S.I.S.S.” type O2 gas outlet connector and “N.I.S.T” type air and O2 connectors and must be supplied with colour- coded air and oxygen hoses with Heyer- type connectors that are compatible with the wall O2 points at end user institution.
2.14 | Mechanical Medical Oxygen / Air Blender
2.15 | Requirements :
2.16 | Deliver Oxygen Range 21 % to 100 %
2.17 | Flow Range 2 to 120 lpm.
2.18 | Maximum Bleed Flow range 0-12 lpm.
2.19 | 2 x Outlet ports
2.2 | The unit must alarm when the inlet pressures differ more than 1.37 bars (20 psi)
2.21 | 3m Colour Coded Oxygen and Air Hoses must be supplied.
2.22 | Appropriate Oxygen and Air Hose probes must be fitted.
3 | STANDARD ACCESSORIES
3.1 | Consumables used for day to day operations should be specified.
3.2 | Price and planned inflationary increase on consumable costs.
3.3 | Price of Optional Accessories ( Brackets or Flow Meters)
5 | Conformity Compliance (Please attach certificate)
5.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
5.2 | OEM must comply and certified on ISO 9001 quality standards, attach proof of compliance
5.2.1 | Please provide unique ref number of the ISO 9001 ceriticate:
5.3 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
5.3.1 | Please provide unique ref number of the ISO 13485 ceriticate:
5.4 | Model quoted for must be EC certified. Attach a copy of certification
5.4.1 | Please provide unique ref number of the EC ceriticate:
6 | COMPREHENSIVE MAINTENANCE
6.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) should be quoted for, as per the SCC document
6.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries, Flow sensors and Oxygen sensor), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
6.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
6.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS, labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
6.3 | The bidder to provide the OEM recommended service interval of the unit:
6.4 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
| BLENDER - AIR/OXYGEN, HIGH FLOW ACCESSORIES | Complies Yes/No
RT2-02-12a | Mobile stand
RT2-02-12b | Mounting bracket
RT2-02-12c | Flow meter
RT2-02-12d | Control knob

## Pulmonary Function Unit
RT2-02-13 | PULMONARY FUNCTION UNIT | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
1 | DESCRIPTION
1.1 | A medical device used to measure and assess lung function on patients.
1a | Product Offered:
1a.1 | Make:
1a.2 | Model:
1a.3 | Initial Manufacturing year of Model:
2 | SPECIFICATIONS
2.1 | The unit offered must be a portable pulmonary function unit, which must provide measuring as follows:
2.1.1 | Slow vital capacity (Vt; ERV; Vcin; Vcex and others. Forced spirometry (inspiratory and expiratory flow-volume curves). Maximal voluntary ventilation (MVV).
3 | The structure of the portable unit being offered must be as follows:
3.1 | The unit must be portable, compact and lightweight for easy transport.
3.2 | Must be easy to clean, disassemble and sterilize all parts in direct contact with patient.
3.3 | The display of the unit must be backlit LCD type, minimum resolution of 320 x 200 pixels.
3.4 | Must have an up to date communication and printer port. Bidders to state type output device
3.5 | The unit must be AC / DC with built-in battery backup, sufficient for a minimum of two (2) hours continuous use.
4 | The signal processing on the unit offered must be as follows:
4.1 | Highly linear, open circuit, differential pressure pneumotachograph, or a bi-directional digital turbine flowmeter.
4.2 | Flow: Range: 0-18 L/sec (bi-directional).
4.3 | Accuracy: + 3% over the full scale.
4.4 | Resolution: 10ml/sec.
4.5 | Volume: Range: 0-10 litre.
4.6 | Accuracy: 3% or 50ml, whichever is greater, over the full range.
4.7 | Resolution: 1ml.
4.8 | Resistance: The resistance of the pneumotachograph and tubing should be less than 1.5cm H2O/L/sec at 12L/sec.
4.9 | Time: The unit should be capable of reading time for at least 15 seconds with an accuracy of 1%.
5 | The data management on the unit offered must be as follows:
5.1 | Calculation of user selectable universal predicted standards. Graphical and numerical display during measurement and in measurement reports is essential.
5.2 | Very fast access to all patient data, including previously performed “pre” data for simultaneous pre and post testing.
5.3 | Backup of at least 10 sets of patient data, together with pre and post graphics in internal database.
5.4 | Software must facilitate adjustments to measurement routines, selectable protocols with a choice of parameters, predicted values and graphics in a pre / post comparison.
5.5 | Saving of measured values and curves in the database.
5.6 | At least three manoevures each for pre and post bronchodilator studies must be obtainable.
5.7 | Automatic saving of the best curve is essential, but must be user selectable as well.
6 | CALIBRATION:
6.1 | Computer assisted calibration of pneumotachograph and automatic Body Temperature and Pressure Saturated (BTPS) correction by means of a 2 or 3 litre calibration syringe.
7 | The parameters offered on the quoted unit must be as follows:
7.1 | Slow spirometry: Vt; BF; MV; ERV; ERV; Vcin; Vcex and others. Flow-volume curve: FVC; FEV1; FEV0,5; FEF2; FEV1% / Fvmax; PEF; MEF95; MEF50; MEF25; MMEF; REF75-85; PIF; MIF50; MEF50% / MIF50 and others. Maximal voluntary ventilation: MVV; RR; Vt and others.
8 | Reports must be provided and available on the unit offered as follows:
8.1 | Selection of defined protocols must be available with individual selection of parameters, predicted values and graphics in a pre / post comparison with output of measuring reports on printer or display. It must be possible to overlay pre and post curves in the printout and on screen.
8.2 | Re-reading of test results for output together with possibility to edit reports.
8.3 | Built-in, high resolution thermal printer, at least 15cm wide.
9 | STANDARD ACCESSORIES
9.1 | spirometry sensor
9.2 | calibration syringe
9.3 | Weather station
9.4 | mouth pieces
9.5 | paper
10 | Conformity Compliance (Please attach certificate)
10.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
10.2 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
10.2.1 | Please provide unique ref number of the ISO 13485 ceriticate:
10.3 | Model on offer must comply with EU Medical Device Regulation (EU) 2017/745 or equvalent from SAHPRA recognized jurisdictions . Attach a copy of certification
10.3.1 | Please provide unique ref number of the ceriticate:
10.4 | Hazardous Substance Act: If this type of equipment / apparatus appears on the schedule of Hazardous substances issued by the Directorate: Health Technology of the Department of Health, a license in terms of the Act on Hazardous Substances (Act 15/1973) must be submitted with the bid / quotation documents. The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
10.5 | All electrical/electronic medical equipment must be licensed by Radiation Control, where applicable a copy of the license must be submitted The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
10.5.1 | Radiation Control License number:
10.6 | The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade (rubber to with 2 screws) type and it must be a minimum length of three (3) meters. NB the mains cable of the unit tendered for must be SABS colour coded.
10.7 | The equipment quoted for must be protected against electromagnetic interference IEC 60601-1-2
11 | COMPREHENSIVE MAINTENANCE
11.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) should be quoted for, as per the SCC document
11.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS, Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
11.2 | The bidder to provide the OEM recommended service interval of the unit:
11.3 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
12 | General
12.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
12.2 | The unit on offer must be the latest technology.
12.2.1 | Initial Manufacturing year of Model:
| PULMONARY FUNCTION UNIT ACCESSORIES | Complies Yes/No
RT2-02-13a | spirometry sensor
RT2-02-13b | calibration syringe
RT2-02-13c | Weather station
RT2-02-13d | mouth pieces
RT2-02-13e | paper

## O2 Flowmeter, Single
RT2-02-14 | Oxygen Flowmeter, single | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
1 | Description
1.1 | A medical device that measures and controls the flow rate of oxygen delivered a patient with a single output.
1 | Product Offered:
1a.1 | Make:
1a.2 | Model:
1a.3 | Initial Manufacturing year of Model:
2 | Technical Specification
2.1 | Flowmeter must be of the pressure compensated type and fit directly into the Oxygen outlet.
2.2 | The flowmeter must have a stainless steel float in a calibrated flow tube
2.3 | Must have a flow accuracy within 10%.
2.4 | The range of 0 – 15 l/m of the flowmeter must be permanent and easy to read.
2.5 | Connector must be Institution specific, 1 flow meter with 1 outlet
2.6 | The supplier must confirm the connection before delivery.
2.7 | Must have a chrome-plated brass body and adjustment knob
2.8 | The flowmeter must operate accurately on a 4 bar pipeline source.
2.9 | The flowmeter hood must be made of crack–resistant polycarbonate.
2.10 | Each flowmeter must be supplied with oxygen outlet ‘tail’ fitting.
2.11 | The flow meter must be easy to service and maintain
2.12 | The unit must be handed over in full operating order
STANDARD ACCESSORIES
Tail piece (Xmas tree-plastic) Pk of 100
3 | Conformity Compliance (Please attach certificate)
3.1 | International OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
3.1.1 | Please provide unique ref number of the ISO 13485 ceriticate:
3.2 | Model on offer must comply with EU Medical Device Regulation (EU) 2017/745 or equvalent from SAHPRA recognized jurisdictions . Attach a copy of certification
3.2.1 | Please provide unique ref number of the ceriticate:
4 | COMPREHENSIVE MAINTENANCE
4.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) should be quoted for, as per the SCC document
4.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries, Flow sensors and Oxygen sensor), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
4.2 | The bidder to provide the OEM recommended service interval of the unit:
4.3 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
| Oxygen Flowmeter, single - Accessories and Consumables | Complies Yes/No
RT2-02-14a | Outer/flow tube (Pk of 10)
RT2-02-14b | Inner tube (Pk of 10)
RT2-02-14c | Flow ball (pk of 10)
RT2-02-14d | Tail piece (Xmas tree-plastic) Pk of 100
RT2-02-14e | Flow Knob (Pk of 10)

## O2 Flowmeter, Double
RT2-02-15 | Oxygen Flowmeter, double | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
1 | Description
1.1 | A medical device that measures and controls the flow rate of oxygen delivered a patient with double output.
1 | Product Offered:
1a.1 | Make:
1a.2 | Model:
1a.3 | Initial Manufacturing year of Model:
2 | Technical Specification
2.1 | Flowmeter must be of the pressure compensated type and fit directly into the Oxygen outlet.
2.2 | The flowmeter must have a stainless steel float in a calibrated flow tube
2.3 | Must have a flow accuracy within 10%.
2.4 | The range of 0 – 15 l/m of the flowmeter must be permanent and easy to read.
2.5 | Connector must be Institution specific, 1 flow meter with 1 outlet
2.6 | The supplier must confirm the connection before delivery.
2.7 | Must have a chrome-plated brass body and adjustment knob
2.8 | The flowmeter must operate accurately on a 4 bar pipeline source.
2.9 | The flowmeter hood must be made of crack–resistant polycarbonate.
2.10 | Each flowmeter must be supplied with oxygen outlet ‘tail’ fitting.
2.11 | The flow meter must be easy to service and maintain
2.12 | The unit must be handed over in full operating order
STANDARD ACCESSORIES
Tail piece (Xmas tree-plastic) Pk of 100
3 | Conformity Compliance (Please attach certificate)
3.1 | International OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
3.1.1 | Please provide unique ref number of the ISO 13485 ceriticate:
3.2 | Model on offer must comply with EU Medical Device Regulation (EU) 2017/745 or equvalent from SAHPRA recognized jurisdictions . Attach a copy of certification
3.2.1 | Please provide unique ref number of the ceriticate:
4 | COMPREHENSIVE MAINTENANCE
4.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) should be quoted for, as per the SCC document
4.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries, Flow sensors and Oxygen sensor), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
4.2 | The bidder to provide the OEM recommended service interval of the unit:
4.3 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
| Oxygen Flowmeter, Double- Accessories and Consumables | Complies Yes/No
RT2-02-15a | Outer/flow tube (Pk of 10)
RT2-02-15b | Inner tube (Pk of 10)
RT2-02-15c | Flow ball (pk of 10)
RT2-02-15d | Tail piece (Xmas tree-plastic) Pk of 100
RT2-02-15e | Flow Knob (Pk of 10)

## Nebulizer, Ultrasonic
RT2-02-16 | NEBULIZER, ULTRASONIC | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
1 | DESCRIPTION
1.1 | a humidifier using high-frequency electricity to power a transducer that vibrates and divides liquid into particles in its nebulizing chamber; used in inhalation therapy.
1a | Product Offered:
1a.1 | Make:
1a.2 | Model:
1a.3 | Initial Manufacturing year of Model:
1a.4 | Radiation Control License number:
2 | TECHNICAL SPECIFICATION
2.1 | This unit must be the ultrasonic oscillator type.
2.2 | The unit must operate off 220V-240V 50hz A.C. single phase power supply.
2.3 | Most of the droplets must be within the 1 to 6 micron range.
2.4 | Output volumes must be a minimum of 2,5ml of aerosol per minute at the maximum output setting.
2.5 | The aerosol flow must be adjustable.
2.6 | The unit must have a self-contained air supply delivering of upto 20l/min.
2.7 | The nebulizer must be manufactured from rust free or plastic material.
2.8 | The unit must have a feed system in the reservoir for continuous operation.
2.9 | The nebulizer must not resist any patient coughing.
2.10 | The operating frequency must be between 1,6 and 1,8MHz.
2.11 | The nebulizer must not weigh more than 5Kg.
2.12 | Circuit hoses and chamber must come complete with the nebulizer.
3 | STANDARD ACCESSORIES
3.1 | Adult and Paeds Mask
3.2 | Mouth piece
3.3 | Hose
3.4 | Medication cups (pk of 5)
4 | Conformity Compliance (Please attach certificate)
4.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
4.2 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
4.2.1 | Please provide unique ref number of the ISO 13485 ceriticate:
4.3 | Model on offer must comply with EU Medical Device Regulation (EU) 2017/745 or equvalent from SAHPRA recognized jurisdictions . Attach a copy of certification
4.3.1 | Please provide unique ref number of the ceriticate:
4.4 | Hazardous Substance Act: If this type of equipment / apparatus appears on the schedule of Hazardous substances issued by the Directorate: Health Technology of the Department of Health, a license in terms of the Act on Hazardous Substances (Act 15/1973) must be submitted with the bid / quotation documents. The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
4.5 | All electrical/electronic medical equipment must be licensed by Radiation Control, where applicable a copy of the license must be submitted The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
4.5.1 | Radiation Control License number:
4.6 | The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade (rubber to with 2 screws) type and it must be a minimum length of three (3) meters. NB the mains cable of the unit tendered for must be SABS colour coded.
4.7 | The equipment quoted for must be protected against electromagnetic interference IEC 60601-1-2
5 | COMPREHENSIVE MAINTENANCE
5.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) should be quoted for, as per the SCC document
5.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS, Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
5.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
5.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS, labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
5.3 | The bidder to provide the OEM recommended service interval of the unit:
5.4 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
6 | General
6.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
6.2 | The unit on offer must be the latest technology.
6.2.1 | Initial Manufacturing year of Model:
| NEBULIZER -ACCESSORIES and CONSUMABLES | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
RT2-02-16a | Adult and Paeds Mask
RT2-02-16b | Mouth piece
RT2-02-16c | Hose
RT2-02-16d | Medication cups (pk of 5)

## EIT, Lung
RT2-02-17 | Electrical Impedance Tomography Unit, Lung Monitoring | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
1 | DESCRIPTION
1.1 | A medical device that uses electrical impulses to create images of the internal structures of the body, specifically the lungs.
2 | Product Offered:
1a.1 | Make:
1a.2 | Model:
1a.3 | Initial Manufacturing year of Model:
2 | SPECIFICATIONS
2.1 | The unit offered must be able to function well without the consideration of the Brand/Make of ventilator used to ventilate the patient
2.2 | The unit must be presented and supplied mounted on a customised cart/trolley
2.3 | The unit offered must be able to be used on the following patients:
2.3.1 | Neonates
2.3.2 | Paediatrics
2.3.3 | Adults
2.4 | The following paramaters must be monitored:
2.4.1 | Dynamic range
2.4.2 | Ventilation maps
2.4.3 | Compliance maps
2.4.4 | Plethysmogram
2.5 | The unit must have touchscreen display size of at least 30cm
2.6 | The frame rate shall not be less than 50 frames per second. Indicate the frame rate that the unit offers
2.7 | The unit must offered with both disposable and re-usable electrode belts options for paediatrics and adults.
2.8 | The unit must offered with only disposable electrode belts for neonates.
2.9 | The unit must be capable of providing measurements for 32 electrodes
2.10 | The unit must operate from both the 220 V, 50 Hz mains supply and internal rechargeable batteries
2.11 | The batteries must have the capacity for full operation for at least 2 hours when fully charged, under continuous operation
3 | Standard Accessories and Conusmables
3.1 | Trunk cable
3.2 | Electrode belt, disposable, small
3.3 | Electrode belt, disposable, medium
3.4 | Electrode belt, disposable, large
3.5 | Patient cable, small
3.6 | Patient cable, medium
3.7 | Patient cable, large
3.8 | Electrode belt, re-usable, medium
3.9 | Electrode belt, re-usable, large
3.10 | ECG electrodes (Pk of 100)
4 | Conformity Compliance (Please attach certificate)
4.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
4.2 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
4.2.1 | Please provide unique ref number of the ISO 13485 ceriticate:
4.3 | Model on offer must comply with EU Medical Device Regulation (EU) 2017/745 or equvalent from SAHPRA recognized jurisdictions . Attach a copy of certification
4.3.1 | Please provide unique ref number of the ceriticate:
4.4 | Hazardous Substance Act: If this type of equipment / apparatus appears on the schedule of Hazardous substances issued by the Directorate: Health Technology of the Department of Health, a license in terms of the Act on Hazardous Substances (Act 15/1973) must be submitted with the bid / quotation documents. The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
4.5 | All electrical/electronic medical equipment must be licensed by Radiation Control, where applicable a copy of the license must be submitted The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
4.5.1 | Radiation Control License number:
5 | COMPREHENSIVE MAINTENANCE
5.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) must be included in the bid price
5.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
6 | General
6.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
6.2 | The unit on offer must be the latest technology.
| Electrical Impedance Tomography Unit -Accessories and Consumables | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point
RT2-02-17a | Battery
RT2-02-17b | Trunk cable
RT2-02-17c | Electrode belt, disposable, small
RT2-02-17d | Electrode belt, disposable, medium
RT2-02-17e | Electrode belt, disposable, large
RT2-02-17f | Patient cable, small
RT2-02-17g | Patient cable, medium
RT2-02-17h | Patient cable, large
RT2-02-17i | Electrode belt, re-usable, medium
RT2-02-17j | Electrode belt, re-usable, large
RT2-02-17k | Flow sensor neonates
RT2-02-17l | Flow sensor Adults
RT2-02-17m | Disposable Electrode belt cover, Size S (box of 10)
RT2-02-17n | Disposable Electrode belt cover, Size M (box of 10)
RT2-02-17o | Disposable Electrode belt cover, Size LL (box of 10)

## Monitor, NIV
RT2-02-18 | Non-Invasive Ventilation Monitor | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
1 | Description
1.1 | A medical device that tracks and displays the vital signs and ventilation parameters in patients receiving non-invasive ventilation therapy.
1a | Product Offered:
1a.1 | Make:
1a.2 | Model:
1a.3 | Initial Manufacturing year of Model:
1a.4 | Radiation Control License number:
2 | Technical Specification
2.1 | Preconfigured, state of the art technology, suitable for Paeds to adult
2.2 | Microprocessor controlled with automatic self-test and calibration
2.3 | The following parameters must be measured and displayed :
2.3.1 | Minute Volume
2.3.2 | Tidal Volume
2.3.3 | Respiratory Rate
2.4 | The monitor must operate off both 220 V, 50 Hz single phase AC power supply and an internal rechargeable battery.
2.5 | If mains power fails the monitor must operate off battery power for at least 2 hours.
2.6 | Visible indication to show whether the monitor is operating off the mains or battery power.
2.7 | All control keys must be splashproof and resistant to the entry of fluids into the unit.
2.8 | The front panel must be easy to clean and disinfect.
2.9 | The monitor must have a carry handle.
2.10 | High resolution LCD colour display, at least 20cm diagonal
2.10.1 | Digits: large: 10 mm, small: 5 mm
2.11 | Trending for 24 hours must be possible on the monitor
2.12 | Defibrillator-proof up to 360 Joules
2.13 | Protected against HF interference
2.14 | Patient isolated from all potentially harmful circuits
2.15 | Must comply with IEC 60601-1
2.16 | Minute Volume
2.16.1 | Range: 0.2 to 50 litres per minute
2.16.2 | Accuracy: + 10%
2.17 | Tidal Volume
2.17.1 | Range: 40 to 2500ml
2.17.2 | Accuracy: + 10%
2.18 | Respiratory Rate
2.18.1 | Range: 4 to 38 Beats per minute
2.18.2 | Accuracy: + 2 b/pm
2.19 | ALARMS
2.19.1 | Visible and audio alarms
2.19.2 | The monitor must automatically set default alarms when switched on.
2.19.3 | The alarms must be programmable
2.19.4 | The following conditions must trigger alarm:
2.19.4.1 | Low Minute Volume (MV)
2.19.4.2 | High Minute Volume (MV)
2.19.4.3 | Low Tidal Volume (TV)
2.19.4.4 | High Tidal Volume (TV)
2.19.4.5 | Low Respitatory Rate (RR)
2.19.4.6 | High Respitatory Rate (RR)
2.19.4.7 | Low Impedance
2.19.4.8 | High Impedance
2.19.4.9 | Low battery alarm
2.19.4.10 | Loose pads
2.20 | Violation of limits must result in both audible and visible alarms and indicate which parameter is affected.
2.21 | The sound level of the audible alarm must be adjustable.
2.22 | It must be possible to temporarily silence the audible alarm but not remove the visible alarm.
2.23 | Duration of the silenced period: 2 minutes.
2.24 | Must operate from 220V 50 Hz mains power
2.25 | Mains cable at least four metres long with SABS approved 15A three prong plug.
2.26 | All cables and sockets must be colour coded, labelled or otherwise prevented from being incorrectly coupled
2.27 | The unit must be handed over in full operating order
2.28 | A starter-pack of consumables must be supplied with the unit.
3 | Standard Accessories & Consumables
3.1 | Roll stand/trolley
3.2 | Patient cable
3.3 | Disposable paeds sensors/probes (Pk of 20)
3.4 | Disposable Adult sensors/probes (Pk of 20)
3.5 | Optional Items, quote prices incl VAT
3.5.1 | Wall bracket, height adjustable on which the monitor can be mounted.
4 | Conformity Compliance (Please attach certificate)
4.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
4.2 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
4.2.1 | Please provide unique ref number of the ISO 13485 ceriticate:
4.3 | Model on offer must comply with EU Medical Device Regulation (EU) 2017/745 or equvalent from SAHPRA recognized jurisdictions . Attach a copy of certification
4.3.1 | Please provide unique ref number of the ceriticate:
4.4 | Hazardous Substance Act: If this type of equipment / apparatus appears on the schedule of Hazardous substances issued by the Directorate: Health Technology of the Department of Health, a license in terms of the Act on Hazardous Substances (Act 15/1973) must be submitted with the bid / quotation documents. The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
4.5 | All electrical/electronic medical equipment must be licensed by Radiation Control, where applicable a copy of the license must be submitted The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
4.5.1 | Radiation Control License number:
5 | COMPREHENSIVE MAINTENANCE
5.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) must be included in the bid price
5.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
6 | General
6.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
6.2 | The unit on offer must be the latest technology.
| Non-Invasive Ventilation Monitor Accessories & Consumables | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
RT2-02-18a | Roll stand/trolley
RT2-02-18b | Patient cable
RT2-02-18c | Disposable paeds sensors/probes (Pk of 20)
RT2-02-18d | Disposable Adult sensors/probes (Pk of 20)
RT2-02-18e | Wall bracket, height adjustable on which the monitor can be mounted.

## Spirometer, portable
RT2-02-19 | SPIROMETER SYSTEM, portable | Complies Yes/No | Provide your answers in this Column. You are advised to be transparent with deviation.
1 | DESCRIPTION
1.1 | A compact, lightweight medical device that measures lung function and capacity.
2 | Product Offered:
2.1 | Make:
2.2 | Model:
2.3 | Initial Manufacturing year of Model:
3 | SPECIFICATIONS (10% tolerance)
3.1 | A portable spirometer unit for the diagnosis of asthma, chronic obstructive pulmonary disease and other conditions that affect breathing.
3.2 | Volume: +/-3% or 50ml, whichever is greater for FVC and FEV1
3.3 | Flow:
3.3.1 | FEF +/-5% or +/-200ml/sec, whichever is greater
3.3.2 | PEF +/-10% or 400ml/sec, whichever greater
3.4 | Flow range: -12 to 16l/sec
3.5 | Parameters Measured:
3.5.1 | FVC,
3.5.2 | FEV.5,
3.5.3 | FEV1,
3.5.4 | FEV3,
3.5.5 | FEV6,
3.5.6 | FEV1/FVC,
3.5.7 | FEV1/FEV6,
3.5.8 | FEV.5/FVC,
3.5.9 | FEV3/FVC,
3.5.10 | PEF,
3.5.11 | FEF25,
3.5.12 | FEF50,
3.5.13 | FEF75,
3.5.14 | FEF25-75,
3.5.15 | FEF 75-85,
3.5.16 | FET,
3.5.17 | FEV1/FEV5,
3.5.18 | FEV1/PEF,
3.5.19 | MIF50,
3.5.20 | FIVC,
3.5.21 | FIV1,
3.5.22 | FIV1/FVC,
3.5.23 | FEV1/FIV1,
3.5.24 | PIF,
3.5.25 | MTT,
3.5.26 | PEV/PIF
3.5.27 | Lung Age
3.5.28 | VC,
3.5.29 | TV,
3.5.30 | ERV,
3.5.31 | IRV,
3.5.32 | IC,
3.5.33 | Ti,
3.5.34 | Te,
3.5.35 | Tt,
3.5.36 | Ti/Tt
3.6 | Hardware Specifications
3.6.1 | Dimension: 55mm x 160mm x 25mm (± 10% Tolerance)
3.6.2 | Weight: 180 g (± 10% Tolerance)
3.6.3 | Display: Alphanumeric LCD touchscreen
3.6.4 | Storage capacity: at least 1000 measurements
3.6.5 | Power Supply: Rechargeable Lithium-Ion 3.7V batteries
3.6.6 | The unit must be able to communicate via USB
4 | Graph options should be flow-volume and volume-time format
4.1 | Mouthpieces x4
4.2 | Battery pack x2
4.3 | Carry bag
4.4 | AC adapter
4.5 | USB Cable
5 | Conformity Compliance (Please attach certificate)
5.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 60601 for medical equipment, attached certification
5.2 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
5.2.1 | Please provide unique ref number of the ISO 13485 ceriticate:
5.3 | Model on offer must comply with EU Medical Device Regulation (EU) 2017/745 or equvalent from SAHPRA recognized jurisdictions . Attach a copy of certification
5.3.1 | Please provide unique ref number of the ceriticate:
5.4 | Hazardous Substance Act: If this type of equipment / apparatus appears on the schedule of Hazardous substances issued by the Directorate: Health Technology of the Department of Health, a license in terms of the Act on Hazardous Substances (Act 15/1973) must be submitted with the bid / quotation documents. The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.5 | All electrical/electronic medical equipment must be licensed by Radiation Control, where applicable a copy of the license must be submitted The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.5.1 | Radiation Control License number:
6 | COMPREHENSIVE MAINTENANCE
6.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) must be included in the bid price
6.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
7 | General
7.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
7.2 | The unit on offer must be the latest technology.