## Profile
STERILIZATION & DISINFECTION EQUIPMENT (Sterilizers & Instrument Washers)
# | Item number | Item Description
1 | RT2-10-01 | Autoclave, Tabletop, 24-30L
2 | RT2-10-02 | Autoclave, Tabletop, 32-38L
3 | RT2-10-03 | Autoclave, Tabletop, 40-50L
4 | RT2-10-04 | Ultrasonic Washer (30L)
5 | RT2-10-05 | Heat Sealer
6 | RT2-10-06 | Plasma Sterilizer 35-50L
7 | RT2-10-07 | Plasma Sterilizer 85-110L
8 | RT2-10-08 | Plasma Sterilizer 145-160L
9 | RT2-10-09 | Endoscope Washer
Cold Chain & Warm Storage Equipment
# | Item number | Item Description
1 | RT2-11-01 | Fridge, Vaccine (Chest), 80 - 90L
2 | RT2-11-02 | Fridge, Vaccine (Chest), 100 - 150L
3 | RT2-11-03 | Fridge, Vaccine (Chest), 240 - 260L
4 | RT2-11-04 | Fridge, Vaccine (Upright), 100 - 150L
5 | RT2-11-05 | Fridge, Vaccine (Upright), 200 - 250L
6 | RT2-11-06 | Fridge, Vaccine (Upright), 300 - 350L
7 | RT2-11-07 | Fridge, Vaccine (Solar-Powered), 120 - 150L
8 | RT2-11-08 | Fridge, Vaccine (Solar-Powered), 200 - 250L
9 | RT2-11-09 | Fridge, Pharmacy, 900 - 1100L
10 | RT2-11-10 | Fluid Warmer, Cabinet, 40-60L
11 | RT2-11-11 | Fluid Warmer, Cabinet, 80-150L
12 | RT2-11-12 | Fluid Warmer, Cabinet, 200-300L
INSTALLATION ALTERATIONS REQUIRED
# | Item number | Item Description
1 | RT2-13-01 | Instrument Washer, Under-counter
2 | RT2-13-02 | Instrument Washer, Single door
3 | RT2-13-03 | Instrument Washer, Double door
4 | RT2-13-04 | Theater Light, Ceiling-Mount
5 | RT2-13-05 | Dialysis Machine (CRRT)
6 | RT2-13-06 | Intermittent Hemodialysis
7 | RT2-13-07 | Mobile RO system
8 | RT2-13-08 | Complete RO plant
9 | RT2-13-09 | Nurse Call System, 8 Bed
10 | RT2-13-10 | Nurse Call System, 16 Bed
11 | RT2-13-11 | Nurse Call System, 24 Bed

## Category 10 - Sterilization Equ
STERILIZATION & DISINFECTION EQUIPMENT (Sterilizers & Instrument Washers)
# | Item number | Item Description
1 | RT2-10-01 | Autoclave, Tabletop, 24-30L
2 | RT2-10-02 | Autoclave, Tabletop, 32-38L
3 | RT2-10-03 | Autoclave, Tabletop, 40-50L
4 | RT2-10-04 | Ultrasonic Washer (30L)
5 | RT2-10-05 | Heat Sealer
6 | RT2-10-06 | Plasma Sterilizer 35-50L
7 | RT2-10-07 | Plasma Sterilizer 85-110L
8 | RT2-10-08 | Plasma Sterilizer 145-160L
9 | RT2-10-09 | Endoscope Washer

## Autoclave, Tabletop, 24-30L
RT2-10-01 | AUTOCLAVE - TABLETOP 24-30 liters | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
1 | DESCRIPTION
1.1 | A strong sealed chamber in which surgical instruments, towels, dressings, and other items can be sterilized by steam under raised pressure
2 | Product Offered:
2.1 | Make:
2.2 | Model:
2.3 | Initial Manufacturing year of Model:
3 | TECHNICAL SPECIFICATION
3.1 | The autoclave must be fully automatic with a computerized control system, which monitors all the physical parameters and must be equipped with a built-in printer to provide a traceable reference documentation of all the sterilization parameters.
3.2 | The chamber volume must be at least 24 to 30 Litre capacity.
3.3 | The autoclave must be fitted with a high vacuum silent running vacuum pump. To ensure perfect pre and post vacuum conditions. (SPLIT vacuum).
3.4 | CLEAN WATER SUPPLY: An external water supply or an internal RESERVOIR of sufficient capacity will be acceptable.Please state capacity of RESERVOIR.
3.5 | WASTE WATER DRAINS: An external waste water drain or an internal RESERVOIR of sufficient capacity will be acceptable.Please state capacity of RESERVOIR.
3.6 | The autoclave offered must be capable of performing the following sterilization cycles: Six sterilization cycles i.e. C1-wrapped,C2-wrapped,C3-unwrapped,C4-unwrapped, C5-Prion cycle (Creutzfeldt-Jakob disease), C6-Flash Cycle.
3.7 | The autoclave offered must be capable of performing the following 3 test cycles:
3.7.1 | 1)Vacuum Test: Test to check if the pressure is correct.
3.7.2 | 2) Bowie & Dick Test: Test to check if the steam correctly penetrates a porous load.
3.7.3 | 3) Helix Test: Test to check if the steam correctly penetrates a hollow load.
3.7.4 | Sterilizing temperature: C1-134⁰ , C2-121⁰, C3-134⁰, C4-121⁰, C5-134⁰, C6-134⁰
3.7.5 | Programmable cycles : Wrapped & Unwrapped automatic.
3.7.6 | Temperature range : 100- 134⁰ with a heat shutoff function.
3.7.7 | Controls should be sealed against fluid penetration.
3.7.8 | Recorder/Printer
3.7.9 | Bidder to state sterilazation time and drying time offered.
3.7.10 | The unit should have chamber temperature monitoring, accurate to ± 1⁰C over the operating range.
3.7.11 | A suitable circuit interrupter to match the load must be fitted.
3.7.12 | Power consumption not to exceed 3,5kW (16 Amps).
3.7.13 | Supply voltage 230 V.A.C. 50Hz.
3.8 | Please specify other safety devices incorporated in the autoclave offered.
3.9 | The autoclave’s drum trays and door must be constructed from stainless steel. (State grade of stainless steel used).
3.10 | State the maximum safe working pressure of the autoclave (in Kpa).
3.11 | State the test pressure of the autoclave drum.
4.9 | The body of the autoclave must be finished off in a baked enamel or be epoxy powder coated.
3.12 | A bacteria filter must be fitted to the vacuum break circuit.
3.19 | Alarms
3.19.1 | Water supply failure
3.19.2 | Cycle failure
3.20 | Safety
3.20.1 | Pressure safety valve.
3.20.2 | Door safety catch.
3.20.3 | Error display warning indicator.
3.20.4 | Overheat protection with warning indicator.
3.20.5 | Door interlock switch.
3.20.6 | Pressure door lock.
3.21 | Water Conditioning System
3.21.1 | The bidder must indicate whether they can offer a water conditioning system for the autoclave for example: An R.O. system, Filter system, Water distiller.
4 | STANDARD ACCESSORIES and CONSUMABLES
4.1 | DIN tray
4.2 | Bacteria filters
4.3 | Printing paper
5 | Conformity Compliance (Please attach certificate)
5.1 | The unit must comply with an acceptable international safety standard such as EN 61010-2-041 for Medical Autoclaves, attached certification
5.1.1 | Please provide unique ref number of the EN 61010-2-041 ceriticate:
5.2 | OEM must comply and certified on ISO 4126-1 quality standards, attach proof of compliance
5.2.1 | Please provide unique ref number of the ISO 4126-1 ceriticate:
5.3 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
5.3.1 | Please provide unique ref number of the ISO 13485 ceriticate:
5.4 | Model on offer must comply with EU Medical Device Regulation (EU) 2017/745 or equvalent from SAHPRA recognized jurisdictions . Attach a copy of certification
5.4.1 | Please provide unique ref number of the ceriticate:
5.5 | Model quoted for must comply with EN 50081-2 and EN 50082-2. Attach a copy of certification
5.5.1 | Please provide unique ref number of the EN 50081-2 and EN 50082-2 ceriticate:
5.6 | Model quoted for must comply with EN 13060/1-4. Attach a copy of certification
5.6.1 | Please provide unique ref number of the EN 13060/1-4 ceriticate:
6 | COMPREHENSIVE MAINTENANCE
6.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) should be quoted for, as per the SCC document
6.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
6.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
6.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
6.3 | The bidder to provide the OEM recommended service interval of the unit:
6.4 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
7 | General
7.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
7.2 | The unit on offer must be the latest technology.
7.2.1 | Initial Manufacturing year of Model:
| AUTOCLAVE - TABLETOP 24-30 liters - Accessories and Consumables | Complies Yes/No
RT2-10-01a | DIN tray
RT2-10-01b | Bacteria filters
RT2-10-01c | Printing paper

## Autoclave, Tabletop, 32-38L
RT2-10-02 | AUTOCLAVE - TABLETOP 32-38 liters | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
1 | DESCRIPTION
1.1 | A strong sealed chamber in which surgical instruments, towels, dressings, and other items can be sterilized by steam under raisedpressure
2 | Product Offered:
2.1 | Make:
2.2 | Model:
2.3 | Initial Manufacturing year of Model:
3 | TECHNICAL SPECIFICATION
3.1 | The autoclave must be fully automatic with a computerized control system, which monitors all the physical parameters and must be equipped with a built-in printer to provide permanent documentation of all the sterilization parameters.
3.2 | The chamber volume must be at least 32 to 38 Litre capacity.
3.3 | The autoclave must be fitted with a high vacuum silent running vacuum pump. To ensure perfect pre and post vacuum conditions. (SPLIT vacuum).
3.4 | CLEAN WATER SUPPLY: An external water supply or an internal RESERVOIR of sufficient capacity will be acceptable.Please state capacity of RESERVOIR.
3.5 | WASTE WATER DRAINS: An external waste water drain or an internal RESERVOIR of sufficient capacity will be acceptable.Please state capacity of RESERVOIR.
3.6 | The autoclave offered must be capable of performing the following sterilization cycles: Six sterilization cycles i.e. C1-wrapped,C2-wrapped,C3-unwrapped,C4-unwrapped, C5-Prion cycle (Creutzfeldt-Jakob disease), C6-Flash Cycle.
3.7 | The autoclave offered must be capable of performing the following 3 test cycles:
3.7.1 | 1)Vacuum Test: Test to check if the pressure is correct.
3.7.2 | 2) Bowie & Dick Test: Test to check if the steam correctly penetrates a porous load.
3.7.3 | 3) Helix Test: Test to check if the steam correctly penetrates a hollow load.
3.7.4 | Sterilizing temperature: C1-134⁰ , C2-121⁰, C3-134⁰, C4-121⁰, C5-134⁰, C6-134⁰
3.7.5 | Programmable cycles : Wrapped & Unwrapped automatic.
3.7.6 | Temperature range : 100- 134⁰ with a heat shutoff function.
3.7.7 | Controls should be sealed against fluid penetration.
3.7.8 | Recorder/Printer
3.7.9 | Bidder to state sterilazation time and drying time offered.
3.7.10 | The unit should have chamber temperature monitoring, accurate to ± 1⁰C over the operating range.
3.7.11 | A suitable circuit interrupter to match the load must be fitted.
3.7.12 | Power consumption not to exceed 3,5kW (16 Amps).
3.7.13 | Supply voltage 230 V.A.C. 50Hz.
3.8 | Please specify other safety devices incorporated in the autoclave offered.
3.9 | The autoclave’s drum trays and door must be constructed from stainless steel. (State grade of stainless steel used).
3.10 | State the maximum safe working pressure of the autoclave (in Kpa).
3.11 | State the test pressure of the autoclave drum.
4.9 | The body of the autoclave must be finished off in a baked enamel or be epoxy powder coated.
3.12 | A bacteria filter must be fitted to the vacuum break circuit.
3.19 | Alarms
3.19.1 | Water supply failure
3.19.2 | Cycle failure
3.20 | Safety
3.20.1 | Pressure safety valve.
3.20.2 | Door safety catch.
3.20.3 | Error display warning indicator.
3.20.4 | Overheat protection with warning indicator.
3.20.5 | Door interlock switch.
3.20.6 | Pressure door lock.
3.21 | Water Conditioning System
3.21.1 | The bidder must indicate whether they can offer a water conditioning system for the autoclave for example: An R.O. system, Filter system, Water distiller.
4 | STANDARD ACCESSORIES and CONSUMABLES
4.1 | DIN tray
4.2 | Bacteria filters
4.3 | Printing paper
5 | Conformity Compliance (Please attach certificate)
5.1 | The unit must comply with an acceptable international safety standard such as EN 61010-2-041 for Medical Autoclaves, attached certification
5.1.1 | Please provide unique ref number of the EN 61010-2-041 ceriticate:
5.2 | OEM must comply and certified on ISO 4126-1 quality standards, attach proof of compliance
5.2.1 | Please provide unique ref number of the ISO 4126-1 ceriticate:
5.3 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
5.3.1 | Please provide unique ref number of the ISO 13485 ceriticate:
5.4 | Model on offer must comply with EU Medical Device Regulation (EU) 2017/745 or equvalent from SAHPRA recognized jurisdictions . Attach a copy of certification
5.4.1 | Please provide unique ref number of the ceriticate:
5.5 | Model quoted for must comply with EN 50081-2 and EN 50082-2. Attach a copy of certification
5.5.1 | Please provide unique ref number of the EN 50081-2 and EN 50082-2 ceriticate:
5.6 | Model quoted for must comply with EN 13060/1-4. Attach a copy of certification
5.6.1 | Please provide unique ref number of the EN 13060/1-4 ceriticate:
6 | COMPREHENSIVE MAINTENANCE
6.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) should be quoted for, as per the SCC document
6.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
6.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
6.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
6.3 | The bidder to provide the OEM recommended service interval of the unit:
6.4 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
7 | General
7.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
7.2 | The unit on offer must be the latest technology.
7.2.1 | Initial Manufacturing year of Model:
| AUTOCLAVE - TABLETOP 32-38 liters - Accessories and Consumables | Complies Yes/No | Price
RT2-10-02a | DIN tray
RT2-10-02b | Bacteria filters
RT2-10-02c | Printing paper

## Autoclave, Tabletop, 40-50L
RT2-10-03 | AUTOCLAVE - TABLETOP 40-50 liters | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
1 | DESCRIPTION
1.1 | A strong sealed chamber in which surgical instruments, towels, dressings, and other items can be sterilized by steam under raisedpressure
2 | Product Offered:
2.1 | Make:
2.2 | Model:
2.3 | Initial Manufacturing year of Model:
3 | TECHNICAL SPECIFICATION
3.1 | The autoclave must be fully automatic with a computerized control system, which monitors all the physical parameters and must be equipped with a built-in printer to provide permanent documentation of all the sterilization parameters.
3.2 | The chamber volume must be at least 40 to 50 Litre capacity.
3.3 | The autoclave must be fitted with a high vacuum silent running vacuum pump. To ensure perfect pre and post vacuum conditions. (SPLIT vacuum).
3.4 | CLEAN WATER SUPPLY: An external water supply or an internal RESERVOIR of sufficient capacity will be acceptable.Please state capacity of RESERVOIR.
3.5 | WASTE WATER DRAINS: An external waste water drain or an internal RESERVOIR of sufficient capacity will be acceptable.Please state capacity of RESERVOIR.
3.6 | The autoclave offered must be capable of performing the following sterilization cycles: Six sterilization cycles i.e. C1-wrapped,C2-wrapped,C3-unwrapped,C4-unwrapped, C5-Prion cycle (Creutzfeldt-Jakob disease), C6-Flash Cycle.
3.7 | The autoclave offered must be capable of performing the following 3 test cycles:
3.7.1 | 1)Vacuum Test: Test to check if the pressure is correct.
3.7.2 | 2) Bowie & Dick Test: Test to check if the steam correctly penetrates a porous load.
3.7.3 | 3) Helix Test: Test to check if the steam correctly penetrates a hollow load.
3.7.4 | Sterilizing temperature: C1-134⁰ , C2-121⁰, C3-134⁰, C4-121⁰, C5-134⁰, C6-134⁰
3.7.5 | Programmable cycles : Wrapped & Unwrapped automatic.
3.7.6 | Temperature range : 100- 134⁰ with a heat shutoff function.
3.7.7 | Controls should be sealed against fluid penetration.
3.7.8 | Recorder/Printer
3.7.9 | Bidder to state sterilazation time and drying time offered.
3.7.10 | The unit should have chamber temperature monitoring, accurate to ± 1⁰C over the operating range.
3.7.11 | A suitable circuit interrupter to match the load must be fitted.
3.7.12 | Power consumption not to exceed 3,5kW (16 Amps).
3.7.13 | Supply voltage 230 V.A.C. 50Hz.
3.8 | Please specify other safety devices incorporated in the autoclave offered.
3.9 | The autoclave’s drum trays and door must be constructed from stainless steel. (State grade of stainless steel used).
3.10 | State the maximum safe working pressure of the autoclave (in Kpa).
3.11 | State the test pressure of the autoclave drum.
4.9 | The body of the autoclave must be finished off in a baked enamel or be epoxy powder coated.
3.12 | A bacteria filter must be fitted to the vacuum break circuit.
3.19 | Alarms
3.19.1 | Water supply failure
3.19.2 | Cycle failure
3.20 | SAFETY
3.20.1 | Pressure safety valve.
3.20.2 | Door safety catch.
3.20.3 | Error display warning indicator.
3.20.4 | Overheat protection with warning indicator.
3.20.5 | Door interlock switch.
3.20.6 | Pressure door lock.
3.21 | Water Conditioning System
3.21.1 | The bidder must indicate whether they can offer a water conditioning system for the autoclave for example: An R.O. system, Filter system, Water distiller.
4 | STANDARD ACCESSORIES and CONSUMABLES
4.1 | DIN tray
4.2 | Bacteria filters
4.3 | Printing paper
5 | Conformity Compliance (Please attach certificate)
5.1 | The unit must comply with an acceptable international safety standard such as EN 61010-2-041 for Medical Autoclaves, attached certification
5.1.1 | Please provide unique ref number of the EN 61010-2-041 ceriticate:
5.2 | OEM must comply and certified on ISO 4126-1 quality standards, attach proof of compliance
5.2.1 | Please provide unique ref number of the ISO 4126-1 ceriticate:
5.3 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
5.3.1 | Please provide unique ref number of the ISO 13485 ceriticate:
5.4 | Model on offer must comply with EU Medical Device Regulation (EU) 2017/745 or equvalent from SAHPRA recognized jurisdictions . Attach a copy of certification
5.4.1 | Please provide unique ref number of the ceriticate:
5.5 | Model quoted for must comply with EN 50081-2 and EN 50082-2. Attach a copy of certification
5.5.1 | Please provide unique ref number of the EN 50081-2 and EN 50082-2 ceriticate:
5.6 | Model quoted for must comply with EN 13060/1-4. Attach a copy of certification
5.6.1 | Please provide unique ref number of the EN 13060/1-4 ceriticate:
6 | COMPREHENSIVE MAINTENANCE
6.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) should be quoted for, as per the SCC document
6.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
6.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
6.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
6.3 | The bidder to provide the OEM recommended service interval of the unit:
6.4 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
7 | General
7.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
7.2 | The unit on offer must be the latest technology.
7.2.1 | Initial Manufacturing year of Model:
| AUTOCLAVE - TABLETOP 40-50 liters -Accessories and Consumables | Complies Yes/No | Price
RT2-10-03a | DIN tray
RT2-10-03b | Bacteria filters
RT2-10-03c | Printing paper

## Ultrasonic cleaner 30L
RT2-10-04 | ULTRASONIC CLEANER | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
1 | SCOPE
1.1 | This specification establishes the requirements for the supply, delivery, installation, demonstration and commissioning of an ultrasonic cleaner for use in theatres and CSSD's
2 | Product Offered:
2.1 | Make:
2.2 | Model:
2.3 | Initial Manufacturing year of Model:
3 | TECHNICAL SPECIFICATION
3.1 | Intended use: The device will be used for pre-cleaning of instruments, especially hollow instruments, and other necessities before sterilization.
3.2 | The ultrasonic cleaner shall be used for the cleaning of tough dirt, blood, bacteria etc on instruments used during surgical procedures.
3.3 | The unit shall be equipped with a temperature controller to set water temperature up to approximately 70 degrees. (Although it is not advisable to heat solution in medical cleaning, to prevent protein coagulation, therefore it shall also be possible to use with solution at ambient temperature)
3.4 | The unit shall be equipped with a mechanical or electronic timer to enable setting of ultrasonic output time from 0 to 20 minutes.
3.5 | The ultrasound frequency shall be approximately 42 KHz. Bidders to state ultrasound frequency, as well as the advantage of the frequency the unit operates at.
3.6 | The tank internal capacity shall be approximately 30 litres and the size approximately 600 X 300 X 200 (LXWXH) – Longest instrument is 580mm, therefore the requirement for length. The closest offer that can satisfy the requirement for the long instruments will be favourably taken into account.
3.7 | The offered unit shall have an ultrasonic output power of approximately 500Watt.
3.8 | The unit shall have a wire basket and a lid.
3.9 | The unit shall be manufactured of good quality stainless steel 304L, and have a smooth, polished surface.
3.10 | The unit shall be equipped with a bottom mounted drain tap to enable draining of solution for cleaning purposes.
4 | Standard Accessories
4.1 | Wire basket
4.2 | Lid
5 | Conformity Compliance (Please attach certificate)
5.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
5.3 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
5.3.1 | Please provide unique ref number of the ISO 13485 ceriticate:
5.4 | Model on offer must comply with EU Medical Device Regulation (EU) 2017/745 or equvalent from SAHPRA recognized jurisdictions . Attach a copy of certification
5.4.1 | Please provide unique ref number of the ceriticate:
5.5 | The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade (rubber to with 2 screws) type and it must be a minimum length of three (3) meters. NB the mains cable of the unit tendered for must be SABS colour coded.
5.6 | The equipment quoted for must be protected against electromagnetic interference IEC 60601-1-2
6 | COMPREHENSIVE MAINTENANCE
6.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) should be quoted for, as per the SCC document
6.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
7 | General
7.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
7.2 | The unit on offer must be the latest technology.
7.2.1 | Initial Manufacturing year of Model:

## Heat sealer
RT2-10-05 | HEAT SEALING MACHINE | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
1 | SCOPE
1.1 | This specification establishes the requirements for the supply, delivery, installation, demonstration and commissioning of a heat sealing machine
2 | Product Offered:
2.1 | Make:
2.2 | Model:
2.3 | Initial Manufacturing year of Model:
3 | TECHNICAL SPECIFICATION
3.1 | Intended use: The device will be used for packaging of instruments and other necessities before sterilization and maintaining this sterility for a fixed period of time after the sterilization process.
3.2 | The device offered shall be trustworthy, provide a solid seal and operational reliability.
3.3 | The heat sealer shall be both internally and externally made of stainless steel in order to reduce maintenance to a minimum and enhance hygiene and infection control.
3.4 | The device shall be suitable for laminated pouches as well as paper pouches
3.5 | Bidder to state type of seal which the device will produce.
3.6 | The device shall have an automatic transportation start as well as a reverse transportation feature which ensures safety and prevents the instruments from being damaged.
3.7 | An accurate sealing temperature setting by means of a keyboard which has an LCD display with a good visual read out.
3.8 | Temperature shall range between 100-199°C
3.9 | Temperature tolerance: ± 2°C
3.10 | The machine shall have a built-in mechanism to protect it from overheating.
3.11 | It shall be possible to do Seal edge adjustment.
3.12 | Seal width shall at least be 15mm
3.13 | The gradual feeder adjustment shall be 10-35mm
3.14 | Bidder to state the transportation speed (m/min).
3.15 | The device shall be able to self-adjust the pressure needed during the sealing process
3.16 | The device shall be equipped with a production counter with audible alarm
3.17 | The device shall be equipped with a safety measure on the auto insert feed which stops the machine from sealing when there is a temperature difference of 5°C from the set point.
3.18 | The unit shall be controlled by a microprocessor. Bidder to state control mechanism.
3.19 | Bidder to indicate whether the device offers a standby system and specify the type of standby to be provided.
3.20 | The device shall be resettable
3.21 | Dimensions: ± 610 x 180mm
3.22 | Mains voltage: 220/240 VAC
3.23 | Capacity: Max: 330W
3.24 | Weight: ± 12 kg
4 | standard Accessories
4.1 | Front table / roller front table
5 | Conformity Compliance (Please attach certificate)
5.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
5.3 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
5.3.1 | Please provide unique ref number of the ISO 13485 ceriticate:
5.4 | Model on offer must comply with EU Medical Device Regulation (EU) 2017/745 or equvalent from SAHPRA recognized jurisdictions . Attach a copy of certification
5.4.1 | Please provide unique ref number of the ceriticate:
5.5 | The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade (rubber to with 2 screws) type and it must be a minimum length of three (3) meters. NB the mains cable of the unit tendered for must be SABS colour coded.
5.6 | The equipment quoted for must be protected against electromagnetic interference IEC 60601-1-2
6 | COMPREHENSIVE MAINTENANCE
6.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) should be quoted for, as per the SCC document
6.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
7 | General
7.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
7.2 | The unit on offer must be the latest technology.
7.2.1 | Initial Manufacturing year of Model:

## Plasma Sterilizer 35-50L
RT2-10-06 | PLASMA STERILIZER, LOW TEMPERATURE, 35- 50L | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
1 | DESCRIPTION
1.1 | Low Temperature Sterilizer that deliver rapid and safe sterilization for delicate intruments that are sensitive to heat and moisture
2 | Product Offered:
2.1 | Make:
2.2 | Model:
2.3 | Initial Manufacturing year of Model:
3 | TECHNICAL SPECIFICATION
3.1 | The unit offered must be fully automatic, microprocessor controlled and feature advanced technology for safe low-temperature plasma sterilization.
3.2 | The unit's frames, external panels, valves, piping and internal shelving must be constructed from 316 grade stainless steel material
3.3 | The Plasma Sterilizer shall use hydrogen peroxide (H2O2) as sterilization agent
3.4 | The unit shall have processing cycles, which are pre-programmed with the recommended sterilizing cycles as follows (+5 mins)
3.4.1 | Fast Cycle time: 40 mins
3.4.2 | Standard Cycle time: 50 Mins
3.4.3 | Advanced Cycle time: 60 Mins
3.5 | The sterilization temperature must be within + 45OC to 50OC
3.6 | The unit offered shall have auto interlock and open pressure monitoring system that ensures all chamber pressure has been released prior to allowing the door to open
3.7 | The chamber must be able to withstand pressure from vacuum up to 3.5 bars
3.8 | The sterilizing chamber shall be in stainless steel and rectangular with stainless steel instrument loading trays to hold pouches in the correct position.
3.9 | The usable volume of the chamber shall be between 35 Litres to 50 Litres
3.10 | The unit offered shall have a built in thermal printer that automatically and continuously monitor and record dates, time of day, load, identification no. alarms, error display and operating parameters.
3.11 | The front panel on the unit shall have user interface control panel, clear menu system icons, visual messages and warning alarms.
3.12 | Display
3.12.1 | It must be at least 7" touch screen display
3.12.2 | The follwing fuctions must visible:
3.12.3 | Date and time
3.12.4 | Temperature
3.12.5 | Set point follow up
3.12.6 | Cycle selected
3.12.7 | Phase
3.12.8 | System errors
3.13 | The unit shall have a built in wheels to provide mobility
3.14 | The following alarms must be available in both audio and visual:
3.14.1 | Power failure
3.14.2 | Low/High Temp
3.14.3 | Door lock
3.14.4 | Vacuum & Pressure alarms
3.15 | The unit must be a single manual door
3.16 | The unit must have a built-in thermal printer
3.17 | USB and Ethernet connectivity must be available in the unit
3.18 | Power
3.18.1 | The unit must utilize 230v mains supply with 50Hz
3.18.2 | Power consumption of the unit must not exceed 2.5 kVA
4 | STANDARD ACCESSORIES and CONSUMABLES
4.1 | 10 x thermal printing paper rolls
4.2 | 2 x Instrument loading trays
4.3 | 10 x Hydrogen Peroxide cassettes or Cartridges
4.4 | Biological and Chemical Indicators
5 | Conformity Compliance (Please attach certificate)
5.1 | The unit must comply with an acceptable international safety standard such as EN 61010-2-041 for Medical Autoclaves, attached certification
5.1.1 | Please provide unique ref number of the EN 61010-2-041 ceriticate:
5.2 | OEM must comply and certified on ISO 4126-1 quality standards, attach proof of compliance
5.2.1 | Please provide unique ref number of the ISO 4126-1 ceriticate:
5.3 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
5.3.1 | Please provide unique ref number of the ISO 13485 ceriticate:
5.4 | Model on offer must comply with EU Medical Device Regulation (EU) 2017/745 or equvalent from SAHPRA recognized jurisdictions . Attach a copy of certification
5.4.1 | Please provide unique ref number of the ceriticate:
5.5 | Model quoted for must comply with EN 50081-2 and EN 50082-2. Attach a copy of certification
5.5.1 | Please provide unique ref number of the EN 50081-2 and EN 50082-2 ceriticate:
5.6 | Model quoted for must comply with EN 13060/1-4. Attach a copy of certification
5.6.1 | Please provide unique ref number of the EN 13060/1-4 ceriticate:
5.7 | Model quoted for must comply with ISO 17665-1. Attach a copy of certification
5.7.1 | Please provide unique ref number of the ISO 17665-1 ceriticate:
5.8 | Local Manufacturers must provide certification that unit bidded for comply with SANS 982: 2009
6 | COMPREHENSIVE MAINTENANCE
6.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) should be quoted for, as per the SCC document
6.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
6.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
6.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
6.3 | The bidder to provide the OEM recommended service interval of the unit:
6.4 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
7 | General
7.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
7.2 | The unit on offer must be the latest technology.
7.2.1 | Initial Manufacturing year of Model:
| PLASMA STERILIZER, LOW TEMPERATURE, 35- 50L - Accessories and Consumables | Complies Yes/No
RT2-10-06a | Thermal printing paper rolls (Pk of 10)
RT2-10-06b | Instrument loading tray
RT2-10-06c | Hydrogen Peroxide cassettes or Cartridges (Pk of 5)
RT2-10-06d | Biological Indicators
RT2-10-06e | Chemical Indicators

## Plasma Sterilizer 85-110L
RT2-10-07 | PLASMA STERILIZER, LOW TEMPERATURE, 85-110L | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
1 | DESCRIPTION
1.1 | Low Temperature Sterilizer that deliver rapid and safe sterilization for delicate intruments that are sensitive to heat and moisture
2 | Product Offered:
2.1 | Make:
2.2 | Model:
2.3 | Initial Manufacturing year of Model:
3 | TECHNICAL SPECIFICATION
3.1 | The unit offered must be fully automatic, microprocessor controlled and feature advanced technology for safe low-temperature plasma sterilization.
3.2 | The unit's frames, external panels, valves, piping and internal shelving must be constructed from 316 grade stainless steel material
3.3 | The Plasma Sterilizer shall use hydrogen peroxide (H2O2) as sterilization agent
3.4 | The unit shall have processing cycles, which are pre-programmed with the recommended sterilizing cycles as follows (+5 mins)
3.4.1 | Fast Cycle time: 40 mins
3.4.2 | Standard Cycle time: 50 Mins
3.4.3 | Advanced Cycle time: 60 Mins
3.5 | The sterilization temperature must be within + 45OC to 50OC
3.6 | The unit offered shall have auto interlock and open pressure monitoring system that ensures all chamber pressure has been released prior to allowing the door to open
3.7 | The chamber must be able to withstand pressure from vacuum up to 3.5 bars
3.8 | The sterilizing chamber shall be in stainless steel and rectangular with stainless steel instrument loading trays to hold pouches in the correct position.
3.9 | The usable volume of the chamber shall be between 85 Litres to 110 Litres
3.10 | The unit offered shall have a built in thermal printer that automatically and continuously monitor and record dates, time of day, load, identification no. alarms, error display and operating parameters.
3.11 | The front panel on the unit shall have user interface control panel, clear menu system icons, visual messages and warning alarms.
3.12 | Display
3.12.1 | It must be at least 7" touch screen display
3.12.2 | The follwing fuctions must visible:
3.12.3 | Date and time
3.12.4 | Temperature
3.12.5 | Set point follow up
3.12.6 | Cycle selected
3.12.7 | Phase
3.12.8 | System errors
3.13 | The unit shall have a built in wheels to provide mobility
3.14 | The following alarms must be available in both audio and visual:
3.14.1 | Power failure
3.14.2 | Low/High Temp
3.14.3 | Door lock
3.14.4 | Vacuum & Pressure alarms
3.15 | The unit must be a Single auto sliding door and two shelves
3.16 | The unit must have a built-in thermal printer
3.17 | USB and Ethernet connectivity must be available in the unit
3.18 | Power
3.18.1 | The unit must utilize 230v mains supply with 50Hz
3.18.2 | Power consumption of the unit must not exceed 2.5 kVA
4 | STANDARD ACCESSORIES
4.1 | 10 x thermal printing paper rolls
4.2 | 4 x Instrument loading trays
4.3 | 10 x Hydrogen Peroxide cassettes or Cartridges
4.4 | Biological and Chemical Indicators
5 | Conformity Compliance (Please attach certificate)
5.1 | The unit must comply with an acceptable international safety standard such as EN 61010-2-041 for Medical Autoclaves, attached certification
5.1.1 | Please provide unique ref number of the EN 61010-2-041 ceriticate:
5.2 | OEM must comply and certified on ISO 4126-1 quality standards, attach proof of compliance
5.2.1 | Please provide unique ref number of the ISO 4126-1 ceriticate:
5.3 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
5.3.1 | Please provide unique ref number of the ISO 13485 ceriticate:
5.4 | Model on offer must comply with EU Medical Device Regulation (EU) 2017/745 or equvalent from SAHPRA recognized jurisdictions . Attach a copy of certification
5.4.1 | Please provide unique ref number of the ceriticate:
5.5 | Model quoted for must comply with EN 50081-2 and EN 50082-2. Attach a copy of certification
5.5.1 | Please provide unique ref number of the EN 50081-2 and EN 50082-2 ceriticate:
5.6 | Model quoted for must comply with EN 13060/1-4. Attach a copy of certification
5.6.1 | Please provide unique ref number of the EN 13060/1-4 ceriticate:
5.7 | Model quoted for must comply with ISO 17665-1. Attach a copy of certification
5.7.1 | Please provide unique ref number of the ISO 17665-1 ceriticate:
5.8 | Local Manufacturers must provide certification that unit bidded for comply with SANS 982: 2009
6 | COMPREHENSIVE MAINTENANCE
6.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) should be quoted for, as per the SCC document
6.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
6.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
6.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
6.3 | The bidder to provide the OEM recommended service interval of the unit:
6.4 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
7 | General
7.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
7.2 | The unit on offer must be the latest technology.
7.2.1 | Initial Manufacturing year of Model:
| PLASMA STERILIZER, LOW TEMPERATURE, 85- 110L - Accessories and Consumables | Complies Yes/No | Price
RT2-10-07a | Thermal printing paper rolls (Pk of 10)
RT2-10-07b | Instrument loading tray
RT2-10-07c | Hydrogen Peroxide cassettes or Cartridges (Pk of 5)
RT2-10-07d | Biological Indicators
RT2-10-07e | Chemical Indicators

## Plasma Sterilizer 145-160L
RT2-10-08 | PLASMA STERILIZER, LOW TEMPERATURE, 145-160L | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
1 | DESCRIPTION
1.1 | Low Temperature Sterilizer that deliver rapid and safe sterilization for delicate intruments that are sensitive to heat and moisture
2 | Product Offered:
2.1 | Make:
2.2 | Model:
2.3 | Initial Manufacturing year of Model:
3 | TECHNICAL SPECIFICATION
3.1 | The unit offered must be fully automatic, microprocessor controlled and feature advanced technology for safe low-temperature plasma sterilization.
3.2 | The unit's frames, external panels, valves, piping and internal shelving must be constructed from 316 grade stainless steel material
3.3 | The Plasma Sterilizer shall use hydrogen peroxide (H2O2) as sterilization agent
3.4 | The unit shall have processing cycles, which are pre-programmed with the recommended sterilizing cycles as follows (+5 mins)
3.4.1 | Fast Cycle time: 40 mins
3.4.2 | Standard Cycle time: 50 Mins
3.4.3 | Advanced Cycle time: 60 Mins
3.5 | The sterilization temperature must be within + 45OC to 50OC
3.6 | The unit offered shall have auto interlock and open pressure monitoring system that ensures all chamber pressure has been released prior to allowing the door to open
3.7 | The chamber must be able to withstand pressure from vacuum up to 3.5 bars
3.8 | The sterilizing chamber shall be in stainless steel and rectangular with stainless steel instrument loading trays to hold pouches in the correct position.
3.9 | The usable volume of the chamber shall be between 145 Litres to 160 Litres
3.10 | The unit offered shall have a built in thermal printer that automatically and continuously monitor and record dates, time of day, load, identification no. alarms, error display and operating parameters.
3.11 | The front panel on the unit shall have user interface control panel, clear menu system icons, visual messages and warning alarms.
3.12 | Display
3.12.1 | It must be at least 7" touch screen display
3.12.2 | The follwing fuctions must visible:
3.12.3 | Date and time
3.12.4 | Temperature
3.12.5 | Set point follow up
3.12.6 | Cycle selected
3.12.7 | Phase
3.12.8 | System errors
3.13 | The unit shall have a built in wheels to provide mobility
3.14 | The following alarms must be available in both audio and visual:
3.14.1 | Power failure
3.14.2 | Low/High Temp
3.14.3 | Door lock
3.14.4 | Vacuum & Pressure alarms
3.15 | The unit must be a Single auto sliding door and two shelves
3.16 | The unit must have a built-in thermal printer
3.17 | USB and Ethernet connectivity must be available in the unit
3.18 | Power
3.18.1 | The unit must utilize 230v mains supply with 50Hz
3.18.2 | Power consumption of the unit must not exceed 2.5 kVA
4 | STANDARD ACCESSORIES
4.1 | 10 x thermal printing paper rolls
4.2 | 4 x Instrument loading trays
4.3 | 10 x Hydrogen Peroxide cassettes or Cartridges
4.4 | Biological and Chemical Indicators
5 | Conformity Compliance (Please attach certificate)
5.1 | The unit must comply with an acceptable international safety standard such as EN 61010-2-041 for Medical Autoclaves, attached certification
5.1.1 | Please provide unique ref number of the EN 61010-2-041 ceriticate:
5.2 | OEM must comply and certified on ISO 4126-1 quality standards, attach proof of compliance
5.2.1 | Please provide unique ref number of the ISO 4126-1 ceriticate:
5.3 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
5.3.1 | Please provide unique ref number of the ISO 13485 ceriticate:
5.4 | Model on offer must comply with EU Medical Device Regulation (EU) 2017/745 or equvalent from SAHPRA recognized jurisdictions . Attach a copy of certification
5.4.1 | Please provide unique ref number of the ceriticate:
5.5 | Model quoted for must comply with EN 50081-2 and EN 50082-2. Attach a copy of certification
5.5.1 | Please provide unique ref number of the EN 50081-2 and EN 50082-2 ceriticate:
5.6 | Model quoted for must comply with EN 13060/1-4. Attach a copy of certification
5.6.1 | Please provide unique ref number of the EN 13060/1-4 ceriticate:
5.7 | Model quoted for must comply with ISO 17665-1. Attach a copy of certification
5.7.1 | Please provide unique ref number of the ISO 17665-1 ceriticate:
5.8 | Local Manufacturers must provide certification that unit bidded for comply with SANS 982: 2009
6 | COMPREHENSIVE MAINTENANCE
6.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) should be quoted for, as per the SCC document
6.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
6.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
6.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
6.3 | The bidder to provide the OEM recommended service interval of the unit:
6.4 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
7 | General
7.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
7.2 | The unit on offer must be the latest technology.
7.2.1 | Initial Manufacturing year of Model:
| PLASMA STERILIZER, LOW TEMPERATURE, 145- 160L - Accessories and Consumables | Complies Yes/No
RT2-10-08a | Thermal printing paper rolls (Pk of 10)
RT2-10-08b | Instrument loading tray
RT2-10-08c | Hydrogen Peroxide cassettes or Cartridges (Pk of 5)
RT2-10-08d | Biological Indicators
RT2-10-08e | Chemical Indicators

## Scope Washer
RT2-10-09 | ENDOSCOPE WASHER | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
1 | DESCRIPTION
1.1 | An automatic machine used for the reprocessing of flexible endospes
2 | Product Offered:
2.1 | Make:
2.2 | Model:
2.3 | Initial Manufacturing year of Model:
3 | TECHNICAL SPECIFICATION
3.1 | The unit is required for the washing and disinfectant of flexible fiber- and videoscopes. The system shall be able to accommodate two scopes per cycle
3.2 | Bidders to state the functions of the washer as far as amount of scopes is concerned. Preference will be given to units that can accommodate two scopes per cycle, but have the option run two separate, independent cycles
3.3 | It shall be possible to wash and disinfect all makes and all types (gastro- colono- etc) of scopes currently available on the SA market
3.4 | The unit offered shall be able to use all makers’ disinfectant and detergent
3.5 | Bidders shall clearly, on a separate page, show what detergents and disinfectants should be used per cycle, and show the cost of these detergents and disinfectants. All couplers and tubing if required shall also be included in the calculations. It shall be possible to determine the running cost of the washer from these calculations, and these running costs shall be taken into account when adjudicating the offers
3.6 | The washer shall disinfect all channels of the scopes to ensure acceptable international standards of hygiene
3.7 | Bidders to clearly and fully describe the functioning of the washer, and how the abovementioned standards of hygiene are achieved with the functioning of the offered unit
3.8 | The washer shall have audio and visual alarms for all monitored parameters. The alarms shall be accurate, and no nuisance alarms shall be sounded
3.9 | The washer shall monitor the presence as well as volumes of detergent and disinfectant
3.1 | The washer shall be able to perform a fully controlled thermal self disinfectant cycle
3.11 | The washer shall monitor fluid pathways with continuous leak testing to prevent fluid ingress during any part of the cleansing cycle
3.12 | The washer shall monitor for channel blockages, and also indicate when sufficient fluid is pumped through the channels
3.13 | The washer shall incorporate high fluid circulation speed to enhance disinfectant removal and prevent bio film production
3.14 | The washer shall be able to flush air through the separate channels to remove water
3.15 | The washer shall be supplied with all accessories, ready for operation, as well as a starter pack consisting of enough consumables to last for a month when 6 cycles are done per day
3.16 | A trough that can be used, and is dedicated for the brushing of the scopes shall be included in the offer. Bidders to give full data of the offered unit
3.17 | Bidders to clearly indicate the water and electricity requirements for the offered unit
4 | Conformity Compliance (Please attach certificate)
4.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
4.2 | The unit must comply and certified on ISO 15883-4:2018 or latest, attach proof of compliance
4.2.1 | Please provide unique ref number of the ISO 15883-4:2018 ceriticate:
4.3 | The unit must comply and certified on ISO 14937, attach proof of compliance
4.3.1 | Please provide unique ref number of the ISO 14937 ceriticate:
4.5 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
4.5.1 | Please provide unique ref number of the ISO 13485 ceriticate:
4.6 | Model on offer must comply with EU Medical Device Regulation (EU) 2017/745 or equvalent from SAHPRA recognized jurisdictions . Attach a copy of certification
4.6.1 | Please provide unique ref number of the ceriticate:
4.7 | The equipment quoted for must be protected against electromagnetic interference IEC 60601-1-2
4.7.1 | Please provide unique ref number of the IEC ceriticate:
5 | COMPREHENSIVE MAINTENANCE
5.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) should be quoted for, as per the SCC document
5.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
5.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
5.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
5.3 | The bidder to provide the OEM recommended service interval of the unit:
5.4 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
6 | General
6.1 | Has the product on offer or any of its components being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
13.2 | The unit on offer must be the latest technology.
13.2.1 | Initial Manufacturing year of Model:

## Category-11 Cold Chain Storage
Cold Chain & Warm Storage Equipment
# | Item number | Item Description
1 | RT2-11-01 | Fridge, Vaccine (Chest), 80 - 90L
2 | RT2-11-02 | Fridge, Vaccine (Chest), 100 - 150L
3 | RT2-11-03 | Fridge, Vaccine (Chest), 240 - 260L
4 | RT2-11-04 | Fridge, Vaccine (Upright), 100 - 150L
5 | RT2-11-05 | Fridge, Vaccine (Upright), 200 - 250L
6 | RT2-11-06 | Fridge, Vaccine (Upright), 300 - 350L
7 | RT2-11-07 | Fridge, Vaccine, Solar-Powered (Upright), 120 - 160L
8 | RT2-11-08 | Fridge, Vaccine, Solar-Powered (Upright), 200 - 250L
9 | RT2-11-09 | Fridge, Pharmacy (Upright), 900 - 1100L
10 | RT2-11-10 | Fluid Warmer, Cabinet, 40-60L
11 | RT2-11-11 | Fluid Warmer, Cabinet, 80-150L
12 | RT2-11-12 | Fluid Warmer, Cabinet, 200-300L

## CFridge, Vaccine, 80-90L
RT2-11-01 | CHEST FRIDGE, VACCINE, 80 - 90L | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
1 | DESCRIPTION
1.1 | Vaccine fridge: For vaccine storage and ice pack freezing area with a capacity from 80 to at least 90 litres
2 | Product Offered:
2.1 | Make:
2.2 | Model:
2.3 | Initial Manufacturing year of Model:
3 | TECHNICAL SPECIFICATION
3.1 | Fridge temperature: +2°C to +8°C
3.2 | Gross refridgerating volume atleast 80 - 90L
3.3 | Vaccine storage capacity of at least 60L
3.5 | Interior lighting should be On when the door opens
3.6 | Minimum of 3 adjustable stainless steel trays with perforated design
3.7 | The door with internal illumination with control panel
3.8 | Must have an integral digital control panel on the outside of the fridge to monitor and display both the fridge and freezer internal temperatures
3.9 | Must have a fan to deliver optimum temperature stability and a faster temperature recovery after door openings
3.10 | Must include an automatic defrost to ensures more efficient cooling
3.11 | Complete with door locking system
3.12 | The unit should include individual freezer door compartment
3.13 | Must incorporate a temperature alarm which sounds if the temperature deviates from the set parameters for more than 15 minutes
3.14 | Stainless steel inner liner
3.15 | Adjustable feet
3.16 | Must have atleast 120 hour holdover facility during power failures at 43oC ambient temperature
3.17 | Chlorofluorocarbon (CFC) and Hydro Chlorofluorocarbon (HCFC) Free
3.18 | Mains cable must be in excess of 3 meters, complete with a standard South African 220/230 V Ac plug top.
3.19 | The unit must be fitted with voltage stabilizer, integrated or separately placed along side.
4 | The alarms must be audible and visual to signal the following:
4.1 | High and low temperature alarm
4.2 | Power failure
4.3 | Door open
5 | Conformity Compliance (Please attach certificate)
5.1 | OEM must comply and certified on ISO 9001 or ISO 13485 quality standards, attach proof of compliance.
5.1.1 | Please provide unique ref number of the ISO ceriticate:
5.2 | Model quoted for must have SANAS accredited body certificate for stainless steel test. Attach a copy of certification
5.2.1 | Please provide unique ref number of the stainless steel test ceriticate:
6 | COMPREHENSIVE MAINTENANCE
6.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) should be quoted for, as per the SCC document
6.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
6.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
6.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
6.3 | The bidder to provide the OEM recommended service interval of the unit:
6.4 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
7 | General
7.1 | The unit on offer must be the latest technology.
7.1.1 | Initial Manufacturing year of Model:

## CFridge, Vaccine, 100-150L
RT2-11-02 | CHEST FRIDGE, VACCINE, 100 - 150L | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
1 | DESCRIPTION
1.1 | Vaccine fridge: For vaccine storage and ice pack freezing area with a capacity from 100 to 150 litres
2 | Product Offered:
2.1 | Make:
2.2 | Model:
2.3 | Initial Manufacturing year of Model:
3 | TECHNICAL SPECIFICATION
3.1 | Fridge temperature: +2°C to +8°C
3.2 | Gross refridgerating volume atleast 100 - 150L
3.3 | Vaccine storage capacity of at least 80L
3.5 | Interior lighting should be On when the door opens
3.6 | Minimum of 3 adjustable stainless steel trays with perforated design
3.7 | The door with internal illumination with control panel
3.8 | Must have an integral digital control panel on the outside of the fridge to monitor and display both the fridge and freezer internal temperatures
3.9 | Must have a fan to deliver optimum temperature stability and a faster temperature recovery after door openings
3.10 | Must include an automatic defrost to ensures more efficient cooling
3.11 | Complete with door locking system
3.12 | The unit should include individual freezer door compartment
3.13 | Must incorporate a temperature alarm which sounds if the temperature deviates from the set parameters for more than 15 minutes
3.14 | Stainless steel inner liner
3.15 | Adjustable feet
3.16 | Must have atleast 120 hour holdover facility during power failures at 43oC ambient temperature
3.17 | Chlorofluorocarbon (CFC) and Hydro Chlorofluorocarbon (HCFC) Free
3.18 | Mains cable must be in excess of 3 meters, complete with a standard South African 220/230 V Ac plug top.
3.19 | The unit must be fitted with voltage stabilizer, integrated or separately placed along side.
4 | The alarms must be audible and visual to signal the following:
4.1 | High and low temperature alarm
4.2 | Power failure
4.3 | Door open
5 | Conformity Compliance (Please attach certificate)
5.1 | OEM must comply and certified on ISO 9001 or ISO 13485 quality standards, attach proof of compliance.
5.1.1 | Please provide unique ref number of the ISO ceriticate:
5.2 | Model quoted for must have SANAS accredited body certificate for stainless steel test. Attach a copy of certification
5.2.1 | Please provide unique ref number of the stainless steel test ceriticate:
6 | COMPREHENSIVE MAINTENANCE
6.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) should be quoted for, as per the SCC document
6.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
6.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
6.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
6.3 | The bidder to provide the OEM recommended service interval of the unit:
6.4 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
7 | General
7.1 | The unit on offer must be the latest technology.
7.1.1 | Initial Manufacturing year of Model:

## CFridge, Vaccine, 240-260L
RT2-11-03 | CHEST FRIDGE, VACCINE, 240 - 260L | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
1 | DESCRIPTION
1.1 | Vaccine fridge: For vaccine storage and ice pack freezing area with a capacity from 240 to 260 litres
2 | Product Offered:
2.1 | Make:
2.2 | Model:
2.3 | Initial Manufacturing year of Model:
3 | TECHNICAL SPECIFICATION
3.1 | Fridge temperature: +2°C to +8°C
3.2 | Gross refridgerating volume atleast 240 - 260L
3.3 | Vaccine storage capacity of at least 200L
3.5 | Interior lighting should be On when the door opens
3.6 | Minimum of 3 adjustable stainless steel trays with perforated design
3.7 | The door with internal illumination with control panel
3.8 | Must have an integral digital control panel on the outside of the fridge to monitor and display both the fridge and freezer internal temperatures
3.9 | Must have a fan to deliver optimum temperature stability and a faster temperature recovery after door openings
3.10 | Must include an automatic defrost to ensures more efficient cooling
3.11 | Complete with door locking system
3.12 | The unit should include individual freezer door compartment
3.13 | Must incorporate a temperature alarm which sounds if the temperature deviates from the set parameters for more than 15 minutes
3.14 | Stainless steel inner liner
3.15 | Adjustable feet
3.16 | Must have atleast 120 hour holdover facility during power failures at 43oC ambient temperature
3.17 | Chlorofluorocarbon (CFC) and Hydro Chlorofluorocarbon (HCFC) Free
3.18 | Mains cable must be in excess of 3 meters, complete with a standard South African 220/230 V Ac plug top.
3.19 | The unit must be fitted with voltage stabilizer, integrated or separately placed along side.
4 | The alarms must be audible and visual to signal the following:
4.1 | High and low temperature alarm
4.2 | Power failure
4.3 | Door open
5 | Conformity Compliance (Please attach certificate)
5.1 | OEM must comply and certified on ISO 9001 or ISO 13485 quality standards, attach proof of compliance.
5.1.1 | Please provide unique ref number of the ISO ceriticate:
5.2 | Model quoted for must have SANAS accredited body certificate for stainless steel test. Attach a copy of certification
5.2.1 | Please provide unique ref number of the stainless steel test ceriticate:
6 | COMPREHENSIVE MAINTENANCE
6.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) should be quoted for, as per the SCC document
6.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
6.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
6.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
6.3 | The bidder to provide the OEM recommended service interval of the unit:
6.4 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
7 | General
7.1 | The unit on offer must be the latest technology.
7.1.1 | Initial Manufacturing year of Model:

## UFridge, Vaccine, 100-150L
RT2-11-04 | UPRIGHT FRIDGE, VACCINE, 100 - 150L | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
1 | DESCRIPTION
1.1 | Vaccine fridge: For vaccine storage and ice pack freezing area with a capacity from 100 to at least 150 litres
2 | Product Offered:
2.1 | Make:
2.2 | Model:
2.3 | Initial Manufacturing year of Model:
3 | TECHNICAL SPECIFICATION
3.1 | Fridge temperature: +2°C to +8°C
3.2 | Freezer temperature: -14°C to -10°C
3.3 | Fridge capacity: 100 to 150 Litres
3.4 | Freezer capacity: 30 -50 Litres
3.5 | Interior lighting should be On when the door opens
3.6 | Minimum of 3 adjustable stainless steel trays with perforated design
3.7 | The door with internal illumination with control panel
3.8 | Must have an integral digital control panel on the outside of the fridge to monitor and display both the fridge and freezer internal temperatures
3.9 | Must have a fan to deliver optimum temperature stability and a faster temperature recovery after door openings
3.10 | Must include an automatic defrost to ensures more efficient cooling
3.11 | Complete with door locking system
3.12 | The unit should include individual freezer door compartment
3.13 | Must incorporate a temperature alarm which sounds if the temperature deviates from the set parameters for more than 15 minutes
3.14 | Stainless steel inner liner
3.15 | Adjustable feet
3.16 | Must have atleast 120 hour holdover facility during power failures.
3.17 | Chlorofluorocarbon (CFC) and Hydro Chlorofluorocarbon (HCFC) Free
3.18 | Mains cable must be in excess of 3 meters, complete with a standard South African 220/230 V Ac plug top.
3.19 | The unit must be fitted with voltage stabilizer, integrated or separately placed along side.
4 | The alarms must be audible and visual to signal the following:
4.1 | High and low temperature alarm
4.2 | Power failure
4.3 | Door open
5 | Conformity Compliance (Please attach certificate)
5.1 | OEM must comply and certified on ISO 9001 or ISO 13485 quality standards, attach proof of compliance.
5.1.1 | Please provide unique ref number of the ISO ceriticate:
5.2 | Model quoted for must have SANAS accredited body certificate for stainless steel test. Attach a copy of certification
5.2.1 | Please provide unique ref number of the stainless steel test ceriticate:
6 | COMPREHENSIVE MAINTENANCE
6.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) should be quoted for, as per the SCC document
6.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
6.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
6.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
6.3 | The bidder to provide the OEM recommended service interval of the unit:
6.4 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
7 | General
7.1 | The unit on offer must be the latest technology.
7.1.1 | Initial Manufacturing year of Model:

## UFridge, Vaccine, 200-250L
RT2-11-05 | UPRIGHT FRIDGE, VACCINE, 200 - 250L | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
1 | DESCRIPTION
1.1 | Vaccine fridge: For vaccine storage and ice pack freezing area with a capacity from 200 to at least 250 litres
2 | Product Offered:
2.1 | Make:
2.2 | Model:
2.3 | Initial Manufacturing year of Model:
3 | TECHNICAL SPECIFICATION
3.1 | Fridge temperature: +2°C to +8°C
3.2 | Freezer temperature: -14°C to -10°C
3.3 | Fridge capacity: 200 to 250 Litres
3.4 | Freezer capacity: 30 -50 Litres
3.5 | Interior lighting should be On when the door opens or separate switch
3.6 | Minimum of 3 adjustable stainless steel trays with perforated design
3.7 | The door with internal illumination with control panel
3.8 | Must have an integral digital control panel on the outside of the fridge to monitor and display both the fridge and freezer internal temperatures
3.9 | Must have a fan to deliver optimum temperature stability and a faster temperature recovery after door openings. Bidder to state how consistant temperature is achieved throughtout the whole fridge
3.10 | Must include an automatic defrost to ensures more efficient cooling
3.11 | Complete with door locking system
3.12 | The unit should include individual freezer door compartment
3.13 | Must incorporate a temperature alarm which sounds if the temperature deviates from the set parameters for more than 15 minutes
3.14 | Stainless steel inner liner
3.15 | Adjustable feet
3.16 | Must have atleast 120 hour holdover facility during power failures at an ambient temperature of 32o C.
3.17 | Chlorofluorocarbon (CFC) and Hydro Chlorofluorocarbon (HCFC) Free
3.18 | Mains cable must be in excess of 3 meters, complete with a standard South African 220/230 V Ac plug top.
3.19 | The unit must be fitted with voltage stabilizer, integrated or separately placed along side.
4 | The alarms must be audible and visual to signal the following:
4.1 | High and low temperature alarm
4.2 | Power failure
4.3 | Door open
5 | Conformity Compliance (Please attach certificate)
5.1 | OEM must comply and certified on ISO 9001 or ISO 13485 quality standards, attach proof of compliance.
5.1.1 | Please provide unique ref number of the ISO ceriticate:
5.2 | Model quoted for must have SANAS accredited body certificate for stainless steel test. Attach a copy of certification
5.2.1 | Please provide unique ref number of the stainless steel test ceriticate:
6 | COMPREHENSIVE MAINTENANCE
6.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) should be quoted for, as per the SCC document
6.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
6.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
6.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
6.3 | The bidder to provide the OEM recommended service interval of the unit:
6.4 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
7 | General
7.1 | The unit on offer must be the latest technology.
7.1.1 | Initial Manufacturing year of Model:

## UFridge, Vaccine, 300-350L
RT2-11-06 | UPRIGHT FRIDGE, VACCINE, 300 - 350 litres | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
1 | DESCRIPTION
1.1 | Vaccine fridge: For vaccine storage and ice pack freezing area with a capacity of 290 to 350 litres
2 | Product Offered:
2.1 | Make:
2.2 | Model:
2.3 | Initial Manufacturing year of Model:
3 | TECHNICAL SPECIFICATION
3.1 | Fridge temperature: +2°C to +8°C
3.2 | Freezer temperature: -14°C to -10°C
3.3 | Fridge capacity: 300 to 350 Litres
3.4 | Freezer capacity: 30 -50 Litres
3.5 | Interior lighting should be On when the door opens
3.6 | Minimum of 3 adjustable stainless steel trays with perforated design
3.7 | The door with internal illumination with control panel
3.8 | Must have an integral digital control panel on the outside of the fridge to monitor and display both the fridge and freezer internal temperatures
3.9 | Must have a fan to deliver optimum temperature stability and a faster temperature recovery after door openings. Bidder to state how consistant temperature is achieved throughtout the whole fridge
3.10 | Must include an automatic defrost to ensures more efficient cooling
3.11 | Complete with door locking system
3.12 | The unit should include individual freezer door compartment
3.13 | Must incorporate a temperature alarm which sounds if the temperature deviates from the set parameters for more than 15 minutes
3.14 | Stainless steel inner liner
3.15 | Adjustable feet
3.16 | Must have atleast 120 hour holdover facility during power failures at an ambient temperature of 32o C.
3.17 | Chlorofluorocarbon (CFC) and Hydro Chlorofluorocarbon (HCFC) Free
3.18 | Mains cable must be in excess of 3 meters, complete with a standard South African 220/230 V Ac plug top.
3.19 | The unit must be fitted with voltage stabilizer, integrated or separately placed along side.
4 | The alarms must be audible and visual to signal the following:
4.1 | High and low temperature alarm
4.2 | Power failure
4.3 | Door open
5 | Conformity Compliance (Please attach certificate)
5.1 | OEM must comply and certified on ISO 9001 or ISO 13485 quality standards, attach proof of compliance.
5.1.1 | Please provide unique ref number of the ISO ceriticate:
5.2 | Model quoted for must have SANAS accredited body certificate for stainless steel test. Attach a copy of certification
5.2.1 | Please provide unique ref number of the stainless steel test ceriticate:
6 | COMPREHENSIVE MAINTENANCE
6.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) should be quoted for, as per the SCC document
6.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
6.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
6.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
6.3 | The bidder to provide the OEM recommended service interval of the unit:
6.4 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
7 | General
7.1 | The unit on offer must be the latest technology.
7.1.1 | Initial Manufacturing year of Model:

## UFridge, Solar, 120-160L
RT2-11-07 | UPRIGHT FRIDGE, VACCINE, Solar Powered, 120 - 160L | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
1 | DESCRIPTION
1.1 | Solar direct drive vaccine fridge: For vaccine storage and ice pack freezing area with a capacity from 120 to at least 160 litres
2 | Product Offered:
2.1 | Make:
2.2 | Model:
2.3 | Initial Manufacturing year of Model:
3 | TECHNICAL SPECIFICATION
3.1 | Fridge temperature: +2°C to +8°C
3.2 | Freezer temperature: -14°C to -10°C
3.3 | Vaccine storage capacity: at least 100 Litres
3.4 | Freezer capacity: 30 -50 Litres
3.5 | Interior lighting should be On when the door opens
3.6 | Minimum of 3 adjustable stainless steel trays with perforated design
3.7 | The door with internal illumination with control panel
3.8 | Must have an integral digital display panel on the outside of the fridge to monitor and display both the fridge and freezer internal temperatures.
3.9 | Must have a fan to deliver optimum temperature stability and a faster temperature recovery after door openings
3.10 | Must include an automatic defrost to ensures more efficient cooling
3.11 | Complete with door locking system
3.12 | The unit should include individual freezer door compartment
3.13 | Must incorporate a temperature alarm which sounds if the temperature deviates from the set parameters for more than 15 minutes
3.14 | Stainless steel inner liner
3.15 | Adjustable feet
3.16 | Must have atleast 120 hour holdover facility during power failures at 43oC.
3.17 | Chlorofluorocarbon (CFC) and Hydro Chlorofluorocarbon (HCFC) Free
3.18 | Solar panel to provide sufficient stable output voltage to power the unit.
3.19 | The length of the panel output cable to the unit should be at least 25m
3.20 | The panel bracket must be adjustable to the desired angle and mountable.
3.21 | The unit must be fitted with voltage stabilizer, integrated or separately placed along side.
4 | The alarms must be audible and visual to signal the following:
4.1 | High and low temperature alarm
4.2 | Power failure
4.3 | Door open
5 | Conformity Compliance (Please attach certificate)
5.1 | OEM must comply and certified on ISO 9001 or ISO 13485 quality standards, attach proof of compliance.
5.1.1 | Please provide unique ref number of the ISO ceriticate:
5.2 | Model quoted for must have SANAS accredited body certificate for stainless steel test. Attach a copy of certification
5.2.1 | Please provide unique ref number of the stainless steel test ceriticate:
6 | COMPREHENSIVE MAINTENANCE
6.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) should be quoted for, as per the SCC document
6.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
6.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
6.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
6.3 | The bidder to provide the OEM recommended service interval of the unit:
6.4 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
7 | General
7.1 | The unit on offer must be the latest technology.
7.1.1 | Initial Manufacturing year of Model:

## UFridge, Solar, 200-250L
RT2-11-08 | UPRIGHT FRIDGE, VACCINE, Solar Powered, 200 - 250L | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
1 | DESCRIPTION
1.1 | Solar direct drive vaccine fridge: For vaccine storage and ice pack freezing area with a capacity from 200 to at least 250 litres
2 | Product Offered:
2.1 | Make:
2.2 | Model:
2.3 | Initial Manufacturing year of Model:
3 | TECHNICAL SPECIFICATION
3.1 | Fridge temperature: +2°C to +8°C
3.2 | Freezer temperature: -14°C to -10°C
3.3 | Vaccine storage capacity: at least 180 Litres
3.4 | Freezer capacity: 30 -50 Litres
3.5 | Interior lighting should be On when the door opens
3.6 | Minimum of 3 adjustable stainless steel trays with perforated design
3.7 | The door with internal illumination with control panel
3.8 | Must have an integral digital control panel on the outside of the fridge to monitor and display both the fridge and freezer internal temperatures
3.9 | Must have a fan to deliver optimum temperature stability and a faster temperature recovery after door openings
3.10 | Must include an automatic defrost to ensures more efficient cooling
3.11 | Complete with door locking system
3.12 | The unit should include individual freezer door compartment
3.13 | Must incorporate a temperature alarm which sounds if the temperature deviates from the set parameters for more than 15 minutes
3.14 | Stainless steel inner liner
3.15 | Adjustable feet
3.16 | Must have atleast 120 hour holdover facility during power failures at 43oC.
3.17 | Chlorofluorocarbon (CFC) and Hydro Chlorofluorocarbon (HCFC) Free
3.18 | Solar panel to provide sufficient stable output voltage to power the unit.
3.19 | The length of the panel output cable to the unit should be at least 25m
3.20 | The panel bracket must be adjustable to the desired angle and mountable.
3.21 | The unit must be fitted with voltage stabilizer, integrated or separately placed along side.
4 | The alarms must be audible and visual to signal the following:
4.1 | High and low temperature alarm
4.2 | Power failure
4.3 | Door open
5 | Conformity Compliance (Please attach certificate)
5.1 | OEM must comply and certified on ISO 9001 or ISO 13485 quality standards, attach proof of compliance.
5.1.1 | Please provide unique ref number of the ISO ceriticate:
5.2 | Model quoted for must have SANAS accredited body certificate for stainless steel test. Attach a copy of certification
5.2.1 | Please provide unique ref number of the stainless steel test ceriticate:
6 | COMPREHENSIVE MAINTENANCE
6.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) should be quoted for, as per the SCC document
6.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
6.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
6.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
6.3 | The bidder to provide the OEM recommended service interval of the unit:
6.4 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
7 | General
7.1 | The unit on offer must be the latest technology.
7.1.1 | Initial Manufacturing year of Model:

## Fridge, Pharmacy, 900-1100L
RT2-11-09 | UPRIGHT FRIDGE, PHARMACY, 900-1100L | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
1 | DESCRIPTION
1.1 | Double glass door fridge: To store Pharmaceutical stock correctly, especially in relation to those medicines which require storage at low temperatures.
2 | Product Offered:
2.1 | Make:
2.2 | Model:
2.3 | Initial Manufacturing year of Model:
3 | TECHNICAL SPECIFICATION
3.1 | Double glass door fridge: To store Pharmaceutical stock correctly, especially in relation to those medicines which require storage at low temperatures.
3.2 | Approximate size: DEPTH 540mm to 700mm ; WIDTH 1100mm to 1400mm; HEIGHT 1300mm to 2000mm
3.3 | Must display Temperature: +2°C to +8°C
3.4 | Capacity: 900 to 1100 Litres
3.5 | Interior lighting should be On when the door is open or have separate switch
3.6 | Minimum of 6 adjustable stainless steel trays with perforated design
3.7 | The unit should have three (3) layer tempered glass doors with internal illumination with external control switch
3.8 | Must have an integral digital control panel on the outside of the fridge to monitor the internal temperature
3.9 | Must have a fan to deliver optimum temperature stability and a faster temperature recovery after door openings. Bidder to state how consistant temperature is achieved throughtout the whole fridge
3.10 | Must include an automatic defrost to ensures more efficient cooling
3.11 | Complete with door locking system
3.12 | Must incorporate a temperature alarm which sounds if the temperature deviates from the set parameters for more than 15 minutes
3.13 | Stainless steel inner liner
3.14 | Must have at least 2 adjustable levelling legs and 4 castors
3.15 | Must have atleast 4 hour holdover facility during power failures at 43oC ambient temp
3.16 | Chlorofluorocarbon (CFC) and Hydro Chlorofluorocarbon (HCFC) free
3.17 | Mains cable must be in excess of 3 meters, complete with a standard South African 220/230 V Ac plug top.
3.18 | The unit must be fitted with voltage stabilizer, integrated or separately placed along side.
4 | The alarms must be audible and visual to signal the following:
4.1 | High and low temperature alarm
4.2 | Power failure
4.3 | Door open
5 | Conformity Compliance (Please attach certificate)
5.1 | OEM must comply and certified on ISO 9001 or ISO 13485 quality standards, attach proof of compliance.
5.1.1 | Please provide unique ref number of the ISO ceriticate:
5.2 | Model quoted for must have SANAS accredited body certificate for stainless steel test. Attach a copy of certification
5.2.1 | Please provide unique ref number of the stainless steel test ceriticate:
6 | COMPREHENSIVE MAINTENANCE
6.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) should be quoted for, as per the SCC document
6.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
6.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
6.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
6.3 | The bidder to provide the OEM recommended service interval of the unit:
6.4 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
7 | General
7.1 | The unit on offer must be the latest technology.
7.1.1 | Initial Manufacturing year of Model:

## Fluid Warmer, Cabinet, 40-60L
RT2-11-10 | FLUID WARMING CABINET, SMALL | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
1 | DESCRIPTION
1.1 | Stationary cabinet used to store and warm fluids and blood (from 40 - 60 Litres)
2 | Product Offered:
2.1 | Make:
2.2 | Model:
2.3 | Initial Manufacturing year of Model:
3 | TECHNICAL SPECIFICATION
3.1 | The offered unit shall have variable compartment sizes with detachable/interchangeable atleast 1 shelve at standard waist height
3.2 | The exterior of the unit must be stainless steel 304
3.3 | The chamber or inside of the unit offered must be stainless steel 304 including detachable shelves.
3.4 | The door of the offered unit should be built in with a toughened and tempered glass door/window to allow visualisation of chamber contents without opening
3.5 | The offered unit should have lighting for easy viewing as well as a light switch on the interface panel
3.6 | The unit should offer temperature ranges of not less than 30˚C and not above 42˚C. The working range should be protected by a locking mechanism which needs specific authority to prevent accidental overheating.
3.7 | The offered unit shall have an LED/LCD user interface panel to allow user to set and monitor temperatures
3.8 | Displayed temperatures should have an accuracy of ±2 ̊C
3.9 | Unit shall display distinctively SET Temperatures as well as ACTUAL chamber temperature at any given time
3.10 | Displayed temperature must be in ˚C
4 | Alarms and Safety features
4.1 | Over Temperature: High - ± 2°C
4.2 | Low Temperature: Low - ± 2°C
4.3 | Power loss
4.4 | Malfunction
4.5 | Unit shall be built with an over temperature shut off system
5 | Conformity Compliance (Please attach certificate)
5.1 | OEM must comply and certified on ISO 9001 or ISO 13485 quality standards, attach proof of compliance.
5.1.1 | Please provide unique ref number of the ISO ceriticate:
5.2 | Model quoted for must have SANAS accredited body certificate for stainless steel test. Attach a copy of certification
5.2.1 | Please provide unique ref number of the stainless steel test ceriticate:
6 | COMPREHENSIVE MAINTENANCE
6.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) should be quoted for, as per the SCC document
6.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
6.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
6.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
6.3 | The bidder to provide the OEM recommended service interval of the unit:
6.4 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
7 | General
7.1 | The unit on offer must be the latest technology.
7.1.1 | Initial Manufacturing year of Model:

## Fluid Warmer, Cabinet, 80-150L
RT2-11-11 | FLUID WARMING CABINET, MEDIUM | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
1 | DESCRIPTION
1.1 | Stationary cabinet used to store and warm fluids and blood (from 80 - 150 Litres)
2 | Product Offered:
2.1 | Make:
2.2 | Model:
2.3 | Initial Manufacturing year of Model:
3 | TECHNICAL SPECIFICATION
3.1 | The offered unit shall have variable compartment sizes with detachable/interchangeable atleast 2 shelves at standard waist height
3.2 | The exterior of the unit must be stainless steel 304
3.3 | The chamber or inside of the unit offered must be stainless steel 304 including detachable shelves.
3.4 | The door of the offered unit should be built in with a toughened and tempered glass door/window to allow visualisation of chamber contents without opening
3.5 | The offered unit should have lighting for easy viewing as well as a light switch on the interface panel
3.6 | The unit should offer temperature ranges of not less than 30˚C and not above 42˚C. The working range should be protected by a locking mechanism which needs specific authority to prevent accidental overheating.
3.7 | The offered unit shall have an LED/LCD user interface panel to allow user to set and monitor temperatures
3.8 | Displayed temperatures should have an accuracy of ±2 ̊C
3.9 | Unit shall display distinctively SET Temperatures as well as ACTUAL chamber temperature at any given time
3.10 | Displayed temperature must be in ˚C
4 | Alarms and Safety features
4.1 | Over Temperature: High - ± 2°C
4.2 | Low Temperature: Low - ± 2°C
4.3 | Power loss
4.4 | Malfunction
4.5 | Unit shall be built with an over temperature shut off system
5 | Conformity Compliance (Please attach certificate)
5.1 | OEM must comply and certified on ISO 9001 or ISO 13485 quality standards, attach proof of compliance.
5.1.1 | Please provide unique ref number of the ISO ceriticate:
5.2 | Model quoted for must have SANAS accredited body certificate for stainless steel test. Attach a copy of certification
5.2.1 | Please provide unique ref number of the stainless steel test ceriticate:
6 | COMPREHENSIVE MAINTENANCE
6.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) should be quoted for, as per the SCC document
6.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
6.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
6.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
6.3 | The bidder to provide the OEM recommended service interval of the unit:
6.4 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
7 | General
7.1 | The unit on offer must be the latest technology.
7.1.1 | Initial Manufacturing year of Model:

## Fluid Warmer, Cabinet, 200-300L
RT2-11-12 | FLUID WARMING CABINET, LARGE | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
1 | DESCRIPTION
1.1 | Stationary cabinet used to store and warm fluids and blood (from 200 - 300 Litres)
2 | Product Offered:
2.1 | Make:
2.2 | Model:
2.3 | Initial Manufacturing year of Model:
3 | TECHNICAL SPECIFICATION
3.1 | The offered unit shall have variable compartment sizes with detachable/interchangeable atleast 3 shelves at standard waist height
3.2 | The exterior of the unit must be stainless steel 304
3.3 | The chamber or inside of the unit offered must be stainless steel 304 including detachable shelves.
3.4 | The door of the offered unit should be built in with a toughened and tempered glass door/window to allow visualisation of chamber contents without opening
3.5 | The offered unit should have lighting for easy viewing as well as a light switch on the interface panel
3.6 | The unit should offer temperature ranges of not less than 30˚C and not above 42˚C. The working range should be protected by a locking mechanism which needs specific authority to prevent accidental overheating.
3.7 | The offered unit shall have an LED/LCD user interface panel to allow user to set and monitor temperatures
3.8 | Displayed temperatures should have an accuracy of ±2 ̊C
3.9 | Unit shall display distinctively SET Temperatures as well as ACTUAL chamber temperature at any given time
3.10 | Displayed temperature must be in ˚C
4 | Alarms and Safety features
4.1 | Over Temperature: High - ± 2°C
4.2 | Low Temperature: Low - ± 2°C
4.3 | Power loss
4.4 | Malfunction
4.5 | Unit shall be built with an over temperature shut off system
5 | Conformity Compliance (Please attach certificate)
5.1 | OEM must comply and certified on ISO 9001 or ISO 13485 quality standards, attach proof of compliance.
5.1.1 | Please provide unique ref number of the ISO ceriticate:
5.2 | Model quoted for must have SANAS accredited body certificate for stainless steel test. Attach a copy of certification
5.2.1 | Please provide unique ref number of the stainless steel test ceriticate:
6 | COMPREHENSIVE MAINTENANCE
6.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) should be quoted for, as per the SCC document
6.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
6.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
6.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
6.3 | The bidder to provide the OEM recommended service interval of the unit:
6.4 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
7 | General
7.1 | The unit on offer must be the latest technology.
7.1.1 | Initial Manufacturing year of Model:

## Category13-summary
INSTALLATION ALTERATIONS REQUIRED
# | Item number | Item Description
1 | RT2-13-01 | Instrument Washer, Under-counter
2 | RT2-13-02 | Instrument Washer, Single door
3 | RT2-13-03 | Instrument Washer, Double door
4 | RT2-13-04 | Theater Light, Ceiling-Mount
5 | RT2-13-05 | Dialysis Machine (CRRT)
6 | RT2-13-06 | Intermittent Hemodialysis
7 | RT2-13-07 | Mobile RO system
8 | RT2-13-08 | Complete RO plant
9 | RT2-13-09 | Nurse Call System, 8 Bed
10 | RT2-13-10 | Nurse Call System, 16 Bed
11 | RT2-13-11 | Nurse Call System, 24 Bed

## Instrument Washer, Undercounter
RT2-13-01 | INSTRUMENT WASHER, UNDER COUNTER | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
1 | DESCRIPTION
1.1 | Under counter Intrument Washer for cleaning, thermal disinfecting and drying surgical instruments, glassware and anesthesia material
1a | Product Offered:
1a.1 | Make:
1a.2 | Model:
1a.3 | Initial Manufacturing year of Model:
2 | TECHNICAL SPECIFICATION
2.1 | The unit offered must be fully automatic and microprocessor controlled.
2.2 | The unit shall have processing cycles, which are pre-programmed with the recommended treatment parameters for pre-washing, cleaning, rinsing, thermal disinfection and drying.
2.3 | The instrument washer shall have detergents and rinse agents, which are automatically dispensed during cycle
2.4 | The instrument washer shall have glass window on the load chamber
2.5 | The exterior height of the washer should be designed to fit under counter/bench
2.6 | The instrument chamber shall be fitted with at least two (2) removable stainless steel shelves/levels each able to accommodate at least 2 full size DIN trays of (WxHxD) 480mm x 250mm x 64mm.
2.8 | The instrument washer shall have rotary spray arms at the top and bottom of the chamber.
2.9 | The front panel on the instrument washer shall have a display, displaying all cycles and errors.
2.1 | The spray system and recirculation piping of the instrument washer shall be stainless steel.
2.11 | The cabinet casing of the instrument washer should be constructed of 316 L stainless steel grade.
2.12 | The instrument washer shall have a water heating system.
2.13 | The instrument washer shall have temperature control and monitoring.
2.14 | The drain capacity of the instrument washer shall be no less than 30 L/min.
2.15 | The instrument washer electrical sump heater shall be no less than 7.4kW.
2.16 | The instrument washer electrical dryer heater shall be no less than 1.5kW.
2.17 | The noise level of the instrument washer shall be no greater than 80 db in the washing and drying cycle.
2.18 | The unit must be fitted with an intergral thermal printer.
2.19 | The instrument washer pump capacities shall be as follows:
2.19.1 | a) Circulation pump – no less than 350L/min
2.19.2 | b) Pump purified water (Optional)
2.19.3 | c) Pump alkaline detergent : +_ 20 L/min.
2.2 | The instrument washer water consumption shall be no less than 12 L/min in the wash phase.
2.21 | The instrument washer shall have an audible alarm.
2.22 | The instrument washer shall have safety interlocking system to prevent operation when the door is open
2.23 | The instrument washer shall have safety interlocking system to prevent door from opening when the washer is in operation
2.24 | The instrument washer shall be supplied with drainage hose, hot and cold water hoses with approximately 1.5m in length
3 | STANDARD ACCESSORIES
3.1 | Heat resistant gloves
3.2 | 10 x thermal printing paper rolls
3.3 | 4x DIN Instrument trays
3.4 | 2L cleaning detergent
3.5 | 2L Rinse
3.6 | 1x Loading trolley
4 | Conformity Compliance (Please attach certificate)
4.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
4.2 | The unit must comply and certified on ISO 15883-4:2018 or latest, attach proof of compliance
4.2.1 | Please provide unique ref number of the ISO 15883-4:2018 ceriticate:
4.4 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
4.4.1 | Please provide unique ref number of the ISO 13485 ceriticate:
4.5 | Model on offer must comply with EU Medical Device Regulation (EU) 2017/745 or equvalent from SAHPRA recognized jurisdictions . Attach a copy of certification
4.5.1 | Please provide unique ref number of the ceriticate:
4.6 | The equipment quoted for must be protected against electromagnetic interference IEC 60601-1-2
4.6.1 | Please provide unique ref number of the IEC ceriticate:
5 | COMPREHENSIVE MAINTENANCE
5.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) should be quoted for, as per the SCC document
5.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
5.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
5.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
5.3 | The bidder to provide the OEM recommended service interval of the unit:
5.4 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
6 | General
6.1 | Has the product on offer or any of its components being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
6.2 | The unit on offer must be the latest technology.
6.2.1 | Initial Manufacturing year of Model:
| INSTRUMENT WASHER, UNDER COUNTER REPLACEMENT ACCESSORIES | Complies Yes/No | Price
RT2-13-01a | Water Pressure booster
RT2-13-01b | Water Purifier
RT2-13-01c | Heat resistant gloves (Pair)
RT2-13-01d | Thermal printing paper rolls (Pk of 10)
RT2-13-01e | DIN Instrument tray
RT2-13-01f | 2L cleaning detergent
RT2-13-01g | 2L Rinse
RT2-13-01h | Loading trolley

## Instrument Washer, Single door
RT2-13-02 | INSTRUMENT WASHER, SINGLE DOOR | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
1 | DESCRIPTION
1.1 | Free standing intrument washer for cleaning, thermal disinfecting and drying surgical instruments, glassware and anesthesia material
1a | Product Offered:
1a.1 | Make:
1a.2 | Model:
1a.3 | Initial Manufacturing year of Model:
2 | TECHNICAL SPECIFICATION
2.1 | The unit offered must be fully automatic and microprocessor controlled.
2.2 | The unit shall have processing cycles, which are pre-programmed with the recommended treatment parameters for pre-washing, cleaning, rinsing, thermal disinfection and drying.
2.3 | The instrument washer shall have detergents and rinse agents, which are automatically dispensed during cycle
2.4 | The instrument washer shall have glass window on the load chamber
2.5 | The exterior height of the washer should be sufficient to accomodate the load height and load chamber height with drying option
2.6 | The instrument chamber shall be fitted with at least four (4) removable stainless steel shelves/levels each able to accommodate at least 2 full size DIN trays of (WxHxD) 480mm x 250mm x 64mm.
2.7 | The instrument washer must have a load height of + 780mm (tolerance 50mm)
2.8 | The instrument washer shall have rotary spray arms at the top and bottom of the chamber.
2.9 | The front panel on the instrument washer shall have a display, displaying all cycles and errors.
2.1 | The spray system and recirculation piping of the instrument washer shall be stainless steel.
2.11 | The cabinet casing of the instrument washer should be constructed of stainless steel.
2.12 | The instrument washer shall have a water heating system.
2.13 | The instrument washer shall have temperature control and monitoring.
2.14 | The drain capacity of the instrument washer shall be no less than 30 L/min.
2.15 | The instrument washer electrical sump heater shall be no less than 7.4kW.
2.16 | The instrument washer electrical dryer heater shall be no less than 1.5kW.
2.17 | The noise level of the instrument washer shall be no greater than 80 db in the washing and drying cycle.
2.18 | The unit must be fitted with an intergral thermal printer.
2.19 | The instrument washer pump capacities shall be as follows:
2.19.1 | a) Circulation pump – no less than 350L/min
2.19.2 | b) Pump purified water (Optional)
2.19.3 | c) Pump alkaline detergent : +_ 20 L/min.
2.2 | The instrument washer water consumption shall be no less than 12 L/min in the wash phase.
2.21 | The instrument washer shall have an audible alarm.
2.22 | The instrument washer shall have safety interlocking system to prevent operation when the door is open
2.23 | The instrument washer shall have safety interlocking system to prevent door from opening when the washer is in operation
2.24 | The instrument washer shall be supplied with drainage hose, hot and cold water hoses with approximately 1.5m in length
2.25 | The trolley should have antistatic castors with at least 2 equipped with locking mechanism/brakes
3 | STANDARD ACCESSORIES
3.1 | Heat resistant gloves
3.2 | 10 x thermal printing paper rolls
3.3 | 8x DIN Instrument trays
3.4 | 10L cleaning detergent
3.5 | 10L Rinse
3.6 | 1x Loading trolley
4 | Conformity Compliance (Please attach certificate)
4.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
4.2 | The unit must comply and certified on ISO 15883-4:2018 or latest, attach proof of compliance
4.2.1 | Please provide unique ref number of the ISO 15883-4:2018 ceriticate:
4.4 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
4.4.1 | Please provide unique ref number of the ISO 13485 ceriticate:
4.5 | Model on offer must comply with EU Medical Device Regulation (EU) 2017/745 or equvalent from SAHPRA recognized jurisdictions . Attach a copy of certification
4.5.1 | Please provide unique ref number of the ceriticate:
4.6 | The equipment quoted for must be protected against electromagnetic interference IEC 60601-1-2
4.6.1 | Please provide unique ref number of the IEC ceriticate:
5 | COMPREHENSIVE MAINTENANCE
5.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) should be quoted for, as per the SCC document
5.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
5.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
5.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
5.3 | The bidder to provide the OEM recommended service interval of the unit:
5.4 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
6 | General
6.1 | Has the product on offer or any of its components being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
6.2 | The unit on offer must be the latest technology.
6.2.1 | Initial Manufacturing year of Model:
| INSTRUMENT WASHER, SINGLE DOOR REPLACEMENT ACCESSORIES | Complies Yes/No | Price
RT2-13-02a | Water Pressure booster
RT2-13-02b | Water Purifier
RT2-13-02c | Heat resistant gloves (Pair)
RT2-13-02d | Thermal printing paper rolls (Pk of 10)
RT2-13-02e | DIN Instrument tray
RT2-13-02f | 5L cleaning detergent
RT2-13-02g | 10L cleaning detergent
RT2-13-02h | 5L Rinse
RT2-13-02i | 10L Rinse
RT2-13-02j | Loading trolley

## Instrument Washer, Double door
RT2-13-03 | INSTRUMENT WASHER, DOUBLE DOOR | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
1 | DESCRIPTION
1.1 | Free standing intrument washer for cleaning, thermal disinfecting and drying surgical instruments, glassware and anesthesia material
1a | Product Offered:
1a.1 | Make:
1a.2 | Model:
1a.3 | Initial Manufacturing year of Model:
2 | TECHNICAL SPECIFICATION
2.1 | The unit offered must be fully automatic and microprocessor controlled.
2.2 | The unit shall have processing cycles, which are pre-programmed with the recommended treatment parameters for pre-washing, cleaning, rinsing, thermal disinfection and drying.
2.3 | The instrument washer shall have detergents and rinse agents, which are automatically dispensed during cycle
2.4 | The instrument washer shall have glass windows on both doors
2.5 | The exterior height of the washer should be sufficient to accomodate the load height and load chamber height with drying option
2.6 | The instrument chamber shall be fitted with at least four (4) removable stainless steel shelves/levels each able to accommodate at least 2 full size DIN trays of (WxHxD) 480mm x 250mm x 64mm.
2.7 | The instrument washer must have a load height of + 780mm (tolerance 50mm)
2.8 | The instrument washer shall have rotary spray arms at the top and bottom of the chamber.
2.9 | The front panel on the instrument washer shall have a display, displaying all cycles and errors.
2.1 | The spray system and recirculation piping of the instrument washer shall be stainless steel.
2.11 | The cabinet casing of the instrument washer should be constructed of stainless steel.
2.12 | The instrument washer shall have a water heating system.
2.13 | The instrument washer shall have temperature control and monitoring.
2.14 | The drain capacity of the instrument washer shall be no less than 30 L/min.
2.15 | The instrument washer electrical sump heater shall be no less than 7.4kW.
2.16 | The instrument washer electrical dryer heater shall be no less than 1.5kW.
2.17 | The noise level of the instrument washer shall be no greater than 80 db in the washing and drying cycle.
2.18 | The unit must be fitted with an intergral thermal printer.
2.19 | The instrument washer pump capacities shall be as follows:
2.19.1 | a) Circulation pump – no less than 350L/min
2.19.2 | b) Pump purified water (Optional)
2.19.3 | c) Pump alkaline detergent : +_ 20 L/min.
2.2 | The instrument washer water consumption shall be no less than 12 L/min in the wash phase.
2.21 | The instrument washer shall have an audible alarm.
2.22 | The instrument washer shall have safety interlocking system to prevent operation when the door is open
2.23 | The instrument washer shall have safety interlocking system to prevent door from opening when the washer is in operation
2.24 | The instrument washer shall be supplied with drainage hose, hot and cold water hoses with approximately 1.5m in length
2.25 | The trolley should have antistatic castors with at least 2 equipped with locking mechanism/brakes
3 | STANDARD ACCESSORIES
3.1 | Heat resistant gloves
3.2 | 10 x thermal printing paper rolls
3.3 | 8x DIN Instrument trays
3.4 | 10L cleaning detergent
3.5 | 10L Rinse
3.6 | 2x Loading trolleys (Clean & Dirty side)
4 | Conformity Compliance (Please attach certificate)
4.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
4.2 | The unit must comply and certified on ISO 15883-4:2018 or latest, attach proof of compliance
4.2.1 | Please provide unique ref number of the ISO 15883-4:2018 ceriticate:
4.4 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
4.4.1 | Please provide unique ref number of the ISO 13485 ceriticate:
4.5 | Model on offer must comply with EU Medical Device Regulation (EU) 2017/745 or equvalent from SAHPRA recognized jurisdictions . Attach a copy of certification
4.5.1 | Please provide unique ref number of the ceriticate:
4.6 | The equipment quoted for must be protected against electromagnetic interference IEC 60601-1-2
4.6.1 | Please provide unique ref number of the IEC ceriticate:
5 | COMPREHENSIVE MAINTENANCE
5.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) should be quoted for, as per the SCC document
5.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
5.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
5.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
5.3 | The bidder to provide the OEM recommended service interval of the unit:
5.4 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
6 | General
6.1 | Has the product on offer or any of its components being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
6.2 | The unit on offer must be the latest technology.
6.2.1 | Initial Manufacturing year of Model:
| INSTRUMENT WASHER, DOUBLE DOOR REPLACEMENT ACCESSORIES | Complies Yes/No
RT2-13-03a | Water Pressure booster
RT2-13-03b | Water Purifier
RT2-13-03c | Heat resistant gloves (Pair)
RT2-13-03d | Thermal printing paper rolls (Pk of 10)
RT2-13-03e | DIN Instrument tray
RT2-13-03f | 5L cleaning detergent
RT2-13-03g | 10L cleaning detergent
RT2-13-03h | 5L Rinse
RT2-13-03i | 10L Rinse
RT2-13-03j | Loading trolley

## Theater Light, Ceiling-mount
RT2-13-04 | Theater Light, Ceiling-Mount | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
1 | DESCRIPTION
1.1
1a | Product Offered:
1a.1 | Make:
1a.2 | Model:
1a.3 | Initial Manufacturing year of Model:
2 | TECHNICAL SPECIFICATION
2.1 | The requirement is for a double head system, with a main lamp with output of at least 160 000 lux, and a satellite lamp system with output of at least 110 000 lux. All measurements at 1 metre
2.2 | The lamps shall be ceiling suspended, and all movements required during a normal surgical procedure shall be possible. During the site visit to an installed base the movements of the lamps shall be used as evaluation criteria for the adjudication of the offers
2.3 | The ceiling mounted suspension shall be so suspended that any position may be selected with ease, for illuminating any required operating area on the patient, with virtually unlimited rotatability
2.4 | The lamp shall be spring balanced in the supporting arm, be easy to move, have no backlash, and shall have friction clutches or similar devices for adjustment
2.5 | The lamp shall be capable of shining both vertically downwards for normal use, and horizontally for lithotomy position operations
2.6 | Good, positive yet easy flexibility of height, lateral and tilt movement is required
2.7 | The lamp head shall be equipped with scratch-resistant material to protect the LED’s from accidental damage. The lamp head shall be of modern design combined in one housing
2.8 | The unit shall have easy, positive manoeuvrability
2.9 | The unit shall have excellent laminar flow features, and smooth surfaces that offer low resistance to airflow
2.1 | The light head and the positioning of the LED’s shall ensure a shadowless working area, regardless of the surgeon and supports staff’s position
2.11 | The focusable size of the light field for the main lamp shall be approximately 200 to 300 mm, and for the second lamp approximately 180 to 230mm. Bidders to state the way focusing is achieved. It shall not be necessary to move the lamp away from the patient to achieve a bigger focal spot
2.12 | Depth of illumination shall be more than 1 000mm without refocusing
2.13 | LED life shall be guaranteed to be at least 40 000 hours
2.14 | The lamp head shall have smooth, flat surfaces for easy cleaning
2.15 | The lights shall be electrically dimmable from approximately 30 to 100% of full illumination. The lamps shall have an “Endoscopy” setting to dim the lights to low light condition. Bidders to state the % of full light achieved during “Endoscopy” dimming
2.16 | The light shall have a colour temperature of approximately 4 300K and a colour rendering index of 95%.
2.17 | Bidders to state the composition of the LED’s in the lamphead, and what replacement of the LED’s entail – per each, per block etc, as well as the current cost for replacement of LED’s
2.18 | UPS backup is required for each lamp. The UPS shall keep the lamp functional on full load for at least one hour after power failure
2.19 | A removable sterilisable handle on the lamp head shall enable the surgeon to easily adjust the position of the light, as well as setting the light field or focus. A spare handle shall be supplied with each unit
2.2 | If ceilings or any other construction is in any way damaged during the installation, repair and making good shall be done immediately
3 | STANDARD ACCESSORIES
3.1 | Handles
4 | Conformity Compliance (Please attach certificate)
4.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
4.3 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
4.3.1 | Please provide unique ref number of the ISO 13485 ceriticate:
4.4 | Model on offer must comply with EU Medical Device Regulation (EU) 2017/745 or equvalent from SAHPRA recognized jurisdictions . Attach a copy of certification
4.4.1 | Please provide unique ref number of the ceriticate:
4.5 | The equipment quoted for must be protected against electromagnetic interference IEC 60601-1-2
4.5.1 | Please provide unique ref number of the IEC ceriticate:
5 | COMPREHENSIVE MAINTENANCE
5.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) should be quoted for, as per the SCC document
5.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
5.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
5.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
5.3 | The bidder to provide the OEM recommended service interval of the unit:
5.4 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
6 | General
6.1 | Has the product on offer or any of its components being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
6.2 | The unit on offer must be the latest technology.
6.2.1 | Initial Manufacturing year of Model:
| Theater Light, Ceiling-Mount- Accessories and Consumables | Complies Yes/No
RT2-13-04a | LED replacement

## Dialysis machine (CRRT)
RT2-13-05 | Dialysis Machine (CRRT) | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
1 | Product Offered:
1.1 | Make:
1.2 | Model:
1.3 | Initial Manufacturing year of Model:
1.4 | Radiation Control License number:
2 | SPECIFICATIONS
2.1 | The unit must be able to perform the following therapies:
2.1.1 | Slow Continuous Ultrafiltration (SCUF)
2.1.2 | Continuous Veno-Venous Hemofiltration (CVVH)
2.1.3 | Continuous Veno-Venous Haemodialysis (CVVHD)
2.1.4 | Continuous Veno-Venous Hemodiafiltration (CVVHDF)
2.1.5 | Hemoperfusion
2.1.6 | Plasmapheresis
2.2 | It must be mobile to be moved between wards by a single person.
2.3 | The unit must have a user-friendly communications interface screen
2.4 | All therapy parameters and information must be displayed on the screen
2.5 | All alarm conditions must be audible and visually displayed
2.6 | Remedy or guided remedy for alarm conditions must be displayed on screen
2.7 | The screen must be splash proof
2.8 | The machine must be able to retain data for at least 15 minutes In case of a power failure via back-up battery
2.9 | The machine must be able to switch between the different dialysis procedures without changing the disposables
2.1 | The unit must incorporate a built-in dose-controlled heparin infusion pump capable of continuously variable infusion rates, using standard disposable syringes
2.11 | State if the machine is capable of automated calcium-citrate regional anticoagulation
2.12 | Weighing System
2.12.1 | The unit must be volumetric controlled and include a weighing system to ensure a volumetric balance loop-back control to the speed of the ultrafiltration pump
2.12.2 | The weighing system must be able to handle weights of at least 10 kilograms
2.12.3 | The accuracy of the weighing system must be ± 30g (balance) and cyclical test during therapy
2.12.4 | Alarm protection must be audible and visual
2.12.5 | Machine must automatically go into "bypass” (stop balancing system) in alarm conditions related to the weighing system
2.12.6 | It must not be possible to silence the alarm interval for more than 60 seconds
2.13 | Blood Pump
2.13.1 | The Blood pump must be a roller pump (with backup manual operation) and digitally display flow
2.13.2 | Blood flow range must be at least: 50 - 300 ml/min
2.14 | Arterial Pressure:
2.14.1 | Arterial pressure must be electronically measured by means of a pressure transducer and digitally displayed on screen
2.14.2 | Measurement range must be approximately: -400 to ±600 mmHg
2.14.3 | General Technical And Safety Specifications
2.14.4 | The unit must comply with an acceptable international electrical safety standard such as IEC 60601-1 and 60601-1-2 for medical equipment
2.14.5 | Units being quoted for must be CE or SABS certified. Attach a copy of certification
2.14.6 | Hazardous Substance Act: If this type of equipment / apparatus appears on the schedule of Hazardous substances issued by the Directorate: Health Technology of the Department of Health, a license in terms of the Act on Hazardous Substances (Act 15/1973) must be submitted with the bid / quotation documents. The license must be registered under the bidders name or a letter of joint venture must be submitted by the licence holder where the licence is not in the name of the bidder. Bidders that neglect to submit a licence will not be considered
2.14.7 | The mains cable of the unit being quoted for must be the hospital grade type and it must be a minimum length of three (3) meters. NB the mains cable of the unit tendered for must be SABS colour coded.
2.14.8 | The equipment quoted for must be protected against electromagnetic interference
2.14.9 | Where equipment quoted for operates off 220 Volt, 50Hz ac supply bidders must ensure that the product quoted for is fitted with a 15 Amp SABS approved mains plug top, which is held together by two screws
2.14.10 | The bidder must be prepared to provide a unit for technical evaluation and clinical assessment on request, at short notice i.e. five working days.
2.14.11 | Will the technology change in the near future? 1 - 6 months, 6-12 months, 12-24 months, 24 - 48 months
2.14.12 | Must be the latest model - state date of initial manufacture of the model range offered
2.14.13 | Bidder's must state the lifespan of the equipment
2.14.14 | Spare parts must be guaranteed available for the specified life of the equipment, with a minimum of ten years.
2.14.15 | Can this equipment be upgraded? If so, give details and a breakdown of the costs on the bid / quotation schedule
2.14.16 | Supply a reference of machines currently in Government or Non Government institutions
2.14.17 | The bidder must guarantee that no additional equipment, parts or software, excluding consumables, will be required for the successful operation of the equipment quoted on in this RFP. A starter pack of all essential accessories must be supplied so that the unit can be put into immediate operation. The cost of the starter pack must be included in the bid price.
2.14.18 | Optional accessories must be quoted for separately on the quotation schedule.
2.14.19 | Equipment Guarantees and Service Specifications
2.14.20 | No part shall be second hand or refurbished
2.14.21 | Tender price to include delivery, installation, commissioning, two year warranty and training. Acceptance of the equipment will only take place after comissioning of the equipment.
2.14.22 | Bidders must supply a twenty four month guarantee against poor workmanship and latent defects and parts. This must be all inclusive and include,BUT NOT LIMITED TO, amongst others, ALL PARTS (including, where approppriate, Consumables, X-Ray tubes and other glassware), labour, travelling and accommodation.The two year warrantee must also include all quality check and quality assurance requirements, including all required calibrations. A fully comprehensive preventative maintenance, service and repair plan including all costs must be included in the two year warranty.Software updates and upgrades to be included. This two year guarantee will commence after formal acceptance and handover of the equipment and successful testing and acceptance by the Health Technology Services - Clinical Engineering..
2.14.23 | The successful bidder must arrange with both the respective Hospital / Institution and the Health Technology Services - Clinical engineering before commissioning the equipment at the respective hospital / institution.
2.14.24 | The successful bidder must arrange for an acceptance test of the equipment. A completed certified copy of this specification must be submitted with the equipment when the acceptance test(s) is to be undertaken. A copy of the acceptance test must be forwarded to the Health Technology Unit of the particular institution
2.14.25 | The recommended number of services per annum must be included during and up until the end of the guarantee period and all costs related to the provisions of such service/s will be for the bidders account. The bidder must state the number of services that will be provided during and up to the end of the guarantee period.
2.14.26 | Bidders must supply an all inclusive, fully comprehensive five year preventative maintenance, service and repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS( including, where approppriate, Consumables, X-Ray tubes and other glassware), labour, travelling, accommodation, service and maintenance. The five year maintenance plan must also include all quality check and quality assurance requirements, including all required calibrations.This contract will commence after the two year warranty period has expired. Software updates and upgrades to be included.
2.14.27 | Software changes to the equipment which are corrective in nature and initiated due to software errors, regulatory requirements or safety reasons, shall be delivered and installed at no charge for the life of the equipment
2.14.28 | Bidder's must have an established service and repair facility in the province of Gauteng. If the service is sub contracted to a local agent, a signed copy of the letter of appointment and acceptance must be submitted with this bid / quotation
2.14.29 | Supply the name, address and telephone number/s of the service department
2.14.30 | Technician(s) must be qualified and factory equivalent trained to deal with service, repair and calibration of the equipment quoted on. NB Proof of factory (or equivalent) training must be submitted with this bid / quotation offer.
2.14.31 | The bidder must state how many technicians they have permanently employed and their names and contact numbers must be listed
2.14.32 | State if the technician(s) are in the direct employ of the bidder
2.14.33 | If the equipment is taken away for repairs during the guarantee period, will a loan set be supplied for use by the institution?
2.14.34 | The up-time of the unit must be better than 98%, excluding scheduled preventative maintenance and software upgrades, measured on a quarterly basis.The percentage lower than 98% will be added to the warranty period. A sliding scale penalty clause will form part of the service contract. This will result in the maintenance payment being reduced by a pro rata amount that the up-time is less than 98%.
2.14.35 | Up-Time is defined as follows: 24/7; i.e. 365days times 24 hours = 8760 Hours. A down time of 2% relates to 175 hours per annum.
2.14.36 | User And Technical Training
2.14.37 | User training must be provided by the successful bidder in the operation of the unit at no extra cost to the final bid price to this department. Bidders must detail the training that would be offered and indicate who would offer the training
2.14.38 | The successful bidder must provide local training for GDoH technicians on the calibration, maintenance, service and repair of the quoted product. The quality and level of the training must be equivalent to the manufacturer's original factory training and any costs incurred to provide this training will be for the bidder's account. A certificate of competency must be issued on successful completion of the training. The training must be provided by the successful bidder within three (3) months from date of initial supply and delivery of the quoted equipment to the customer.
2.14.39 | Bidders must note that dedicated test equipment, spare parts and any special tooling required for the upkeep and maintenance of equipment quoted on must be made available to the Health Technology Services - Clinical Engineering.
2.14.40 | All the necessary calibration and maintenance software, where applicable, required to maintain and calibrate the equipment, must be supplied with the equipment to the Health Technology Unit(s) at no extra cost to the final bid price
2.14.41 | Complete original service / repair manual book or CD copies in English must be supplied. The manual should include the following information: fault finding guide, circuit diagrams / schematics, circuit descriptions and PCB layouts, calibration guide, part numbers and exploded diagram of mechanical parts / panels
2.14.3 | Alarm limits must be approximately -250 to +100 mmHg
2.14.4 | Acoustic alarm silence must not be longer than 60 seconds
2.15 | Pre-Filter Pressure
2.15.1 | Machine must be able to measure pre-filter entry pressure in order to assist the user to extend the life of the filter through timeous warning of build-up in filter
2.15.2 | The pressure must be electronically measured by means of a pressure transducer and digitally displayed on the screen
2.15.3 | The measurement range must be approximately: -400 to ±600 mmHg
2.15.4 | Alarm limits must be adjustable 0 to ± 400 mmHg.
2.15.5 | Adjustable range for Hemofiltration Therapies approximately: 0 to ±600 mmHg
2.15.6 | Acoustic alarm silence interval not greater than 60 seconds
2.16 | Venous Pressure Monitor
2.16.1 | Must be electronically measured by means of a pressure transducer and digitally displayed on the screen
2.16.2 | Measuring range must be approximately: -400 to ±600 mmHg
2.16.3 | Acoustic alarm silence interval must not exceed 60 seconds
2.17 | Safety Air Detector
2.17.1 | The Safety Air Detector must be an ultrasonic detector
2.17.2 | It must automatically do cyclical checks during therapy
2.17.3 | It must be able to detect the presence of micro-bubbles as well as micro-foam during therapy, or if the blood level drops below the sensor in the tubing
2.17.4 | The placement of the Safety Air Detector is such that it will prevent false alarms caused by the turbulence from the venous chamber
2.17.5 | State the alarm sensitivity
2.17.6 | The safety clamp must form part of the unit to close the re- Infusion blood line to the patient in case of alarm caused by the Safety Air detector or in all "Blood pump stop" conditions
2.18 | Fluid Circuit (Air Detector):
2.18.1 | The unit must have air detection system in the substitution/dialysate solution flow path or another type of system in place to warn for empty bag
2.18.2 | In alarm conditions the unit must go automatically into "Bypass" mode and Substitution and Ultrafiltration pumps must stop
2.18.3 | Alarms must be both audible and visual
2.18.4 | Acoustic alarm silence interval shall not exceed 60 seconds
2.19 | Substitution/Dialysate Pump:
2.19.1 | Must have a 2 roller (or more) peristaltic pump and flow rate must be digitally displayed on screen
2.19.2 | The pump should automatically calibrate in preparation phase
2.19.3 | Flow range for Continuous Therapy:
2.19.4 | Approximate minimal flow range of 0 to 6000 ml/hr (CVVH)
2.19.5 | Approximate minimal flow range of 0 to 12 000 ml/hr (CVVHD, CVVHDF)
2.2 | Filter Inlet Pressure
2.20.1 | This pressure must monitor for any obstruction in the dialysate flow into the filter
2.20.2 | Must be electronically measured by means of a pressure transducer and digitally displayed
2.20.3 | Acoustic alarm silence Interval shall not exceed 60 seconds
2.21 | Warmer
2.21.1 | The unit must include a warming system to prevent hypothermia in patients. Please state if it is a blood or dialysate warmer
2.21.2 | State the adjustable temperature range
2.21.3 | Please state the temperature alarm limits and duration of condition required prior to alarm activation. (Example: alarm will activate after 10 seconds at 41oC or 1 sec at 45oC).
2.21.4 | Warmer should automatically switch-off in high temperature alarm conditions but allow therapy to continue
2.21.5 | Alarms to be both audible and visually displayed on the screen
2.21.6 | Acoustic alarm silence interval should not exceed 60 seconds
2.22 | Blood Leak Detector
2.22.1 | A photometric red sensitive blood leak detector should form part of the filter outlet system of the unit
2.22.2 | State sensitivity
2.22.3 | Unit should automatically perform zero calibration before every therapy session
2.22.4 | Thereafter cyclical self-testing during therapy
2.22.5 | Acoustic alarm silence not longer than 60 seconds
2.23 | Filter Outlet Pressure
2.23.1 | The unit should incorporate a detection system in the ultrafiltration circuit in order to prevent obstruction of the ultrafiltration circuit and/or full waste bag, thus preventing possible spillage and environment contamination
2.23.2 | State the system used
2.23.3 | Acoustic alarm silence interval should not exceed 60 seconds
2.24 | Ultrafiltration Pump
2.24.1 | Must be a two or more roller peristaltic pump and flow rate digitally displayed on screen
2.24.2 | Exact pump speed must be regulated: Please state means whereby this is achieved and how verification thereof by alternative means is done
2.24.3 | State UF flow ranges for the treatment
3 | Standard Accessories & Consumables
3.1 | Dialyser filter (Artificial kidney), medium, Adult (1.5 - 2.1m2)
3.2 | Dialyser filter (Artificial kidney), medium, Peads/Children (1.1 - 1.3m2)
3.3 | Dialyser filter (Artificial kidney), High Flux, Adult (1.5 - 2.1m2)
3.4 | Dialyser filter (Artificial kidney), High Flux, Peads/Children (1.1 - 1.3m2)
3.5 | Dialyser filter (Artificial kidney), High Performance, Adult (1.5 - 2.1m2)
3.6 | Dialyser filter (Artificial kidney), High Performance, Peads/Children (1.1 - 1.3m2)
3.7 | Dialysis Solution (Acatate & Bicarb cartridge), Mixing ratio of 1:34, 5L
3.8 | Dialysis Solution (Acatate & Bicarb cartridge), Mixing ratio of 1:35, 5L
3.9 | Dialysis Solution (Acatate & Bicarb cartridge), Normal Calcium, 1.5mmol/l
3.10 | Dialysis Solution (Acatate & Bicarb cartridge), Low Calcium, 1.25mmol/l
3.11 | Dialysis Solution (Acatate & Bicarb cartridge), Low Glucose, 5.5mmol/l
3.12 | Dialysis Solution (Acatate & Bicarb cartridge), High Potasium, 3.5mmol/l
3.13 | Dry Bicarb 650g cartridge
3.14 | Dry Bicarb 760g cartridge
3.15 | Sterilising Agent, 5L
3.16 | Decalcifying Agent, 5L
3.17 | Test strip, Chlorine (pk of 50)
3.18 | Test strip, Water Hardness (pk of 50)
3.19 | Endotoxins filter
3.20 | AV Fistula Needle (ultra-sharp three bevel, turntable colour-coded, packed with 30cm tubing, Latex and BPA free), 15G
3.21 | AV Fistula Needle (ultra-sharp three bevel, turntable colour-coded, packed with 30cm tubing, Latex and BPA free), 16G
3.22 | AV Fistula Needle (ultra-sharp three bevel, turntable colour-coded, packed with 30cm tubing, Latex and BPA free), 17G
3.23 | Catheter lock cap
3.24 | IV Blood line, Adult (Pk of 50)
3.25 | IV Blood line, Peads/Child (Pk of 50)
3.26 | K-shield (sterile single-use iodine cap), Pk of 50
3.27 | K-shield (sterile single-use iodine cap), Pk of 100
3.28 | Prepacked two (2) sterile syringe set, prefilled with trisodium citrate locking solution at 30% and 46.7% concentrate, 2.5ml
4 | Conformity Compliance (Please attach certificate)
4.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
4.2 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
4.2.1 | Please provide unique ref number of the ISO 13485 ceriticate:
4.3 | Model on offer must comply with EU Medical Device Regulation (EU) 2017/745 or equvalent from SAHPRA recognized jurisdictions . Attach a copy of certification
4.3.1 | Please provide unique ref number of the ceriticate:
5 | COMPREHENSIVE MAINTENANCE
5.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) must be included in the bid price
5.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
5.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
5.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
6 | General
6.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
6.2 | The unit on offer must be the latest technology.
| Dialysis Machine (CRRT), Accessoried & Consumables | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
RT2-13-05a | Dialyser filter (Artificial kidney), medium, Adult (1.5 - 2.1m2)
RT2-13-05b | Dialyser filter (Artificial kidney), medium, Peads/Children (1.1 - 1.3m2)
RT2-13-05c | Dialyser filter (Artificial kidney), High Flux, Adult (1.5 - 2.1m2)
RT2-13-05d | Dialyser filter (Artificial kidney), High Flux, Peads/Children (1.1 - 1.3m2)
RT2-13-05e | Dialyser filter (Artificial kidney), High Performance, Adult (1.5 - 2.1m2)
RT2-13-05f | Dialyser filter (Artificial kidney), High Performance, Peads/Children (1.1 - 1.3m2)
RT2-13-05g | Dialysis Solution (Acatate & Bicarb cartridge), Mixing ratio of 1:34, 5L
RT2-13-05h | Dialysis Solution (Acatate & Bicarb cartridge), Mixing ratio of 1:35, 5L
RT2-13-05i | Dialysis Solution (Acatate & Bicarb cartridge), Normal Calcium, 1.5mmol/l
RT2-13-05j | Dialysis Solution (Acatate & Bicarb cartridge), Low Calcium, 1.25mmol/l
RT2-13-05k | Dialysis Solution (Acatate & Bicarb cartridge), Low Glucose, 5.5mmol/l
RT2-13-05l | Dialysis Solution (Acatate & Bicarb cartridge), High Potasium, 3.5mmol/l
RT2-13-05m | Dry Bicarb 650g cartridge
RT2-13-05n | Dry Bicarb 760g cartridge
RT2-13-05o | Sterilising Agent, 5L
RT2-13-05p | Decalcifying Agent, 5L
RT2-13-05q | Test strip, Chlorine (pk of 50)
RT2-13-05r | Test strip, Water Hardness (pk of 50)
RT2-13-05s | Endotoxins filter
RT2-13-05t | AV Fistula Needle (ultra-sharp three bevel, turntable colour-coded, packed with 30cm tubing, Latex and BPA free), 15G
RT2-13-05u | AV Fistula Needle (ultra-sharp three bevel, turntable colour-coded, packed with 30cm tubing, Latex and BPA free), 16G
RT2-13-05v | AV Fistula Needle (ultra-sharp three bevel, turntable colour-coded, packed with 30cm tubing, Latex and BPA free), 17G
RT2-13-05x | Catheter lock cap
RT2-13-05y | IV Blood line, Adult (Pk of 50)
RT2-13-05z | IV Blood line, Peads/Child (Pk of 50)
RT2-13-05aa | K-shield (sterile single-use iodine cap), Pk of 50
RT2-13-05ab | K-shield (sterile single-use iodine cap), Pk of 100
RT2-13-05ac | Prepacked two (2) sterile syringe set, prefilled with trisodium citrate locking solution at 30% and 46.7% concentrate, 2.5ml

## Intermittent Hemodialysis
RT2-13-06 | Intermittent Hemodialysis Unit | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
1 | Product Offered:
1.1 | Make:
1.2 | Model:
1.3 | Initial Manufacturing year of Model:
1.4 | Radiation Control License number:
2 | SPECIFICATIONS
2.1 | The unit must be able to perform the following therapies:
2.1.1 | Intermittent Hemodialysis
2.2 | It must be mobile to be moved between wards by a single person
2.3 | The unit must have a user-friendly communications interface screen
2.4 | All therapy parameters and information must be displayed on the screen
2.5 | All alarm conditions must be audible and visually displayed
2.6 | Remedy or guided remedy for alarm conditions must be displayed on screen
2.7 | The screen must be splash proof
2.8 | The machine must be able to retain data for at least 15 minutes in case of a power failure via back-up battery
2.9 | The unit must incorporate a built-in dose-controlled heparin infusion pump capable of continuously variable infusion rates, using standard disposable syringes
2.1 | Blood Pump
2.10.1 | The Blood pump must be a roller pump (with backup manual operation) and digitally display flow
2.10.2 | Blood flow range must be at least: 50 - 600 ml/min
2.11 | Arterial Pressure:
2.11.1 | Arterial pressure must be electronically measured by means of a pressure transducer and digitally displayed on screen
2.11.2 | Measurement range must be approximately: -400 to ±600 mmHg
2.11.3 | General Technical And Safety Specifications
2.11.4 | The unit must comply with an acceptable international electrical safety standard such as IEC 60601-1 and 60601-1-2 for medical equipment
2.11.5 | Units being quoted for must be CE or SABS certified. Attach a copy of certification
2.11.6 | Hazardous Substance Act: If this type of equipment / apparatus appears on the schedule of Hazardous substances issued by the Directorate: Health Technology of the Department of Health, a license in terms of the Act on Hazardous Substances (Act 15/1973) must be submitted with the bid / quotation documents. The license must be registered under the bidders name or a letter of joint venture must be submitted by the licence holder where the licence is not in the name of the bidder. Bidders that neglect to submit a licence will not be considered
2.11.7 | The mains cable of the unit being quoted for must be the hospital grade type and it must be a minimum length of three (3) meters. NB the mains cable of the unit tendered for must be SABS colour coded.
2.11.8 | The equipment quoted for must be protected against electromagnetic interference
2.11.9 | Where equipment quoted for operates off 220 Volt, 50Hz ac supply bidders must ensure that the product quoted for is fitted with a 15 Amp SABS approved mains plug top, which is held together by two screws
2.11.10 | The bidder must be prepared to provide a unit for technical evaluation and clinical assessment on request, at short notice i.e. five working days.
2.11.11 | Will the technology change in the near future? 1 - 6 months, 6-12 months, 12-24 months, 24 - 48 months
2.11.12 | Must be the latest model - state date of initial manufacture of the model range offered
2.11.13 | Bidder's must state the lifespan of the equipment
2.11.14 | Spare parts must be guaranteed available for the specified life of the equipment, with a minimum of ten years.
2.11.15 | Can this equipment be upgraded? If so, give details and a breakdown of the costs on the bid / quotation schedule
2.11.16 | Supply a reference of machines currently in Government or Non Government institutions
2.11.17 | The bidder must guarantee that no additional equipment, parts or software, excluding consumables, will be required for the successful operation of the equipment quoted on in this RFP. A starter pack of all essential accessories must be supplied so that the unit can be put into immediate operation. The cost of the starter pack must be included in the bid price.
2.11.18 | Optional accessories must be quoted for separately on the quotation schedule.
2.11.19 | Equipment Guarantees and Service Specifications
2.11.20 | No part shall be second hand or refurbished
2.11.21 | Tender price to include delivery, installation, commissioning, two year warranty and training. Acceptance of the equipment will only take place after comissioning of the equipment.
2.11.22 | Bidders must supply a twenty four month guarantee against poor workmanship and latent defects and parts. This must be all inclusive and include,BUT NOT LIMITED TO, amongst others, ALL PARTS (including, where approppriate, Consumables, X-Ray tubes and other glassware), labour, travelling and accommodation.The two year warrantee must also include all quality check and quality assurance requirements, including all required calibrations. A fully comprehensive preventative maintenance, service and repair plan including all costs must be included in the two year warranty.Software updates and upgrades to be included. This two year guarantee will commence after formal acceptance and handover of the equipment and successful testing and acceptance by the Health Technology Services - Clinical Engineering..
2.11.23 | The successful bidder must arrange with both the respective Hospital / Institution and the Health Technology Services - Clinical engineering before commissioning the equipment at the respective hospital / institution.
2.11.24 | The successful bidder must arrange for an acceptance test of the equipment. A completed certified copy of this specification must be submitted with the equipment when the acceptance test(s) is to be undertaken. A copy of the acceptance test must be forwarded to the Health Technology Unit of the particular institution
2.11.25 | The recommended number of services per annum must be included during and up until the end of the guarantee period and all costs related to the provisions of such service/s will be for the bidders account. The bidder must state the number of services that will be provided during and up to the end of the guarantee period.
2.11.26 | Bidders must supply an all inclusive, fully comprehensive five year preventative maintenance, service and repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS( including, where approppriate, Consumables, X-Ray tubes and other glassware), labour, travelling, accommodation, service and maintenance. The five year maintenance plan must also include all quality check and quality assurance requirements, including all required calibrations.This contract will commence after the two year warranty period has expired. Software updates and upgrades to be included.
2.11.27 | Software changes to the equipment which are corrective in nature and initiated due to software errors, regulatory requirements or safety reasons, shall be delivered and installed at no charge for the life of the equipment
2.11.28 | Bidder's must have an established service and repair facility in the province of Gauteng. If the service is sub contracted to a local agent, a signed copy of the letter of appointment and acceptance must be submitted with this bid / quotation
2.11.29 | Supply the name, address and telephone number/s of the service department
2.11.30 | Technician(s) must be qualified and factory equivalent trained to deal with service, repair and calibration of the equipment quoted on. NB Proof of factory (or equivalent) training must be submitted with this bid / quotation offer.
2.11.31 | The bidder must state how many technicians they have permanently employed and their names and contact numbers must be listed
2.11.32 | State if the technician(s) are in the direct employ of the bidder
2.11.33 | If the equipment is taken away for repairs during the guarantee period, will a loan set be supplied for use by the institution?
2.11.34 | The up-time of the unit must be better than 98%, excluding scheduled preventative maintenance and software upgrades, measured on a quarterly basis.The percentage lower than 98% will be added to the warranty period. A sliding scale penalty clause will form part of the service contract. This will result in the maintenance payment being reduced by a pro rata amount that the up-time is less than 98%.
2.11.35 | Up-Time is defined as follows: 24/7; i.e. 365days times 24 hours = 8760 Hours. A down time of 2% relates to 175 hours per annum.
2.11.36 | User And Technical Training
2.11.37 | User training must be provided by the successful bidder in the operation of the unit at no extra cost to the final bid price to this department. Bidders must detail the training that would be offered and indicate who would offer the training
2.11.38 | The successful bidder must provide local training for GDoH technicians on the calibration, maintenance, service and repair of the quoted product. The quality and level of the training must be equivalent to the manufacturer's original factory training and any costs incurred to provide this training will be for the bidder's account. A certificate of competency must be issued on successful completion of the training. The training must be provided by the successful bidder within three (3) months from date of initial supply and delivery of the quoted equipment to the customer.
2.11.39 | Bidders must note that dedicated test equipment, spare parts and any special tooling required for the upkeep and maintenance of equipment quoted on must be made available to the Health Technology Services - Clinical Engineering.
2.11.40 | All the necessary calibration and maintenance software, where applicable, required to maintain and calibrate the equipment, must be supplied with the equipment to the Health Technology Unit(s) at no extra cost to the final bid price
2.11.41 | Complete original service / repair manual book or CD copies in English must be supplied. The manual should include the following information: fault finding guide, circuit diagrams / schematics, circuit descriptions and PCB layouts, calibration guide, part numbers and exploded diagram of mechanical parts / panels
2.11.3 | Alarm limits must be approximately -250 to +100 mmHg
2.11.4 | Acoustic alarm silence must not be longer than 60 seconds
2.12 | Pre-Filter Pressure
2.12.1 | Machine must be able to measure pre-filter entry pressure in order to assist the user to extend the life of the filter through timeous warning of build-up in filter
2.12.2 | The pressure must be electronically measured by means of a pressure transducer and digitally displayed on the screen
2.12.3 | The measurement range must be approximately: -400 to ±600 mmHg
2.12.4 | Alarm limits must be adjustable 0 to ± 400 mmHg.
2.12.5 | Adjustable range for Hemofiltration Therapies approximately: 0 to ±600 mmHg
2.12.6 | Acoustic alarm silence interval not greater than 60 seconds
2.13 | Venous Pressure Monitor
2.13.1 | Must be electronically measured by means of a pressure transducer and digitally displayed on the screen
2.13.2 | Measuring range must be approximately: -400 to ±600 mmHg
2.13.3 | Acoustic alarm silence interval must not exceed 60 seconds
2.14 | Safety Air Detector
2.14.1 | The Safety Air Detector must be an ultrasonic detector
2.14.2 | It must automatically do cyclical checks during therapy
2.14.3 | It must be able to detect the presence of micro-bubbles as well as micro-foam during therapy, or if the blood level drops below the sensor in the tubing
2.14.4 | The placement of the Safety Air Detector is such that it will prevent false alarms caused by the turbulence from the venous chamber
2.14.5 | State the alarm sensitivity
2.14.6 | The safety clamp must form part of the unit to close the re- Infusion blood line to the patient in case of alarm caused by the Safety Air detector or in all "Blood pump stop" conditions
2.15 | Fluid Circuit (Air Detector):
2.15.1 | The unit must have air detection system in the substitution/dialysate solution flow path or another type of system in place to warn for empty bag
2.15.2 | In alarm conditions the unit must go automatically into "Bypass" mode and Substitution and Ultrafiltration pumps must stop
2.15.3 | Alarms must be both audible and visual
2.15.4 | Acoustic alarm silence interval shall not exceed 60 seconds
2.16 | Filter Inlet Pressure
2.16.1 | This pressure must monitor for any obstruction in the dialysate flow into the filter
2.16.2 | Must be electronically measured by means of a pressure transducer and digitally displayed
2.16.3 | Acoustic alarm silence Interval shall not exceed 60 seconds
2.17 | Warmer
2.17.1 | The unit must include a warming system to prevent hypothermia in patients. Please state if it is a blood or dialysate warmer
2.17.2 | State the adjustable temperature range
2.17.3 | Please state the temperature alarm limits and duration of condition required prior to alarm activation. (Example: alarm will activate after 10 seconds at 41oC or 1 sec at 45oC).
2.17.4 | Warmer should automatically switch-off in high temperature alarm conditions but allow therapy to continue
2.17.5 | Alarms to be both audible and visually displayed on the screen
2.17.6 | Acoustic alarm silence interval should not exceed 60 seconds
2.18 | Blood Leak Detector
2.18.1 | A photometric red sensitive blood leak detector should form part of the filter outlet system of the unit
2.18.2 | State sensitivity
2.18.3 | Unit should automatically perform zero calibration before every therapy session
2.18.4 | Thereafter cyclical self-testing during therapy
2.18.5 | Acoustic alarm silence not longer than 60 seconds
2.19 | Ultrafiltration Pump
2.19.1 | Must be a two or more roller peristaltic pump and flow rate digitally displayed on screen
2.19.2 | Exact pump speed must be regulated: Please state means whereby this is achieved and how verification thereof by alternative means is done
2.19.2 | UF flow ranges of 0-2000ml/hr
2.19.3 | Ability to perform UF and sodium profiling
2.2 | Dialysate Section
2.20.1 | Electronic and hydraulic compartments must be isolated to eliminate the possibility of electrical shock
2.20.2 | Acetate and bicarbonate dialysate should be possible
2.20.3 | Continuous dialysate mixing should be monitored continuously in both acetate and bicarbonate modes by a temperature compensated conductivity measuring system before supplied to the dialuser
2.20.4 | Conductivity reading should be visually displayed with visual and audio alarms when deviations of +- 0.5ms/cm occur
2.20.5 | Deviations should lead to automatic bypass of the dialyser
2.20.6 | Integral final water filter should be present
2.20.7 | De-earation hould be adequate
2.21 | Sterilization
2.21.1 | Must be a chemical cycle
2.21.2 | The various times for the cleaning programs must be adjustable
2.21.3 | Equipped with decalcification program
2.21.4 | A complete description of cleaning and sterilization programs must be provided
3 | Standard Accessories & Consumables
3.1 | Dialyser filter (Artificial kidney), medium, Adult (1.5 - 2.1m2)
3.2 | Dialyser filter (Artificial kidney), medium, Peads/Children (1.1 - 1.3m2)
3.3 | Dialyser filter (Artificial kidney), High Flux, Adult (1.5 - 2.1m2)
3.4 | Dialyser filter (Artificial kidney), High Flux, Peads/Children (1.1 - 1.3m2)
3.5 | Dialyser filter (Artificial kidney), High Performance, Adult (1.5 - 2.1m2)
3.6 | Dialyser filter (Artificial kidney), High Performance, Peads/Children (1.1 - 1.3m2)
3.7 | Dialysis Solution (Acatate & Bicarb cartridge), Mixing ratio of 1:34, 5L
3.8 | Dialysis Solution (Acatate & Bicarb cartridge), Mixing ratio of 1:35, 5L
3.9 | Dialysis Solution (Acatate & Bicarb cartridge), Normal Calcium, 1.5mmol/l
3.10 | Dialysis Solution (Acatate & Bicarb cartridge), Low Calcium, 1.25mmol/l
3.11 | Dialysis Solution (Acatate & Bicarb cartridge), Low Glucose, 5.5mmol/l
3.12 | Dialysis Solution (Acatate & Bicarb cartridge), High Potasium, 3.5mmol/l
3.13 | Dry Bicarb 650g cartridge
3.14 | Dry Bicarb 760g cartridge
3.15 | Sterilising Agent, 5L
3.16 | Decalcifying Agent, 5L
3.17 | Test strip, Chlorine (pk of 50)
3.18 | Test strip, Water Hardness (pk of 50)
3.19 | Endotoxins filter
3.20 | AV Fistula Needle (ultra-sharp three bevel, turntable colour-coded, packed with 30cm tubing, Latex and BPA free), 15G
3.21 | AV Fistula Needle (ultra-sharp three bevel, turntable colour-coded, packed with 30cm tubing, Latex and BPA free), 16G
3.22 | AV Fistula Needle (ultra-sharp three bevel, turntable colour-coded, packed with 30cm tubing, Latex and BPA free), 17G
3.23 | Catheter lock cap
3.24 | IV Blood line, Adult (Pk of 50)
3.25 | IV Blood line, Peads/Child (Pk of 50)
3.26 | K-shield (sterile single-use iodine cap), Pk of 50
3.27 | K-shield (sterile single-use iodine cap), Pk of 100
3.28 | Prepacked two (2) sterile syringe set, prefilled with trisodium citrate locking solution at 30% and 46.7% concentrate, 2.5ml
4 | Conformity Compliance (Please attach certificate)
4.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
4.2 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
4.2.1 | Please provide unique ref number of the ISO 13485 ceriticate:
4.3 | Model on offer must comply with EU Medical Device Regulation (EU) 2017/745 or equvalent from SAHPRA recognized jurisdictions . Attach a copy of certification
4.3.1 | Please provide unique ref number of the ceriticate:
5 | COMPREHENSIVE MAINTENANCE
5.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) must be included in the bid price
5.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
5.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
5.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
6 | General
6.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
6.2 | The unit on offer must be the latest technology.
| Intermittent Hemodialysis Unit, Accessoried & Consumables | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
RT2-13-06a | Dialyser filter (Artificial kidney), medium, Adult (1.5 - 2.1m2)
RT2-13-06b | Dialyser filter (Artificial kidney), medium, Peads/Children (1.1 - 1.3m2)
RT2-13-06c | Dialyser filter (Artificial kidney), High Flux, Adult (1.5 - 2.1m2)
RT2-13-06d | Dialyser filter (Artificial kidney), High Flux, Peads/Children (1.1 - 1.3m2)
RT2-13-06e | Dialyser filter (Artificial kidney), High Performance, Adult (1.5 - 2.1m2)
RT2-13-06f | Dialyser filter (Artificial kidney), High Performance, Peads/Children (1.1 - 1.3m2)
RT2-13-06g | Dialysis Solution (Acatate & Bicarb cartridge), Mixing ratio of 1:34, 5L
RT2-13-06h | Dialysis Solution (Acatate & Bicarb cartridge), Mixing ratio of 1:35, 5L
RT2-13-06i | Dialysis Solution (Acatate & Bicarb cartridge), Normal Calcium, 1.5mmol/l
RT2-13-06j | Dialysis Solution (Acatate & Bicarb cartridge), Low Calcium, 1.25mmol/l
RT2-13-06k | Dialysis Solution (Acatate & Bicarb cartridge), Low Glucose, 5.5mmol/l
RT2-13-06l | Dialysis Solution (Acatate & Bicarb cartridge), High Potasium, 3.5mmol/l
RT2-13-06m | Dry Bicarb 650g cartridge
RT2-13-06n | Dry Bicarb 760g cartridge
RT2-13-06o | Sterilising Agent, 5L
RT2-13-06p | Decalcifying Agent, 5L
RT2-13-06q | Test strip, Chlorine (pk of 50)
RT2-13-06r | Test strip, Water Hardness (pk of 50)
RT2-13-06s | Endotoxins filter
RT2-13-06t | AV Fistula Needle (ultra-sharp three bevel, turntable colour-coded, packed with 30cm tubing, Latex and BPA free), 15G
RT2-13-06u | AV Fistula Needle (ultra-sharp three bevel, turntable colour-coded, packed with 30cm tubing, Latex and BPA free), 16G
RT2-13-06v | AV Fistula Needle (ultra-sharp three bevel, turntable colour-coded, packed with 30cm tubing, Latex and BPA free), 17G
RT2-13-06x | Catheter lock cap
RT2-13-06y | IV Blood line, Adult (Pk of 50)
RT2-13-06z | IV Blood line, Peads/Child (Pk of 50)
RT2-13-06aa | K-shield (sterile single-use iodine cap), Pk of 50
RT2-13-06ab | K-shield (sterile single-use iodine cap), Pk of 100
RT2-13-06ac | Prepacked two (2) sterile syringe set, prefilled with trisodium citrate locking solution at 30% and 46.7% concentrate, 2.5ml

## RO, Mobile
RT2-13-07 | Mobile Reverse Osmosis | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
1 | Product Offered:
1.1 | Make:
1.2 | Model:
1.3 | Initial Manufacturing year of Model:
2 | SPECIFICATIONS
2.1 | Capacity: Able to supply purified water to two HD machines simultaneously
2.2 | Lightweight medical grade material housing with wheels that is durable and easy to move around within the hospital.
2.3 | The unit must be user-friendly
2.4 | Easy to clean case and easily transportable
2.5 | Readily visible pressure gauges, which include pipeline water pressure and membrane pressure
2.6 | The following minimum rejection rate should be obtained
2.7 | Total dissolved salts >96%
2.8 | Bacteria and pyrogens > 99%
2.9 | Particles 100%
2.1 | Minimum pure water outlet of 2litre/min @ 10 bar
2.11 | Ultra Violet light sterilizer
2.12 | Particle and carbon filters
3 | Conformity Compliance (Please attach certificate)
3.1 | The unit must comply with an acceptable AAMI Standards, attached certification
3.2 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
3.2.1 | Please provide unique ref number of the ISO 13485 ceriticate:
3.3 | Model on offer must comply with EU Medical Device Regulation (EU) 2017/745 or equvalent from SAHPRA recognized jurisdictions . Attach a copy of certification
3.3.1 | Please provide unique ref number of the ceriticate:
4 | COMPREHENSIVE MAINTENANCE
4.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) must be included in the bid price
4.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
4.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
4.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
4.3 | The Following must be performed and fully covered during 2-year Warranty and 5-year Comprehensive Maintenance;
4.3.1 | Micron Filter must be replaced during servicing.
4.3.2 | Cleaning of Sand filter.
4.3.3 | Cleaning of Activated Carbon Filter
4.3.4 | Anti scaling dosing chemical (As per requirement)
4.3.5 | Raw water pump (Checking/servicing, if required)
4.3.6 | R.O. Membrane (Checking/ Replacement, if required)
4.3.7 | High pressure pump (Checking/servicing, if required)
4.3.8 | Checking of flow rate of membrane & TDS on site
4.3.9 | Endotoxin test of RO will be bidders responsibility.The bidder submit the test reports/documentations to user department.
4.4 | Bidder must provide Water Analysis test Six monthly as per AAMI standard or any relevant standard during 2-year Warranty & 5-year Comprehensive Maintenance Period by bidder. Also submit the test reports/documentations to user department.
4.5 | Product water testing will be responsibility of supplier.
5 | General
5.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
5.2 | The unit on offer must be the latest technology.

## RO Plant
RT2-13-08 | R. O. Plant | Complies Yes/No | Provide your answers in this Column. You are advised to be transparent with deviation.
1 | DESCRIPTION
1.1 | A device/system used to treat water intended for use in the delivery of haemodialysis and related therapies
2 | Product Offered:
2.1 | Make:
2.2 | Model:
2.3 | Initial Manufacturing year of Model:
3 | SPECIFICATIONS (10% tolerance)
3.1 | PLANT REQUIREMENTS
3.1.1 | This specification is for the design, supply, delivery, installation, testing and commissioning and maintenance of a complete water purification system for ultra-pure water for haemodialysis applications with the following main criteria and equipment:
3.1.2 | Using varying water quality from the Local Municipalities generate water for dialysis applications conforming to the requirements of ANSI/AAMI RD62:2006 as amended.
3.1.3 | The water flow required is for simultaneous operation of 10 dialysis machines each with a normal flow requirement of 500ml/minute and 800ml/minute during the cleaning and rinsing cycle. The cleaning and rinsing cycle for each machine generally occur twice per day for 45 minutes per cycle per machine.
3.1.4 | 10 000liter raw water storage tanks with automatic level control, level indication, overflow, drain and alarm.
3.1.5 | Boost pressure pump taking water from the raw water tank with adequate pressure to pump water through the pre-filter system.
3.1.6 | Duplex depth filtration system complete with automatic back wash system and controls.
3.1.7 | Particle filtration unit with duplex 10µm absolute stainless steel (grade 316) mesh or wedge wire filter cartridges.
3.1.8 | Duplex ion-exchange water softener completes with automatic controls and brine tank.
3.1.9 | Duplex carbon adsorption filters each with an empty bed contact time (EBCT) of more than 6 minutes giving a total EBCT of at least 12 minutes at the maximum water flow rate through the filters.
3.1.10 | Duplex 1 µm absolute fine filter cartridge system.
3.1.11 | Complete Reverse osmosis filtration system.
3.1.12 | One set of endotoxin and bacteria retentive filters after the reverse osmosis plant before entering the renal unit.
3.1.13 | Chemical and heat sterilization systems.
3.1.14 | All control, monitoring, recording and alarm systems with additional remote alarm panel.
3.1.15 | All pumps, motors and electrical controls to ensure adequate flow and pressures through the water treatment and circulation system.
3.1.16 | All electrical work including the provision of a supply cable from the nearest distribution board located 15 meters from the location of the new filtration installation.
3.1.17 | Cumulative water flow meters. One measuring the total raw water flow and one measuring reverse osmosis reject flow.
3.1.18 | Sterilize the complete water system and storage tanks and conduct the necessary test to confirm sterilization. a) Bacterial count b) Endotoxin levels
3.1.19 | Ultrasonic cleaning bath adequately sized and capable of simultaneously cleaning at least two 10µ stainless steel mesh pre-filter elements as used for the water purification system.
3.1.20 | Water analysis instruments for daily/weekly tests to be performed by the user. The minimum testing equipment must be an ISO certified TDS meter and a free and total chlorine meter. Both instruments must be temperature compensated and complete with calibration equipment and reagent sets adequate for at least one calendar year. Bidders to recommend additional instruments separately.
3.1.21 | The specification clause 10.29 calls for the bunded area to be cleaned out and sealed with an epoxy sealer. It is the bidder’s choice to either supply a fully packaged plant complete with fully supported stainless steel drip tray and place the equipment adjacent to the existing position or use the existing position. In both cases must the existing bunded area and area below packaged plant be cleaned and treated and made water proof as part of the scope of supply.
3.1.22 | Maintenance and operating manuals, parts lists, manufacturers data sheets, maintenance schedules and list of recommended spares for all equipment. In triplicate
3.1.23 | Testing and commissioning of the installation and training of staff in the use, care and maintenance of the equipment.
3.1.24 | Full maintenance during the 24-month guarantee period. All chemicals, filter elements and tests required for 12 months must be included in the initial price. The tests must include twelve (12) full water analysis tests on the raw water supply and purified water supply to the dialysis machines to determine the chemical, bacterial and endotoxin levels as specified by ANSI/AAMI RD62:2006 a). All other items and requirements, modifications to existing piping and tank, removal of existing equipment, whether specifically mentioned or not, for a complete, functional and safe water purification system complying with all the relevant codes and specifications.
3.1.25 | A complete first aid kit suitable for the treatment associated with chemical injury associated with the purification and sterilization processes must be provided. a). All safety notices and safety equipment. b). Removal of the existing system and handing over to the hospital maintenance section is included in this contract. c). A dust free equipment room/ cubicle must be constructed in the plant room to house all the equipment of the water purifying system. d). It is a condition of bid that all bidders submit with their bid document the details and contact persons of at least three similar water treatment projects completed successfully by them. e) If a bidder intends using a PLC (programmable logic controller) to control the water treatment plant, the software package to program, download and communicate with the PLC as well as the cable to connect the PLC to a PC or note book computer, must form part of the scope of supply for this bid.
3.1.26 | The normal operating hours of the renal unit, to which the proposed new water treatment plant supply water for dialysis, is from 07h00 to 19h00 Monday to Saturday inclusive. During emergency situations, water for dialysis may be required at times not falling within the mentioned time
3.1.27 | The allowable floor loading where the water treatment plant is to be installed is according to the National building regulations as indicated in SANS 0400 which in turn refers to SANS 01060: The general procedures and loadings to be adopted in the design of buildings. In summary, the imposed floor loading must be limited to 1.5kN/m2
3.2 | The equipment and system supplied must be complete and functional and must be as a minimum according to this specification and the latest revisions of the following standards and specifications:
3.2.1.1 | Output water quality should match AAMI (Association for the Advancement of Medical Instrumentation) standards for Hemodialysis water (Electrolytes and heavy metals as stipulated in AAMI guidelines; Bacteria <100cfu/ml and endotoxin level <0.25).
3.2.1.2 | SANS 10400: The Application of the National Building Regulations
3.2.1.3 | SANS 10142: The Wiring of Premises Part 1: Low Voltage Installations
3.2.1.4 | Regulations governing Hazardous Biological Agents R1390
3.2.1.5 | Department of Public Works specification PW371: Specification of Materials and Methods to be used and is available from National Public Works website
3.2.1.6 | Occupational Health and Safety Act and Regulations
3.2.1.7 | All other relevant specifications, standards and documents whether referenced in the above documents or not.
3.2.1.8 | This specification is supplemental to the above and any conflicting information must be referred to the engineer for clarification.
3.2.2 | The water purification system must be able to function irrespective of the site, supply water conditions, plant room ambient temperatures and weather conditions.
3.2.3 | The existing renal unit must remain in operation while the new system is installed. The change over from the existing system to the new system must be done during a period from 14h00 on a Saturday to 19h00 on a Sunday evening. The changeover times and dates must be arranged with the head of the renal unit in conjunction with the hospital maintenance superintendent.
3.2.4 | It is compulsory that bidders visit the site. Bidders must also check all data and satisfy themselves to the nature and extent of the work to be done and make allowance for any items omitted from the specification but obviously required and necessary for the completion of the work and proper functioning of the installation
3.2.5 | The raw water feeding the proposed new water treatment plant is drawn from the hospital water supply. The hospital is in general supplied with water from the Municipality water system. As the water quality vary as a result of pipe breakages, variable purification processes and alternative water sources, the system offered must be capable of producing the required water under conditions of variable supply water quality without causing damage to the membranes.
3.3 | Site:
3.3.1 | The equipment offered is intended for use on a daily basis. The equipment must be installed in a plant room. The general site temperatures in which the water treatment system must be able to operate successfully are as follows:
3.3.1.1 | Summer maximum dry bulb temperature: 45° C
3.3.1.2 | Winter minimum dry bulb temperature: -3° C
3.3.2 | The water purification system must however be able to function irrespective of the site and weather conditions.
3.3.3 | The new water purification system will be replacing an existing system housed in the same location as the proposed new system. Provision for moving the old one should be arranged.
3.4.1 | Duplex type cartridge filters from 316 stainless steel mesh with absolute filtration down to 10µ.
3.4.1.1 | The filter housings must be capable of holding the stainless-steel mesh filters as well as throwaway fibre filter cartridges.
3.4.1.2 | The filter cartridges must be sized in conjunction with the depth filter system to ensure that cleaning of the mesh filter and/or discarding of the fibre filters are limited to not less than six months.
3.4.1.3 | It must be possible to remove one filter cartridge for cleaning without disrupting the water flow through the system.
3.4.1.4 | If screw on filter housings are offered, dedicated spanners to loosen and tighten the housings must be included. The spanner lever length must be limited to reduce the risk of over tightening and causing damage to the filter housing.
3.4.1.5 | The mesh cartridges must be cleanable and must be able to withstand ultrasonic cleaning.
3.4.2 | Duplex water softener completes with one common brine tank, automatic regeneration controls, safety interlocks and water monitoring system to prevent high concentration of salt entering the supply water system during regeneration.
3.4.2.1 | One section of the duplex water softener must be capable of lowering the hardness to below 10 PPM by removing as a minimum the Ca and Mg to protect the reverse osmosis membranes.
3.4.2.2 | Each softener vessel size must be adequate to limit the regeneration cycle to once per day and only to take place after hours when the dialysis machines do not call for treated water.
3.4.2.3 | An automatic flushing and monitoring system must be in place to prevent brine from passing into the water system towards the RO equipment.
3.4.2.4 | The regeneration controls must be interlocked with a master real time clock to ensure regeneration cannot take place when system is in use.
3.4.2.5 | The brine tank must be adequately sized to carry at least two 50kg bags of salt with free space to facilitate water flow for regeneration.
3.4.2.6 | The brine tank must be complete with secure lid with closures.
3.4.3 | Duplex activated carbon filters complete with sampling points must be provided.
3.4.3.1 | The EBCT (Extended bed contact time) of each carbon tank must be at least 6 (six) minutes during maximum water flow rate through the system.
3.4.3.2 | If the carbon tanks are too large to achieve the specified EBCT, four carbon tanks may be installed, two parallel pairs in series and each parallel pair providing a minimum EBCT of 6 minutes at maximum flow.
3.4.3.3 | Balancing valves and rota-meter flow meters must be installed to ensure equal flow through each carbon tank.
3.4.3.4 | The two tanks and/or two pairs of tanks must be easily interchanged and must be rechargeable.
3.4.3.5 | Regenerated carbon in the carbon tanks must not be allowed.
3.4.3.6 | Warning labels to this effect must also be fixed to the carbon tanks.
3.4.3.7 | The iodine number of the activated carbon must be such that with the specified EBCT, the chloramines in the water must be reduced to below 0.1mg/L with an activated carbon life in the first carbon tank of at least 6 months.
3.4.3.8 | The activated carbon ash content must be a minimum to reduce the risk of algae growth due to ferric oxide leaching into the water.
3.4.3.9 | All carbon tanks must be fitted with cleanable mesh filters to prevent carbon dust from entering the water system.
3.4.3.10 | The back-wash system, if fitted, must prevent any chlorinated water entering the system on the RO plant side of the carbon filters.
3.4.4 | A duplex fine filter system capable of removing all particles larger than 1µ must be installed between the activated carbon filters and reverse osmosis unit.
3.4.4.1 | The filter capacity must be matched to the particle load so that the filter life is more than 6 months during continuous operation at full flow.
3.4.4.2 | The general requirements as specified for the 10µ mesh filter and housings must also apply to the fine filter.
3.4.5 | Reverse osmosis unit complete with:
3.4.5.1 | Membranes with membrane vessels
3.4.5.2 | Automatic and manual shut off
3.4.5.3 | Pressure regulating valves and pumps
3.4.5.4 | Control and alarm system
3.4.5.5 | Automatic and manual flow and pressure monitoring and indication systems capable of producing purified water
3.4.5.6 | The membranes must be sized to limit replacement to once per year.
3.4.5.7 | The membrane vessels must be piped in such a manner that individual membranes can be removed and replaced without interrupting the filtered water flow to the dialysis machines.
3.4.5.8 | The bypass water must be limited to a maximum of 40% of the total flow.
3.4.5.9 | The bypass flow must be piped to a 500-liter storage tank complete with chlorinator and pump to pump the bypass water to the hospital water storage tanks.
3.4.5.10 | The tank and bypass water inlet must be separated with an air break/back flow preventer.
3.4.5.11 | The pump head must be a minimum of 30m.
3.4.5.12 | The conductivity of the water supplied to the RO plant as well as the product water must be automatically monitored and integrated with the alarm system.
3.4.5.13 | The conductivity monitoring systems must include an alarm for the % rejection rate and product water conductivity levels suitable for dialysis use.
3.4.5.14 | The conductivity and rejection rate alarm levels must be established by doing the purity tests
3.4.5.15 | The control system and piping must ensure that a continuous flow is maintained through the membranes.
3.4.5.16 | If the flow should fall away, a no flow alarm must be activated.
3.4.5.17 | The circulating pump flow rate must be such that the flow velocity in the pipes never go below 1 m/s.
3.4.6.1 | Ultra filters for bacterial protection must be provided at the outlet of the RO plant and the supply to the dialysis machines.
3.4.6.2 | The ultra-filters must remove from the water supply all endotoxins to a level 80% below the action level of 1 EU/ml and remove all the bacteria to a level 80% below the action level of 25 CFU.
3.4.7 | The existing water treatment system includes a stainless-steel purified water storage tank and 25mm diameter stainless steel pipe system for circulating water to the dialysis machines and back to the plant room.
3.4.7.1 | Bidders must allow for incorporating the storage tank and circulating water pipe system into the new filtration system and supply and install the circulating pump and ultra-filters between the tank and supply to the dialysis machines
3.4.7.2 | The remainder of the existing water filtration system must be removed and handed to the hospital for disposal via the accepted disposal route.
3.4.7.3 | Any alterations, modifications, inspections, improvements etc. required to the tank and circulating pipe system must form part of this contract.
3.4.7.4 | Bidders must allow for all such work to enable them to accept responsibility for the complete system.
3.4.7.5 | The circulating pump flow rate must be such that the minimum pipe velocity never goes below 1 m/s in the return line even when maximum draw off to the dialysis machines occur.
3.5 | MATERIALS USED
3.5.1 | All materials used in the water treatment system must not react in any way with the water and/or chemicals used.
3.5.2 | The materials must be capable of withstanding the high temperatures that is required for heat sterilization.
3.5.3 | As a minimum, the return water temperature during heat sterilization must be 80°C.
3.5.4 | If metallic materials are in contact with the water, stainless steel equal or better than 316L must be used.
3.5.4.1 | All welds must be done according to AWS D1.1 with proven weld procedures for stainless steel.
3.5.4.2 | Welding must be done by a welder qualified for the relevant weld procedure.
3.5.4.3 | Welds must be x-rayed to verify no pinholes and/or crevices exist where bacteria can attach.
3.5.5 | All fittings, connections, bends etc. must be pharmacological quality and must be such that no crevices are formed where bacteria can attach.
3.5.6 | Surface roughness of all surfaces in contact with the water must also be such that bacteria growth is minimized.
3.5.7 | Bidders must allow to survey all existing piping used for the water distribution to the dialysis machines, fittings and storage vessel to be re-used. Any recommended rectification of existing system must be done on a proven cost basis.
3.5.8 | All pipe and other brackets, bolts and nuts, drip trays, filter brackets and fittings not in contact with the water must be from at least 304 stainless steel. Where salt is used, brackets etc. must be from composite materials suitable for the application.
3.5.9 | Contact between dissimilar metals where galvanic corrosion can take place is not allowed. The degree of dissimilarity must be such that the electrode potential between the two metals is less than 100 mV.
3.6 | PUMPS
3.6.1 | All pumps must be with 316L stainless steel shafts, non-ferrous impellers and corrosion free bodies.
3.6.2 | Pump materials must comply with MATERIAL USED section in this document.
3.6.3 | Pump motors must be sized to be non-overloading, standard frame size, minimum IP rating of IP54 and 50Hz 3 phase TEFC squirrel cage induction motors to suit the electrical supply available on site.
3.6.4 | Motor insulation must be class F or better with a class B temperature rise.
3.6.5 | Motors larger than 3.5kW, where inverters are not used, must be connected to a star/delta starter with mechanical interlock between contactors.
3.6.6 | Mechanical seals with stainless steel components and silicon carbide or similar wear surfaces must be used. All ‘O’ rings and rubber seals must be suitable for the water and chemicals used in the process.
3.6.7 | All bolts and nuts must be 316 stainless steel.
3.6.8 | Dosing pumps must be of the metered type and compatible with the relevant chemicals.
3.6.9 | All pumps must be mounted and secured with corrosion resistant brackets.
3.6.10 | Cables and wiring must be run on corrosion resistant cable baskets.
3.6.11 | All motors must be protected with motor protection units and must be with lockable local isolators.
3.7 | PIPES AND FITTINGS
3.7.1 | All piping, fittings, valves, instrumentation etc. must comply with the requirements of the MATERIALS USED section in this document.
3.7.2 | The spacing of pipe supports for non-metallic piping where heat sterilization is used must be not more than 650mm.
3.7.3 | For normal operating temperatures, pipe supports for all pipes must be not more than 800mm.
3.7.4 | Pipe supports must be such that free movement caused by expansion and contraction is not impeded.
3.7.4.1 | All supports must also be such that pipe damage cannot occur due to continuous movement in the brackets.
3.7.4.2 | Anchor points must be strategically placed to allow free movement due to expansion and contraction but adequately secure the piping to prevent movement caused by normal flow and change in flow direction.
3.7.5 | All filter housings, control valves, instrumentation, sample taps etc. must be separately bracketed and must not be supported from the pipe work.
3.7.6 | All valves must be of the ball type with replaceable seals and union couplings
3.7.7.1 | All pipes must be installed vertically and horizontally within the accuracy of a normal builder’s level. Horizontal piping must include at least a 2° fall to a drain point.
3.7.7.2 | Neat manifold type brackets must be used in place of ad-hoc brackets.
3.7.8.1 | Piping and fittings must be arranged that no “dead leg” can be formed.
3.7.8.2 | Union fittings must be installed to allow easy removal of filters, instruments etc., without disturbing the pipe work.
3.7.8.3 | Excessive fittings and change of direction must be avoided.
3.7.9.1 | All pipe brackets must be corrosion proof. All brackets, bolts, nuts, screws etc. must be from stainless steel.
3.7.9.2 | As stainless steel is not compatible with chlorine, the chlorine dosing pump and fittings must be mounted with corrosion proof brackets from composite materials suitable for the application.
3.7.10.1 | Adequate allowance must be made for pipe expansion.
3.7.10.2 | It must not be possible to enclose any pipe section to the extent that pipe damage can occur when the temperature increases cause thermal expansion of the water and the consequential pressure increase exceed the operating pressure of the pipes or attachments.
3.7.11 | The pressure rating of all pipes and fittings must be compatible with the operating pressure and temperature but must as a minimum be class 16.
3.7.12.1 | All pipe systems must be hydraulically pressure tested to a pressure 1.5 X normal operating pressure.
3.7.12.2 | The pressure must be held for 24 hours while the pump is un coupled from the system.
3.7.12.3 | Pressure gauges used for the pressure test must be complete with valid test certificates
3.7.12.4 | The pressure must be monitored and manually recorded while the system is tested over a 24-hour period.
3.8 | ELECTRICAL WORK
3.8.1.1 | All electrical work must conform to SANS 10142 and a certificate of conformance (COC) must be issued for the installation.
3.8.1.2 | The COC must include the additional requirements for medical areas.
3.8.2.1 | All cables, horizontal and vertical, must be secured to corrosion proof cable baskets.
3.8.2.2 | Instrumentation and power cables must be on separate cable baskets.
3.8.2.3 | Cable baskets must be sized such that at least 20% free space is available.
3.8.2.4 | Cables must not be bunched and must be laid side by side and individually fixed to the cable baskets with cable ties with minimum width of 5mm for power cables and 3mm for control cables.
3.8.3.1 | The supply, installation and connecting up of the main 3 phase 380V four wire cable and external earth conductor from a DB situated near the location of the new water treatment plant also forms part of this contract.
3.8.3.2 | Bidders to allow 15 meters of cabling from the nearest DB to the water purification plant control panel.
3.8.3.3 | Connection to the existing DB must be coordinated with the hospital maintenance superintendent.
3.8.4 | All equipment must be capable of operating under the varying conditions of national grid and emergency generator.
3.8.5 | All equipment and controls must be capable of operating error free in the presence of voltage surges, electromagnetic fields and other anomalies such as radio frequencies.
3.8.6 | All equipment and materials used must comply with the relevant SANS code or equivalent BS or IEC codes.
3.8.7 | All cabling for voltages higher than 24V must be 1000V SWA PVC multi-core with copper conductors. Cable sizes must be determined according to the requirements of SANS 10142.
3.8.8 | Control cables must be insulated screened cable correctly rated for the application
3.8.9 | All equipment, frames etc. must be complete with clearly identified earth studs and must be bonded to the earth system with separate earth cables of the anti-theft type. Mechanical connections must not be relied upon to provide earth continuity.
3.8.10 | Earth conductors must be bonded together where loops are connected to equipment. Bonding must be such that if a specific item of equipment is disconnected, the earth loop cannot be broken.
3.8.11 | All power cables must be made off with plated bronze glands on a corrosion proof metallic gland plate that is bonded to the panel earth bar. To avoid mechanical stress, cable tails must be looped before connected to terminals
3.8.12 | All wire ends must be fitted with color coded cable ends of the correct size applicable to the wire size
3.8.13 | All cables and wires must be clearly marked at both connected ends as well as externally on the cable baskets. The identification markings must be the same as shown on the control panel drawings and must comply with the requirements of IEC 61343-2
3.8.14 | Wiring colors must conform to the requirements of SANS 10142. Control wiring colors must be consistent and must be indicated on the drawings.
3.8.15 | No joints must be allowed in cables and control wiring. All cables penetrating equipment and/or panels must pass through and be secured with glands as specified for the power cables.
3.8.16 | The electric control panel must be manufactured from galvanized steel and powder coated electric orange. All instrumentation must be visible on the face of the control panel and must be fitted with mounting bezels
3.8.17 | All push buttons, indicator lamps and warning lamps must be appropriately colored as per IEC 60204-1.
3.8.18 | The control panel for the water treatment plant must be fitted with surge protectors. The current capacity of the surge protectors must be rated at least 50% higher than the rating of the circuit breaker protecting the control panel cable
3.8.19.1 | All equipment, instrumentation, indicator lights, switches etc. must be clearly marked with engraved labels of which the lettering is at least 6mm high.
3.8.19.2 | Engraved lettering must be black on a white background
3.8.19.3 | All labels must be the same length and height and must be fixed to the equipment/panel face with minimum two M3 stainless steel round heat screws and hexagon nuts
3.8.19.4 | Screws not to extend more than 5mm beyond the nut. The nuts must be kept in place with a low strength thread locking compound.
3.9 | CONTROL AND MONITORING SYSTEM
3.9.1 | The water treatment plant must be complete with a control and monitoring system
3.9.1.1 | Sampling points must also be provided before and after each item of equipment and filter
3.9.1.2 | The sampling points must be from stainless steel and must be of such a nature that easy sterilization can be done prior to taking a sample
3.9.1.3 | The common drip tray below all equipment must be such that any spillage during the sampling process is taken to the drain point without accumulating in the drip tray.
3.9.1.4 | All materials must conform to the requirements as specified under the MATERIALS USED section.
3.9.2 | As a minimum, the following control and monitoring systems must be included:
3.9.2.1 | High- and low-level indicators with extreme alarms for the raw water tank.
3.9.2.2 | Booster pump flow and pressure
3.9.2.3 | Pressure before and after the depth filters i.e. ΔP
3.9.2.4 | Total hardness after water softener and pressure drop across water softeners
3.9.2.5 | Pressure drop across each carbon tank, flow rate through each tank and total chlorine content after each tank.
3.9.2.6 | Pressure drop across all filter cartridge
3.9.2.7 | Pressure drop and flow across RO membranes, product and bypass water flow rate and TDS of product water
3.9.2.8 | Temperature and flow rate through water distribution piping
3.9.2.9 | Cumulative raw water feed flow and RO plant bypass water flow.
3.9.2.10 | Regeneration and back wash controls on depth filters, water softener and carbon filters.
3.9.3 | All monitoring equipment must be temperature compensated and come complete with test certificates confirming accuracy and reliability. The accuracy of monitoring and alarm equipment must be appropriate for the application.
3.9.4 | System critical alarms must also be indicating on a clearly marked remote panel in the renal unit supplied and installed as part of this contract.
3.9.5.1 | Pressure gauges must be from stainless steel with a minimum dial size of 100mm.
3.9.5.2 | Normal operating range must be marked with a blue line, maximum and minimum operating pressure must be marked with a red line.
3.9.5.3 | All alarms and monitoring systems must comply with the requirement standard
3.10 | STERILISATION SYSTEM
3.10.1.1 | The water treatment system must be complete with thermal and chemical sterilization units.
3.10.1.2 | All materials must be compatible with both sterilization systems. Bidders to include measures and systems to prevent sterilization fluid from entering the dialysis machines.
3.10.2 | This system must include complete rinsing cycles. Clear indication and warning signals must indicate sterilization is in progress.
3.10.3.1 | The thermal sterilization system must include all electrical heaters, safety devices, controls and temperature monitoring and indication systems.
3.10.3.2 | The temperature of the return water from the circulating system and associated time must be used as the indication for completion of the sterilization process.
3.10.3.3 | Biders to allow for full sterilizing and testing of purified water on a monthly basis for twelve months during the guarantee period
3.11 | EQUIPMENT ROOM/ CUBICLE
3.11.1 | The rooms/ cubicles will be constructed in the existing plant rooms.
3.11.2 | The rooms will be dust free and the temperature will be mechanically controlled.
3.11.3 | The rooms/ cubicles must have a bunded area and the drainage will be connected to the existing drainage.
3.11 | DRAWINGS AND DOCUMENTS
3.11.1 | The successful bidder must provide three copies of full maintenance and operating manuals both in file with plastic sleeves and in electronic format. The minimum contents of the manuals must be:
3.11.1.1 | Schematic diagrams, control diagrams and circuits, PLC program listing with comments, PID’s and as build drawings.
3.11.1.2 | Basic start up, shut down and running procedures.
3.11.1.3 | Fault check list in flow diagram format
3.11.1.4 | Water safety plan including basic test procedures to ensure safe water quality and logging system recording readings taken on a daily basis.
3.11.1.5 | Routine maintenance plan and monitoring to be performed by hospital staff.
3.11.1.6 | Flow rates and pressure settings to be monitored and recorded in log book and on electronic program if possible.
3.11.1.7 | List of critical spares and recommended spares to be kept on site complete with names and contact details of suppliers.
3.11.1.8 | Material Safety Data sheets for chemicals used in process.
3.11.1.9 | Emergency containment procedure in the event of a chemical spill.
3.11.1.10 | Emergency treatment in the event of accidental contact with the chemicals.
3.11.1.11 | Manufacturers data sheets and equipment brochures.
3.11.1.12 | Hydrostatic test certificates for piping and vessels.
3.11.1.13 | Monthly water test certificates confirming conformance to ANSI/AAMI RD62:2006
3.11.1.14 | Any other information regarded as pertinent to the installation and operation of the system.
4 | Conformity Compliance (Please attach certificate)
4.1 | The unit must comply with AAMI Standards, attach certification
4.2 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
4.2.1 | Please provide unique ref number of the ISO 13485 ceriticate:
4.3 | Model on offer must comply with EU Medical Device Regulation (EU) 2017/745 or equvalent from SAHPRA recognized jurisdictions . Attach a copy of certification
4.3.1 | Please provide unique ref number of the ceriticate:
5 | COMPREHENSIVE MAINTENANCE
5.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) must be included in the bid price
5.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
5.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be addressed as per the SCC document
5.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
5.3 | The Following must be performed and fully covered during 2-year Warranty and 5-year Comprehensive Maintenance;
5.3.1 | Micron Filter must be replaced during servicing.
5.3.2 | Cleaning of Sand filter.
5.3.3 | Cleaning of Activated Carbon Filter
5.3.4 | Anti scaling dosing chemical (As per requirement)
5.3.5 | Raw water pump (Checking/servicing, if required)
5.3.6 | R.O. Membrane (Checking/ Replacement, if required)
5.3.7 | High pressure pump (Checking/servicing, if required)
5.3.8 | Checking of flow rate of membrane & TDS on site
5.3.9 | Endotoxin test of RO will be bidders responsibility upon installation of the RO.The bidder must submit the test reports/documentations to user department.
5.4 | Bidder must provide Water Analysis test Six monthly as per AAMI standard or any relevant standard during 2-year Warranty & 5-year Comprehensive Maintenance Period by bidder. Also submit the test reports/documentations to user department.
5.5 | Product water testing will be the responsibility of the supplier.
6 | General
6.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
6.2 | The unit on offer must be the latest technology.

## Nurse Call System 8 Bed
RT2-13-09 | Nurse Call System, 8-bed Package | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
1 | DESCRIPTION
1.1 | A communication and alert system to enable patients to promptly summon assistance from nursing staff
1.2 | The nurse call system must be provided for 8-bedded ward with complete components.
2 | Product Offered:
2.1 | Make:
2.2 | Model:
2.3 | Initial Manufacturing year of Model:
3 | SPECIFICATION
3.1 | System Components:
3.1.1 | Call Points: Devices installed at patient bedsides, bathrooms, and other strategic locations, allowing patients or staff to trigger an alert.
3.1.2 | Indicator Panels: Located at the nurses' station, these panels display the location of the call (room number, bed number) and the type of call (emergency, routine).
3.1.3 | Overhead Lights: Visual indicators installed above patient doors or in hallways, which light up to show where the call was made.
3.1.4 | Central Controller/Server: Manages the system's data flow, integrates with hospital information systems, and stores call logs for reporting.
3.1.5 | Intercom/Voice Communication Units: Allow two-way communication between the patient and the nursing staff directly from the bedside unit. (Optional)
3.2 | Call Types:
3.2.1 | Routine Call: Non-emergency requests such as needing water or assistance with personal care.
3.2.2 | Emergency Call: For urgent situations requiring immediate attention, such as medical distress.
3.2.3 | Code Blue Call: Specifically for cardiac arrest or other critical conditions, triggering immediate response protocols.
3.2.4 | Staff Assist Call: Allows staff to request assistance from other team members.
3.3 | Operational Features:
3.3.1 | Alert Prioritization: Differentiates between emergency and non-emergency calls, ensuring high-priority alerts are addressed first.
3.3.2 | Call Escalation: If a call is not acknowledged within a set time, it escalates to other staff members or supervisory personnel.
3.3.3 | Real-time Monitoring: Continuous tracking and display of call status, including active, acknowledged, or cleared calls.
3.3.4 | Logging and Reporting: Records all call data, including timestamps, response times, and types of calls, for performance analysis and compliance.
3.3.5 | User Interface:
3.3.5.1 | Nurse Station Display: Intuitive touchscreen or computer-based interface showing real-time status of all calls.
3.3.5.2 | Patient Interface: Simple, user-friendly interface for patients, with large, clearly labeled buttons for various call types.
3.4 | Technical requirements:
3.4.1 | Power Supply: Uninterruptible power supply (UPS) integration to ensure system functionality during power outages.
3.4.2 | Connectivity: Wired (Ethernet, PoE) connectivity for various system components.
3.4.3 | Optionally capable of supporting expansion to accommodate more call points, rooms, and integration with additional systems.
3.5 | Security and Privacy:
3.5.1 | Data Encryption: Ensures secure communication between devices and protects patient data.
3.5.2 | Access Control: Restricts access to system controls and sensitive information to authorized personnel only.
3.5.3 | Audit Trails: Maintains logs of system access and changes for security and compliance purposes.
3.6 | Emergency Backup:
3.6.1 | Redundancy: Backup systems for critical components to ensure functionality during failures.
3.6.2 | Emergency Override: Manual override options for emergency situations, ensuring patient safety at all times.
3.7 | Customization Options:
3.7.1 | Alert Tones and Visual Indicators: Customizable sounds and light patterns for different call types.
3.7.2 | Language Support: Multilingual support for patient and staff interfaces.
3.7.3 | Integration Options: Customizable APIs to interface with various hospital systems (e.g., fire alarms, CCTV).
3.8 | The unit must be in full operation order following commissioning.
4 | Standard Accessories
4.1 | Patient Remotes/call buttons
4.2 | Nurse Console/Central Monitoring
5 | Conformity Compliance (Please attach certificate)
5.1 | OEM must comply and certified on ISO 9001 or ISO 13485 quality standards, attach proof of compliance.
5.1.1 | Please provide unique ref number of the ISO ceriticate:
5.2 | Model on offer must comply with EU Medical Device Regulation (EU) 2017/745 or equvalent from SAHPRA recognized jurisdictions . Attach a copy of certification
5.2.1 | Please provide unique ref number of the ceriticate:
5.3 | Meets healthcare standards such as HIPAA (Health Insurance Portability and Accountability Act), UL1069, and other local regulatory requirements.
6 | COMPREHENSIVE MAINTENANCE
6.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) should be quoted for, as per the SCC document
6.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
6.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
6.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
6.3 | The bidder to provide the OEM recommended service interval of the unit:
6.4 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
7 | General
7.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
7.2 | The unit on offer must be the latest technology.
7.2.1 | Initial Manufacturing year of Model:

## Nurse Call System 16 Bed
RT2-13-010 | Nurse Call System, 16-bed Package | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
1 | DESCRIPTION
1.1 | A communication and alert system to enable patients to promptly summon assistance from nursing staff
1.2 | The nurse call system must be provided for 16-bedded ward with complete components.
2 | Product Offered:
2.1 | Make:
2.2 | Model:
2.3 | Initial Manufacturing year of Model:
3 | SPECIFICATION
3.1 | System Components:
3.1.1 | Call Points: Devices installed at patient bedsides, bathrooms, and other strategic locations, allowing patients or staff to trigger an alert.
3.1.2 | Indicator Panels: Located at the nurses' station, these panels display the location of the call (room number, bed number) and the type of call (emergency, routine).
3.1.3 | Overhead Lights: Visual indicators installed above patient doors or in hallways, which light up to show where the call was made.
3.1.4 | Central Controller/Server: Manages the system's data flow, integrates with hospital information systems, and stores call logs for reporting.
3.1.5 | Intercom/Voice Communication Units: Allow two-way communication between the patient and the nursing staff directly from the bedside unit. (Optional)
3.2 | Call Types:
3.2.1 | Routine Call: Non-emergency requests such as needing water or assistance with personal care.
3.2.2 | Emergency Call: For urgent situations requiring immediate attention, such as medical distress.
3.2.3 | Code Blue Call: Specifically for cardiac arrest or other critical conditions, triggering immediate response protocols.
3.2.4 | Staff Assist Call: Allows staff to request assistance from other team members.
3.3 | Operational Features:
3.3.1 | Alert Prioritization: Differentiates between emergency and non-emergency calls, ensuring high-priority alerts are addressed first.
3.3.2 | Call Escalation: If a call is not acknowledged within a set time, it escalates to other staff members or supervisory personnel.
3.3.3 | Real-time Monitoring: Continuous tracking and display of call status, including active, acknowledged, or cleared calls.
3.3.4 | Logging and Reporting: Records all call data, including timestamps, response times, and types of calls, for performance analysis and compliance.
3.3.5 | User Interface:
3.3.5.1 | Nurse Station Display: Intuitive touchscreen or computer-based interface showing real-time status of all calls.
3.3.5.2 | Patient Interface: Simple, user-friendly interface for patients, with large, clearly labeled buttons for various call types.
3.4 | Technical requirements:
3.4.1 | Power Supply: Uninterruptible power supply (UPS) integration to ensure system functionality during power outages.
3.4.2 | Connectivity: Wired (Ethernet, PoE) connectivity for various system components.
3.4.3 | Optionally capable of supporting expansion to accommodate more call points, rooms, and integration with additional systems.
3.5 | Security and Privacy:
3.5.1 | Data Encryption: Ensures secure communication between devices and protects patient data.
3.5.2 | Access Control: Restricts access to system controls and sensitive information to authorized personnel only.
3.5.3 | Audit Trails: Maintains logs of system access and changes for security and compliance purposes.
3.6 | Emergency Backup:
3.6.1 | Redundancy: Backup systems for critical components to ensure functionality during failures.
3.6.2 | Emergency Override: Manual override options for emergency situations, ensuring patient safety at all times.
3.7 | Customization Options:
3.7.1 | Alert Tones and Visual Indicators: Customizable sounds and light patterns for different call types.
3.7.2 | Language Support: Multilingual support for patient and staff interfaces.
3.7.3 | Integration Options: Customizable APIs to interface with various hospital systems (e.g., fire alarms, CCTV).
3.8 | The unit must be in full operation order following commissioning.
4 | Standard Accessories
4.1 | Patient Remotes/call buttons
4.2 | Nurse Console/Central Monitoring
5 | Conformity Compliance (Please attach certificate)
5.1 | OEM must comply and certified on ISO 9001 or ISO 13485 quality standards, attach proof of compliance.
5.1.1 | Please provide unique ref number of the ISO ceriticate:
5.2 | Model on offer must comply with EU Medical Device Regulation (EU) 2017/745 or equvalent from SAHPRA recognized jurisdictions . Attach a copy of certification
5.2.1 | Please provide unique ref number of the ceriticate:
5.3 | Meets healthcare standards such as HIPAA (Health Insurance Portability and Accountability Act), UL1069, and other local regulatory requirements.
6 | COMPREHENSIVE MAINTENANCE
6.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) should be quoted for, as per the SCC document
6.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
6.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
6.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
6.3 | The bidder to provide the OEM recommended service interval of the unit:
6.4 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
7 | General
7.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
7.2 | The unit on offer must be the latest technology.
7.2.1 | Initial Manufacturing year of Model:

## Nurse Call System 24 Bed
RT2-13-011 | Nurse Call System, 24-bed Package | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
1 | DESCRIPTION
1.1 | A communication and alert system to enable patients to promptly summon assistance from nursing staff
1.2 | The nurse call system must be provided for 24-bedded ward with complete components.
2 | Product Offered:
2.1 | Make:
2.2 | Model:
2.3 | Initial Manufacturing year of Model:
3 | SPECIFICATION
3.1 | System Components:
3.1.1 | Call Points: Devices installed at patient bedsides, bathrooms, and other strategic locations, allowing patients or staff to trigger an alert.
3.1.2 | Indicator Panels: Located at the nurses' station, these panels display the location of the call (room number, bed number) and the type of call (emergency, routine).
3.1.3 | Overhead Lights: Visual indicators installed above patient doors or in hallways, which light up to show where the call was made.
3.1.4 | Central Controller/Server: Manages the system's data flow, integrates with hospital information systems, and stores call logs for reporting.
3.1.5 | Intercom/Voice Communication Units: Allow two-way communication between the patient and the nursing staff directly from the bedside unit. (Optional)
3.2 | Call Types:
3.2.1 | Routine Call: Non-emergency requests such as needing water or assistance with personal care.
3.2.2 | Emergency Call: For urgent situations requiring immediate attention, such as medical distress.
3.2.3 | Code Blue Call: Specifically for cardiac arrest or other critical conditions, triggering immediate response protocols.
3.2.4 | Staff Assist Call: Allows staff to request assistance from other team members.
3.3 | Operational Features:
3.3.1 | Alert Prioritization: Differentiates between emergency and non-emergency calls, ensuring high-priority alerts are addressed first.
3.3.2 | Call Escalation: If a call is not acknowledged within a set time, it escalates to other staff members or supervisory personnel.
3.3.3 | Real-time Monitoring: Continuous tracking and display of call status, including active, acknowledged, or cleared calls.
3.3.4 | Logging and Reporting: Records all call data, including timestamps, response times, and types of calls, for performance analysis and compliance.
3.3.5 | User Interface:
3.3.5.1 | Nurse Station Display: Intuitive touchscreen or computer-based interface showing real-time status of all calls.
3.3.5.2 | Patient Interface: Simple, user-friendly interface for patients, with large, clearly labeled buttons for various call types.
3.4 | Technical requirements:
3.4.1 | Power Supply: Uninterruptible power supply (UPS) integration to ensure system functionality during power outages.
3.4.2 | Connectivity: Wired (Ethernet, PoE) connectivity for various system components.
3.4.3 | Optionally capable of supporting expansion to accommodate more call points, rooms, and integration with additional systems.
3.5 | Security and Privacy:
3.5.1 | Data Encryption: Ensures secure communication between devices and protects patient data.
3.5.2 | Access Control: Restricts access to system controls and sensitive information to authorized personnel only.
3.5.3 | Audit Trails: Maintains logs of system access and changes for security and compliance purposes.
3.6 | Emergency Backup:
3.6.1 | Redundancy: Backup systems for critical components to ensure functionality during failures.
3.6.2 | Emergency Override: Manual override options for emergency situations, ensuring patient safety at all times.
3.7 | Customization Options:
3.7.1 | Alert Tones and Visual Indicators: Customizable sounds and light patterns for different call types.
3.7.2 | Language Support: Multilingual support for patient and staff interfaces.
3.7.3 | Integration Options: Customizable APIs to interface with various hospital systems (e.g., fire alarms, CCTV).
3.8 | The unit must be in full operation order following commissioning.
4 | Standard Accessories
4.1 | Patient Remotes/call buttons
4.2 | Nurse Console/Central Monitoring
5 | Conformity Compliance (Please attach certificate)
5.1 | OEM must comply and certified on ISO 9001 or ISO 13485 quality standards, attach proof of compliance.
5.1.1 | Please provide unique ref number of the ISO ceriticate:
5.2 | Model on offer must comply with EU Medical Device Regulation (EU) 2017/745 or equvalent from SAHPRA recognized jurisdictions . Attach a copy of certification
5.2.1 | Please provide unique ref number of the ceriticate:
5.3 | Meets healthcare standards such as HIPAA (Health Insurance Portability and Accountability Act), UL1069, and other local regulatory requirements.
6 | COMPREHENSIVE MAINTENANCE
6.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) should be quoted for, as per the SCC document
6.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
6.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
6.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
6.3 | The bidder to provide the OEM recommended service interval of the unit:
6.4 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
7 | General
7.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
7.2 | The unit on offer must be the latest technology.
7.2.1 | Initial Manufacturing year of Model: