## Category-5 Monitors & General
GENERAL MEDICAL EQUIPMENT
# | Item number | Item Description
1 | RT2-05-001 | Bloodgas Analyzer, Single-use
2 | RT2-05-002 | Bloodgas Analyzer, Cartridge Type
3 | RT2-05-003 | Bloodgas Analyzer, Solutions
4 | RT2-05-004 | Defibrilator, ECG&Pacing
5 | RT2-05-005 | Defibrilator, ECG&SpO2
6 | RT2-05-006 | ECG Machine, 12-lead
7 | RT2-05-007 | Ecmo Machine
8 | RT2-05-008 | Examination Light, LED, Mobile
9 | RT2-05-009 | Examination Light, LED, Mobile with battery
10 | RT2-05-010 | Examination Light, LED, Wall-mount
11 | RT2-05-011 | Infusion Pump, Neo
12 | RT2-05-012 | Infusion Pump, Paeds-Adult
13 | RT2-05-013 | Monitor, Basic
14 | RT2-05-014 | Monitor, ICU
15 | RT2-05-015 | Monitor, ICU with Transport Module
16 | RT2-05-016 | Central Monitoring, 8-bed package
17 | RT2-05-017 | Central Monitoring Station, ICU
18 | RT2-05-018 | Vital Signs Monitor with oximetery & Temp
19 | RT2-05-019 | Pulse Oximeter, Tabletop
20 | RT2-05-020 | Transcutaneous CO2 Monitor, Non-Invasive
21 | RT2-05-021 | Pulse Wave Monitor, Non-Invasive
22 | RT2-05-022 | Theater Light, Mobile
23 | RT2-05-023 | Dermatoscope
24 | RT2-05-024 | Infusion Pump, MRI Compatible
25 | RT2-05-025 | Syringe Driver, Neonatal
26 | RT2-05-026 | Syringe Driver, Paeds-Adult
27 | RT2-05-027 | Pressure Relief Pump with Mattress (180kg), Overlay
28 | RT2-05-028 | Pressure Relief Pump with Mattress (270kg), Overlay
29 | RT2-05-029 | Pressure Relief Pump with Mattress (180kg), Replacement
30 | RT2-05-030 | Pressure Relief Pump with Mattress (270kg), Replacement
31 | RT2-05-031 | Deep Vein Thrombosis (DVT) Pump
32 | RT2-05-032 | Ultrasound, Cardiology
33 | RT2-05-033 | Cardiac Output Monitor
34 | RT2-05-034 | Suction unit, wall-mount
35 | RT2-05-035 | Suction Unit, Mobile with Battery
36 | RT2-05-036 | Central Patient Monitoring, 12-Bed ICU Package
37 | RT2-05-037 | Central Patient Monitoring, 16-Bed ICU Package

## Analyzer, BloodGas, Single-use
RT2-05-001 | BLOOD GAS ANALYZER - SINGLE-USE CARTRIDGE (LOW UTILIZATION RATE) | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
1 | DESCRIPTION
1.1 | Blood Gas Analyser used for Arterial blood sample analysis and venous blood and other fluids like urine and effusions
1.2 | The analyzer must have Validation Report conducted by a local (RSA) health institution. Please provide contactable details of where it was conducted (Please attach a copy at the back of this sheet)
2 | Product Offered:
2.1 | Make:
2.2 | Model:
2.3 | Initial Manufacturing year of Model:
2.4 | Radiation Control License number:
3 | TECHNICAL SPECIFICATION
3.1 | The requirement is for a blood gas analyser that can perform various tests, and in low volumes
3.2 | The unit shall make use of a single cartridge to perform a test
3.3 | The unit shall provide lab-accurate results. If any correlation with lab reports are available, bidders to include a copy in the offer
3.4 | The unit shall operate by acquiring blood from syringe sample, capillary tube or a few drops from a finger-prick
3.5 | Test results shall be available in no longer than two minutes
3.6 | The total cost of ownership shall be clearly described in the offer. Bidders shall therefore state the price of the machine offered, as well as the cost per test for various tests. Typically blood gas and electrolyte tests will be done, but it shall also be possible to perform a range of other tests on the machine
3.7 | Blood Gas.
3.7.1 | pH
3.7.2 | pCO2 (carbon dioxide partial pressure)
3.7.3 | pO2 (oxygen partial pressure)
3.8 | Electrolytes
3.8.1 | Na+ (sodium)
3.8.2 | K+ (potassium)
3.8.3 | Ca2+ (ionised calcium)
3.8.4 | Cl- (chloride)
3.9 | Metabolites
3.9.1 | Glucose
3.10 | Oximetry
3.10.1 | /ltHb (haemoglobin)
3.10.2 | sO2 (saturation of haemoglobin)
3.10.3 | Hematocrit
3.11 | To be able to compare test costs, the bidder shall price the following, most common test (to be available from one test sample – therefore all these tests on one cartridge). Bidders to clearly show that these values are available in one cartridge, and show the cost per cartridge
3.12 | The offered system shall be user friendly, and it shall be able to operate with minimal user interaction
3.13 | No units that require daily operator maintenance will be considered. The offered unit shall include QA and verification systems. Bidders to clearly describe the way it is done
3.14 | The unit shall perform QA checks daily. Bidders to indicate if this will be done automatically, or if user intervention is required
3.15 | The unit shall be equipped with a printer to keep record of all results
3.16 | The unit shall have the ability to store patient test data, and at least 50 patient records shall be able to be stored
3.17 | When the results are printed, it shall be possible to print either the patient name or Hospital ID number on the result
3.18 | Bidders to include a box of at least 10 X Bloodgas and Electrolyte test cartridges in the offered price
4 | STANDARD ACCESSORIES
4.1 | Thermal printing paper x 10
4.2 | A starter pack consumables/accessories must be supplied with the initial supply and unit must be operational (for atleast 30 days) after commissioning. Box of at least 10 X Bloodgas and Electrolyte test cartridges
4.3 | Internal rechargable battery
5 | Conformity Compliance (Please attach certificate)
5.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
5.2 | OEM must comply and be certified on ISO 13485 quality standards, attach proof of compliance
5.2.1 | Please provide unique ref number of the ISO 13485 ceriticate:
5.3 | Model on offer must comply with EU Medical Device Regulation (EU) 2017/745 or equvalent from SAHPRA recognized jurisdictions . Attach a copy of certification
5.3.1 | Please provide unique ref number of the ceriticate:
5.4 | Hazardous Substance Act: If this type of equipment / apparatus appears on the schedule of Hazardous substances issued by the Directorate: Health Technology of the Department of Health, a license in terms of the Act on Hazardous Substances (Act 15/1973) must be submitted with the bid / quotation documents. The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.5 | All electrical/electronic medical equipment must be licensed by Radiation Control, where applicable a copy of the license must be submitted The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.5.1 | Radiation Control License number:
5.6 | The unit shall operate from 220V AC/50Hz mains supply. The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade (rubber to with 2 screws) type and it must be a minimum length of three (3) meters. NB the mains cable of the unit tendered for must be SABS colour coded.
5.7 | The equipment quoted for must be protected against electromagnetic interference IEC 60601-1-2
6 | COMPREHENSIVE MAINTENANCE
6.1 | A 24-month warranty inclusive maintenance and service (Year 1 to Year 2) should be quoted for, as per the SCC document
6.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
6.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
6.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
6.3 | The bidder to provide the OEM recommended service interval of the unit:
6.4 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
7 | General
7.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
7.2 | The unit on offer must be the latest technology.
7.2.1 | Initial Manufacturing year of Model:
RT2-05-001 | BLOOD GAS ANALYZER - SINGLE-USE CARTRIDGE ACCESSORIES AND CONSUMABLES | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
RT2-05-001a | Thermal printing paper x 10
RT2-05-001b | Bloodgas & Electrolytes cartridge
RT2-05-001c | Quality control
RT2-05-001d | Internal rechargable battery
| 9 | General (Please attach certificates)
| 9.1 | All components must bear the CE mark or equivalent.
| 9.2 | The unit must comply with the following quality & general safety standards: ISO 9001/ ISO 13485; IEC 60601-1-1
| 9.3 | The Unit must meet EMC safety standards of IEC60601 -1-2

## Analyzer, BloodGas, Cartridge
RT2-05-002 | BLOOD GAS ANALYZER - CARTRIDGE TYPE | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
1 | DESCRIPTION
1.1 | A device used to measure combinations of pH, blood gas (i.e. pCO2 and pO2), electrolytes, and metabolites parameters from whole blood samples
1.2 | The analyzer must have Validation Report conducted by local (RSA) health institution, please provide contactable details of where it was conducted (Please attach a copy)
2 | Product Offered:
2.1 | Make:
2.2 | Model:
2.3 | Initial Manufacturing year of Model:
2.4 | Radiation Control License number:
3 | TECHNICAL SPECIFICATION
3.1 | The unit's single cartridge/cassette/solution pack must be able to deliver no less than 300 tests
3.2 | Patient Data input ( i.e. ID Number, Temperature and FiO2) must be displayed.
3.3 | It must be an integrated unit requiring a single sample insertion. The blood should be automatically aspirated from either a syringe or capillary tube. The volume of blood required to produce a full analysis should not exceed 200ul (0.2ml) for a syringe sample or 100ul (0.1ml) for a capillary tube sample.
3.4 | The units used to report the results must be able to be set according to the end-users requirements.
3.5 | The following parameters must be able to be measured by the machine. The results must be accurate within the specified range (+/- 5%):
3.6 | Blood Gas.
3.6.1 | pH
3.6.2 | pCO2 (carbon dioxide partial pressure)
3.6.3 | pO2 (oxygen partial pressure)
3.7 | Electrolytes
3.7.1 | Na+ (sodium)
3.7.2 | K+ (potassium)
3.7.3 | Ca2+ (ionised calcium)
3.7.4 | Cl- (chloride)
3.8 | Metabolites
3.8.1 | Glucose
3.8.2 | Lactate
3.9 | The machine must be able to calculate the following parameters:
3.9.1 | Bicarbonate standard
3.9.2 | Bicarbonate actual
3.9.3 | Base excess
3.9.4 | Temperature adjusted values (Operator must be able to input patient’s temperature)
3.10 | Measurement Cartridges, Quality control and Waste management
3.10.1 | The machine must have measurement cartridge, quality control and waste management for analysis of all the specified parameters and must not require the use of reagent bottles or gas tanks.
3.10.2 | The cartridge(s) must be all inclusive. There must be no need to change fluids, electrodes or reagents separately
3.10.3 | The insertion and removal of the cartridge must be end-user-friendly and not require special expertise.
3.10.4 | The machine must utilize cartridges that have an expiry period (use life) not less than 28 days after installation into the machine.
3.10.5 | Cartridges of variable tests must be available for the machine so the end-user can purchase according to use.
3.10.6 | There must be a range of cartridges with sample size greater than 100 available.
3.10.7 | QC must be fully automated, with all necessary reagents contained within cartridges, and must not require the use of separate reagents or gas tanks.
3.10.8 | The number of Quality Control / Calibrations per day must be end-user adjustable to account for different usage patterns.
3.10.9 | Waste and wash must be fully contained – either within the measurement cartridge or in a separate cartridge/pack to minimize the risk of contamination of the end-user.
3.10.10 | The machine must have a system to remove clots in the unit. Please state how it is achieved
3.10.11 | The machine must take no longer than 120 seconds from time of aspiration to printing of the result. Analysis Time ≤ 100 seconds.
3.10.12 | A printer for results must be an integral part of the machine
3.10.13 | There must be a memory storage facility to store previous test results, which must be able to be printed.
3.10.14 | The unit must have integrated barcode scanner for entering required data
3.10.15 | The unit offered must be supplied with a UPS (Uninterrupted Power Supply) with backup power of at least 1 hour, the cost of which must be included in the final bid price.
3.10.16 | The final bid price must include all consumables required for the full operation of the machine for 28 days or 500 analyses inclusive of QC, whichever is the greater. These must be supplied on delivery so that the machine can be put into full service immediately on delivery.
3.10.17 | Bicarbonate and Carbonic acid content total (ctCO2) measurement
3.10.18 | Functional O2 Saturation (SO2) measurement
3.10.19 | Sample Volume
3.10.19a | Micro ≤ 100 μL
3.10.19b | Normal ≤ 200 μL
3.10.20 | Analysis Time < 100 seconds
3.10.21 | Lactate measurement
3.10.22 | Bilirubin measurement (optional)
3.10.23 | Glucose measurement
3.10.24 | Met-haemoglobin measurement (optional)
3.11 | Printer
3.12 | Display
3.13 | Automatic Calibration
3.14 | Automatic Cleaning cycle
3.15 | Patient Data input ( i.e. ID Number, Temperature and FiO2)
3.16 | Data Management
3.17 | Interface capability
4 | STANDARD ACCESSORIES
4.1 | Thermal printing paper x 10
4.2 | The unit must be offered with a 2 KVA online UPS and the cost included in the bid amount
4.3 | A starter pack consumables/accessories must be supplied with the initial supply and unit must be operational (for atleast 30 days) after commissioning
5 | Conformity Compliance (Please attach certificate)
5.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
5.2 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
5.2.1 | Please provide unique ref number of the ISO 13485 ceriticate:
5.3 | Model on offer must comply with EU Medical Device Regulation (EU) 2017/745 or equvalent from SAHPRA recognized jurisdictions . Attach a copy of certification
5.3.1 | Please provide unique ref number of the EC ceriticate:
5.4 | Hazardous Substance Act: If this type of equipment / apparatus appears on the schedule of Hazardous substances issued by the Directorate: Health Technology of the Department of Health, a license in terms of the Act on Hazardous Substances (Act 15/1973) must be submitted with the bid / quotation documents. The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.5 | All electrical/electronic medical equipment must be licensed by Radiation Control, where applicable a copy of the license must be submitted The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.5.1 | Radiation Control License number:
5.6 | The unit shall operate from 220V AC/50Hz mains supply. The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade (rubber to with 2 screws) type and it must be a minimum length of three (3) meters. NB the mains cable of the unit tendered for must be SABS colour coded.
5.7 | The equipment quoted for must be protected against electromagnetic interference IEC 60601-1-2
6 | COMPREHENSIVE MAINTENANCE
6.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) should be quoted for, as per the SCC document
6.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries and Quality control sensor), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
6.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
6.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries and Quality control sensor), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
6.3 | The bidder to provide the OEM recommended service interval of the unit:
6.4 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
7 | General
7.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
7.2 | The unit on offer must be the latest technology.
7.2.1 | Initial Manufacturing year of Model:
RT2-05-002 | BLOOD GAS ANALYZER - CARTRIDGE TYPE ACCESSORIES AND CONSUMABLES | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
RT2-05-002a | Thermal printing paper x 10
RT2-05-002b | 2 KVA online UPS
RT2-05-002c | Quality control
RT2-05-002d | Cartridge (300 test)
RT2-05-002e | Cartridge (400 test)
RT2-05-002f | Cartridge (600 test)
RT2-05-002g | CO-Oximetry (UPGRADE) - 3 End-user choice parameters
| 9 | General (Please attach certificates)
| 9.1 | All components must bear the CE mark or equivalent.
| 9.2 | The unit must comply with the following quality & general safety standards: ISO 9001/ ISO 13485; IEC 60601-1-1
| 9.3 | The Unit must meet EMC safety standards of IEC60601 -1-2

## Analyzer, BloodGas, Solutions
RT2-05-003 | BLOOD GAS ANALYZER - SOLUTIONS | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
1 | DESCRIPTION
1.1 | A device used to measure combinations of pH, blood gas (i.e. pCO2 and pO2), electrolytes, and metabolites parameters from whole blood samples
2 | Product Offered:
2.1 | Make:
2.2 | Model:
2.3 | Initial Manufacturing year of Model:
2.4 | Radiation Control License number:
3 | TECHNICAL SPECIFICATION
3.1 | A blood gas analyser, which must include bloodgas, electrolyte, glucometry and hematocrit analysis.
3.2 | The unit must employ the latest state of the art technology, which will include microprocessor control.
3.3 | The unit must be user friendly and be easy to operate with replacement of individual solutions.
3.4 | The analyser must incorporate a reasonable size display of atleast 17 cm, and must be fitted with a built-in thermal array printer.
3.5 | The analyser must utilize a scanning mechanism for easy replacements of solutions
3.6 | The following parameters must be able to be measured by the machine. The results must be accurate within the specified range (+/- 5%):
3.6.1 | Measured parameters must be as follows: Please state the range
3.6.1.1 | pH
3.6.1.2 | Pa CO2
3.6.1.3 | Pa O2
3.6.1.4 | Hemoglobin
3.6.1.5 | Barometric Pressure
3.6.1.6 | Na+
3.6.1.7 | K+
3.6.1.8 | Glucose
3.6.1.9 | Chloride
3.6.1.10 | HB
3.6.1.11 | Co-oximetry
3.6.2 | Calculated parameters must be as follows:
3.6.2.1 | pH, temp.
3.6.2.2 | Pa CO2, temp.
3.6.2.3 | Pa O2, temp.
3.6.2.4 | Sa O2.
3.6.2.5 | Hematocrit
3.6.2.6 | HCO3.
3.6.2.7 | SBC.
3.6.2.8 | ABE.
3.6.2.9 | SBE.
3.6.2.10 | TCO2 (P).
3.6.2.11 | TCO2 (b).
3.6.2.12 | AaDO2.
3.6.2.13 | PaSO
3.6.2.14 | PaSO, temp.
3.6.2.15 | HCT
3.6.2.16 | Shunt.
3.6.2.17 | Shunt, temp.
3.6.2.18 | Aniongap.
3.6.2.19 | Aniongap, k+.
3.6.3 | The analyzer must have Validation Report conducted by local (RSA) health institution. Please provide contactable details of where it was conducted (Please attach a copy)
3.7 | The unit must be capable of the following input parameters:
3.7.1 | Patient identification.
3.7.2 | Patient temperature.
3.7.3 | FiO2.
3.7.4 | Sampling time.
3.7.5 | Sample type.
3.7.6 | Department.
3.7.7 | Patient age.
3.7.8 | Patient height.
3.7.9 | Patient weight.
3.7.10 | Patient sex.
3.7.11 | tHB.
3.7.12 | Operators identification.
3.7.13 | Accession number.
3.8 | The unit must be able to perform analysis on:
3.8.1 | Whole blood.
3.8.2 | Plasma.
3.8.3 | Cerebrospinal fluid.
3.8.4 | Urine.
3.8.5 | Pleural fluid.
3.9 | The analyser must be able to accommodate samples from:
3.9.1 | Syringes.
3.9.2 | Capillary tubes.
3.9.3 | Test tubes.
3.10 | The required sample volume must not exceed 200ul
3.11 | The analyser must have a micro-mode option.
3.11.1 | Bidder to specify the micro-mode volume.
3.12 | The analyser being bidded for must include a comprehensive Internal Data Management System that is:
3.12.1 | Easy to operate.
3.12.2 | Store at least 1200 patient results.
3.13 | The Internal Data Management System must include:
3.13.1 | Set up program.
3.13.2 | Patient reports.
3.13.3 | Calibration program.
3.13.4 | Quality assurance program.
3.13.5 | Troubleshooting.
3.13.6 | In the event of any deviation in the normal range of patient results, a warning should be displayed on the screen and on the printed results.
3.13.7 | The unit will utilize touchscreen for the entry of data.
3.13.8 | The unit must be easy to maintain and must require minimum maintenance.
3.13.9 | The unit must perform an auto-cleaning, daily, which must include the cleaning of the inlet.
3.13.10 | The unit must incorporate the facility for remote troubleshooting and maintenance.
3.13.11 | The unit must be able to perform calibrations automatically at pre-selected intervals and/or on demand.
3.13.12 | All calibration results must be displayed on the screen and logged onto the Data Management System.
3.13.13 | The unit gasses must be precalibrated and supplied as consumables without additional calibration required.
3.13.14 | The unit must have the necessary ports for connection to an external information system and modem.
3.13.15 | The unit must be upgradeable to provide future needs if, and when required by the Hospital. Please specify in detail of all available upgrades.
3.13.16 | The unit must be able to operate off 220V 50hz single phase A.C. power supply and be fused in both the Live and Neutral.
3.13.17 | All electronic components on the unit must be of modular design.
3.13.18 | Bidders must provide a complete running cost of the unit and it must be based on twenty samples per day. These must include all solutions, gases, electrodes, membrances, tubing etc. required per annum.
4 | STANDARD ACCESSORIES
4.1 | Thermal printing paper x 10
4.2 | A starter pack consumables/accessories must be supplied with the initial supply and unit must be operational (for atleast 30 days) after commissioning
4.3 | The unit must be offered with 2 KVA online UPS and the cost must be included in the bid amount
5 | Conformity Compliance (Please attach certificate)
5.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
5.2 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
5.2.1 | Please provide unique ref number of the ISO 13485 ceriticate:
5.3 | Model on offer must comply with EU Medical Device Regulation (EU) 2017/745 or equvalent from SAHPRA recognized jurisdictions . Attach a copy of certification
5.3.1 | Please provide unique ref number of the EC ceriticate:
5.4 | Hazardous Substance Act: If this type of equipment / apparatus appears on the schedule of Hazardous substances issued by the Directorate: Health Technology of the Department of Health, a license in terms of the Act on Hazardous Substances (Act 15/1973) must be submitted with the bid / quotation documents. The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.5 | All electrical/electronic medical equipment must be licensed by Radiation Control, where applicable a copy of the license must be submitted The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.5.1 | Radiation Control License number:
5.6 | The unit shall operate from 220V AC/50Hz mains supply. The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade (rubber to with 2 screws) type and it must be a minimum length of three (3) meters. NB the mains cable of the unit tendered for must be SABS colour coded.
5.7 | The equipment quoted for must be protected against electromagnetic interference IEC 60601-1-2
6 | COMPREHENSIVE MAINTENANCE
6.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) should be quoted for, as per the SCC document
6.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries and Quality control sensor), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
6.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
6.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries and Quality control sensor), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
6.3 | The bidder to provide the OEM recommended service interval of the unit:
6.4 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
7 | General
7.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
7.2 | The unit on offer must be the latest technology.
7.2.1 | Initial Manufacturing year of Model:
RT2-05-003 | BLOOD GAS ANALYZER - SOLUTIONS, ACCESSORIES and CONSUMABLES | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
RT2-05-003a | Cleaning solution
RT2-05-003b | Calibration solution
RT2-05-003c | Calibration Gas bottle
RT2-05-003d | Rinse solution
RT2-05-003e | Thermal printing paper x 10
RT2-05-003f | 2 KVA online UPS
RT2-05-003g | Heparinized Syringe (100/box)

## Defib with ECG
RT2-05-004 | Defibrillator with ECG | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
1 | Description
1.1 | Defibrillator to reverse fibrillation to a normal rhythm with ECG monitoring
2 | Product Offered:
2.1 | Make:
2.2 | Model:
2.3 | Initial Manufacturing year of Model:
2.4 | Radiation Control License number:
3 | Technical Specification
3.1 | The defibrillator offered must provide ECG monitoring, with a printer
3.2 | The defibrillator, monitor, printer, battery and battery-charger must be integrated into a single unit.
3.3 | The defibrillator must deliver a bi-phasic impulse
3.4 | The Defib paddles must be detacheable from the cable or machine to exchange with AED pads
3.5 | The defibrillator must have manual and automatic (AED) modes
3.6 | Automatic mode must measure patient impedance and deliver an appropriate impulse automatically generated, between 0 and 250 J
3.7 | Audible/voice and on-screen messages for use during AED mode
3.8 | Manual mode energy selection: 5 - 270 J
3.9 | Selection of manual energy must be simple with a clear unambiguous indication of the energy selected
3.10 | The defibrillator must have the following selection of defibrillation:
3.10.1 | a. Non-synchronous defibrillation
3.10.2 | b. Synchronous defibrillation
3.11 | The defibrillator must indicate if the synchronous mode has been selected.
3.12 | The defibrillator must default to non-synchronous mode at initial switch on.
3.13 | Synchronous defibrillation detection must be possible via ECG patient cable and via paddles
3.14 | The synch marker must be clearly indicated on the ECG trace.
3.15 | The defibrillator must contain a rechargeable battery which charges automatically when the defibrillator is plugged into the 220 V, 50 Hz supply.
3.16 | The battery must charge whether the defibrillator is switched on not
3.17 | The defibrillator must have a visual indicator to indicate that the mains supply is switched on.
3.18 | The internal rechargeable battery must be able to deliver at least 50 defibrillations at the maximum setting and provide a minimum of 2 hours ECG monitoring in the event of a mains power failure
3.19 | Battery status indicator with audible and visual alarm for low battery
3.20 | Audible tone while charging
3.21 | Audible tone, different from above, to indicate full charge
3.22 | Visual indication of fully charged state
3.23 | Defibrillating charge time to maximum setting must not exceed 10 seconds.
3.24 | In the event of a completely depleted battery, the defibrillator must be fully operable on the 220 V, 50 Hz supply.
3.25 | The defibrillator must do an automatic self-test when switched on.
3.26 | Override of self test must be possible
3.27 | Operation of the defibrillator must be user-friendly with clear indications for:
3.27.1 | Switch on;
3.27.2 | Select energy;
3.27.3 | Charge;
3.27.4 | Discharge.
3.28 | The defibrillator must be portable and weigh less than 7 kg with all accessories attached
3.29 | The unit must have a carry handle
3.30 | The defibrillator must incorporate an internal 50 ohms test load to allow a check for satisfactory functioning.
3.31 | The defibrillator must incorporate an emergency dump circuit that will automatically provide a safe current path for unused energy when the paddles are charged but not discharged within sixty seconds.
3.32 | The defibrillator must be able to do both internal and external defibrillation.
3.33 | Energy selection must accommodate patients from paeds to adults for both internal and external defibrillation.
3.34 | The defibrillator must automatically sense when internal paddles are in use and restrict the maximum selectable energy to 50 Joules.
3.35 | Apex paddle to incorporate a push button to charge the defibrillator.
3.36 | Energy discharge via defibrillator paddles to patient or test load must be activated by pressing buttons on both apex and sternum paddles simultaneously.
3.37 | Energy discharge via AED pads or internal electrodes by means of a button on the panel
3.38 | ECG monitoring must be possible through either the ECG leads or the paddles and unit must display the ECG lead selected (I, II or III)/(aVl, aVR or aVF) or indicate whether the paddles are the source of the ECG
3.39 | Size of ECG: 0,5; 1; 2 and 4 cm/mV
3.40 | Sweep rate: 25 mm/s
3.41 | 50 Hz filter
3.42 | Heart rate: 30 - 300 bpm
3.43 | Analysis time: maximum 15 s
3.44 | Artefact detection
3.45 | Pacemaker detection
3.46 | The following alarms must be displayed:
3.46.1 | low battery,
3.46.2 | no paper (printer)
3.46.3 | ECG lead off/fault.
3.48 | Device must alarm audibly and visually on violation of adjustable setting of LOW and HIGH heart rate
3.49 | Alarm limit range: adjustable from 30 to 300 bpm (approximately) and 4 seconds of asystole.
3.50 | The audible alarm should have an mute function to temporarily silence the audible alarm.
3.51 | Audible alarm-silencing period: 120 s (max)
3.52 | The screen must be visible under all lighting conditions.
3.53 | Display size: Approximately 115 x 85 mm.
3.54 | Display resolution: Approximately 320 x 240 pixels
3.55 | Display to indicate the following:
3.55.1 | Alarm messages
3.55.1 | Heart rate
3.55.2 | Systole blinker
3.55.3 | Alarm limits
3.55.4 | ECG, with freeze capability
3.56 | A high-resolution, approximately 200 dpi, printer must be incorporated in the unit
3.57 | Two channel printer for ECG and SpO2 waveform
3.58 | Printer sensitivity range 1 mV = 5 mm to 1 mV = 25 mm
3.59 | Muscle artifact filter
3.60 | Print-out of the following:
3.60.1 | ECG trace
3.60.2 | Date
3.60.3 | Time
3.60.4 | Energy selected
3.60.5 | Energy delivered
3.60.6 | SpO2 wavefrom (Optional)
3.61 | Printer paper: 50 - 100 mm
3.62 | Paper speed: 25 to 50 mm/s .
3.63 | Manual and automatic printer modes.
3.64 | Automatic mode must print ECG trace before and after alarm/defibrillation conditions
3.65 | Memory/storage: At least 1 hour of ECG recording with all significant events
3.66 | SpO2 monitoring (Optional)
3.66.1 | Alarm for max and min SpO2
3.66.2 | Saturation range: 80 - 100%, accuracy within 2% to 3% for readings between 70 - 90%
3.66.3 | Pulse rate range: 30 - 300 bpm, ± 5 bpm
3.67 | NIBP monitoring (Optional)
3.67.1 | Oscillometric measurement method
3.67.2 | Ranges (approximate) :
3.67.2.1 | Systolic : 30 - 270 mmHg. Resolution : 1 mmHg
3.67.2.2 | Diastolic : 10 - 150 mmHg. Resolution : 1 mmHg
3.67.2.3 | Mean : 20 - 250 mmHg. Resolution : 1 mmHg
3.67.3 | Cuff pressure : 0 - 280 mmHg
3.67.4 | Cuff inflation time < 10 s
3.67.5 | All available adult reusable cuffs must be supplied with the monitor.
3.68 | The unit must be handed over in full operating order
3.69 | A starter-pack of consumables must be supplied with the unit.
4 | Standard Accessories and Consumbles
4.1 | 3-Lead ECG patient cable (x1)
4.2 | ECG electrodes (25/Pack)
4.3 | AED pads (x2 pairs - Paeds & Adults)
4.4 | Cable for AED pads (x1)
4.5 | Electrode gel (x1: 250ml)
4.6 | Printer paper (x 10 rolls)
5 | Conformity Compliance (Please attach certificate)
5.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
5.2 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
5.2.1 | Please provide unique ref number of the ISO 13485 ceriticate:
5.3 | Model on offer must comply with EU Medical Device Regulation (EU) 2017/745 or equvalent from SAHPRA recognized jurisdictions . Attach a copy of certification
5.3.1 | Please provide unique ref number of the EC ceriticate:
5.4 | Hazardous Substance Act: If this type of equipment / apparatus appears on the schedule of Hazardous substances issued by the Directorate: Health Technology of the Department of Health, a license in terms of the Act on Hazardous Substances (Act 15/1973) must be submitted with the bid / quotation documents. The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.5 | All electrical/electronic medical equipment must be licensed by Radiation Control, where applicable a copy of the license must be submitted The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.5.1 | Radiation Control License number:
5.6 | The unit shall operate from 220V AC/50Hz mains supply. The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade (rubber to with 2 screws) type and it must be a minimum length of three (3) meters. NB the mains cable of the unit tendered for must be SABS colour coded.
5.7 | The equipment quoted for must be protected against electromagnetic interference IEC 60601-1-2
5.8 | The unit must comply with electrical safety and performance standard such as IEC 60601-2-4, attached certification
5.8.1 | Please provide unique ref number of the IEC ceriticate:
6 | COMPREHENSIVE MAINTENANCE
6.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) should be quoted for, as per the SCC document
6.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
6.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
6.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
6.3 | The bidder to provide the OEM recommended service interval of the unit:
6.4 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
7 | General
7.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
7.2 | The unit on offer must be the latest technology.
7.2.1 | Initial Manufacturing year of Model:
RT2-05-004 | Defibrillator with ECG | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
RT2-05-004a | Reusable external adult defibrillator paddles
RT2-05-004b | Reusable external paediatric defibrillator paddles
RT2-05-004c | 3-Lead ECG patient cable (x1)
RT2-05-004d | 5-Lead ECG patient cable (x1)
RT2-05-004e | 5-Lead ECG patient cable (x1)
RT2-05-004f | ECG electrodes (25/Pack)
RT2-05-004g | AED pads (x2 pairs - Paeds & Adults)
RT2-05-004h | Cable for AED pads (x1)
RT2-05-004i | Electrode gel (x1: 250ml)
RT2-05-004j | Printer paper (x 10 rolls)
RT2-05-004k | SpO2 sensor
RT2-05-004l | SpO2 cable
RT2-05-004m | BP hose
RT2-05-004n | BP cuff (Adult)
RT2-05-004o | Dedicated trolley
RT2-05-004p | Internal paddles
RT2-05-004q | Rechargable battery

## Defib, wit ECG&Pacing
RT2-05-005 | Defibrillator, with ECG & Pacing | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
1 | Description
1.1 | Defibrillator to reverse fibrillation to a normal rhythm with pacing capabilties and ECG monitoring
2 | Product Offered:
2.1 | Make:
2.2 | Model:
2.3 | Initial Manufacturing year of Model:
2.4 | Radiation Control License number:
3 | Technical Specification
3.1 | The defibrillator offered must provide ECG monitoring, a printer and external pacing
3.2 | The defibrillator, monitor, printer, pacing unit, battery and battery-charger must be integrated into a single unit.
3.3 | The defibrillator must deliver a bi-phasic impulse
3.4 | The Defib paddles must be detacheable from the cable or machine to exchange with AED pads
3.5 | The defibrillator must have manual and automatic (AED) modes
3.6 | Automatic mode must measure patient impedance and deliver an appropriate impulse automatically generated to a maximum of 250 J
3.7 | Audible/voice and on-screen messages for use during AED mode
3.8 | Manual mode energy selection: 5 - 270 J
3.9 | Selection of manual energy must be simple with a clear unambiguous indication of the energy selected
3.10 | The defibrillator must have the following selection of defibrillation:
3.11 | a. Non-synchronous defibrillation
3.12 | b. Synchronous defibrillation
3.13 | The defibrillator must indicate if the synchronous mode has been selected.
3.14 | The defibrillator must default to non-synchronous mode at initial switch on.
3.15 | Synchronous defibrillation detection must be possible via ECG patient cable and via paddles
3.16 | The synch marker must be clearly indicated on the ECG trace.
3.17 | The defibrillator must contain a rechargeable battery which charges automatically when the defibrillator is plugged into the 220 V, 50 Hz supply.
3.18 | The battery must charge whether the defibrillator is switched on or not
3.19 | The defibrillator must have a visual indicator to indicate that the mains supply is switched on.
3.20 | The internal rechargeable battery must be able to deliver at least 50 defibrillations at the maximum setting and provide a minimum of 2 hours ECG monitoring in the event of a mains power failure
3.21 | Battery status indicator with audible and visual alarm for low battery
3.22 | Audible tone while charging
3.23 | Audible tone, different from above, to indicate full charge
3.24 | Visual indication of fully charged state
3.25 | Defibrillating charge time to maximum setting must not exceed 10 seconds.
3.26 | In the event of a completely depleted battery, the defibrillator must be fully operable on the 220 V, 50 Hz supply.
3.27 | The defibrillator must do an automatic self-test when switched on.
3.28 | Override of self test must be possible
3.29 | Operation of the defibrillator must be user-friendly with clear indications for:
3.29.1 | Switch on;
3.29.2 | Select energy;
3.29.3 | Charge;
3.29.4 | Discharge.
3.30 | The defibrillator must be portable and weigh less than 7 kg with all accessories attached
3.31 | The unit must have a carry handle
3.32 | The defibrillator must incorporate an internal 50 ohms test load to allow a check for satisfactory functioning.
3.33 | The defibrillator must incorporate an emergency dump circuit that will automatically provide a safe current path for unused energy when the paddles are charged but not discharged within sixty seconds.
3.34 | The defibrillator must be able to do both internal and external defibrillation.
3.35 | Energy selection must accommodate patients from paeds to adults for both internal and external defibrillation.
3.36 | The defibrillator must automatically sense when internal paddles are in use and restrict the maximum selectable energy to 50 Joules.
3.37 | Apex paddle to incorporate a push button to charge the defibrillator.
3.38 | Energy discharge via defibrillator paddles to patient or test load must be activated by pressing buttons on both apex and sternum paddles simultaneously.
3.39 | Energy discharge via AED pads or internal electrodes by means of a button on the panel
3.40 | ECG monitoring must be possible through either the ECG leads or the paddles and unit must display the ECG lead selected (I, II or III)/(aVl, aVR or aVF) or indicate whether the paddles are the source of the ECG
3.40.1 | Size of ECG: 0,5; 1; 2 and 4 cm/mV
3.40.2 | Sweep rate: 25 mm/s
3.40.3 | 50 Hz filter
3.40.4 | Heart rate: 30 - 300 bpm
3.40.5 | Analysis time: maximum 15 s
3.40.6 | Artifact detection
3.40.7 | Pacemaker detection
3.41 | The following alarms must be displayed:
3.41.1 | low battery,
3.41.2 | no paper (printer)
3.41.3 | ECG lead off/fault.
3.42 | Device must alarm audibly and visually on violation of adjustable setting of LOW and HIGH heart rate
3.43 | Alarm limit range: adjustable from 30 to 300 bpm (approximately) and 4 seconds of asystole.
3.44 | The audible alarm should have a mute function to temporarily silence the audible alarm.
3.45 | Audible alarm-silencing period: 120 s (max)
3.46 | The screen must be visible under all lighting conditions.
3.47 | Display size: Approximately 115 x 85 mm.
3.48 | Display resolution: Approximately 320 x 240 pixels
3.49 | Display to indicate the following:
3.49.1 | Alarm messages
3.49.2 | Heart rate
3.49.3 | Systole blinker
3.49.4 | Alarm limits
3.49.5 | ECG, with freeze capability
3.50 | A high-resolution, approximately 200 dpi, printer must be incorporated in the unit
3.51 | Two channel printer for ECG and SpO2 waveform
3.52 | Printer sensitivity range 1 mV = 5 mm to 1 mV = 25 mm
3.53 | Muscle artifact filter
3.54 | Print-out of the following:
3.54.1 | ECG trace
3.54.2 | SpO2 wavefrom
3.54.3 | Date
3.54.4 | Time
3.54.5 | Energy selected
3.54.6 | Energy delivered
3.55 | Printer paper: 50 -100mm
3.56 | Paper speed: 25 to 50 mm/s .
3.57 | Manual and automatic printer modes.
3.58 | Automatic mode must print ECG trace before and after alarm/defibrillation conditions
3.59 | Memory/storage: At least 1 hour of ECG recording with all significant events
3.60 | External pacemaker
3.60.1 | Modes:
3.60.1.1 | Fixed
3.60.1.2 | Demand
3.60.2 | 30 - 180 paces/min (Fixed, demand)
3.60.3 | Impulse current: 10 - 180 mA, in 1 mA steps
3.60.4 | Impulse shape: Square: 20 -40 ms
3.60.5 | Refractory period (Demand): 240 ms (HR > 100 bpm), 240 ms (HR < 100 bpm)
3.60.6 | QRS marker in Demand mode
3.60.7 | There must be a continuous display of the pacing mode, the selected current and the selected pacing rate.
3.61 | The unit must be handed over in full operating order
3.62 | A starter-pack of consumables must be supplied with the unit.
3.63 | SpO2 monitoring (Optional)
3.63.1 | Alarm for max and min SpO2
3.63.2 | Saturation range: 80 - 100%, accuracy within 2% to 3% for readings between 70 - 90%
3.63.3 | Pulse rate range: 30 - 300 bpm, ± 5 bpm
3.64 | NIBP monitoring (Optional)
3.64.1 | Oscillometric measurement method
3.64.2 | Ranges (approximate) :
3.64.2.1 | Systolic : 30 - 270 mmHg. Resolution : 1 mmHg
3.64.2.2 | Diastolic : 10 - 150 mmHg. Resolution : 1 mmHg
3.64.2.3 | Mean : 20 - 250 mmHg. Resolution : 1 mmHg
3.64.3 | Cuff pressure : 0 - 280 mmHg
3.64.4 | Cuff inflation time < 10 s
3.64.5 | All available adult reusable cuffs must be supplied with the monitor.
4 | Standard Accessories and Consumables
4.1 | 3-Lead ECG patient cable (x1)
4.2 | ECG electrodes (25/pack)
4.3 | AED pads (x2 pairs- Paeds & Adults)
4.4 | Cable for AED pads (x1)
4.5 | Pacing electrodes
4.6 | Pacing cable
4.7 | Electrode gel (x1: 250ml)
4.8 | Printer paper (x 10 rolls)
5 | Conformity Compliance (Please attach certificate)
5.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
5.2 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
5.2.1 | Please provide unique ref number of the ISO 13485 ceriticate:
5.3 | Model on offer must comply with EU Medical Device Regulation (EU) 2017/745 or equvalent from SAHPRA recognized jurisdictions . Attach a copy of certification
5.3.1 | Please provide unique ref number of the EC ceriticate:
5.4 | Hazardous Substance Act: If this type of equipment / apparatus appears on the schedule of Hazardous substances issued by the Directorate: Health Technology of the Department of Health, a license in terms of the Act on Hazardous Substances (Act 15/1973) must be submitted with the bid / quotation documents. The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.5 | All electrical/electronic medical equipment must be licensed by Radiation Control, where applicable a copy of the license must be submitted The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.5.1 | Radiation Control License number:
5.6 | The unit shall operate from 220V AC/50Hz mains supply. The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade (rubber to with 2 screws) type and it must be a minimum length of three (3) meters. NB the mains cable of the unit tendered for must be SABS colour coded.
5.7 | The equipment quoted for must be protected against electromagnetic interference IEC 60601-1-2
5.8 | The unit must comply with electrical safety and performance standard such as IEC 60601-2-4, attached certification
5.8.1 | Please provide unique ref number of the IEC ceriticate:
6 | COMPREHENSIVE MAINTENANCE
6.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) should be quoted for, as per the SCC document
6.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
6.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
6.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
6.3 | The bidder to provide the OEM recommended service interval of the unit:
6.4 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
7 | General
7.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
7.2 | The unit on offer must be the latest technology.
7.2.1 | Initial Manufacturing year of Model:
RT2-05-005 | Defibrillator Accessories and Consumables | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
RT2-05-005a | Reusable external adult defibrillator paddles
RT2-05-005b | Reusable external paediatric defibrillator paddles
RT2-05-005c | 3-Lead ECG patient cable (x1)
RT2-05-005d | 5-Lead ECG patient cable (x1)
RT2-05-005e | 5-Lead ECG patient cable (x1)
RT2-05-005f | ECG electrodes (25/Pack)
RT2-05-005g | AED pads (x2 pairs - Paeds & Adults)
RT2-05-005h | Cable for AED pads (x1)
RT2-05-005i | Electrode gel (x1: 250ml)
RT2-05-005j | Printer paper (x 10 rolls)
RT2-05-005k | SpO2 sensor
RT2-05-005l | SpO2 cable
RT2-05-005m | BP hose
RT2-05-005n | BP cuff (Adult)
RT2-05-005o | Dedicated trolley
RT2-05-005p | Internal paddles
RT2-05-005q | Rechargable battery

## ECG, 12-Lead
RT2-05-006 | ELECTROCARDIOGRAPH (12 CHANNEL) on trolley. | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
1 | Description
1.1 | A machine that checks the fuctioning of the Heart, recording how often the heart beats and how regularly it beats.
2 | Product Offered:
2.1 | Make:
2.2 | Model:
2.3 | Initial Manufacturing year of Model:
2.4 | Radiation Control License number:
3 | Technical Specification
3.1 | Twelve channel portable electrocardiograph
3.2 | The unit must contain a recorder/printer.
3.3 | The unit must have at least 7 inch colour TFT display
3.5 | The unit must do a self-test at power-on
3.6 | The following must be displayed:
3.6.1 | Heart rate
3.6.2 | Patient ID
3.6.3 | 12 lead waveforms
3.6.4 | Lead identification
3.6.5 | ECG gain
3.6.6 | Paper speed
3.6.7 | Filters
3.6.8 | Electrode connection indicator
3.6.9 | Messages and prompts
3.6.10 | Clock with date and time
3.6.11 | Battery power status
3.7 | R-R analysis for 30 to 300 sec
3.8 | Pacemaker detection
3.9 | The display screen must be capable of displaying two lines of text and a minimum of three ECG waveforms.
3.10 | The unit must be supplied with a 10 lead ECG cable and be able to record 3, 6 and 12 lead configurations.
3.11 | For the 12 lead configuration:
3.11.1 | Lead I, Lead II, Lead III
3.11.2 | AVR, AVL, AVF
3.11.3 | V1, V2, V3, V4, V5, V6
3.12 | Manual and automatic modes
3.13 | The unit must be capable of ECG interpretation for paediatric to adult.
3.14 | Full alphanumeric keyboard or touchscreen
3.15 | The unit must have facility for entering patient data and text
3.16 | The unit must offer a minimum of six pre-set ECG programs.
3.17 | The unit must have Event markers
3.18 | Automatic indication of disconnected leads/electrodes
3.19 | Automatic check for incorrect lead placement
3.20 | Full screen print preview of ECG waveforms prior to printing
3.21 | Heart rate recognition: peak-to peak detection. 30 - 300 bpm, accuracy 1 bpm
3.22 | Electrode connection quality check.
3.23 | ECG amplifier specs
3.23.1 | Simultaneous acquisition of 12 leads
3.23.2 | Frequency response: 0.05 - 150 Hz
3.23.3 | Sensitivity: Automatic or selectable: 2.5 - 20 mm/mV
3.23.4 | Automatic amplitude adjustment in the range of 2.5 to 40 mm/mV.
3.23.5 | Sensitivity threshold: 20 µV
3.23.6 | Input impedance: > 5 MΩ
3.23.7 | Patient leakage current: < 10 µA
3.23.8 | Calibration pulse: 1 mV
3.23.9 | Filters:
3.23.9a | 50 Hz
3.23.9b | Low pass: 25, 35, 75, 100, 150 Hz
3.23.10 | CMRR: > 100 dB
3.23.11 | Sampling frequency: 1 kHz
3.23.12 | Artifact rejection
3.23.13 | Baseline stabiliser
3.24 | Printer
3.24.1 | Automatic and manual printout mode.
3.24.2 | In addition to the required ECG waveforms the printer must be able to print:
3.24.2a | device settings
3.24.2b | heart rate
3.24.2c | date and time
3.24.2d | lead identification
3.24.2e | calibration pulse
3.24.2f | patient data
3.24.3 | The ECG printout must be available on A4 size, Z fold paper.
3.24.4 | Printer quality: 8 dots/mm (vertical), 20 dots/mm (horiz, 25 mm/s)
3.24.5 | Paper speed must be selectable as follows: 25 and 50 mm/s.
3.25 | Internal memory capable of storing at least 200 patient files.
3.26 | External storage .bmo SD card (or other) for 2 000 patient files
3.27 | Internal protection against damage from 360 J defibrillator discharge.
3.28 | The unit must have an internal battery with integrated charger
3.29 | Battery capacity: at least 2 hours of continuous recording
3.30 | The unit must be fully operable from mains power regardless of the internal battery condition.
3.31 | A sturdy trolley to house the ECG machine and accessories must be supplied.
3.32 | Trolley must have at least one drawer
3.33 | Trolley must have at least four castors, two with brakes
3.34 | The unit must be handed over in full operating order
3.35 | A starter-pack of consumables must be supplied with the unit.
4 | Standard Accessories and Consumables
4.1 | ECG cable with adaptor clips
4.2 | ECG electrodes (25/Pack)
4.3 | Printing paper (x5)
4.4 | Dedicated trolley
5 | Conformity Compliance (Please attach certificate)
5.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
5.2 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
5.2.1 | Please provide unique ref number of the ISO 13485 ceriticate:
5.3 | Model on offer must comply with EU Medical Device Regulation (EU) 2017/745 or equvalent from SAHPRA recognized jurisdictions . Attach a copy of certification
5.3.1 | Please provide unique ref number of the EC ceriticate:
5.4 | Hazardous Substance Act: If this type of equipment / apparatus appears on the schedule of Hazardous substances issued by the Directorate: Health Technology of the Department of Health, a license in terms of the Act on Hazardous Substances (Act 15/1973) must be submitted with the bid / quotation documents. The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.5 | All electrical/electronic medical equipment must be licensed by Radiation Control, where applicable a copy of the license must be submitted The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.5.1 | Radiation Control License number:
5.6 | The unit shall operate from 220V AC/50Hz mains supply. The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade (rubber to with 2 screws) type and it must be a minimum length of three (3) meters. NB the mains cable of the unit tendered for must be SABS colour coded.
5.7 | The equipment quoted for must be protected against electromagnetic interference IEC 60601-1-2
5.8 | The unit must comply with electrical safety and performance standard such as IEC 60601-2-25, attached certification
5.8.1 | Please provide unique ref number of the IEC ceriticate:
5.9 | The unit must comply with electrical safety and performance standard such as IEC 60601-2-51, attached certification
5.9.1 | Please provide unique ref number of the IEC ceriticate:
6 | COMPREHENSIVE MAINTENANCE
6.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) should be quoted for, as per the SCC document
6.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
6.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
6.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
6.3 | The bidder to provide the OEM recommended service interval of the unit:
6.4 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
7 | General
7.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
7.2 | The unit on offer must be the latest technology.
7.2.1 | Initial Manufacturing year of Model:
RT2-05-006 | ELECTROCARDIOGRAPH (12 CHANNEL) ACCESSORIES and CONSUMABLES | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
RT2-05-006a | ECG cable with adaptor clips
RT2-05-006b | ECG electrodes (25/Pack)
RT2-05-006c | Printing paper (x5)
RT2-05-006d | Dedicated trolley
RT2-05-006e | Rechargable battery

## ECMO
RT2-05-007 | ECMO MACHINE (Extracorporeal Membrane Oxygenation Unit) | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
1 | DESCRIPTION
1.1 | A portable heart-lung cardiac bypass support system unit for use in the fields of Critical care, Emergency Medicine, cardiology and cardiac surgery to provide circulatory and/or pulmonary support
2 | Product Offered:
2.1 | Make:
2.2 | Model:
2.3 | Initial Manufacturing year of Model:
3 | TECHNICAL SPECIFICATION
3.1 | It must be suitable for use in the operating theatre, ICU, Accident and Emergency departments
3.2 | It must be suitable for both paediatric and adult use
3.3 | The unit being offered must comprise of the latest technology. It must be user friendly and incorporate reliable safety features
3.4 | The unit being quoted for must operate from 220V, 50Hz A.C. mains supply. The unit must be fused in both LIVE and NEUTRAL.
3.5 | Compact and lightweight design to ensure portability during patient transfer and transportation
3.6 | The unit to have a user-friendly touchscreen interface with a large LCD display for ease of use
3.7 | The unit must include the following intergrated Sensors :
3.7.1 | 4x external pressures
3.7.2 | 3x internal pressures
3.7.3 | 2x external temperatures
3.7.4 | 2x internal temperatures
3.7.5 | 1x venous oxygen saturation
3.7.6 | 1x haemoglobin
3.7.7 | 1x haematocrit
3.7.8 | 1x flow sensor with intergrated bubble sensor
3.7.9 | 1x bubble sensor
3.7.10 | 1x level sensor
3.8 | Must include a battery pack with a minimum of 90 minutes battery life (when fully charged) to allow for independent usage.
3.9 | The unit must have an indicator to display battery status
3.10 | Interfaces:
3.10.1 | 1 x USB port type A (for data export on USB stick)
3.10.2 | 1 x USB port type B (for data recording system and service purposes)
3.10.3 | 1x Connection for alarm output (ward call)
3.11 | An Intergrated disposable heart-lung support set to be provided and include:
3.11.1 | gas exchanger with diffusion membrane
3.11.2 | highly efficient heat exchanger for precise temperature management
3.11.3 | incorporated low-trauma centrifugal pump
3.11.4 | infrared sensor to measure venous oxygen saturations, haematocrit, haemoglobin and venous temperature
3.12 | This intergrated disposable set must come in a sterile packaging containing all the necessary components and tubing to operate the unit
3.13 | The tubing must have a low priming volume to ensure less haemodilution (<700ml)
3.14 | The tubing should be colour-coded to ensure safe patient connection
3.15 | The tubing should have a BIOLINE heparin coating (SOFTLINE for certain sets where HIT is encountered) and be DEHP-plasticizer free
3.16 | The set should provide a maximum blood flow of 0.5 - 7l/min
3.17 | Surface area gas exchange membrane range of 1.3 m2 - 1.8 m2 and a surface area heat exchange 0.3 m2 - 0.4 m2
3.18 | Disposable Insertion kits to be provided with each heart-lung set for cannulation of patient's vessels and to include:
3.18.1 | Cannulae of various sizes: outside diameters (from 15 to 29Fr) and lengths (from 15 to 55 cm), allowing individual selection based on the patient’s needs.
3.18.2 | Cannulae to be manufactured out of biocompatible polyurethane and coated with heparin BIOLINE
3.18.3 | All cannulae to be wire-reinforced to prevent kinking
3.18.4 | All cannulae to be thin-walled to ensure best flow with minimal pressure drop
4 | STANDARD ACCESSORIES
4.1 | Mobile stand/cart with at least 4 castors and 2 fitted with brakes. Stand to include a height adjustable infusion pole and a retractable drawer for accessories
4.2 | Disposable Insertion kits (pack of 5)
4.3 | Guidewires of different lengths to be part of each insertion kit to fulfil requirements of both arterial and venous cannulation
4.4 | Multi-step dilators to be part of each insertion kit
4.5 | Puncture needle as a component of the insertion kit
4.6 | A scalpel and syringe to be part of each insertion kit
5 | Conformity Compliance (Please attach certificate)
5.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
5.2 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
5.2.1 | Please provide unique ref number of the ISO 13485 ceriticate:
5.3 | Model on offer must comply with EU Medical Device Regulation (EU) 2017/745 or equvalent from SAHPRA recognized jurisdictions . Attach a copy of certification
5.3.1 | Please provide unique ref number of the EC ceriticate:
5.4 | Hazardous Substance Act: If this type of equipment / apparatus appears on the schedule of Hazardous substances issued by the Directorate: Health Technology of the Department of Health, a license in terms of the Act on Hazardous Substances (Act 15/1973) must be submitted with the bid / quotation documents. The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.5 | All electrical/electronic medical equipment must be licensed by Radiation Control, where applicable a copy of the license must be submitted The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.5.1 | Radiation Control License number:
5.6 | The unit shall operate from 220V AC/50Hz mains supply. The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade (rubber to with 2 screws) type and it must be a minimum length of three (3) meters. NB the mains cable of the unit tendered for must be SABS colour coded.
5.7 | The equipment quoted for must be protected against electromagnetic interference IEC 60601-1-2
6 | COMPREHENSIVE MAINTENANCE
6.1 | A 24-month warranty inclusive maintenance and service (Year 1 to Year 2) should be quoted for, as per the SCC document
6.1.1 | THIS ENTAILS: A 24-month warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 24-month warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 24-month warranty will commence after formal Commissioning and handover of the equipment.
6.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
6.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
6.3 | The bidder to provide the OEM recommended service interval of the unit:
6.4 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
7 | General
7.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
7.2 | The unit on offer must be the latest technology.
7.2.1 | Initial Manufacturing year of Model:
RT2-05-007 | ECMO MACHINE (Extracorporeal Membrane Oxygenation Unit) Accessories and Consumables | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
RT2-05-007a | Gas exchanger with diffusion membrane
RT2-05-007b | Heat exchanger
RT2-05-007c | Low-trauma centrifugal pump
RT2-05-007d | Infrared sensor
RT2-05-007e | Disposable Insertion kits (Multi-step dilators, Puncture needle, scalpel and syringe)
RT2-05-007f | Cannulae of various sizes: outside diameters (from 15 to 29Fr) and lengths (from 15 to 55 cm)
RT2-05-007g | Guidewires of different lengths
RT2-05-007h | Heart lung set

## Exam light, LED, Mobile
RT2-05-008 | Examination light, LED, Mobile | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
1 | Description
1.1 | A mobile device that directly illuminates specific areas of the body during diagnostic procedures and general examinations, using LED technology.
2 | Product Offered:
2.1 | Make:
2.2 | Model:
2.3 | Initial Manufacturing year of Model:
3 | Technical Specification
3.1 | The examination light must be a single head LED with suitable reflector.
3.2 | The lamp must be mounted on a mobile stand.
3.3 | The lamp must have sufficient movement to ensure effective lighting over the bed with a minimum length of 1m
3.4 | The extension must be spring balanced with smooth and easy manoeuvring with one hand
3.5 | No movement (drift) must take place after positioning the lamp.
3.6 | Must have sturdy upright pole at least 1,5 m with enough counter-weight to support arm at full reach at 90o
3.7 | Must have friction joints to allow for mobility and positioning.
3.8 | Must have a handle for positioning the light
3.9 | Power supply must be hidden in the base of the unit, or neatly installed out of sight.
3.10 | Light intensity of the lamp must not be less than 30 000 lux measured at 1,0 metre.
3.11 | The luminaire must be cool to touch, even during multiple examinations.
3.12 | The light must be separately switched either on the column or on the luminaire.
3.13 | Colour temp, approximately 4 000K
3.14 | Colour rendering index: > 90
3.15 | Lamp life: > 30 000 continuous hours
3.16 | Spot size (maximum):
3.16.1 | Working distance of 50 cm = 15 cm
3.16.2 | Working distance of 100 cm = 25 cm
3.17 | Base diameter: at least 35 cm
3.18 | Must operate from 230V, 50 Hz mains power
3.19 | Mobile stand: lockable castors + hook for mains cable
3.20 | Durable castors
3.21 | Smooth, easily cleaned surface
3.22 | The unit must be handed over in full operating order
4 | Conformity Compliance (Please attach certificate)
4.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
4.2 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
4.2.1 | Please provide unique ref number of the ISO 13485 ceriticate:
4.3 | Model on offer must comply with EU Medical Device Regulation (EU) 2017/745 or equvalent from SAHPRA recognized jurisdictions . Attach a copy of certification
4.3.1 | Please provide unique ref number of the EC ceriticate:
4.4 | Hazardous Substance Act: If this type of equipment / apparatus appears on the schedule of Hazardous substances issued by the Directorate: Health Technology of the Department of Health, a license in terms of the Act on Hazardous Substances (Act 15/1973) must be submitted with the bid / quotation documents. The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
4.4.1 | All electrical/electronic medical equipment must be licensed by Radiation Control, where applicable a copy of the license must be submitted The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
4.5 | Radiation Control License number:
4.6 | The unit shall operate from 220V AC/50Hz mains supply. The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade (rubber to with 2 screws) type and it must be a minimum length of three (3) meters. NB the mains cable of the unit tendered for must be SABS colour coded.
4.7 | The equipment quoted for must be protected against electromagnetic interference IEC 60601-1-2
5 | COMPREHENSIVE MAINTENANCE
5.1 | A 24-month warranty inclusive maintenance and service (Year 1 to Year 2) should be quoted for, as per the SCC document
5.1.1 | THIS ENTAILS: A 24-month warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 24-month warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 24-month warranty will commence after formal Commissioning and handover of the equipment.
5.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
5.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
5.3 | The bidder to provide the OEM recommended service interval of the unit:
5.4 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
6 | General
6.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
6.2 | The unit on offer must be the latest technology.
6.2.1 | Initial Manufacturing year of Model:
RT2-05-008 | Examination light, LED, Mobile - Accessories | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
RT2-05-008a | Bidders to state the replacement cost per LED or range of LEDs.

## Exam light, LED, Mobile+Battery
RT2-05-009 | Examination light, LED, mobile, with battery | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
1 | Description
1.1 | A mobile device that directly illuminates specific areas of the body during diagnostic procedures and general examinations, using LED technology, with battery power backup
2 | Product Offered:
2.1 | Make:
2.2 | Model:
2.3 | Initial Manufacturing year of Model:
3 | Technical Specification
3.1 | The examination light must be a single head LED with suitable reflector.
3.2 | The lamp must be mounted on a mobile stand
3.3 | The lamp must have sufficient movement to ensure effective lighting over the bed with a minimum length of 1m
3.4 | The extension must be spring balanced with smooth and easy manoeuvring with one hand
3.5 | No movement (drift) must take place after positioning the lamp.
3.6 | Must have sturdy upright pole at least 1,5 m with enough counter-weight to support arm at full reach at 90o
3.7 | Must have friction joints to allow for mobility and positioning.
3.8 | Must have a handle for positioning the light
3.9 | Power supply and battery must be hidden in the base of the unit, or neatly installed out of sight.
3.10 | Light intensity of the lamp must not be less than 30 000 lux measured at 1,0 metre.
3.11 | The luminaire must be cool to touch, even during multiple examinations.
3.12 | The light must be separately switched either on the column or on the luminaire.
3.13 | Colour temp, approximately 4 000K
3.14 | Colour rendering index: > 90
3.15 | Lamp life: > 30 000 continuous hours
3.16 | Spot size (maximum):
3.16.1 | Working distance of 50 cm = 15 cm
3.16.2 | Working distance of 100 cm = 25 cm
3.17 | Base diameter: at least 35 cm
3.18 | Must operate from 230V, 50 Hz mains power
3.19 | Must contain a battery charger
3.20 | Battery must be capable for at least 2 hours of operation
3.21 | Mobile stand: lockable castors + hook for mains cable
3.22 | Durable castors
3.23 | Smooth, easily cleaned surface
3.24 | The unit must be handed over in full operating order
4 | Conformity Compliance (Please attach certificate)
4.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
4.2 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
4.2.1 | Please provide unique ref number of the ISO 13485 ceriticate:
4.3 | Model on offer must comply with EU Medical Device Regulation (EU) 2017/745 or equvalent from SAHPRA recognized jurisdictions . Attach a copy of certification
4.3.1 | Please provide unique ref number of the EC ceriticate:
4.4 | Hazardous Substance Act: If this type of equipment / apparatus appears on the schedule of Hazardous substances issued by the Directorate: Health Technology of the Department of Health, a license in terms of the Act on Hazardous Substances (Act 15/1973) must be submitted with the bid / quotation documents. The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
4.4.1 | All electrical/electronic medical equipment must be licensed by Radiation Control, where applicable a copy of the license must be submitted The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
4.5 | Radiation Control License number:
4.6 | The unit shall operate from 220V AC/50Hz mains supply. The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade (rubber to with 2 screws) type and it must be a minimum length of three (3) meters. NB the mains cable of the unit tendered for must be SABS colour coded.
4.7 | The equipment quoted for must be protected against electromagnetic interference IEC 60601-1-2
5 | COMPREHENSIVE MAINTENANCE
5.1 | A 24-month warranty inclusive maintenance and service (Year 1 to Year 2) should be quoted for, as per the SCC document
5.1.1 | THIS ENTAILS: A 24-month warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 24-month warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 24-month warranty will commence after formal Commissioning and handover of the equipment.
5.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
5.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
5.3 | The bidder to provide the OEM recommended service interval of the unit:
5.4 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
6 | General
6.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
6.2 | The unit on offer must be the latest technology.
6.2.1 | Initial Manufacturing year of Model:
RT2-05-009 | Examination light, LED, Mobile + Battery - Accessories | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
RT2-05-009a | Bidders to state the replacement cost per LED or range of LEDs.
RT2-05-009b | Rechargable battery

## Exam light, LED, Wall-mount
RT2-05-010 | Examination light, LED, wall mount | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
1 | Description
1.1 | A mobile device that directly illuminates specific areas of the body during diagnostic procedures and general examinations, using LED technology, that is wall-mounted
2 | Product Offered:
2.1 | Make:
2.2 | Model:
2.3 | Initial Manufacturing year of Model:
3 | Technical Specification
3.1 | The examination light must be a single head LED with suitable reflector.
3.2 | The lamp must have sufficient movement to ensure effective lighting over the bed with a minimum length of 1m
3.3 | The extension must be spring balanced with smooth and easy manoeuvring with one hand. No "goose-necks" will be acceptable
3.4 | Must be able to push completely out of the way against the wall for storage.
3.5 | The unit must have Restriction to prevent bashing against the wall
3.6 | No movement (drift) must take place after positioning the lamp.
3.7 | Must have friction joints to allow for mobility and positioning.
3.8 | Must have a handle for positioning the light
3.9 | Light intensity of the lamp must not be less than 30 000 lux measured at 1,0 metre.
3.10 | The luminaire must be cool to touch, even during multiple examinations.
3.11 | Colour temp, approximately 4 000K
3.12 | Colour rendering index: > 90
3.13 | Lamp life: > 30 000 continuous hours
3.14 | Spot size (maximum):
3.15 | Working distance of 50 cm = 15 cm
3.16 | Working distance of 100 cm = 25 cm
3.17 | Power supply to be built in and mounted on the wall
3.18 | Smooth, easily cleaned surface
3.19 | Bidder must be responsible for installation/mounting and electrical connection at no extra cost. End-user consulation will be required prior installation.
3.20 | The unit must be handed over in full operating order
4 | Conformity Compliance (Please attach certificate)
4.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
4.2 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
4.2.1 | Please provide unique ref number of the ISO 13485 ceriticate:
4.3 | Model on offer must comply with EU Medical Device Regulation (EU) 2017/745 or equvalent from SAHPRA recognized jurisdictions . Attach a copy of certification
4.3.1 | Please provide unique ref number of the EC ceriticate:
4.4 | Hazardous Substance Act: If this type of equipment / apparatus appears on the schedule of Hazardous substances issued by the Directorate: Health Technology of the Department of Health, a license in terms of the Act on Hazardous Substances (Act 15/1973) must be submitted with the bid / quotation documents. The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
4.4.1 | All electrical/electronic medical equipment must be licensed by Radiation Control, where applicable a copy of the license must be submitted The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
4.5 | Radiation Control License number:
4.6 | The unit shall operate from 220V AC/50Hz mains supply. The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade (rubber to with 2 screws) type and it must be a minimum length of three (3) meters. NB the mains cable of the unit tendered for must be SABS colour coded.
4.7 | The equipment quoted for must be protected against electromagnetic interference IEC 60601-1-2
5 | COMPREHENSIVE MAINTENANCE
5.1 | A 24-month warranty inclusive maintenance and service (Year 1 to Year 2) should be quoted for, as per the SCC document
5.1.1 | THIS ENTAILS: A 24-month warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 24-month warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 24-month warranty will commence after formal Commissioning and handover of the equipment.
5.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
5.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
5.3 | The bidder to provide the OEM recommended service interval of the unit:
5.4 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
6 | General
6.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
6.2 | The unit on offer must be the latest technology.
6.2.1 | Initial Manufacturing year of Model:
RT2-05-010 | Examination light, LED, wall mount- Accessories | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
RT2-05-010a | Bidders to state the replacement cost per LED or range of LEDs.

## Infusion Pump, Neo
RT2-05-011 | INFUSION PUMP, Neonatal | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
1 | Description
1.1 | Volumetric infusion pump dedicated for neonatal patients
2 | Product Offered:
2.1 | Make:
2.2 | Model:
2.3 | Initial Manufacturing year of Model:
2.4 | Radiation Control License number:
3 | Technical Specification
3.1 | The unit must use a linear peristaltic drive mechanism. Other drive mechanisms may be offered and details must be provided.
3.2 | The unit must provide user selectable infusion rate from 1 to 999 ml/h, selectable in 1 ml/h increments.
3.3 | The unit must provide user selectable infusion volume from 1 to 9 999 ml, selectable in 1 ml increments.
3.4 | The unit must deliver the selected volume with an accuracy better than 5% throughout the whole range of operation/infusion.
3.5 | Bolus rate, user selectable: 10 - 999 ml/hr
3.6 | Bolus volume, user selectable: 1 - 1000 ml
3.7 | The unit must also be capable of carrying out blood administration.
3.8 | It must not be possible to change the rate while the infusion is in progress, without confirming change.
3.9 | The unit must have an LCD display
3.10 | The following must be clearly displayed on the front panel under all lighting conditions:
3.10.1 | Pump running
3.10.2 | AC mains supply operation
3.10.3 | AC power failure/battery operation
3.10.4 | Infusion rate selected
3.10.5 | Volume selected to be infused
3.10.6 | Volume infused
3.10.7 | Alarm condition and alarm/error messages
3.11 | When an infusion is completed an audible tone must sound.
3.12 | On completion of an infusion there must be a Keep Vein Open (KVO) rate of 1 - 5 ml/h.
3.13 | The unit must provide the following alarms:
3.13.1 | Air in line detection
3.13.2 | Occlusion during infusion (downstream and upstream)
3.13.3 | Occlusion pressure alarms
3.13.3a | Occlusion pressure alarm user selectable: 50 - 150 mmHg
3.13.3b | The unit must be neonatal dedicated pump and occlusion pressure must not exceed the 3.13.3a specified maximum
3.13.4 | Low battery (alarm must sound when there is still approximately 10 minutes battery operation available).
3.13.5 | Open door
3.13.6 | Infusion set removed
3.13.7 | Wrong infusion set detection
3.13.8 | Completion of selected volume to be infused
3.13.9 | Zero infusion rate
3.13.10 | Zero volume selected for infusion
3.13.11 | Internal malfunction
3.14 | All alarm conditions must be accompanied by an audible tone.
3.15 | Free flow protection must be available
3.16 | Dedicated neonatal infusion sets must be used to achieve volumetric accuracy.
3.16.1 | In case of universal sets, the packaging must be clearly labeled that they cover full patient range.
3.16.2 | Generic infusion sets must be OEM approved. A report from the OEM must be attached supporting claim of approval.
3.17 | The unit must have a memory, which stores alarm messages that can be recalled by service technicians when carrying out preventative maintenance, repairs or servicing.
3.18 | The unit must have a service mode, to be accessed by service technicians to check important parameters of the unit, without having to first strip the unit for checking/ testing.
3.19 | The casing of the unit must be impact resistant at 1m height.
3.20 | The unit must be provided with a mechanism for attaching to either a mobile drip stand or a gabler rail. NB: whichever is requested by the hospital/institution.
3.21 | The 220 V, 50 Hz AC supply to the unit must be fused in live or both live and neutral.
3.22 | The unit must be fitted with an internal rechargeable battery.
3.23 | In the event of a 220 V mains failure the battery must automatically provide continuity of operation.
3.24 | The battery must have capacity to operate continuously for a minimum of 3 hours at an infusion rate of 125 ml/h.
3.25 | The battery charger for the internal rechargeable battery must be an integral part of the infusion pump.
3.26 | The infusion pump must be provided with circuitry to ensure that the internal rechargeable battery will be protected against over-charging and over-discharging.
3.27 | The infusion pump must be fully operable from the mains supply regardless of the condition of the battery
3.28 | The battery must be automatically charged when the unit is connected to the mains supply.
3.29 | Purge rate (1 - 999 ml/h) Purge volume (1 - 99 ml) (Sci Gr)
3.30 | Unit can be mounted on standard bed/wall rail or mobile pole/stand
3.31 | Suitable for all intravenous infusions of fluids
3.32 | Display/alarm for low/high flow, occlusion, open door, end-of-infusion and built-in battery status
3.33 | Audio visual alarm with silencing feature
3.34 | Stand-by mode when not in use
3.35 | The unit must have mounting bracket for standard bed/wall rail or mobile pole/stand
3.36 | The unit must be handed over in full operating order
3.37 | It must not be possible to run when the infusion set is inserted incorrectly
3.38 | This neonatal dedicated pump must be clearly marked NEONATAL or have a clear distinction from an adult pump
4 | STANDARD ACCESSORIES
4.1 | Clear fluids (Primary administration set) minimum 3 needle free access ports (100/box)
5 | Conformity Compliance (Please attach certificate)
5.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
5.2 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
5.2.1 | Please provide unique ref number of the ISO 13485 ceriticate:
5.3 | Model on offer must comply with EU Medical Device Regulation (EU) 2017/745 or equvalent from SAHPRA recognized jurisdictions . Attach a copy of certification
5.3.1 | Please provide unique ref number of the EC ceriticate:
5.4 | Hazardous Substance Act: If this type of equipment / apparatus appears on the schedule of Hazardous substances issued by the Directorate: Health Technology of the Department of Health, a license in terms of the Act on Hazardous Substances (Act 15/1973) must be submitted with the bid / quotation documents. The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.5 | All electrical/electronic medical equipment must be licensed by Radiation Control, where applicable a copy of the license must be submitted The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.5.1 | Radiation Control License number:
5.6 | The unit shall operate from 220V AC/50Hz mains supply. The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade (rubber to with 2 screws) type and it must be a minimum length of three (3) meters. NB the mains cable of the unit tendered for must be SABS colour coded.
5.7 | The equipment quoted for must be protected against electromagnetic interference IEC 60601-1-2
5.8 | The unit must comply with electrical safety and performance standard such as IEC 60601-2-24 for Infusion Pumps, attached certification
5.8.1 | Please provide unique ref number of the IEC ceriticate:
6 | COMPREHENSIVE MAINTENANCE
6.1 | A 24-month warranty inclusive maintenance and service (Year 1 to Year 2) should be quoted for, as per the SCC document
6.1.1 | THIS ENTAILS: A 24-month warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 24-month warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 24-month warranty will commence after formal Commissioning and handover of the equipment.
6.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
6.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
6.3 | The bidder to provide the OEM recommended service interval of the unit:
6.4 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
7 | General
7.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
7.2 | The unit on offer must be the latest technology.
7.2.1 | Initial Manufacturing year of Model:
RT2-05-011 | INFUSION PUMP ACCESSORIES and CONSUMABLES. | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point
RT2-05-011a | Clear fluids (Primary administration set) minimum 3 needle free access ports (100/box)
RT2-05-011b | Blood administration sets (100/box)
RT2-05-011c | TPN sets (with filter) -100/box
RT2-05-011d | Light resistant (sensetive) sets (100/box)
RT2-05-011e | Low sorb sets (100/box)
RT2-05-011f | Burette sets (100/box)
RT2-05-011g | Secondary infusion sets (Rider/ Piggy bag set) - 100/box
RT2-05-011h | Triport/Tripot needle free access
RT2-05-011i | Dosing error software (UPGRADE)
RT2-05-011j | Capablity for labeling drug names for display on the pump (UPGRADE)
RT2-05-011k | Docking/stackable station - 4 slots/tiles/ports
RT2-05-011l | Docking/stackable station - 6 slots/tiles/ports
RT2-05-011m | Docking/stackable station - 8 slots/tiles/ports
RT2-05-011n | Rechargable battery

## Infusion Pump, Peads-Adult
RT2-05-012 | INFUSION PUMP, Paeds/Adult | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
1 | Description
Volumetric infusion pump dedicated for paediatric and adult patients
2 | Product Offered:
2.1 | Make:
2.2 | Model:
2.3 | Initial Manufacturing year of Model:
2.4 | Radiation Control License number:
3 | Technical Specification
3.1 | The unit must use a linear peristaltic drive mechanism. Other drive mechanisms may be offered details to be provided.
3.2 | The unit must provide user selectable infusion rate from 1 to 999 ml/h, selectable in 1 ml/h increments.
3.3 | The unit must provide user selectable infusion volume from 1 to 9 999 ml, selectable in 1 ml increments.
3.4 | The unit must deliver the selected volume with an accuracy better than 5% throughout the whole range of operation/infusion.
3.5 | Bolus rate, user selectable: 10 - 999 ml/hr
3.6 | Bolus volume, user selectable: 1 - 1000 ml
3.7 | The unit must also be capable of carrying out blood administration.
3.8 | It must not be possible to change the rate while the infusion is in progress, without confirming change.
3.9 | The unit must have an LCD display
3.10 | The following must be clearly displayed on the front panel under all lighting conditions:
3.10.1 | Pump running
3.10.2 | AC mains supply operation
3.10.3 | AC power failure/battery operation
3.10.4 | Infusion rate selected
3.10.5 | Volume selected to be infused
3.10.6 | Volume infused
3.10.7 | Alarm condition and alarm/error messages
3.11 | When an infusion is completed an audible tone must sound.
3.12 | On completion of an infusion there must be a Keep Vein Open (KVO) rate of 1 - 5 ml/h.
3.13 | The unit must provide the following alarms:
3.13.1 | Air in line detection
3.13.2 | Occlusion during infusion (downstream and upstream)
3.13.3 | Occlusion pressure alarms
3.13.3a | Occlusion pressure alarm user selectable (Paeds): 150 - 250 mmHg
3.13.3b | Occlusion pressure alarm user selectable (Adult): 250 - 900 mmHg
3.13.4 | Low battery (alarm must sound when there is still approximately 10 minutes battery operation available).
3.13.5 | Open door
3.13.6 | Infusion set removed
3.13.7 | Wrong infusion set detection
3.13.8 | Completion of selected volume to be infused
3.13.9 | Zero infusion rate
3.13.10 | Zero volume selected for infusion
3.13.11 | Internal malfunction
3.14 | All alarm conditions must be accompanied by an audible tone.
3.15 | Free flow protection must be available
3.16 | State clearly if a dedicated infusion set is required to achieve volumetric accuracy.
3.17 | The unit must have a memory, which stores alarm messages that can be recalled by service technicians when carrying out preventative maintenance, repairs or servicing.
3.18 | The unit must have a service mode, to be accessed by service technicians to check important parameters of the unit, without having to first strip the unit for checking/ testing.
3.19 | The casing of the unit must be impact resistant at 1m height.
3.20 | The unit must be provided with a mechanism for attaching to either a mobile drip stand or a gabler rail. NB: whichever is requested by the hospital/institution.
3.21 | The 220 V, 50 Hz AC supply to the unit must be fused in the live or both live and neutral.
3.22 | The unit must be fitted with an internal rechargeable battery.
3.23 | In the event of a 220 V mains failure the battery must automatically provide continuity of operation.
3.24 | The battery must have capacity to operate continuously for a minimum of 3 hours at an infusion rate of 125 ml/h.
3.25 | The battery charger for the internal rechargeable battery must be an integral part of the infusion pump.
3.26 | The infusion pump must be provided with circuitry to ensure that the internal rechargeable battery will be protected against over-charging and over-discharging.
3.27 | The infusion pump must be fully operable from the mains supply regardless of the condition of the battery
3.28 | The battery must be automatically charged when the unit is connected to the mains supply.
3.29 | Purge rate (1 - 999 ml/h) Purge volume (1 - 99 ml) (Sci Gr)
3.30 | Unit can be mounted on standard bed/wall rail or mobile pole/stand
3.31 | Suitable for all intravenous infusions of fluids
3.32 | Display/alarm for low/high flow, occlusion, open door, end-of-infusion and built-in battery status
3.33 | Audio visual alarm with silencing feature
3.34 | Stand-by mode when not in use
3.35 | The unit must have mounting bracket for standard bed/wall rail or mobile pole/stand
3.36 | The unit must be handed over in full operating order
3.37 | It must not be possible to run when the infusion set is inserted incorrectly
4 | STANDARD ACCESSORIES
4.1 | Clear fluids (Primary administration set) minimum 3 needle free access ports (100/box)
5 | Conformity Compliance (Please attach certificate)
5.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
5.2 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
5.2.1 | Please provide unique ref number of the ISO 13485 ceriticate:
5.3 | Model on offer must comply with EU Medical Device Regulation (EU) 2017/745 or equvalent from SAHPRA recognized jurisdictions . Attach a copy of certification
5.3.1 | Please provide unique ref number of the EC ceriticate:
5.4 | Hazardous Substance Act: If this type of equipment / apparatus appears on the schedule of Hazardous substances issued by the Directorate: Health Technology of the Department of Health, a license in terms of the Act on Hazardous Substances (Act 15/1973) must be submitted with the bid / quotation documents. The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.5 | All electrical/electronic medical equipment must be licensed by Radiation Control, where applicable a copy of the license must be submitted The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.5.1 | Radiation Control License number:
5.6 | The unit shall operate from 220V AC/50Hz mains supply. The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade (rubber to with 2 screws) type and it must be a minimum length of three (3) meters. NB the mains cable of the unit tendered for must be SABS colour coded.
5.7 | The equipment quoted for must be protected against electromagnetic interference IEC 60601-1-2
5.8 | The unit must comply with electrical safety and performance standard such as IEC 60601-2-24 for Infusion Pumps, attached certification
5.8.1 | Please provide unique ref number of the IEC ceriticate:
6 | COMPREHENSIVE MAINTENANCE
6.1 | A 24-month warranty inclusive maintenance and service (Year 1 to Year 2) should be quoted for, as per the SCC document
6.1.1 | THIS ENTAILS: A 24-month warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 24-month warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 24-month warranty will commence after formal Commissioning and handover of the equipment.
6.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
6.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
6.3 | The bidder to provide the OEM recommended service interval of the unit:
6.4 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
7 | General
7.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
7.2 | The unit on offer must be the latest technology.
7.2.1 | Initial Manufacturing year of Model:
RT2-05-012 | INFUSION PUMP ACCESSORIES and CONSUMABLES. | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point
RT2-05-012a | Clear fluids (Primary administration set) minimum 3 needle free access ports (100/box)
RT2-05-012b | Blood administration sets (100/box)
RT2-05-012c | TPN sets (with filter) -100/box
RT2-05-012d | Light resistant (sensetive) sets (100/box)
RT2-05-012e | Low sorb sets (100/box)
RT2-05-012f | Burette sets (100/box)
RT2-05-012g | Secondary infusion sets (Rider/ Piggy bag set) - 100/box
RT2-05-012h | Triport/Tripot needle free access
RT2-05-012i | Dosing error software (UPGRADE)
RT2-05-012j | Capablity for labeling drug names for display on the pump (UPGRADE)
RT2-05-012k | Docking/stackable station - 4 slots/tiles/ports
RT2-05-012l | Docking/stackable station - 6 slots/tiles/ports
RT2-05-012m | Docking/stackable station - 8 slots/tiles/ports
RT2-05-012n | Rechargable battery

## Monitor, Basic
RT2-05-013 | MONITOR, BASIC | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
1 | Description
1.1 | A medical device that continuously tracks and displays a patient's vital signs and other physiological parameters
2 | Product Offered:
2.1 | Make:
2.2 | Model:
2.3 | Initial Manufacturing year of Model:
2.4 | Radiation Control License number:
3 | Technical Specification
3.1 | A state of the art technology monitor, suitable for neonates to adult including mounting solution suitable for the facility
3.2 | A microprocessor controlled device with automatic self-test and calibration
3.3 | The following parameters must be measured and displayed :
3.3.1 | ECG waveform
3.3.2 | Oxygen saturation: waveform + numerical
3.3.3 | Heart rate: numerical
3.3.4 | Respiration: waveform + numerical
3.3.5 | NIBP: numerical
3.3.6 | Temperature: numerical
3.4 | The monitor must operate off both 220 V, 50 Hz single phase AC power supply and an internal rechargeable battery.
3.5 | If mains power fails the monitor must operate off battery power for at least 2 hours.
3.6 | Visible indication to show whether the monitor is operating off the mains or battery power.
3.7 | All control keys must be splashproof and resistant to the entry of fluids into the unit.
3.8 | The front panel must be easy to clean and disinfect.
3.9 | The monitor must have a carry handle.
3.10 | High resolution LCD colour display, at least 13" diagonal
3.10.1 | Digits: large: 10 mm, small: 5 mm
3.10.2 | Sweep speeds : 25 and 50 mm/s
3.10.3 | It must be possible to freeze the display. This must not affect the QRS signal.
3.11 | Trending for 24 hours must be possible on the monitor
3.12 | Defibrillator-proof up to 360 Joules
3.13 | Protected against HF interference
3.14 | Patient isolated from all potentially harmful circuits
3.15 | ECG
3.15.1 | 3 lead ECG with arrhythmia detection
3.15.2 | The amplitude of the ECG trace must be adjustable from 1 mV = 5 mm to 1 mV = 20 mm.
3.15.3 | Digital display of heart rate : 30 - 300 bpm
3.15.4 | Accuracy : 3 bpm
3.15.5 | Resolution : 1 bpm
3.15.6 | Differential input with CMRR > 80 dB
3.15.7 | Input impedance > 5 MOhm
3.15.8 | Respiration monitoring through ECG electrodes
3.15.9 | Pacemaker detection/rejection : 2 to 500 mV; 0,1 to 2 ms
3.16 | OXYGEN SATURATION
3.16.1 | It must be the latest technology with clinical research support based on accuracy and real-time monitoring. Please submit the published peer reviewed/researched paper as supporting documents.
3.16.2 | The unit should indicate whether the signal from the pulse oximeter sensor is adequate for reliable measurements. If the signal strength or quality is inadequate, this must be clearly displayed.
3.16.3 | An audible bleep must sound for each heart beat. The pitch must vary with the oxygen saturation.
3.16.4 | Measurement range : 0 - 100%
3.16.5 | Accuracy:
3.16.5a | 80 - 100% : 1,5%
3.16.5a | 60 - 80% : 3,5%
3.16.6 | Continuous waveform display
3.16.7 | Digital display of SpO2%
3.16.8 | The sensor lead must be at least 3 metres long.
3.16.9 | Automatic calibration for skin colour
3.16.10 | Pulse rate: 25 - 250 bpm
3.17 | NON-INVASIVE BLOOD PRESSURE
3.17.1 | Oscillometric measurement method
3.17.2 | Automatic, manual and stat modes
3.17.3 | Measure and display :
3.17.3a | - systolic pressure
3.17.3b | - diastolic pressure
3.17.3c | - mean pressure
3.17.4 | The interval between measurements must be selectable from 1 to 60 minutes.
3.17.5 | Ranges (approximate) :
3.17.5a | Systolic : 30 - 270 mmHg. Resolution : 1 mmHg
3.17.5b | Diastolic : 10 - 150 mmHg. Resolution : 1 mmHg
3.17.5c | Mean : 20 - 250 mmHg. Resolution : 1 mmHg
3.17.6 | Cuff pressure : 0 - 280 mmHg
3.17.7 | Cuff inflation time < 10 s
3.17.8 | Automatic deflation if measurement exceeds 3 min or if mains power fails
3.18 | TEMPERATURE
3.18.1 | Range : 10 - 45 ˚C
3.18.2 | Digital display
3.18.3 | Accuracy: 0.3 ˚C
3.18.4 | Resolution : 0.1 ˚C
3.19 | ALARMS
3.19.1 | Visible and audio alarms
3.19.2 | The monitor must automatically set default alarms when switched on.
3.19.3 | The alarms must be programmable
3.19.4 | Alarm limits must be displayed with the parameter digit
3.19.5 | Heart rate alarm : 30 - 300 bpm
3.19.6 | Asystole alarm
3.19.7 | Arrhythmia alarms
3.19.8 | Apnea alarm
3.19.9 | Alarms for faulty/not-connected electrode/probe
3.19.10 | User selectable alarm limits for systolic, diastolic and mean pressures
3.19.11 | Alarm for incorrect NIBPcuff connection
3.19.12 | Adjustable SpO2 high and low alarms:
3.19.12a | High : 70 - 100%
3.19.12b | Low : 50 - 95%
3.19.13 | Adjustable high and low temperature alarms
3.19.14 | Violation of limits must result in both audible and visible alarms and indicate which parameter is affected.
3.19.15 | The sound level of the audible alarm must be adjustable.
3.19.16 | It must be possible to temporarily silence the audible alarm but not remove the visible alarm.
3.19.17 | Duration of the silenced period: 2 minutes.
3.19.18 | Low battery alarm
3.20 | GENERAL
3.20.3 | All cables and sockets must be colour coded, labelled or otherwise prevented from being incorrectly coupled
3.20.4 | The unit must be handed over in full operating order
3.20.5 | A starter-pack of consumables must be supplied with the unit, as per end-user requirement
4 | Standard Accessories and Consumables
4.1 | Neonatal Pack
4.1.1 | Neonatal NIBP hose
4.1.2 | Disposable neonate cuff - pack of 20
4.1.3 | Spo2 interface cable
4.1.4 | Finger wrap-around probe for neonates - Disposable pack of 20
4.1.5 | Ankle wrap-around probe for neonates - Disposable pack of 20
4.1.6 | Reusable temperature skin probe
4.1.7 | ECG interface cable
4.1.8 | Paediatric 3 lead ECG cable
AND/OR
4.2 | Paediatric Pack
4.2.1 | Adult/Paeds NIBP hose
4.2.2 | Reusable latex-free paediatric cuff
4.2.3 | ECG interface cable
4.2.4 | Paediatric 3 lead ECG cable
4.2.5 | Spo2 interface cable
4.2.6 | Finger sensor for paeds
4.2.7 | Reusable Paeds temperature probe
AND/OR
4.3 | Adult Pack
4.3.1 | NIBP hose
4.3.2 | Reusable latex-free adult cuff
4.3.3 | Reusable latex-free obese cuff
4.3.4 | Spo2 interface cable
4.3.5 | Finger sensor for adults
4.3.6 | Adult 3 lead ECG cable
4.3.7 | ECG interface cable
4.3.8 | Reusable Adult temperature probe
AND
4.4 | Rechargeable Battery
4.5 | Wall bracket or Roll stand as per end-user requirement
5 | Conformity Compliance (Please attach certificate)
5.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
5.2 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
5.2.1 | Please provide unique ref number of the ISO 13485 ceriticate:
5.3 | Model on offer must comply with EU Medical Device Regulation (EU) 2017/745 or equvalent from SAHPRA recognized jurisdictions . Attach a copy of certification
5.3.1 | Please provide unique ref number of the EC ceriticate:
5.4 | Hazardous Substance Act: If this type of equipment / apparatus appears on the schedule of Hazardous substances issued by the Directorate: Health Technology of the Department of Health, a license in terms of the Act on Hazardous Substances (Act 15/1973) must be submitted with the bid / quotation documents. The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.5 | All electrical/electronic medical equipment must be licensed by Radiation Control, where applicable a copy of the license must be submitted The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.5.1 | Radiation Control License number:
5.6 | The unit shall operate from 220V AC/50Hz mains supply. The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade (rubber to with 2 screws) type and it must be a minimum length of three (3) meters. NB the mains cable of the unit tendered for must be SABS colour coded.
5.7 | The equipment quoted for must be protected against electromagnetic interference IEC 60601-1-2
5.8 | The unit must comply with electrical safety and performance standard such as IEC 60601-2-27 for Electrocardiographic Monitoring, attached certification
5.8.1 | Please provide unique ref number of the IEC ceriticate:
5.9 | The unit must comply with electrical safety and performance standard such as IEC 60601-2-30 for Automated Non-invasive Sphygmomanometers, attached certification
5.9.1 | Please provide unique ref number of the IEC ceriticate:
5.10 | The unit must comply with electrical safety and performance standard such as IEC 60601-2-49 for Patient Monitoring, attached certification
5.10.1 | Please provide unique ref number of the IEC ceriticate:
6 | COMPREHENSIVE MAINTENANCE
6.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) should be quoted for, as per the SCC document
6.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
6.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
6.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
6.3 | The bidder to provide the OEM recommended service interval of the unit:
6.4 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
7 | General
7.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
7.2 | The unit on offer must be the latest technology.
7.2.1 | Initial Manufacturing year of Model:
RT2-05-013 | BASIC MONITOR ACCESSORIES and CONSUMABLES | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point
RT2-05-013a | Paediatric 3 lead ECG cable
RT2-05-013b | Paediatric 5 lead ECG cable
RT2-05-013c | Adult 3 lead ECG cable
RT2-05-013d | Adult 5 lead ECG cable
RT2-05-013e | ECG interface cable
RT2-05-013f | Reusable latex-free paediatric cuff
RT2-05-013g | Reusable latex-free adult cuff
RT2-05-013h | Reusable latex-free neonate cuff
RT2-05-013i | Reusable latex-free obese cuff
RT2-05-013j | Adult/Paeds NIBP hose
RT2-05-013k | Neonatal NIBP hose
RT2-05-013l | Finger sensor for paeds
RT2-05-013m | Finger sensor for adults
RT2-05-013n | Finger wrap-around probe for neonates
RT2-05-013o | Ankle wrap-around probe for neonates
RT2-05-013p | Spo2 interface cable
RT2-05-013q | Reusable Adult temperature probe
RT2-05-013r | Reusable Paeds temperature probe
RT2-05-013s | Monitor Front bezel
RT2-05-013t | Rechargeable Battery
RT2-05-013u | NIBP pump module
RT2-05-013v | Wall bracket, height adjustable on which the monitor can be mounted.
RT2-05-013w | Roll stand
RT2-05-013x | Finger wrap-around probe for neonates - disposable
RT2-05-013y | Ankle wrap-around probe for neonates - disposable
RT2-05-013z | Disposable neonate cuff
RT2-05-013aa | Adult Esophageal/Rectal temperature probe
RT2-05-013ab | Paediatric Esophageal/Rectal temperature probe

## Monitor, ICU
RT2-05-014 | MONITOR, ICU | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
1 | Description
1.1 | A medical device that continuously tracks and displays a patient's vital signs and other physiological parameters for use in intensive care settings
2 | Product Offered:
2.1 | Make:
2.2 | Model:
2.3 | Initial Manufacturing year of Model:
2.4 | Radiation Control License number:
3 | Technical Specification
3.1 | The monitor must be compatible with various brands of Central Station monitoring. Bidder to list compatible brands
3.2 | Modular, state of the art technology, suitable for neonates to adults
3.3 | Microprocessor controlled with automatic self-test and calibration
3.4 | The following parameters must be measured and displayed :
3.4.1 | ECG waveform
3.4.2 | Oxygen saturation: waveform + numerical
3.4.3 | Heart rate: numerical
3.4.4 | Respiration: waveform + numerical
3.4.5 | 3X Invasive pressure: waveform + numerical (systolic, diastolic, mean)
3.4.6 | NIBP: numerical (systolic, diastolic, mean)
3.4.7 | Temperature: numerical
3.4.8 | Side-stream Capnography highly desireable. Provide full information
3.5 | The monitor must operate off both 220 V, 50 Hz single phase AC power supply and an internal rechargeable battery.
3.6 | If mains power fails the monitor must operate off battery power for at least 2 hours.
3.7 | Visible indication to show whether the monitor is operating from the mains or battery power.
3.8 | All control keys must be splashproof and resistant to the entry of fluids into the unit.
3.9 | The front panel must be easy to clean and disinfect.
3.10 | High resolution LCD colour display, 45 cm diagonal minimum
3.11 | At least 6 waveforms must be displayed
3.11.1 | Digits: large: 10 mm, small: 5 mm
3.11.2 | Sweep speeds : 25 and 50 mm/s
3.11.3 | It must be possible to freeze the display. This must not affect the QRS signal.
3.12 | Trending for 72 hours must be possible on the monitor,
3.13 | Unit to provide at least 72 hours of full disclosure wave data for review
3.14 | Defibrillator-proof up to 360 Joules
3.15 | Protected against HF interference
3.16 | Patient isolated from all potentially harmful circuits
3.17 | ECG
3.17.2 | ECG 12 lead with arrhythmia detection
3.17.3 | The amplitude of the ECG trace must be adjustable from 1 mV = 5 mm to 1 mV = 20 mm.
3.17.4 | Digital display of heart rate : 30 - 300 bpm
3.17.5 | Accuracy : 3 bpm
3.17.6 | Resolution : 1 bpm
3.17.7 | Differential input with CMRR > 80 dB
3.17.8 | Input impedance > 5 MOhm
3.17.9 | Respiration monitoring through ECG electrodes
3.17.10 | Pacemaker detection/rejection : 2 to 500 mV; 0,1 to 2 ms
3.17.11 | Synch pulse for defibrillator
3.18 | OXYGEN SATURATION
3.18.1 | It must be the latest technology with clinical research support based on accuracy and real-time monitoring. Please submit the published peer reviewed/researched paper as supporting documents.
3.18.2 | The unit should indicate whether the signal from the pulse oximeter sensor is adequate for reliable measurements. If the signal strength or quality is inadequate, this must be clearly displayed.
3.18.3 | An audible bleep must sound for each heart beat. The pitch must vary with the oxygen saturation.
3.18.4 | Measurement range : 0 - 100%
3.18.5 | Accuracy:
3.18.5a | 80 - 100% : 1,5%
3.18.5b | 60 - 80% : 3,5%
3.18.6 | Continuous waveform display
3.18.7 | Digital display of SpO2%
3.18.8 | The sensor lead must be at least 3 metres long.
3.18.9 | Automatic calibration for skin colour
3.18.10 | Pulse rate: 25 - 250 bpm
3.19 | NON-INVASIVE BLOOD PRESSURE
3.19.1 | Oscillometric measurement method
3.19.2 | Automatic, manual and stat modes
3.19.3 | Measure and display :
3.19.3a | - systolic pressure
3.19.3b | - diastolic pressure
3.19.3c | - mean pressure
3.19.5 | The interval between measurements must be selectable from 1 to 60 minutes.
3.19.6 | Ranges (approximate) :
3.19.6a | Systolic : 30 - 270 mmHg. Resolution : 1 mmHg
3.19.6b | Diastolic : 10 - 150 mmHg. Resolution : 1 mmHg
3.19.6c | Mean : 20 - 250 mmHg. Resolution : 1 mmHg
3.19.7 | Cuff pressure : 0 - 280 mmHg
3.19.8 | Cuff inflation time < 10 s
3.19.9 | Automatic deflation if measurement exceeds 3 min or if mains power fails
3.20 | INVASIVE BLOOD PRESSURE
3.20.2 | Measure and display :
3.20.2a | - systolic pressure
3.20.2b | - diastolic pressure
3.20.2c | - mean pressure
3.20.3 | Systolic : 0 - 300 mmHg. Resolution : 1 mmHg
3.20.4 | Diastolic : 0 - 250 mmHg. Resolution : 1 mmHg
3.20.5 | Mean : 0 - 250 mmHg. Resolution : 1 mmHg
3.21 | TEMPERATURE
3.21.1 | Range : 10 - 45 ˚C
3.21.2 | Digital display
3.21.3 | Accuracy: 0.3 ˚C
3.21.4 | Resolution : 0.1 ˚C
3.22 | ALARMS
3.22.1 | Visible and audio alarms
3.22.2 | The monitor must automatically set default alarms when switched on.
3.22.3 | The alarms must be programmable
3.22.4 | Alarm limits must be displayed with the parameter digit
3.22.5 | Heart rate alarm : 30 - 300 bpm
3.22.6 | Asystole alarm
3.22.7 | Arrhythmia alarms
3.22.8 | Apnea alarm
3.22.9 | Alarms for faulty/not-connected electrode/probe
3.22.10 | User selectable alarm limits for systolic, diastolic and mean pressures
3.22.11 | Alarm for incorrect NIBPcuff connection
3.22.12 | Adjustable SpO2 high and low alarms:
3.22.12a | High : 70 - 100%
3.22.12b | Low : 50 - 95%
3.22.13 | Adjustable high and low temperature alarms
3.22.14 | Violation of limits must result in both audible and visible alarms and indicate which parameter is affected.
3.22.15 | The sound level of the audible alarm must be adjustable.
3.22.16 | It must be possible to temporarily silence the audible alarm but not remove the visible alarm.
3.22.17 | Duration of the silenced period: 2 minutes.
3.22.18 | Low battery alarm
3.23 | GENERAL
3.23.2 | All cables and sockets must be colour coded, labelled or otherwise prevented from being incorrectly coupled
3.23.3 | The unit must be handed over in full operating order
3.23.4 | A starter-pack of consumables must be supplied with the unit, as per end-user requirement
4 | Standard Accessories and Consumables
4.1 | Neonatal Pack
4.1.1 | Neonatal NIBP hose
4.1.2 | Disposable neonate cuff - pack of 20
4.1.3 | Spo2 interface cable
4.1.4 | Finger wrap-around probe for neonates - Disposable pack of 20
4.1.5 | Ankle wrap-around probe for neonates - Disposable pack of 20
4.1.6 | Reusable temperature skin probe
4.1.7 | ECG interface cable
4.1.8 | Paediatric 3 lead ECG cable
AND/OR
4.2 | Paediatric Pack
4.2.1 | Adult/Paeds NIBP hose
4.2.2 | Reusable latex-free paediatric cuff
4.2.3 | ECG interface cable
4.2.4 | Paediatric 3 lead ECG cable
4.2.5 | Spo2 interface cable
4.2.6 | Finger sensor for paeds
4.2.7 | Reusable Paeds temperature probe
AND/OR
4.3 | Adult Pack
4.3.1 | NIBP hose
4.3.2 | Reusable latex-free adult cuff
4.3.3 | Reusable latex-free obese cuff
4.3.4 | Spo2 interface cable
4.3.5 | Finger sensor for adults
4.3.6 | Adult 3 lead ECG cable
4.3.7 | ECG interface cable
4.3.8 | Reusable Adult temperature probe
AND
4.4 | Invasive blood pressure cable X2
4.5 | Invasive blood pressure tranceducers X2
4.6 | Rechargeable Battery
4.7 | Wall bracket or Roll stand as per end-user requirement
5 | Conformity Compliance (Please attach certificate)
5.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
5.2 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
5.2.1 | Please provide unique ref number of the ISO 13485 ceriticate:
5.3 | Model on offer must comply with EU Medical Device Regulation (EU) 2017/745 or equvalent from SAHPRA recognized jurisdictions . Attach a copy of certification
5.3.1 | Please provide unique ref number of the EC ceriticate:
5.4 | Hazardous Substance Act: If this type of equipment / apparatus appears on the schedule of Hazardous substances issued by the Directorate: Health Technology of the Department of Health, a license in terms of the Act on Hazardous Substances (Act 15/1973) must be submitted with the bid / quotation documents. The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.5 | All electrical/electronic medical equipment must be licensed by Radiation Control, where applicable a copy of the license must be submitted The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.5.1 | Radiation Control License number:
5.6 | The unit shall operate from 220V AC/50Hz mains supply. The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade (rubber to with 2 screws) type and it must be a minimum length of three (3) meters. NB the mains cable of the unit tendered for must be SABS colour coded.
5.7 | The equipment quoted for must be protected against electromagnetic interference IEC 60601-1-2
5.8 | The unit must comply with electrical safety and performance standard such as IEC 60601-2-27 for Electrocardiographic Monitoring, attached certification
5.8.1 | Please provide unique ref number of the IEC ceriticate:
5.9 | The unit must comply with electrical safety and performance standard such as IEC 60601-2-30 for Automated Non-invasive Sphygmomanometers, attached certification
5.9.1 | Please provide unique ref number of the IEC ceriticate:
5.10 | The unit must comply with electrical safety and performance standard such as IEC 60601-2-34 for Invasive Blood Pressure monitoring, attached certification
5.10.1 | Please provide unique ref number of the IEC ceriticate:
5.11 | The unit must comply with electrical safety and performance standard such as IEC 60601-2-49 for Patient Monitoring, attached certification
5.11.1 | Please provide unique ref number of the IEC ceriticate:
6 | COMPREHENSIVE MAINTENANCE
6.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) should be quoted for, as per the SCC document
6.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
6.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
6.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
6.3 | The bidder to provide the OEM recommended service interval of the unit:
6.4 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
7 | General
7.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
7.2 | The unit on offer must be the latest technology.
7.2.1 | Initial Manufacturing year of Model:
RT2-05-014 | ICU MONITOR ACCESSORIES and CONSUMABLES | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point
RT2-05-014a | Paediatric 3 lead ECG cable
RT2-05-014b | Paediatric 5 lead ECG cable
RT2-05-014c | Adult 3 lead ECG cable
RT2-05-014d | Adult 5 lead ECG cable
RT2-05-014e | ECG interface cable
RT2-05-014f | Reusable latex-free paediatric cuff
RT2-05-014g | Reusable latex-free adult cuff
RT2-05-014h | Reusable latex-free neonate cuff
RT2-05-014i | Reusable latex-free obese cuff
RT2-05-014j | Adult/Paeds NIBP hose
RT2-05-014k | Neonatal NIBP hose
RT2-05-014l | Finger sensor for paeds
RT2-05-014m | Finger sensor for adults
RT2-05-014n | Finger wrap-around probe for neonates
RT2-05-014o | Ankle wrap-around probe for neonates
RT2-05-014p | Spo2 interface cable
RT2-05-014q | Reusable Adult temperature probe
RT2-05-014r | Reusable Paeds temperature probe
RT2-05-014s | Monitor Front bezel
RT2-05-014t | Rechargeable Battery
RT2-05-014u | NIBP pump module
RT2-05-014v | Wall bracket, height adjustable on which the monitor can be mounted.
RT2-05-014w | Roll stand
RT2-05-014x | EEG (UPGRADE)
RT2-05-014y | Intra-cranial pressure (UPGRADE)
RT2-05-014z | Cardiac output (UPGRADE)
RT2-05-014aa | CO2/EtCO2 (UPGRADE)
RT2-05-014ab | Built-in Printer (UPGRADE)
RT2-05-014ac | Apnea monitor for neonatal use (UPGRADE)
RT2-05-014ad | Invasive blood pressure cable
RT2-05-014ae | Invasive blood pressure tranceducer
RT2-05-014af | Finger wrap-around probe for neonates - disposable
RT2-05-014ag | Ankle wrap-around probe for neonates - disposable
RT2-05-014ah | Disposable neonate cuff
RT2-05-014ai | Adult Esophageal/Rectal temperature probe
RT2-05-014aj | Paediatric Esophageal/Rectal temperature probe

## Monitor, ICU, wit Module
RT2-05-015 | MONITOR, ICU, with Transport Module | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
1 | Description
1.1 | A medical device that continuously tracks and displays a patient's vital signs and other physiological parameters for use in intensive care settings, with detachable module for patient transportation without losing trends
2 | Product Offered:
2.1 | Make:
2.2 | Model:
2.3 | Transport module Model:
2.4 | Initial Manufacturing year of Model:
2.5 | Radiation Control License number:
3 | Technical Specification
3.1 | The monitor must be compatible with various brands of Central Station monitoring. Bidder to list compatible brands
3.2 | Modular, state of the art technology, suitable for neonates to adults
3.3 | Microprocessor controlled with automatic self-test and calibration
3.4 | The following parameters must be measured and displayed :
3.4.1 | ECG waveform
3.4.2 | Oxygen saturation: waveform + numerical
3.4.3 | Heart rate: numerical
3.4.4 | Respiration, waveform + digital
3.4.5 | 3X Invasive pressure, waveform + digital (systolic, diastolic, mean)
3.4.6 | NIBP, digital (systolic, diastolic, mean)
3.4.7 | Temperature, digital
3.4.8 | Side-stream Capnography highly desireable. Provide full information
3.5 | The monitor must operate off both 220 V, 50 Hz single phase AC power supply and an internal rechargeable battery.
3.6 | If mains power fails the monitor must operate off battery power for at least 2 hours.
3.7 | Visible indication to show whether the monitor is operating from the mains or battery power.
3.8 | All control keys must be splashproof and resistant to the entry of fluids into the unit.
3.9 | The front panel must be easy to clean and disinfect.
3.10 | High resolution LCD colour display, 45 cm diagonal minimum
3.11 | At least 6 waveforms must be displayed
3.11.1 | Digits: large: 10 mm, small: 5 mm
3.11.2 | Sweep speeds : 25 and 50 mm/s
3.11.3 | It must be possible to freeze the display. This must not affect the QRS signal.
3.12 | Trending for 72 hours must be possible on the monitor,
3.13 | Unit to provide at least 72 hours of full disclosure wave data for review
3.14 | Defibrillator-proof up to 360 Joules
3.15 | Protected against HF interference
3.16 | Patient isolated from all potentially harmful circuits
3.17 | The monitor must have a detachable module
3.17.1 | Module must have a minimum 10cm screen
3.17.2 | Parameter ports must be located on the module to avoid any reconnection for transportation
3.17.3 | Automatic display of parameters on the monitor screen once attached
3.17.4 | Module must have a battery capacity of at least 4 hours
3.18 | ECG
3.18.1 | ECG 12 lead with arrhythmia detection
3.18.2 | The amplitude of the ECG trace must be adjustable from 1 mV = 5 mm to 1 mV = 20 mm.
3.18.3 | Digital display of heart rate : 30 - 300 bpm
3.18.4 | Accuracy : 3 bpm
3.18.5 | Resolution : 1 bpm
3.18.6 | Differential input with CMRR > 80 dB
3.18.7 | Input impedance > 5 MOhm
3.18.8 | Respiration monitoring through ECG electrodes
3.18.9 | Pacemaker detection/rejection : 2 to 500 mV; 0,1 to 2 ms
3.18.10 | Synch pulse for defibrillator
3.19 | OXYGEN SATURATION
3.19.1 | It must be the latest technology with clinical research support based on accuracy and real-time monitoring. Please submit the published peer reviewed/researched paper as supporting documents.
3.19.2 | The unit should indicate whether the signal from the pulse oximeter sensor is adequate for reliable measurements. If the signal strength or quality is inadequate, this must be clearly displayed.
3.19.3 | An audible bleep must sound for each heart beat. The pitch must vary with the oxygen saturation.
3.19.4 | Measurement range : 0 - 100%
3.19.5 | Accuracy:
3.19.5a | 80 - 100% : 1,5%
3.19.5b | 60 - 80% : 3,5%
3.19.6 | Continuous waveform display
3.19.7 | Digital display of SpO2%
3.19.8 | The sensor lead must be at least 3 metres long.
3.19.9 | Automatic calibration for skin colour
3.19.10 | Pulse rate: 25 - 250 bpm
3.20 | NON-INVASIVE BLOOD PRESSURE
3.20.1 | Oscillometric measurement method
3.20.2 | Automatic, manual and stat modes
3.20.3 | Measure and display :
3.20.3a | - systolic pressure
3.20.3b | - diastolic pressure
3.20.3c | - mean pressure
3.20.4 | The interval between measurements must be selectable from 1 to 60 minutes.
3.20.5 | Ranges (approximate) :
3.20.5a | Systolic : 30 - 270 mmHg. Resolution : 1 mmHg
3.20.5b | Diastolic : 10 - 150 mmHg. Resolution : 1 mmHg
3.20.5c | Mean : 20 - 250 mmHg. Resolution : 1 mmHg
3.20.6 | Cuff pressure : 0 - 280 mmHg
3.20.7 | Cuff inflation time < 10 s
3.20.8 | Automatic deflation if measurement exceeds 3 min or if mains power fails
3.21 | INVASIVE BLOOD PRESSURE
3.21.1 | Measure and display :
3.21.1a | - systolic pressure
3.21.1b | - diastolic pressure
3.21.1c | - mean pressure
3.21.2 | Systolic : 0 - 300 mmHg. Resolution : 1 mmHg
3.21.3 | Diastolic : 0 - 250 mmHg. Resolution : 1 mmHg
3.21.4 | Mean : 0 - 250 mmHg. Resolution : 1 mmHg
3.22 | TEMPERATURE
3.22.1 | Range : 10 - 45 ˚C
3.22.2 | Digital display
3.22.3 | Accuracy: 0.3 ˚C
3.22.4 | Resolution : 0.1 ˚C
3.23 | ALARMS
3.23.1 | Visible and audio alarms
3.23.2 | The monitor must automatically set default alarms when switched on.
3.23.3 | The alarms must be programmable
3.23.4 | Alarm limits must be displayed with the parameter digit
3.23.5 | Heart rate alarm : 30 - 300 bpm
3.23.6 | Asystole alarm
3.23.7 | Arrhythmia alarms
3.23.8 | Apnea alarm
3.23.9 | Alarms for faulty/not-connected electrode/probe
3.23.10 | User selectable alarm limits for systolic, diastolic and mean pressures
3.23.11 | Alarm for incorrect NIBPcuff connection
3.23.12 | Adjustable SpO2 high and low alarms:
3.23.12a | High : 70 - 100%
3.23.12b | Low : 50 - 95%
3.23.13 | Adjustable high and low temperature alarms
3.23.14 | Violation of limits must result in both audible and visible alarms and indicate which parameter is affected.
3.23.15 | The sound level of the audible alarm must be adjustable.
3.23.16 | It must be possible to temporarily silence the audible alarm but not remove the visible alarm.
3.23.17 | Duration of the silenced period: 2 minutes.
3.23.18 | Low battery alarm
3.24 | GENERAL
3.24.1 | All cables and sockets must be colour coded, labelled or otherwise prevented from being incorrectly coupled
3.24.2 | The unit must be handed over in full operating order
3.24.3 | A starter-pack of consumables must be supplied with the unit, as per end-user requirement
4 | Standard Accessories and Consumables
4.1 | Neonatal Pack
4.1.1 | Neonatal NIBP hose
4.1.2 | Reusable latex-free neonate cuff
4.1.3 | Spo2 interface cable
4.1.4 | Finger wrap-around probe for neonates
4.1.5 | Ankle wrap-around probe for neonates
4.1.6 | Reusable temperature skin probe
4.1.7 | ECG interface cable
4.1.8 | Paediatric 3 lead ECG cable
AND/OR
4.2 | Paediatric Pack
4.2.1 | Adult/Paeds NIBP hose
4.2.2 | Reusable latex-free paediatric cuff
4.2.3 | ECG interface cable
4.2.4 | Paediatric 3 lead ECG cable
4.2.5 | Spo2 interface cable
4.2.6 | Finger sensor for paeds
4.2.7 | Reusable Paeds temperature probe
AND/OR
4.3 | Adult Pack
4.3.1 | NIBP hose
4.3.2 | Reusable latex-free adult cuff
4.3.3 | Reusable latex-free obese cuff
4.3.4 | Spo2 interface cable
4.3.5 | Finger sensor for adults
4.3.6 | Adult 3 lead ECG cable
4.3.7 | ECG interface cable
4.3.8 | Reusable Adult temperature probe
AND
4.4 | Invasive blood pressure cable X2
4.5 | Invasive blood pressure tranceducers X2
4.6 | Rechargeable Battery
4.7 | Wall bracket or Roll stand as per end-user requirement
5 | Conformity Compliance (Please attach certificate)
5.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
5.2 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
5.2.1 | Please provide unique ref number of the ISO 13485 ceriticate:
5.3 | Model on offer must comply with EU Medical Device Regulation (EU) 2017/745 or equvalent from SAHPRA recognized jurisdictions . Attach a copy of certification
5.3.1 | Please provide unique ref number of the EC ceriticate:
5.4 | Hazardous Substance Act: If this type of equipment / apparatus appears on the schedule of Hazardous substances issued by the Directorate: Health Technology of the Department of Health, a license in terms of the Act on Hazardous Substances (Act 15/1973) must be submitted with the bid / quotation documents. The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.5 | All electrical/electronic medical equipment must be licensed by Radiation Control, where applicable a copy of the license must be submitted The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.5.1 | Radiation Control License number:
5.6 | The unit shall operate from 220V AC/50Hz mains supply. The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade (rubber to with 2 screws) type and it must be a minimum length of three (3) meters. NB the mains cable of the unit tendered for must be SABS colour coded.
5.7 | The equipment quoted for must be protected against electromagnetic interference IEC 60601-1-2
5.8 | The unit must comply with electrical safety and performance standard such as IEC 60601-2-27 for Electrocardiographic Monitoring, attached certification
5.8.1 | Please provide unique ref number of the IEC ceriticate:
5.9 | The unit must comply with electrical safety and performance standard such as IEC 60601-2-30 for Automated Non-invasive Sphygmomanometers, attached certification
5.9.1 | Please provide unique ref number of the IEC ceriticate:
5.10 | The unit must comply with electrical safety and performance standard such as IEC 60601-2-34 for Invasive Blood Pressure monitoring, attached certification
5.10.1 | Please provide unique ref number of the IEC ceriticate:
5.11 | The unit must comply with electrical safety and performance standard such as IEC 60601-2-49 for Patient Monitoring, attached certification
5.11.1 | Please provide unique ref number of the IEC ceriticate:
6 | COMPREHENSIVE MAINTENANCE
6.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) should be quoted for, as per the SCC document
6.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
6.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
6.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
6.3 | The bidder to provide the OEM recommended service interval of the unit:
6.4 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
7 | General
7.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
7.2 | The unit on offer must be the latest technology.
7.2.1 | Initial Manufacturing year of Model:
RT2-05-015 | ICU MONITOR with Transport Module ACCESSORIES and CONSUMABLES | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point
RT2-05-015a | Paediatric 3 lead ECG cable
RT2-05-015b | Paediatric 5 lead ECG cable
RT2-05-015c | Adult 3 lead ECG cable
RT2-05-015d | Adult 5 lead ECG cable
RT2-05-015e | ECG interface cable
RT2-05-015f | Reusable latex-free paediatric cuff
RT2-05-015g | Reusable latex-free adult cuff
RT2-05-015h | Reusable latex-free neonate cuff
RT2-05-015i | Reusable latex-free obese cuff
RT2-05-015j | Adult/Paeds NIBP hose
RT2-05-015k | Neonatal NIBP hose
RT2-05-015l | Finger sensor for paeds
RT2-05-015m | Finger sensor for adults
RT2-05-015n | Finger wrap-around probe for neonates
RT2-05-015o | Ankle wrap-around probe for neonates
RT2-05-015p | Spo2 interface cable
RT2-05-015q | Reusable Adult temperature probe
RT2-05-015r | Reusable Paeds temperature probe
RT2-05-015s | Monitor Front bezel
RT2-05-015t | Rechargeable Battery
RT2-05-015u | NIBP pump module
RT2-05-015v | Wall bracket, height adjustable on which the monitor can be mounted.
RT2-05-015w | Roll stand
RT2-05-015x | EEG (UPGRADE)
RT2-05-015y | Intra-cranial pressure (UPGRADE)
RT2-05-015z | Cardiac output (UPGRADE)
RT2-05-015aa | CO2/EtCO2 (UPGRADE)
RT2-05-015ab | Built-in Printer (UPGRADE)
RT2-05-015ac | Apnea monitor for neonatal use (UPGRADE)
RT2-05-015ad | Invasive blood pressure cable
RT2-05-015ae | Invasive blood pressure tranceducer

## Central Monitoring, 8-bed
RT2-05-016 | Central Patient Monitoring, 8-Bed ICU Package | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
1 | Description
1.1 | A centralized monitoring system for continuous, real-time observation of patients' vital signs other physiological parameters for use in intensive care settings.
2.1 | Product Offered (Central Station):
2.1.1 | Make:
2.1.2 | Model:
2.1.3 | Brief Description:
2.2 | Product Offered (Monitor):
2.2.1 | Make:
2.2.2 | Model:
2.2.3 | Initial Manufacturing year of Model:
2.2.4 | Radiation Control License number:
3 | Technical Specification
3.1 | Central Station
3.1.1 | The central station must have a Single display, and 8 ICU monitors with complete accessories.
3.1.2 | The unit must be network configurable and wireless.
3.1.3 | Must have a viewing screen which will also do patient management and reviews and size must be at least 48 cm (19") colour LCD
3.1.4 | Touchscreen (Optional)
3.1.5 | The unit must be able to remotely manage patient bedside monitors, including viewing active or historic data.
3.1.6 | Password protection must be available to protect the critical menus and certain settings.
3.1.7 | All waveforms and alphanumeric displays must be available at the central station depending on the parameters of the bedside monitor
3.1.8 | Bi-directional communication between monitors and central station
3.1.9 | The central station must be able to download all the patient vitals after the patient has been disconnected and reconnected
3.1.10 | Display at least 8 beds information
3.1.11 | Full arrhythmia recall and ST segment analysis with 12 lead ECG interpretation
3.1.12 | Screen configurations must be able to be saved as favourites for easy access
3.1.13 | Up to 10 arrhythmia events per patient monitoring session must be available
3.1.14 | Must be able to review full disclosure on the central station for the duration of the disclosure time. At least 120 hours
3.1.15 | The central station must allow access to patient demographic information; information arrhythmias, ST-limit violations, and ECG samples in event review; parameter data in graphic and numeric trends; full disclosure data in strip view, page view with zoom, or in calipers.
3.1.16 | Alarms
3.1.16.1 | All alarms must be able to be muted only at the bedside
3.1.16.2 | 3-level alarm must be available audible & visual.
3.1.16.3 | The sound levels and colour to differentiate alarm type
3.1.16.4 | High and low limits of alarm must be offered
3.1.16.5 | Monitor alarm limits must be adjustable from the central station. This setting must be enabled/disabled via a password protected menu.
3.2 | Monitor
3.2.1 | Modular, state of the art technology, suitable for neonates to adult
3.2.2 | Microprocessor controlled with automatic self-test and calibration
3.2.3 | The following parameters must be measured and displayed :
3.2.3.1 | ECG waveform
3.2.3.2 | Oxygen saturation, waveform + digital
3.2.3.3 | Heart rate, digital
3.2.3.4 | Respiration, waveform + digital
3.2.3.5 | Minimum 3 Invasive pressure, waveform + digital (systolic, diastolic, mean)
3.2.3.6 | NIBP, digital (systolic, diastolic, mean)
3.2.3.7 | Temperature, digital
3.2.3.8 | Side-stream Capnography highly desireable. Provide full information
3.2.4 | The monitor must operate off both 220 V, 50 Hz single phase AC power supply and an internal rechargeable battery.
3.2.5 | If mains power fails the monitor must operate off battery power for at least 2 hours.
3.2.6 | Visible indication to show whether the monitor is operating from the mains or battery power.
3.2.7 | All control keys must be splashproof and resistant to the entry of fluids into the unit.
3.2.8 | The front panel must be easy to clean and disinfect.
3.2.9 | High resolution LCD colour display, 35 cm diagonal minimum
3.2.10 | At least 6 waveforms must be displayed
3.2.10.1 | Digits: large: 10 mm, small: 5 mm
3.2.10.2 | Sweep speeds : 25 and 50 mm/s
3.2.10.3 | It must be possible to freeze the display. This must not affect the QRS signal.
3.2.11 | Trending for 72 hours must be possible on the monitor,
3.2.12 | Unit to provide at least 72 hours of full disclosure wave data for review
3.2.13 | Defibrillator-proof up to 360 Joules
3.2.14 | Protected against HF interference
3.2.15 | Patient isolated from all potentially harmful circuits
3.3 | ECG
3.3.1 | ECG 12 lead with arrhythmia detection
3.3.2 | Adult ECG lead must be supplied
3.3.3 | The amplitude of the ECG trace must be adjustable from 1 mV = 5 mm to 1 mV = 20 mm.
3.3.4 | Digital display of heart rate : 30 - 300 bpm
3.3.5 | Accuracy : 3 bpm
3.3.6 | Resolution : 1 bpm
3.3.7 | Differential input with CMRR > 80 dB
3.3.8 | Input impedance > 5 MOhm
3.3.9 | Respiration monitoring through ECG electrodes
3.3.10 | Pacemaker detection/rejection : 2 to 500 mV; 0,1 to 2 ms
3.3.11 | Synch pulse for defibrillator
3.4 | OXYGEN SATURATION
3.4.1 | It must be the latest technology with clinical research support based on accuracy and real-time monitoring. Please submit the published peer reviewed/researched paper as supporting documents.
3.4.2 | The unit should indicate whether the signal from the pulse oximeter sensor is adequate for reliable measurements. If the signal strength or quality is inadequate, this must be clearly displayed.
3.4.3 | An audible bleep must sound for each heart beat. The pitch must vary with the oxygen saturation.
3.4.4 | Measurement range : 0 - 100%
3.4.5 | Accuracy:
3.4.5a | 80 - 100% : 1,5%
3.4.5b | 60 - 80% : 3,5%
3.4.6 | Continuous waveform display
3.4.7 | Digital display of SpO2%
3.4.8 | The sensor lead must be at least 3 metres long.
3.4.9 | Automatic calibration for skin colour
3.4.10 | Pulse rate: 25 - 250 bpm
3.5 | NON-INVASIVE BLOOD PRESSURE
3.5.1 | Oscillometric measurement method
3.5.2 | Automatic, manual and stat modes
3.5.3 | Measure and display :
3.5.4a | - systolic pressure
3.5.4b | - diastolic pressure
3.5.4c | - mean pressure
3.5.6 | The interval between measurements must be selectable from 1 to 60 minutes.
3.5.7 | Ranges (approximate) :
3.5.7a | Systolic : 30 - 270 mmHg. Resolution : 1 mmHg
3.5.7b | Diastolic : 10 - 150 mmHg. Resolution : 1 mmHg
3.5.7c | Mean : 20 - 250 mmHg. Resolution : 1 mmHg
3.5.8 | Cuff pressure : 0 - 280 mmHg
3.5.9 | Cuff inflation time < 10 s
3.5.10 | Automatic deflation if measurement exceeds 3 min or if mains power fails
3.6 | INVASIVE BLOOD PRESSURE
3.6.1 | Measure and display :
3.6.1a | - systolic pressure
3.6.1b | - diastolic pressure
3.6.1c | - mean pressure
3.6.2 | Systolic : 0 - 300 mmHg. Resolution : 1 mmHg
3.6.3 | Diastolic : 0 - 250 mmHg. Resolution : 1 mmHg
3.6.4 | Mean : 0 - 250 mmHg. Resolution : 1 mmHg
3.7 | TEMPERATURE
3.7.1 | Range : 10 - 45 ˚C
3.7.2 | Digital display
3.7.3 | Accuracy: 0.3 ˚C
3.7.4 | Resolution : 0.1 ˚C
3.8 | ALARMS
3.8.1 | Visible and audio alarms
3.8.2 | The monitor must automatically set default alarms when switched on.
3.8.3 | The alarms must be programmable
3.8.4 | Alarm limits must be displayed with the parameter digit
3.8.5 | Heart rate alarm : 30 - 300 bpm
3.8.6 | Asystole alarm
3.8.7 | Arrhythmia alarms
3.8.8 | Apnea alarm
3.8.9 | Alarms for faulty/not-connected electrode/probe
3.8.10 | User selectable alarm limits for systolic, diastolic and mean pressures
3.8.11 | Alarm for incorrect NIBPcuff connection
3.8.12 | Adjustable SpO2 high and low alarms:
3.8.12a | High : 70 - 100%
3.8.12b | Low : 50 - 95%
3.8.13 | Adjustable high and low temperature alarms
3.8.14 | Violation of limits must result in both audible and visible alarms and indicate which parameter is affected.
3.8.15 | The sound level of the audible alarm must be adjustable.
3.8.16 | It must be possible to temporarily silence the audible alarm but not remove the visible alarm.
3.8.17 | Duration of the silenced period: 2 minutes.
3.8.18 | Low battery alarm
3.9 | Printer
3.9.1 | Paper size must be standard A4 off the shelf paper
3.9.2 | With the capability of full disclosure waveform printing, multi-wave printing, tabular trend printing and trend graph printing.
3.9.3 | Printing must be possible at both bedside monitor and central station
3.10 | Data Storage and Review
3.10.1 | Trends of at least 120 hours must be available
3.10.2 | Trends must be displayed graphically and in tabular format
3.10.3 | Full disclosure of at least 144 hours to be offered
3.10.4 | Waveform storage must be available for at least 144 hours post discharge
3.10.5 | Events must be able to be sorted by time and type
3.10.6 | All patient events, graphical and numerical data and alarms must be available for viewing on a first in first out basis.
3.11 | General
3.11.1 | Network installation and cabling will be done by the bidder and included in the tender price.
3.11.2 | Height adjustable brackets. Monitor to be adjustable to a lower level than the bottom of the wall mounted baseplate to accommodate shorter staff members.
3.11.3 | The bidder shall provide fixtures and wall fittings for the entire system including network cables and switches.
3.11.4 | The bid includes the license and any software upgrades
3.11.5 | The unit must be handed over in full operating order
4 | The following accessories must be suitable for the patient size identified and supplied for each monitor:
4.1 | Neonatal Pack
4.1.1 | Neonatal NIBP hose
4.1.2 | Reusable latex-free neonate cuff
4.1.3 | Spo2 interface cable
4.1.4 | Finger wrap-around probe for neonates
4.1.5 | Ankle wrap-around probe for neonates
4.1.6 | Reusable temperature skin probe
4.1.7 | ECG interface cable
4.1.8 | Paediatric 3 lead ECG cable
AND/OR
4.2 | Paediatric Pack
4.2.1 | Adult/Paeds NIBP hose
4.2.2 | Reusable latex-free paediatric cuff
4.2.3 | ECG interface cable
4.2.4 | Paediatric 3 lead ECG cable
4.2.5 | Spo2 interface cable
4.2.6 | Finger sensor for paeds
4.2.7 | Reusable Paeds temperature probe
AND/OR
4.3 | Adult Pack
4.3.1 | NIBP hose
4.3.2 | Reusable latex-free adult cuff
4.3.3 | Reusable latex-free obese cuff
4.3.4 | Spo2 interface cable
4.3.5 | Finger sensor for adults
4.3.6 | Adult 3 lead ECG cable
4.3.7 | ECG interface cable
4.3.8 | Reusable Adult temperature probe
AND
4.4 | Invasive blood pressure cable X2
4.5 | Invasive blood pressure tranceducers X2
4.6 | Rechargeable Battery
4.7 | Wall bracket or Roll stand as per end-user requirement
5 | Conformity Compliance (Please attach certificate)
5.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
5.2 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
5.2.1 | Please provide unique ref number of the ISO 13485 ceriticate:
5.3 | Model on offer must comply with EU Medical Device Regulation (EU) 2017/745 or equvalent from SAHPRA recognized jurisdictions . Attach a copy of certification
5.3.1 | Please provide unique ref number of the EC ceriticate:
5.4 | Hazardous Substance Act: If this type of equipment / apparatus appears on the schedule of Hazardous substances issued by the Directorate: Health Technology of the Department of Health, a license in terms of the Act on Hazardous Substances (Act 15/1973) must be submitted with the bid / quotation documents. The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.5 | All electrical/electronic medical equipment must be licensed by Radiation Control, where applicable a copy of the license must be submitted The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.5.1 | Radiation Control License number:
5.6 | The unit shall operate from 220V AC/50Hz mains supply. The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade (rubber to with 2 screws) type and it must be a minimum length of three (3) meters. NB the mains cable of the unit tendered for must be SABS colour coded.
5.7 | The equipment quoted for must be protected against electromagnetic interference IEC 60601-1-2
5.8 | The unit must comply with electrical safety and performance standard such as IEC 60601-2-27 for Electrocardiographic Monitoring, attached certification
5.8.1 | Please provide unique ref number of the IEC ceriticate:
5.9 | The unit must comply with electrical safety and performance standard such as IEC 60601-2-30 for Automated Non-invasive Sphygmomanometers, attached certification
5.9.1 | Please provide unique ref number of the IEC ceriticate:
5.10 | The unit must comply with electrical safety and performance standard such as IEC 60601-2-34 for Invasive Blood Pressure monitoring, attached certification
5.10.1 | Please provide unique ref number of the IEC ceriticate:
5.11 | The unit must comply with electrical safety and performance standard such as IEC 60601-2-49 for Patient Monitoring, attached certification
5.11.1 | Please provide unique ref number of the IEC ceriticate:
6 | COMPREHENSIVE MAINTENANCE
6.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) should be quoted for, as per the SCC document
6.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
6.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
6.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
6.3 | The bidder to provide the OEM recommended service interval of the unit:
6.4 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
7 | General
7.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
7.2 | The unit on offer must be the latest technology.
7.2.1 | Initial Manufacturing year of Model:

## Central Monitoring, 12-bed
RT2-05-036 | Central Patient Monitoring, 12-Bed ICU Package | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
1 | Description
1.1 | A centralized monitoring system for continuous, real-time observation of patients' vital signs other physiological parameters for use in intensive care settings.
2.1 | Product Offered (Central Station):
2.1.1 | Make:
2.1.2 | Model:
2.1.3 | Brief Description:
2.2 | Product Offered (Monitor):
2.2.1 | Make:
2.2.2 | Model:
2.2.3 | Initial Manufacturing year of Model:
2.2.4 | Radiation Control License number:
3 | Technical Specification
3.1 | Central Station
3.1.1 | The central station must have a Single display, and 12 ICU monitors with complete accessories.
3.1.2 | The unit must be network configurable and wireless.
3.1.3 | Must have a viewing screen which will also do patient management and reviews and size must be at least 48 cm (19") colour LCD
3.1.4 | Touchscreen (Optional)
3.1.5 | The unit must be able to remotely manage patient bedside monitors, including viewing active or historic data.
3.1.6 | Password protection must be available to protect the critical menus and certain settings.
3.1.7 | All waveforms and alphanumeric displays must be available at the central station depending on the parameters of the bedside monitor
3.1.8 | Bi-directional communication between monitors and central station
3.1.9 | The central station must be able to download all the patient vitals after the patient has been disconnected and reconnected
3.1.10 | Display at least 8 beds information
3.1.11 | Full arrhythmia recall and ST segment analysis with 12 lead ECG interpretation
3.1.12 | Screen configurations must be able to be saved as favourites for easy access
3.1.13 | Up to 10 arrhythmia events per patient monitoring session must be available
3.1.14 | Must be able to review full disclosure on the central station for the duration of the disclosure time. At least 120 hours
3.1.15 | The central station must allow access to patient demographic information; information arrhythmias, ST-limit violations, and ECG samples in event review; parameter data in graphic and numeric trends; full disclosure data in strip view, page view with zoom, or in calipers.
3.1.16 | Alarms
3.1.16.1 | All alarms must be able to be muted only at the bedside
3.1.16.2 | 3-level alarm must be available audible & visual.
3.1.16.3 | The sound levels and colour to differentiate alarm type
3.1.16.4 | High and low limits of alarm must be offered
3.1.16.5 | Monitor alarm limits must be adjustable from the central station. This setting must be enabled/disabled via a password protected menu.
3.2 | Monitor
3.2.1 | Modular, state of the art technology, suitable for neonates to adult
3.2.2 | Microprocessor controlled with automatic self-test and calibration
3.2.3 | The following parameters must be measured and displayed :
3.2.3.1 | ECG waveform
3.2.3.2 | Oxygen saturation, waveform + digital
3.2.3.3 | Heart rate, digital
3.2.3.4 | Respiration, waveform + digital
3.2.3.5 | Minimum 3 Invasive pressure, waveform + digital (systolic, diastolic, mean)
3.2.3.6 | NIBP, digital (systolic, diastolic, mean)
3.2.3.7 | Temperature, digital
3.2.3.8 | Side-stream Capnography highly desireable. Provide full information
3.2.4 | The monitor must operate off both 220 V, 50 Hz single phase AC power supply and an internal rechargeable battery.
3.2.5 | If mains power fails the monitor must operate off battery power for at least 2 hours.
3.2.6 | Visible indication to show whether the monitor is operating from the mains or battery power.
3.2.7 | All control keys must be splashproof and resistant to the entry of fluids into the unit.
3.2.8 | The front panel must be easy to clean and disinfect.
3.2.9 | High resolution LCD colour display, 35 cm diagonal minimum
3.2.10 | At least 6 waveforms must be displayed
3.2.10.1 | Digits: large: 10 mm, small: 5 mm
3.2.10.2 | Sweep speeds : 25 and 50 mm/s
3.2.10.3 | It must be possible to freeze the display. This must not affect the QRS signal.
3.2.11 | Trending for 72 hours must be possible on the monitor,
3.2.12 | Unit to provide at least 72 hours of full disclosure wave data for review
3.2.13 | Defibrillator-proof up to 360 Joules
3.2.14 | Protected against HF interference
3.2.15 | Patient isolated from all potentially harmful circuits
3.3 | ECG
3.3.1 | ECG 12 lead with arrhythmia detection
3.3.2 | Adult ECG lead must be supplied
3.3.3 | The amplitude of the ECG trace must be adjustable from 1 mV = 5 mm to 1 mV = 20 mm.
3.3.4 | Digital display of heart rate : 30 - 300 bpm
3.3.5 | Accuracy : 3 bpm
3.3.6 | Resolution : 1 bpm
3.3.7 | Differential input with CMRR > 80 dB
3.3.8 | Input impedance > 5 MOhm
3.3.9 | Respiration monitoring through ECG electrodes
3.3.10 | Pacemaker detection/rejection : 2 to 500 mV; 0,1 to 2 ms
3.3.11 | Synch pulse for defibrillator
3.4 | OXYGEN SATURATION
3.4.1 | It must be the latest technology with clinical research support based on accuracy and real-time monitoring. Please submit the published peer reviewed/researched paper as supporting documents.
3.4.2 | The unit should indicate whether the signal from the pulse oximeter sensor is adequate for reliable measurements. If the signal strength or quality is inadequate, this must be clearly displayed.
3.4.3 | An audible bleep must sound for each heart beat. The pitch must vary with the oxygen saturation.
3.4.4 | Measurement range : 0 - 100%
3.4.5 | Accuracy:
3.4.5a | 80 - 100% : 1,5%
3.4.5b | 60 - 80% : 3,5%
3.4.6 | Continuous waveform display
3.4.7 | Digital display of SpO2%
3.4.8 | The sensor lead must be at least 3 metres long.
3.4.9 | Automatic calibration for skin colour
3.4.10 | Pulse rate: 25 - 250 bpm
3.5 | NON-INVASIVE BLOOD PRESSURE
3.5.1 | Oscillometric measurement method
3.5.2 | Automatic, manual and stat modes
3.5.3 | Measure and display :
3.5.4a | - systolic pressure
3.5.4b | - diastolic pressure
3.5.4c | - mean pressure
3.5.6 | The interval between measurements must be selectable from 1 to 60 minutes.
3.5.7 | Ranges (approximate) :
3.5.7a | Systolic : 30 - 270 mmHg. Resolution : 1 mmHg
3.5.7b | Diastolic : 10 - 150 mmHg. Resolution : 1 mmHg
3.5.7c | Mean : 20 - 250 mmHg. Resolution : 1 mmHg
3.5.8 | Cuff pressure : 0 - 280 mmHg
3.5.9 | Cuff inflation time < 10 s
3.5.10 | Automatic deflation if measurement exceeds 3 min or if mains power fails
3.6 | INVASIVE BLOOD PRESSURE
3.6.1 | Measure and display :
3.6.1a | - systolic pressure
3.6.1b | - diastolic pressure
3.6.1c | - mean pressure
3.6.2 | Systolic : 0 - 300 mmHg. Resolution : 1 mmHg
3.6.3 | Diastolic : 0 - 250 mmHg. Resolution : 1 mmHg
3.6.4 | Mean : 0 - 250 mmHg. Resolution : 1 mmHg
3.7 | TEMPERATURE
3.7.1 | Range : 10 - 45 ˚C
3.7.2 | Digital display
3.7.3 | Accuracy: 0.3 ˚C
3.7.4 | Resolution : 0.1 ˚C
3.8 | ALARMS
3.8.1 | Visible and audio alarms
3.8.2 | The monitor must automatically set default alarms when switched on.
3.8.3 | The alarms must be programmable
3.8.4 | Alarm limits must be displayed with the parameter digit
3.8.5 | Heart rate alarm : 30 - 300 bpm
3.8.6 | Asystole alarm
3.8.7 | Arrhythmia alarms
3.8.8 | Apnea alarm
3.8.9 | Alarms for faulty/not-connected electrode/probe
3.8.10 | User selectable alarm limits for systolic, diastolic and mean pressures
3.8.11 | Alarm for incorrect NIBPcuff connection
3.8.12 | Adjustable SpO2 high and low alarms:
3.8.12a | High : 70 - 100%
3.8.12b | Low : 50 - 95%
3.8.13 | Adjustable high and low temperature alarms
3.8.14 | Violation of limits must result in both audible and visible alarms and indicate which parameter is affected.
3.8.15 | The sound level of the audible alarm must be adjustable.
3.8.16 | It must be possible to temporarily silence the audible alarm but not remove the visible alarm.
3.8.17 | Duration of the silenced period: 2 minutes.
3.8.18 | Low battery alarm
3.9 | Printer
3.9.1 | Paper size must be standard A4 off the shelf paper
3.9.2 | With the capability of full disclosure waveform printing, multi-wave printing, tabular trend printing and trend graph printing.
3.9.3 | Printing must be possible at both bedside monitor and central station
3.10 | Data Storage and Review
3.10.1 | Trends of at least 120 hours must be available
3.10.2 | Trends must be displayed graphically and in tabular format
3.10.3 | Full disclosure of at least 144 hours to be offered
3.10.4 | Waveform storage must be available for at least 144 hours post discharge
3.10.5 | Events must be able to be sorted by time and type
3.10.6 | All patient events, graphical and numerical data and alarms must be available for viewing on a first in first out basis.
3.11 | General
3.11.1 | Network installation and cabling will be done by the bidder and included in the tender price.
3.11.2 | Height adjustable brackets. Monitor to be adjustable to a lower level than the bottom of the wall mounted baseplate to accommodate shorter staff members.
3.11.3 | The bidder shall provide fixtures and wall fittings for the entire system including network cables and switches.
3.11.4 | The bid includes the license and any software upgrades
3.11.5 | The unit must be handed over in full operating order
4 | The following accessories must be suitable for the patient size identified and supplied for each monitor:
4.1 | Neonatal Pack
4.1.1 | Neonatal NIBP hose
4.1.2 | Reusable latex-free neonate cuff
4.1.3 | Spo2 interface cable
4.1.4 | Finger wrap-around probe for neonates
4.1.5 | Ankle wrap-around probe for neonates
4.1.6 | Reusable temperature skin probe
4.1.7 | ECG interface cable
4.1.8 | Paediatric 3 lead ECG cable
AND/OR
4.2 | Paediatric Pack
4.2.1 | Adult/Paeds NIBP hose
4.2.2 | Reusable latex-free paediatric cuff
4.2.3 | ECG interface cable
4.2.4 | Paediatric 3 lead ECG cable
4.2.5 | Spo2 interface cable
4.2.6 | Finger sensor for paeds
4.2.7 | Reusable Paeds temperature probe
AND/OR
4.3 | Adult Pack
4.3.1 | NIBP hose
4.3.2 | Reusable latex-free adult cuff
4.3.3 | Reusable latex-free obese cuff
4.3.4 | Spo2 interface cable
4.3.5 | Finger sensor for adults
4.3.6 | Adult 3 lead ECG cable
4.3.7 | ECG interface cable
4.3.8 | Reusable Adult temperature probe
AND
4.4 | Invasive blood pressure cable X2
4.5 | Invasive blood pressure tranceducers X2
4.6 | Rechargeable Battery
4.7 | Wall bracket or Roll stand as per end-user requirement
5 | Conformity Compliance (Please attach certificate)
5.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
5.2 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
5.2.1 | Please provide unique ref number of the ISO 13485 ceriticate:
5.3 | Model on offer must comply with EU Medical Device Regulation (EU) 2017/745 or equvalent from SAHPRA recognized jurisdictions . Attach a copy of certification
5.3.1 | Please provide unique ref number of the EC ceriticate:
5.4 | Hazardous Substance Act: If this type of equipment / apparatus appears on the schedule of Hazardous substances issued by the Directorate: Health Technology of the Department of Health, a license in terms of the Act on Hazardous Substances (Act 15/1973) must be submitted with the bid / quotation documents. The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.5 | All electrical/electronic medical equipment must be licensed by Radiation Control, where applicable a copy of the license must be submitted The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.5.1 | Radiation Control License number:
5.6 | The unit shall operate from 220V AC/50Hz mains supply. The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade (rubber to with 2 screws) type and it must be a minimum length of three (3) meters. NB the mains cable of the unit tendered for must be SABS colour coded.
5.7 | The equipment quoted for must be protected against electromagnetic interference IEC 60601-1-2
5.8 | The unit must comply with electrical safety and performance standard such as IEC 60601-2-27 for Electrocardiographic Monitoring, attached certification
5.8.1 | Please provide unique ref number of the IEC ceriticate:
5.9 | The unit must comply with electrical safety and performance standard such as IEC 60601-2-30 for Automated Non-invasive Sphygmomanometers, attached certification
5.9.1 | Please provide unique ref number of the IEC ceriticate:
5.10 | The unit must comply with electrical safety and performance standard such as IEC 60601-2-34 for Invasive Blood Pressure monitoring, attached certification
5.10.1 | Please provide unique ref number of the IEC ceriticate:
5.11 | The unit must comply with electrical safety and performance standard such as IEC 60601-2-49 for Patient Monitoring, attached certification
5.11.1 | Please provide unique ref number of the IEC ceriticate:
6 | COMPREHENSIVE MAINTENANCE
6.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) should be quoted for, as per the SCC document
6.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
6.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
6.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
6.3 | The bidder to provide the OEM recommended service interval of the unit:
6.4 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
7 | General
7.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
7.2 | The unit on offer must be the latest technology.
7.2.1 | Initial Manufacturing year of Model:

## Central Monitoring, 16-bed
RT2-05-037 | Central Patient Monitoring, 16-Bed ICU Package | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
1 | Description
1.1 | A centralized monitoring system for continuous, real-time observation of patients' vital signs other physiological parameters for use in intensive care settings.
2.1 | Product Offered (Central Station):
2.1.1 | Make:
2.1.2 | Model:
2.1.3 | Brief Description:
2.2 | Product Offered (Monitor):
2.2.1 | Make:
2.2.2 | Model:
2.2.3 | Initial Manufacturing year of Model:
2.2.4 | Radiation Control License number:
3 | Technical Specification
3.1 | Central Station
3.1.1 | The central station must have a Single display, and 16 ICU monitors with complete accessories.
3.1.2 | The unit must be network configurable and wireless.
3.1.3 | Must have a viewing screen which will also do patient management and reviews and size must be at least 48 cm (19") colour LCD
3.1.4 | Touchscreen (Optional)
3.1.5 | The unit must be able to remotely manage patient bedside monitors, including viewing active or historic data.
3.1.6 | Password protection must be available to protect the critical menus and certain settings.
3.1.7 | All waveforms and alphanumeric displays must be available at the central station depending on the parameters of the bedside monitor
3.1.8 | Bi-directional communication between monitors and central station
3.1.9 | The central station must be able to download all the patient vitals after the patient has been disconnected and reconnected
3.1.10 | Display at least 8 beds information
3.1.11 | Full arrhythmia recall and ST segment analysis with 12 lead ECG interpretation
3.1.12 | Screen configurations must be able to be saved as favourites for easy access
3.1.13 | Up to 10 arrhythmia events per patient monitoring session must be available
3.1.14 | Must be able to review full disclosure on the central station for the duration of the disclosure time. At least 120 hours
3.1.15 | The central station must allow access to patient demographic information; information arrhythmias, ST-limit violations, and ECG samples in event review; parameter data in graphic and numeric trends; full disclosure data in strip view, page view with zoom, or in calipers.
3.1.16 | Alarms
3.1.16.1 | All alarms must be able to be muted only at the bedside
3.1.16.2 | 3-level alarm must be available audible & visual.
3.1.16.3 | The sound levels and colour to differentiate alarm type
3.1.16.4 | High and low limits of alarm must be offered
3.1.16.5 | Monitor alarm limits must be adjustable from the central station. This setting must be enabled/disabled via a password protected menu.
3.2 | Monitor
3.2.1 | Modular, state of the art technology, suitable for neonates to adult
3.2.2 | Microprocessor controlled with automatic self-test and calibration
3.2.3 | The following parameters must be measured and displayed :
3.2.3.1 | ECG waveform
3.2.3.2 | Oxygen saturation, waveform + digital
3.2.3.3 | Heart rate, digital
3.2.3.4 | Respiration, waveform + digital
3.2.3.5 | Minimum 3 Invasive pressure, waveform + digital (systolic, diastolic, mean)
3.2.3.6 | NIBP, digital (systolic, diastolic, mean)
3.2.3.7 | Temperature, digital
3.2.3.8 | Side-stream Capnography highly desireable. Provide full information
3.2.4 | The monitor must operate off both 220 V, 50 Hz single phase AC power supply and an internal rechargeable battery.
3.2.5 | If mains power fails the monitor must operate off battery power for at least 2 hours.
3.2.6 | Visible indication to show whether the monitor is operating from the mains or battery power.
3.2.7 | All control keys must be splashproof and resistant to the entry of fluids into the unit.
3.2.8 | The front panel must be easy to clean and disinfect.
3.2.9 | High resolution LCD colour display, 35 cm diagonal minimum
3.2.10 | At least 6 waveforms must be displayed
3.2.10.1 | Digits: large: 10 mm, small: 5 mm
3.2.10.2 | Sweep speeds : 25 and 50 mm/s
3.2.10.3 | It must be possible to freeze the display. This must not affect the QRS signal.
3.2.11 | Trending for 72 hours must be possible on the monitor,
3.2.12 | Unit to provide at least 72 hours of full disclosure wave data for review
3.2.13 | Defibrillator-proof up to 360 Joules
3.2.14 | Protected against HF interference
3.2.15 | Patient isolated from all potentially harmful circuits
3.3 | ECG
3.3.1 | ECG 12 lead with arrhythmia detection
3.3.2 | Adult ECG lead must be supplied
3.3.3 | The amplitude of the ECG trace must be adjustable from 1 mV = 5 mm to 1 mV = 20 mm.
3.3.4 | Digital display of heart rate : 30 - 300 bpm
3.3.5 | Accuracy : 3 bpm
3.3.6 | Resolution : 1 bpm
3.3.7 | Differential input with CMRR > 80 dB
3.3.8 | Input impedance > 5 MOhm
3.3.9 | Respiration monitoring through ECG electrodes
3.3.10 | Pacemaker detection/rejection : 2 to 500 mV; 0,1 to 2 ms
3.3.11 | Synch pulse for defibrillator
3.4 | OXYGEN SATURATION
3.4.1 | It must be the latest technology with clinical research support based on accuracy and real-time monitoring. Please submit the published peer reviewed/researched paper as supporting documents.
3.4.2 | The unit should indicate whether the signal from the pulse oximeter sensor is adequate for reliable measurements. If the signal strength or quality is inadequate, this must be clearly displayed.
3.4.3 | An audible bleep must sound for each heart beat. The pitch must vary with the oxygen saturation.
3.4.4 | Measurement range : 0 - 100%
3.4.5 | Accuracy:
3.4.5a | 80 - 100% : 1,5%
3.4.5b | 60 - 80% : 3,5%
3.4.6 | Continuous waveform display
3.4.7 | Digital display of SpO2%
3.4.8 | The sensor lead must be at least 3 metres long.
3.4.9 | Automatic calibration for skin colour
3.4.10 | Pulse rate: 25 - 250 bpm
3.5 | NON-INVASIVE BLOOD PRESSURE
3.5.1 | Oscillometric measurement method
3.5.2 | Automatic, manual and stat modes
3.5.3 | Measure and display :
3.5.4a | - systolic pressure
3.5.4b | - diastolic pressure
3.5.4c | - mean pressure
3.5.6 | The interval between measurements must be selectable from 1 to 60 minutes.
3.5.7 | Ranges (approximate) :
3.5.7a | Systolic : 30 - 270 mmHg. Resolution : 1 mmHg
3.5.7b | Diastolic : 10 - 150 mmHg. Resolution : 1 mmHg
3.5.7c | Mean : 20 - 250 mmHg. Resolution : 1 mmHg
3.5.8 | Cuff pressure : 0 - 280 mmHg
3.5.9 | Cuff inflation time < 10 s
3.5.10 | Automatic deflation if measurement exceeds 3 min or if mains power fails
3.6 | INVASIVE BLOOD PRESSURE
3.6.1 | Measure and display :
3.6.1a | - systolic pressure
3.6.1b | - diastolic pressure
3.6.1c | - mean pressure
3.6.2 | Systolic : 0 - 300 mmHg. Resolution : 1 mmHg
3.6.3 | Diastolic : 0 - 250 mmHg. Resolution : 1 mmHg
3.6.4 | Mean : 0 - 250 mmHg. Resolution : 1 mmHg
3.7 | TEMPERATURE
3.7.1 | Range : 10 - 45 ˚C
3.7.2 | Digital display
3.7.3 | Accuracy: 0.3 ˚C
3.7.4 | Resolution : 0.1 ˚C
3.8 | ALARMS
3.8.1 | Visible and audio alarms
3.8.2 | The monitor must automatically set default alarms when switched on.
3.8.3 | The alarms must be programmable
3.8.4 | Alarm limits must be displayed with the parameter digit
3.8.5 | Heart rate alarm : 30 - 300 bpm
3.8.6 | Asystole alarm
3.8.7 | Arrhythmia alarms
3.8.8 | Apnea alarm
3.8.9 | Alarms for faulty/not-connected electrode/probe
3.8.10 | User selectable alarm limits for systolic, diastolic and mean pressures
3.8.11 | Alarm for incorrect NIBPcuff connection
3.8.12 | Adjustable SpO2 high and low alarms:
3.8.12a | High : 70 - 100%
3.8.12b | Low : 50 - 95%
3.8.13 | Adjustable high and low temperature alarms
3.8.14 | Violation of limits must result in both audible and visible alarms and indicate which parameter is affected.
3.8.15 | The sound level of the audible alarm must be adjustable.
3.8.16 | It must be possible to temporarily silence the audible alarm but not remove the visible alarm.
3.8.17 | Duration of the silenced period: 2 minutes.
3.8.18 | Low battery alarm
3.9 | Printer
3.9.1 | Paper size must be standard A4 off the shelf paper
3.9.2 | With the capability of full disclosure waveform printing, multi-wave printing, tabular trend printing and trend graph printing.
3.9.3 | Printing must be possible at both bedside monitor and central station
3.10 | Data Storage and Review
3.10.1 | Trends of at least 120 hours must be available
3.10.2 | Trends must be displayed graphically and in tabular format
3.10.3 | Full disclosure of at least 144 hours to be offered
3.10.4 | Waveform storage must be available for at least 144 hours post discharge
3.10.5 | Events must be able to be sorted by time and type
3.10.6 | All patient events, graphical and numerical data and alarms must be available for viewing on a first in first out basis.
3.11 | General
3.11.1 | Network installation and cabling will be done by the bidder and included in the tender price.
3.11.2 | Height adjustable brackets. Monitor to be adjustable to a lower level than the bottom of the wall mounted baseplate to accommodate shorter staff members.
3.11.3 | The bidder shall provide fixtures and wall fittings for the entire system including network cables and switches.
3.11.4 | The bid includes the license and any software upgrades
3.11.5 | The unit must be handed over in full operating order
4 | The following accessories must be suitable for the patient size identified and supplied for each monitor:
4.1 | Neonatal Pack
4.1.1 | Neonatal NIBP hose
4.1.2 | Reusable latex-free neonate cuff
4.1.3 | Spo2 interface cable
4.1.4 | Finger wrap-around probe for neonates
4.1.5 | Ankle wrap-around probe for neonates
4.1.6 | Reusable temperature skin probe
4.1.7 | ECG interface cable
4.1.8 | Paediatric 3 lead ECG cable
AND/OR
4.2 | Paediatric Pack
4.2.1 | Adult/Paeds NIBP hose
4.2.2 | Reusable latex-free paediatric cuff
4.2.3 | ECG interface cable
4.2.4 | Paediatric 3 lead ECG cable
4.2.5 | Spo2 interface cable
4.2.6 | Finger sensor for paeds
4.2.7 | Reusable Paeds temperature probe
AND/OR
4.3 | Adult Pack
4.3.1 | NIBP hose
4.3.2 | Reusable latex-free adult cuff
4.3.3 | Reusable latex-free obese cuff
4.3.4 | Spo2 interface cable
4.3.5 | Finger sensor for adults
4.3.6 | Adult 3 lead ECG cable
4.3.7 | ECG interface cable
4.3.8 | Reusable Adult temperature probe
AND
4.4 | Invasive blood pressure cable X2
4.5 | Invasive blood pressure tranceducers X2
4.6 | Rechargeable Battery
4.7 | Wall bracket or Roll stand as per end-user requirement
5 | Conformity Compliance (Please attach certificate)
5.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
5.2 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
5.2.1 | Please provide unique ref number of the ISO 13485 ceriticate:
5.3 | Model on offer must comply with EU Medical Device Regulation (EU) 2017/745 or equvalent from SAHPRA recognized jurisdictions . Attach a copy of certification
5.3.1 | Please provide unique ref number of the EC ceriticate:
5.4 | Hazardous Substance Act: If this type of equipment / apparatus appears on the schedule of Hazardous substances issued by the Directorate: Health Technology of the Department of Health, a license in terms of the Act on Hazardous Substances (Act 15/1973) must be submitted with the bid / quotation documents. The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.5 | All electrical/electronic medical equipment must be licensed by Radiation Control, where applicable a copy of the license must be submitted The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.5.1 | Radiation Control License number:
5.6 | The unit shall operate from 220V AC/50Hz mains supply. The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade (rubber to with 2 screws) type and it must be a minimum length of three (3) meters. NB the mains cable of the unit tendered for must be SABS colour coded.
5.7 | The equipment quoted for must be protected against electromagnetic interference IEC 60601-1-2
5.8 | The unit must comply with electrical safety and performance standard such as IEC 60601-2-27 for Electrocardiographic Monitoring, attached certification
5.8.1 | Please provide unique ref number of the IEC ceriticate:
5.9 | The unit must comply with electrical safety and performance standard such as IEC 60601-2-30 for Automated Non-invasive Sphygmomanometers, attached certification
5.9.1 | Please provide unique ref number of the IEC ceriticate:
5.10 | The unit must comply with electrical safety and performance standard such as IEC 60601-2-34 for Invasive Blood Pressure monitoring, attached certification
5.10.1 | Please provide unique ref number of the IEC ceriticate:
5.11 | The unit must comply with electrical safety and performance standard such as IEC 60601-2-49 for Patient Monitoring, attached certification
5.11.1 | Please provide unique ref number of the IEC ceriticate:
6 | COMPREHENSIVE MAINTENANCE
6.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) should be quoted for, as per the SCC document
6.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
6.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
6.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
6.3 | The bidder to provide the OEM recommended service interval of the unit:
6.4 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
7 | General
7.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
7.2 | The unit on offer must be the latest technology.
7.2.1 | Initial Manufacturing year of Model:

## Central Station
RT2-05-017 | Central Station [16 bed license] | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
1 | Description
1.1 | A centralized monitoring system for continuous, real-time observation of patients' vital signs other physiological parameters for use in intensive care settings.
2 | Product Offered (Central Station):
2.1 | Make:
2.2 | Model:
2.3 | Brief Description:
3 | Technical Specification
3.1 | Central Stations that are Universal/Vendor Neutral allowing end-user choice of Network configurable Monitor brands, will be given preference
3.2 | The central station must have 2 screens.
3.3 | The unit must be network configurable and wireless.
3.4 | Must have a viewing screen as well as a review screen for measurements and both screens sizes must be at least 48 cm (19") colour LCD
3.5 | The second screen must be able to do patient management and reviews without interruption of the live patient view on the first screen
3.6 | Touchscreen (Optional)
3.7 | The unit must be able to remotely manage patient bedside monitors, including viewing active or historic data.
3.8 | Password protection must be available to protect the critical menus and certain settings.
3.9 | All waveforms and alphanumeric displays must be available at the central station depending on the parameters of the bedside monitor
3.10 | Bi-directional communication between monitors and central station
3.11 | The central station must be able to download all the patient vitals after the patient has been disconnected and reconnected
3.12 | Display at least 8 beds information
3.13 | Full arrhythmia recall and ST segment analysis with 12 lead ECG interpretation
3.14 | Screen configurations must be able to be saved as favourites for easy access
3.15 | Up to to 10 arrhythmia events per patient monitoring session must be available
3.16 | Must be able to review full disclosure on the central station for the duration of the disclosure time. At least 120 hours
3.17 | The central station must allow access to patient demographic information; information arrhythmias, ST-limit violations, and ECG samples in event review; parameter data in graphic and numeric trends; full disclosure data in strip view, page view with zoom, or in calipers.
3.18 | The station must be provided complete with a 16-bed licence [once off], switch and installation (Monitor configuration ready) and any software upgrades.
3.19 | Alarms
3.19.1 | All alarms must be able to be muted only at the bedside
3.19.2 | 3-level alarm must be available audible & visual.
3.19.3 | The sound levels and colour to differentiate alarm type
3.19.4 | High and low limits of alarm must be offered
3.19.5 | Monitor alarm limits must be adjustable from the central station. This setting must be enabled/disabled via a password protected menu.
3.20 | Printer
3.20.1 | Paper size must be standard A4 off the shelf paper
3.20.2 | With the capability of full disclosure waveform printing, multi-wave printing, tabular trend printing and trend graph printing.
3.20.3 | Printing must be possible at both bedside monitor and central station
3.21 | Data Storage and Review
3.21.1 | Trends of at least 120 hours must be available
3.21.2 | Trends must be displayed graphically and in tabular format
3.21.3 | Full disclosure of at least 144 hours to be offered
3.21.4 | Waveform storage must be available for at least 144 hours post discharge
3.21.5 | Events must be able to be sorted by time and type
3.21.6 | All patient events, graphical and numerical data and alarms must be available for viewing on a first in first out basis.
3.22 | General
3.22.1 | Network installation and cabling will be done by the bidder and included in the tender price.
3.22.2 | The bidder shall provide fixtures and wall fittings for the entire system including network cables and switches.
3.22.3 | The central station must be compatible with different brands.
3.22.4 | The unit to be supplied with the UPS (uninterrupted power supply)
4 | Conformity Compliance (Please attach certificate)
4.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
4.2 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
4.2.1 | Please provide unique ref number of the ISO 13485 ceriticate:
4.3 | Model on offer must comply with EU Medical Device Regulation (EU) 2017/745 or equvalent from SAHPRA recognized jurisdictions . Attach a copy of certification
4.3.1 | Please provide unique ref number of the EC ceriticate:
4.4 | The unit shall operate from 220V AC/50Hz mains supply. The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade (rubber to with 2 screws) type and it must be a minimum length of three (3) meters. NB the mains cable of the unit tendered for must be SABS colour coded.
4.5 | The equipment quoted for must be protected against electromagnetic interference IEC 60601-1-2
5 | COMPREHENSIVE MAINTENANCE
5.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) should be quoted for, as per the SCC document
5.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
6 | General
6.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
6.2 | The unit on offer must be the latest technology.
6.2.1 | Initial Manufacturing year of Model:

## Vital Signs, NIBP+SpO2
RT2-05-018 | Vital Signs Monitor, on trolley | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
1 | Description
1.1 | A medical device that measures and displays a patient's vital signs, that is on a mobile trolley
2 | Product Offered:
2.1 | Make:
2.2 | Model:
2.3 | Initial Manufacturing year of Model:
2.4 | Radiation Control License number:
3 | Technical Specification
3.1 | The unit must measure NIBP, Temperature and SpO2 suitable for adults, paeds and neonates
3.2 | Self test procedure at switch-on
3.3 | The unit must have trending for at least 24 hours
3.4 | NIBP
3.4.1 | The unit must measure blood pressure non-invasively.
3.4.3 | The unit must use the oscillometric method of measurement.
3.4.4 | Modes: Manual, automatic and STAT
3.4.5 | Pressure measurements:
3.4.5a | Adult: Systolic: 30 - 270 mm Hg, Diastolic: 20 - 240 mm Hg, MAP: 20 - 250 mm Hg
3.4.5b | Paeds: Systolic: 30 - 160 mm Hg, Diastolic: 20 - 120 mm Hg, MAP: 20 - 120 mm Hg
3.4.5c | Neonate: Systolic: 30 - 130 mm Hg, Diastolic: 20 - 100 mm Hg, MAP: 20 - 120 mm Hg
3.4.6 | Auto zero when unit is switched on
3.4.7 | The interval between measurements must be selectable from 1 – 120 minutes.
3.4.8 | Over pressure protection: Automatic deflation at 300 mm Hg adult, 150 mm Hg neonate
3.4.9 | Average measurement time: approximately 30 s
3.4.10 | If measurement exceeds two minutes the cuff must deflate, an alarm must sound and likely faults indicated.
3.4.11 | The air hose must be at least two metres long.
3.5 | PULSE OXIMETER
3.5.2 | The unit must continuously measure and display SpO2 %
3.5.3 | Oxygen saturation (SpO2 %): Range of measurement: 0 - 100%
3.5.4 | The accuracy must be better than 2,5 % (70 - 100%) and 5% (50 - 70%).
3.5.5 | Pulse rate (beats/min): Range of measurement: 20 - 300 bpm
3.5.6 | The pulse rate accuracy must be better than ± 2 %.
3.5.7 | An audible bleep should sound for each heartbeat, the pitch of which should vary with the oxygen saturation.
3.5.8 | Automatic rejection of motion artifacts
3.6 | SpO2 SENSOR:
3.6.1 | Sensor: Dual LED, Infra-red/visible red, which does not heat the point of application to the patient.
3.6.2 | The sensor lead must be approximately 3 metres long.
3.7 | LCD Display:
3.7.1 | The display must offer clear viewing under all lighting conditions.
3.7.2 | The following must be displayed:
3.7.2a | Systolic, diastolic and mean arterial pressures (numeric)
3.7.2b | Oxygen saturation (SpO2 %)
3.7.2c | Pulse rate (beats/min)
3.7.2d | Alarm limits
3.7.2e | SpO2 waveform or bar-graph
3.7.2f | Trend information
3.7.2g | Battery status
3.7.2h | Indication of mains or battery operation
3.7.3 | The unit must indicate whether the signal to the oximeter sensor is adequate for reliable measurements. If the signal strength or quality is inadequate, this must be clearly displayed.
3.8 | TEMPERATURE
3.8.1 | Range : 10 - 45 ˚C
3.8.2 | Digital display
3.8.3 | Accuracy: 0.3 ˚C
3.8.4 | Resolution : 0.1 ˚C
3.9 | ALARMS:
3.9.1 | All alarms must have default settings
3.9.2 | The alarm parameters must be user-adjustable:
3.9.3 | Adjustable high and low alarm limits for systolic, diastolic, mean arterial pressure and heart rate.
3.9.4 | SpO2 LOW: 50 % - 99 %
3.9.5 | Pulse rate LOW: 40 – 150 b.p.m.
3.9.6 | Pulse rate HIGH: 70 – 250 b.p.m.
3.9.7 | SpO2 probe not attached to patient or the pulse oximeter
3.9.8 | Violation of limits must result in both audible and visual alarms and indication of the parameters affected.
3.9.9 | Adjustable audio alarm volume
3.9.10 | Audio alarm mute: 2 minute.
3.9.11 | Low battery: Alarm at least 10 minutes before battery power is exhausted
3.9.12 | Temperature high and low alarms
3.10 | Other system requirements
3.10.1 | The unit must include an internal rechargeable battery
3.10.2 | The unit must be able to function from the mains supply regardless of the battery condition.
3.10.3 | Battery must have the capacity for the unit to operate fully for at least five hours
3.10.4 | The battery must be charged while the unit is operating from the mains supply.
3.10.5 | The battery must automatically take over in the event of a mains power failure.
3.10.6 | The power supply must be an integral part of the unit.
3.10.7 | The case of the unit must be robust and impact resistant.
3.10.8 | The design must be such that spilled liquids do not gain access to the internal working of the unit.
3.10.9 | The monitor must be mounted on a trolley with basket for accessories
3.10.10 | The unit must be handed over in full operating order
3.10.11 | A starter-pack of consumables must be supplied with the unit.
4 | Standard Accessories and Consumables
4.1 | Neonatal Pack
4.1.1 | Neonatal NIBP hose
4.1.2 | Disposable neonate cuff - pack of 20
4.1.3 | Spo2 interface cable
4.1.4 | Finger wrap-around probe for neonates - disposable pack of 20
4.1.5 | Ankle wrap-around probe for neonates - disposable pack of 20
4.1.6 | Reusable temperature skin probe
AND/OR
4.2 | Paediatric Pack
4.2.1 | Adult/Paeds NIBP hose
4.2.2 | Reusable latex-free paediatric cuff
4.2.5 | Spo2 interface cable
4.2.6 | Finger sensor for paeds
4.2.7 | Reusable Paeds temperature probe
AND/OR
4.3 | Adult Pack
4.3.1 | NIBP hose
4.3.2 | Reusable latex-free adult cuff
4.3.3 | Reusable latex-free obese cuff
4.3.4 | Spo2 interface cable
4.3.5 | Finger sensor for adults
4.3.8 | Reusable Adult temperature probe
AND
4.4 | Rechargeable Battery
4.5 | Roll stand
5 | Conformity Compliance (Please attach certificate)
5.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
5.2 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
5.2.1 | Please provide unique ref number of the ISO 13485 ceriticate:
5.3 | Model on offer must comply with EU Medical Device Regulation (EU) 2017/745 or equvalent from SAHPRA recognized jurisdictions . Attach a copy of certification
5.3.1 | Please provide unique ref number of the EC ceriticate:
5.4 | Hazardous Substance Act: If this type of equipment / apparatus appears on the schedule of Hazardous substances issued by the Directorate: Health Technology of the Department of Health, a license in terms of the Act on Hazardous Substances (Act 15/1973) must be submitted with the bid / quotation documents. The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.5 | All electrical/electronic medical equipment must be licensed by Radiation Control, where applicable a copy of the license must be submitted The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.5.1 | Radiation Control License number:
5.6 | The unit shall operate from 220V AC/50Hz mains supply. The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade (rubber to with 2 screws) type and it must be a minimum length of three (3) meters. NB the mains cable of the unit tendered for must be SABS colour coded.
5.7 | The equipment quoted for must be protected against electromagnetic interference IEC 60601-1-2
5.8 | The unit must comply with electrical safety and performance standard such as IEC 60601-2-30, attached certification
5.8.1 | Please provide unique ref number of the IEC ceriticate:
6 | COMPREHENSIVE MAINTENANCE
6.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) should be quoted for, as per the SCC document
6.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
6.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
6.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
6.3 | The bidder to provide the OEM recommended service interval of the unit:
6.4 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
7 | General
7.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
7.2 | The unit on offer must be the latest technology.
7.2.1 | Initial Manufacturing year of Model:
RT2-05-018 | Vital Signs Monitor, on trolley | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point
RT2-05-018a | Reusable latex-free paediatric cuff
RT2-05-018b | Reusable latex-free adult cuff
RT2-05-018c | Reusable latex-free neonate cuff
RT2-05-018d | Reusable latex-free obese cuff
RT2-05-018e | Adult/Paeds NIBP hose
RT2-05-018f | Neonatal NIBP hose
RT2-05-018g | Finger sensor for paeds
RT2-05-018h | Finger sensor for adults
RT2-05-018i | Finger wrap-around probe for neonates
RT2-05-018j | Ankle wrap-around probe for neonates
RT2-05-018k | Spo2 interface cable
RT2-05-018l | Reusable Adult temperature probe
RT2-05-018m | Reusable Paeds temperature probe
RT2-05-018n | Monitor Front bezel
RT2-05-018o | Rechargeable Battery
RT2-05-018p | NIBP pump module
RT2-05-018q | Roll stand
RT2-05-018r | Finger wrap-around probe for neonates - disposable
RT2-05-018s | Ankle wrap-around probe for neonates - disposable
RT2-05-018t | Disposable neonate cuff
RT2-05-018u | Adult Esophageal/Rectal temperature probe
RT2-05-018v | Paediatric Esophageal/Rectal temperature probe

## Pulse Oximeter, Tabletop
RT2-05-019 | PULSE OXIMETER, Tabletop | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
1 | Description
1.1 | A tabletop, medical device used to non-invasively monitor oxygen saturation (SpO2) and pulse rate in patients.
2 | Product Offered:
2.1 | Make:
2.2 | Model:
2.3 | Initial Manufacturing year of Model:
2.4 | Radiation Control License number:
3 | Technical Specification
3.1 | Dedicated oxygen saturation monitor for paediatric and neonatal use.
3.2 | The design must be flat with rubber footing to provide stability from any disturbance or movements on flat surface or desktop. Therefore any other design that could be confused as a Vital Signs or Patient monitor will not be acceptable
3.3 | The unit must continuously measure and display SpO2 %
3.4 | Oxygen saturation (SpO2 %): Range of measurement: 0 - 100%
3.5 | The accuracy must be better than 2,5 % (70 - 100%) and 5% (50 - 70%).
3.6 | Pulse rate (beats/min): Range of measurement: 20 - 300 bpm
3.7 | The pulse rate accuracy must be better than ± 2 %.
3.8 | An audible bleep should sound for each heartbeat, the pitch of which should vary with the oxygen saturation.
3.9 | Automatic rejection of motion artifacts
3.10 | The unit must have trending for at least 72 hours
3.11 | Self test procedure at switch-on
3.12 | The display must offer clear viewing under all lighting conditions.
3.13 | The following must be displayed:
3.13.1 | Oxygen saturation (SpO2 %)
3.13.2 | Pulse rate (beats/min)
3.13.3 | Alarm limits
3.13.4 | Pulse waveform
3.13.5 | Trend information
3.13.6 | Battery status
3.13.7 | Indication of mains or battery operation
3.14 | The unit must indicate whether the signal to the oximeter sensor is adequate for reliable measurements. If the signal strength or quality is inadequate, this must be clearly displayed.
3.15 | SENSOR:
3.15.1 | Sensor: Dual LED, Infra-red/visible red, which does not heat the point of application to the patient.
3.15.2 | The sensor lead must be approximately 3 metres long.
3.16 | ALARMS:
3.16.1 | All alarms must have default settings
3.16.2 | The alarm parameters must be user-adjustable:
3.16.2a | SpO2 LOW: 50 % - 99 %
3.16.2b | Pulse rate LOW: 40 – 150 b.p.m.
3.16.2c | Pulse rate HIGH: 70 – 250 b.p.m.
3.16.2d | Probe not attached to patient or the pulse oximeter
3.16.3 | Violation of these limits must result in both audible and visual alarms and indication of the parameters affected.
3.16.4 | Adjustable audio alarm volume
3.16.5 | Audio alarm mute: 1 minute. The unit must resume normal alarm functions after the minute has expired.
3.16.6 | Low battery: Alarm at least 10 minutes before battery power is exhausted
3.17 | General
3.17.1 | The unit must include an internal rechargeable battery
3.17.2 | Battery charger to be incorporated in the unit
3.17.3 | The unit must be able to function from the mains supply regardless of the battery condition.
3.17.4 | Battery must have the capacity for the unit to operate fully for at least five hours
3.17.5 | The battery must be charged while the unit is operating from the mains supply.
3.17.6 | The battery must automatically take over in the event of a mains power failure.
3.17.7 | Automatic switch-off after a period of non-use
3.17.8 | The case of the unit must be robust and impact resistant.
3.17.9 | The design must be such that spilled liquids do not gain access to the internal working of the unit.
3.17.10 | The unit must be handed over in full operating order
3.17.11 | A starter-pack of consumables must be supplied with the unit.
4 | Standard Accessories and Consumables
4.1 | Reusable Finger sensor for paeds
4.2 | Reusable Ankle/wrist wrap-around probe for neonates
4.3 | Reusable Finger wrap-around probe for neonates
4.4 | Spo2 interface cable
4.5 | Rechargable Battery
4.6 | Disposable Ankle/wrist wrap-around probe for neonates - pack of 20
5 | Conformity Compliance (Please attach certificate)
5.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
5.2 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
5.2.1 | Please provide unique ref number of the ISO 13485 ceriticate:
5.3 | Model on offer must comply with EU Medical Device Regulation (EU) 2017/745 or equvalent from SAHPRA recognized jurisdictions . Attach a copy of certification
5.3.1 | Please provide unique ref number of the EC ceriticate:
5.4 | Hazardous Substance Act: If this type of equipment / apparatus appears on the schedule of Hazardous substances issued by the Directorate: Health Technology of the Department of Health, a license in terms of the Act on Hazardous Substances (Act 15/1973) must be submitted with the bid / quotation documents. The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.5 | All electrical/electronic medical equipment must be licensed by Radiation Control, where applicable a copy of the license must be submitted The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.5.1 | Radiation Control License number:
5.6 | The unit shall operate from 220V AC/50Hz mains supply. The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade (rubber to with 2 screws) type and it must be a minimum length of three (3) meters. NB the mains cable of the unit tendered for must be SABS colour coded.
5.7 | The equipment quoted for must be protected against electromagnetic interference IEC 60601-1-2
6 | COMPREHENSIVE MAINTENANCE
6.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) should be quoted for, as per the SCC document
6.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
6.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
6.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
6.3 | The bidder to provide the OEM recommended service interval of the unit:
6.4 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
7 | General
7.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
7.2 | The unit on offer must be the latest technology.
7.2.1 | Initial Manufacturing year of Model:
RT2-05-019 | PULSE OXIMETER, Tabletop Accessories and Consumables | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point
RT2-05-019a | Reusable Finger sensor for paeds
RT2-05-019b | Reusable Ankle/wrist wrap-around probe for neonates
RT2-05-019c | Reusable Finger wrap-around probe for neonates
RT2-05-019d | Spo2 interface cable
RT2-05-019e | Rechargable Battery
RT2-05-019f | Front bezel
RT2-05-019g | Roll stand
RT2-05-019h | Disposable Finger sensor for paeds
RT2-05-019i | Disposable Ankle/wrist wrap-around probe for neonates

## CO2 Monitor, NIV
RT2-05-020 | Transcutaneous CO2 Monitor, Non-Invasive | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
1 | Description
1.1 | A device used to measure the level of carbon dioxide (CO2) in the blood without needing to draw blood samples.
2 | Product Offered:
2.1 | Make:
2.2 | Model:
2.3 | Initial Manufacturing year of Model:
2.4 | Radiation Control License number:
3 | Technical Specification
3.1 | Suitable for adult, paediatric and neonatal use.
3.2 | The design must be flat with rubber footing to provide stability from any disturbance or movements on flat surface or desktop.
3.3 | The unit must continuously measure and display PCO2, SpO2 and Pulse Rate
3.4 | Transcutaneous CO2 measurements range 0 - 200 mmHg
3.5 | Unit must filter artifact and other noises from movements
3.4 | Oxygen saturation (SpO2 %) range of measurement must be between 0 - 100%
3.5 | The accuracy must be better than 2,5 % (70 - 100%) and 5% (50 - 70%).
3.6 | Pulse rate (beats/min) range of measurement must be between 20 - 300 bpm
3.7 | The pulse rate accuracy must be better than ± 2 %.
3.8 | An audible bleep should sound for each heartbeat, the pitch of which should vary with the oxygen saturation
3.9 | Automatic rejection of motion artefacts
3.10 | The unit must have trending of parameters for at least 72 hours
3.11 | Self test procedure at switch-on
3.12 | The display must offer clear viewing under all lighting conditions
3.13 | The following must be displayed:
3.13.1 | Oxygen saturation (SpO2 %)
3.13.2 | Pulse rate (beats/min)
3.13.3 | Transcutaneous Carbon Dioxide (tcCO2)
3.13.4 | Alarm limits
3.13.5 | Pulse waveform
3.13.6 | CO2 waveform
3.13.7 | Trend information
3.13.8 | Battery status
3.13.9 | Indication of mains or battery operation
3.14 | The unit must indicate whether the signal to the oximeter sensor is adequate for reliable measurements. If the signal strength or quality is inadequate, this must be clearly displayed.
3.15 | SENSOR:
3.15.1 | Sensor: Digital, measure all parameters and does not heat the point of application to the patient.
3.15.2 | Sensor calibration must be automatic to ensure unit readiness at all times
3.16 | ALARMS:
3.16.1 | All alarms must have default settings
3.16.2 | The alarm parameters must be user-adjustable as follows:
3.16.2.1 | SpO2 LOW: 50 % - 99 %
3.16.2.2 | Pulse rate LOW
3.16.2.3 | Pulse rate HIGH
3.16.2.4 | Probe not attached to patient
3.16.3 | Violation of these limits must result in both audible and visual alarms and indication of the parameters affected.
3.16.4 | Adjustable audio alarm volume
3.16.5 | Audio alarm mute: 1 minute. The unit must resume normal alarm functions after the minute has expired.
3.16.6 | Low battery: Alarm at least 10 minutes before battery power is exhausted
3.17 | General
3.17.1 | The unit must include an internal rechargeable battery
3.17.2 | Battery charger to be built-in to the unit
3.17.3 | The unit must be able to function from the mains supply regardless of the battery condition.
3.17.4 | Battery must have the capacity for the unit to operate fully for at least five hours
3.17.5 | The battery must be charged while the unit is operating from the mains supply.
3.17.6 | The battery must automatically take over in the event of a mains power failure.
3.17.7 | Automatic switch-off after a period of non-use
3.17.8 | The casing of the unit must be robust and impact resistant.
3.17.9 | The design must be such that spilled liquids do not gain access to the internal electronic components of the unit.
3.17.10 | The unit must be handed over in full operating order
3.17.11 | A starter-pack of consumables must be supplied with the unit.
4 | Standard Accessories and Consumables
4.1 | CO2 sensor
4.2 | Finger sensor for paeds
4.3 | Finger sensor for adults
4.4 | Finger wrap-around probe for neonates
4.5 | Spo2 interface cable
4.6 | Adhesive sensor holders (Pk)
4.7 | Rechargeable battery
4.8 | Calibration gas bottle
5 | Conformity Compliance (Please attach certificate)
5.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
5.2 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
5.2.1 | Please provide unique ref number of the ISO 13485 ceriticate:
5.3 | Model on offer must comply with EU Medical Device Regulation (EU) 2017/745 or equvalent from SAHPRA recognized jurisdictions . Attach a copy of certification
5.3.1 | Please provide unique ref number of the EC ceriticate:
5.4 | Hazardous Substance Act: If this type of equipment / apparatus appears on the schedule of Hazardous substances issued by the Directorate: Health Technology of the Department of Health, a license in terms of the Act on Hazardous Substances (Act 15/1973) must be submitted with the bid / quotation documents. The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.5 | All electrical/electronic medical equipment must be licensed by Radiation Control, where applicable a copy of the license must be submitted The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.5.1 | Radiation Control License number:
5.6 | The unit shall operate from 220V AC/50Hz mains supply. The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade (rubber to with 2 screws) type and it must be a minimum length of three (3) meters. NB the mains cable of the unit tendered for must be SABS colour coded.
5.7 | The equipment quoted for must be protected against electromagnetic interference IEC 60601-1-2
5.8 | The unit must comply with electrical safety and performance standard such as IEC 60601-2-23 for Transcutaneous partial pressure monitoring, attached certification
5.8.1 | Please provide unique ref number of the IEC ceriticate:
6 | COMPREHENSIVE MAINTENANCE
6.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) should be quoted for, as per the SCC document
6.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
6.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
6.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
6.3 | The bidder to provide the OEM recommended service interval of the unit:
6.4 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
7 | General
7.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
7.2 | The unit on offer must comprise of the latest technology.
7.2.1 | Initial Manufacturing year of Model:
RT2-05-020 | Transcutaneous CO2 Monitor, Non-Invasive | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point
RT2-05-020a | CO2 Sensor
RT2-05-020b | Finger sensor for paeds
RT2-05-020c | Finger sensor for adults
RT2-05-020d | Finger wrap-around probe for neonates
RT2-05-020e | Spo2 interface cable
RT2-05-020f | CO2 interface cable
RT2-05-020g | Battery
RT2-05-020h | Front bezel
RT2-05-020i | Roll stand
RT2-05-020j | Wall bracket

## Pulse Wave Monitor, NIV
RT2-05-021 | Pulse Wave Monitor, Non-Invasive | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
1 | Description
1.1 | Device used to measure elastic properties of vascular system (Arterial stiffness)
2 | Product Offered:
2.1 | Make:
2.2 | Model:
2.3 | Initial Manufacturing year of Model:
2.4 | Radiation Control License number:
3 | Technical Specification
3.1 | The unit must measure Aortic Pulse Wave Velocity, Central Systolic Blood Pressure and Augmentation Index
3.2 | Additional parameters include:
3.2.1 | Systolic/Diastolic Blood Pressure
3.2.2 | Systolic/Diastolic Area Index
3.2.3 | Diastolic Reflection Area
3.2.4 | Aortic Pulse Pressure
3.2.5 | Ankle-Brachial Index
3.2.6 | Heart Rate
3.2.7 | Mean Arterial Pressure
3.3 | The unit must have trending for at least 24 hours
3.4 | BP
3.4.1 | The unit must measure blood pressure non-invasively.
3.4.2 | Must be suitable for adults, paeds and neonates
3.4.3 | The unit must use the oscillometric method of measurement.
3.4.4 | Modes: Manual, automatic and STAT
3.4.5 | Pressure measurements:
3.4.5.1 | Adult: 30 - 270 mm Hg
3.4.9 | Average measurement time: approximately 30 s
3.4.10 | If measurement exceeds 90 seconds the cuff must deflate, an alarm must sound and likely faults indicated.
3.4.11 | The air hose must be at least 1 metre long.
3.5 | Pressure:
3.4.3 | The unit must use the Piezo-resistive method of measurement.
3.7 | LCD Display:
3.7.1 | The display must be clear.
3.7.2 | The following must be displayed in waveform and numeric:
3.7.2.1 | Aortic Pulse Wave Velocity
3.7.2.2 | Central Systolic Blood Pressure
3.7.2.3 | Augmentation Index
3.7.2.4 | Alarm limits
3.7.2.5 | Trend information
3.7.2.6 | Battery status
3.7.2.7 | Indication of mains or battery operation
3.7.3 | The unit must indicate whether the signal to the oximeter sensor is adequate for reliable measurements. If the signal strength or quality is inadequate, this must be clearly displayed.
3.8 | ALARMS:
3.8.1 | All alarms must have default settings
3.8.2 | The alarm parameters must be user-adjustable:
3.8.3 | Adjustable high and low alarm limits for systolic, diastolic, mean arterial pressure and heart rate.
3.8.8 | Violation of limits must result in both audible and visual alarms and indication of the parameters affected.
3.8.9 | Adjustable audio alarm volume
3.8.10 | Audio alarm mute: 2 minute.
3.8.11 | Low battery: Alarm at least 10 minutes before battery power is exhausted
3.9 | Other system requirements
3.9.1 | The unit must operate from alkaline battery(ies)
3.9.3 | The unit must be offered with quality spare battery(ies)
3.9.10 | The case of the unit must be robust and impact resistant.
3.9.11 | The design must be such that spilled liquids do not gain access to the internal working of the unit.
3.9.12 | The monitor must have a carry case
3.9.13 | The unit must be handed over in full operating order
3.9.14 | A starter-pack of consumables must be supplied with the unit.
4 | Standard Accessories and Consumables
4.1 | Reusable latex-free paediatric cuff
4.2 | Reusable latex-free adult cuff
4.3 | Reusable latex-free neonate cuff
4.4 | Reusable latex-free obese cuff
4.5 | NIBP hose
4.6 | Alkaline Batteries
5 | Conformity Compliance (Please attach certificate)
5.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
5.2 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
5.2.1 | Please provide unique ref number of the ISO 13485 ceriticate:
5.3 | Model on offer must comply with EU Medical Device Regulation (EU) 2017/745 or equvalent from SAHPRA recognized jurisdictions . Attach a copy of certification
5.3.1 | Please provide unique ref number of the EC ceriticate:
5.4 | Hazardous Substance Act: If this type of equipment / apparatus appears on the schedule of Hazardous substances issued by the Directorate: Health Technology of the Department of Health, a license in terms of the Act on Hazardous Substances (Act 15/1973) must be submitted with the bid / quotation documents. The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.5 | All electrical/electronic medical equipment must be licensed by Radiation Control, where applicable a copy of the license must be submitted The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.5.1 | Radiation Control License number:
6 | COMPREHENSIVE MAINTENANCE
6.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) should be quoted for, as per the SCC document
6.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
6.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
6.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
6.3 | The bidder to provide the OEM recommended service interval of the unit:
6.4 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
7 | General
7.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
7.2 | The unit on offer must be the latest technology.
7.2.1 | Initial Manufacturing year of Model:
RT2-05-021 | Vital Signs Monitor, on trolley | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point
RT2-05-021a | Reusable latex-free paediatric cuff
RT2-05-021b | Reusable latex-free adult cuff
RT2-05-021c | Reusable latex-free neonate cuff
RT2-05-021d | Reusable latex-free obese cuff
RT2-05-021e | NIBP hose
RT2-05-021f | Finger sensor for paeds
RT2-05-021g | Finger sensor for adults
RT2-05-021h | Finger wrap-around probe for neonates
RT2-05-021i | Spo2 interface cable
RT2-05-021j | Two reusable temperature probes
RT2-05-021k | Front bezel
RT2-05-021l | Back cover
RT2-05-021m | Battery
RT2-05-021n | NIBP pump/module

## Theater Light, Mobile
RT2-05-022 | THEATRE LIGHT MOBILE | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
1 | DESCRIPTION
1.1 | A mobile, single-head surgical light designed for optimal illumination during surgical procedures.
2 | Product Offered:
2.1 | Make:
2.2 | Model:
2.3 | Initial Manufacturing year of Model:
3 | TECHNICAL SPECIFICATION
3.1 | Unit must be of LED technology.
3.2 | The unit must be rubust, constructed from a corrosion resistant material i.e. Alluminium or equivalent.
3.3 | The unit must comply with the laminar flow thus controlling contamination.
3.4 | Once positioned the light must not drift away.
3.5 | The unit must provide low heat radiation, efficient and stable illuminination.
3.6 | ILLUMINATION TO BE AS FOLLOWS:
3.6.1 | Concentrated spot: At 1 meter distance: The unit must have central illumination of 130 000 LUX - 160 000 LUX.
3.6.2 | The LED life-span should be 40 000 hours and above.
3.6.3 | Diameter of lighthead to be between 60 – 100 cm.
3.6.4 | Depth of the illumination to be ± 80 cm3.
3.6.5 | Colour rendering index not be less than 95 RA
3.7 | ELECTRICAL DATA:
3.7.1 | The unit must have a bulit in electrical protection system .
3.7.2 | The unit must have a built in battery backup which will allow for continuous operation for a minumum period of 4 hrs.
3.8 | MECHANICAL DATA:
3.8.1 | The unit must have a central handle to adjust the positioning of the headlight.
3.8.2 | Handle must be sterilisable. If the bidded item uses a different positioning mechanism, bidder to state how sterilization is achieved
3.8.3 | The unit must be balanced in all positions of the headlight.
3.8.4 | Braking system on all movement parts must be available.
3.9 | CONTROLS
3.9.1 | The lighthead must have controls to allow the user to adjust the light focus/intensity
4 | STANDARD ACCESSORIES
4.1 | Sterilisable light handle
5 | Conformity Compliance (Please attach certificate)
5.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
5.2 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
5.2.1 | Please provide unique ref number of the ISO 13485 ceriticate:
5.3 | Model on offer must comply with EU Medical Device Regulation (EU) 2017/745 or equvalent from SAHPRA recognized jurisdictions . Attach a copy of certification
5.3.1 | Please provide unique ref number of the EC ceriticate:
5.4 | Hazardous Substance Act: If this type of equipment / apparatus appears on the schedule of Hazardous substances issued by the Directorate: Health Technology of the Department of Health, a license in terms of the Act on Hazardous Substances (Act 15/1973) must be submitted with the bid / quotation documents. The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.5 | All electrical/electronic medical equipment must be licensed by Radiation Control, where applicable a copy of the license must be submitted The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.5.1 | Radiation Control License number:
5.6 | The unit shall operate from 220V AC/50Hz mains supply. The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade (rubber to with 2 screws) type and it must be a minimum length of three (3) meters. NB the mains cable of the unit tendered for must be SABS colour coded.
5.7 | The equipment quoted for must be protected against electromagnetic interference IEC 60601-1-2
5.8 | The unit must comply with electrical safety and performance standard such as IEC 60601-2-41 for Transcutaneous partial pressure monitoring, attached certification
5.8.1 | Please provide unique ref number of the IEC ceriticate:
6 | COMPREHENSIVE MAINTENANCE
6.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) should be quoted for, as per the SCC document
6.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
6.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
6.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
6.3 | The bidder to provide the OEM recommended service interval of the unit:
6.4 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
7 | General
7.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
7.2 | The unit on offer must be the latest technology.
7.2.1 | Initial Manufacturing year of Model:
RT2-05-022 | THEATER LIGHT MOBILE Accessories | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point
RT2-05-022a | light handle
RT2-05-022b | Battery
RT2-05-022c | LED replacement

## Dermatoscope
RT2-05-023 | DERMATOSCOPE | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
1 | DESCRIPTION
1.1 | A device for clinical inspection of skin lesions and to prevent dIseases
2 | Product Offered:
2.1 | Make:
2.2 | Model:
2.3 | Initial Manufacturing year of Model:
3 | TECHNICAL SPECIFICATION
3.1 | The device must clearly dispaly patterns of pigmentation, blood vessels, and other features that aren't visible to the naked eye.
3.2 | The focusing system of the unit offered must provide the following:
3.3 | 10x - 16x magnification for recognition of even the smallest details and structures
3.4 | LED illumination with a lifetime rating of not less than 50000 hours
3.5 | Color rendering index must be greater than 80 on a scale of 100
3.6 | Must have both polarised and non-polarised illumination modes without the exchange of contact plates
3.7 | The unit must have an indicator to display battery status
3.8 | When fully charged, the handle must provide minimum of 4 hours continouos operation
3.9 | The unit must be supplied with 2 handles to avoid service interruption
4 | STANDARD ACCESSORIES
4.1 | Dematoscope head
4.2 | 2x Contact plates with scale (Diameter: Large- 23mm and small- 8mm)
4.3 | Dematoscope oil atleast 10ml
4.4 | Hard case
4.5 | 2x Li-Ion rechargeble handle
4.6 | Charger for handles
5 | Conformity Compliance (Please attach certificate)
5.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
5.2 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
5.2.1 | Please provide unique ref number of the ISO 13485 ceriticate:
5.3 | Model on offer must comply with EU Medical Device Regulation (EU) 2017/745 or equvalent from SAHPRA recognized jurisdictions . Attach a copy of certification
5.3.1 | Please provide unique ref number of the EC ceriticate:
5.4 | Hazardous Substance Act: If this type of equipment / apparatus appears on the schedule of Hazardous substances issued by the Directorate: Health Technology of the Department of Health, a license in terms of the Act on Hazardous Substances (Act 15/1973) must be submitted with the bid / quotation documents. The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.5 | All electrical/electronic medical equipment must be licensed by Radiation Control, where applicable a copy of the license must be submitted The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.5.1 | Radiation Control License number:
5.6 | The unit shall operate from 220V AC/50Hz mains supply. The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade (rubber to with 2 screws) type and it must be a minimum length of three (3) meters. NB the mains cable of the unit tendered for must be SABS colour coded.
5.7 | The equipment quoted for must be protected against electromagnetic interference IEC 60601-1-2
6 | COMPREHENSIVE MAINTENANCE
6.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) should be quoted for, as per the SCC document
6.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
6.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
6.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
6.3 | The bidder to provide the OEM recommended service interval of the unit:
6.4 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
7 | General
7.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
7.2 | The unit on offer must be the latest technology.
7.2.1 | Initial Manufacturing year of Model:
RT2-05-023 | DERMATOSCOPE | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point
RT2-05-023a | Dematoscope head
RT2-05-023b | Contact plates with scale (Diameter: Large- 23mm)
RT2-05-023c | Contact plates with scale - 8mm diameter
RT2-05-023d | Dematoscope oil- 10ml
RT2-05-023e | Hard case
RT2-05-023f | Li-Ion rechargeble handle
RT2-05-023g | Charger for handles
RT2-05-023h | Photoadaptor
RT2-05-023i | Adaptor cord/cable
RT2-05-023j | Lens Covers

## Infusion Pump, MRI compatible
RT2-05-024 | INFUSION PUMP, MRI compatible | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
1 | Description
1.1 | Volumetric infusion pump that can be used in MRI scanner rooms with magnets of 1.5 and 3 Tesla with a field strength between 300-400 gauss.
2 | Product Offered:
2.1 | Make:
2.2 | Model:
2.3 | Initial Manufacturing year of Model:
2.4 | Radiation Control License number:
3 | Technical Specification
3.1 | The unit must use a linear peristaltic drive mechanism and non-magnetic. Other drive mechanisms may be offered and details to be provided.
3.2 | The unit must provide user selectable infusion rate from 1 to 999 ml/h, selectable in 1 ml/h increments.
3.3 | The unit must provide user selectable infusion volume from 1 to 9 999 ml, selectable in 1 ml increments.
3.4 | The unit must deliver the selected volume with an accuracy better than 5% throughout the whole range of operation/infusion.
3.5 | Bolus rate, user selectable: 10 - 999 ml/hr
3.6 | Bolus volume, user selectable: 1 - 1000 ml
3.7 | The unit must also be capable of carrying out blood administration.
3.8 | It must not be possible to change the rate while the infusion is in progress, without confirming change.
3.9 | LCD display
3.10 | The following must be clearly displayed on the front panel under all lighting conditions:
3.10.1 | Pump running
3.10.2 | AC mains supply operation
3.10.3 | AC power failure/battery operation
3.10.4 | Infusion rate selected
3.10.5 | Volume selected to be infused
3.10.6 | Volume infused
3.10.7 | Alarm condition and alarm/error messages
3.11 | When an infusion is completed an audible tone must sound.
3.12 | On completion of an infusion there must be a Keep Vein Open (KVO) rate of 1 - 5 ml/h.
3.13 | The unit must provide the following alarms:
3.13.1 | Air in line detection
3.13.2 | Occlusion during infusion (downstream and upstream)
3.13.3 | Occlusion pressure alarms
3.13.3a | Occlusion pressure alarm user selectable (Neonatal): 50 - 150 mmHg
3.13.3b | Occlusion pressure alarm user selectable (Adult): 200 - 500 mmHg
3.13.4 | Low battery (alarm must sound when there is still approximately 10 minutes battery operation available).
3.13.5 | Open door
3.13.6 | Infusion set removed
3.13.7 | Wrong infusion set detection
3.13.8 | Completion of selected volume to be infused
3.13.9 | Zero infusion rate
3.13.10 | Zero volume selected for infusion
3.13.11 | Internal malfunction
3.14 | All alarm conditions must be accompanied by an audible tone.
3.15 | Free flow protection
3.16 | State clearly if a dedicated infusion set is required to achieve volumetric accuracy.
3.17 | The unit must have a memory, which stores alarm messages that can be recalled by service technicians when carrying out preventative maintenance, repairs or servicing.
3.18 | The unit must have a service mode, to be accessed by service technicians to check important parameters of the unit, without having to first strip the unit for checking/ testing.
3.19 | The casing of the unit must be impact resistant at 1m height.
3.20 | The unit must be provided with a mechanism for attaching to either a mobile drip stand or a gabler rail: NB whichever is requested by the hospital/institution.
3.21 | The 220 V, 50 Hz AC supply to the unit must be fused in the live or both live and neutral.
3.22 | The unit must be fitted with an internal rechargeable battery.
3.23 | In the event of a 220 V mains failure the battery must automatically provide continuity of operation.
3.24 | The battery must have capacity to operate continuously for a minimum of 3 hours at an infusion rate of 125 ml/h.
3.25 | The battery charger for the internal rechargeable battery must be an integral part of the infusion pump.
3.26 | The infusion pump must be provided with circuitry to ensure that the internal rechargeable battery will be protected against over-charging and over-discharging.
3.27 | The infusion pump must be fully operable from the mains supply regardless of the condition of the battery
3.28 | The battery must be automatically charged when the unit is connected to the mains supply.
3.29 | Purge rate (1 - 999 ml/h) Purge volume (1 - 99 ml) (Sci Gr)
3.30 | Unit can be mounted on standard bed/wall rail or mobile pole/stand
3.31 | Suitable for all intravenous infusions of fluids
3.32 | Display/alarm for low/high flow, occlusion, open door, end-of-infusion and built-in battery status
3.33 | Audio visual alarm with silencing feature
3.34 | Stand-by mode when not in use
3.35 | The unit must have mounting bracket for standard bed/wall rail or mobile pole/stand
3.36 | The unit must be handed over in full operating order
4 | STANDARD ACCESSORIES
4.1 | Clear fluids (Primary administration set) minimum 3 needle free access ports (100/box)
5 | Conformity Compliance (Please attach certificate)
5.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
5.2 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
5.2.1 | Please provide unique ref number of the ISO 13485 ceriticate:
5.3 | Model on offer must comply with EU Medical Device Regulation (EU) 2017/745 or equvalent from SAHPRA recognized jurisdictions . Attach a copy of certification
5.3.1 | Please provide unique ref number of the EC ceriticate:
5.4 | Hazardous Substance Act: If this type of equipment / apparatus appears on the schedule of Hazardous substances issued by the Directorate: Health Technology of the Department of Health, a license in terms of the Act on Hazardous Substances (Act 15/1973) must be submitted with the bid / quotation documents. The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.5 | All electrical/electronic medical equipment must be licensed by Radiation Control, where applicable a copy of the license must be submitted The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.5.1 | Radiation Control License number:
5.6 | The unit shall operate from 220V AC/50Hz mains supply. The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade (rubber to with 2 screws) type and it must be a minimum length of three (3) meters. NB the mains cable of the unit tendered for must be SABS colour coded.
5.7 | The equipment quoted for must be protected against electromagnetic interference IEC 60601-1-2
5.8 | The unit must comply with electrical safety and performance standard such as IEC 60601-2-24 for Infusion Pumps, attached certification
5.8.1 | Please provide unique ref number of the IEC ceriticate:
6 | COMPREHENSIVE MAINTENANCE
6.1 | A 24-month warranty inclusive maintenance and service (Year 1 to Year 2) should be quoted for, as per the SCC document
6.1.1 | THIS ENTAILS: A 24-month warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 24-month warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 24-month warranty will commence after formal Commissioning and handover of the equipment.
6.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
6.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
6.3 | The bidder to provide the OEM recommended service interval of the unit:
6.4 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
7 | General
7.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
7.2 | The unit on offer must be the latest technology.
7.2.1 | Initial Manufacturing year of Model:
RT2-05-024 | INFUSION PUMP ACCESSORIES and CONSUMABLES. | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point
RT2-05-024a | Clear fluids (Primary administration set) minimum 3 needle free access ports (100/box)
RT2-05-024b | Blood administration sets (100/box)
RT2-05-024c | TPN sets (with filter) -100/box
RT2-05-024d | Light resistant (sensetive) sets (100/box)
RT2-05-024e | Low sorb sets (100/box)
RT2-05-024f | Burette sets (100/box)
RT2-05-024g | Secondary infusion sets (Rider/ Piggy bag set) - 100/box
RT2-05-024h | Triport/Tripot needle free access
RT2-05-024i | Dosing error software (UPGRADE)
RT2-05-024j | Capablity for labeling drug names for display on the pump (UPGRADE)

## Syringe Driver, Neonatal
RT2-05-025 | SYRINGE DRIVER, Neo | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
1 | DESCRIPTION
1.1 | A programmable infusion device capable of infusing blood or IV administration via a syringe. Suitable for neonatal infusions and occlussion pressure limits
2 | Product Offered:
2.1 | Make:
2.2 | Model:
2.3 | Initial Manufacturing year of Model:
2.4 | Radiation Control License number:
3 | TECHNICAL SPECIFICATION
3.1 | The unit must be distinguished by colour or model name from the adult unit.
3.2 | The unit must be for general drug administration application i.e. not dedicated for any specific drug.
3.3 | A general purpose Syringe Infusion Device must be supplied with all names of the different brands of syringes and extension sets that the unit has been calibrated on. The listed syringe sizes must range from 5-60ml.
3.4 | Must be capable of an adjustable bolus administration up to rate of 1200 ml/hr
3.5 | Overall volumetric accuracy of 2% or better
3.6 | Bidders must state whether this unit can be pre-programmed for a specified drug. If yes, state how many?
3.6.1 | The bidder must state if unit being quoted on is supplied with built in Dose / Volume calculation.
3.6.2 | If it is available ONLY as an optional feature, the bidder must clearly state the price inclusive of V.A.T. of the unit OFFERED with and without this optional feature.
3.6.3 | The unit must automatically sense the syringe size and display visual indication on the front panel of the size of the syringe loaded. It must not be possible to operate the unit without it being loaded with a syringe.
3.6.4 | The Syringe Driver offered must be provided with circuitry, which must ensure that the internal rechargeable battery will be protected against over-charging and over-discharging
3.7 | Rate settings and Controls
3.7.1 | The infusion rate must be selectable in a minimum range of 0.1 to 999.9 ml/h and the infusion rate must be well displayed on the display mounted on the front panel and which must offer excellent visibility under all lighting conditions.
3.7.2 | The unit must also include the delivery of a bolus dose when required during the normal infusion delivery.
3.7.3 | The unit must revert to the previously selected normal infusion delivery after administering a bolus.
3.7.4 | It must not be possible to change the flow rate while the pump is running.
3.7.5 | The design must be such that the pump must first be stopped, flow rate changed and then restarted or alternatively there must be ample safety precaution against unauthorized tampering of any settings.
3.7.6 | If no flow rate / zero flow rate is selected, it must not be possible to start the infusion delivery. The unit must provide a visible alarm within a preset period should the infusion delivery not be started. State this time interval.
3.7.7 | The control panel must be well sealed preventing entry of fluids to the internal workings and circuits of the unit.
3.7.8 | The Control panel must indicate whether 220V-240V, 50Hz A.C. power supply and internal rechargeable battery power is being used.
3.7.9 | An essential feature must be “an infusion in progress” indicator, which will provide the operator with visual indication that the pump is carrying out infusion.
3.7.10 | Alarms
3.7.10.1 | Internal malfunction
3.7.10.2 | Continuous audible and visual alarm for end of travel of syringe plunger
3.7.10.3 | An audible and visible pre-alarm to indicate that the syringe travel plunger is nearing its end
3.7.10.4 | Syringe not properly loaded or removed
3.7.10.5 | The unit must produce a continuous audible and visual alarm for end of travel of syringe plunger.
3.7.10.6 | The unit must also be equipped with an audible and visible pre-alarm and its function must be such that it will indicate when the syringe travel plunger is nearing its end and/or when the set Volume to be infused is approaching final completion.
3.7.10.7 | The maximum occlusion PRESSURE at which the unit will provide an alarm, must be stated by the bidder for the unit being offered.
3.7.10.8 | The TIME to alarm following occlusion at 1 ml/h and all higher rates must be stated by the bidder for the unit being offered.
3.7.10.9 | The bolus following release of occlusion at 1 ml/h and all higher rates (at alarm), must not exceed 0,6 ml.
3.7.10.10 | The unit must have a Low Battery alarm to warn of impending depletion of the internal rechargeable battery power.
3.7.10.11 | It must be possible to temporarily silence the alarm for a period of time not exceeding two minutes at which point in time if the fault condition still exists the unit must automatically provide an audible alarm again.
3.7.10.12 | There must preferably be separate switch controls and indicators for starting and stopping the pump. When the unit is turned off, the previous settings such as volume and dose delivery must be retained in memory and displayed at next switch on.
3.8 | There must be a fast purge mode of operation, which will allow priming of IV line. It must not be possible to operate the unit continuously in this mode.
3.9 | The unit must be designed to operate from both the 220V-240V, 50Hz A.C. single-phase power supply and an internal rechargeable battery. The unit must be fused in both the LIVE and NEUTRAL.
3.10 | The unit must be supplied with an internal rechargeable backup battery, capable of operating the pump for a minimum of two hours in the event of A.C. power supply failure. It is essential that the changeover to battery operation on mains failure be automatic
3.11 | The battery charger and power supply must be built in internally to the unit and units that are offered with battery charger and power supply, as a separate item will not be considered.
3.12 | The unit must be provided complete with a universal attachment for mounting the unit to both a drip stand and Rail or equivalent.
3.13 | The accuracy (long term) measured over 60 min at 1ml/h and all higher rates must be better than ± 5% for the unit being offered.
3.14 | The accuracy (short term) at 5 ml/h must be better than ± 5% of mean on 2 minutes observation window for the unit being offered.
3.15 | Time from start-up to attain 95% of rate set at 1 ml/h and all higher rates must be stated by the bidder for the unit being offered.
3.16 | The unit must be capable of displaying total Volume Infused. Any bolus administered by the unit during normal infusion must be added to the total volume infused and displayed.
3.17 | The unit being offered shall have been subjected to a clinical evaluation. A clinical evaluation report shall accompany the bid offer. If the unit has not been subjected to a clinical evaluation, the bidder must arrange for a clinical evaluation to be carried out, and an evaluation report must be forwarded.
3.18 | The unit must include a service mode that could be accessed by a service technician to enable key parameters of the unit to be checked without dismantling the unit.
3.19 | There must be a test routine for the verification of operation of alarms on the unit being offered.
4 | STANDARD ACCESSORIES
4.1 | Compatible Syringes LUER LOCK Size: 2ml (Pk of 10)
4.2 | Compatible Syringes LUER LOCK Size: 3ml (Pk of 10)
4.3 | Compatible Syringes LUER LOCK Size: 5ml (Pk of 10)
4.4 | Compatible Syringes LUER LOCK Size: 10ml (Pk of 10)
4.5 | Extension sets Clear Fluid (Straight) - 100/box
5 | Conformity Compliance (Please attach certificate)
5.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
5.2 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
5.2.1 | Please provide unique ref number of the ISO 13485 ceriticate:
5.3 | Model on offer must comply with EU Medical Device Regulation (EU) 2017/745 or equvalent from SAHPRA recognized jurisdictions . Attach a copy of certification
5.3.1 | Please provide unique ref number of the EC ceriticate:
5.4 | Hazardous Substance Act: If this type of equipment / apparatus appears on the schedule of Hazardous substances issued by the Directorate: Health Technology of the Department of Health, a license in terms of the Act on Hazardous Substances (Act 15/1973) must be submitted with the bid / quotation documents. The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.5 | All electrical/electronic medical equipment must be licensed by Radiation Control, where applicable a copy of the license must be submitted The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.5.1 | Radiation Control License number:
5.6 | The unit shall operate from 220V AC/50Hz mains supply. The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade (rubber to with 2 screws) type and it must be a minimum length of three (3) meters. NB the mains cable of the unit tendered for must be SABS colour coded.
5.7 | The equipment quoted for must be protected against electromagnetic interference IEC 60601-1-2
5.8 | The unit must comply with electrical safety and performance standard such as IEC 60601-2-24 for Infusion Pumps and controllers, attached certification
5.8.1 | Please provide unique ref number of the IEC ceriticate:
6 | COMPREHENSIVE MAINTENANCE
6.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) must be included in the bid price
6.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
6.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
6.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
6.3 | The bidder to provide the OEM recommended service interval of the unit:
6.4 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
7 | General
7.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
7.2 | The unit on offer must be the latest technology.
7.2.1 | Initial Manufacturing year of Model:
RT2-05-025 | SYRINGE DRIVER -ACCESSORIES and CONSUMABLES | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point
RT2-05-025a | Extension sets Clear Fluid (Straight) - 100/box
RT2-05-025b | Extension sets (Spiral)- 100/box
RT2-05-025c | Extension sets (Light resistant/sensitive)- 100/box
RT2-05-025d | Extension sets Low Sorb- 100/box
RT2-05-025e | Extension sets Blood- 100/box
RT2-05-025f | Neonatal Closed blood set with Filter - 100/box
RT2-05-025g | Compatible Syringes LUER LOCK Size: 2ml (Pk of 10)
RT2-05-025h | Compatible Syringes LUER LOCK Size: 3ml (Pk of 10)
RT2-05-025i | Compatible Syringes LUER LOCK Size: 5ml (Pk of 10)
RT2-05-025j | Compatible Syringes LUER LOCK Size: 10ml (Pk of 10)
RT2-05-025k | Compatible Syringes LUER LOCK Size: 20ml (Pk of 10)
RT2-05-025l | Compatible Syringes LUER LOCK Size: 30ml (Pk of 10)
RT2-05-025m | Compatible Syringes LUER LOCK Size: 50ml (Pk of 10)
RT2-05-025n | Compatible Syringes LUER LOCK Size: 60ml (Pk of 10)

## Syringe Driver, Adult
RT2-05-026 | SYRINGE DRIVER, Paeds,Adult | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
1 | DESCRIPTION
1.1 | A programmable infusion device capable of infusing blood or IV administration via a syringe. Suitable for paediatric and adult infusions and occlussion pressure limits
2 | Product Offered:
2.1 | Make:
2.2 | Model:
2.3 | Initial Manufacturing year of Model:
2.4 | Radiation Control License number:
3 | TECHNICAL SPECIFICATION
3.1 | The unit must be for general drug administration application i.e. not dedicated for any specific drug.
3.2 | A general purpose Syringe Infusion Device must be supplied with all names of the different brands of syringes and extension sets that the unit has been calibrated on. The listed syringe sizes must range from 5-60ml.
3.3 | Must be capable of an adjustable bolus administration up to rate of 1200 ml/hr
3.4 | Overall volumetric accuracy of 2% or better
3.5 | Bidders must state whether this unit can be pre-programmed for a specified drug. If yes, state how many?
3.5.1 | The bidder must state if unit being quoted on is supplied with built in Dose / Volume calculation.
3.5.2 | If it is available ONLY as an optional feature, the bidder must clearly state the price inclusive of V.A.T. of the unit OFFERED with and without this optional feature.
3.5.3 | The unit must automatically sense the syringe size and display visual indication on the front panel of the size of the syringe loaded. It must not be possible to operate the unit without it being loaded with a syringe.
3.5.4 | The Syringe Driver offered must be provided with circuitry, which must ensure that the internal rechargeable battery will be protected against over-charging and over-discharging
3.6 | Rate settings and Controls
3.6.1 | The infusion rate must be selectable in a minimum range of 0.1 to 999.9 ml/h and the infusion rate must be well displayed on the display mounted on the front panel and which must offer excellent visibility under all lighting conditions.
3.6.2 | The unit must also include the delivery of a bolus dose when required during the normal infusion delivery.
3.6.3 | The unit must revert to the previously selected normal infusion delivery after administering a bolus.
3.6.4 | It must not be possible to change the flow rate while the pump is running.
3.6.5 | The design must be such that the pump must first be stopped, flow rate changed and then restarted or alternatively there must be ample safety precaution against unauthorized tampering of any settings.
3.6.6 | If no flow rate / zero flow rate is selected, it must not be possible to start the infusion delivery. The unit must provide a visible alarm within a preset period should the infusion delivery not be started. State this time interval.
3.6.7 | The control panel must be well sealed preventing entry of fluids to the internal workings and circuits of the unit.
3.6.8 | The Control panel must indicate whether 220V-240V, 50Hz A.C. power supply and internal rechargeable battery power is being used.
3.6.9 | An essential feature must be “an infusion in progress” indicator, which will provide the operator with visual indication that the pump is carrying out infusion.
3.7 | Alarms
3.7.1 | Internal malfunction
3.7.2 | Continuous audible and visual alarm for end of travel of syringe plunger
3.7.3 | An audible and visible pre-alarm to indicate that the syringe travel plunger is nearing its end
3.7.4 | Syringe not properly loaded or removed
3.7.5 | The unit must produce a continuous audible and visual alarm for end of travel of syringe plunger.
3.7.6 | The unit must also be equipped with an audible and visible pre-alarm and its function must be such that it will indicate when the syringe travel plunger is nearing its end and/or when the set Volume to be infused is approaching final completion.
3.7.7 | The maximum occlusion PRESSURE at which the unit will provide an alarm, must be stated by the bidder for the unit being offered.
3.7.8 | The TIME to alarm following occlusion at 1 ml/h and all higher rates must be stated by the bidder for the unit being offered.
3.7.9 | The bolus following release of occlusion at 1 ml/h and all higher rates (at alarm), must not exceed 0,6 ml.
3.7.10 | The unit must have a Low Battery alarm to warn of impending depletion of the internal rechargeable battery power.
3.7.11 | It must be possible to temporarily silence the alarm for a period of time not exceeding two minutes. At which point in time if the fault condition still exists, the unit must automatically provide an audible alarm again.
3.7.12 | There must preferably be separate switch controls and indicators for starting and stopping the pump. When the unit is turned off, the previous settings such as volume and dose delivery must be retained in memory and displayed at next switch on.
3.8 | There must be a fast purge mode of operation, which will allow priming of IV line. It must not be possible to operate the unit continuously in this mode.
3.9 | The unit must be supplied with an internal rechargeable backup battery, capable of operating the pump for a minimum of two hours in the event of A.C. power supply failure. It is essential that the changeover to battery operation on mains failure must be automatic
3.10 | The battery charger and power supply must be built in internally to the unit and units that are offered with battery charger and power supply, as a separate item will not be considered.
3.11 | The unit must be provided complete with a universal attachment for mounting the unit to both a drip stand and Rail or equivalent.
3.12 | The accuracy (long term) measured over 60 min at 1ml/h and all higher rates must be better than ± 5% for the unit being offered.
3.13 | The accuracy (short term) at 5 ml/h must be better than ± 5% of mean on 2 minutes observation window for the unit being offered.
3.14 | Time from start-up to attain 95% of rate set at 1 ml/h and all higher rates must be stated by the bidder for the unit being offered.
3.15 | The unit must be capable of displaying Total Volume Infused. Any bolus administered by the unit during normal infusion must be added to the total volume infused and displayed.
3.16 | The unit being offered must have been subjected to a clinical evaluation. A clinical evaluation report must accompany the bid offer. If the unit has not been subjected to a clinical evaluation, the bidder must arrange for a clinical evaluation to be carried out, and an evaluation report must be forwarded.
3.17 | The unit must include a service mode that could be accessed by a service technician to enable key parameters of the unit to be checked without dismantling the unit.
3.18 | There must be a test routine for the verification of operation of alarms on the unit being offered.
4 | STANDARD ACCESSORIES
4.1 | Compatible Syringes LUER LOCK Size: 2ml (Pk of 10)
4.2 | Compatible Syringes LUER LOCK Size: 3ml (Pk of 10)
4.3 | Compatible Syringes LUER LOCK Size: 5ml (Pk of 10)
4.4 | Compatible Syringes LUER LOCK Size: 10ml (Pk of 10)
4.5 | Extension sets Clear Fluid (Straight) - 100/box
5 | Conformity Compliance (Please attach certificate)
5.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
5.2 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
5.2.1 | Please provide unique ref number of the ISO 13485 ceriticate:
5.3 | Model on offer must comply with EU Medical Device Regulation (EU) 2017/745 or equvalent from SAHPRA recognized jurisdictions . Attach a copy of certification
5.3.1 | Please provide unique ref number of the EC ceriticate:
5.4 | Hazardous Substance Act: If this type of equipment / apparatus appears on the schedule of Hazardous substances issued by the Directorate: Health Technology of the Department of Health, a license in terms of the Act on Hazardous Substances (Act 15/1973) must be submitted with the bid / quotation documents. The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.5 | All electrical/electronic medical equipment must be licensed by Radiation Control, where applicable a copy of the license must be submitted The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.5.1 | Radiation Control License number:
5.6 | The unit shall operate from 220V AC/50Hz mains supply. The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade (rubber to with 2 screws) type and it must be a minimum length of three (3) meters. NB the mains cable of the unit tendered for must be SABS colour coded.
5.7 | The equipment quoted for must be protected against electromagnetic interference IEC 60601-1-2
5.8 | The unit must comply with electrical safety and performance standard such as IEC 60601-2-24 for Infusion Pumps and controllers, attached certification
5.8.1 | Please provide unique ref number of the IEC ceriticate:
6 | COMPREHENSIVE MAINTENANCE
6.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) must be included in the bid price
6.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
6.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
6.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
6.3 | The bidder to provide the OEM recommended service interval of the unit:
6.4 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
7 | General
7.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
7.2 | The unit on offer must be the latest technology.
7.2.1 | Initial Manufacturing year of Model:
RT2-05-026 | SYRINGE DRIVER -ACCESSORIES and CONSUMABLES | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point
RT2-05-026a | Extension sets Clear Fluid (Straight) - 100/box
RT2-05-026b | Extension sets (Spiral)- 100/box
RT2-05-026c | Extension sets (Light resistant/sensitive)- 100/box
RT2-05-026d | Extension sets Low Sorb- 100/box
RT2-05-026e | Extension sets Blood- 100/box
RT2-05-026f | Neonatal Closed blood set with Filter - 100/box
RT2-05-026g | Compatible Syringes LUER LOCK Size: 2ml (Pk of 10)
RT2-05-026h | Compatible Syringes LUER LOCK Size: 3ml (Pk of 10)
RT2-05-026i | Compatible Syringes LUER LOCK Size: 5ml (Pk of 10)
RT2-05-026j | Compatible Syringes LUER LOCK Size: 10ml (Pk of 10)
RT2-05-026k | Compatible Syringes LUER LOCK Size: 20ml (Pk of 10)
RT2-05-026l | Compatible Syringes LUER LOCK Size: 30ml (Pk of 10)
RT2-05-026m | Compatible Syringes LUER LOCK Size: 50ml (Pk of 10)
RT2-05-026n | Compatible Syringes LUER LOCK Size: 60ml (Pk of 10)

## Relief Pump, Overlay-180kg
RT2-05-027 | Pressure Relief pump with Mattress (180kg), Overlay | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
1 | DESCRIPTION
1.1 | A medical device designed to provide adequate pressure relief, comfort and support for patients who spend extended periods in bed.
2 | Product Offered:
2.1 | Make:
2.2 | Model:
2.3 | Initial Manufacturing year of Model:
3 | Technical Specifications
3.1 | Mattress
3.1.1 | Mattress shall support patients up to 180 kg.
3.1.2 | Mattress cell configuration shall consist of:
3.1.2.1 | One or more static head cells that must be replaceable
3.1.2.2 | Alternating cells along the length of the mattress, that must be replaceable
3.1.2.3 | The mattress shall provide a lift of approximately 120 mm
3.1.3 | The mattress shall be of modular design allowing for easy and cost effective repairs
3.1.4 | Matress fit any bed with dimensions as follow
3.1.4.1 | Length: 2000mm
3.1.4.2 | Max height at edges must be no less than 115mm
3.1.4.3 | Width: 900mm
3.1.4.4 | Cell height inflated: 115mm
3.1.5 | The alternating operating cycle shall be variable up to 10 minutes (approximately)
3.1.6 | The mattress shall automatically adjust pressures to optimal level according to patient weight, build and position as set by dial on pump
3.1.7 | The mattress shall have a single hand action CPR release and deflate within 20 sec
3.2 | Cover
3.2.1 | The mattress cover shall be two way stretch and loose fitting
3.2.2 | The cover shall be water resistant
3.2.3 | The cover shall be vapour permeable
3.2.4 | The cover shall be bacteriostatic, fungistatic and antimicrobial
3.2.5 | The cover shall resist development of mould and mildew
3.2.6 | The cover shall be completely removable
3.2.7 | The cover shall be fire retardant
3.3 | Pump
3.3.1 | The pump shall be ergonomically designed, slim line and unobtrusive with approximate size of 150 x 270 x 160 (mm)
3.3.2 | The pump shall be whisper quiet and vibration free
3.3.3 | The pump shall have a carry handle
3.3.4 | The pump shall have a pressure control dial for setting of patient weight and position
3.3.5 | The pump shall have a hanging bracket ensuring pump can be fixed to most beds
3.3.6 | The pump shall have a recessed control panel to prevent accidental operations
3.3.7 | Interface pressure measured shall be below 20 mmHg for 50% of the cycle time
3.3.8 | Highest interface pressure shall be 45 mmHg reducing to 5 mmHg during cycle time
3.3.9 | Pump shall be able to distinguish between genuine alarms and transitory patient movement
3.3.10 | Operating cycle: variable, 0-10 minutes
3.4 | Tube Connectors
3.4.1 | Must be a snap fit tube connectors
3.4.2 | Connecting and disconnecting of tubing shall be a single handed operation
3.4.3 | Tubing shall be designed to prevent accidental disconnection
3.4.4 | Tube set shall be designed to eliminate incorrect connection
3.4.5 | Tubing shall be anti-kinking preventing blockage
3.5 | Alarms
3.5.1 | Alarm conditions to be recorded in memory for diagnostic purposes
3.5.2 | Must have a Mute button to silence audible alarms
3.5.3 | The following alarms must be available in both audible and visual:
3.5.4 | Low pressure alarm
3.5.5 | High pressure alarm
3.5.6 | Power failure alarm
3.5.7 | Pump fault alarm
3.6 | Power Supply
3.6.1 | The unit shall have a built in rechargeable battery for alarm indication only
3.6.2 | The unit shall have a built in rechargeable battery for pump operation in case of power failure. The battery must have a 1-hour operation when fully charged. (Preferred)
4 | Standard Accessories
4.1 | Mattress Cell
4.2 | Mattress
4.3 | Mattress cover
4.4 | Pump
5 | Conformity Compliance (Please attach certificate)
5.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
5.2 | OEM must comply and be certified on ISO 13485 quality standards, attach proof of compliance
5.2.1 | Please provide unique ref number of the ISO 13485 ceriticate:
5.3 | Model on offer must comply with EU Medical Device Regulation (EU) 2017/745 or equvalent from SAHPRA recognized jurisdictions . Attach a copy of certification
5.3.1 | Please provide unique ref number of the ceriticate:
5.4 | Hazardous Substance Act: If this type of equipment / apparatus appears on the schedule of Hazardous substances issued by the Directorate: Health Technology of the Department of Health, a license in terms of the Act on Hazardous Substances (Act 15/1973) must be submitted with the bid / quotation documents. The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.5 | All electrical/electronic medical equipment must be licensed by Radiation Control, where applicable a copy of the license must be submitted The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.5.1 | Radiation Control License number:
5.6 | The unit must operate from 220V AC/50Hz mains supply. The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade (rubber to with 2 screws) type and it must be a minimum length of three (3) meters. NB the mains cable of the unit tendered for must be SABS colour coded.
5.7 | The equipment quoted for must be protected against electromagnetic interference IEC 60601-1-2
6 | COMPREHENSIVE MAINTENANCE
6.1 | A 24-month warranty inclusive maintenance and service (Year 1 to Year 2) should be quoted for, as per the SCC document
6.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
6.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
6.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
6.3 | The bidder to provide the OEM recommended service interval of the unit:
6.4 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
7 | General
7.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
7.2 | The unit on offer must be the latest technology.
7.2.1 | Initial Manufacturing year of Model:
RT2-05-027 | Replacement ACCESSORIES and CONSUMABLES (To be priced individually) | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point
RT2-05-027a | Mattress Cell
RT2-05-027b | Mattress
RT2-05-027c | Mattress cover
RT2-05-027d | Pump
RT2-05-027e | Tube connector

## Relief Pump, Replace-180kg
RT2-05-028 | Pressure Relief pump with Mattress (180kg), Replacement | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
1 | DESCRIPTION
1.1 | A medical device designed to provide adequate pressure relief, comfort and support for patients who spend extended periods in bed.
1.2 | The device to be used as replacement to the hospital mattress.
2 | Product Offered:
2.1 | Make:
2.2 | Model:
2.3 | Initial Manufacturing year of Model:
3 | Technical Specifications
3.1 | Mattress
3.1.1 | Mattress must support patients up to 180 kg.
3.1.2 | Mattress cell configuration must consist of:
3.1.2.1 | One or more static head cells that must be replaceable
3.1.2.2 | Alternating cells along the length of the mattress, that must be replaceable
3.1.2.3 | The mattress must provide a lift of approximately 120 mm
3.1.3 | The mattress must be of modular design allowing for easy and cost effective repairs
3.1.4 | Matress fit any bed with dimensions as follow
3.1.4.1 | Length: 2000mm
3.1.4.2 | Max height at edges must be no less than 200mm
3.1.4.3 | Width: 900mm
3.1.4.4 | Cell height inflated: 200mm
3.1.5 | The alternating operating cycle must be variable up to 10 minutes (approximately)
3.1.6 | The mattress must automatically adjust pressures to optimal level according to patient weight, build and position as set by dial on pump
3.1.7 | The mattress must have a single hand action CPR release and deflate within 20 sec
3.2 | Cover
3.2.1 | The mattress cover must be two way stretch and loose fitting
3.2.2 | The cover must be water resistant
3.2.3 | The cover must be vapour permeable
3.2.4 | The cover must be bacteriostatic, fungistatic and antimicrobial
3.2.5 | The cover must resist development of mould and mildew
3.2.6 | The cover must be completely removable
3.2.7 | The cover must be fire retardant
3.3 | Pump
3.3.1 | The pump must be ergonomically designed, slim line and unobtrusive with approximate size of 150 x 270 x 160 (mm)
3.3.2 | The pump must be whisper quiet and vibration free
3.3.3 | The pump must have a carry handle
3.3.4 | The pump must have a pressure control dial for setting of patient weight and position
3.3.5 | The pump must have a hanging bracket ensuring pump can be fixed to most beds
3.3.6 | The pump must have a recessed control panel to prevent accidental operations
3.3.7 | Interface pressure measured must be below 20 mmHg for 50% of the cycle time
3.3.8 | Highest interface pressure must be 45 mmHg reducing to 5 mmHg during cycle time
3.3.9 | Pump must be able to distinguish between genuine alarms and transitory patient movement
3.3.10 | Operating cycle: variable, 0-10 minutes
3.4 | Tube Connectors
3.4.1 | It must be a snap fit tube connectors
3.4.2 | Connecting and disconnecting of tubing must be a single handed operation
3.4.3 | Tubing must be designed to prevent accidental disconnection
3.4.4 | Tube set must be designed to eliminate incorrect connection
3.4.5 | Tubing must be anti-kinking preventing blockage
3.5 | Alarms
3.5.1 | Alarm conditions to be recorded in memory for diagnostic purposes
3.5.2 | Must have a Mute button to silence audible alarms
3.5.3 | The following alarms must be available in both audible and visual:
3.5.3.1 | Low pressure alarm
3.5.3.2 | High pressure alarm
3.5.3.3 | Power failure alarm
3.5.3.4 | Pump fault alarm
3.6 | Power Supply
3.6.1 | The unit must have a built in rechargeable battery for alarm indication only
3.6.2 | The unit must have a built in rechargeable battery for pump operation in case of power failure. The battery must have a 2-hour operation when fully charged. (Preferred)
4 | Standard Accessories
4.1 | Mattress Cell
4.2 | Mattress
4.3 | Mattress cover
4.4 | Pump
5 | Conformity Compliance (Please attach certificate)
5.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
5.2 | OEM must comply and be certified on ISO 13485 quality standards, attach proof of compliance
5.2.1 | Please provide unique ref number of the ISO 13485 ceriticate:
5.3 | Model on offer must comply with EU Medical Device Regulation (EU) 2017/745 or equvalent from SAHPRA recognized jurisdictions . Attach a copy of certification
5.3.1 | Please provide unique ref number of the ceriticate:
5.4 | Hazardous Substance Act: If this type of equipment / apparatus appears on the schedule of Hazardous substances issued by the Directorate: Health Technology of the Department of Health, a license in terms of the Act on Hazardous Substances (Act 15/1973) must be submitted with the bid / quotation documents. The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.5 | All electrical/electronic medical equipment must be licensed by Radiation Control, where applicable a copy of the license must be submitted The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.5.1 | Radiation Control License number:
5.6 | The unit must operate from 220V AC/50Hz mains supply. The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade (rubber to with 2 screws) type and it must be a minimum length of three (3) meters. NB the mains cable of the unit tendered for must be SABS colour coded.
5.7 | The equipment quoted for must be protected against electromagnetic interference IEC 60601-1-2
6 | COMPREHENSIVE MAINTENANCE
6.1 | A 24-month warranty inclusive maintenance and service (Year 1 to Year 2) should be quoted for, as per the SCC document
6.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
6.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
6.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
6.3 | The bidder to provide the OEM recommended service interval of the unit:
6.4 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
7 | General
7.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
7.2 | The unit on offer must be the latest technology.
7.2.1 | Initial Manufacturing year of Model:
RT2-05-028 | Replacement ACCESSORIES and CONSUMABLES (To be priced individually) | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point
RT2-05-028a | Mattress Cell
RT2-05-028b | Mattress
RT2-05-028c | Mattress cover
RT2-05-028d | Pump
RT2-05-028e | Tube connector

## Relief Pump, Overlay- 270kg
RT2-05-029 | Pressure Relief pump with Mattress (270kg), Overlay | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
1 | DESCRIPTION
1.1 | A medical device designed to provide adequate pressure relief, comfort and support for patients who spend extended periods in bed.
2 | Product Offered:
2.1 | Make:
2.2 | Model:
2.3 | Initial Manufacturing year of Model:
3 | Technical Specifications
3.1 | Mattress:
3.1.1 | Mattress must support patients up to 270 kg.
3.1.2 | Mattress cell configuration must consist of:
3.1.2.1 | One or more static head cells that must be replaceable
3.1.2.2 | Alternating cells along the length of the mattress, that must be replaceable
3.1.3 | The mattress must provide a lift of approximately 135 mm
3.1.4 | The mattress must be of modular design allowing for easy and cost effective repairs
3.1.5 | The mattress must fit on any hospital bed
3.1.6 | The alternating operating cycle must be variable up to 10 minutes (approximately)
3.1.7 | The mattress must automatically adjust pressures to optimal level according to patient weight, build and position as set by dial on pump
3.1.8 | The mattress must have a single hand action CPR release and deflate within 15 sec
3.1.9 | Dimensions:
3.1.9.1 | Length: 2000mm
3.1.9.2 | Max height at edges: 135mm
3.1.9.3 | Width: 850mm
3.1.9.4 | Cell height inflated: 135mm
3.2 | Cover:
3.2.1 | The mattress cover must be two way stretch and loose fitting
3.2.2 | The cover must be water resistant
3.2.3 | The cover must be vapour permeable
3.2.4 | The cover must be bacteriostatic, fungistatic and antimicrobial
3.2.5 | The cover must resist development of mould and mildew
3.2.6 | The cover must be completely removable
3.2.7 | The cover must be fire retardant
3.3 | Pump:
3.3.1 | The pump must be ergonomically designed, slim line and unobtrusive with approximate size of 350 x 200 x 100 (mm) {LxHxD}
3.3.2 | The pump must be whisper quiet and vibration free
3.3.3 | The pump must have a carry handle
3.3.4 | The pump must have a pressure control dial for setting of patient weight and position
3.3.5 | The pump must have a hanging bracket ensuring pump can be fixed to most beds
3.3.6 | The pump must have a recessed control panel to prevent accidental operations
3.3.7 | Interface pressure measured must be below 20 mmHg for 50% of the cycle time
3.3.8 | Highest interface pressure must be 45 mmHg reducing to 5 mmHg during cycle time
3.3.9 | Pump must be able to distinguish between genuine alarms and transitory patient movement
3.3.10 | Operating cycle: variable, 0-10 minutes
3.4 | Multiway Tube-set:
3.4.1 | Connecting and disconnecting of tubing must be a single handed operation
3.4.2 | Tubing must be anti-kinking preventing blockage
3.4.3 | Tube-set must be interchangeable thus eliminating incorrect connection
3.5 | Alarms:
3.5.1 | Alarm conditions to be recorded in memory for diagnostic purposes
3.5.2 | Must have a Mute button to silence audible alarms
3.5.3 | The following alarms must be available in both audible and visual:
3.5.4 | Low pressure alarm
3.5.5 | High pressure alarm
3.5.6 | Power failure alarm
3.5.7 | Pump fault alarm
3.6 | Power Supply:
3.6.1 | The unit must have a built in rechargeable battery for alarm indication only
3.6.2 | The unit must have a built in rechargeable battery for pump operation in case of power failure. The battery must have a 1-hour operation when fully charged. (Preferred)
4 | Standard Accessories
4.1 | Mattress Cell
4.2 | Mattress
4.3 | Mattress cover
4.4 | Pump
5 | Conformity Compliance (Please attach certificate)
5.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
5.2 | OEM must comply and be certified on ISO 13485 quality standards, attach proof of compliance
5.2.1 | Please provide unique ref number of the ISO 13485 ceriticate:
5.3 | Model on offer must comply with EU Medical Device Regulation (EU) 2017/745 or equvalent from SAHPRA recognized jurisdictions . Attach a copy of certification
5.3.1 | Please provide unique ref number of the ceriticate:
5.4 | Hazardous Substance Act: If this type of equipment / apparatus appears on the schedule of Hazardous substances issued by the Directorate: Health Technology of the Department of Health, a license in terms of the Act on Hazardous Substances (Act 15/1973) must be submitted with the bid / quotation documents. The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.5 | All electrical/electronic medical equipment must be licensed by Radiation Control, where applicable a copy of the license must be submitted The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.5.1 | Radiation Control License number:
5.6 | The unit must operate from 220V AC/50Hz mains supply. The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade (rubber to with 2 screws) type and it must be a minimum length of three (3) meters. NB the mains cable of the unit tendered for must be SABS colour coded.
5.7 | The equipment quoted for must be protected against electromagnetic interference IEC 60601-1-2
6 | COMPREHENSIVE MAINTENANCE
6.1 | A 24-month warranty inclusive maintenance and service (Year 1 to Year 2) should be quoted for, as per the SCC document
6.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
6.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
6.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
6.3 | The bidder to provide the OEM recommended service interval of the unit:
6.4 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
7 | General
7.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
7.2 | The unit on offer must be the latest technology.
7.2.1 | Initial Manufacturing year of Model:
RT2-05-029 | Replacement ACCESSORIES and CONSUMABLES (To be priced individually) | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point
RT2-05-029a | Mattress Cell
RT2-05-029b | Mattress
RT2-05-029c | Mattress cover
RT2-05-029d | Pump
RT2-05-029e | Tube connector

## Relief Pump, Replace- 270kg
RT2-05-030 | Pressure Relief pump with Mattress (270kg), Replacement | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
1 | DESCRIPTION
1.1 | A medical device designed to provide adequate pressure relief, comfort and support for patients who spend extended periods in bed.
1.2 | The device to be used as replacement to the hospital mattress.
2 | Product Offered:
2.1 | Make:
2.2 | Model:
2.3 | Initial Manufacturing year of Model:
3 | Technical Specifications
3.1 | Mattress:
3.1.1 | Mattress must support patients up to 270 kg.
3.1.2 | Mattress cell configuration must consist of:
3.1.2.1 | One or more static head cells that must be replaceable
3.1.2.2 | Alternating cells along the length of the mattress, that must be replaceable
3.1.3 | The mattress must provide a lift of approximately 135 mm
3.1.4 | The mattress must be of modular design allowing for easy and cost effective repairs
3.1.5 | The mattress must fit on any hospital bed
3.1.6 | The alternating operating cycle must be variable up to 10 minutes (approximately)
3.1.7 | The mattress must automatically adjust pressures to optimal level according to patient weight, build and position as set by dial on pump
3.1.8 | The mattress must have a single hand action CPR release and deflate within 15 sec
3.1.9 | Dimensions:
3.1.9.1 | Length: 2000mm
3.1.9.2 | Max height at edges: 200mm
3.1.9.3 | Width: 850mm
3.1.9.4 | Cell height inflated: 200mm
3.2 | Cover:
3.2.1 | The mattress cover must be two way stretch and loose fitting
3.2.2 | The cover must be water resistant
3.2.3 | The cover must be vapour permeable
3.2.4 | The cover must be bacteriostatic, fungistatic and antimicrobial
3.2.5 | The cover must resist development of mould and mildew
3.2.6 | The cover must be completely removable
3.2.7 | The cover must be fire retardant
3.3 | Pump:
3.3.1 | The pump must be ergonomically designed, slim line and unobtrusive with approximate size of 350 x 200 x 100 (mm) {LxHxD}
3.3.2 | The pump must be whisper quiet and vibration free
3.3.3 | The pump must have a carry handle
3.3.4 | The pump must have a pressure control dial for setting of patient weight and position
3.3.5 | The pump must have a hanging bracket ensuring pump can be fixed to most beds
3.3.6 | The pump must have a recessed control panel to prevent accidental operations
3.3.7 | Interface pressure measured must be below 20 mmHg for 50% of the cycle time
3.3.8 | Highest interface pressure must be 45 mmHg reducing to 5 mmHg during cycle time
3.3.9 | Pump must be able to distinguish between genuine alarms and transitory patient movement
3.3.10 | Pump must comply with internationally recognized safety standards
3.3.11 | Operating cycle: variable, 0-10 minutes
3.4 | Multiway Tube-set:
3.4.1 | Connecting and disconnecting of tubing must be a single handed operation
3.4.2 | Tubing must be anti-kinking preventing blockage
3.4.3 | Tube-set must be interchangeable thus eliminating incorrect connection
3.5 | Alarms:
3.5.1 | Alarm conditions to be recorded in memory for diagnostic purposes
3.5.2 | Mute button to silence audible alarms
3.5.3 | The following alarms must be available in both audible and visual:
3.5.4 | Low pressure alarm
3.5.5 | High pressure alarm
3.5.6 | Power failure alarm
3.5.7 | Pump fault alarm
3.6 | Power Supply:
3.6.1 | The unit must have a built in rechargeable battery for alarm indication only
3.6.2 | The unit must have a built in rechargeable battery for pump operation in case of power failure. The battery must have a 1-hour operation when fully charged. (Preferred)
4 | Standard Accessories
4.1 | Mattress Cell
4.2 | Mattress
4.3 | Mattress cover
4.4 | Pump
5 | Conformity Compliance (Please attach certificate)
5.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
5.2 | OEM must comply and be certified on ISO 13485 quality standards, attach proof of compliance
5.2.1 | Please provide unique ref number of the ISO 13485 ceriticate:
5.3 | Model on offer must comply with EU Medical Device Regulation (EU) 2017/745 or equvalent from SAHPRA recognized jurisdictions . Attach a copy of certification
5.3.1 | Please provide unique ref number of the ceriticate:
5.4 | Hazardous Substance Act: If this type of equipment / apparatus appears on the schedule of Hazardous substances issued by the Directorate: Health Technology of the Department of Health, a license in terms of the Act on Hazardous Substances (Act 15/1973) must be submitted with the bid / quotation documents. The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.5 | All electrical/electronic medical equipment must be licensed by Radiation Control, where applicable a copy of the license must be submitted The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.5.1 | Radiation Control License number:
5.6 | The unit must operate from 220V AC/50Hz mains supply. The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade (rubber to with 2 screws) type and it must be a minimum length of three (3) meters. NB the mains cable of the unit tendered for must be SABS colour coded.
5.7 | The equipment quoted for must be protected against electromagnetic interference IEC 60601-1-2
6 | COMPREHENSIVE MAINTENANCE
6.1 | A 24-month warranty inclusive maintenance and service (Year 1 to Year 2) should be quoted for, as per the SCC document
6.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
6.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
6.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
6.3 | The bidder to provide the OEM recommended service interval of the unit:
6.4 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
7 | General
7.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
7.2 | The unit on offer must be the latest technology.
7.2.1 | Initial Manufacturing year of Model:
RT2-05-030 | Replacement ACCESSORIES and CONSUMABLES (To be priced individually) | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point
RT2-05-030a | Mattress Cell
RT2-05-030b | Mattress
RT2-05-030c | Mattress cover
RT2-05-030d | Pump
RT2-05-030e | Tube connector

## DVT Pump
RT2-05-031 | Deep Vein Thrombosis (DVT) Pump | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
1 | DESCRIPTION
1.1 | A medical device designed to prevent and treat Deep Vein Thrombosis (DVT) by improving blood circulation and reducing blood stasis.
2 | Product Offered:
2.1 | Make:
2.2 | Model:
2.3 | Initial Manufacturing year of Model:
3 | SPECIFICATIONS
3.1 | The unit offered must be robust, lightweight, compact and user friendly.
3.2 | Pump weight must not exceed 5 kg
3.3 | The unit must offer at least uniform or sequential calf and foot compressions
3.4 | The unit must have different pressures to suit the type of compression (Please provide pressures of the unit):
3.4.1 | Calf compression pressure of at least 40 mmHg
3.4.2 | Foot compression pressure of at least 130 mmHg
3.5 | Provide the compression cycle in terms of time it takes to inflate and deflate for:
3.5.1 | Uniform/sequential calf (in sec)
3.5.2 | Foot (in sec)
3.6 | The unit on offer must have a tube of no less than 2m length
3.7 | The unit must have a Li-ion back-up battery
3.8 | The battery should run for at least 4 hours when fully charged
3.9 | The pump shall have a carry handle
3.10 | The pump shall have a hanging bracket ensuring pump can be fixed to most beds
3.11 | Alarms
3.11.1 | Mute button to silence audible alarms
3.11.2 | The following alarms must be available in both audible and visual:
3.11.2.1 | Low pressure alarm
3.11.2.2 | High pressure alarm
3.11.2.3 | Power failure alarm
3.11.2.4 | Pump fault alarm
4 | Standard Accessories and Conusmables
4.1 | Standard foot cuff
4.2 | Large foot cuff
4.3 | Small leg sleeve
4.4 | Medium leg sleeve
4.5 | Large leg sleeve
4.6 | Hose
4.7 | Rechargeable Battery
5 | Conformity Compliance (Please attach certificate)
5.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
5.2 | OEM must comply and be certified on ISO 13485 quality standards, attach proof of compliance
5.2.1 | Please provide unique ref number of the ISO 13485 ceriticate:
5.3 | Model on offer must comply with EU Medical Device Regulation (EU) 2017/745 or equvalent from SAHPRA recognized jurisdictions . Attach a copy of certification
5.3.1 | Please provide unique ref number of the ceriticate:
5.4 | Hazardous Substance Act: If this type of equipment / apparatus appears on the schedule of Hazardous substances issued by the Directorate: Health Technology of the Department of Health, a license in terms of the Act on Hazardous Substances (Act 15/1973) must be submitted with the bid / quotation documents. The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.5 | All electrical/electronic medical equipment must be licensed by Radiation Control, where applicable a copy of the license must be submitted The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.5.1 | Radiation Control License number:
5.6 | The unit must operate from 220V AC/50Hz mains supply. The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade (rubber to with 2 screws) type and it must be a minimum length of three (3) meters. NB the mains cable of the unit tendered for must be SABS colour coded.
5.7 | The equipment quoted for must be protected against electromagnetic interference IEC 60601-1-2
6 | COMPREHENSIVE MAINTENANCE
6.1 | A 24-month warranty inclusive maintenance and service (Year 1 to Year 2) should be quoted for, as per the SCC document
6.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
6.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
6.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
6.3 | The bidder to provide the OEM recommended service interval of the unit:
6.4 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
7 | General
7.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
7.2 | The unit on offer must be the latest technology.
7.2.1 | Initial Manufacturing year of Model:
7.2 | The unit on offer must be the latest technology.
RT2-05-031 | DVT -Accessories and Consumables | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point
RT2-05-031a | Standard foot cuff
RT2-05-031b | Large foot cuff
RT2-05-031c | Small leg sleeve
RT2-05-031d | Medium leg sleeve
RT2-05-031e | Large leg sleeve
RT2-05-031f | Small calf sleeve
RT2-05-031g | Medium calf sleeve
RT2-05-031h | Large calf sleeve
RT2-05-031i | Small thigh sleeve
RT2-05-031j | Medium thigh sleeve
RT2-05-031k | Large thigh sleeve
RT2-05-031l | Hose
RT2-05-031m | Rechargeable Battery

## Ultrasound, Cardiology
RT2-05-032 | Ultrasound machine, 3D or better- Cardiology | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
1 | DESCRIPTION
1.1 | A Point-of-care Ultrasound machine for Adult and Paediatric Cardiac , including transoesofageal and transthoracic cardiac, as well as superficial vascular and nerve imaging, including appropriate ultrasound probes
2 | Product Offered:
2.1 | Make:
2.2 | Model:
2.3 | Initial Manufacturing year of Model:
2.4 | Radiation Control License number:
3 | TECHNICAL SPECIFICATION
3.1 | Basic Machine requirements
3.1.1 | A High-end Cardiac Ultrasound system needs to be offered
3.1.2 | Digital broadband beam-former
3.1.3 | Real-Time Compounding
3.1.4 | Adaptive Image Processing for noise and artifact reduction
3.1.5 | 2D as well as 3D imaging
3.1.6 | Real time 4D (Live 3D) TEE using matrix transducer technology for instantaneous volume rendering of cardiac anatomy .
3.1.7 | Live Biplane Imaging using Matrix transducers (infinite plane adjustments)
3.1.8 | New Single crystal technology or Multiple crystal on transthoracic and Transoesophageal transducers
3.1.9 | The equipment must be able to function in the physical surroundings of theatre in temperature range of at least 10 to 35 degrees Celsius
3.1.10 | Ergonomic transducers designs with lightweight flexible cables
3.1.11 | High-resolution interactive graphical colour touch panel with adjustment for various ambient light conditions
3.1.12 | 3 Active transducer connectors and one non Imaging transducer port
3.1.13 | Intelligent one-button optimization in 2D and Doppler modes
3.2 | IMAGING MODALITIES
3.2.1 | B/Mode 2D
3.2.1.1 | Scan methods:
· Linear scan (state if transducers can perform oblique scanning)
· Sector scan
· Convex scan
3.2.1.2 | Viewing depth:
· Maximum >30 cm, state scanning depth for specific probes
3.2.1.3 | Line density:
· The line density differs depending on the transducer used. Several line densities are available for selection for each transducer.
3.2.1.4 | Scan angle and scan width:
· Adjustment of scan angle (scan width) is possible.
· Adjustment of scan steering (scanning position adjustment) is possible.
· Adjustment of linear scan steering is possible.
· Trapezoid scanning providing convex view on linear probes .
3.2.1.5 | PAN/Zoom:
· Scale can be enlarged or reduced by the encoder.
· The panning position can be arbitrarily selected using the trackball.
· HD zoom: Available (real-time only)
· Spot zoom: Available
· High Definition Write Zoom and Read Zoom with pan feature
3.2.1.6 | Focus:
· Transmission: Maximum 8 steps
· Reception: Continuous
· Auto-focusing: Available
3.2.1.7 | Transmission frequency:
· Image frequency needs to be 0 to 20 MHz
· It differs for different Ultrasound probes- eg Multi Array Cardiac probe 1-5 MHz, Transoesophageal probe 2-7 MHz, Linear hockey stick-like probe 7-15 MHz
3.2.1.8 | Gain:
· Display brightness for B-mode display can be changed.
3.2.1.9 | Dynamic range above 270dB
3.2.1.10 | System Grey scale level needs to be± 256 shades
3.2.1.11 | 2D loop up to 2 000 frames must be recorded
3.2.1.12 | Indicate maximum length of captured loop in heart beats and minutes
3.2.1.13 | Real time triplex mode must be standard
3.2.1.14 | Frame Rate: supports high frame rate
3.2.1.15 | Gain control one button
3.2.1.16 | STC (TGC ): 8-step slide controls (common to B and M modes) (Time Gain Compensation – lateral gain)
3.2.1.17 | Acoustic output adjustment: Adjustable up to 100%
3.2.1.18 | Colour Palette (2D MAP): Available
3.2.1.19 | THI (Tissue Harmonic Imaging):state if available on all probes
· Pulse Subtraction : Available
· Differential Harmonic Imaging: Available
3.2.1.20 | Compounding Imaging: state if supported
· Frequency and/or Spatial Compounding imaging in real-time is supported.
3.2.1.21 | Display orientation:
· Top/bottom reversal is possible
· Left/right reversal is possible
3.2.1.22 | Panoramic View: state if supported
· Panoramic View is supported
· Measurement is available
3.2.1.23 | One button optimization:
· Enables automatic gain and STC/TGC control for B mode as well as Velocity range and Base line shift control for Spectrum Doppler
3.2.2 | Spectral Doppler
3.2.2.1 | Doppler mode:
· PWD (Pulsed-Wave Doppler) and Colour Doppler available on all probes
· CWD (Continuous-Wave Doppler) available on all cardiac sector probes
3.2.2.2 | Doppler pulse repetition frequency (PRF)
· State PRF frequency
3.2.2.3 | Doppler scan:
· B/C/D simultaneous scan
· State if simultaneous display of Doppler image , side by side, in real-time is possible
3.2.2.4 | Doppler sampling volume:
· The Doppler range gate width can be changed.
· Sample volume size ± 0,5 – 20 mm
3.2.2.5 | Sample shift:
· From 0 cm to the maximum depth
3.2.2.5 | Doppler filter:
· Doppler cut-off frequency can be changed.
3.2.2.6 | Doppler gain:
· Doppler image display brightness can be changed
· Independent gain control for PW and CW
3.2.2.7 | Indication of Doppler spectrum direction:
· Display of the forward and backward velocity spectrum can be reversed
3.2.2.8 | Doppler baseline shift (zero-shift):
· The baseline (zero-velocity line) of Doppler images can be shifted.
· The baseline shift setting can also be adjusted when images that were frozen are displayed.
3.2.2.9 | Doppler audio:
· Stereo output (forward flow and backward flow)
3.2.2.10 | Colour Palette (Doppler MAP): Available to be selected when frozen
3.2.2.11 | Image memory to operate in both dual and quad screen format
3.2.2.12 | Display of Doppler scale:
· 2 Types (velocity and Doppler shift frequency)
3.2.2.13 | Doppler focus:
· Automatically follows the sample position
3.2.2.14 | Doppler angle display
3.2.2.15 | Oblique Doppler scan:
· Oblique scan is possible with some linear transducers.
3.2.2.16 | Doppler multi-frequency:
· Doppler transmission and reception frequencies are selectable in PWD.
3.2.2.17 | Doppler sweep speed:
· Sweep speed can be changed
3.2.2.18 | Post processing must be available on frozen images
3.2.2.19 | Doppler Auto trace:
· Waveform Peak, Mean, Peak + Mean
· Trace area specified: Toward, Reverse, All
· Measurement item: Max, Min, Mean, Pl, RI, etc.
· One button optimization technology.
3.2.2.20 | Tissue Doppler- TDI
3.2.2.21 | M-mode must be available, with Anatomical M-Mode
3.2.2.22 | Live (on screen) automatic Doppler Analysis
3.2.3 | COLOUR DOPPLER
3.2.3.1 | Colour Doppler mode:
· Flow velocity
· Flow velocity/variance
3.2.3.2 | Power Imaging mode: state if available
· On all transducers
· Advanced Dynamic Flow mode (high-resolution flow imaging)
3.2.3.3 | Colour Doppler:
· State maximum colour frame rate
· M-mode must be available
3.2.3.4 | Doppler velocity range: to be stated
3.2.3.5 | Colour Doppler frame interpolation (C Frame Interpolation) Automatically set according to the transducer information and selected exam conditions .
3.2.3.6 | Colour Doppler baseline (C Baseline):
· The zero-velocity line of colour Doppler images can be shifted.
· Zero-shift setting can also be adjusted when an image is frozen or an image in the cine memory is displayed.
3.2.3.7 | Colour reversal is available
3.2.3.8 | Balance between B/W and colour display. Colour weighting to B/W image can be set by comparing the colour Doppler images and B/W images.
3.2.3.9 | Parallel colour Doppler processing. Simultaneous signal reception in a maximum of 4 directions (QSP) is possible.
3.2.3.10 | Colour gain:
3.2.3.11 | · Display brightness of colour Doppler images can be changed.
3.2.3.12 | Colour Doppler multi-frequency
3.2.3.13 | Transmission frequency for colour Doppler image acquisition can be changed.
3.2.3.14 | Colour Doppler line density
3.2.3.15 | Line density of colour Doppler images can be changed.
3.2.3.16 | Adjustment of colour ROI position, size, and angle is possible. Must be fully operational through the total depth of image
3.2.3.17 | Colour Doppler transmission focus
3.2.3.18 | Automatically follows the colour Doppler ROI position.
3.2.3.19 | Colour Doppler filter
3.2.3.20 | Filter cut-off can be changed
3.2.3.21 | ROI filter function
3.2.3.22 | Simultaneous dual-screen display with B mode
3.2.3.23 | Cineloop review possible
3.2.3.24 | On screen measurements must be possible from real time examination
3.2.3.25 | Unit to offer programmable quick annotation
3.2.3.26 | Indicate if Directional Colour Angio is available
3.2.3.27 | Trapezoid scanning providing convex view on linear probes.
3.3 | MEASUREMENT and ANALYSIS
3.3.1 | ±16 callipers (8 channels) for measuring must be offered
3.3.2 | The operating modes must include: 2D, M-Mode, Anatomic M-mode, Colour Flow Doppler, Pulsed Wave Doppler (PW), High PRF PW, Continuous Wave Doppler, Chroma Imaging, Tissue Doppler Imaging, Pulse Inversion, Cineloop Image, M-mode and Doppler Review, 4D / Live 3D, triplex and Harmonic with pulse inversion technology.
3.3.3 | Basic measurements:
3.3.3.1 | · Measurement tools including: distance, depth, area, and circumference
3.3.3.2 | · B-mode measurements
3.3.3.3 | · Distance
3.3.3.4 | · Area: using ellipse or trace
3.3.3.5 | · Angle: intersection of two lines
3.3.3.6 | · Volume flow Application measurements auto labelling on touch screen.
3.3.3.7 | · Comprehensive reports generated automatically.
3.3.3.8 | · Maximum Dynamic Range of 280dB
3.3.3.9 | · Real-Time Compounding
3.3.3.10 | · Adaptive Image Processing for noise and artifact reduction
3.3.3.11 | · The latest version of vascular, adult & child cardiac- (transthoracic and transoesophageal - preferably 4D / Live 3D Transesophageal Echocardiography), lung- and regional anaesthesia software packages must be included
3.3.4 | Spectral Doppler measurements:
3.3.4.1 | · Velocity: manual
3.3.4.2 | · Acceleration
3.3.4.3 | · Time /slope
3.3.4.4 | · Heart rate
3.3.4.5 | · PI
3.3.4.6 | · RI
3.3.4.7 | · Flow volume
3.3.4.8 | · Doppler trace
3.3.4.9 | · Automatic Doppler analysis
3.3.5 | Cardiac imaging & measurement
3.3.5.1 | · Tissue specific imaging software for adult and paediatric echo ultrasound applications
3.3.5.2 | · Display optimisation software with tissue specific pre-sets for adult and paediatric echo imaging and Doppler application
3.3.5.3 | · Analysis software for full package of adult and paediatric echo imaging protocol and report (to include standard and 3D quantification on TEE)
3.3.5.4 | · Tissue Harmonics and Pulse Inversion Harmonic Imaging
3.3.5.5 | · High Definition Write Zoom and Read Zoom with pan features
3.3.5.6 | · Chroma Imaging
3.3.5.7 | · Measurement tools including: distance, depth, area, and circumference
3.3.5.8 | · Allows operation with all the desired transducers
3.3.5.9 | · Speckle tracking must be offered as an optional functionality
3.3.5.10 | · For 3/4D:
3.3.5.11 | · 4D / Live 3D Volume Imaging using matrix transducer
3.3.5.12 | · 4D / Live 3D Zoom
3.3.5.13 | · 4D / Live 3D Full volume
3.3.5.14 | · 4D / Live 3D Quantification with MPR and islice.
3.3.5.15 | · 4D / Live 3D display of dynamic 3D rendering and LV true volumes.
3.3.5.16 | · Full Quantification package for 3D cardiac volume datasets
3.3.6 | Vascular Access & Regional anaesthesia
3.3.6.1 | · Tissue pre-sets for vascular access and regional anaesthesia must be available
3.3.6.2 | · Imaging software for vascular access (eg jugular vein, radial artery, femoral vessels).
3.3.6.3 | · The imaging must also be suitable for regional anaesthesia
3.3.7 | Lung imaging
3.3.8 | · Software must be included for optimising lung imaging
3.3.9 | Tissue Doppler:
3.3.8.1 | · Tissue Doppler with high frame rate
3.3.8.2 | · M mode colour Doppler and M mode tissue Doppler
3.3.9 | Worksheet/ Report functions:
3.3.9.1 | · Fully editable results datasheet
3.3.9.2 | · Integrated patient exam report, editable
3.3.9.3 | · Embed images in report
3.3.9.4 | · Customizable to workflow
3.3.9.5 | · On-screen recall, measurement, text editing
3.3.9.6 | · Exam directory
3.3.9.7 | · Selection between mean value display and latest value display is possible.
3.3.9.8 | · Comment entry is possible
3.3.9.9 | · The created reports must be printable
3.3.9.10 | · The created reports must preferably be output as PDF files. This must also be available for cardiac, lung, and regional anaesthesia reports
3.3.10 | Software. As described elsewhere, but to include:
3.3.10.1 | · Adult cardiac including TTE & TEE 3D & live 3D TEE as well as lung ultrasound
3.3.10.2 | · Paediatric cardiac, including TTE
3.3.10.3 | · Vascular access & regional anaesthesia
3.3.10.4 | · Dicom 3 software must be included
3.3.10.5 | Software must be ugradeable and future upgrades included in the offer
3.3.10.6 | Note: Paediatric 3D cardiac echography is desirable and an alternative system with appropriate software and probes must also be tendered if available.
3.4 | MEMORY
3.4.1 | Cine/ storage:
3.4.1.1 | · Memory capacity: (Minimum 2200 frames)
3.4.1.2 | · Internal HDD of at least 512GB
3.4.1.3 | · Record/playback modes
3.4.1.4 | · Loop playback is possible
3.4.1.5 | · Frame-advance playback is possible
3.4.1.6 | · Cine playback is possible for Doppler images and M-mode images. (Minimum 30 s )
3.4.1.7 | · Clips: Live images can be recorded
3.4.2 | Archiving:
3.4.2.1 | · Hard disk drive to exceed 512 MB
3.4.2.2 | · Recorded data on CD\DVD must be readable on standard laptop\computer (with additional software from company included if necessary
3.4.2.3 | · Integrated DVD/CD burning capability for storage of DICOM images or export in JPEG and avi for PC compatibility
3.4.2.4 | · USB flash drive
3.4.2.5 | · Network: DICOM : must be compliant and must be included
3.4.2.6 | · Advanced DICOM with full work list must be offered and included
3.4.2.7 | · DICOM 3.0 Print and Store capability to internal drive or DVD/CD
3.4.2.8 | · Export of images from system must be possible, state methods
3.5 | TRANSDUCERS
All transducers required must be included in the main offer as part of the total price
3.5.1 | Transducer selection:
3.5.1.1 | · 4 Active imaging transducer connectors and one non Imaging transducer port
3.5.1.2 | · Electronic switching between transducers using at least 4 universal connectors
3.5.1.3 | · Automatic pre-set optimization of each transducer
3.5.1.4 | · High-resolution interactive graphical colour touch panel with adjustment for various ambient light conditions
3.5.2 | Transducer design:
3.5.2.1 | · Ergonomic design with lightweight flexible cables
3.5.2.2 | · Advanced penetration with less artifacts
3.5.2.3 | · Broadband frequency
3.5.2.4 | · Support for high frequencies
3.5.2.5 | · Selected frequency to be displayed on screen
3.5.2.6 | · All transducers be digitally controlled so that continuous beam forming is possible
3.5.2.7 | · Multi array cardiac transducer 1-5 MHz for Cardiac Echography as well as lung ultrasound
3.5.2.8 | · A 2 to 7 MHz Transoesophageal 3D echocardiography probe with live 3D software (adults)
3.5.2.9 | · Sector array with 48 elements
3.5.2.10 | · Manually rotatable 0-180°
3.5.2.11 | · SD steerable with PW Doppler, CW Doppler, colour Doppler, XRES and harmonic imaging
3.5.2.12 | · 5-12 MHz broadband phased array with small footprint for paediatric cardiac and general applications
3.5.2.13 | · SD steerable with PW Doppler, CW Doppler, colour Doppler, XRES and harmonic imaging
3.5.2.14 | · A 3 to 7 MHz paediatric transoesophagial (TEE) probe
3.5.2.15 | · Manually rotatable 0-180°
3.5.2.16 | · SD steerable with PW Doppler, CW Doppler, colour Doppler, XRES and harmonic imaging
3.5.3 | Broadband linear array transducer:
3.5.3.1 | · Hockey stick-like design
3.5.3.2 | · Supports vascular carotid, arterial, venous and regional anaesthesia superficial imaging applications.
3.5.3.3 | · Minimum 7 to 15 MHz frequency range
3.5.3.4 | · Supports PW, colour Doppler, power Doppler
3.5.3.5 | · Supports Strain & Strain Rate and Speckle Tracking
3.5.3.6 | Paediatric 3D Cardiac echocardiography is desirable and an alternative system capable of this must be tendered with appropriate probes
3.6 | SYSTEM CONTROLS\ DESIGN
3.6.1 | Fully digital, high end, colour Doppler , ultrasound unit
3.6.2 | On-board workstation-class data management with thumbnail previews and storage of images, loops, and reports
3.6.3 | System must have 12 “ or larger touch command screen
3.6.4 | Display screen:
3.6.4.1 | · Nearly infinite positioning adjustments for optimum scanning, height, swivel, tilt
3.6.4.2 | · Up and down
3.6.4.3 | · Side to side movement
3.6.4.4 | · 20” wide
3.6.4.5 | · flicker free
3.6.5 | Control panel:
3.6.5.1 | · Retractable alphanumeric keyboard
3.6.5.2 | · Palm rest
3.6.5.3 | · Reduced number of hard controls
3.6.5.4 | · Primary controls clustered for convenience
3.6.5.5 | · Ambient lighting for different environments
3.6.5.6 | · Intelligent one-button optimization in 2D and Doppler modes eg Iscan
3.6.6 | Touch Screen:
3.6.6.1 | · Workflow related controls
3.6.6.2 | · Quick access to hidden controls
3.6.6.3 | · Supports alphanumeric keyboard option
3.6.6.4 | · High-resolution interactive graphical colour touch panel with adjustment for various ambient light conditions
3.6.7 | Pre-sets:
3.6.7.1 | · Single selection to choose pre-set, transducers
3.6.7.2 | · Completely editable pre-sets for specific exams
3.7 | PHYSICAL SPECIFICATIONS
3.7.1 | System Cart:
3.7.1.1 | · State of the art ergonomic design for comfort and convenience
3.7.1.2 | · Easy manoeuvrability and mobility
3.7.1.3 | · Transducer and gel bottle holders
3.7.1.4 | · High quality, shock absorbing castors with wheel lock facility and 4 wheel swivel
3.7.1.5 | · Automatic cooling capacity with minimal noise
3.7.2 | Monitor:
3.7.2.1 | · Flat screen LCD monitor at least 20 inch
3.7.2.2 | · Extended viewing angle
3.7.2.3 | · Mounted on fully articulating extension arm
3.7.3 | ECG:
3.7.3.1 | · A 3 lead ECG cable must be provided
3.7.3.2 | · Lead l is the standard connection
3.7.3.3 | · External input is possible
3.7.3.4 | · Lead selection is possible
3.7.3.5 | · Automatic heart rate calculation and display
3.7.4 | Peripherals:
3.7.4.1 | · System supports up to two on-board peripheral devices such as DVD, colour printer (USB), B/W printer (USB)
3.7.4.2 | · Support various colour printers externally mounted
3.7.4.3 | · Automated measurement tool for Speckle tracking (optional)
3.7.5 | Input/output ports:
3.7.5.1 | · Transducer Connectors: minimum 4 imaging
3.7.5.2 | · DVD input/output signals
3.7.5.3 | · USB ports
3.7.5.4 | · External video output signal
3.7.6 | Printer
3.8 | Power and Batteries
3.8.1 | The monitor must operate off both the 220VV , 50Hz single phase AC power supply and an internally fitted rechargeable battery. The unit must be fused in both the LIVE and NEUTRAL.
3.8.2 | The unit offered must be supplied with an internal rechargeable battery.
3.8.3 | The rechargeable battery must be automatically charged when the unit is connected to a live A.C. supply
3.8.4 | In the event of A.C. mains failure the battery must automatically provide continuous operation
3.8.5 | The unit must be provided with visible indication to reflect whether the unit is being operated off the AC power or the rechargeable battery power
3.8.6 | The unit offered must be provided with circuitry, which must ensure that the rechargeable battery will be protected against over-charging and over-discharging
3.8.7 | Units that are quoted on and employ external battery chargers and power supply will not be considered.
3.8.8 | Battery back-up option must be provided (battery allowing unit to be placed in a sleep mode while being moved from one location to the next so as to provide rapid resumption of operation)
3.8.9 | The rechargeable battery must be of a reasonable capacity, such that with battery power the unit must be able to operate continuously for a minimum of four (4)hours
3.8.9.1 | UPS option in the event that offered unit's battery capacity does not meet the 4-hour requirement
3.8.9.2 | Please state the following:
3.8.9.2.1 | Capacity
3.8.9.2.2 | Voltage
3.8.9.2.3 | Type of battery used
3.8.9.2.4 | Length of battery operating time
4 | STANDARD ACCESSORIES
4.1 | 3 lead ECG cable
4.2 | Multi array cardiac transducer 1-5 MHz for Cardiac Echography as well as lung ultrasound
4.3 | 2 to 7 MHz Transoesophageal 3D echocardiography probe with live 3D software (adults)
4.4 | 5-12 MHz broadband phased array with small footprint for paediatric cardiac and general applications
4.5 | Ultrasound Gel
4.6 | Printer
4.7 | Printing paper (1 box)
5 | Conformity Compliance (Please attach certificate)
5.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
5.2 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
5.2.1 | Please provide unique ref number of the ISO 13485 ceriticate:
5.3 | Model on offer must comply with EU Medical Device Regulation (EU) 2017/745 or equvalent from SAHPRA recognized jurisdictions . Attach a copy of certification
5.3.1 | Please provide unique ref number of the EC ceriticate:
5.4 | Hazardous Substance Act: If this type of equipment / apparatus appears on the schedule of Hazardous substances issued by the Directorate: Health Technology of the Department of Health, a license in terms of the Act on Hazardous Substances (Act 15/1973) must be submitted with the bid / quotation documents. The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.5 | All electrical/electronic medical equipment must be licensed by Radiation Control, where applicable a copy of the license must be submitted The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.5.1 | Radiation Control License number:
5.6 | The unit shall operate from 220V AC/50Hz mains supply. The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade (rubber to with 2 screws) type and it must be a minimum length of three (3) meters. NB the mains cable of the unit tendered for must be SABS colour coded.
5.7 | The equipment quoted for must be protected against electromagnetic interference IEC 60601-1-2
5.9 | The unit must comply with electrical safety and performance standard such as IEC 60601-2-37, attached certification
5.9.1 | Please provide unique ref number of the IEC ceriticate:
6 | COMPREHENSIVE MAINTENANCE
6.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) must be quoted for, as per the SCC document
6.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
6.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
6.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
6.3 | The bidder to provide the OEM recommended service interval of the unit:
6.4 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
7 | General
7.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
7.2 | The unit on offer must be the latest technology.
7.2.1 | Initial Manufacturing year of Model:
RT2-05-032 | Ultrasound machine, 3D or better- Cardiology Accessories and Consumables | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point
RT2-05-032a | Uninterrupted Power Supply (UPS)
RT2-05-032b | Hockey stick-like broadband linear array transducer 7-15 MHz
RT2-05-032c | 3 to 7 MHz paediatric transoesophagial (TEE) probe
RT2-05-032d | 3 lead ECG cable
RT2-05-032e | Multi array cardiac transducer 1-5 MHz for Cardiac Echography as well as lung ultrasound
RT2-05-032f | 2 to 7 MHz Transoesophageal 3D echocardiography probe with live 3D software (adults)
RT2-05-032g | 5-12 MHz broadband phased array with small footprint for paediatric cardiac and general applications
RT2-05-032h | Printing paper
RT2-05-032i | Ultrasound Gel

## CO Monitor
RT2-05-033 | CARDIAC OUTPUT MONITOR | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
1 | DESCRIPTION
1.1 | A dedicated monitor to measure cardiac output in patients and facilitate advanced haemodynamic monitoring
2 | Product Offered:
2.1 | Make:
2.2 | Model:
2.3 | Initial Manufacturing year of Model:
2.4 | Radiation Control License number:
3 | TECHNICAL SPECIFICATION
3.1 | A monitor that can measure cardiac output (and oximetry when invasive) to facilitate advanced haemodynamic monitoring
3.2 | The monitor must have noninvasive and/or invasive technology options (quoted separately with indication of interchangeability)
3.3 | The monitor must be light weight, not more than 1kg. Specity weight
3.4 | The monitor must have a colour display with software to make physiological interpretation easy
3.5 | The monitor must have an appropriate stand/pole/trolley with all necessary components (eg brackets)
3.6 | Display
3.6.1 | LCD/LED colour display screen of at least 15 cm diagonal diameter
3.6.2 | Resolution 600 x 800 pixels
3.6.3 | Viewable at at least 80 degree angle
3.6.4 | Must be durable and cleanable
3.6.5 | Speaker count: 2
3.6.6 | Data security clearance certificate - attach
3.7 | Environmental requirements
3.7.1 | Must perform at temperatures at least 5 to 35 ⁰C
3.7.2 | Operational humidity: up to 90% at 40 ⁰C
3.8 | Parameters to be measured with non-invasive unit (Specify all components necessary for functional non-invasive measurement)
3.8.1 | Cardiac Output (CO) Continuous measurement of the volume of blood pumped by the heart measured in liters per minute (0-12 L/min)
3.8.2 | Cardiac Index (CI) Cardiac output relative to body surface area (BSA) (0-12 L/min/m2)
3.8.3 | Diastolic Pressure (DIA) Diastolic blood pressure (50-110 mmHg)
3.8.4 | Mean Arterial Pressure (MAP) Averaged systemic blood pressure over one cardiac cycle (50-130 mmHg)
3.8.5 | Pulse Rate (PR) Number of arterial blood pressure pulses per minute (40-130 bpm)
3.8.6 | Stroke Volume (SV) Volume of blood pumped with each heart beat (0-160 mL/b)
3.8.7 | Stroke Volume Index (SVI) Stroke volume relative to body surface area (BSA) (0-8 mL/b/m2)
3.8.8 | Systemic Vascular Resistance (SVR) The resistance that the left ventricle must overcome to eject stroke volume with each beat (500-1500 dyne-sec/cm5)
3.8.9 | Systemic Vascular Resistance Index (SVRI) SVR relative to body surface area (500-3000 dyne-sec-m2/cm5)
3.8.10 | Stroke Volume Variation (SVV) The percent difference between SVmin, max and mean (0-50%)
3.8.11 | Systolic Pressure (SYS) Systolic blood pressure (80-160 mmHg)
3.9 | Parameters to be measured with invasive unit (Specify all components necessary for functional invasive measurement)
3.9.1 | Cardiac Output (CO) Continuous measurement of the volume of blood pumped by the heart measured in liters per minute (0-12 L/min)
3.9.2 | Cardiac Index (CI) Cardiac output relative to body surface area (BSA) (0-12 L/min/m2)
3.9.3 | Diastolic Pressure (DIA) Diastolic blood pressure (50-110 mmHg)
3.9.4 | Mean Arterial Pressure (MAP) Averaged systemic blood pressure over one cardiac cycle (50-130 mmHg)
3.9.5 | Pulse Rate (PR) Number of arterial blood pressure pulses per minute (40-130 bpm)
3.9.6 | Stroke Volume (SV) Volume of blood pumped with each heart beat (0-160 mL/b)
3.9.7 | Stroke Volume Index (SVI) Stroke volume relative to body surface area (BSA) (0-8 mL/b/m2)
3.9.8 | Systemic Vascular Resistance (SVR) The resistance that the left ventricle must overcome to eject stroke volume with each beat (500-1500 dyne-sec/cm5)
3.9.9 | Systemic Vascular Resistance Index (SVRI) SVR relative to body surface area (500-3000 dyne-sec-m2/cm5)
3.9.10 | Stroke Volume Variation (SVV) The percent difference between SVmin, max and mean (0-50%)
3.9.11 | Systolic Pressure (SYS) Systolic blood pressure (80-160 mmHg)
3.9.12 | Oxygen delivery and consumption
3.10 | Both the invasive and non-invasive options must provide advanced sofware for physiologic interpretation and diagnostics
3.10.1 | Input must be available for further calculations including:
3.10.1.1 | Gender
3.10.1.2 | Age 2-120
3.10.1.3 | Height in cm (30-250)
3.10.1.4 | Weight (1-400 kg)
3.10.1.5 | BSA (0.8-5 m2)
3.10.1.6 | ID (12 digits)
3.10.2 | Graphic and picture display to represent physiology in time
3.10.3 | Further parameter calculation:
3.10.3.1 | Cardiac power index (CPI)
3.10.3.2 | Cardiac power output (CPO)
3.10.3.3 | Oxygen delivery (DO2)
3.10.3.4 | Oxygen delivery index (DO2I)
3.10.3.5 | Oxygen consumption (VO2)
3.10.3.6 | Oxygen consumption index (VO2I)
3.11 | User changeable Alarms
3.11.1 | CO (1-20 L/min)
3.11.2 | CI (0-20 L/min/m2)
3.11.3 | SV (0-30 mL/b )
3.11.4 | SVI (0-200 mL/b/m2)
3.11.5 | SVV (0-99%)
3.11.6 | SVR (0-5000 dyne-sec/cm5)
3.11.7 | SVRI (0-9900 dyne-sec-m2/cm5)
3.11.8 | MAP (20-300 mmHg)
3.11.9 | SYS (20-300 mmHg)
3.11.10 | DIA (20-300 mmHg)
3.11.11 | PR (0-200 Bpm)
3.12 | Onboard review and storage: : must have
3.12.1 | Trends
3.12.1.1. | Selectable 15, 20, 30 interval collection
3.12.1.1. | retained at least 48 hours
3.12.2 | Data download
3.12.2.1 | The user to export monitored patient data to a USB device in Windows Excel XML 2003 format
3.12.2.2 | Case report in Pdf format is desireable
3.13 | User interface and controls: Must have
3.13.1 | Softkeys to drive advanced features
3.13.2 | Low profile keyboard sealed to the edge or touch screen for maximum infection control
3.13.3 | Touch screen preferable
3.13.4 | Ports: USB, Ethernet, VGA must be available
3.13.5 | An on-screen troubleshooting facility must be available
3.14 | Power
3.14.1 | The unit offered must be supplied with an internal rechargeable battery.
3.14.2 | The rechargeable battery must be automatically charged when the unit is connected to a live A.C. supply
3.14.3 | In the event of A.C. mains failure the battery must automatically provide continuous operation
3.14.4 | The unit must be provided with visible indication to reflect whether the unit is being operated off the AC power or the rechargeable battery power
3.14.5 | The unit offered must be provided with circuitry, which must ensure that the rechargeable battery will be protected against over-charging and over-discharging
3.14.6 | Units that are quoted on and employ external battery chargers and power supply will not be considered.
3.14.7 | The rechargeable battery must be of a reasonable capacity, such that with battery power the unit must be able to operate continuously for a minimum of two (2)hours
3.15 | The unit must be handed over in full operating order
3.16 | A starter-pack of consumables must be supplied with the unit.
4 | STANDARD ACCESSORIES
4.1 | Arterial line kit
4.2 | Pulmonary artery catheter
4.3 | Thermodilution catheter
4.4 | Rechargeable battery
4.5 | A stand/trolley with basket/drawer for accessories
5 | Conformity Compliance (Please attach certificate)
5.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
5.2 | OEM must comply and be certified on ISO 13485 quality standards, attach proof of compliance
5.2.1 | Please provide unique ref number of the ISO 13485 ceriticate:
5.3 | Model on offer must comply with EU Medical Device Regulation (EU) 2017/745 or equvalent from SAHPRA recognized jurisdictions . Attach a copy of certification
5.3.1 | Please provide unique ref number of the ceriticate:
5.4 | Hazardous Substance Act: If this type of equipment / apparatus appears on the schedule of Hazardous substances issued by the Directorate: Health Technology of the Department of Health, a license in terms of the Act on Hazardous Substances (Act 15/1973) must be submitted with the bid / quotation documents. The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.5 | All electrical/electronic medical equipment must be licensed by Radiation Control, where applicable a copy of the license must be submitted The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.5.1 | Radiation Control License number:
5.6 | The unit must operate from 220V AC/50Hz mains supply. The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade (rubber to with 2 screws) type and it must be a minimum length of three (3) meters. NB the mains cable of the unit tendered for must be SABS colour coded.
5.7 | The equipment quoted for must be protected against electromagnetic interference IEC 60601-1-2
6 | COMPREHENSIVE MAINTENANCE
6.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) must be quoted for, as per the SCC document
6.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
6.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
6.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
6.3 | The bidder to provide the OEM recommended service interval of the unit:
6.4 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
7 | General
7.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
7.2 | The unit on offer must be the latest technology.
7.2.1 | Initial Manufacturing year of Model:
RT2-05-033 | CARDIAC OUTPUT MONITOR | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point
RT2-05-033a | CO2 sensor
RT2-05-033b | A stand/trolley with basket/drawer for accessories
RT2-05-033c | All essential disposables/or reusables must be quoted
RT2-05-033d | Proximal and distal Pressure sensors
RT2-05-033e | Pulmonary artery catheter
RT2-05-033f | Thermodilution catheter

## Suction Unit, Wall-mount, Doubl
RT2-05-034 | Suction unit, wall-mount | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
1 | Description
1.1 | A medical device used for fluid removal from a patient's airway, surgical site or other body areas, that is wall-mounted
2 | Product Offered:
2.1 | Make:
2.2 | Model:
2.3 | Initial Manufacturing year of Model:
3 | Technical Specification
3.1 | Must be a complete suction unit with an accessory rail mounting clip.
3.2 | The unit must have a Heyer connection
3.3 | The unit must be of robust design.
3.4 | Two receiver jars, each of 2 litre capacity, graduated in 100 ml steps
3.5 | Receiver jars must be high impact crack resistant polycarbonate or similar material (No glass)
3.6 | The receiving jar with lid must be fully autoclavable at 1034o C.
3.7 | Receiver jar lid must have overflow protection
3.8 | The unit offered must be able to be used with liners.
3.9 | The suction inlets and connectors must be plastic.
3.10 | Suction power must be controllable by means of a vacuum control
3.11 | The unit must have a durable on/off switch
3.12 | Operating pressure: 0 to 80 kPa
3.13 | The gauge must be marked in kPa and mmHg
3.14 | The gauge must be fully removable for quick maintenance.
3.15 | The pressure hose must be colour coded yellow.
3.16 | Length of the pressure hose: 2 m
3.17 | The suction unit must be easy to service and maintain.
3.18 | The suction capacity must be at least 25 litres per minute.
3.19 | Water trap/safety trap
3.20 | The unit must be handed over in full operating order
3.21 | A starter-pack of consumables must be supplied with the unit.
4 | Standard Accessories and Consumables
4.1 | Silicone tubes (soft) set 8 x 14 mm
4.2 | Tube connectors
4.3 | Probe connector
4.4 | Aspiration probe.
4.5 | Particle filter
4.6 | Suction tube with canula, reusable
4.7 | Vacuum probe
4.8 | Hose vacuum probe
4.9 | Suction Regulator
4.10 | Suction Gauge
4.11 | 2 liter canister x2
5 | Conformity Compliance (Please attach certificate)
5.1 | OEM must comply and certified on ISO 9001 quality standards, attach proof of compliance
5.1.1 | Please provide unique ref number of the ISO 9001 ceriticate:
5.2 | OEM must comply and certified on ISO 103485 quality standards, attach proof of compliance
5.2.1 | Please provide unique ref number of the ISO 103485 ceriticate:
5.3 | Model quoted for must be EC or SABS certified. Attach a copy of certification
5.3.1 | Please provide unique ref number of the EC or SABS ceriticate:
6 | COMPREHENSIVE MAINTENANCE
6.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) should be quoted for, as per the SCC document
6.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS, Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
7 | General
7.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
7.2 | The unit on offer must be the latest technology.
7.2.1 | Initial Manufacturing year of Model:
RT2-05-0034 | Suction unit, wall-mount Accessories and Consumables | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point
RT2-05-0034a | Silicone tubes (soft) set 8 x 14 mm
RT2-05-034b | Tube connectors
RT2-05-034c | Probe connector
RT2-05-034d | Aspiration probe.
RT2-05-034e | Particle filter
RT2-05-034f | Suction tube with canula, reusable
RT2-05-034g | Vacuum probe
RT2-05-034h | Hose vacuum probe
RT2-05-034i | Suction Regulator
RT2-05-034j | Suction Gauge
RT2-05-034k | 3 liter canister
RT2-05-034l | Mobile stand (with double canister slots)

## Suction Unit, Mobile
RT2-05-035 | Suction Unit, Mobile with Battery | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
1 | Description
1.1 | A medical device used for fluid removal from a patient's airway, surgical site or other body areas, that is mobile with a backup battery
2 | Product Offered:
2.1 | Make:
2.2 | Model:
2.3 | Initial Manufacturing year of Model:
3 | Technical Specification
3.1 | Unit must be mounted on a mobile stand with at least 4 x castor wheels.
3.2 | Must contain a rechargable battery back-up with built-in charger
3.3 | Two of the castors must be lockable
3.4 | Operating pressure: 0 to -60 kPa
3.5 | The gauge must be marked in kPa and mmHg
3.6 | The gauge must be fully removable for maintenance.
3.7 | The suction capacity must be at least 25 litres per minute.
3.8 | On-off switch must be illuminated and have a splash-proof cover
3.9 | Must have suction control for accurate regulation of the vacuum
3.10 | Two receiver jars, each of 2 litre capacity, graduated in 100 ml steps
3.11 | Receiver jars must be high impact crack resistant polycarbonate or similar material (No glass)
3.12 | Receiver jars and lids must be removable and autoclavable at 134o C
3.13 | Receiver jar lid must have overflow protection
3.14 | Receiver jar must operate with and without liners.
3.15 | Vacuum pump regulator must be fitted with a hydrophobic bacterial filter to prevent aspirated fluid from entering the suction pump.
3.16 | Working noise less than 50 dB
3.17 | The suction outlets must be stainless steel.
3.18 | The pump must be oil free and maintenance free
3.19 | The mobile stand must have a handle for pushing
3.22 | Battery status indicator
3.23 | Battery capacity: at least 1 hour operation
3.24 | Battery must be sealed, maintenance free
3.25 | The suction unit must be easy to service and maintain.
3.26 | The unit must be handed over in full operating order
3.27 | A starter-pack of consumables must be supplied with the unit.
4 | Standard Accessories and Consumables
4.1 | Silicone tubes (soft) set 8 x 14 mm
4.2 | Tube connectors
4.3 | Probe connector
4.4 | Aspiration probe.
4.5 | Particle filter
4.6 | Suction tube with canula, reusable
4.7 | Vacuum probe
4.8 | Hose vacuum probe
4.9 | Suction Regulator
4.10 | Suction Gauge
4.11 | Rechargeable Battery
5 | Conformity Compliance (Please attach certificate)
5.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
5.2 | OEM must comply and be certified on ISO 13485 quality standards, attach proof of compliance
5.2.1 | Please provide unique ref number of the ISO 13485 ceriticate:
5.3 | Model on offer must comply with EU Medical Device Regulation (EU) 2017/745 or equvalent from SAHPRA recognized jurisdictions . Attach a copy of certification
5.3.1 | Please provide unique ref number of the ceriticate:
5.4 | Hazardous Substance Act: If this type of equipment / apparatus appears on the schedule of Hazardous substances issued by the Directorate: Health Technology of the Department of Health, a license in terms of the Act on Hazardous Substances (Act 15/1973) must be submitted with the bid / quotation documents. The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.5 | All electrical/electronic medical equipment must be licensed by Radiation Control, where applicable a copy of the license must be submitted The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
5.5.1 | Radiation Control License number:
5.6 | The unit must operate from 220V AC/50Hz mains supply. The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade (rubber to with 2 screws) type and it must be a minimum length of three (3) meters. NB the mains cable of the unit tendered for must be SABS colour coded.
5.7 | The equipment quoted for must be protected against electromagnetic interference IEC 60601-1-2
6 | COMPREHENSIVE MAINTENANCE
6.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) should be quoted for, as per the SCC document
6.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS, Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
6.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
6.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
6.3 | The bidder to provide the OEM recommended service interval of the unit:
6.4 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
7 | General
7.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
7.2 | The unit on offer must be the latest technology.
7.2.1 | Initial Manufacturing year of Model:
RT2-05-035 | Mobile Suction unit, Accessories and Consumables | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point
RT2-05-035a | Silicone tubes (soft) set 8 x 14 mm
RT2-05-035b | Tube connectors
RT2-05-035c | Probe connector
RT2-05-035d | Aspiration probe.
RT2-05-035e | Particle filter
RT2-05-035f | Suction tube with canula, reusable
RT2-05-035g | Suction Regulator
RT2-05-035h | Suction Gauge
RT2-05-035i | Rechargeable Battery