## Category-6 Gen Surgery
GENERAL SURGERY EQUIPMENT
# | Item number | Item Description
1 | RT2-06-001 | Doppler, Vascular
2 | RT2-06-002 | Electrosurgical Unit, Basic
3 | RT2-06-003 | Electrosurgery unit, general surgery (350W)
4 | RT2-06-004 | Electrosurgical Unit, Advanced
5 | RT2-06-005 | Electrosurgical Unit, Argon

## Doppler, Vascular
RT2-06-001 | DOPPLER, VASCULAR | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
1 | DESCRIPTION
1.1 | A handheld, battery operated vascular doppler for detection of bloodstream status from smaller and superficial blood vessels to large and deeper blood vessels.
1a | Product Offered:
1a.1 | Make:
1a.2 | Model:
1a.3 | Initial Manufacturing year of Model:
2 | TECHNICAL SPECIFICATION
2.1 | The unit offered should detect the blood flow average velocity.
2.2 | The unit offered should have audio and visual display(LCD or LED)
2.3 | The following must be displayed:
2.3.1 | Heart rate
2.3.2 | Heartbeat synchronous indicator
2.3.3 | Low battery indicator
2.4 | Accuracy + 3 bpm
2.5 | Power source: Rechargerble battery and main supply of 220 volts/ 50Hz with adaptor.
2.6 | The unit offered should automatically shut off when not in use
2.7 | The unit offered should have interchangeable probes of various frequencies.
2.8 | The unit offered must be in a protective case
2.9 | Vascular Probe
2.9.1 | The probe offered must be waterproof
2.9.2 | The probe of the unit offered should fully operate at 4MHz to 8MHz for different vascular extremities.
3 | STANDARD ACCESSORIES
3.1 | Vascular probe
3.2 | Rechargeble Battery
3.3 | Power adapter
4 | Conformity Compliance (Please attach certificate)
4.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
4.2 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
4.2.1 | Please provide unique ref number of the ISO 13485 ceriticate:
4.4 | Model quoted for must be EC certified. Attach a copy of certification
4.4.1 | Please provide unique ref number of the EC ceriticate:
5 | COMPREHENSIVE MAINTENANCE
5.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) should be quoted for, as per the SCC document
5.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
6 | General
6.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
6.2 | The unit on offer must be the latest technology.
6.2.1 | Initial Manufacturing year of Model:
| DOPPLER, VASCULAR | Complies Yes/No
RT2-06-001a | Vascular probe
RT2-06-001b | Rechargeble Battery
RT2-06-001c | Power adapter

## Electrosurgery basic
RT2-06-002 | Electrosurgical unit basic | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
1 | DESCRIPTION
1.1 | An Electrosurgical Unit (ESU) is a medical device used to cut, coagulate and remove tissue during surgical procedures.
1a | Product Offered:
1a.1 | Make:
1a.2 | Model:
1a.3 | Initial Manufacturing year of Model:
1a.4 | Radiation Control License number:
2 | TECHNICAL SPECIFICATION
2.1 | The unit must be basic electrosurgical unit for cutting and coagulation.
2.2 | The unit must be Microprocessor controlled generator with mono-polar and bipolar output
2.3 | The unit must employ solid state technology for the generation of the various outputs required
2.4 | The unit must do a self-test at switch-on
2.5 | The unit must be able to perform the following functions in monopolar mode:
2.5.1 | Pure cut
2.5.2 | Coagulation
2.5.3 | Blend
2.5.4 | Spray coagulation
2.6 | The output power selected must be numerically displayed
2.8 | Output frequency: approximately 400 kHz
2.9 | Maximum output power, mono-polar: up to 100 W (cutting)
2.10 | Bipolar output power: 50 W (coagulation)
2.11 | Double function foot switch (mono and bipolar)
2.12 | The unit must automatically switch between the foot and hand switch when required.
2.13 | The unit must comprise of both visual and audible activation indicators
2.14 | Audible and visual Warning/Alarm must be given when:
2.14.1 | Cut function is activated
2.14.2 | Coagulation function is activated
2.14.3 | Neutral electrode (patient plate) is disconnected from the unit
2.14.4 | The unit has been connected with a faulty neutral electrode or cable
2.15 | When an alarm associated within the neutral electrode occurs, the power output of the unit must be reduced to zero.
2.16 | The unit must monitor patient contact with the neutral electrode and activate safety measure should there be insufficient contact with the elecrode.
2.17 | The unit must have a safety measure should the patient make insufficient contact with the neutral electrode.
2.18 | The unit musty be powered from 220 V, 50 Hz supply.
2.19 | The unit must be mounted on a sturdy trolley with 4 antistatic swivel castors, 2 with brakes
2.20 | Trolley must have a drawer and storage for foot pedal and cable
2.21 | Mains cable must be at least 3 (three) metres long with SABS approved 15A three prong plug and colour code.
2.22 | The unit must be handed over in full operating order
2.23 | A starter-pack of consumables must be supplied with the unit.
3 | Standard Accessories and Consumables
3.1 | Foot switch, double, with cable (1x)
3.2 | Patient electrode, adult, stainless steel or conductive rubber, reusable
3.3 | Patient electrode cable
3.4 | Patient electrode, split, suitable for monitoring patient contact, reusable, autoclavable
3.5 | Cable for split patient electrode
3.6 | Electrode handle, mono-polar, reusable, foot switch controlled, (2x) with connecting cable
3.7 | Electrode handle, mono-polar, reusable, finger switch controlled, (2x) with connecting cable
3.8 | A set of different mono-polar reusable electrodes (needle, blade, ball and loop) (1x)
3.9 | Bipolar forceps, reusable, foot switch controlled, with connecting cable (1x)
3.10 | Dedicated trolley
4 | Conformity Compliance (Please attach certificate)
4.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
4.2 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
4.2.1 | Please provide unique ref number of the ISO 13485 ceriticate:
4.3 | Model on offer must comply with EU Medical Device Regulation (EU) 2017/745 or equvalent from SAHPRA recognized jurisdictions . Attach a copy of certification
4.3.1 | Please provide unique ref number of the EC ceriticate:
4.4 | Hazardous Substance Act: If this type of equipment / apparatus appears on the schedule of Hazardous substances issued by the Directorate: Health Technology of the Department of Health, a license in terms of the Act on Hazardous Substances (Act 15/1973) must be submitted with the bid / quotation documents. The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
4.5 | All electrical/electronic medical equipment must be licensed by Radiation Control, where applicable a copy of the license must be submitted The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
4.5.1 | Radiation Control License number:
4.6 | The unit shall operate from 220V AC/50Hz mains supply. The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade (rubber to with 2 screws) type and it must be a minimum length of three (3) meters. NB the mains cable of the unit tendered for must be SABS colour coded.
4.7 | The equipment quoted for must be protected against electromagnetic interference IEC 60601-1-2
4.8 | The unit must comply with electrical safety and performance standard such as IEC 60601-2-2, attach certification
4.8.1 | Please provide unique ref number of the IEC ceriticate:
5 | COMPREHENSIVE MAINTENANCE
5.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) should be quoted for, as per the SCC document
5.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
5.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
5.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
5.3 | The bidder to provide the OEM recommended service interval of the unit:
5.4 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
6 | General
6.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
6.2 | The unit on offer must be the latest technology.
6.2.1 | Initial Manufacturing year of Model:
| Electrosurgical unit basic - Accessories and Consumables | Complies Yes/No
RT2-06-002a | Foot switch, double, with cable
RT2-06-002b | Patient electrode, adult, stainless steel or conductive rubber, reusable
RT2-06-002c | Patient electrode cable
RT2-06-002d | Patient electrode, split, suitable for monitoring patient contact, reusable, autoclavable
RT2-06-002e | Cable for split patient electrode
RT2-06-002f | Electrode handle, mono-polar, reusable, foot switch controlled, (2x) with connecting cable
RT2-06-002g | Electrode handle, mono-polar, reusable, finger switch controlled, (2x) with connecting cable
RT2-06-002h | A set of different mono-polar reusable electrodes (needle, blade, ball and loop) (1x)
RT2-06-002i | Bipolar forceps, reusable, foot switch controlled, with connecting cable (1x)
RT2-06-002j | Disposable single neutral electrode pad (Pack of 5)
RT2-06-002k | Disposable split neutral electrode pad (Pack of 5)

## Electrosurgery general
RT2-06-003 | Electrosurgery unit, general surgery (350W) | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
1 | DESCRIPTION
1.1 | An Electrosurgical Unit (ESU) is a medical device used to cut, coagulate and remove tissue during surgical procedures.
1a | Product Offered:
1a.1 | Make:
1a.2 | Model:
1a.3 | Initial Manufacturing year of Model:
1a.4 | Radiation Control License number:
2 | TECHNICAL SPECIFICATION
2.1 | The unit must employ solid state technology for the generation of the various outputs required for general surgery.
2.2 | The unit must do a self-test at switch-on
2.3 | The unit must have the following modes of operations:
2.3.1 | Monopolar mode
2.3.2 | Bipolar mode
2.4 | The unit must be able to perform the following functions in monopolar mode:
2.4.1 | Pure cut
2.4.2 | Coagulation
2.4.3 | Blend
2.4.4 | Spray coagulation
2.5 | The pure cut function must have variable power from 0 to 350 W when delivered into a 500 Ohm load
2.6 | The output power selected must be numerically displayed
2.7 | Power output must remain constant over a wide range of tissue impedance.
2.8 | The electrodes must be electrically isolated from earthing.
2.9 | The coagulation function must have the following output:
2.9.1 | Variable power from 0 to 120 W delivered into a 500 Ohm load in the coagulation pin point mode
2.9.2 | Variable power from 0 to 100 W delivered into a 500 Ohm load in the coagulation spray mode
2.10 | The blend function must have a variable control of the degree of haemostasis.
2.11 | The Bipolar mode must deliver variable power from 0 to 50 W into a 125 Ohm load in both the Bipolar Cut and Bipolar Coagulation modes.
2.12 | Frequency must be approximately 500 kHz.
2.13 | The unit must comprise of both visual and audible activation indicators
2.14 | Audible and visual warning/Alarm must be given when:
2.14.1 | Cut function is activated
2.14.2 | Coagulation function is activated
2.14.3 | Neutral electrode (patient plate) disconnected from the unit
2.14.4 | The unit has been connected with a faulty neutral electrode or cable
2.15 | When an alarm associated with the neutral electrode occurs, the power output of the unit must be reduced to zero.
2.16 | The unit must monitor patient contact with the neutral electrode
2.17 | The unit must have a safety measure should the patient make insufficient contact with the neutral electrode.
2.18 | The output power to the active electrode must be activated by either foot pedal control or a finger-operated switch on the handpiece of the active electrode
2.19 | There must be no possibility of activating both hand and foot operated accessories simultaneously.
2.20 | The foot switch must have a barrier between the two pedals
2.21 | The foot pedal cable must be detachable from both the electrosurgical unit and the foot switch
2.22 | The foot pedal must be splash proof.
2.23 | The unit’s power ON switch must be enclosed to prevent accidental switch On or Off;
2.24 | The unit must be powered from 220 V, 50 Hz supply.
2.25 | Mains cable must be at least 3 (three) metres long with SABS approved 15A three prong plug and colour code.
2.26 | The unit must be supplied with a trolley with brakes on antistatic castors with a drawer to store footswitch and accessories.
2.27 | The unit must be handed over in full operating order
2.28 | A starter-pack of consumables must be supplied with the unit.
3 | Standard Accessories and Consumables
3.1 | Foot switch, double, with cable
3.2 | Patient electrode, adult, stainless steel or conductive rubber, reusable
3.3 | Patient electrode cable
3.4 | Patient electrode, split, suitable for monitoring patient contact, reusable, autoclavable
3.5 | Cable for split patient electrode
3.6 | Electrode handle, mono-polar, reusable, foot switch controlled, (2x) with connecting cable
3.7 | Electrode handle, mono-polar, reusable, finger switch controlled, (2x) with connecting cable
3.8 | Set different mono-polar reusable electrodes (needle, blade, ball and loop) (1x)
3.9 | Bipolar forceps, reusable, foot switch controlled, with connecting cable (1x)
3.10 | Trolley
4 | Conformity Compliance (Please attach certificate)
4.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
4.2 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
4.2.1 | Please provide unique ref number of the ISO 13485 ceriticate:
4.3 | Model on offer must comply with EU Medical Device Regulation (EU) 2017/745 or equvalent from SAHPRA recognized jurisdictions . Attach a copy of certification
4.3.1 | Please provide unique ref number of the EC ceriticate:
4.4 | Hazardous Substance Act: If this type of equipment / apparatus appears on the schedule of Hazardous substances issued by the Directorate: Health Technology of the Department of Health, a license in terms of the Act on Hazardous Substances (Act 15/1973) must be submitted with the bid / quotation documents. The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
4.5 | All electrical/electronic medical equipment must be licensed by Radiation Control, where applicable a copy of the license must be submitted The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
4.5.1 | Radiation Control License number:
4.6 | The unit shall operate from 220V AC/50Hz mains supply. The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade (rubber to with 2 screws) type and it must be a minimum length of three (3) meters. NB the mains cable of the unit tendered for must be SABS colour coded.
4.7 | The equipment quoted for must be protected against electromagnetic interference IEC 60601-1-2
4.8 | The unit must comply with electrical safety and performance standard such as IEC 60601-2-2, attach certification
4.8.1 | Please provide unique ref number of the IEC ceriticate:
5 | COMPREHENSIVE MAINTENANCE
5.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) should be quoted for, as per the SCC document
5.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
5.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
5.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
5.3 | The bidder to provide the OEM recommended service interval of the unit:
5.4 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
6 | General
6.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
6.2 | The unit on offer must be the latest technology.
6.2.1 | Initial Manufacturing year of Model:
RT2-06-003 | Electrosurgery unit, general surgery (350 W)- Accessories and Consumables | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
RT2-06-003a | Foot switch, double, with cable
RT2-06-003b | Patient electrode, adult, stainless steel or conductive rubber, reusable
RT2-06-003c | Patient electrode cable
RT2-06-003d | Patient electrode, split, suitable for monitoring patient contact, reusable, autoclavable
RT2-06-003e | Cable for split patient electrode
RT2-06-003f | Electrode handle, mono-polar, reusable, foot switch controlled, (2x) with connecting cable
RT2-06-003g | Electrode handle, mono-polar, reusable, finger switch controlled, (2x) with connecting cable
RT2-06-003h | Set different mono-polar reusable electrodes (needle, blade, ball and loop) (1x)
RT2-06-003i | Bipolar forceps, reusable, foot switch controlled, with connecting cable (1x)
RT2-06-003j | Disposable single neutral electrode pad (Pack of 5)
RT2-06-003k | Disposable split neutral electrode pad (Pack of 5)
RT2-06-003l | Vessel sealing
RT2-06-003m | Memory for at least 10 programs with their waveforms and power levels
RT2-06-003n | Smoke evacuation

## Electrosurgery advanced
RT2-06-004 | Electrosurgery unit, advanced | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
1 | DESCRIPTION
1.1 | An Electrosurgical Unit (ESU) is a medical device used to cut, coagulate and remove tissue during surgical procedures.
1a | Product Offered:
1a.1 | Make:
1a.2 | Model:
1a.3 | Initial Manufacturing year of Model:
1a.4 | Radiation Control License number:
2 | TECHNICAL SPECIFICATION
2.1 | Must be programmable for the full range of the following electrosurgical requirements:
2.1.1 | General surgery
2.1.2 | Cardiology
2.1.3 | ENT
2.1.4 | Gastroenterology
2.1.5 | Gynaecology
2.1.6 | Laparoscopy
2.1.7 | Neurosurgery
2.1.8 | Thoracic surgery
2.1.9 | Urology (TURP, TUVP)
2.2 | User must be able to enter and save preference settings
2.3 | Power output must be at least 400 W with output isolated from ground
2.4 | The unit must offer automatic power adjustment for variation in tissue impedance
2.5 | The unit must offer automatic recognition of the electrode used and adjustment of the output parameters accordingly
2.6 | The unit must employ solid state technology for the generation of the various outputs required
2.7 | The unit must do a self-test at switch-on
2.8 | The unit must have the following modes of operations:
2.8.1 | Monopolar mode
2.8.2 | Bipolar mode
2.8.3 | Suitable for two operators with active outputs simultaneously
2.8.4 | Two footswitches active simultaneously
2.8.5 | One footswitch and one handswitch active simultaneously
2.8.6 | Two handswitches active simultaneously
2.8.7 | Monopolar and bipolar active simultaneously
2.9 | The unit must be able to perform the following functions in monopolar mode:
2.9.1 | Pure cut
2.9.2 | Coagulation
2.9.3 | Blend
2.9.4 | Spray coagulation
2.10 | The pure cut function must have variable power from 0 up to 400 W with a 500 Ohm load
2.11 | The output power selected must be numerically displayed
2.12 | The coagulation function must have the following output:
2.12.1 | Variable power from 0 to 120 W delivered into a 500 Ohm load in the coagulation pin point mode
2.12.2 | Variable power from 0 to 100 W delivered into a 500 Ohm load in the coagulation spray mode
2.13 | The blend function must have a variable control of the degree of haemostasis.
2.14 | The Bipolar mode must deliver variable power from 0 to 50 W into a 125 Ohm load in both the Bipolar Cut and Bipolar Coagulation modes.
2.15 | LigaSure vessel sealing technology upgradable
2.16 | Frequency must be approximately 500 kHz.
2.17 | Memory for at least 10 programs with their waveforms and power levels
2.18 | The unit must comprise of both visual and audible activation indicators
2.19 | Audible and visual warning/Alarm must be given when:
2.19.1 | Cut function is activated
2.19.2 | Coagulation function is activated
2.19.3 | Neutral electrode (patient plate) disconnected from the unit
2.19.4 | The unit has been connected with a faulty neutral electrode or cable
2.20 | When an alarm associated with the neutral electrode occurs, the power output of the unit must be reduced to zero.
2.21 | The unit must monitor patient contact with the neutral electrode
2.22 | The unit must have a safety measure should the patient make insufficient contact with the neutral electrode.
2.23 | The output power to the active electrode must be activated by either foot pedal control or a finger-operated switch on the handpiece of the active electrode
2.24 | The foot switch must have a barrier between the two pedals
2.25 | The foot pedal cable must be detachable from both the electrosurgical unit and the foot switch
2.26 | The foot pedal must be splash proof.
2.27 | The unit’s power ON switch must be enclosed to prevent accidental switch On or Off
2.28 | The unit must be powered from 220 V, 50 Hz supply.
2.29 | Mains cable must be at least 3 (three) metres long with SABS approved 15A three prong plug and colour code.
2.30 | The unit must be supplied with a trolley on antistatic castors, with brakes and a drawer to store footswitch and accessories.
2.31 | The unit must be handed over in full operating order
2.32 | A starter-pack of consumables must be supplied with the unit.
3 | Standard Accessories and Consumables
3.1 | Foot switch, double, with cable (1x)
3.2 | Patient electrode, adult, stainless steel or conductive rubber, reusable
3.3 | Patient electrode cable
3.4 | Patient electrode, split, suitable for monitoring patient contact, reusable, sterlizable
3.5 | Cable for split patient electrode
3.6 | Electrode handles with cables, reusable, autoclavable (2x)
3.7 | Blade/knife electrodes, straight, reusable (2x)
3.8 | Needle electrodes (thick) (2x)
3.9 | Needle electrode (thin) (1x)
3.10 | Ball electrodes, straight, 4 mm diameter ((2x)
3.11 | Ball electrode, straight, 2,5 mm diameter (1x)
3.12 | Bipolar electrodes reusable, autoclavable (2x)
4 | Conformity Compliance (Please attach certificate)
4.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
4.2 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
4.2.1 | Please provide unique ref number of the ISO 13485 ceriticate:
4.3 | Model on offer must comply with EU Medical Device Regulation (EU) 2017/745 or equvalent from SAHPRA recognized jurisdictions . Attach a copy of certification
4.3.1 | Please provide unique ref number of the EC ceriticate:
4.4 | Hazardous Substance Act: If this type of equipment / apparatus appears on the schedule of Hazardous substances issued by the Directorate: Health Technology of the Department of Health, a license in terms of the Act on Hazardous Substances (Act 15/1973) must be submitted with the bid / quotation documents. The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
4.5 | All electrical/electronic medical equipment must be licensed by Radiation Control, where applicable a copy of the license must be submitted The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
4.5.1 | Radiation Control License number:
4.6 | The unit shall operate from 220V AC/50Hz mains supply. The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade (rubber to with 2 screws) type and it must be a minimum length of three (3) meters. NB the mains cable of the unit tendered for must be SABS colour coded.
4.7 | The equipment quoted for must be protected against electromagnetic interference IEC 60601-1-2
4.8 | The unit must comply with electrical safety and performance standard such as IEC 60601-2-2, attach certification
4.8.1 | Please provide unique ref number of the IEC ceriticate:
5 | COMPREHENSIVE MAINTENANCE
5.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) should be quoted for, as per the SCC document
5.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
5.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
5.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
5.3 | The bidder to provide the OEM recommended service interval of the unit:
5.4 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
6 | General
6.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
6.2 | The unit on offer must be the latest technology.
6.2.1 | Initial Manufacturing year of Model:
| Electrosurgery unit, Advanced- Accessories and Consumables | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
RT2-06-004a | Foot switch, double, with cable
RT2-06-004b | Patient electrode, adult, stainless steel or conductive rubber, reusable
RT2-06-004c | Patient electrode cable
RT2-06-004d | Patient electrode, split, suitable for monitoring patient contact, reusable, autoclavable
RT2-06-004e | Cable for split patient electrode
RT2-06-004f | Electrode handle, mono-polar, reusable, foot switch controlled, (2x) with connecting cable
RT2-06-004g | Electrode handle, mono-polar, reusable, finger switch controlled, (2x) with connecting cable
RT2-06-004h | Set different mono-polar reusable electrodes (needle, blade, ball and loop) (1x)
RT2-06-004i | Bipolar forceps, reusable, foot switch controlled, with connecting cable (1x)
RT2-06-004j | Disposable single neutral electrode pad (Pack of 5)
RT2-06-004k | Disposable split neutral electrode pad (Pack of 5)
RT2-06-004l | Vessel sealing
RT2-06-004m | Memory for at least 10 programs with their waveforms and power levels
RT2-06-004n | Smoke evacuation
RT2-06-004o | LigaSure vessel sealing technology upgrade

## Electrosurgery with Argon
RT2-06-005 | Electrosurgery unit, with Argon Module | Complies Yes/No | Provide your answers in this Column. You are advised to be straight to the point.
1 | DESCRIPTION
1.1
1a | Product Offered:
1a.1 | Make:
1a.2 | Model:
1a.3 | Initial Manufacturing year of Model:
1a.4 | Radiation Control License number:
2 | TECHNICAL SPECIFICATION
2.1 | Must be programmable for the full range of electrosurgical requirements:
2.1.1 | General surgery
2.1.2 | Cardiology
2.1.3 | ENT
2.1.4 | Gastroenterology
2.1.5 | Gynaecology
2.1.6 | Laparoscopy
2.1.7 | Neurosurgery
2.1.8 | Thoracic surgery
2.1.9 | Urology (TURP, TUVP)
2.2 | User must be able to enter and save preference settings
2.3 | Power output: at least 400 W with output isolated from ground
2.4 | Automatic power adjustment for variation in tissue impedance
2.5 | Automatic recognition of the electrode used and adjustment of the output parameters accordingly
2.6 | The unit must employ solid state technology for the generation of the various outputs required
2.7 | The unit must do a self-test at switch-on
2.8 | The unit must have the following modes of operations:
2.8.1 | Monopolar mode
2.8.2 | Bipolar mode
2.8.3 | Suitable for two operators with active outputs simultaneously
2.8.4 | Two footswitches active simultaneously
2.8.5 | One footswitch and one handswitch active simultaneously
2.8.6 | Two handswitches active simultaneously
2.8.7 | Monopolar and bipolar active simultaneously
2.9 | The unit must be able to perform the following functions in monopolar mode:
2.9.1 | Pure cut
2.9.2 | Coagulation
2.9.3 | Blend
2.9.4 | Spray coagulation
2.9.5 | Argon cutting & coagulation
2.9.6 | Argon pulse coagulation
2.10 | The pure cut function must have variable power from up to 400 W with a 500 Ohm load
2.11 | The output power selected must be digitally displayed
2.12 | The coagulation function must have the following output:
2.12.1 | Variable power from 0 to 120 W delivered into a 500 Ohm load in the coagulation pin point mode
2.12.2 | Variable power from 0 to 100 W delivered into a 500 Ohm load in the coagulation spray mode
2.12.3 | The blend function must have a variable control of the degree of haemostasis.
2.12.4 | The Bipolar mode must deliver variable power from 0 to 50 W into a 125 Ohm load in both the Bipolar Cut and Bipolar Coagulation modes.
2.13 | Frequency must be approximately 500 kHz.
2.14 | Memory for at least 10 programs with their waveforms and power levels
2.15 | Visual and audible activation indicators
2.16 | Audible and visual warnings must be given when:
2.16.1 | Cut function is activated
2.16.2 | Coagulation function is activated
2.16.3 | Neutral electrode (patient plate) disconnected from the unit
2.16.4 | The unit has been connected with a faulty neutral electrode or cable
2.17 | When an alarm associated within the neutral electrode occurs, the power output of the unit must be reduced to zero.
2.18 | The unit must monitor patient contact with the neutral electrode
2.19 | The unit must fail-safe should the patient make insufficient contact with the neutral electrode.
2.20 | The output power to the active electrode must be activated by either foot pedal control or a finger-operated switch on the handpiece of the active electrode
2.21 | There must be no possibility of activating both hand and foot operated accessories simultaneously.
2.22 | The foot switch must have a barrier between the two pedals
2.23 | The foot pedal cable must be detachable from both the electrosurgical unit and the foot switch
2.24 | The foot pedal must be splash proof.
2.25 | The unit’s power On switch must be enclosed to prevent accidental switch On or Off
2.26 | Argon
2.26.1 | The unit must utilize pure argon with inlet gas pressure 3 - 5 bar
2.26.2 | Argon cylinder must have pressure reducer
2.26.3 | Gas Flow range between 0.1 - 10 litre per minutes
2.26.4 | Gas filtration must be possible. Please state how it is achieved and filtration range
2.26.5 | The unit must have overpressure monitoring
2.27 | The equipment must function off 220 V, 50 Hz single phase mains supply
2.28 | Mains cable must be at least 3 (three) metres long with SABS approved 15A three prong plug and colour code.
2.29 | The unit must be supplied with a trolley on antistatic castors, with brakes and a drawer to store footswitch and accessories.
2.30 | The unit must be handed over in full operating order
2.31 | A starter-pack of consumables must be supplied with the unit.
3 | Standard Accessories and Consumables
3.1 | Foot switch, double, with cable (1x)
3.2 | Patient electrode, adult, stainless steel or conductive rubber, reusable
3.3 | Patient electrode cable
3.4 | Patient electrode, split, suitable for monitoring patient contact, reusable, autoclavable
3.5 | Cable for split patient electrode
3.6 | Electrode handles with cables, reusable, autoclavable (2x)
3.7 | Blade/knife electrodes, straight, reusable (2x)
3.8 | Needle electrodes (thick) (2x)
3.9 | Needle electrode (thin) (1x)
3.10 | Ball electrodes, straight, 4 mm diameter ((2x)
3.11 | Ball electrode, straight, 2,5 mm diameter (1x)
3.12 | Bipolar electrodes reusable, sterilizable (2x)
3.13 | Flexible argon probe
3.14 | Argon electrode connection
4 | Conformity Compliance (Please attach certificate)
4.1 | The unit must comply with an acceptable international electrical safety standard such as IEC 601-1 for medical equipment, attached certification
4.2 | OEM must comply and certified on ISO 13485 quality standards, attach proof of compliance
4.2.1 | Please provide unique ref number of the ISO 13485 ceriticate:
4.3 | Model on offer must comply with EU Medical Device Regulation (EU) 2017/745 or equvalent from SAHPRA recognized jurisdictions . Attach a copy of certification
4.3.1 | Please provide unique ref number of the EC ceriticate:
4.4 | Hazardous Substance Act: If this type of equipment / apparatus appears on the schedule of Hazardous substances issued by the Directorate: Health Technology of the Department of Health, a license in terms of the Act on Hazardous Substances (Act 15/1973) must be submitted with the bid / quotation documents. The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
4.5 | All electrical/electronic medical equipment must be licensed by Radiation Control, where applicable a copy of the license must be submitted The license must be registered under the bidders name or a letter of joint venture must be submitted by the license holder where the license is not in the name of the bidder. Bidders that neglect to submit a license will not be considered
4.5.1 | Radiation Control License number:
4.6 | The unit shall operate from 220V AC/50Hz mains supply. The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade (rubber to with 2 screws) type and it must be a minimum length of three (3) meters. NB the mains cable of the unit tendered for must be SABS colour coded.
4.7 | The equipment quoted for must be protected against electromagnetic interference IEC 60601-1-2
4.8 | The unit must comply with electrical safety and performance standard such as IEC 60601-2-2, attach certification
4.8.1 | Please provide unique ref number of the IEC ceriticate:
5 | COMPREHENSIVE MAINTENANCE
5.1 | A 24 month warranty inclusive maintenance and service (Year 1 to Year 2) should be quoted for, as per the SCC document
5.1.1 | THIS ENTAILS: A 2-year warranty against poor workmanship and latent defects and parts. This must be all inclusive and include, BUT NOT LIMITED TO, amongst others, ALL PARTS (including Batteries), Labour, Traveling and Accommodation. The 2-year warranty must also include all quality check, quality assurance requirements, Preventative Maintenance/Calibrations, Software updates and upgrades to be included. This 2-year warranty will commence after formal Commissioning and handover of the equipment.
5.2 | A 5-year Comprehensive extended warranty (Year 3 to Year 7) must be quoted for, as per the SCC document
5.2.1 | THIS ENTAILS: A Fully Comprehensive 5-year Preventative Maintenance/Service and Corrective Maintenance/Repair contract covering all equipment, hardware and software. This contract would cover, but not be limited to the following: ALL PARTS (including Batteries), labour, travelling, accommodation, service and repair. The 5-year Maintenance Plan must also include all quality check, quality assurance requirements, calibrations , Software updates and upgrades.This contract will commence after the 2-year warranty period has expired.
5.3 | The bidder to provide the OEM recommended service interval of the unit:
5.4 | Bidder to provider details of technical work (Spare parts & Calibrations) involving the unit bidded for
6 | General
6.1 | Has the product on offer being on safety recall by the Regulatory Authority in the last 5 years? If yes, please provide further details of the recall and how was it addressed
6.2 | The unit on offer must be the latest technology.
6.2.1 | Initial Manufacturing year of Model:
RT2-06-005 | Electrosurgery unit, with Argon Module - Accessories and Consumables | Complies Yes/No
RT2-06-005a | Foot switch, double, with cable
RT2-06-005b | Patient electrode, adult, stainless steel or conductive rubber, reusable
RT2-06-005c | Patient electrode cable
RT2-06-005d | Patient electrode, split, suitable for monitoring patient contact, reusable, sterlizable
RT2-06-005e | Cable for split patient electrode
RT2-06-005f | Electrode handles with cables, reusable, autoclavable
RT2-06-005g | Blade/knife electrodes, straight, reusable
RT2-06-005h | Needle electrodes (thick)
RT2-06-005i | Needle electrode (thin)
RT2-06-005j | Ball electrodes, straight, 4 mm diameter
RT2-06-005k | Ball electrode, straight, 2,5 mm diameter
RT2-06-005l | Bipolar electrodes reusable, autoclavable
RT2-06-005m | Flexible argon probe
RT2-06-005n | Argon electrode connection
RT2-06-005o | Disposable single and split neutral electrode pads
RT2-06-005p | Vessel sealing
RT2-06-005q | Smoke evacuation