## RT274-1-2023 Technical Specific
Line # | Item Number | Category and other requirements | Item Description | SAHPRA Licence Required Yes/No | Type Approval Certificate required Yes/No | Unit of Measure | Compliant with specification (Yes/No) | Indicate the Deviation if applicable
1 | RT274-01-001 | Category 1 | Pre-cut cleaning floss kit for OAE machines & tympanometer probes pack of (Large) at minimum of 30 pieces of thick floss | No | No | Pack of 30 pieces
2 | RT274-01-002 | Category 1 | Pre-cut cleaning floss kit for OAE machines & tympanometers probes pack of (Small and Large) minimum of 50 pieces of floss | No | No | Pack of 100 pieces
3 | RT274-01-003 | Category 1 | Pre-cut cleaning floss kit for OAE machines & tympanometers probes (small) pack of minimum of 30 pieces of thin floss | No | No | Pack of 100 pieces
4 | RT274-01-009 | Category 1 | Hearing aid management database, single license, for measurement and fitting of hearing aids to combine management of hearing aids from a range of suppliers into an integrated system | No | No | Each
5 | RT274-01-010 | Category 1 | Hearing aid management database, multi-license, for measurement and fitting of hearing aids to combine management of hearing aids from a range of suppliers into an integrated system | No | No | Each
6 | RT274-01-065 | Category 1 | Disposable electrodes for electrophysiologic diagnostic testing. Bidder to specify number in pack and price per unit. Supplier to specify equipment compatibility | Yes | No | Each
7 | RT274-01-066 | Category 1 | Disposable electrodes for electrophysiologic screening ABR testing. Bidder to specify number in pack and price per unit. Supplier to specify equipment compatibility | Yes | No | Each
8 | RT274-01-067 | Category 1 | Electrodes, disposable EP/ENG snap, wet gel, Supplier to specify number in pack and price per unit. | Yes | No | Each
9 | RT274-01-019 | Category 1 | Base unit: A base unit to which screening and diagnostic electrophysiology modules and their necessary hardware can be added or connected; must have capability for multiple modules to be added. Must be portable and must come with protective carrycase. Bidder must provide pricing for the maintenance plan as provided for on the pricing schedule. | Yes | No | Each
10 | RT274-01-078 | Category 1 | Diagnostic high frequency DPOAE and TEOAE base unit: A diagnostic unit capable of measuring DPOAEs and TEOAE’s, with at least six configurable protocols testing from 1.5-16 000Hz, must store at least 50 tests and display normative data to assist with interpretation of results, must have software for download to PC based database, must include the unit, cradle, if relevant, OAE probe, probe calibration test cavity, database software, micro USB charging / transfer cable, carry case, ear tip kit and replacement probe tubes and calibration certificate. Base unit: A base unit to which screening modules and their necessary hardware can be added or connected, must have capability for multiple modules to added, must allow for subsequent addition of additional modules. Bidder must provide pricing for the maintenance plan as provided for on the pricing schedule. | Yes | No | Each
11 | RT274-01-081 | Category 1 | Replacement probe tip OAEs compatible with item RT274-01- 41, 42, 54 and previous awarded tenders | Yes | No | Each
12 | RT274-01-082 | Category 1 | Replacement cables OAEs compatible with item RT274-01- 41, 42, 54 and previous awarded tenders | Yes | No | Each
13 | RT274-01-083 | Category 1 | Replacement power supply OAEs compatible with item RT274-01- 25, 36, 41, 42, 54 and previous awarded tenders | Yes | No | Each
14 | RT274-01-030 | Category 1 | Replacement cables for AABR compatible with item RT274-01- 36, 49 and previous awarded tenders | Yes | No | Each
15 | RT274-01-084 | Category 1 | Replacement power supply for AABR compatible with item RT274-01- 36, 49 and previous awarded tenders | Yes | No | Each
16 | RT274-01-032 | Category 1 | Replacement callibration cavity for AABR compatible with item RT274-01- 36, 49 and previous awarded tenders | Yes | No | Each
17 | RT274-01-033 | Category 1 | Replacement lead electrodes for AABR compatible with item RT274-01- 36, 49 and previous awarded tenders | Yes | No | Each
18 | RT274-01-034 | Category 1 | Replacement insert earphones 30 ohms | Yes | No | Each
19 | RT274-01-035 | Category 1 | Replacement insert earphones 50 ohms | Yes | No | Each
20 | RT274-01-085 | Category 1 | Replacement trigger cable for electrophysiology system | Yes | No | Each
21 | RT274-01-107 | Category 1 | Electrophysiology Tip trode electrode | Yes | No | Each
22 | RT274-01-108 | Category 1 | Gold foil electrode (large) for ECochG testing | Yes | No | Each
23 | RT274-01-109 | Category 1 | Gold foil electrode (small) for ECochG testing | Yes | No | Each
24 | RT274-01-110 | Category 1 | Replacement Accessory biologic check test Box for Electrophysiology compatible with current and previous tenders | Yes | No | Each
25 | RT274-01-111 | Category 1 | Portable suction unit for wax removal | Yes | No | Each
26 | RT274-01-112 | Category 1 | Hearing aid repair restoration system. A Unit that Cleans and Repairs Hearing Instruments, includes Drying Chamber Wicks Away Moisture and can Use the Vacuum and Pressure Wands Simultaneously | Yes | No | Each
27 | RT274-01-113 | Category 1 | Ear Wax Removal, Video/ Wireless Ear Cleaner, Ear Wax Removal Kit, Rechargeable battery may be an option. Ear Cleaner with Camera and Light, Ear Cleaning Kit. Carry case must be provided | Yes | No | Each
28 | RT274-01-114 | Category 1 | Wireless handheld video otoscope. A minimum of 12 speculi to be supplied.Rechargeable battery may be an option. Carry case must be provided. | Yes | No | Each
29 | RT274-01-115 | Category 1 | Replacement contralateral reflex probe assembly | Yes | No | Each
30 | RT274-01-116 | Category 1 | Replacement otolight lightbulbs. Must be compatible with otolights awarded on RT274-02-001 on RT274-2023 and RT274 contracts since 2012 | No | No | Each
31 | RT274-01-117 | Category 1 | LED replacement otolight lightbulbs. Must be compatible with otolights awarded on RT274-02-001 on RT274-2023 and RT274 contracts since 2012 | No | No | Each
32 | RT274-01-118 | Category 1 | Retubing kit in a zip durable pouch complete with a minimum of stainless steel tube expander, scalpel, hearing aid screw driver, clear strong instant glue, air blower for earmould, needle nose plier, a tube threader | No | No | Each
33 | RT274-01-119 | Category 1 | Replacement Tube expander | No | No | Each
34 | RT274-01-120 | Category 1 | Replacement Scalpel | Yes | No | Each
35 | RT274-01-121 | Category 1 | Replacement set of hearing aid screw drivers | No | No | Each
36 | RT274-01-122 | Category 1 | Replacement needle nose plier | No | No | Each
37 | RT274-01-123 | Category 1 | Replacement Tube Threader | No | No | Each
38 | RT274-01-124 | Category 1 | Small electric hand drill complete with drill bits suitable for drilling earmoulds, pack of a minumum of 4 different sizes such as 0.5, 1, 1.5, 2 and 2.5mm drill bits, buffer ball, grindstone and sand paper | No | No | Each
39 | RT274-01-125 | Category 1 | Set of drillbits and accessories for a small electric earmould drill consisting of a minimum of 4 different sizes such as 0.5, 1, 1.5, 2 and 2.5mm drill bits, buffer ball, grindstone and sand paper, for ear mould. To be compatible with awarded drill | No | No | Each
40 | RT274-01-126 | Category 1 | Set of drillbit adaptors for a small electric earmould drill compatible with item RT274-02-006 | No | No | Each
41 | RT274-01-127 | Category 1 | Power supply for electric hand drill, compatible with item RT274-02-006 | No | No | Each
42 | RT274-01-128 | Category 1 | Kidney bowl, stainless steel specialized ear irrigation bowl, approximately 15-25cm in length. Multiple award | No | No | Each
43 | RT274-01-129 | Category 1 | Kidney bowl, plastic specialized ear irrigation bowl, approximately 15-25cm in length. Multiple award | No | No | Each
44 | RT274-01-130 | Category 1 | Cut-out cup, ear syringing,stainless steel, receiver minimum 500ml | No | No | Each
45 | RT274-01-131 | Category 1 | Cut-out cup, ear syringing, plastic, receiver minimum 500ml | No | No | Each
46 | RT274-01-132 | Category 1 | Ear irrigation tips, screw onto special ear irrigation syringes, flared tips and produce divergent streams of water, supplier to specify quantity per pack and unit price. Compatible with RT274-03-004 awarded on current tender (RT274-2023) and all previous RT274 tenders. | Yes | No | Each
47 | RT274-01-133 | Category 1 | Syringe, ear irrigation, metal | Yes | No | Each
48 | RT274-01-134 | Category 1 | Curette, cerumen, paediatric, stainless steel | Yes | No | Each
49 | RT274-01-135 | Category 1 | Curette, cerumen, adult, stainless steel | Yes | No | Each
50 | RT274-01-136 | Category 1 | Hook, cerumen, stainless steel, for extraction of hard wax and other obstacles from the ear canal | Yes | No | Each
51 | RT274-01-137 | Category 1 | Hook, cerumen, plastic for extraction of hard wax and other obstacles from the ear canal | Yes | No | Each
52 | RT274-01-138 | Category 1 | Wax removal Helix tool | Yes | No | Each
53 | RT274-01-139 | Category 1 | Tool for removal of wax from earmould. Multiple award. | No | No | Each
54 | RT274-01-140 | Category 1 | Ear loop, billeau, small, stainless steel | Yes | No | Each
55 | RT274-01-141 | Category 1 | Ear loop, billeau, small, plastic | Yes | No | Each
56 | RT274-01-142 | Category 1 | Day hook, stainless steel for extraction of cerumen and other obstructions from the ear canal | Yes | No | Each
57 | RT274-01-143 | Category 1 | Curettes, plastic cerumen, paediatric, supplier to specify quantity per pack and unit price | Yes | No | Each
58 | RT274-01-144 | Category 1 | Day hook, stainless steel for extraction of curemen and other obstructions from the ear canal | Yes | No | Each
59 | RT274-01-145 | Category 1 | Day hook, plastic, for extraction of curemen and other obstructions from the ear canal | Yes | No | Each
60 | RT274-01-146 | Category 1 | Headlight, focused, battery operated, cordless, portable, for conducting various audiological procedures | No | No | Each
61 | RT274-01-147 | Category 1 | Detacheable, passive digital non-surgical bone conduction sound processor suitable for conductive hearing loss that does not require a softband and headband. May have telephone compatibility, direct audio input, datalogging capabilities, FM system compatibility where applicable. A pack of 6 approriately sized disposable batteries/appropriate rechargeable battery where indicated, activity/everyday case. A 2 year warranty on the sound processor must provided by the supplier. Processor must have reliability reporting and manufacturer certification (as per submitted requirements). | Yes | Yes | Each
62 | RT274-01-148 | Category 1 | Frame warmer/ heater for bone conduction spectacle hearing aid to adjust the temple of the spectacle frame for improve bone conduction fitting | No | No | Each
63 | RT274-01-149 | Category 1 | Replacement of the dummy bone conduction spectacle hearing aid temple in various colours and sizes. Must be compatible with all bone conduction hearing aids spectacles awarded | Yes | No | Each
64 | RT274-01-150 | Category 1 | Replacement plano lenses for digital bone conduction hearing aid spectacles, must be compatible with item RT274-04-001 awarded on the RT274-2023 tender. | No | No | Each
65 | RT274-01-151 | Category 1 | Replacement bone conduction transducer for all bone conduction hearing aids. Compatible with devices awarded on all previous RT274 tenders. | Yes | No | Each
66 | RT274-01-152 | Category 1 | Mountable/adhesive connector/pad/adaptor, hypo allergic, medical grade, water resistant, for passive non-surgical bone conduction sound processor. Compatible with devices awarded on all previous RT274 tenders. | Yes | No | Each
67 | RT274-01-153 | Category 1 | Super power digital rechargeable BTE for profound hearing loss (90+dB), MPO: 140+dB, gain 80+dB, with Bluetooth compatibility, with tinnitus management, with frequency-lowering. | Yes | Yes | Each
68 | RT274-01-154 | Category 1 | Super power digital rechargeable BTE for profound hearing loss (90+dB), MPO: 140+dB, gain 80+dB, without Bluetooth compatibility, with tinnitus management, with frequency-lowering. | Yes | No | Each
69 | RT274-01-155 | Category 1 | Super power digital rechargeable BTE for profound hearing loss (90+dB), MPO: 140+dB, gain 80+dB, with Bluetooth compatibility, without tinnitus management, with frequency-lowering. | Yes | Yes | Each
70 | RT274-01-156 | Category 1 | Super power digital rechargeable BTE for profound hearing loss (90+dB), MPO: 140+dB, gain 80+dB, without Bluetooth compatibility, without tinnitus management, with frequency-lowering. | Yes | No | Each
71 | RT274-01-157 | Category 1 | Super power digital rechargeable BTE for profound hearing loss (90+dB), MPO: 140+dB, gain 80+dB, with Bluetooth compatibility, with tinnitus management, without frequency-lowering. | Yes | Yes | Each
72 | RT274-01-158 | Category 1 | Super power digital rechargeable BTE for profound hearing loss (90+dB), MPO: 140+dB, gain 80+dB, without Bluetooth compatibility, with tinnitus management, without frequency-lowering. | Yes | No | Each
73 | RT274-01-159 | Category 1 | Replacement tamper proof battery doors for BTEs. Multiple Award | Yes | No | Each
74 | RT274-01-160 | Category 1 | Replacement battery doors for BTEs. Multiple Award | Yes | No | Each
75 | RT274-01-161 | Category 1 | Replacement slim tubes for for all BTEs awarded on RT274-2023 and previous RT274 transversal contracts. Multiple Award | Yes | No | Each
76 | RT274-01-162 | Category 1 | Replacement slim tube adaptors for for all BTEs awarded on RT274-2023 and previous RT274 transversal contracts. Multiple Award | Yes | No | Each
77 | RT274-01-163 | Category 1 | Replacement domes for for all BTEs awarded on RT274-2023 and previous RT274 transversal contracts. Multiple Award | Yes | No | Each
78 | RT274-01-164 | Category 1 | Tool for removing tone hooks (pins) to convert into slim tubes for all BTEs awarded on RT274-2023 and previous RT274 transversal contracts. Multiple Award | No | No | Each
79 | RT274-02-002 | Category 2 | Brush Tool for removing wax from Hearing Aids awarded on RT274-2023 and previous RT274 transversal contracts. Multiple Award | No | No | Each
80 | RT274-02-003 | Category 2 | Wax Tool for removing wax from slim tube Hearing Aids awarded on RT274-2023 and previous RT274 transversal contracts. Multiple Award | No | No | Each
81 | RT274-02-005 | Category 2 | Digital non-rechargeable ITE for moderate-severe hearing loss (56-70dB), MPO: 116-135dB gain 46-70dB, with Bluetooth compatibility, without tinnitus management, with frequency-lowering. | Yes | Yes | Each
82 | RT274-02-015 | Category 2 | Digital non-rechargeable ITE for severe-profound hearing loss (71-90dB), MPO: MPO: 126-145dB gain 61-85dB, with Bluetooth compatibility, with tinnitus management, with frequency-lowering. | Yes | Yes | Each
83 | RT274-02-016 | Category 2 | Digital non-rechargeable ITE for severe-profound hearing loss (71-90dB), MPO: MPO: 126-145dB gain 61-85dB, with Bluetooth compatibility, without tinnitus management, with frequency-lowering. | Yes | Yes | Each
84 | RT274-02-017 | Category 2 | Digital non-rechargeable ITC for moderate-severe hearing loss (56-70dB), MPO: 116-135dB gain 46-70dB, with Bluetooth compatibility, without tinnitus management, with frequency-lowering. | Yes | Yes | Each
85 | RT274-02-018 | Category 2 | Digital non-rechargeable ITC for moderate-severe hearing loss (56-70dB), MPO: 116-135dB gain 46-70dB, with Bluetooth compatibility, with tinnitus management, without frequency-lowering. | Yes | Yes | Each
86 | RT274-02-019 | Category 2 | Digital non-rechargeable ITC for severe-profound hearing loss (71-90dB), MPO: MPO: 126-145dB gain 61-85dB, with Bluetooth compatibility, with tinnitus management, with frequency-lowering. | Yes | Yes | Each
87 | RT274-02-006 | Category 2 | Digital non-rechargeable ITC for severe-profound hearing loss (71-90dB), MPO: MPO: 126-145dB gain 61-85dB, with Bluetooth compatibility, with tinnitus management, without frequency-lowering. | Yes | Yes | Each
88 | RT274-02-007 | Category 2 | Digital rechargeable CIC for mild-moderate hearing loss (25-55dB), MPO: 75-125dB gain 40-60dB, with Bluetooth compatibility, with tinnitus management, with frequency-lowering. | Yes | Yes | Each
89 | RT274-02-020 | Category 2 | Digital rechargeable CIC for mild-moderate hearing loss (25-55dB), MPO: 75-125dB gain 40-60dB, without Bluetooth compatibility, with tinnitus management, with frequency-lowering. | Yes | No | Each
90 | RT274-02-021 | Category 2 | Digital rechargeable CIC for mild-moderate hearing loss (25-55dB), MPO: 75-125dB gain 40-60dB, with Bluetooth compatibility, without tinnitus management, with frequency-lowering. | Yes | Yes | Each
91 | RT274-03-002 | Category 3 | Digital rechargeable CIC for moderate-severe hearing loss (56-70dB), MPO: 116-135dB gain 46-70dB, with Bluetooth compatibility, with tinnitus management, without frequency-lowering. | Yes | Yes | Each
92 | RT274-03-003 | Category 3 | Digital rechargeable CIC for moderate-severe hearing loss (56-70dB), MPO: 116-135dB gain 46-70dB, without Bluetooth compatibility, with tinnitus management, without frequency-lowering. | Yes | No | Each
93 | RT274-03-022 | Category 3 | Digital rechargeable CIC for moderate-severe hearing loss (56-70dB), MPO: 116-135dB gain 46-70dB, with Bluetooth compatibility, without tinnitus management, without frequency-lowering. | Yes | Yes | Each
94 | RT274-03-005 | Category 3 | Digital rechargeable CIC for severe-profound hearing loss (71-90dB), MPO: MPO: 126-145dB gain 61-85dB, with Bluetooth compatibility, with tinnitus management, with frequency-lowering. | Yes | Yes | Each
95 | RT274-03-006 | Category 3 | Digital rechargeable CIC for severe-profound hearing loss (71-90dB), MPO: MPO: 126-145dB gain 61-85dB, without Bluetooth compatibility, with tinnitus management, with frequency-lowering. | Yes | No | Each
96 | RT274-03-008 | Category 3 | Digital rechargeable CIC for severe-profound hearing loss (71-90dB), MPO: MPO: 126-145dB gain 61-85dB, without Bluetooth compatibility, without tinnitus management, with frequency-lowering. | Yes | No | Each
97 | RT274-03-009 | Category 3 | Digital rechargeable CIC for severe-profound hearing loss (71-90dB), MPO: MPO: 126-145dB gain 61-85dB, with Bluetooth compatibility, with tinnitus management, without frequency-lowering. | Yes | Yes | Each
98 | RT274-03-010 | Category 3 | Digital rechargeable CIC for severe-profound hearing loss (71-90dB), MPO: MPO: 126-145dB gain 61-85dB, without Bluetooth compatibility, with tinnitus management, without frequency-lowering. | Yes | No | Each
99 | RT274-03-023 | Category 3 | Digital rechargeable CIC for severe-profound hearing loss (71-90dB), MPO: MPO: 126-145dB gain 61-85dB, with Bluetooth compatibility, without tinnitus management, without frequency-lowering. | Yes | Yes | Each
100 | RT274-03-024 | Category 3 | Digital rechargeable CIC for severe-profound hearing loss (71-90dB), MPO: MPO: 126-145dB gain 61-85dB, without Bluetooth compatibility, without tinnitus management, without frequency-lowering. | Yes | No | Each
101 | RT274-03-012 | Category 3 | Batteries, dry cell, size AA, minimum card of 4-6 | Yes | No | Each
102 | RT274-03-014 | Category 3 | Alkaline 9V battery, IEC type 6lr61, flat type, 9V-500mal | Yes | No | Each
103 | RT274-03-025 | Category 3 | Rechargeable lithium-ion batteries size 13, for hearing aid | Yes | No | Each
104 | RT274-03-029 | Category 3 | Syringe, 10-15cm in length, double plunge, suitable for making earmould impressions | Yes | No | Each
105 | RT274-03-020 | Category 3 | Replacement mini sleeves for RIC. Multiple Awarded | Yes | No | Each
106 | RT274-03-030 | Category 3 | Replacement mini ear tips for RIC. Multiple Awarded | Yes | No | Each
107 | RT274-03-021 | Category 3 | CROS BTE, Digital rechargeable, wireless, behind the ear hearing aid, cross fitting, for contra lateral routing of signal, used to fit unilateral hearing loss. Complete with options of earmould, slim tube and dome, dry aid kit, | Yes | Yes | Each
108 | RT274-04-020 | Category 4 | Digital rechargeable BTE for bi- cross fitting, for contra lateral routing of signal, used to fit asymmetric hearing loss, where the one ear is dead and the other ear has a mild-moderate hearing loss (25-55dB), MPO: 75-120dB gain 40-60dB, without Bluetooth compatibility, with tinnitus management, without frequency-lowering. | Yes | No | Each
109 | RT274-04-026 | Category 4 | Digital rechargeable BTE for bi- cross fitting, for contra lateral routing of signal, used to fit asymmetric hearing loss, where the one ear is dead and the other ear has a moderate-severe hearing loss (56-70dB), MPO: 116-130dB gain 46-65dB, without Bluetooth compatibility, without tinnitus management, with frequency-lowering. | Yes | No | Each
110 | RT274-04-027 | Category 4 | Digital rechargeable BTE for bi- cross fitting, for contra lateral routing of signal, used to fit asymmetric hearing loss, where the one ear is dead and the other ear has a moderate-severe hearing loss (56-70dB), MPO: 116-130dB gain 46-65dB, with Bluetooth compatibility, with tinnitus management, without frequency-lowering. | Yes | Yes | Each
111 | RT274-04-033 | Category 4 | Digital rechargeable BTE for bi- cross fitting, for contra lateral routing of signal, used to fit asymmetric hearing loss, where the one ear is dead and the other ear has a severe-profound hearing loss (71-100dB), MPO: MPO: 120-145dB gain 60-85dB, with Bluetooth compatibility, without tinnitus management, with frequency-lowering. | Yes | Yes | Each
112 | RT274-04-034 | Category 4 | Digital rechargeable BTE for bi- cross fitting, for contra lateral routing of signal, used to fit asymmetric hearing loss, where the one ear is dead and the other ear has a severe-profound hearing loss (71-100dB), MPO: MPO: 120-145dB gain 60-85dB, without Bluetooth compatibility, without tinnitus management, with frequency-lowering. | Yes | No | Each
113 | RT274-04-035 | Category 4 | Digital rechargeable BTE for bi- cross fitting, for contra lateral routing of signal, used to fit asymmetric hearing loss, where the one ear is dead and the other ear has a severe-profound hearing loss (71-100dB), MPO: MPO: 120-145dB gain 60-85dB, with Bluetooth compatibility, with tinnitus management, without frequency-lowering. | Yes | Yes | Each
114 | RT274-05-049 | Category 5 | Replacement battery doors for bone conduction sound processors. NB: All dB values in Category 5 specifications compy to the below requrements: Must be according to the 2cc coupler readings, All BTE hearing aids must contain the following as minimum specifications: Technology obsolescence: not older than 18 months from date of publication of tender, Active noise reduction/speech enhancement, Active noise reduction/speech enhancement, Minimum of 4 programs, Minimum of 8 channels, Minimum of IP57 rating, Programmable directionality, Data logging, Feedback management, Volume control, Telecoil, All hearing aids to include an ear mould (acrylic or silicone) or 2 slim tubes and domes, All non-rechargeable hearing aids must include 6 appropriately sized zinc air batteries and dry aid kit, All re-chargeable hearing aids must be provided with a charger and case and dry aid kit | Yes | No | Each
115 | RT274-05-051 | Category 5 | Retention devices for all bone conduction hearing devices issued within South Africa over the past 5 years. NB: All dB values in Category 5 specifications compy to the below requrements: Must be according to the 2cc coupler readings, All BTE hearing aids must contain the following as minimum specifications: Technology obsolescence: not older than 18 months from date of publication of tender, Active noise reduction/speech enhancement, Active noise reduction/speech enhancement, Minimum of 4 programs, Minimum of 8 channels, Minimum of IP57 rating, Programmable directionality, Data logging, Feedback management, Volume control, Telecoil, All hearing aids to include an ear mould (acrylic or silicone) or 2 slim tubes and domes, All non-rechargeable hearing aids must include 6 appropriately sized zinc air batteries and dry aid kit, All re-chargeable hearing aids must be provided with a charger and case and dry aid kit | Yes | No | Each
116 | RT274-05-053 | Category 5 | Magnetic digital bone conduction sound processor (one processor only; no surgical kit) compatible with all active transcutaneous bone conduction hearing implants awarded on the tender as well as all previous devices issued in South Africa over the past 5 years. The sound processor should be suitable for single sided deafness, conductive and mixed hearing losses with bone conduction thresholds of up to 45 dBHL (normal to moderate hearing thresholds). Sound processor kit to comprise of a minimum of: one sound processing unit, one test rod/skull simulator where indicated, pack of 6 appropriately sized disposable batteries/appropriate rechargeable battery where indicated, drying capsule/kit, sound processor cover (optional), an activity/everyday case, tamperproof locks and indicator lights. The sound processor should with have the option of both omnidirectional and directional microphones, have capability to be programmed as a standalone or via an interfaced software, have telephone compatibility, direct audio input, datalogging and FM system compatibility and streaming compatibility. Sound processor to be available in a range of colours, water resistance rating of IP54 for temporary emersion. Processor must have reliability reporting and manufacturer certification (as per submitted requirements), a 2-year warranty on sound processor and aftercare service must be guaranteed by the manufacturer. | Yes | Yes | Each
117 | RT274-05-055 | Category 5 | Surgical Vibrating ossicular prosthesis/ active middle ear implant system: one external sound processor, one vibrating ossicular or round window prosthesis surgically implantable in the middle ear with a wire that connects to, one partially implantable receptor and transducer capable of providing piezoelectric, electromagnetic or electromechanical vibrational energy. Suitable for the treatment of certain sensorineural hearing losses, conductive hearing losses and mixed hearing losses. Sound processor kit to comprise of a minimum: one sound processing unit, one test rod where indicated, magnetic spacers with variable strength to ensure patient comfort and secure fit, pack of 6 appropriately sized disposable batteries/appropriate rechargeable battery where indicated, drying capsule/kit, sound processor cover (optional), an activity/everyday case, tamperproof locks and/or indicator lights. The sound processor should with have the option of both omnidirectional and directional microphones, have capability to be programmed as a standalone or via an interfaced software, have telephone compatibility, ability to do insitu audiometry/threshold, datalogging capabilities, FM system compatibility and/streaming capabilities. Sound processor to be available in a range of colours, water resistance rating of IP54 for temporary emersion. Processor must have reliability reporting and manufacturer certification (as per submitted requirements), a 2-year warranty on sound processor and aftercare service must be guaranteed by the manufacturer. | Yes | No | Each
118 | RT274-05-057 | Category 5 | Magnetic digital sound processor (one processor only; no surgical kit) compatible with all surgical vibrating ossicular/round window prosthesis/active middle ear implants awarded on the tender as well as all previous devices issued in South Africa over the past 5 years. Suitable for the treatment of certain sensorineural hearing losses, conductive hearing losses and mixed hearing losses. Sound processor kit to comprise of a minimum: one sound processing unit, one test rod where indicated, magnetic spacers with variable strength to ensure patient comfort and secure fit, pack of 6 appropriately sized disposable batteries/appropriate rechargeable battery where indicated, drying capsule/kit, sound processor cover (optional), an activity/everyday case, tamperproof locks and/or indicator lights. The sound processor should with have the option of both omnidirectional and directional microphones, have capability to be programmed as a standalone or via an interfaced software, have telephone compatibility, ability to do insitu audiometry/threshold, datalogging capabilities, FM system compatibility and/streaming capabilities. Sound processor to be available in a range of colours, water resistance rating of IP54 for temporary emersion. Processor must have reliability reporting and manufacturer certification (as per submitted requirements), a 2-year warranty on sound processor and aftercare service must be guaranteed by the manufacturer. NB: All dB values in Category 5 specifications compy to the below requrements: Must be according to the 2cc coupler readings, All BTE hearing aids must contain the following as minimum specifications: Technology obsolescence: not older than 18 months from date of publication of tender, Active noise reduction/speech enhancement, Active noise reduction/speech enhancement, Minimum of 4 programs, Minimum of 8 channels, Minimum of IP57 rating, Programmable directionality, Data logging, Feedback management, Volume control, Telecoil, All hearing aids to include an ear mould (acrylic or silicone) or 2 slim tubes and domes, All non-rechargeable hearing aids must include 6 appropriately sized zinc air batteries and dry aid kit, All re-chargeable hearing aids must be provided with a charger and case and dry aid kit | Yes | Yes | Each
119 | RT274-05-065 | Category 5 | Replacement tamper proof battery doors for BTEs. Multiple Award. NB: All dB values in Category 5 specifications compy to the below requrements: Must be according to the 2cc coupler readings, All BTE hearing aids must contain the following as minimum specifications: Technology obsolescence: not older than 18 months from date of publication of tender, Active noise reduction/speech enhancement, Active noise reduction/speech enhancement, Minimum of 4 programs, Minimum of 8 channels, Minimum of IP57 rating, Programmable directionality, Data logging, Feedback management, Volume control, Telecoil, All hearing aids to include an ear mould (acrylic or silicone) or 2 slim tubes and domes, All non-rechargeable hearing aids must include 6 appropriately sized zinc air batteries and dry aid kit, All re-chargeable hearing aids must be provided with a charger and case and dry aid kit | Yes | No | Each
120 | RT274-05-066 | Category 5 | Replacement battery doors for BTEs. Multiple Award. NB: All dB values in Category 5 specifications compy to the below requrements: Must be according to the 2cc coupler readings, All BTE hearing aids must contain the following as minimum specifications: Technology obsolescence: not older than 18 months from date of publication of tender, Active noise reduction/speech enhancement, Active noise reduction/speech enhancement, Minimum of 4 programs, Minimum of 8 channels, Minimum of IP57 rating, Programmable directionality, Data logging, Feedback management, Volume control, Telecoil, All hearing aids to include an ear mould (acrylic or silicone) or 2 slim tubes and domes, All non-rechargeable hearing aids must include 6 appropriately sized zinc air batteries and dry aid kit, All re-chargeable hearing aids must be provided with a charger and case and dry aid kit | Yes | No | Each
121 | RT274-05-067 | Category 5 | Replacement slim tubes for for all BTEs awarded on RT274-2023 and previous RT274 transversal contracts. Multiple Award. NB: All dB values in Category 5 specifications compy to the below requrements: Must be according to the 2cc coupler readings, All BTE hearing aids must contain the following as minimum specifications: Technology obsolescence: not older than 18 months from date of publication of tender, Active noise reduction/speech enhancement, Active noise reduction/speech enhancement, Minimum of 4 programs, Minimum of 8 channels, Minimum of IP57 rating, Programmable directionality, Data logging, Feedback management, Volume control, Telecoil, All hearing aids to include an ear mould (acrylic or silicone) or 2 slim tubes and domes, All non-rechargeable hearing aids must include 6 appropriately sized zinc air batteries and dry aid kit, All re-chargeable hearing aids must be provided with a charger and case and dry aid kit | Yes | No | Each
122 | RT274-05-068 | Category 5 | Replacement slim tube adaptors for for all BTEs awarded on RT274-2023 and previous RT274 transversal contracts. Multiple Award. NB: All dB values in Category 5 specifications compy to the below requrements: Must be according to the 2cc coupler readings, All BTE hearing aids must contain the following as minimum specifications: Technology obsolescence: not older than 18 months from date of publication of tender, Active noise reduction/speech enhancement, Active noise reduction/speech enhancement, Minimum of 4 programs, Minimum of 8 channels, Minimum of IP57 rating, Programmable directionality, Data logging, Feedback management, Volume control, Telecoil, All hearing aids to include an ear mould (acrylic or silicone) or 2 slim tubes and domes, All non-rechargeable hearing aids must include 6 appropriately sized zinc air batteries and dry aid kit, All re-chargeable hearing aids must be provided with a charger and case and dry aid kit | Yes | No | Each
123 | RT274-05-069 | Category 5 | Replacement domes for for all BTEs awarded on RT274-2023 and previous RT274 transversal contracts. Multiple Award. NB: All dB values in Category 5 specifications compy to the below requrements: Must be according to the 2cc coupler readings, All BTE hearing aids must contain the following as minimum specifications: Technology obsolescence: not older than 18 months from date of publication of tender, Active noise reduction/speech enhancement, Active noise reduction/speech enhancement, Minimum of 4 programs, Minimum of 8 channels, Minimum of IP57 rating, Programmable directionality, Data logging, Feedback management, Volume control, Telecoil, All hearing aids to include an ear mould (acrylic or silicone) or 2 slim tubes and domes, All non-rechargeable hearing aids must include 6 appropriately sized zinc air batteries and dry aid kit, All re-chargeable hearing aids must be provided with a charger and case and dry aid kit | Yes | No | Each
124 | RT274-05-070 | Category 5 | Tool for removing tone hooks (pins) to convert into slim tubes for all BTEs awarded on RT274-2023 and previous RT274 transversal contracts. Multiple Award. NB: All dB values in Category 5 specifications compy to the below requrements: Must be according to the 2cc coupler readings, All BTE hearing aids must contain the following as minimum specifications: Technology obsolescence: not older than 18 months from date of publication of tender, Active noise reduction/speech enhancement, Active noise reduction/speech enhancement, Minimum of 4 programs, Minimum of 8 channels, Minimum of IP57 rating, Programmable directionality, Data logging, Feedback management, Volume control, Telecoil, All hearing aids to include an ear mould (acrylic or silicone) or 2 slim tubes and domes, All non-rechargeable hearing aids must include 6 appropriately sized zinc air batteries and dry aid kit, All re-chargeable hearing aids must be provided with a charger and case and dry aid kit | No | No | Each
125 | RT274-05-071 | Category 5 | Magnet Tool for removing battery from battery door for all Hearing Aids awarded on RT274-2023 and previous RT274 transversal contracts. Multiple Award. NB: All dB values in Category 5 specifications compy to the below requrements: Must be according to the 2cc coupler readings, All BTE hearing aids must contain the following as minimum specifications: Technology obsolescence: not older than 18 months from date of publication of tender, Active noise reduction/speech enhancement, Active noise reduction/speech enhancement, Minimum of 4 programs, Minimum of 8 channels, Minimum of IP57 rating, Programmable directionality, Data logging, Feedback management, Volume control, Telecoil, All hearing aids to include an ear mould (acrylic or silicone) or 2 slim tubes and domes, All non-rechargeable hearing aids must include 6 appropriately sized zinc air batteries and dry aid kit, All re-chargeable hearing aids must be provided with a charger and case and dry aid kit | No | No | Each
126 | RT274-05-072 | Category 5 | Brush Tool for removing wax from Hearing Aids awarded on RT274-2023 and previous RT274 transversal contracts. Multiple Award. NB: All dB values in Category 5 specifications compy to the below requrements: Must be according to the 2cc coupler readings, All BTE hearing aids must contain the following as minimum specifications: Technology obsolescence: not older than 18 months from date of publication of tender, Active noise reduction/speech enhancement, Active noise reduction/speech enhancement, Minimum of 4 programs, Minimum of 8 channels, Minimum of IP57 rating, Programmable directionality, Data logging, Feedback management, Volume control, Telecoil, All hearing aids to include an ear mould (acrylic or silicone) or 2 slim tubes and domes, All non-rechargeable hearing aids must include 6 appropriately sized zinc air batteries and dry aid kit, All re-chargeable hearing aids must be provided with a charger and case and dry aid kit | No | No | Each
127 | RT274-05-073 | Category 5 | Wax Tool for removing wax from slim tube Hearing Aids awarded on RT274-2023 and previous RT274 transversal contracts. Multiple Award. NB: All dB values in Category 5 specifications compy to the below requrements: Must be according to the 2cc coupler readings, All BTE hearing aids must contain the following as minimum specifications: Technology obsolescence: not older than 18 months from date of publication of tender, Active noise reduction/speech enhancement, Active noise reduction/speech enhancement, Minimum of 4 programs, Minimum of 8 channels, Minimum of IP57 rating, Programmable directionality, Data logging, Feedback management, Volume control, Telecoil, All hearing aids to include an ear mould (acrylic or silicone) or 2 slim tubes and domes, All non-rechargeable hearing aids must include 6 appropriately sized zinc air batteries and dry aid kit, All re-chargeable hearing aids must be provided with a charger and case and dry aid kit | No | No | Each
128 | RT274-05-074 | Category 5 | Water-resistant stickers designed specifically to attach to hearing aids, to include paediatric-friendly options. NB: All dB values in Category 5 specifications compy to the below requrements: Must be according to the 2cc coupler readings, All BTE hearing aids must contain the following as minimum specifications: Technology obsolescence: not older than 18 months from date of publication of tender, Active noise reduction/speech enhancement, Active noise reduction/speech enhancement, Minimum of 4 programs, Minimum of 8 channels, Minimum of IP57 rating, Programmable directionality, Data logging, Feedback management, Volume control, Telecoil, All hearing aids to include an ear mould (acrylic or silicone) or 2 slim tubes and domes, All non-rechargeable hearing aids must include 6 appropriately sized zinc air batteries and dry aid kit, All re-chargeable hearing aids must be provided with a charger and case and dry aid kit | Yes | No | Each
129 | RT274-06-022 | Category 6 | Digital non-rechargeable ITE for moderate-severe hearing loss (56-70dB), MPO: 116-135dB gain 46-70dB, with Bluetooth compatibility, without tinnitus management, with frequency-lowering. NB: all dB values in the category 6 must comply to the following specifications: Must be according to the 2cc coupler readings, Minimum specification for all Custom hearing aids, Digital signal processing, Technology obsolescence: not older than 18 months from date of publication of tender, Active noise reduction/speech enhancement, Minimum of 4 programs, Minimum of 8 channels, Minimum of IP57 rating, Data logging, Feedback management, Volume control wheel or button for ITE/ITC, Non-rechargeable hearing aids must come with 6 appropriately sized zinc air batteries and dryaid kit and waxguard set if applicable, All re-chargeable hearing aids must be provided with a charger and case and dryaid kit or capsule and waxguard set if applicable | Yes | Yes | Each
130 | RT274-06-034 | Category 6 | Digital non-rechargeable ITE for severe-profound hearing loss (71-90dB), MPO: MPO: 126-145dB gain 61-85dB, with Bluetooth compatibility, with tinnitus management, with frequency-lowering. NB: all dB values in the category 6 must comply to the following specifications: Must be according to the 2cc coupler readings, Minimum specification for all Custom hearing aids, Digital signal processing, Technology obsolescence: not older than 18 months from date of publication of tender, Active noise reduction/speech enhancement, Minimum of 4 programs, Minimum of 8 channels, Minimum of IP57 rating, Data logging, Feedback management, Volume control wheel or button for ITE/ITC, Non-rechargeable hearing aids must come with 6 appropriately sized zinc air batteries and dryaid kit and waxguard set if applicable, All re-chargeable hearing aids must be provided with a charger and case and dryaid kit or capsule and waxguard set if applicable | Yes | Yes | Each
131 | RT274-06-038 | Category 6 | Digital non-rechargeable ITE for severe-profound hearing loss (71-90dB), MPO: MPO: 126-145dB gain 61-85dB, with Bluetooth compatibility, without tinnitus management, with frequency-lowering. NB: all dB values in the category 6 must comply to the following specifications: Must be according to the 2cc coupler readings, Minimum specification for all Custom hearing aids, Digital signal processing, Technology obsolescence: not older than 18 months from date of publication of tender, Active noise reduction/speech enhancement, Minimum of 4 programs, Minimum of 8 channels, Minimum of IP57 rating, Data logging, Feedback management, Volume control wheel or button for ITE/ITC, Non-rechargeable hearing aids must come with 6 appropriately sized zinc air batteries and dryaid kit and waxguard set if applicable, All re-chargeable hearing aids must be provided with a charger and case and dryaid kit or capsule and waxguard set if applicable | Yes | Yes | Each
132 | RT274-06-070 | Category 6 | Digital non-rechargeable ITC for moderate-severe hearing loss (56-70dB), MPO: 116-135dB gain 46-70dB, with Bluetooth compatibility, without tinnitus management, with frequency-lowering. NB: all dB values in the category 6 must comply to the following specifications: Must be according to the 2cc coupler readings, Minimum specification for all Custom hearing aids, Digital signal processing, Technology obsolescence: not older than 18 months from date of publication of tender, Active noise reduction/speech enhancement, Minimum of 4 programs, Minimum of 8 channels, Minimum of IP57 rating, Data logging, Feedback management, Volume control wheel or button for ITE/ITC, Non-rechargeable hearing aids must come with 6 appropriately sized zinc air batteries and dryaid kit and waxguard set if applicable, All re-chargeable hearing aids must be provided with a charger and case and dryaid kit or capsule and waxguard set if applicable | Yes | Yes | Each
133 | RT274-06-074 | Category 6 | Digital non-rechargeable ITC for moderate-severe hearing loss (56-70dB), MPO: 116-135dB gain 46-70dB, with Bluetooth compatibility, with tinnitus management, without frequency-lowering. NB: all dB values in the category 6 must comply to the following specifications: Must be according to the 2cc coupler readings, Minimum specification for all Custom hearing aids, Digital signal processing, Technology obsolescence: not older than 18 months from date of publication of tender, Active noise reduction/speech enhancement, Minimum of 4 programs, Minimum of 8 channels, Minimum of IP57 rating, Data logging, Feedback management, Volume control wheel or button for ITE/ITC, Non-rechargeable hearing aids must come with 6 appropriately sized zinc air batteries and dryaid kit and waxguard set if applicable, All re-chargeable hearing aids must be provided with a charger and case and dryaid kit or capsule and waxguard set if applicable | Yes | Yes | Each
134 | RT274-06-082 | Category 6 | Digital non-rechargeable ITC for severe-profound hearing loss (71-90dB), MPO: MPO: 126-145dB gain 61-85dB, with Bluetooth compatibility, with tinnitus management, with frequency-lowering. NB: all dB values in the category 6 must comply to the following specifications: Must be according to the 2cc coupler readings, Minimum specification for all Custom hearing aids, Digital signal processing, Technology obsolescence: not older than 18 months from date of publication of tender, Active noise reduction/speech enhancement, Minimum of 4 programs, Minimum of 8 channels, Minimum of IP57 rating, Data logging, Feedback management, Volume control wheel or button for ITE/ITC, Non-rechargeable hearing aids must come with 6 appropriately sized zinc air batteries and dryaid kit and waxguard set if applicable, All re-chargeable hearing aids must be provided with a charger and case and dryaid kit or capsule and waxguard set if applicable | Yes | Yes | Each
135 | RT274-06-091 | Category 6 | Digital non-rechargeable ITC for severe-profound hearing loss (71-90dB), MPO: MPO: 126-145dB gain 61-85dB, with Bluetooth compatibility, with tinnitus management, without frequency-lowering. NB: all dB values in the category 6 must comply to the following specifications: Must be according to the 2cc coupler readings, Minimum specification for all Custom hearing aids, Digital signal processing, Technology obsolescence: not older than 18 months from date of publication of tender, Active noise reduction/speech enhancement, Minimum of 4 programs, Minimum of 8 channels, Minimum of IP57 rating, Data logging, Feedback management, Volume control wheel or button for ITE/ITC, Non-rechargeable hearing aids must come with 6 appropriately sized zinc air batteries and dryaid kit and waxguard set if applicable, All re-chargeable hearing aids must be provided with a charger and case and dryaid kit or capsule and waxguard set if applicable | Yes | Yes | Each
136 | RT274-06-098 | Category 6 | Digital rechargeable CIC for mild-moderate hearing loss (25-55dB), MPO: 75-125dB gain 40-60dB, with Bluetooth compatibility, with tinnitus management, with frequency-lowering. NB: all dB values in the category 6 must comply to the following specifications: Must be according to the 2cc coupler readings, Minimum specification for all Custom hearing aids, Digital signal processing, Technology obsolescence: not older than 18 months from date of publication of tender, Active noise reduction/speech enhancement, Minimum of 4 programs, Minimum of 8 channels, Minimum of IP57 rating, Data logging, Feedback management, Volume control wheel or button for ITE/ITC, Non-rechargeable hearing aids must come with 6 appropriately sized zinc air batteries and dryaid kit and waxguard set if applicable, All re-chargeable hearing aids must be provided with a charger and case and dryaid kit or capsule and waxguard set if applicable | Yes | Yes | Each
137 | RT274-06-100 | Category 6 | Digital rechargeable CIC for mild-moderate hearing loss (25-55dB), MPO: 75-125dB gain 40-60dB, without Bluetooth compatibility, with tinnitus management, with frequency-lowering. NB: all dB values in the category 6 must comply to the following specifications: Must be according to the 2cc coupler readings, Minimum specification for all Custom hearing aids, Digital signal processing, Technology obsolescence: not older than 18 months from date of publication of tender, Active noise reduction/speech enhancement, Minimum of 4 programs, Minimum of 8 channels, Minimum of IP57 rating, Data logging, Feedback management, Volume control wheel or button for ITE/ITC, Non-rechargeable hearing aids must come with 6 appropriately sized zinc air batteries and dryaid kit and waxguard set if applicable, All re-chargeable hearing aids must be provided with a charger and case and dryaid kit or capsule and waxguard set if applicable | Yes | No | Each
138 | RT274-06-102 | Category 6 | Digital rechargeable CIC for mild-moderate hearing loss (25-55dB), MPO: 75-125dB gain 40-60dB, with Bluetooth compatibility, without tinnitus management, with frequency-lowering. NB: all dB values in the category 6 must comply to the following specifications: Must be according to the 2cc coupler readings, Minimum specification for all Custom hearing aids, Digital signal processing, Technology obsolescence: not older than 18 months from date of publication of tender, Active noise reduction/speech enhancement, Minimum of 4 programs, Minimum of 8 channels, Minimum of IP57 rating, Data logging, Feedback management, Volume control wheel or button for ITE/ITC, Non-rechargeable hearing aids must come with 6 appropriately sized zinc air batteries and dryaid kit and waxguard set if applicable, All re-chargeable hearing aids must be provided with a charger and case and dryaid kit or capsule and waxguard set if applicable | Yes | Yes | Each
139 | RT274-06-104 | Category 6 | Digital rechargeable CIC for mild-moderate hearing loss (25-55dB), MPO: 75-125dB gain 40-60dB, without Bluetooth compatibility, without tinnitus management, with frequency-lowering. NB: all dB values in the category 6 must comply to the following specifications: Must be according to the 2cc coupler readings, Minimum specification for all Custom hearing aids, Digital signal processing, Technology obsolescence: not older than 18 months from date of publication of tender, Active noise reduction/speech enhancement, Minimum of 4 programs, Minimum of 8 channels, Minimum of IP57 rating, Data logging, Feedback management, Volume control wheel or button for ITE/ITC, Non-rechargeable hearing aids must come with 6 appropriately sized zinc air batteries and dryaid kit and waxguard set if applicable, All re-chargeable hearing aids must be provided with a charger and case and dryaid kit or capsule and waxguard set if applicable | Yes | No | Each
140 | RT274-06-106 | Category 6 | Digital rechargeable CIC for mild-moderate hearing loss (25-55dB), MPO: 75-125dB gain 40-60dB, with Bluetooth compatibility, with tinnitus management, without frequency-lowering. NB: all dB values in the category 6 must comply to the following specifications: Must be according to the 2cc coupler readings, Minimum specification for all Custom hearing aids, Digital signal processing, Technology obsolescence: not older than 18 months from date of publication of tender, Active noise reduction/speech enhancement, Minimum of 4 programs, Minimum of 8 channels, Minimum of IP57 rating, Data logging, Feedback management, Volume control wheel or button for ITE/ITC, Non-rechargeable hearing aids must come with 6 appropriately sized zinc air batteries and dryaid kit and waxguard set if applicable, All re-chargeable hearing aids must be provided with a charger and case and dryaid kit or capsule and waxguard set if applicable | Yes | Yes | Each
141 | RT274-06-108 | Category 6 | Digital rechargeable CIC for mild-moderate hearing loss (25-55dB), MPO: 75-125dB gain 40-60dB, without Bluetooth compatibility, with tinnitus management, without frequency-lowering. NB: all dB values in the category 6 must comply to the following specifications: Must be according to the 2cc coupler readings, Minimum specification for all Custom hearing aids, Digital signal processing, Technology obsolescence: not older than 18 months from date of publication of tender, Active noise reduction/speech enhancement, Minimum of 4 programs, Minimum of 8 channels, Minimum of IP57 rating, Data logging, Feedback management, Volume control wheel or button for ITE/ITC, Non-rechargeable hearing aids must come with 6 appropriately sized zinc air batteries and dryaid kit and waxguard set if applicable, All re-chargeable hearing aids must be provided with a charger and case and dryaid kit or capsule and waxguard set if applicable | Yes | No | Each
142 | RT274-06-110 | Category 6 | Digital rechargeable CIC for mild-moderate hearing loss (25-55dB), MPO: 75-125dB gain 40-60dB, with Bluetooth compatibility, without tinnitus management, without frequency-lowering. NB: all dB values in the category 6 must comply to the following specifications: Must be according to the 2cc coupler readings, Minimum specification for all Custom hearing aids, Digital signal processing, Technology obsolescence: not older than 18 months from date of publication of tender, Active noise reduction/speech enhancement, Minimum of 4 programs, Minimum of 8 channels, Minimum of IP57 rating, Data logging, Feedback management, Volume control wheel or button for ITE/ITC, Non-rechargeable hearing aids must come with 6 appropriately sized zinc air batteries and dryaid kit and waxguard set if applicable, All re-chargeable hearing aids must be provided with a charger and case and dryaid kit or capsule and waxguard set if applicable | Yes | Yes | Each
143 | RT274-06-112 | Category 6 | Digital rechargeable CIC for mild-moderate hearing loss (25-55dB), MPO: 75-125dB gain 40-60dB, without Bluetooth compatibility, without tinnitus management, without frequency-lowering. NB: all dB values in the category 6 must comply to the following specifications: Must be according to the 2cc coupler readings, Minimum specification for all Custom hearing aids, Digital signal processing, Technology obsolescence: not older than 18 months from date of publication of tender, Active noise reduction/speech enhancement, Minimum of 4 programs, Minimum of 8 channels, Minimum of IP57 rating, Data logging, Feedback management, Volume control wheel or button for ITE/ITC, Non-rechargeable hearing aids must come with 6 appropriately sized zinc air batteries and dryaid kit and waxguard set if applicable, All re-chargeable hearing aids must be provided with a charger and case and dryaid kit or capsule and waxguard set if applicable | Yes | No | Each
144 | RT274-06-114 | Category 6 | Digital rechargeable CIC for moderate-severe hearing loss (56-70dB), MPO: 116-135dB gain 46-70dB, with Bluetooth compatibility, with tinnitus management, with frequency-lowering. NB: all dB values in the category 6 must comply to the following specifications: Must be according to the 2cc coupler readings, Minimum specification for all Custom hearing aids, Digital signal processing, Technology obsolescence: not older than 18 months from date of publication of tender, Active noise reduction/speech enhancement, Minimum of 4 programs, Minimum of 8 channels, Minimum of IP57 rating, Data logging, Feedback management, Volume control wheel or button for ITE/ITC, Non-rechargeable hearing aids must come with 6 appropriately sized zinc air batteries and dryaid kit and waxguard set if applicable, All re-chargeable hearing aids must be provided with a charger and case and dryaid kit or capsule and waxguard set if applicable | Yes | Yes | Each
145 | RT274-06-116 | Category 6 | Digital rechargeable CIC for moderate-severe hearing loss (56-70dB), MPO: 116-135dB gain 46-70dB, without Bluetooth compatibility, with tinnitus management, with frequency-lowering. NB: all dB values in the category 6 must comply to the following specifications: Must be according to the 2cc coupler readings, Minimum specification for all Custom hearing aids, Digital signal processing, Technology obsolescence: not older than 18 months from date of publication of tender, Active noise reduction/speech enhancement, Minimum of 4 programs, Minimum of 8 channels, Minimum of IP57 rating, Data logging, Feedback management, Volume control wheel or button for ITE/ITC, Non-rechargeable hearing aids must come with 6 appropriately sized zinc air batteries and dryaid kit and waxguard set if applicable, All re-chargeable hearing aids must be provided with a charger and case and dryaid kit or capsule and waxguard set if applicable | Yes | No | Each
146 | RT274-06-118 | Category 6 | Digital rechargeable CIC for moderate-severe hearing loss (56-70dB), MPO: 116-135dB gain 46-70dB, with Bluetooth compatibility, without tinnitus management, with frequency-lowering. NB: all dB values in the category 6 must comply to the following specifications: Must be according to the 2cc coupler readings, Minimum specification for all Custom hearing aids, Digital signal processing, Technology obsolescence: not older than 18 months from date of publication of tender, Active noise reduction/speech enhancement, Minimum of 4 programs, Minimum of 8 channels, Minimum of IP57 rating, Data logging, Feedback management, Volume control wheel or button for ITE/ITC, Non-rechargeable hearing aids must come with 6 appropriately sized zinc air batteries and dryaid kit and waxguard set if applicable, All re-chargeable hearing aids must be provided with a charger and case and dryaid kit or capsule and waxguard set if applicable | Yes | Yes | Each
147 | RT274-06-120 | Category 6 | Digital rechargeable CIC for moderate-severe hearing loss (56-70dB), MPO: 116-135dB gain 46-70dB, without Bluetooth compatibility, without tinnitus management, with frequency-lowering. NB: all dB values in the category 6 must comply to the following specifications: Must be according to the 2cc coupler readings, Minimum specification for all Custom hearing aids, Digital signal processing, Technology obsolescence: not older than 18 months from date of publication of tender, Active noise reduction/speech enhancement, Minimum of 4 programs, Minimum of 8 channels, Minimum of IP57 rating, Data logging, Feedback management, Volume control wheel or button for ITE/ITC, Non-rechargeable hearing aids must come with 6 appropriately sized zinc air batteries and dryaid kit and waxguard set if applicable, All re-chargeable hearing aids must be provided with a charger and case and dryaid kit or capsule and waxguard set if applicable | Yes | No | Each
148 | RT274-06-122 | Category 6 | Digital rechargeable CIC for moderate-severe hearing loss (56-70dB), MPO: 116-135dB gain 46-70dB, with Bluetooth compatibility, with tinnitus management, without frequency-lowering. NB: all dB values in the category 6 must comply to the following specifications: Must be according to the 2cc coupler readings, Minimum specification for all Custom hearing aids, Digital signal processing, Technology obsolescence: not older than 18 months from date of publication of tender, Active noise reduction/speech enhancement, Minimum of 4 programs, Minimum of 8 channels, Minimum of IP57 rating, Data logging, Feedback management, Volume control wheel or button for ITE/ITC, Non-rechargeable hearing aids must come with 6 appropriately sized zinc air batteries and dryaid kit and waxguard set if applicable, All re-chargeable hearing aids must be provided with a charger and case and dryaid kit or capsule and waxguard set if applicable | Yes | Yes | Each
149 | RT274-06-124 | Category 6 | Digital rechargeable CIC for moderate-severe hearing loss (56-70dB), MPO: 116-135dB gain 46-70dB, without Bluetooth compatibility, with tinnitus management, without frequency-lowering. NB: all dB values in the category 6 must comply to the following specifications: Must be according to the 2cc coupler readings, Minimum specification for all Custom hearing aids, Digital signal processing, Technology obsolescence: not older than 18 months from date of publication of tender, Active noise reduction/speech enhancement, Minimum of 4 programs, Minimum of 8 channels, Minimum of IP57 rating, Data logging, Feedback management, Volume control wheel or button for ITE/ITC, Non-rechargeable hearing aids must come with 6 appropriately sized zinc air batteries and dryaid kit and waxguard set if applicable, All re-chargeable hearing aids must be provided with a charger and case and dryaid kit or capsule and waxguard set if applicable | Yes | No | Each
150 | RT274-06-126 | Category 6 | Digital rechargeable CIC for moderate-severe hearing loss (56-70dB), MPO: 116-135dB gain 46-70dB, with Bluetooth compatibility, without tinnitus management, without frequency-lowering. NB: all dB values in the category 6 must comply to the following specifications: Must be according to the 2cc coupler readings, Minimum specification for all Custom hearing aids, Digital signal processing, Technology obsolescence: not older than 18 months from date of publication of tender, Active noise reduction/speech enhancement, Minimum of 4 programs, Minimum of 8 channels, Minimum of IP57 rating, Data logging, Feedback management, Volume control wheel or button for ITE/ITC, Non-rechargeable hearing aids must come with 6 appropriately sized zinc air batteries and dryaid kit and waxguard set if applicable, All re-chargeable hearing aids must be provided with a charger and case and dryaid kit or capsule and waxguard set if applicable | Yes | Yes | Each
151 | RT274-06-128 | Category 6 | Digital rechargeable CIC for moderate-severe hearing loss (56-70dB), MPO: 116-135dB gain 46-70dB, without Bluetooth compatibility, without tinnitus management, without frequency-lowering. NB: all dB values in the category 6 must comply to the following specifications: Must be according to the 2cc coupler readings, Minimum specification for all Custom hearing aids, Digital signal processing, Technology obsolescence: not older than 18 months from date of publication of tender, Active noise reduction/speech enhancement, Minimum of 4 programs, Minimum of 8 channels, Minimum of IP57 rating, Data logging, Feedback management, Volume control wheel or button for ITE/ITC, Non-rechargeable hearing aids must come with 6 appropriately sized zinc air batteries and dryaid kit and waxguard set if applicable, All re-chargeable hearing aids must be provided with a charger and case and dryaid kit or capsule and waxguard set if applicable | Yes | No | Each
152 | RT274-06-130 | Category 6 | Digital rechargeable CIC for severe-profound hearing loss (71-90dB), MPO: MPO: 126-145dB gain 61-85dB, with Bluetooth compatibility, with tinnitus management, with frequency-lowering. NB: all dB values in the category 6 must comply to the following specifications: Must be according to the 2cc coupler readings, Minimum specification for all Custom hearing aids, Digital signal processing, Technology obsolescence: not older than 18 months from date of publication of tender, Active noise reduction/speech enhancement, Minimum of 4 programs, Minimum of 8 channels, Minimum of IP57 rating, Data logging, Feedback management, Volume control wheel or button for ITE/ITC, Non-rechargeable hearing aids must come with 6 appropriately sized zinc air batteries and dryaid kit and waxguard set if applicable, All re-chargeable hearing aids must be provided with a charger and case and dryaid kit or capsule and waxguard set if applicable | Yes | Yes | Each
153 | RT274-06-132 | Category 6 | Digital rechargeable CIC for severe-profound hearing loss (71-90dB), MPO: MPO: 126-145dB gain 61-85dB, without Bluetooth compatibility, with tinnitus management, with frequency-lowering. NB: all dB values in the category 6 must comply to the following specifications: Must be according to the 2cc coupler readings, Minimum specification for all Custom hearing aids, Digital signal processing, Technology obsolescence: not older than 18 months from date of publication of tender, Active noise reduction/speech enhancement, Minimum of 4 programs, Minimum of 8 channels, Minimum of IP57 rating, Data logging, Feedback management, Volume control wheel or button for ITE/ITC, Non-rechargeable hearing aids must come with 6 appropriately sized zinc air batteries and dryaid kit and waxguard set if applicable, All re-chargeable hearing aids must be provided with a charger and case and dryaid kit or capsule and waxguard set if applicable | Yes | No | Each
154 | RT274-06-134 | Category 6 | Digital rechargeable CIC for severe-profound hearing loss (71-90dB), MPO: MPO: 126-145dB gain 61-85dB, with Bluetooth compatibility, without tinnitus management, with frequency-lowering. NB: all dB values in the category 6 must comply to the following specifications: Must be according to the 2cc coupler readings, Minimum specification for all Custom hearing aids, Digital signal processing, Technology obsolescence: not older than 18 months from date of publication of tender, Active noise reduction/speech enhancement, Minimum of 4 programs, Minimum of 8 channels, Minimum of IP57 rating, Data logging, Feedback management, Volume control wheel or button for ITE/ITC, Non-rechargeable hearing aids must come with 6 appropriately sized zinc air batteries and dryaid kit and waxguard set if applicable, All re-chargeable hearing aids must be provided with a charger and case and dryaid kit or capsule and waxguard set if applicable | Yes | Yes | Each
155 | RT274-06-136 | Category 6 | Digital rechargeable CIC for severe-profound hearing loss (71-90dB), MPO: MPO: 126-145dB gain 61-85dB, without Bluetooth compatibility, without tinnitus management, with frequency-lowering. NB: all dB values in the category 6 must comply to the following specifications: Must be according to the 2cc coupler readings, Minimum specification for all Custom hearing aids, Digital signal processing, Technology obsolescence: not older than 18 months from date of publication of tender, Active noise reduction/speech enhancement, Minimum of 4 programs, Minimum of 8 channels, Minimum of IP57 rating, Data logging, Feedback management, Volume control wheel or button for ITE/ITC, Non-rechargeable hearing aids must come with 6 appropriately sized zinc air batteries and dryaid kit and waxguard set if applicable, All re-chargeable hearing aids must be provided with a charger and case and dryaid kit or capsule and waxguard set if applicable | Yes | No | Each
156 | RT274-06-138 | Category 6 | Digital rechargeable CIC for severe-profound hearing loss (71-90dB), MPO: MPO: 126-145dB gain 61-85dB, with Bluetooth compatibility, with tinnitus management, without frequency-lowering. NB: all dB values in the category 6 must comply to the following specifications: Must be according to the 2cc coupler readings, Minimum specification for all Custom hearing aids, Digital signal processing, Technology obsolescence: not older than 18 months from date of publication of tender, Active noise reduction/speech enhancement, Minimum of 4 programs, Minimum of 8 channels, Minimum of IP57 rating, Data logging, Feedback management, Volume control wheel or button for ITE/ITC, Non-rechargeable hearing aids must come with 6 appropriately sized zinc air batteries and dryaid kit and waxguard set if applicable, All re-chargeable hearing aids must be provided with a charger and case and dryaid kit or capsule and waxguard set if applicable | Yes | Yes | Each
157 | RT274-06-140 | Category 6 | Digital rechargeable CIC for severe-profound hearing loss (71-90dB), MPO: MPO: 126-145dB gain 61-85dB, without Bluetooth compatibility, with tinnitus management, without frequency-lowering. NB: all dB values in the category 6 must comply to the following specifications: Must be according to the 2cc coupler readings, Minimum specification for all Custom hearing aids, Digital signal processing, Technology obsolescence: not older than 18 months from date of publication of tender, Active noise reduction/speech enhancement, Minimum of 4 programs, Minimum of 8 channels, Minimum of IP57 rating, Data logging, Feedback management, Volume control wheel or button for ITE/ITC, Non-rechargeable hearing aids must come with 6 appropriately sized zinc air batteries and dryaid kit and waxguard set if applicable, All re-chargeable hearing aids must be provided with a charger and case and dryaid kit or capsule and waxguard set if applicable | Yes | No | Each
158 | RT274-06-142 | Category 6 | Digital rechargeable CIC for severe-profound hearing loss (71-90dB), MPO: MPO: 126-145dB gain 61-85dB, with Bluetooth compatibility, without tinnitus management, without frequency-lowering. NB: all dB values in the category 6 must comply to the following specifications: Must be according to the 2cc coupler readings, Minimum specification for all Custom hearing aids, Digital signal processing, Technology obsolescence: not older than 18 months from date of publication of tender, Active noise reduction/speech enhancement, Minimum of 4 programs, Minimum of 8 channels, Minimum of IP57 rating, Data logging, Feedback management, Volume control wheel or button for ITE/ITC, Non-rechargeable hearing aids must come with 6 appropriately sized zinc air batteries and dryaid kit and waxguard set if applicable, All re-chargeable hearing aids must be provided with a charger and case and dryaid kit or capsule and waxguard set if applicable | Yes | Yes | Each
159 | RT274-06-144 | Category 6 | Digital rechargeable CIC for severe-profound hearing loss (71-90dB), MPO: MPO: 126-145dB gain 61-85dB, without Bluetooth compatibility, without tinnitus management, without frequency-lowering. NB: all dB values in the category 6 must comply to the following specifications: Must be according to the 2cc coupler readings, Minimum specification for all Custom hearing aids, Digital signal processing, Technology obsolescence: not older than 18 months from date of publication of tender, Active noise reduction/speech enhancement, Minimum of 4 programs, Minimum of 8 channels, Minimum of IP57 rating, Data logging, Feedback management, Volume control wheel or button for ITE/ITC, Non-rechargeable hearing aids must come with 6 appropriately sized zinc air batteries and dryaid kit and waxguard set if applicable, All re-chargeable hearing aids must be provided with a charger and case and dryaid kit or capsule and waxguard set if applicable | Yes | No | Each
160 | RT274-06-147 | Category 6 | Wax guards, product specific, small plastic attachment for in-the-ear, in-the-canal and complete-in-canal hearing instruments for the prevention of wax blockages. Compatible with all awarded RIC, ITE, ITC and CIC hearing aids on RT274-2023 and previous RT274 tenders. Multiple award. NB: all dB values in the category 6 must comply to the following specifications: Must be according to the 2cc coupler readings, Minimum specification for all Custom hearing aids, Digital signal processing, Technology obsolescence: not older than 18 months from date of publication of tender, Active noise reduction/speech enhancement, Minimum of 4 programs, Minimum of 8 channels, Minimum of IP57 rating, Data logging, Feedback management, Volume control wheel or button for ITE/ITC, Non-rechargeable hearing aids must come with 6 appropriately sized zinc air batteries and dryaid kit and waxguard set if applicable, All re-chargeable hearing aids must be provided with a charger and case and dryaid kit or capsule and waxguard set if applicable | Yes | No | Each
161 | RT274-07-001 | Category 7 | Batteries, dry cell, size AA, minimum card of 4-6 | No | No | Pack of 5
162 | RT274-07-002 | Category 7 | Batteries, dry cell, size AAA, minimum card of 4- 6 | No | No | Pack of 5
163 | RT274-07-003 | Category 7 | Alkaline 9V battery, IEC type 6lr61, flat type, 9V-500mal | No | No | Each
164 | RT274-07-008 | Category 7 | Rechargeable lithium-ion batteries size 13, for hearing aid | No | No | Each
165 | RT274-07-037 | Category 7 | Sachet of silica gel, with saturation indicator on sachet, supplier to specify quantity per pack and unit price | No | No | Pack of 10
166 | RT274-07-055 | Category 7 | Earmould impression starter kit consisting of a mimimum of 1 pck of small otostops, 1 pck of medium otostops, 1 pck of large otostops, 1 single plunge syringe, 1 set of silicon (including catalyst) based impression material, 1 otolight to be made available in portable hard case/ storage box carrier | Yes | No | Each
167 | RT274-07-046 | Category 7 | Syringe, 10-15cm in length, double plunge, suitable for making earmould impressions | Yes | No | Each
168 | RT274-07-056 | Category 7 | Replacement mini receiver for RIC. Multiple Awarded | Yes | No | Each
169 | RT274-07-057 | Category 7 | Replacement mini sleeves for RIC. Multiple Awarded | Yes | No | Each
170 | RT274-07-058 | Category 7 | Replacement mini ear tips for RIC. Multiple Awarded | Yes | No | Each
171 | RT274-07-059 | Category 7 | Otoscope, complete with 12 speculae and compatible batteries complete with hard casing | Yes | No | Each
172 | RT274-07-060 | Category 7 | Replacement lightbulb for otoscope, Multiple award; to be compatible with 42182005-00008, as well as the otoscopes awarded in the 2017 and 2014 RT274 contracts | Yes | No | Each
173 | RT274-09-001 | Category 9 | CROS BTE, Digital rechargeable, wireless, behind the ear hearing aid, cross fitting, for contra lateral routing of signal, used to fit unilateral hearing loss. Complete with options of earmould, slim tube and dome, dry aid kit, NB: All dB values for BTE Cros and Bicros in Category 9 specifications must comply to the following requirements:Must be according to the 2cc coupler readings, All BTE hearing aids must contain the following as minimum specifications: Technology obsolescence: not older than 18 months from date of publication of tender, Active noise reduction/speech enhancement, Active noise reduction/speech enhancement, Minimum of 4 programs, Minimum of 8 channels, Minimum of IP57 rating, Programmable directionality, Data logging, Feedback management, Volume control, Telecoil, All hearing aids to include an ear mould (acrylic or silicone) or 2 slim tubes and domes, All non-rechargeable hearing aids must include 6 appropriately sized zinc air batteries and dry aid kit, All re-chargeable hearing aids must be provided with a charger and case and dry aid kit and - All dB values for RIC Cros and Bicros in tcategory 9 must comply to the following specifications: Must be according to the 2cc coupler readings, Minimum specification for all RIC hearing aids, Digital signal processing, Technology obsolescence: not older than 18 months from date of publication of tender, Active noise reduction/speech enhancement, Minimum of 4 programs, Minimum of 8 channels, Minimum of IP57 rating, Programmable directionality, Data logging, Feedback management, Volume control, Telecoil, All hearing aids to come with a single receiver and dome, or flex tips, or micro mould or a RIC mould for Power RIC, wax guard. Non-rechargeable hearing aids must come with 6 appropriately sized zinc air batteries. All re-chargeable hearing aids must be provided with a charger and case | Yes | Yes | Each
174 | RT274-09-003 | Category 9 | CROS RIC, Digital rechargeable, wireless, receiver in the ear hearing aid, cross fitting, for contra lateral routing of signal, used to fit unilateral hearing loss. Complete with RIC receiver, option of flex tips or micro mould, slim tube and dome, dry aid kit, wax guard NB: All dB values for BTE Cros and Bicros in Category 9 specifications must comply to the following requirements:Must be according to the 2cc coupler readings, All BTE hearing aids must contain the following as minimum specifications: Technology obsolescence: not older than 18 months from date of publication of tender, Active noise reduction/speech enhancement, Active noise reduction/speech enhancement, Minimum of 4 programs, Minimum of 8 channels, Minimum of IP57 rating, Programmable directionality, Data logging, Feedback management, Volume control, Telecoil, All hearing aids to include an ear mould (acrylic or silicone) or 2 slim tubes and domes, All non-rechargeable hearing aids must include 6 appropriately sized zinc air batteries and dry aid kit, All re-chargeable hearing aids must be provided with a charger and case and dry aid kit and - All dB values for RIC Cros and Bicros in tcategory 9 must comply to the following specifications: Must be according to the 2cc coupler readings, Minimum specification for all RIC hearing aids, Digital signal processing, Technology obsolescence: not older than 18 months from date of publication of tender, Active noise reduction/speech enhancement, Minimum of 4 programs, Minimum of 8 channels, Minimum of IP57 rating, Programmable directionality, Data logging, Feedback management, Volume control, Telecoil, All hearing aids to come with a single receiver and dome, or flex tips, or micro mould or a RIC mould for Power RIC, wax guard. Non-rechargeable hearing aids must come with 6 appropriately sized zinc air batteries. All re-chargeable hearing aids must be provided with a charger and case | Yes | Yes | Each
175 | RT274-09-005 | Category 9 | Digital rechargeable BTE for bi- cross fitting, for contra lateral routing of signal, used to fit asymmetric hearing loss, where the one ear is dead and the other ear has a mild-moderate hearing loss (25-55dB), MPO: 75-120dB gain 40-60dB, with Bluetooth compatibility, with tinnitus management, with frequency-lowering. NB: All dB values for BTE Cros and Bicros in Category 9 specifications must comply to the following requirements:Must be according to the 2cc coupler readings, All BTE hearing aids must contain the following as minimum specifications: Technology obsolescence: not older than 18 months from date of publication of tender, Active noise reduction/speech enhancement, Active noise reduction/speech enhancement, Minimum of 4 programs, Minimum of 8 channels, Minimum of IP57 rating, Programmable directionality, Data logging, Feedback management, Volume control, Telecoil, All hearing aids to include an ear mould (acrylic or silicone) or 2 slim tubes and domes, All non-rechargeable hearing aids must include 6 appropriately sized zinc air batteries and dry aid kit, All re-chargeable hearing aids must be provided with a charger and case and dry aid kit and - All dB values for RIC Cros and Bicros in tcategory 9 must comply to the following specifications: Must be according to the 2cc coupler readings, Minimum specification for all RIC hearing aids, Digital signal processing, Technology obsolescence: not older than 18 months from date of publication of tender, Active noise reduction/speech enhancement, Minimum of 4 programs, Minimum of 8 channels, Minimum of IP57 rating, Programmable directionality, Data logging, Feedback management, Volume control, Telecoil, All hearing aids to come with a single receiver and dome, or flex tips, or micro mould or a RIC mould for Power RIC, wax guard. Non-rechargeable hearing aids must come with 6 appropriately sized zinc air batteries. All re-chargeable hearing aids must be provided with a charger and case | Yes | Yes | Each
176 | RT274-09-007 | Category 9 | Digital rechargeable BTE for bi- cross fitting, for contra lateral routing of signal, used to fit asymmetric hearing loss, where the one ear is dead and the other ear has a mild-moderate hearing loss (25-55dB), MPO: 75-120dB gain 40-60dB, without Bluetooth compatibility, with tinnitus management, with frequency-lowering. NB: All dB values for BTE Cros and Bicros in Category 9 specifications must comply to the following requirements:Must be according to the 2cc coupler readings, All BTE hearing aids must contain the following as minimum specifications: Technology obsolescence: not older than 18 months from date of publication of tender, Active noise reduction/speech enhancement, Active noise reduction/speech enhancement, Minimum of 4 programs, Minimum of 8 channels, Minimum of IP57 rating, Programmable directionality, Data logging, Feedback management, Volume control, Telecoil, All hearing aids to include an ear mould (acrylic or silicone) or 2 slim tubes and domes, All non-rechargeable hearing aids must include 6 appropriately sized zinc air batteries and dry aid kit, All re-chargeable hearing aids must be provided with a charger and case and dry aid kit and - All dB values for RIC Cros and Bicros in tcategory 9 must comply to the following specifications: Must be according to the 2cc coupler readings, Minimum specification for all RIC hearing aids, Digital signal processing, Technology obsolescence: not older than 18 months from date of publication of tender, Active noise reduction/speech enhancement, Minimum of 4 programs, Minimum of 8 channels, Minimum of IP57 rating, Programmable directionality, Data logging, Feedback management, Volume control, Telecoil, All hearing aids to come with a single receiver and dome, or flex tips, or micro mould or a RIC mould for Power RIC, wax guard. Non-rechargeable hearing aids must come with 6 appropriately sized zinc air batteries. All re-chargeable hearing aids must be provided with a charger and case | Yes | No | Each
177 | RT274-09-009 | Category 9 | Digital rechargeable BTE for bi- cross fitting, for contra lateral routing of signal, used to fit asymmetric hearing loss, where the one ear is dead and the other ear has a mild-moderate hearing loss (25-55dB), MPO: 75-120dB gain 40-60dB, with Bluetooth compatibility, without tinnitus management, with frequency-lowering. NB: All dB values for BTE Cros and Bicros in Category 9 specifications must comply to the following requirements:Must be according to the 2cc coupler readings, All BTE hearing aids must contain the following as minimum specifications: Technology obsolescence: not older than 18 months from date of publication of tender, Active noise reduction/speech enhancement, Active noise reduction/speech enhancement, Minimum of 4 programs, Minimum of 8 channels, Minimum of IP57 rating, Programmable directionality, Data logging, Feedback management, Volume control, Telecoil, All hearing aids to include an ear mould (acrylic or silicone) or 2 slim tubes and domes, All non-rechargeable hearing aids must include 6 appropriately sized zinc air batteries and dry aid kit, All re-chargeable hearing aids must be provided with a charger and case and dry aid kit and - All dB values for RIC Cros and Bicros in tcategory 9 must comply to the following specifications: Must be according to the 2cc coupler readings, Minimum specification for all RIC hearing aids, Digital signal processing, Technology obsolescence: not older than 18 months from date of publication of tender, Active noise reduction/speech enhancement, Minimum of 4 programs, Minimum of 8 channels, Minimum of IP57 rating, Programmable directionality, Data logging, Feedback management, Volume control, Telecoil, All hearing aids to come with a single receiver and dome, or flex tips, or micro mould or a RIC mould for Power RIC, wax guard. Non-rechargeable hearing aids must come with 6 appropriately sized zinc air batteries. All re-chargeable hearing aids must be provided with a charger and case | Yes | Yes | Each
178 | RT274-09-011 | Category 9 | Digital rechargeable BTE for bi- cross fitting, for contra lateral routing of signal, used to fit asymmetric hearing loss, where the one ear is dead and the other ear has a mild-moderate hearing loss (25-55dB), MPO: 75-120dB gain 40-60dB, without Bluetooth compatibility, without tinnitus management, with frequency-lowering. NB: All dB values for BTE Cros and Bicros in Category 9 specifications must comply to the following requirements:Must be according to the 2cc coupler readings, All BTE hearing aids must contain the following as minimum specifications: Technology obsolescence: not older than 18 months from date of publication of tender, Active noise reduction/speech enhancement, Active noise reduction/speech enhancement, Minimum of 4 programs, Minimum of 8 channels, Minimum of IP57 rating, Programmable directionality, Data logging, Feedback management, Volume control, Telecoil, All hearing aids to include an ear mould (acrylic or silicone) or 2 slim tubes and domes, All non-rechargeable hearing aids must include 6 appropriately sized zinc air batteries and dry aid kit, All re-chargeable hearing aids must be provided with a charger and case and dry aid kit and - All dB values for RIC Cros and Bicros in tcategory 9 must comply to the following specifications: Must be according to the 2cc coupler readings, Minimum specification for all RIC hearing aids, Digital signal processing, Technology obsolescence: not older than 18 months from date of publication of tender, Active noise reduction/speech enhancement, Minimum of 4 programs, Minimum of 8 channels, Minimum of IP57 rating, Programmable directionality, Data logging, Feedback management, Volume control, Telecoil, All hearing aids to come with a single receiver and dome, or flex tips, or micro mould or a RIC mould for Power RIC, wax guard. Non-rechargeable hearing aids must come with 6 appropriately sized zinc air batteries. All re-chargeable hearing aids must be provided with a charger and case | Yes | No | Each
179 | RT274-09-013 | Category 9 | Digital rechargeable BTE for bi- cross fitting, for contra lateral routing of signal, used to fit asymmetric hearing loss, where the one ear is dead and the other ear has a mild-moderate hearing loss (25-55dB), MPO: 75-120dB gain 40-60dB, with Bluetooth compatibility, with tinnitus management, without frequency-lowering. NB: All dB values for BTE Cros and Bicros in Category 9 specifications must comply to the following requirements:Must be according to the 2cc coupler readings, All BTE hearing aids must contain the following as minimum specifications: Technology obsolescence: not older than 18 months from date of publication of tender, Active noise reduction/speech enhancement, Active noise reduction/speech enhancement, Minimum of 4 programs, Minimum of 8 channels, Minimum of IP57 rating, Programmable directionality, Data logging, Feedback management, Volume control, Telecoil, All hearing aids to include an ear mould (acrylic or silicone) or 2 slim tubes and domes, All non-rechargeable hearing aids must include 6 appropriately sized zinc air batteries and dry aid kit, All re-chargeable hearing aids must be provided with a charger and case and dry aid kit and - All dB values for RIC Cros and Bicros in tcategory 9 must comply to the following specifications: Must be according to the 2cc coupler readings, Minimum specification for all RIC hearing aids, Digital signal processing, Technology obsolescence: not older than 18 months from date of publication of tender, Active noise reduction/speech enhancement, Minimum of 4 programs, Minimum of 8 channels, Minimum of IP57 rating, Programmable directionality, Data logging, Feedback management, Volume control, Telecoil, All hearing aids to come with a single receiver and dome, or flex tips, or micro mould or a RIC mould for Power RIC, wax guard. Non-rechargeable hearing aids must come with 6 appropriately sized zinc air batteries. All re-chargeable hearing aids must be provided with a charger and case | Yes | Yes | Each
180 | RT274-09-015 | Category 9 | Digital rechargeable BTE for bi- cross fitting, for contra lateral routing of signal, used to fit asymmetric hearing loss, where the one ear is dead and the other ear has a mild-moderate hearing loss (25-55dB), MPO: 75-120dB gain 40-60dB, without Bluetooth compatibility, with tinnitus management, without frequency-lowering. NB: All dB values for BTE Cros and Bicros in Category 9 specifications must comply to the following requirements:Must be according to the 2cc coupler readings, All BTE hearing aids must contain the following as minimum specifications: Technology obsolescence: not older than 18 months from date of publication of tender, Active noise reduction/speech enhancement, Active noise reduction/speech enhancement, Minimum of 4 programs, Minimum of 8 channels, Minimum of IP57 rating, Programmable directionality, Data logging, Feedback management, Volume control, Telecoil, All hearing aids to include an ear mould (acrylic or silicone) or 2 slim tubes and domes, All non-rechargeable hearing aids must include 6 appropriately sized zinc air batteries and dry aid kit, All re-chargeable hearing aids must be provided with a charger and case and dry aid kit and - All dB values for RIC Cros and Bicros in tcategory 9 must comply to the following specifications: Must be according to the 2cc coupler readings, Minimum specification for all RIC hearing aids, Digital signal processing, Technology obsolescence: not older than 18 months from date of publication of tender, Active noise reduction/speech enhancement, Minimum of 4 programs, Minimum of 8 channels, Minimum of IP57 rating, Programmable directionality, Data logging, Feedback management, Volume control, Telecoil, All hearing aids to come with a single receiver and dome, or flex tips, or micro mould or a RIC mould for Power RIC, wax guard. Non-rechargeable hearing aids must come with 6 appropriately sized zinc air batteries. All re-chargeable hearing aids must be provided with a charger and case | Yes | Yes | Each
181 | RT274-09-017 | Category 9 | Digital rechargeable BTE for bi- cross fitting, for contra lateral routing of signal, used to fit asymmetric hearing loss, where the one ear is dead and the other ear has a mild-moderate hearing loss (25-55dB), MPO: 75-120dB gain 40-60dB, with Bluetooth compatibility, without tinnitus management, without frequency-lowering. NB: All dB values for BTE Cros and Bicros in Category 9 specifications must comply to the following requirements:Must be according to the 2cc coupler readings, All BTE hearing aids must contain the following as minimum specifications: Technology obsolescence: not older than 18 months from date of publication of tender, Active noise reduction/speech enhancement, Active noise reduction/speech enhancement, Minimum of 4 programs, Minimum of 8 channels, Minimum of IP57 rating, Programmable directionality, Data logging, Feedback management, Volume control, Telecoil, All hearing aids to include an ear mould (acrylic or silicone) or 2 slim tubes and domes, All non-rechargeable hearing aids must include 6 appropriately sized zinc air batteries and dry aid kit, All re-chargeable hearing aids must be provided with a charger and case and dry aid kit and - All dB values for RIC Cros and Bicros in tcategory 9 must comply to the following specifications: Must be according to the 2cc coupler readings, Minimum specification for all RIC hearing aids, Digital signal processing, Technology obsolescence: not older than 18 months from date of publication of tender, Active noise reduction/speech enhancement, Minimum of 4 programs, Minimum of 8 channels, Minimum of IP57 rating, Programmable directionality, Data logging, Feedback management, Volume control, Telecoil, All hearing aids to come with a single receiver and dome, or flex tips, or micro mould or a RIC mould for Power RIC, wax guard. Non-rechargeable hearing aids must come with 6 appropriately sized zinc air batteries. All re-chargeable hearing aids must be provided with a charger and case | Yes | Yes | Each
182 | RT274-09-019 | Category 9 | Digital rechargeable BTE for bi- cross fitting, for contra lateral routing of signal, used to fit asymmetric hearing loss, where the one ear is dead and the other ear has a mild-moderate hearing loss (25-55dB), MPO: 75-120dB gain 40-60dB, without Bluetooth compatibility, without tinnitus management, without frequency-lowering. NB: All dB values for BTE Cros and Bicros in Category 9 specifications must comply to the following requirements:Must be according to the 2cc coupler readings, All BTE hearing aids must contain the following as minimum specifications: Technology obsolescence: not older than 18 months from date of publication of tender, Active noise reduction/speech enhancement, Active noise reduction/speech enhancement, Minimum of 4 programs, Minimum of 8 channels, Minimum of IP57 rating, Programmable directionality, Data logging, Feedback management, Volume control, Telecoil, All hearing aids to include an ear mould (acrylic or silicone) or 2 slim tubes and domes, All non-rechargeable hearing aids must include 6 appropriately sized zinc air batteries and dry aid kit, All re-chargeable hearing aids must be provided with a charger and case and dry aid kit and - All dB values for RIC Cros and Bicros in tcategory 9 must comply to the following specifications: Must be according to the 2cc coupler readings, Minimum specification for all RIC hearing aids, Digital signal processing, Technology obsolescence: not older than 18 months from date of publication of tender, Active noise reduction/speech enhancement, Minimum of 4 programs, Minimum of 8 channels, Minimum of IP57 rating, Programmable directionality, Data logging, Feedback management, Volume control, Telecoil, All hearing aids to come with a single receiver and dome, or flex tips, or micro mould or a RIC mould for Power RIC, wax guard. Non-rechargeable hearing aids must come with 6 appropriately sized zinc air batteries. All re-chargeable hearing aids must be provided with a charger and case | Yes | No | Each
183 | RT274-09-021 | Category 9 | Digital rechargeable BTE for bi- cross fitting, for contra lateral routing of signal, used to fit asymmetric hearing loss, where the one ear is dead and the other ear has a moderate-severe hearing loss (56-70dB), MPO: 116-130dB gain 46-65dB, with Bluetooth compatibility, with tinnitus management, with frequency-lowering. NB: All dB values for BTE Cros and Bicros in Category 9 specifications must comply to the following requirements:Must be according to the 2cc coupler readings, All BTE hearing aids must contain the following as minimum specifications: Technology obsolescence: not older than 18 months from date of publication of tender, Active noise reduction/speech enhancement, Active noise reduction/speech enhancement, Minimum of 4 programs, Minimum of 8 channels, Minimum of IP57 rating, Programmable directionality, Data logging, Feedback management, Volume control, Telecoil, All hearing aids to include an ear mould (acrylic or silicone) or 2 slim tubes and domes, All non-rechargeable hearing aids must include 6 appropriately sized zinc air batteries and dry aid kit, All re-chargeable hearing aids must be provided with a charger and case and dry aid kit and - All dB values for RIC Cros and Bicros in tcategory 9 must comply to the following specifications: Must be according to the 2cc coupler readings, Minimum specification for all RIC hearing aids, Digital signal processing, Technology obsolescence: not older than 18 months from date of publication of tender, Active noise reduction/speech enhancement, Minimum of 4 programs, Minimum of 8 channels, Minimum of IP57 rating, Programmable directionality, Data logging, Feedback management, Volume control, Telecoil, All hearing aids to come with a single receiver and dome, or flex tips, or micro mould or a RIC mould for Power RIC, wax guard. Non-rechargeable hearing aids must come with 6 appropriately sized zinc air batteries. All re-chargeable hearing aids must be provided with a charger and case | Yes | Yes | Each
184 | RT274-09-023 | Category 9 | Digital rechargeable BTE for bi- cross fitting, for contra lateral routing of signal, used to fit asymmetric hearing loss, where the one ear is dead and the other ear has a moderate-severe hearing loss (56-70dB), MPO: 116-130dB gain 46-65dB, without Bluetooth compatibility, with tinnitus management, with frequency-lowering. NB: All dB values for BTE Cros and Bicros in Category 9 specifications must comply to the following requirements:Must be according to the 2cc coupler readings, All BTE hearing aids must contain the following as minimum specifications: Technology obsolescence: not older than 18 months from date of publication of tender, Active noise reduction/speech enhancement, Active noise reduction/speech enhancement, Minimum of 4 programs, Minimum of 8 channels, Minimum of IP57 rating, Programmable directionality, Data logging, Feedback management, Volume control, Telecoil, All hearing aids to include an ear mould (acrylic or silicone) or 2 slim tubes and domes, All non-rechargeable hearing aids must include 6 appropriately sized zinc air batteries and dry aid kit, All re-chargeable hearing aids must be provided with a charger and case and dry aid kit and - All dB values for RIC Cros and Bicros in tcategory 9 must comply to the following specifications: Must be according to the 2cc coupler readings, Minimum specification for all RIC hearing aids, Digital signal processing, Technology obsolescence: not older than 18 months from date of publication of tender, Active noise reduction/speech enhancement, Minimum of 4 programs, Minimum of 8 channels, Minimum of IP57 rating, Programmable directionality, Data logging, Feedback management, Volume control, Telecoil, All hearing aids to come with a single receiver and dome, or flex tips, or micro mould or a RIC mould for Power RIC, wax guard. Non-rechargeable hearing aids must come with 6 appropriately sized zinc air batteries. All re-chargeable hearing aids must be provided with a charger and case | Yes | No | Each
185 | RT274-09-025 | Category 9 | Digital rechargeable BTE for bi- cross fitting, for contra lateral routing of signal, used to fit asymmetric hearing loss, where the one ear is dead and the other ear has a moderate-severe hearing loss (56-70dB), MPO: 116-130dB gain 46-65dB, with Bluetooth compatibility, without tinnitus management, with frequency-lowering. NB: All dB values for BTE Cros and Bicros in Category 9 specifications must comply to the following requirements:Must be according to the 2cc coupler readings, All BTE hearing aids must contain the following as minimum specifications: Technology obsolescence: not older than 18 months from date of publication of tender, Active noise reduction/speech enhancement, Active noise reduction/speech enhancement, Minimum of 4 programs, Minimum of 8 channels, Minimum of IP57 rating, Programmable directionality, Data logging, Feedback management, Volume control, Telecoil, All hearing aids to include an ear mould (acrylic or silicone) or 2 slim tubes and domes, All non-rechargeable hearing aids must include 6 appropriately sized zinc air batteries and dry aid kit, All re-chargeable hearing aids must be provided with a charger and case and dry aid kit and - All dB values for RIC Cros and Bicros in tcategory 9 must comply to the following specifications: Must be according to the 2cc coupler readings, Minimum specification for all RIC hearing aids, Digital signal processing, Technology obsolescence: not older than 18 months from date of publication of tender, Active noise reduction/speech enhancement, Minimum of 4 programs, Minimum of 8 channels, Minimum of IP57 rating, Programmable directionality, Data logging, Feedback management, Volume control, Telecoil, All hearing aids to come with a single receiver and dome, or flex tips, or micro mould or a RIC mould for Power RIC, wax guard. Non-rechargeable hearing aids must come with 6 appropriately sized zinc air batteries. All re-chargeable hearing aids must be provided with a charger and case | Yes | Yes | Each
186 | RT274-09-027 | Category 9 | Digital rechargeable BTE for bi- cross fitting, for contra lateral routing of signal, used to fit asymmetric hearing loss, where the one ear is dead and the other ear has a moderate-severe hearing loss (56-70dB), MPO: 116-130dB gain 46-65dB, without Bluetooth compatibility, without tinnitus management, with frequency-lowering. NB: All dB values for BTE Cros and Bicros in Category 9 specifications must comply to the following requirements:Must be according to the 2cc coupler readings, All BTE hearing aids must contain the following as minimum specifications: Technology obsolescence: not older than 18 months from date of publication of tender, Active noise reduction/speech enhancement, Active noise reduction/speech enhancement, Minimum of 4 programs, Minimum of 8 channels, Minimum of IP57 rating, Programmable directionality, Data logging, Feedback management, Volume control, Telecoil, All hearing aids to include an ear mould (acrylic or silicone) or 2 slim tubes and domes, All non-rechargeable hearing aids must include 6 appropriately sized zinc air batteries and dry aid kit, All re-chargeable hearing aids must be provided with a charger and case and dry aid kit and - All dB values for RIC Cros and Bicros in tcategory 9 must comply to the following specifications: Must be according to the 2cc coupler readings, Minimum specification for all RIC hearing aids, Digital signal processing, Technology obsolescence: not older than 18 months from date of publication of tender, Active noise reduction/speech enhancement, Minimum of 4 programs, Minimum of 8 channels, Minimum of IP57 rating, Programmable directionality, Data logging, Feedback management, Volume control, Telecoil, All hearing aids to come with a single receiver and dome, or flex tips, or micro mould or a RIC mould for Power RIC, wax guard. Non-rechargeable hearing aids must come with 6 appropriately sized zinc air batteries. All re-chargeable hearing aids must be provided with a charger and case | Yes | No | Each
187 | RT274-09-029 | Category 9 | Digital rechargeable BTE for bi- cross fitting, for contra lateral routing of signal, used to fit asymmetric hearing loss, where the one ear is dead and the other ear has a moderate-severe hearing loss (56-70dB), MPO: 116-130dB gain 46-65dB, with Bluetooth compatibility, with tinnitus management, without frequency-lowering. NB: All dB values for BTE Cros and Bicros in Category 9 specifications must comply to the following requirements:Must be according to the 2cc coupler readings, All BTE hearing aids must contain the following as minimum specifications: Technology obsolescence: not older than 18 months from date of publication of tender, Active noise reduction/speech enhancement, Active noise reduction/speech enhancement, Minimum of 4 programs, Minimum of 8 channels, Minimum of IP57 rating, Programmable directionality, Data logging, Feedback management, Volume control, Telecoil, All hearing aids to include an ear mould (acrylic or silicone) or 2 slim tubes and domes, All non-rechargeable hearing aids must include 6 appropriately sized zinc air batteries and dry aid kit, All re-chargeable hearing aids must be provided with a charger and case and dry aid kit and - All dB values for RIC Cros and Bicros in tcategory 9 must comply to the following specifications: Must be according to the 2cc coupler readings, Minimum specification for all RIC hearing aids, Digital signal processing, Technology obsolescence: not older than 18 months from date of publication of tender, Active noise reduction/speech enhancement, Minimum of 4 programs, Minimum of 8 channels, Minimum of IP57 rating, Programmable directionality, Data logging, Feedback management, Volume control, Telecoil, All hearing aids to come with a single receiver and dome, or flex tips, or micro mould or a RIC mould for Power RIC, wax guard. Non-rechargeable hearing aids must come with 6 appropriately sized zinc air batteries. All re-chargeable hearing aids must be provided with a charger and case | Yes | Yes | Each
188 | RT274-09-031 | Category 9 | Digital rechargeable BTE for bi- cross fitting, for contra lateral routing of signal, used to fit asymmetric hearing loss, where the one ear is dead and the other ear has a moderate-severe hearing loss (56-70dB), MPO: 116-130dB gain 46-65dB, without Bluetooth compatibility, with tinnitus management, without frequency-lowering. NB: All dB values for BTE Cros and Bicros in Category 9 specifications must comply to the following requirements:Must be according to the 2cc coupler readings, All BTE hearing aids must contain the following as minimum specifications: Technology obsolescence: not older than 18 months from date of publication of tender, Active noise reduction/speech enhancement, Active noise reduction/speech enhancement, Minimum of 4 programs, Minimum of 8 channels, Minimum of IP57 rating, Programmable directionality, Data logging, Feedback management, Volume control, Telecoil, All hearing aids to include an ear mould (acrylic or silicone) or 2 slim tubes and domes, All non-rechargeable hearing aids must include 6 appropriately sized zinc air batteries and dry aid kit, All re-chargeable hearing aids must be provided with a charger and case and dry aid kit and - All dB values for RIC Cros and Bicros in tcategory 9 must comply to the following specifications: Must be according to the 2cc coupler readings, Minimum specification for all RIC hearing aids, Digital signal processing, Technology obsolescence: not older than 18 months from date of publication of tender, Active noise reduction/speech enhancement, Minimum of 4 programs, Minimum of 8 channels, Minimum of IP57 rating, Programmable directionality, Data logging, Feedback management, Volume control, Telecoil, All hearing aids to come with a single receiver and dome, or flex tips, or micro mould or a RIC mould for Power RIC, wax guard. Non-rechargeable hearing aids must come with 6 appropriately sized zinc air batteries. All re-chargeable hearing aids must be provided with a charger and case | Yes | No | Each
189 | RT274-09-033 | Category 9 | Digital rechargeable BTE for bi- cross fitting, for contra lateral routing of signal, used to fit asymmetric hearing loss, where the one ear is dead and the other ear has a moderate-severe hearing loss (56-70dB), MPO: 116-130dB gain 46-65dB, with Bluetooth compatibility, without tinnitus management, without frequency-lowering. NB: All dB values for BTE Cros and Bicros in Category 9 specifications must comply to the following requirements:Must be according to the 2cc coupler readings, All BTE hearing aids must contain the following as minimum specifications: Technology obsolescence: not older than 18 months from date of publication of tender, Active noise reduction/speech enhancement, Active noise reduction/speech enhancement, Minimum of 4 programs, Minimum of 8 channels, Minimum of IP57 rating, Programmable directionality, Data logging, Feedback management, Volume control, Telecoil, All hearing aids to include an ear mould (acrylic or silicone) or 2 slim tubes and domes, All non-rechargeable hearing aids must include 6 appropriately sized zinc air batteries and dry aid kit, All re-chargeable hearing aids must be provided with a charger and case and dry aid kit and - All dB values for RIC Cros and Bicros in tcategory 9 must comply to the following specifications: Must be according to the 2cc coupler readings, Minimum specification for all RIC hearing aids, Digital signal processing, Technology obsolescence: not older than 18 months from date of publication of tender, Active noise reduction/speech enhancement, Minimum of 4 programs, Minimum of 8 channels, Minimum of IP57 rating, Programmable directionality, Data logging, Feedback management, Volume control, Telecoil, All hearing aids to come with a single receiver and dome, or flex tips, or micro mould or a RIC mould for Power RIC, wax guard. Non-rechargeable hearing aids must come with 6 appropriately sized zinc air batteries. All re-chargeable hearing aids must be provided with a charger and case | Yes | Yes | Each
190 | RT274-09-035 | Category 9 | Digital rechargeable BTE for bi- cross fitting, for contra lateral routing of signal, used to fit asymmetric hearing loss, where the one ear is dead and the other ear has a moderate-severe hearing loss (56-70dB), MPO: 116-130dB gain 46-65dB, without Bluetooth compatibility, without tinnitus management, without frequency-lowering. NB: All dB values for BTE Cros and Bicros in Category 9 specifications must comply to the following requirements:Must be according to the 2cc coupler readings, All BTE hearing aids must contain the following as minimum specifications: Technology obsolescence: not older than 18 months from date of publication of tender, Active noise reduction/speech enhancement, Active noise reduction/speech enhancement, Minimum of 4 programs, Minimum of 8 channels, Minimum of IP57 rating, Programmable directionality, Data logging, Feedback management, Volume control, Telecoil, All hearing aids to include an ear mould (acrylic or silicone) or 2 slim tubes and domes, All non-rechargeable hearing aids must include 6 appropriately sized zinc air batteries and dry aid kit, All re-chargeable hearing aids must be provided with a charger and case and dry aid kit and - All dB values for RIC Cros and Bicros in tcategory 9 must comply to the following specifications: Must be according to the 2cc coupler readings, Minimum specification for all RIC hearing aids, Digital signal processing, Technology obsolescence: not older than 18 months from date of publication of tender, Active noise reduction/speech enhancement, Minimum of 4 programs, Minimum of 8 channels, Minimum of IP57 rating, Programmable directionality, Data logging, Feedback management, Volume control, Telecoil, All hearing aids to come with a single receiver and dome, or flex tips, or micro mould or a RIC mould for Power RIC, wax guard. Non-rechargeable hearing aids must come with 6 appropriately sized zinc air batteries. All re-chargeable hearing aids must be provided with a charger and case | Yes | No | Each
191 | RT274-09-037 | Category 9 | Digital rechargeable BTE for bi- cross fitting, for contra lateral routing of signal, used to fit asymmetric hearing loss, where the one ear is dead and the other ear has a severe-profound hearing loss (71-100dB), MPO: MPO: 120-145dB gain 60-85dB, with Bluetooth compatibility, with tinnitus management, with frequency-lowering. NB: All dB values for BTE Cros and Bicros in Category 9 specifications must comply to the following requirements:Must be according to the 2cc coupler readings, All BTE hearing aids must contain the following as minimum specifications: Technology obsolescence: not older than 18 months from date of publication of tender, Active noise reduction/speech enhancement, Active noise reduction/speech enhancement, Minimum of 4 programs, Minimum of 8 channels, Minimum of IP57 rating, Programmable directionality, Data logging, Feedback management, Volume control, Telecoil, All hearing aids to include an ear mould (acrylic or silicone) or 2 slim tubes and domes, All non-rechargeable hearing aids must include 6 appropriately sized zinc air batteries and dry aid kit, All re-chargeable hearing aids must be provided with a charger and case and dry aid kit and - All dB values for RIC Cros and Bicros in tcategory 9 must comply to the following specifications: Must be according to the 2cc coupler readings, Minimum specification for all RIC hearing aids, Digital signal processing, Technology obsolescence: not older than 18 months from date of publication of tender, Active noise reduction/speech enhancement, Minimum of 4 programs, Minimum of 8 channels, Minimum of IP57 rating, Programmable directionality, Data logging, Feedback management, Volume control, Telecoil, All hearing aids to come with a single receiver and dome, or flex tips, or micro mould or a RIC mould for Power RIC, wax guard. Non-rechargeable hearing aids must come with 6 appropriately sized zinc air batteries. All re-chargeable hearing aids must be provided with a charger and case | Yes | Yes | Each
192 | RT274-09-039 | Category 9 | Digital rechargeable BTE for bi- cross fitting, for contra lateral routing of signal, used to fit asymmetric hearing loss, where the one ear is dead and the other ear has a severe-profound hearing loss (71-100dB), MPO: MPO: 120-145dB gain 60-85dB, without Bluetooth compatibility, with tinnitus management, with frequency-lowering. NB: All dB values for BTE Cros and Bicros in Category 9 specifications must comply to the following requirements:Must be according to the 2cc coupler readings, All BTE hearing aids must contain the following as minimum specifications: Technology obsolescence: not older than 18 months from date of publication of tender, Active noise reduction/speech enhancement, Active noise reduction/speech enhancement, Minimum of 4 programs, Minimum of 8 channels, Minimum of IP57 rating, Programmable directionality, Data logging, Feedback management, Volume control, Telecoil, All hearing aids to include an ear mould (acrylic or silicone) or 2 slim tubes and domes, All non-rechargeable hearing aids must include 6 appropriately sized zinc air batteries and dry aid kit, All re-chargeable hearing aids must be provided with a charger and case and dry aid kit and - All dB values for RIC Cros and Bicros in tcategory 9 must comply to the following specifications: Must be according to the 2cc coupler readings, Minimum specification for all RIC hearing aids, Digital signal processing, Technology obsolescence: not older than 18 months from date of publication of tender, Active noise reduction/speech enhancement, Minimum of 4 programs, Minimum of 8 channels, Minimum of IP57 rating, Programmable directionality, Data logging, Feedback management, Volume control, Telecoil, All hearing aids to come with a single receiver and dome, or flex tips, or micro mould or a RIC mould for Power RIC, wax guard. Non-rechargeable hearing aids must come with 6 appropriately sized zinc air batteries. All re-chargeable hearing aids must be provided with a charger and case | Yes | No | Each
193 | RT274-09-041 | Category 9 | Digital rechargeable BTE for bi- cross fitting, for contra lateral routing of signal, used to fit asymmetric hearing loss, where the one ear is dead and the other ear has a severe-profound hearing loss (71-100dB), MPO: MPO: 120-145dB gain 60-85dB, with Bluetooth compatibility, without tinnitus management, with frequency-lowering. NB: All dB values for BTE Cros and Bicros in Category 9 specifications must comply to the following requirements:Must be according to the 2cc coupler readings, All BTE hearing aids must contain the following as minimum specifications: Technology obsolescence: not older than 18 months from date of publication of tender, Active noise reduction/speech enhancement, Active noise reduction/speech enhancement, Minimum of 4 programs, Minimum of 8 channels, Minimum of IP57 rating, Programmable directionality, Data logging, Feedback management, Volume control, Telecoil, All hearing aids to include an ear mould (acrylic or silicone) or 2 slim tubes and domes, All non-rechargeable hearing aids must include 6 appropriately sized zinc air batteries and dry aid kit, All re-chargeable hearing aids must be provided with a charger and case and dry aid kit and - All dB values for RIC Cros and Bicros in tcategory 9 must comply to the following specifications: Must be according to the 2cc coupler readings, Minimum specification for all RIC hearing aids, Digital signal processing, Technology obsolescence: not older than 18 months from date of publication of tender, Active noise reduction/speech enhancement, Minimum of 4 programs, Minimum of 8 channels, Minimum of IP57 rating, Programmable directionality, Data logging, Feedback management, Volume control, Telecoil, All hearing aids to come with a single receiver and dome, or flex tips, or micro mould or a RIC mould for Power RIC, wax guard. Non-rechargeable hearing aids must come with 6 appropriately sized zinc air batteries. All re-chargeable hearing aids must be provided with a charger and case | Yes | Yes | Each
194 | RT274-09-043 | Category 9 | Digital rechargeable BTE for bi- cross fitting, for contra lateral routing of signal, used to fit asymmetric hearing loss, where the one ear is dead and the other ear has a severe-profound hearing loss (71-100dB), MPO: MPO: 120-145dB gain 60-85dB, without Bluetooth compatibility, without tinnitus management, with frequency-lowering. NB: All dB values for BTE Cros and Bicros in Category 9 specifications must comply to the following requirements:Must be according to the 2cc coupler readings, All BTE hearing aids must contain the following as minimum specifications: Technology obsolescence: not older than 18 months from date of publication of tender, Active noise reduction/speech enhancement, Active noise reduction/speech enhancement, Minimum of 4 programs, Minimum of 8 channels, Minimum of IP57 rating, Programmable directionality, Data logging, Feedback management, Volume control, Telecoil, All hearing aids to include an ear mould (acrylic or silicone) or 2 slim tubes and domes, All non-rechargeable hearing aids must include 6 appropriately sized zinc air batteries and dry aid kit, All re-chargeable hearing aids must be provided with a charger and case and dry aid kit and - All dB values for RIC Cros and Bicros in tcategory 9 must comply to the following specifications: Must be according to the 2cc coupler readings, Minimum specification for all RIC hearing aids, Digital signal processing, Technology obsolescence: not older than 18 months from date of publication of tender, Active noise reduction/speech enhancement, Minimum of 4 programs, Minimum of 8 channels, Minimum of IP57 rating, Programmable directionality, Data logging, Feedback management, Volume control, Telecoil, All hearing aids to come with a single receiver and dome, or flex tips, or micro mould or a RIC mould for Power RIC, wax guard. Non-rechargeable hearing aids must come with 6 appropriately sized zinc air batteries. All re-chargeable hearing aids must be provided with a charger and case | Yes | No | Each
195 | RT274-09-045 | Category 9 | Digital rechargeable BTE for bi- cross fitting, for contra lateral routing of signal, used to fit asymmetric hearing loss, where the one ear is dead and the other ear has a severe-profound hearing loss (71-100dB), MPO: MPO: 120-145dB gain 60-85dB, with Bluetooth compatibility, with tinnitus management, without frequency-lowering. NB: All dB values for BTE Cros and Bicros in Category 9 specifications must comply to the following requirements:Must be according to the 2cc coupler readings, All BTE hearing aids must contain the following as minimum specifications: Technology obsolescence: not older than 18 months from date of publication of tender, Active noise reduction/speech enhancement, Active noise reduction/speech enhancement, Minimum of 4 programs, Minimum of 8 channels, Minimum of IP57 rating, Programmable directionality, Data logging, Feedback management, Volume control, Telecoil, All hearing aids to include an ear mould (acrylic or silicone) or 2 slim tubes and domes, All non-rechargeable hearing aids must include 6 appropriately sized zinc air batteries and dry aid kit, All re-chargeable hearing aids must be provided with a charger and case and dry aid kit and - All dB values for RIC Cros and Bicros in tcategory 9 must comply to the following specifications: Must be according to the 2cc coupler readings, Minimum specification for all RIC hearing aids, Digital signal processing, Technology obsolescence: not older than 18 months from date of publication of tender, Active noise reduction/speech enhancement, Minimum of 4 programs, Minimum of 8 channels, Minimum of IP57 rating, Programmable directionality, Data logging, Feedback management, Volume control, Telecoil, All hearing aids to come with a single receiver and dome, or flex tips, or micro mould or a RIC mould for Power RIC, wax guard. Non-rechargeable hearing aids must come with 6 appropriately sized zinc air batteries. All re-chargeable hearing aids must be provided with a charger and case | Yes | Yes | Each
196 | RT274-09-047 | Category 9 | Digital rechargeable BTE for bi- cross fitting, for contra lateral routing of signal, used to fit asymmetric hearing loss, where the one ear is dead and the other ear has a severe-profound hearing loss (71-100dB), MPO: MPO: 120-145dB gain 60-85dB, without Bluetooth compatibility, with tinnitus management, without frequency-lowering. NB: All dB values for BTE Cros and Bicros in Category 9 specifications must comply to the following requirements:Must be according to the 2cc coupler readings, All BTE hearing aids must contain the following as minimum specifications: Technology obsolescence: not older than 18 months from date of publication of tender, Active noise reduction/speech enhancement, Active noise reduction/speech enhancement, Minimum of 4 programs, Minimum of 8 channels, Minimum of IP57 rating, Programmable directionality, Data logging, Feedback management, Volume control, Telecoil, All hearing aids to include an ear mould (acrylic or silicone) or 2 slim tubes and domes, All non-rechargeable hearing aids must include 6 appropriately sized zinc air batteries and dry aid kit, All re-chargeable hearing aids must be provided with a charger and case and dry aid kit and - All dB values for RIC Cros and Bicros in tcategory 9 must comply to the following specifications: Must be according to the 2cc coupler readings, Minimum specification for all RIC hearing aids, Digital signal processing, Technology obsolescence: not older than 18 months from date of publication of tender, Active noise reduction/speech enhancement, Minimum of 4 programs, Minimum of 8 channels, Minimum of IP57 rating, Programmable directionality, Data logging, Feedback management, Volume control, Telecoil, All hearing aids to come with a single receiver and dome, or flex tips, or micro mould or a RIC mould for Power RIC, wax guard. Non-rechargeable hearing aids must come with 6 appropriately sized zinc air batteries. All re-chargeable hearing aids must be provided with a charger and case | Yes | No | Each
197 | RT274-09-049 | Category 9 | Digital rechargeable BTE for bi- cross fitting, for contra lateral routing of signal, used to fit asymmetric hearing loss, where the one ear is dead and the other ear has a severe-profound hearing loss (71-100dB), MPO: MPO: 120-145dB gain 60-85dB, with Bluetooth compatibility, without tinnitus management, without frequency-lowering. NB: All dB values for BTE Cros and Bicros in Category 9 specifications must comply to the following requirements:Must be according to the 2cc coupler readings, All BTE hearing aids must contain the following as minimum specifications: Technology obsolescence: not older than 18 months from date of publication of tender, Active noise reduction/speech enhancement, Active noise reduction/speech enhancement, Minimum of 4 programs, Minimum of 8 channels, Minimum of IP57 rating, Programmable directionality, Data logging, Feedback management, Volume control, Telecoil, All hearing aids to include an ear mould (acrylic or silicone) or 2 slim tubes and domes, All non-rechargeable hearing aids must include 6 appropriately sized zinc air batteries and dry aid kit, All re-chargeable hearing aids must be provided with a charger and case and dry aid kit and - All dB values for RIC Cros and Bicros in tcategory 9 must comply to the following specifications: Must be according to the 2cc coupler readings, Minimum specification for all RIC hearing aids, Digital signal processing, Technology obsolescence: not older than 18 months from date of publication of tender, Active noise reduction/speech enhancement, Minimum of 4 programs, Minimum of 8 channels, Minimum of IP57 rating, Programmable directionality, Data logging, Feedback management, Volume control, Telecoil, All hearing aids to come with a single receiver and dome, or flex tips, or micro mould or a RIC mould for Power RIC, wax guard. Non-rechargeable hearing aids must come with 6 appropriately sized zinc air batteries. All re-chargeable hearing aids must be provided with a charger and case | Yes | Yes | Each
198 | RT274-09-051 | Category 9 | Digital rechargeable BTE for bi- cross fitting, for contra lateral routing of signal, used to fit asymmetric hearing loss, where the one ear is dead and the other ear has a severe-profound hearing loss (71-100dB), MPO: MPO: 120-145dB gain 60-85dB, without Bluetooth compatibility, without tinnitus management, without frequency-lowering. NB: All dB values for BTE Cros and Bicros in Category 9 specifications must comply to the following requirements:Must be according to the 2cc coupler readings, All BTE hearing aids must contain the following as minimum specifications: Technology obsolescence: not older than 18 months from date of publication of tender, Active noise reduction/speech enhancement, Active noise reduction/speech enhancement, Minimum of 4 programs, Minimum of 8 channels, Minimum of IP57 rating, Programmable directionality, Data logging, Feedback management, Volume control, Telecoil, All hearing aids to include an ear mould (acrylic or silicone) or 2 slim tubes and domes, All non-rechargeable hearing aids must include 6 appropriately sized zinc air batteries and dry aid kit, All re-chargeable hearing aids must be provided with a charger and case and dry aid kit and - All dB values for RIC Cros and Bicros in tcategory 9 must comply to the following specifications: Must be according to the 2cc coupler readings, Minimum specification for all RIC hearing aids, Digital signal processing, Technology obsolescence: not older than 18 months from date of publication of tender, Active noise reduction/speech enhancement, Minimum of 4 programs, Minimum of 8 channels, Minimum of IP57 rating, Programmable directionality, Data logging, Feedback management, Volume control, Telecoil, All hearing aids to come with a single receiver and dome, or flex tips, or micro mould or a RIC mould for Power RIC, wax guard. Non-rechargeable hearing aids must come with 6 appropriately sized zinc air batteries. All re-chargeable hearing aids must be provided with a charger and case | Yes | No | Each
199 | RT274-10-075 | Category 10 | Cochlear implant sound processor compatible with all cochlear implants and electrode arrays previously issued in South Africa over the past 20 years. Sound processor kit to include: one behind-the-ear or off-the-ear sound processor with the option of both omnidirectional and directional microphones, battery modules sufficient to last a full day with sufficient batteries to support the average patient for 3 years of daily use; battery charger; a range of magnet strength options; coil and coil cable to connect sound processor to internal coil if required for sound processor, with available spares sufficient for the average patient for 3 years microphone protection covers (sufficient for 3 years); an electronic device dry and store kit; an activity/everyday case; FM system compatibility and a remote microphone system with proven outcomes included. Sound processor to be available in a range of colours and wearing options and to include: automatic auditory scene classifier; water resistance rating of IP54 for temporary immersion and an option of accessories to allow for IP68 for full immersion. Processor must have reliability reporting and manufacturer certification (as per submitted requirements), a 3-year warranty on sound processor, minimum 1-year warranty on parts and maintenance plan to support patients with repairs and consumables | Yes | Yes | Each
200 | RT274-10-077 | Category 10 | Cochlear implant sound processor compatible with all cochlear implants and electrode arrays issued in South Africa over the past 20 years. Sound processor kit to include: one behind-the-ear or off-the-ear sound processor with the option of both omnidirectional and directional microphones, battery modules sufficient to last a full day with sufficient batteries to support the average patient for 3 years of daily use; battery charger; a range of magnet strength options; coil and coil cable to connect sound processor to internal coil if required for sound processor, with available spares sufficient for the average patient for 3 years microphone protection covers (sufficient for 3 years); an electronic device dry and store kit; an activity/everyday case; FM system compatibility and a remote microphone system with proven outcomes included. Sound processor to be available in a range of colours and wearing options and to include: automatic auditory scene classifier; water resistance rating of IP54 for temporary immersion and an option of accessories to allow for IP68 for full immersion. Processor must have reliability reporting and manufacturer certification (as per submitted requirements), a 3-year warranty on sound processor, minimum 1-year warranty on parts | Yes | Yes | Each
201 | RT274-10-079 | Category 10 | Cochlear implant sound processors compatible with all cochlear implants and electrode arrays issued in South Africa over the past 20 years. Sound processor kit to include: two behind-the-ear or off-the-ear sound processors with the option of both omnidirectional and directional microphones, battery modules sufficient to last a full day with sufficient batteries to support the average patient for 3 years of daily use; battery charger; a range of magnet strength options; coil and coil cable to connect sound processor to internal coil if required for sound processor, with available spares sufficient for the average patient for 3 years microphone protection covers (sufficient for 3 years); an electronic device dry and store kit; an activity/everyday case; FM system compatibility and a remote microphone system with proven outcomes included. Sound processor to be available in a range of colours and wearing options and to include: automatic auditory scene classifier; water resistance rating of IP54 for temporary immersion and an option of accessories to allow for IP68 for full immersion. Processor must have reliability reporting and manufacturer certification (as per submitted requirements), a 3-year warranty on sound processor, minimum 1-year warranty on parts and maintenance plan to support patients with repairs and consumables. | Yes | Yes | Each
202 | RT274-10-081 | Category 10 | Cochlear implant sound processors compatible with all cochlear implants and electrode arrays issued in South Africa over the past 20 years. Sound processor kit to include: two behind-the-ear or off-the-ear sound processors with the option of both omnidirectional and directional microphones, battery modules sufficient to last a full day with sufficient batteries to support the average patient for 3 years of daily use; battery charger; a range of magnet strength options; coil and coil cable to connect sound processor to internal coil if required for sound processor, with available spares sufficient for the average patient for 3 years microphone protection covers (sufficient for 3 years); an electronic device dry and store kit; an activity/everyday case; FM system compatibility and a remote microphone system with proven outcomes included. Sound processor to be available in a range of colours and wearing options and to include: automatic auditory scene classifier; water resistance rating of IP54 for temporary immersion and an option of accessories to allow for IP68 for full immersion. Processor must have reliability reporting and manufacturer certification (as per submitted requirements), a 3-year warranty on sound processor, minimum 1-year warranty on parts. | Yes | Yes | Each
203 | RT274-10-083 | Category 10 | Cochlear implant sound processor (one processor only; no kit) compatible with all cochlear implants and electrode arrays issued in South Africa over the past 20 years. Sound processor to have the option of both omnidirectional and directional microphones; to be available in a range of colours and wearing options and to include: automatic auditory scene classifier; water resistance rating of IP54 for temporary immersion and an option of accessories to allow for IP68 for full immersion. Processor must have reliability reporting and manufacturer certification (as per submitted requirements), a 3-year warranty on sound processor. | Yes | Yes | Each
204 | RT274-10-020 | Category 10 | USB cable connection for programming a cochlear implant, compatible with all sound processors awarded on the tender as well as all previous devices issued within South Africa over the past 5 years | Yes | No | Each
205 | RT274-10-084 | Category 10 | Programming cable for cochlear implant, compatible with all sound processors issued within South Africa over the past 5 years. Multiple offers. | Yes | No | Each
206 | RT274-10-086 | Category 10 | Programming pod/device compatible with all cochlear implant sound processors issued within South Africa over the past 5 years. | Yes | No | Each
207 | RT274-10-088 | Category 10 | Cochlear implant sound processor rechargeable battery module compatible with all sound processors issued within South Africa over the past 5 years, available in a range of colours including black, brown and beige. Bidder to specify quantity and price per unit. Multiple offers. | Yes | No | Each
208 | RT274-10-090 | Category 10 | Cochlear implant sound processor disposable battery holder/rack compatible with all sound processors issued within South Africa over the past 5 years, available in a range of colours including black, brown and beige. | Yes | No | Each
209 | RT274-10-092 | Category 10 | Cochlear implant sound processor disposable battery cover compatible with all sound processors all previous devices issued within South Africa over the past 5 years, available in a range of colours including black, brown and beige. Multiple offers. | Yes | No | Each
210 | RT274-10-094 | Category 10 | Cochlear implant coils compatible with all sound processors issued within South Africa over the past 5 years available in a range of colours including black, brown and beige. Multiple offers. | Yes | No | Each
211 | RT274-10-095 | Category 10 | Cochlear implant cables compatible with all sound processors issued within South Africa over the past 5 years available in a range of colours including black, brown and beige. Multiple offers. | Yes | No | Each
212 | RT274-10-097 | Category 10 | Magnets compatible with all cochlear implant sound processors issued within South Africa over the past 5 years available in a range of strength options for retention and a range of colours including black, brown and beige. Multiple offers. | Yes | No | Each
213 | RT274-10-028 | Category 10 | Magnet insertion tool compatible with all cochlear implant sound processors awarded on the tender as well as all previous devices issued within South Africa over the past 5 years | Yes | No | Each
214 | RT274-10-098 | Category 10 | Earhooks for retention of cochlear implant sound processor on the ear, compatible with all sound processors issued within South Africa over the past 5 years and in a range of sizes from infant to adult. Multiple offers. | Yes | No | Each
215 | RT274-10-100 | Category 10 | Charger for sound processor rechargeable battery compatible with all cochlear implant sound processors issued within South Africa over the past 5 years. Multiple offers. | Yes | No | Each
216 | RT274-10-102 | Category 10 | Remote control device to allow for control of cochlear implant sound processor, include volume and programme controls, compatible with all sound processors issued within South Africa over the past 5 years | Yes | Yes | Each
217 | RT274-10-104 | Category 10 | Waterproof cover to allow cochlear implant sound processor to be submerged in water with a minimum IP68 rating compatible with all sound processors issued within South Africa over the past 5 years | Yes | No | Each
218 | RT274-10-106 | Category 10 | Retention device for cochlear implant sound processor - to secure sound processor on the ear to prevent damage to or loss of processor, compatible with sound processors issued over the past 5 years. Multiple offers | Yes | No | Each
219 | RT274-10-108 | Category 10 | Retention device for cochlear implant sound processor - to secure sound processor to the head or hair to prevent damage to or loss of processor, compatible with all cochlear implant processors issued in South Africa over the past 5 years. Multiple offers. | Yes | No | Each
220 | RT274-10-110 | Category 10 | Battery charger set with charger and cable compatible with all cochlear implant sound processors pissued within South Africa over the past 5 years | No | No | Each
221 | RT274-10-112 | Category 10 | Battery charger compatible with all cochlear implant sound processors issued within South Africa over the past 5 years | No | No | Each
222 | RT274-10-113 | Category 10 | Microphone covers/protectors compatible with all cochlear implant sound processors issued within South Africa over the past 5 years | Yes | No | Each
223 | RT274-10-037 | Category 10 | Battery charger cable compatible with all cochlear implant sound processors awarded on the tender as well as all previous devices issued within South Africa over the past 5 years | Yes | No | Each
224 | RT274-10-114 | Category 10 | Portable battery charging cable in the event of inability to charge via electricity compatible with all cochlear implant sound processors issued within South Africa over the past 5 year | Yes | No | Each
225 | RT274-10-116 | Category 10 | Power bank for charging cochlear implant sound processor in the event of inability to charge via electricity | No | No | Each
226 | RT274-10-118 | Category 10 | Adaptor to connect rechargeable battery to cochlear implant sound processor | Yes | No | Each
227 | RT274-10-119 | Category 10 | O-ring part for cochlear implant sound processor | Yes | No | Each
228 | RT274-10-120 | Category 10 | Coil cover for cochlear implant sound processor for all magnet strengths. Multiple offers. | Yes | No | Each
229 | RT274-10-125 | Category 10 | Electronic drying kit for cochlear implant sound processor that does not require an additional dessicant component. | No | No | Each
230 | RT274-10-126 | Category 10 | Water-resistant stickers designed specifically for cochlear implant sound processors, to include paediatric-friendly options. | Yes | No | Each
231 | RT274-10-130 | Category 10 | Direct streaming attachment to processor to facilitate streaming from phone or tablet directly to cochlear implant sound processor | Yes | Yes | Each
232 | RT274-10-050 | Category 10 | Device to enable Bluetooth wireless streaming to cochlear implant sound processor where direct streaming is not available in a cochlear implant sound processor | Yes | Yes | Each
233 | RT274-10-131 | Category 10 | Magnetic detachable digital bone conduction sound processor (one processor only; no surgical kit) compatible with all passive transcutaneous bone conduction hearing implants awarded on the tender as well as all previous devices issued in South Africa over the past 10 years. The sound processor should be suitable for conductive and mixed hearing losses with bone conduction thresholds of up to 45 dBHL (normal to moderate hearing thresholds). Sound processor kit to comprise of a minimum of: one sound processing unit, one test rod/skull simulator where indicated, pack of 6 appropriately sized disposable batteries/appropriate rechargeable battery where indicated, drying capsule/kit, sound processor cover (optional), an activity/everyday case, tamperproof locks and indicator lights (optional). The sound processor should with have the option of both omnidirectional and directional microphones, have capability to be programmed as a standalone or via an interfaced software, may have telephone compatibility, direct audio input, datalogging and FM system compatibility and streaming compatibility. Sound processor to be available in a range of colours, water resistance rating of IP54 for temporary emersion. Processor must have reliability reporting and manufacturer certification (as per submitted requirements), a 2-year warranty on sound processor and aftercare service must be guaranteed by the manufacturer. | Yes | Yes | Each
234 | RT274-10-133 | Category 10 | Magnetic detachable digital bone conduction sound processor (one processor only; no surgical kit) compatible with all passive transcutaneous bone conduction hearing implants awarded on the tender as well as all previous devices issued in South Africa over the past 10 years. The sound processor should be suitable for conductive and mixed hearing losses with bone conduction thresholds of up to 65dBHL (power option). Sound processor kit to comprise of a minimum of: one sound processing unit, one test rod/skull simulator where indicated, pack of 6 appropriately sized disposable batteries/appropriate rechargeable battery where indicated, drying capsule/kit, sound processor cover (optional), an activity/everyday case, tamperproof locks and indicator lights (optional). The sound processor should with have the option of both omnidirectional and directional microphones, have capability to be programmed as a standalone or via an interfaced software, may have telephone compatibility, direct audio input, datalogging and FM system compatibility and streaming compatibility. Sound processor to be available in a range of colours, water resistance rating of IP54 for temporary emersion. Processor must have reliability reporting and manufacturer certification (as per submitted requirements), a 2-year warranty on sound processor and aftercare service must be guaranteed by the manufacturer. | Yes | Yes | Each
235 | RT274-10-134 | Category 10 | Replacement battery doors for bone conduction sound processors. | Yes | No | Each
236 | RT274-10-135 | Category 10 | Microphone covers/protectors compatible with all bone conduction hearing devices issued within South Africa over the past 5 years | Yes | No | Each
237 | RT274-10-136 | Category 10 | Retention devices for all bone conduction hearing devices issued within South Africa over the past 5 years. | Yes | No | Each
238 | RT274-10-137 | Category 10 | Water-resistant stickers designed specifically for bone conduction sound processors, to include paediatric-friendly options. | Yes | No | Each
239 | RT274-10-138 | Category 10 | Custom coil cable covers to protect cochlear implant sound processor coils from damage. | Yes | No | Each
240 | RT274-10-140 | Category 10 | Direct streaming device to facilitate ease of listening in difficult listening situations compatible with all implantable bone conduction implant sound processors issued within South Africa over the past 5 years | Yes | Yes | Each
241 | RT274-11-007 | Category 11 | South African Sign language dictionary: Dictionary with the following - Sign and words - Recorded version of signs on electronic medium (optional) | No | No | Each