## RT284-2019 CONTRACT CIRCULAR
RT284-2019: SUPPLY AND DELIVERY OF HYPODERMIC SYRINGES, NEEDLES AND BLOODLETTING DEVICES AND RELATED ITEMS 1 NOVEMBER 2019 TO 31 OCTOBER 2022
ITEM CODE | NO | ITEM DESCRIPTION | OFFER OPTION | AWARDED BIDDER | DELIVERY PRICE IN ZAR (AS PER YOUR BID DOCUMENT) VAT 15% | BRAND NAME | PRODUCT CODE NUMBER | NAME OF MANUFACTURER | COUNTRY OF ORIGIN | LEAD TIME DAY / S | IMPORT COMPONENT OF PRICE AS % | FOREIGN CURRENCY | MINIMUM ORDER QUANTITY | CONFORM TO SPECIFICATION | UOM OFFERED / SPEC DEVIATION | REASONS FOR PASSOVER | TEST RESPORTS (COMPLY / NOT COMPLY) | RECOMMENDATIONS | TOTAL POINTS SCORED
41104102-00002 | 1 | Bloodletting lancet, with safety retractable point, for use independent of any other device. For use in adults Size 21G Sterile, Single use, Outer package must be labelled sterile UNIT: Box (100 per box) |  | Chemical World cc | 46 | SURGI PRO | BLS | YUANDING-XINLE | CHINA | 3 | 85 | USD | 100 | YES | 100/Box | COMPLY |  | Recommended | 100
41104102-00002 | 1 | Bloodletting lancet, with safety retractable point, for use independent of any other device. For use in adults Size 21G Sterile, Single use, Outer package must be labelled sterile UNIT: Box (100 per box) | 1 | CENTRAL MEDICAL PTY LTD | 0.47 | Mejor | BL2118 | Tianjin Huahong Technology Co Ltd | China | 21 | 80 | China yuan | 50 box of 100 unit | yes | box of 100 unit | COMPLY |  |  | 93.8636
41104102-00002 | 1 | Bloodletting lancet, with safety retractable point, for use independent of any other device. For use in adults Size 21G Sterile, Single use, Outer package must be labelled sterile UNIT: Box (100 per box) | Basic Line Offer | BIOCOR HOSPITAL SUPPLIES PTY LTD | 0.49 | Medlance Plus | MLP21-1.8 | HTL Strefa SA | Poland | 70 | 85 | USD | 10 | No | Per Box of 200 | COMPLY |  |  | 100
41104102-00002 | 1 | Bloodletting lancet, with safety retractable point, for use independent of any other device. For use in adults Size 21G Sterile, Single use, Outer package must be labelled sterile UNIT: Box (100 per box) | A | Quality Medical Supplies Pty Ltd | 0.54 | Linkfar | TBA | Shandong Lianfa Medical Plastic Products Co.Ltd | China | 14 | 80% | USD | 20 Box of 100 | Yes | Box | COMPLY pending confirmation of latex free |  |  | 78.2603
41104102-00002 | 1 | Bloodletting lancet, with safety retractable point, for use independent of any other device. For use in adults Size 21G Sterile, Single use, Outer package must be labelled sterile UNIT: Box (100 per box) |  | Chemical World cc | 0.55 | SURGI PLUS | BLS | STERILANCE | CHINA | 3 | 85 | USD | 100 | YES | 100 | COMPLY |  |  | -25.7317
41104102-00002 | 1 | Bloodletting lancet, with safety retractable point, for use independent of any other device. For use in adults Size 21G Sterile, Single use, Outer package must be labelled sterile UNIT: Box (100 per box) |  | GIBELA TRADE AND INVEST 1001 CC | 0.56 | TARA | BLO001 | Shandong Qiaopai Group Co., Ltd | CHINA | 21 | 85 | USD | 10 | YES | box 100 | COMPLY |  |  | -36.5385
41104102-00002 | 1 | Bloodletting lancet, with safety retractable point, for use independent of any other device. For use in adults Size 21G Sterile, Single use, Outer package must be labelled sterile UNIT: Box (100 per box) |  | Chemical World cc | 0.64 | EASY JECT | BLS | MEDHARMONY | CHINA | 3 | 85 | USD | 100 | YES | 100 | COMPLY |  |  | 62.6154
41104102-00002 | 1 | Bloodletting lancet, with safety retractable point, for use independent of any other device. For use in adults Size 21G Sterile, Single use, Outer package must be labelled sterile UNIT: Box (100 per box) | Only one offer submitted | Tara Technology | 0.66 | Tara | TAR809 | HTL-Strafa | Poland | 21 days | 85% | USD | 10 | Yes | 100 per box | No samples sub mitted for evaluation |  |  | 100
41104102-00002 | 1 | Bloodletting lancet, with safety retractable point, for use independent of any other device. For use in adults Size 21G Sterile, Single use, Outer package must be labelled sterile UNIT: Box (100 per box) |  | LECHOBA MEDICAL TECHNOLOGIES (PTY) LTD | 0.79 | Vitrex (Sterilance) Lite III | 377213 | VITREX MEDICAL A/S | DENMARK | 21 DAYS | 75% | EURO | 1000 | YES (PUNCTURE DEPTH OF 1.8MM) | 100/BOX | COMPLY |  |  | 60.8
41104102-00002 | 1 | Bloodletting lancet, with safety retractable point, for use independent of any other device. For use in adults Size 21G Sterile, Single use, Outer package must be labelled sterile UNIT: Box (100 per box) |  | Medi-Core Technologies | 0.79 | ULIFE | 41104102-00002 | JIAXING U-LIFE MEDICAL DEVICE TECHNOLOGY.,CO LTD | CHINA | 21 | 80 | USD | 1 box | YES | BOX | COMPLY pending confirmation of latex free |  |  | 69.4298
41104102-00002 | 1 | Bloodletting lancet, with safety retractable point, for use independent of any other device. For use in adults Size 21G Sterile, Single use, Outer package must be labelled sterile UNIT: Box (100 per box) | N/A | PROMED TECHNOLOGIES (PTY) LTD | 0.8 | U-LIFE | 102002 | JIAXING | CHINA | 14 DAYS | 85% | USD | 1 BOX | YES | BOX | COMPLY pending confirmation of latex free |  |  | 100
41104102-00002 | 1 | Bloodletting lancet, with safety retractable point, for use independent of any other device. For use in adults Size 21G Sterile, Single use, Outer package must be labelled sterile UNIT: Box (100 per box) |  | UNITRADE 1032 CC | 0.89 | UNIJEC | 1 | SHANGHAI KINDLY ENTERPRISE DEVELOPMENT GROUP | ROC | 21 | 85 | USD | 2000 | YES | 100 | COMPLY pending confirmation of latex free |  |  | 100
41104102-00002 | 1 | Bloodletting lancet, with safety retractable point, for use independent of any other device. For use in adults Size 21G Sterile, Single use, Outer package must be labelled sterile UNIT: Box (100 per box) | Offer 1 of 1 | SSEM Mthembu Medical (Pty) Ltd | 0.91 | Solace | LOSL21.2.2 | Solace | Sweden | 56 | 85% | USD | 1 | Yes | 100/Box | COMPLY pending confirmation of latex free |  |  | 100
41104102-00002 | 1 | Bloodletting lancet, with safety retractable point, for use independent of any other device. For use in adults Size 21G Sterile, Single use, Outer package must be labelled sterile UNIT: Box (100 per box) | 1 | Vital Life Medical Distributors (PTY) LTD | 0.92 | Healthease | NL-07 | Neomedic | China | 14 days | 80% | US Dollar | 50 boxes | Yes | 100 per box | COMPLY pending confirmation of latex free |  |  | -32.8235
41104102-00002 | 1 | Bloodletting lancet, with safety retractable point, for use independent of any other device. For use in adults Size 21G Sterile, Single use, Outer package must be labelled sterile UNIT: Box (100 per box) | OFFER 1 | SUPRA HEALTHCARE JOHANNESBURG PTY LTD | 1.01 | Avacare | LANCESAF001 | HEALTHY MEDICAL DEVELOPMENT CO., LTD | CHINA | Initial 60, thereafter 14 | 84 | USD | 1 Box | YES | 100 pcs/box | No brand name. No product code. |  |  | 92.3529
41104102-00002 | 1 | Bloodletting lancet, with safety retractable point, for use independent of any other device. For use in adults Size 21G Sterile, Single use, Outer package must be labelled sterile UNIT: Box (100 per box) | OFFER 2 | SUPRA HEALTHCARE JOHANNESBURG PTY LTD | 1.13 | Avacare | LANCESAF001 | HENAN AILE INDUSTRIAL CO.,LTD | CHINA | Initial 60, thereafter 14 | 84 | USD | 1 Box | YES | 100 pcs/box | No brand name. No product code. |  |  | 100
41104102-00002 | 1 | Bloodletting lancet, with safety retractable point, for use independent of any other device. For use in adults Size 21G Sterile, Single use, Outer package must be labelled sterile UNIT: Box (100 per box) | N/A | SAFARMEX (PTY) LTD | 1.14 | Saclin | To be confirmed | Zhejiang Mediunion Healthcare Group Co. Ltd | China | 3days | 80% | US$ | 50 000 pieces | Yes | Box of 100 | No sample submitted for evaluation |  |  | -32.8235
41104102-00002 | 1 | Bloodletting lancet, with safety retractable point, for use independent of any other device. For use in adults Size 21G Sterile, Single use, Outer package must be labelled sterile UNIT: Box (100 per box) |  | AKACIA MEDICAL (PTY) LTD | 2.02 | AKACIA | MI006400 | HTL STREFA | POLAND | EX STOCK -20DAYS | 67% | USD | BOX 100 | YES | N/A | COMPLY |  |  | 100
41104102-00002 | 1 | Bloodletting lancet, with safety retractable point, for use independent of any other device. For use in adults Size 21G Sterile, Single use, Outer package must be labelled sterile UNIT: Box (100 per box) | 1 | Chemical World cc | 0.44 | MEDIJECT | BLS | KYUAN | CHINA | 3 | 85 | USD | 100 | YES | 100 | Product information is reflected on a removable sticker. LOT number, shelf life and date of production of the sticker and outer package is not the same. Product information of a removable sticker indicates a size 21G and the product information of the outer package indicates size 28G |  |  | 67.9412
41104102-00003 | 2 | Bloodletting lancet, with safety retractable point, for use independent of any other device. For use in paediatric Outer package must be labelled sterile Size: 26G Sterile, Single use, UNIT: Box (100 per box) |  | Chemical World cc | 0.44 | MEDI JECT | BLS | YUANDING XINLE | CHINA | 3 | 85 | USD | 100 | YES | 100 | Offered a size 21G instead of 25G |  |  | 100
41104102-00003 | 2 | Bloodletting lancet, with safety retractable point, for use independent of any other device. For use in paediatric Outer package must be labelled sterile Size: 25G Sterile, Single use, UNIT: Box (100 per box) | 1 | CENTRAL MEDICAL PTY LTD | 46.95 | Mejor | BL2514 | Tianjin Huahong Technology Co Ltd | China | 21 | 80 | China yuan | 50 box of 100 unit | yes | 100/Box | COMPLY |  | Recommended | 100
41104102-00003 | 2 | Bloodletting lancet, with safety retractable point, for use independent of any other device. For use in paediatric Outer package must be labelled sterile Size: 25G Sterile, Single use, UNIT: Box (100 per box) | Basic Line Offer | BIOCOR HOSPITAL SUPPLIES PTY LTD | 0.51 | Medlance Plus | MPL25-1.5 | HTL Strefa SA | Poland | 70 | 85 | USD | 10 | No | Per Box of 200 | COMPLY |  |  | 100
41104102-00003 | 2 | Bloodletting lancet, with safety retractable point, for use independent of any other device. For use in paediatric Outer package must be labelled sterile Size: 25G Sterile, Single use, UNIT: Box (100 per box) |  | Chemical World cc | 0.52 | SURGI PRO | BLS | KYUAN | CHINA | 3 | 85 | USD | 100 | YES | 100 | Offered a size 26G instead of 25G LOT number, shelf life and date of production of the sticker and outer package is not the same. Product information of a removable sticker indicates a size 25G and the product information of the out |  |  | 100
41104102-00003 | 2 | Bloodletting lancet, with safety retractable point, for use independent of any other device. For use in paediatric Outer package must be labelled sterile Size: 25G Sterile, Single use, UNIT: Box (100 per box) | A | Quality Medical Supplies Pty Ltd | 0.54 | Linkfar | TBA | Shandong Lianfa Medical Plastic Products Co.Ltd | China | 14 | 80% | USD | 20 Box of 100 | Yes | Box | Offered a size 26G instead of 25G |  |  | 100
41104102-00003 | 2 | Bloodletting lancet, with safety retractable point, for use independent of any other device. For use in paediatric Outer package must be labelled sterile Size: 26G Sterile, Single use, UNIT: Box (100 per box) |  | Chemical World cc | 0.59 | SURGI PLUS | BLS | STERILANCE | CHINA | 3 | 85 | USD | 100 | YES | 100 | Offered a size 26G instead of 25G |  |  | 100
41104102-00003 | 2 | Bloodletting lancet, with safety retractable point, for use independent of any other device. For use in paediatric Outer package must be labelled sterile Size: 25G Sterile, Single use, UNIT: Box (100 per box) | Only one offer submitted | Tara Technology | 0.66 | Tara | TAR810 | HTL-Strafa | Poland | 21 days | 85% | USD | 10 | Yes | 100 per box | No samples sub mitted for evaluation |  |  | 100
41104102-00003 | 2 | Bloodletting lancet, with safety retractable point, for use independent of any other device. For use in paediatric Outer package must be labelled sterile Size: 25G Sterile, Single use, UNIT: Box (100 per box) |  | GIBELA TRADE AND INVEST 1001 CC | 0.7 | TARA | BLO002 | Shandong Qiaopai Group Co., Ltd | CHINA | 21 | 85 | USD | 10 | YES | box 100 | Size is indicated on a removable sticker refers 25G and the color coding of the packaging for Tara in terms bloodletting devices indicates a size of 26G which is the color blue. Lot no: is indicated on a removable sticker 19021320/03 |  |  | -22.9464
41104102-00003 | 2 | Bloodletting lancet, with safety retractable point, for use independent of any other device. For use in paediatric Outer package must be labelled sterile Size: 25G Sterile, Single use, UNIT: Box (100 per box) | OPTIONAL OFFER OF 26G | LECHOBA MEDICAL TECHNOLOGIES (PTY) LTD | 0.79 | Vitrex (Sterilance) Lite III | 375213 | VITREX MEDICAL A/S | DENMARK | 21 DAYS | 75% | EURO | 1000 | NO (OFFERED 26G .PUNCTURE DEPTH OF 1.8MM) | 100/BOX | Offered a size 26G instead of 25G |  |  | -565
41104102-00003 | 2 | Bloodletting lancet, with safety retractable point, for use independent of any other device. For use in paediatric Outer package must be labelled sterile Size: 25G Sterile, Single use, UNIT: Box (100 per box) |  | Medi-Core Technologies | 0.79 | ULIFE | 41104102-00003 | JIAXING U-LIFE MEDICAL DEVICE TECHNOLOGY.,CO LTD | CHINA | 21 | 80 | USD | 1 box | YES | BOX | COMPLY pending confirmation of latex free |  |  | -307.9577
41104102-00003 | 2 | Bloodletting lancet, with safety retractable point, for use independent of any other device. For use in paediatric Outer package must be labelled sterile Size: 25G Sterile, Single use, UNIT: Box (100 per box) | N/A | PROMED TECHNOLOGIES (PTY) LTD | 0.8 | U-LIFE | 102003 | JIAXING | CHINA | 14 DAYS | 85% | USD | 1 BOX | YES | BOX | COMPLY pending confirmation of latex free |  |  | -386.5075
41104102-00003 | 2 | Bloodletting lancet, with safety retractable point, for use independent of any other device. For use in paediatric Outer package must be labelled sterile Size: 25G Sterile, Single use, UNIT: Box (100 per box) | Offer 1 of 1 | SSEM Mthembu Medical (Pty) Ltd | 0.91 | Solace | LOSL25.2.2 | Solace | Sweden | 56 | 85% | USD | 1 | Yes | 100/Box | No sample submitted for evaluation |  |  | -242.1336
41104102-00003 | 2 | Bloodletting lancet, with safety retractable point, for use independent of any other device. For use in paediatric Outer package must be labelled sterile Size: 25G Sterile, Single use, UNIT: Box (100 per box) | OFFER 1 | SUPRA HEALTHCARE JOHANNESBURG PTY LTD | 1.01 | Avacare | LANCESAF003 | HEALTHY MEDICAL DEVELOPMENT CO., LTD | CHINA | Initial 60, thereafter 14 | 84 | USD | 1 Box | YES | 100 pcs/box | color coding indicates a size 26g instead of size 25. No brand name. No product code. |  |  | 95
41104102-00003 | 2 | Bloodletting lancet, with safety retractable point, for use independent of any other device. For use in paediatric Outer package must be labelled sterile Size: 25G Sterile, Single use, UNIT: Box (100 per box) |  | AKACIA MEDICAL (PTY) LTD | 1.02 | AKACIA | MI006410 | HTL STREFA | POLAND | EX STOCK -20DAYS | 67% | USD | BOX 100 | YES | N/A | COMPLY |  |  | 95
41104102-00003 | 2 | Bloodletting lancet, with safety retractable point, for use independent of any other device. For use in paediatric Outer package must be labelled sterile Size: 25G Sterile, Single use, UNIT: Box (100 per box) | OFFER 2 | SUPRA HEALTHCARE JOHANNESBURG PTY LTD | 1.13 | Avacare | LANCESAF003 | HENAN AILE INDUSTRIAL CO.,LTD | CHINA | Initial 60, thereafter 14 | 84 | USD | 1 Box | YES | 100 pcs/box | No sample submitted for evaluation |  |  | 81.3422
41104102-00003 | 2 | Bloodletting lancet, with safety retractable point, for use independent of any other device. For use in paediatric Outer package must be labelled sterile Size: 25G Sterile, Single use, UNIT: Box (100 per box) | N/A | SAFARMEX (PTY) LTD | 1.14 | Saclin | To be confirmed | Zhejiang Mediunion Healthcare Group Co. Ltd | China | 3days | 80% | US$ | 50 000 pieces | Yes | Box of 100 | No sample submitted for evaluation |  |  | 100
41104102-00003 | 2 | Bloodletting lancet, with safety retractable point, for use independent of any other device. For use in paediatric Outer package must be labelled sterile Size: 25G Sterile, Single use, UNIT: Box (100 per box) | Basic Line offer | Becton Dickinson | 4.75 | BD VAC QUICKHEEL, GREEN | 368103 | Becton Dickinson | United Kingdom | 7 Working Days | 76% | USD | 1 BOX | NO | 1 BOX = 50 UNITS / Blade Lancet, Depth 1mm, Width/Gauge 25mm | To request the company to confirm the size of 1mm X 2.5mm as the specs of 25G |  |  | 100
41104102-00003 | 2 | Bloodletting lancet, with safety retractable point, for use independent of any other device. For use in paediatric Outer package must be labelled sterile Size: 25G Sterile, Single use, UNIT: Box (100 per box) | BASIC LINE OFFER | MALACHITE MEDICAL (PTY) LTD | 0.48 | PRIONGUARD PRIMALANCE | LAN00125 | SUZHOU KYUAN MEDICAL | CHINA | 21 | 74.27 | USD | BOX 100 | YES | 40 BOXES OF 100 PER CASE | The brand name on the removable sticker is prionguarde and the brand name on the outer package is KYUAN. Product information that is reflected on the removable sticker indicates a size 25G and the product information that is on the outer package is 28G which is misleading. The reference number is hand written LAN00125. LOT number, shelf life and date of production of the sticker and outer package is not the same. |  |  | 60
41104102-00004 | 3 | Bloodletting lancet, with safety retractable point, for use independent of any other device. For use in neonates Size: 28G Sterile, Single use Outer package must be labelled sterile UNIT: Box (100 per box) |  | Chemical World cc | 43.13 | MEDIJECT | BLS |  | CHINA | 3 | 85 | USD | 100 | YES | 100/Box | COMPLY |  | Recommended | 100
41104102-00004 | 3 | Bloodletting lancet, with safety retractable point, for use independent of any other device. For use in neonates Size: 28G Sterile, Single use Outer package must be labelled sterile | 1 | CENTRAL MEDICAL PTY LTD | 0.47 | Mejor | BL2816 | Tianjin Huahong Technology Co Ltd | China | 21 | 80 | China yuan | 50 box of 100 unit | yes | box of 100 unit | COMPLY |  |  | 94
41104102-00004 | 3 | Bloodletting lancet, with safety retractable point, for use independent of any other device. For use in neonates Size: 28G Sterile, Single use Outer package must be labelled sterile |  | Chemical World cc | 0.52 | SURGI PRO | BLS |  | CHINA | 3 | 85 | USD | 100 | YES | 100 | COMPLY |  |  | 100
41104102-00004 | 3 | Bloodletting lancet, with safety retractable point, for use independent of any other device. For use in neonates Size: 28G Sterile, Single use Outer package must be labelled sterile | A | Quality Medical Supplies Pty Ltd | 0.54 | Linkfar | TBA | Shandong Lianfa Medical Plastic Products Co.Ltd | China | 14 | 80% | USD | 20 Box of 100 | Yes | Box | COMPLY pending confirmation of latex free |  |  | 93.0769
41104102-00004 | 3 | Bloodletting lancet, with safety retractable point, for use independent of any other device. For use in neonates Size: 28G Sterile, Single use Outer package must be labelled sterile |  | Chemical World cc | 0.59 | SURI PLUS | BLS |  | CHINA | 3 | 85 | USD | 100 | YES | 100 | COMPLY |  |  | 93.0769
41104102-00004 | 3 | Bloodletting lancet, with safety retractable point, for use independent of any other device. For use in neonates Size: 28G Sterile, Single use Outer package must be labelled sterile | Only one offer submitted | Tara Technology | 0.66 | Tara | TAR811 | HTL-Strafa | Poland | 21 days | 85% | USD | 10 | Yes | 100 per box | No samples sub mitted for evaluation |  |  | 93.0769
41104102-00004 | 3 | Bloodletting lancet, with safety retractable point, for use independent of any other device. For use in neonates Size: 28G Sterile, Single use Outer package must be labelled sterile |  | GIBELA TRADE AND INVEST 1001 CC | 0.7 | TARA | BLO003 | Shandong Qiaopai Group Co., Ltd | CHINA | 21 | 85 | USD | 10 | YES | box 100 | COMPLY |  |  | 93.0769
41104102-00004 | 3 | Bloodletting lancet, with safety retractable point, for use independent of any other device. For use in neonates Size: 28G Sterile, Single use Outer package must be labelled sterile |  | LECHOBA MEDICAL TECHNOLOGIES (PTY) LTD | 0.79 | Vitrex (Sterilance) Lite III | 374213 | VITREX MEDICAL A/S (PART OF CTI GROUP) | DENMARK | 21 DAYS | 75% | EURO | 1000 | YES (PUNCTURE DEPTH OF 1.8MM) | 100/BOX | COMPLY |  |  | 93.0769
41104102-00004 | 3 | Bloodletting lancet, with safety retractable point, for use independent of any other device. For use in neonates Size: 28G Sterile, Single use Outer package must be labelled sterile |  | Medi-Core Technologies | 0.79 | ULIFE | 41104102-00004 | JIAXING U-LIFE MEDICAL DEVICE TECHNOLOGY.,CO LTD | CHINA | 21 | 80 | USD | 1 box | yes | BOX | COMPLY pending confirmation of latex free |  |  | 51
41104102-00004 | 3 | Bloodletting lancet, with safety retractable point, for use independent of any other device. For use in neonates Size: 28G Sterile, Single use Outer package must be labelled sterile | N/A | PROMED TECHNOLOGIES (PTY) LTD | 0.8 | U-LIFE | 102004 | JIAXING | CHINA | 14 DAYS | 85% | USD | 1 BOX | YES | BOX | COMPLY pending confirmation of latex free |  |  | 39.5918
41104102-00004 | 3 | Bloodletting lancet, with safety retractable point, for use independent of any other device. For use in neonates Size: 28G Sterile, Single use Outer package must be labelled sterile | Offer 1 of 1 | SSEM Mthembu Medical (Pty) Ltd | 0.91 | Solace | LOSL28.2.2 | Solace | Sweden | 56 | 85% | USD | 1 | Yes | 100/box | COMPLY pending confirmation of latex free |  |  | 46.0784
41104102-00004 | 3 | Bloodletting lancet, with safety retractable point, for use independent of any other device. For use in neonates Size: 28G Sterile, Single use Outer package must be labelled sterile | OFFER 1 | SUPRA HEALTHCARE JOHANNESBURG PTY LTD | 1.01 | Avacare | LANCESAF004 | HEALTHY MEDICAL DEVELOPMENT CO., LTD | CHINA | Initial 60, thereafter 14 | 84 | USD | 1 Box | YES | 100 pcs/box | No brand name. No product code. |  |  | 46.0784
41104102-00004 | 3 | Bloodletting lancet, with safety retractable point, for use independent of any other device. For use in neonates Size: 28G Sterile, Single use Outer package must be labelled sterile |  | AKACIA MEDICAL (PTY) LTD | 1.02 | AKACIA | MI006405 | HTL STREFA | POLAND | EX STOCK -20DAYS | 67% | USD | BOX 100 | YES | N/A | COMPLY |  |  | 6.4
41104102-00004 | 3 | Bloodletting lancet, with safety retractable point, for use independent of any other device. For use in neonates Size: 28G Sterile, Single use Outer package must be labelled sterile | OFFER 2 | SUPRA HEALTHCARE JOHANNESBURG PTY LTD | 1.13 | Avacare | LANCESAF004 | HENAN AILE INDUSTRIAL CO.,LTD | CHINA | Initial 60, thereafter 14 | 84 | USD | 1 Box | YES | 100 pcs/box | No brand name. No product code. |  |  | 100
41104102-00004 | 3 | Bloodletting lancet, with safety retractable point, for use independent of any other device. For use in neonates Size: 28G Sterile, Single use Outer package must be labelled sterile | N/A | SAFARMEX (PTY) LTD | 1.14 | Saclin | To be confirmed | Zhejiang Mediunion Healthcare Group Co. Ltd | China | 3days | 80% | US$ | 50 000 pieces | Yes | Box of 100 | No sample submitted for evaluation |  |  | 91
41104102-00004 | 3 | Bloodletting lancet, with safety retractable point, for use independent of any other device. For use in neonates Size: 28G Sterile, Single use Outer package must be labelled sterile | Basic Line offer | Becton Dickinson | 4.75 | BD VAC QUICKHEEL, PINK | 368102 | Becton Dickinson | United Kingdom | 7 Working Days | 76% | USD | 1 BOX | NO | 1 BOX = 50 UNITS / Blade Lancet, Depth 0.8mm, Width/Gauge 1.75mm | To request the company to confirm the size of 0.85mm X 1.75mm as the specs of 25G |  |  | -22.6087
41104102-00004 | 3 | Bloodletting lancet, with safety retractable point, for use independent of any other device. For use in neonates Size: 28G Sterile, Single use Outer package must be labelled sterile | BASIC LINE OFFER | MALACHITE MEDICAL (PTY) LTD | 0.48 | PRIONGUARD PRIMALANCE | LAN00128 | SUZHOU KYUAN MEDICAL | CHINA | 21 | 74.27 | USD | BOX 100 | YES | 40 BOXES OF 100 PER CASE | The brand name on the removable sticker is prionguarde and the brand name on the outer package is KYUAN. The reference number is hand written LAN00125. LOT number, shelf life and date of production of the sticker and outer package is not the same. |  |  | 100
42221512-00000 | 4 | Intravenous cannula, with introducer needle, radio-opaque, teflon/polyurethane, with double-tapered tip and stylet. 3-facetted J-point bevelled needle Length: Not less than 45mm Size: 14G Colour: Orange Flow rate: 320-330 ml/min Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 327 UNIT: Box (100 per box) |  | Chemical World cc | 2.19 | BIO ON | CIVNP14 | BIOMED | INDIA | 3 | 85 | USD | 1000 | YES | 1 | Offered a flow rate of 270 ml/min instead 320-330 ml/min | NO TEST REPORTS |  | 91
42221512-00000 | 4 | Intravenous cannula, with introducer needle, radio-opaque, teflon/polyurethane, with double-tapered tip and stylet. 3-facetted J-point bevelled needle Length: Not less than 45mm Size: 14G Colour: Orange Flow rate: 320-330 ml/min Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 327 UNIT: Box (100 per box) | OFFER 2 | SUPRA HEALTHCARE JOHANNESBURG PTY LTD | 2.44 | Vennula | IVC-14 | Romsons | INDIA | Initial 60, thereafter 14 | 84 | USD | 1 Box | YES | 100/BOX | Offered a flow rate of 270ml/min instead of 320-330ml/min | NO TEST REPORTS |  | 91
42221512-00000 | 4 | Intravenous cannula, with introducer needle, radio-opaque, teflon/polyurethane, with double-tapered tip and stylet. 3-facetted J-point bevelled needle Length: Not less than 45mm Size: 14G Colour: Orange Flow rate: 320-330 ml/min Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 327 UNIT: Box (100 per box) | Basic Line offer | Becton Dickinson | 500 | BD Insyte 14G | 381267 | Becton Dickinson | San Augustine | 7 Working Days | 76% | USD | 50 units / 1 BOX | NO | 50 units in a box | COMPLY | COMPLY | Recommended | 91
42221512-00000 | 4 | Intravenous cannula, with introducer needle, radio-opaque, teflon/polyurethane, with double-tapered tip and stylet. 3-facetted J-point bevelled needle Length: Not less than 45mm Size: 14G Colour: Orange Flow rate: 320-330 ml/min Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 327 UNIT: Box (100 per box) | Only one offer submitted | Tara Technology | 2.5 | Wellmed / Pencath | TAR812 | Wellmed | India | 21 days | 85% | USD | 10 | Yes | 100 per box | Offered a flow rate of 280ml/min instead of 320-330ml/min | NOT COMPLY |  | 55.7826
42221512-00000 | 4 | Intravenous cannula, with introducer needle, radio-opaque, teflon/polyurethane, with double-tapered tip and stylet. 3-facetted J-point bevelled needle Length: Not less than 45mm Size: 14G Colour: Orange Flow rate: 320-330 ml/min Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 327 UNIT: Box (100 per box) | A | Quality Medical Supplies Pty Ltd | 2.76 | Vennula | TBA | Romsons International | India | 14 | 80% | USD | 100 Units | Yes | Box | Offered a flow rate of 270ml/min instead of 320-330ml/min | NO TEST REPORTS |  | 91.2903
42221512-00000 | 4 | Intravenous cannula, with introducer needle, radio-opaque, teflon/polyurethane, with double-tapered tip and stylet. 3-facetted J-point bevelled needle Length: Not less than 45mm Size: 14G Colour: Orange Flow rate: 320-330 ml/min Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 327 UNIT: Box (100 per box) | Offer 1 of 1 | SSEM Mthembu Medical (Pty) Ltd | 2.76 | Polymed | PM121014 | Polymed | India | 56 | 85% | USD | 500 | Yes | 50/Box | 0ffered a flow rate of 305ml/min instead 320-330 ml/min | NOT COMPLY |  | 59.2353
42221512-00000 | 4 | Intravenous cannula, with introducer needle, radio-opaque, teflon/polyurethane, with double-tapered tip and stylet. 3-facetted J-point bevelled needle Length: Not less than 45mm Size: 14G Colour: Orange Flow rate: 320-330 ml/min Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 327 UNIT: Box (100 per box) |  | CRYSATL PIER TRADING 148cc | 2.98 | MEDIFLOW | [8 empty] | Each | 0ffered a flow rate of 300ml/min instead 320-330 ml/min | NOT COMPLY |  | 88.75
42221512-00000 | 4 | Intravenous cannula, with introducer needle, radio-opaque, teflon/polyurethane, with double-tapered tip and stylet. 3-facetted J-point bevelled needle Length: Not less than 45mm Size: 14G Colour: Orange Flow rate: 320-330 ml/min Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 327 UNIT: Box (100 per box) | OFFER 1 | SUPRA HEALTHCARE JOHANNESBURG PTY LTD | 3.28 | Healcath | IVC-14 | Harsoria | INDIA | Initial 60, thereafter 14 | 84 | USD | 1 Box | YES | 100/BOX | Offered a flow rate of 270ml/min instead of 320-330ml/min |  |  | 91
42221512-00000 | 4 | Intravenous cannula, with introducer needle, radio-opaque, teflon/polyurethane, with double-tapered tip and stylet. 3-facetted J-point bevelled needle Length: Not less than 45mm Size: 14G Colour: Orange Flow rate: 320-330 ml/min Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 327 UNIT: Box (100 per box) BOX (50 PER BOX) |  | JUMLA MEDICAL SUPPLIES (PTY) LTD | 4.01 | JELCO SERIVA | 8358 | Smiths Medical International | INDIA | 21 DAYS | 68.4% | BRITISH POUND | 500 | YES | NIL | Offered a flow rate of 305ml/min320-330 ml/min | NOT COMPLY |  | 36
42221512-00000 | 4 | Intravenous cannula, with introducer needle, radio-opaque, teflon/polyurethane, with double-tapered tip and stylet. 3-facetted J-point bevelled needle Length: Not less than 45mm Size: 14G Colour: Orange Flow rate: 320-330 ml/min Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 327 UNIT: Box (100 per box) | Alternative Offer | BIOCOR HOSPITAL SUPPLIES PTY LTD | 4.6 | Jelco Seriva | 8358 | Smith Medical | India | 21 | 85 | USD | 5 Boxes of 50 | Yes | N/A | offered a flow rate of 305ml/min instead 320-330 ml/min | NOT COMPLY |  | 93.2075
42221512-00000 | 4 | Intravenous cannula, with introducer needle, radio-opaque, teflon/polyurethane, with double-tapered tip and stylet. 3-facetted J-point bevelled needle Length: Not less than 45mm Size: 14G Colour: Orange Flow rate: 320-330 ml/min Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 327 UNIT: Box (100 per box) | BASIC LINE OFFER | A R FREIGHT PRIVATE LIMITED | 4.8 | JELCO SERIVA | 8358 | SMITHS MEDICAL | INDIA | 21 DAYS | 85% | USD | 100 | YES | N/A | Offered a flow rate of 305 ml/min instead of Flow rate:320-330 ml/min | NOT COMPLY |  | 100
42221512-00000 | 4 | Intravenous cannula, with introducer needle, radio-opaque, teflon/polyurethane, with double-tapered tip and stylet. 3-facetted J-point bevelled needle Length: Not less than 45mm Size: 14G Colour: Orange Flow rate: 320-330 ml/min Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 327 UNIT: Box (100 per box) | 1 | B. BRAUN MEDICAL (PTY) LTD | 4.8 | Introcan FEP 14G x 2", 2,2 x 50mm | 4252217B | B. Braun | Germany | 42 | 50% | Euro | 100 | Yes | Each | Offered a flow rate is 345ml/min instead of 320ml – 330ml/min | NOT COMPLY |  | 85.7895
42221512-00000 | 4 | Intravenous cannula, with introducer needle, radio-opaque, teflon/polyurethane, with double-tapered tip and stylet. 3-facetted J-point bevelled needle Length: Not less than 45mm Size: 14G Colour: Orange Flow rate: 320-330 ml/min Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 327 UNIT: Box (100 per box) | OPTIONAL OFFER OF 14G X 45MM WITH 305ML/MIN FLOWRATE | LECHOBA MEDICAL TECHNOLOGIES (PTY) LTD | 4.98 | CHIRAFLEX | CHFX0301445 | CHIRANA INJECTA A/S (PART OF CTI GROUP) | SLOVAK REPUBLIC | 21 DAYS | 75% | EURO | 100 | NO (FLOW RATE IS 305ML/MIN) | EACH (PACKED 50/BOX) | Offered a flow rate 305 ml/min instead 320-330 ml/min | NOT COMPLY |  | 91
42221512-00000 | 4 | Intravenous cannula, with introducer needle, radio-opaque, teflon/polyurethane, with double-tapered tip and stylet. 3-facetted J-point bevelled needle Length: Not less than 45mm Size: 14G Colour: Orange Flow rate: 320-330 ml/min Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 327 UNIT: Box (100 per box) | Basic Line Offer | BIOCOR HOSPITAL SUPPLIES PTY LTD | 5.75 | Jelco Radiopaque | 4038 | Smith Medical | Italy | 21 | 85 | USD | 5 Boxes of 50 | Yes | N/A | No sample submitted for evaluation | NOT COMPLY |  | 91
42221512-00000 | 4 | Intravenous cannula, with introducer needle, radio-opaque, teflon/polyurethane, with double-tapered tip and stylet. 3-facetted J-point bevelled needle Length: Not less than 45mm Size: 14G Colour: Orange Flow rate: 320-330 ml/min Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 327 UNIT: Box (100 per box) |  | ALMIKA TRADING PTY LTD | 5.95 | BD | 381267 | BD | USA | 14 | 0 | NO | 100 | YES | 1 | COMPLY | COMPLY |  | 57.6471
42221512-00000 | 4 | Intravenous cannula, with introducer needle, radio-opaque, teflon/polyurethane, with double-tapered tip and stylet. 3-facetted J-point bevelled needle Length: Not less than 45mm Size: 14G Colour: Orange Flow rate: 320-330 ml/min Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 327 UNIT: Box (100 per box) |  | BAS MEDEXPRESS | 9.78 | Bicakcilap | 14614011 |  | Turkey | 1 - 3 days | 15% | N/A | As per client | Yes | Each | Offered a flow rate of 58ml/min instead of 320-330ml/min. Offered 20G instead of 14G. Offered 32mm instead of not less than 33mm |  |  | -440
42221512-00001 | 5 | Intravenous cannula, with introducer needle, radio-opaque, teflon/polyurethane, with double-tapered tip and stylet. 3-facetted J-point bevelled needle Length: Not less than 45mm Size: 16G Colour: Grey Flow rate: 210-220 ml/min Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 327 UNIT: Box (100 per box) | Only one offer submitted | Tara Technology | 2.05 | Wellmed / Pencath | TAR813 | Wellmed | India | 21 days | 85% | USD | 10 | Yes | 100 per box | Offered a flow rate of 200 ml/min instead of Flow rate:210-220 ml/min. | NOT COMPLY |  | -359
42221512-00001 | 5 | Intravenous cannula, with introducer needle, radio-opaque, teflon/polyurethane, with double-tapered tip and stylet. 3-facetted J-point bevelled needle Length: Not less than 45mm Size: 16G Colour: Grey Flow rate: 210-220 ml/min Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 327 UNIT: Box (100 per box) |  | Chemical World cc | 2.14 | BIO ON | CIVNP16 | BIOMED | INDIA | 3 | 85 | USD | 1000 | YES | 1 | Offered a flow rate of 180 ml/min instead of Flow rate:210-220 ml/min | NOT COMPLY |  | 25.6094
42221512-00001 | 5 | Intravenous cannula, with introducer needle, radio-opaque, teflon/polyurethane, with double-tapered tip and stylet. 3-facetted J-point bevelled needle Length: Not less than 45mm Size: 16G Colour: Grey Flow rate: 210-220 ml/min Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 327 UNIT: Box (100 per box) | OFFER 2 | SUPRA HEALTHCARE JOHANNESBURG PTY LTD | 2.3 | Vennula | IVC-16 | Romsons | INDIA | Initial 60, thereafter 14 | 84 | USD | 1 Box | YES | 100/BOX | Offered a flow rate of 180 ml/min instead of Flow rate:210-220 ml/min. | NO TEST REPORTS |  | -656.7941
42221512-00001 | 5 | Intravenous cannula, with introducer needle, radio-opaque, teflon/polyurethane, with double-tapered tip and stylet. 3-facetted J-point bevelled needle Length: Not less than 45mm Size: 16G Colour: Grey Flow rate: 210-220 ml/min Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 327 UNIT: Box (100 per box) | Basic Line offer | Becton Dickinson | 2.5 | BD Insyte 16G | 381257 | Becton Dickinson | San Augustine | 7 working day | 76% | USD | 50 units / 1 BOX | NO | 50 units in a box | Offered a flow rate of 205 ml/min instead of Flow rate:210-220 ml/min. | NOT COMPLY |  | 94.375
42221512-00001 | 5 | Intravenous cannula, with introducer needle, radio-opaque, teflon/polyurethane, with double-tapered tip and stylet. 3-facetted J-point bevelled needle Length: Not less than 45mm Size: 16G Colour: Grey Flow rate: 210-220 ml/min Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 327 UNIT: Box (100 per box) | A | Quality Medical Supplies Pty Ltd | 2.76 | Vennula | TBA | Romsons International | India | 14 | 80% | USD | 100 Units | Yes | Box | Offered a flow rate of 180 ml/min instead of Flow rate:210-220 ml/min. | NO TEST REPORTS |  | -656.7941
42221512-00001 | 5 | Intravenous cannula, with introducer needle, radio-opaque, teflon/polyurethane, with double-tapered tip and stylet. 3-facetted J-point bevelled needle Length: Not less than 45mm Size: 16G Colour: Grey Flow rate: 210-220 ml/min Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 327 UNIT: Box (100 per box) |  | CRYSATL PIER TRADING 148cc | 2.76 | MEDIFLOW |  | GLOBAL MEDIKIT LIMITED C/O SR MEDICAL CARE | [6 empty] | EACH | Offered a flow rate of 200 ml/min instead of Flow rate:210-220 ml/min. | NOT COMPLY |  | 94.375
42221512-00001 | 5 | Intravenous cannula, with introducer needle, radio-opaque, teflon/polyurethane, with double-tapered tip and stylet. 3-facetted J-point bevelled needle Length: Not less than 45mm Size: 16G Colour: Grey Flow rate: 210-220 ml/min Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 327 UNIT: Box (100 per box) | Offer 1 of 1 | SSEM Mthembu Medical (Pty) Ltd | 2.76 | Polymed | PM121016 | Polymed | India | 56 | 85% | USD | 500 | Yes | 50/Box | Offered a flow rate of 200 ml/min instead of Flow rate:210-220 ml/min. | NOT COMPLY |  | -695.4615
42221512-00001 | 5 | Intravenous cannula, with introducer needle, radio-opaque, teflon/polyurethane, with double-tapered tip and stylet. 3-facetted J-point bevelled needle Length: Not less than 45mm Size: 16G Colour: Grey Flow rate: 210-220 ml/min Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 327 UNIT: Box (100 per box) | OFFER 1 | SUPRA HEALTHCARE JOHANNESBURG PTY LTD | 3.03 | Healcath | IVC-16 | Harsoria | INDIA | Initial 60, thereafter 14 | 84 | USD | 1 Box | YES | 100/BOX | Offered a flow rate of 200 ml/min instead of Flow rate:210-220 ml/min. | NOT COMPLY |  | 91
42221512-00001 | 5 | Intravenous cannula, with introducer needle, radio-opaque, teflon/polyurethane, with double-tapered tip and stylet. 3-facetted J-point bevelled needle Length: Not less than 45mm Size: 16G Colour: Grey Flow rate: 210-220 ml/min Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 327 UNIT: Box (100 per box) BOX (50 PER BOX) |  | JUMLA MEDICAL SUPPLIES (PTY) LTD | 4.01 | JELCO SERIVA | 8351 | Smiths Medical International | INDIA | 21 DAYS | 68.4% | BRITISH POUND | 500 | YES | NIL | Offered a flow rate of 200 ml/min instead of Flow rate:210-220 ml/min. | NOT COMPLY |  | -695.4615
42221512-00001 | 5 | Intravenous cannula, with introducer needle, radio-opaque, teflon/polyurethane, with double-tapered tip and stylet. 3-facetted J-point bevelled needle Length: Not less than 45mm Size: 16G Colour: Grey Flow rate: 210-220 ml/min Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 327 UNIT: Box (100 per box) | Alternative Offer | BIOCOR HOSPITAL SUPPLIES PTY LTD | 4.6 | Jelco Seriva | 8351 | Smith Medical | India | 21 | 85 | USD | 5 Boxes of 50 | Yes | N/A | Offered a flow rate of 200 ml/min instead of Flow rate:210-220 ml/min | NOT COMPLY |  | 91
42221512-00001 | 5 | Intravenous cannula, with introducer needle, radio-opaque, teflon/polyurethane, with double-tapered tip and stylet. 3-facetted J-point bevelled needle Length: Not less than 45mm Size: 16G Colour: Grey Flow rate: 210-220 ml/min Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 327 UNIT: Box (100 per box) | BASIC LINE OFFER | A R FREIGHT PRIVATE LIMITED | 4.8 | JELCO SERIVA | 8351 | SMITHS MEDICAL | INDIA | 21 DAYS | 85% | USD | 100 | YES | N/A | Offered a flow rate of 200 ml/min instead of Flow rate:210-220 ml/min | NOT COMPLY |  | -695.4615
42221512-00001 | 5 | Intravenous cannula, with introducer needle, radio-opaque, teflon/polyurethane, with double-tapered tip and stylet. 3-facetted J-point bevelled needle Length: Not less than 45mm Size: 16G Colour: Grey Flow rate: 210-220 ml/min Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 327 UNIT: Box (100 per box) | 1 | B. BRAUN MEDICAL (PTY) LTD | 480 | Introcan FEP 16G x 2", 1,7 x 50mm | 4252160B | B. Braun | Germany | 42 | 50% | Euro | 100 | Yes | 100 per box | COMPLY | COMPLY | Recommended | 95
42221512-00001 | 5 | Intravenous cannula, with introducer needle, radio-opaque, teflon/polyurethane, with double-tapered tip and stylet. 3-facetted J-point bevelled needle Length: Not less than 45mm Size: 16G Colour: Grey Flow rate: 210-220 ml/min Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 327 UNIT: Box (100 per box) | OPTIONAL OFFER OF 16G X 45MM WITH 305ML/MIN FLOWRATE | LECHOBA MEDICAL TECHNOLOGIES (PTY) LTD | 4.98 | CHIRAFLEX | CHFX0301645 | CHIRANA INJECTA A/S (PART OF CTI GROUP) | SLOVAK REPUBLIC | 21 DAYS | 75% | EURO | 100 | NO (FLOW RATE IS 200ML/MIN) | EACH (PACKED 50/BOX) | Offered a flow rate of 200 ml/min instead of Flow rate:210-220 ml/min. | NOT COMPLY |  | 91
42221512-00001 | 5 | Intravenous cannula, with introducer needle, radio-opaque, teflon/polyurethane, with double-tapered tip and stylet. 3-facetted J-point bevelled needle Length: Not less than 45mm Size: 16G Colour: Grey Flow rate: 210-220 ml/min Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 327 UNIT: Box (100 per box) | Basic Line Offer | BIOCOR HOSPITAL SUPPLIES PTY LTD | 5.75 | Jelco Radiopaque | 4034 | Smith | Italy | 21 | 85 | USD | 5 Boxes of 50 | Yes | N/A | No sample submitted for evaluation | NOT COMPLY |  | 91
42221512-00001 | 5 | Intravenous cannula, with introducer needle, radio-opaque, teflon/polyurethane, with double-tapered tip and stylet. 3-facetted J-point bevelled needle Length: Not less than 45mm Size: 16G Colour: Grey Flow rate: 210-220 ml/min Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 327 UNIT: Box (100 per box) |  | ALMIKA TRADING PTY LTD | 5.95 | BD | 381257 | BD | USA | 14 | 0 | NO | 100 | YES | 1 | Offered a flow rate of 205ml/min instead of 210-220ml/min. | NOT COMPLY |  | 100
42221512-00001 | 5 | Intravenous cannula, with introducer needle, radio-opaque, teflon/polyurethane, with double-tapered tip and stylet. 3-facetted J-point bevelled needle Length: Not less than 45mm Size: 16G Colour: Grey Flow rate: 210-220 ml/min Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 327 UNIT: Box (100 per box) |  | BAS MEDEXPRESS | 7.75 | Bicakcilap | 14616011 |  | Turkey | 1 - 3 days | 15% | N/A | As per client | Yes | Each | Offered a flow rate of 100 ml/min instead of Flow rate:210-220 ml/min. Offered 18G instead of 16G |  |  | 100
42221512-00002 | 6 | Intravenous cannula, with introducer needle, radio-opaque, teflon/polyurethane, with double-tapered tip and stylet. 3-facetted J-point bevelled needle 3-facetted J-point bevelled needle Size: 18G Colour: Green Flow rate: 105-115 ml/min Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 327 UNIT: Box (100 per box) | Only one offer submitted | Tara Technology | 1.95 | Wellmed / Pencath | TAR814 | Wellmed | India | 21 days | 85% | USD | 10 | Yes | 100 per box | Offered a flow rate of 95 ml/min instead of Flow rate:105-115 ml/min. | NOT COMPLY |  | 100
42221512-00002 | 6 | Intravenous cannula, with introducer needle, radio-opaque, teflon/polyurethane, with double-tapered tip and stylet. 3-facetted J-point bevelled needle 3-facetted J-point bevelled needle Size: 18G Colour: Green Flow rate: 105-115 ml/min Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 327 UNIT: Box (100 per box) |  | Chemical World cc | 1.99 | BIO ON | CIVNP18 | BIOMED | INDIA | 3 | 85 | USD | 1000 | YES | 1 | Offered a flow rate of 80 ml/min instead of Flow rate:105-115 ml/min. | NOT COMPLY |  | 100
42221512-00002 | 6 | Intravenous cannula, with introducer needle, radio-opaque, teflon/polyurethane, with double-tapered tip and stylet. 3-facetted J-point bevelled needle 3-facetted J-point bevelled needle Size: 18G Colour: Green Flow rate: 105-115 ml/min Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 327 UNIT: Box (100 per box) | OFFER 2 | SUPRA HEALTHCARE JOHANNESBURG PTY LTD | 2.3 | Vennula | IVC-18 | Romsons | INDIA | Initial 60, thereafter 14 | 84 | USD | 1 Box | YES | 100/BOX | Offered a flow rate of 90 ml/min instead of Flow rate:105-115 ml/min. | NO TEST REPORTS
42221512-00002 | 6 | Intravenous cannula, with introducer needle, radio-opaque, teflon/polyurethane, with double-tapered tip and stylet. 3-facetted J-point bevelled needle 3-facetted J-point bevelled needle Size: 18G Colour: Green Flow rate: 105-115 ml/min Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 327 UNIT: Box (100 per box) | Offer 1 of 1 | SSEM Mthembu Medical (Pty) Ltd | 2.47 | Polymed | PM121018 | Polymed | India | 56 | 85% | USD | 500 | Yes | 50/Box | Offered a flow rate of 95 ml/min instead of Flow rate:105-115 ml/min. | NOT COMPLY
42221512-00002 | 6 | Intravenous cannula, with introducer needle, radio-opaque, teflon/polyurethane, with double-tapered tip and stylet. 3-facetted J-point bevelled needle 3-facetted J-point bevelled needle Size: 18G Colour: Green Flow rate: 105-115 ml/min Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 327 UNIT: Box (100 per box) | Basic Line offer | Becton Dickinson | 2.75 | BD Insyte 18G | 381246 | Becton Dickinson | San Augustine | 7 working days | 76% | USD | 50 units / 1 BOX | NO | 50 units in a box | No sample submitted for evaluation. | NO TEST REPORT
42221512-00002 | 6 | Intravenous cannula, with introducer needle, radio-opaque, teflon/polyurethane, with double-tapered tip and stylet. 3-facetted J-point bevelled needle 3-facetted J-point bevelled needle Size: 18G Colour: Green Flow rate: 105-115 ml/min Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 327 UNIT: Box (100 per box) | A | Quality Medical Supplies Pty Ltd | 2.76 | Vennula | TBA | Romsons International | India | 14 | 80% | USD | 100 Units | Yes | Box | Offered a flow rate of 90 ml/min instead of Flow rate:105-115 ml/min. | NOT COMPLY
42221512-00002 | 6 | Intravenous cannula, with introducer needle, radio-opaque, teflon/polyurethane, with double-tapered tip and stylet. 3-facetted J-point bevelled needle 3-facetted J-point bevelled needle Size: 18G Colour: Green Flow rate: 105-115 ml/min Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 327 UNIT: Box (100 per box) |  | CRYSATL PIER TRADING 148cc | 2.76 | MEDIFLOW/SR CARE |  | GLOBAL MEDIKIT LIMITED C/O SR MEDICAL CARE | [6 empty] | EACH | Offered a flow rate of 90 ml/min instead of Flow rate:105-115 ml/min. | NOT COMPLY
42221512-00002 | 6 | Intravenous cannula, with introducer needle, radio-opaque, teflon/polyurethane, with double-tapered tip and stylet. 3-facetted J-point bevelled needle 3-facetted J-point bevelled needle Size: 18G Colour: Green Flow rate: 105-115 ml/min Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 327 UNIT: Box (100 per box) | OFFER 1 | SUPRA HEALTHCARE JOHANNESBURG PTY LTD | 2.78 | Healcath | IVC-18 | Harsoria | INDIA | Initial 60, thereafter 14 | 84 | USD | 1 Box | YES | 100/BOX | Offered a flow rate of 85 ml/min instead of Flow rate:105-115 ml/min. | NOT COMPLY
42221512-00002 | 6 | Intravenous cannula, with introducer needle, radio-opaque, teflon/polyurethane, with double-tapered tip and stylet. 3-facetted J-point bevelled needle 3-facetted J-point bevelled needle Size: 18G Colour: Green Flow rate: 105-115 ml/min Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 327 UNIT: Box (100 per box) | EACH | MSIMED CC | 3.45 | KANG KANG | CIV01 | ZHENJIANG | CHINA | 25 | 20% | USD | 20000 PCS | YES | EACH | No sample submitted for evaluation | NO TEST REPORTS
42221512-00002 | 6 | Intravenous cannula, with introducer needle, radio-opaque, teflon/polyurethane, with double-tapered tip and stylet. 3-facetted J-point bevelled needle 3-facetted J-point bevelled needle Size: 18G Colour: Green Flow rate: 105-115 ml/min Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 327 UNIT: Box (100 per box) BOX (50 PER BOX) |  | JUMLA MEDICAL SUPPLIES (PTY) LTD | 4.01 | JELCO SERIVA | 8354 | Smiths Medical International | INDIA | 21 DAYS | 68.4% | BRITISH POUND | 500 | YES | NIL | Offered a flow rate of 90 ml/min instead of Flow rate:105-115 ml/min. | NOT COMPLY
42221512-00002 | 6 | Intravenous cannula, with introducer needle, radio-opaque, teflon/polyurethane, with double-tapered tip and stylet. 3-facetted J-point bevelled needle 3-facetted J-point bevelled needle Size: 18G Colour: Green Flow rate: 105-115 ml/min Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 327 UNIT: Box (100 per box) | Alternative Offer | BIOCOR HOSPITAL SUPPLIES PTY LTD | 4.6 | Jelco Seriva | 8354 | Smith Medical | India | 21 | 85 | USD | 5 Boxes of 50 | Yes | N/A | Offered a flow rate of 90 ml/min instead of Flow rate:105-115 ml/min. | NOT COMPLY
42221512-00002 | 6 | Intravenous cannula, with introducer needle, radio-opaque, teflon/polyurethane, with double-tapered tip and stylet. 3-facetted J-point bevelled needle 3-facetted J-point bevelled needle Size: 18G Colour: Green Flow rate: 105-115 ml/min Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 327 UNIT: Box (100 per box) | BASIC LINE OFFER | A R FREIGHT PRIVATE LIMITED | 4.8 | JELCO SERIVA | 8354 | SMITHS MEDICAL | INDIA | 21 DAYS | 85% | USD | 100 | YES | N/A | Offered a flow rate of 90 ml/min instead of Flow rate:105-115 ml/min. | NOT COMPLY
42221512-00002 | 6 | Intravenous cannula, with introducer needle, radio-opaque, teflon/polyurethane, with double-tapered tip and stylet. 3-facetted J-point bevelled needle 3-facetted J-point bevelled needle Size: 18G Colour: Green Flow rate: 105-115 ml/min Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 327 UNIT: Box (100 per box) | 1 | B. BRAUN MEDICAL (PTY) LTD | 480 | Introcan FEP 18G x 1 3/4, 1,3 x 45mm | 4252136B | B. Braun | Germany | 42 | 50% | Euro | 100 | Yes | 100 per box | Offered a flow rate of 100 ml/min instead of Flow rate:105-115 ml/min. | COMPLY | Recommended with deviation | 95
42221512-00002 | 6 | Intravenous cannula, with introducer needle, radio-opaque, teflon/polyurethane, with double-tapered tip and stylet. 3-facetted J-point bevelled needle 3-facetted J-point bevelled needle Size: 18G Colour: Green Flow rate: 105-115 ml/min Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 327 UNIT: Box (100 per box) | OPTIONAL OFFER OF 18G X 45MM WITH 305ML/MIN FLOWRATE | LECHOBA MEDICAL TECHNOLOGIES (PTY) LTD | 4.98 | CHIRAFLEX | CHFX0301845 | CHIRANA INJECTA A/S (PART OF CTI GROUP) | SLOVAK REPUBLIC | 21 DAYS | 75% | EURO | 100 | NO (FLOW RATE IS 90ML/MIN) | EACH (PACKED 50/BOX) | Offered a flow rate of 90 ml/min instead of Flow rate:105-115 ml/min. | NOT COMPLY
42221512-00002 | 6 | Intravenous cannula, with introducer needle, radio-opaque, teflon/polyurethane, with double-tapered tip and stylet. 3-facetted J-point bevelled needle 3-facetted J-point bevelled needle Size: 18G Colour: Green Flow rate: 105-115 ml/min Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 327 UNIT: Box (100 per box) | Basic Line Offer | BIOCOR HOSPITAL SUPPLIES PTY LTD | 5.75 | Jelco Radiopaque | 4032 | Smith Medical | Italy | 21 | 85 | USD | 5 Boxes of 50 | Yes | N/A | No sample submitted for evaluation | NOT COMPLY
42221512-00002 | 6 | Intravenous cannula, with introducer needle, radio-opaque, teflon/polyurethane, with double-tapered tip and stylet. 3-facetted J-point bevelled needle 3-facetted J-point bevelled needle Size: 18G Colour: Green Flow rate: 105-115 ml/min Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 327 UNIT: Box (100 per box) |  | ALMIKA TRADING PTY LTD | 5.95 | BD | 381246 | BD | USA | 14 | 0 | NO | 100 | YES | 1 | No sample submitted for evaluation. | NO TEST REPORTS
42221512-00002 | 6 | Intravenous cannula, with introducer needle, radio-opaque, teflon/polyurethane, with double-tapered tip and stylet. 3-facetted J-point bevelled needle 3-facetted J-point bevelled needle Size: 18G Colour: Green Flow rate: 105-115 ml/min Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 327 UNIT: Box (100 per box) |  | BAS MEDEXPRESS | 7.32 | Bicakcilap | 1468011 |  | Turkey | 1 - 3 days | 15% | N/A | As per client | Yes | Each | Offered a flow rate of 100ml/min instead of 105-115 ml/min
42221512-00003 | 7 | Intravenous cannula, with introducer needle, radio-opaque, teflon/polyurethane, with double-tapered tip and stylet. 3-facetted J-point bevelled needle kLength: Not less than 33mm Size: 20G Colour: Pink Flow rate: 60-66 ml/min Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 327 UNIT: Box (100 per box) |  | CRYSATL PIER TRADING 148cc | 276 | MEDIFLOW |  | SRM 0003 | [6 empty] | 100 per box | COMPLY | COMPLY | Recommended | 100
42221512-00003 | 7 | Intravenous cannula, with introducer needle, radio-opaque, teflon/polyurethane, with double-tapered tip and stylet. 3-facetted J-point bevelled needle kLength: Not less than 33mm Size: 20G Colour: Pink Flow rate: 60-66 ml/min Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 327 UNIT: Box (100 per box) | Only one offer submitted | Tara Technology | 1.95 | Wellmed / Pencath | TAR815 | Wellmed | India | 21 days | 85% | USD | 10 | Yes | 100 per box | Offered a length of 32mm instead of not less than 33ml | NOT COMPLY
42221512-00003 | 7 | Intravenous cannula, with introducer needle, radio-opaque, teflon/polyurethane, with double-tapered tip and stylet. 3-facetted J-point bevelled needle kLength: Not less than 33mm Size: 20G Colour: Pink Flow rate: 60-66 ml/min Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 327 UNIT: Box (100 per box) |  | Chemical World cc | 1.96 | BIO ON | CIVNP20 | BIOMED | INDIA | 3 | 85 | USD | 1000 | YES | 1 | Offered a flow rate of 54 ml/min instead of Flow rate:60-66 ml/min | NOT COMPLY
42221512-00003 | 7 | Intravenous cannula, with introducer needle, radio-opaque, teflon/polyurethane, with double-tapered tip and stylet. 3-facetted J-point bevelled needle kLength: Not less than 33mm Size: 20G Colour: Pink Flow rate: 60-66 ml/min Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 327 UNIT: Box (100 per box) | Basic Line offer | Becton Dickinson | 2 | BD Insyte 20G | 381234 | Becton Dickinson | San Augustine | 7 working days | 76% | USD | 50 units / 1 BOX | NO | 50 units in a box | Offered a length of 30mm instead of not less than 33mm | NOT COMPLY
42221512-00003 | 7 | Intravenous cannula, with introducer needle, radio-opaque, teflon/polyurethane, with double-tapered tip and stylet. 3-facetted J-point bevelled needle kLength: Not less than 33mm Size: 20G Colour: Pink Flow rate: 60-66 ml/min Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 327 UNIT: Box (100 per box) | OFFER 2 | SUPRA HEALTHCARE JOHANNESBURG PTY LTD | 2.3 | Vennula | IVC-20 | Romsons | INDIA | Initial 60, thereafter 14 | 84 | USD | 1 Box | YES | 100/BOX | Offered a flow rate of 55 ml/min instead of Flow rate:60-66 ml/min. Offered a length of 32mm instead of not less than 33mm | NO TEST REPORTS
42221512-00003 | 7 | Intravenous cannula, with introducer needle, radio-opaque, teflon/polyurethane, with double-tapered tip and stylet. 3-facetted J-point bevelled needle kLength: Not less than 33mm Size: 20G Colour: Pink Flow rate: 60-66 ml/min Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 327 UNIT: Box (100 per box) | Offer 1 of 1 | SSEM Mthembu Medical (Pty) Ltd | 2.47 | Polymed | PM121020 | Polymed | India | 56 | 85% | USD | 500 | Yes | 50/Box | Offered a length of 32mm instead of not less than 33mm | NOT COMPLY
42221512-00003 | 7 | Intravenous cannula, with introducer needle, radio-opaque, teflon/polyurethane, with double-tapered tip and stylet. 3-facetted J-point bevelled needle kLength: Not less than 33mm Size: 20G Colour: Pink Flow rate: 60-66 ml/min Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 327 UNIT: Box (100 per box) | A | Quality Medical Supplies Pty Ltd | 2.76 | Vennula | TBA | Romsons International | India | 14 | 80% | USD | 100 Units | Yes | Box | Offered a flow rate of 55 ml/min instead of Flow rate:60-66 ml/min. Offered a length of 32mm instead of not less than 33mm | NOT COMPLY
42221512-00003 | 7 | Intravenous cannula, with introducer needle, radio-opaque, teflon/polyurethane, with double-tapered tip and stylet. 3-facetted J-point bevelled needle kLength: Not less than 33mm Size: 20G Colour: Pink Flow rate: 60-66 ml/min Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 327 UNIT: Box (100 per box) | OFFER 1 | SUPRA HEALTHCARE JOHANNESBURG PTY LTD | 2.78 | Healcath | IVC-14 | Harsoria | INDIA | Initial 60, thereafter 14 | 84 | USD | 1 Box | YES | 100/BOX | Offered a flow rate of 55 ml/min instead of Flow rate:60-66 ml/min. Offered a length of 32mm instead of not less than 33mm | NOT COMPLY
42221512-00003 | 7 | Intravenous cannula, with introducer needle, radio-opaque, teflon/polyurethane, with double-tapered tip and stylet. 3-facetted J-point bevelled needle kLength: Not less than 33mm Size: 20G Colour: Pink Flow rate: 60-66 ml/min Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 327 UNIT: Box (100 per box) | EACH | MSIMED CC | 3.45 | KANG KANG | CIV02 | ZHENJIANG | CHINA | 25 | 20% | USD | 20000 PCS | YES | EACH | Offered a flow rate of 54 ml/min instead of Flow rate:60-66 ml/min. Offered a length of 32mm instead of not less than 33mm | NO TEST REPORTS
42221512-00003 | 7 | Intravenous cannula, with introducer needle, radio-opaque, teflon/polyurethane, with double-tapered tip and stylet. 3-facetted J-point bevelled needle kLength: Not less than 33mm Size: 20G Colour: Pink Flow rate: 60-66 ml/min Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 327 UNIT: Box (100 per box) BOX (50 PER BOX) |  | JUMLA MEDICAL SUPPLIES (PTY) LTD | 4.01 | JELCO SERIVA | 8356 | Smiths Medical International | INDIA | 21 DAYS | 68.4% | BRITISH POUND | 500 | YES | NIL | Offered a flow rate of 56 ml/min instead of Flow rate:60-66 ml/min Offered a length of 32mm instead of not less than 33mm | NOT COMPLY
42221512-00003 | 7 | Intravenous cannula, with introducer needle, radio-opaque, teflon/polyurethane, with double-tapered tip and stylet. 3-facetted J-point bevelled needle kLength: Not less than 33mm Size: 20G Colour: Pink Flow rate: 60-66 ml/min Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 327 UNIT: Box (100 per box) | Alternative Offer | BIOCOR HOSPITAL SUPPLIES PTY LTD | 4.6 | Jelco Seriva | 8356 | Smith Medical | India | 21 | 85 | USD | 5 Boxes Of 50 | Yes | N/A | Offered a flow rate of 56 ml/min instead of Flow rate:60-66 ml/min | NOT COMPLY
42221512-00003 | 7 | Intravenous cannula, with introducer needle, radio-opaque, teflon/polyurethane, with double-tapered tip and stylet. 3-facetted J-point bevelled needle kLength: Not less than 33mm Size: 20G Colour: Pink Flow rate: 60-66 ml/min Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 327 UNIT: Box (100 per box) |  | ALMIKA TRADING PTY LTD | 4.63 | BD | 381234 | BD | USA | 14 | 0 | NO | 100 | YES | 1 | Offered a length of 30mm instead of not less than 33mm | NOT COMPLY
42221512-00003 | 7 | Intravenous cannula, with introducer needle, radio-opaque, teflon/polyurethane, with double-tapered tip and stylet. 3-facetted J-point bevelled needle kLength: Not less than 33mm Size: 20G Colour: Pink Flow rate: 60-66 ml/min Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 327 UNIT: Box (100 per box) | BASIC LINE OFFER | A R FREIGHT PRIVATE LIMITED | 4.8 | JELCO SERIVA | 8356 | SMITHS MEDICAL | INDIA | 21 DAYS | 85% | USD | 100 | YES | N/A | Offered a flow rate of 56 ml/min instead of Flow rate:60-66 ml/min | NOT COMPLY
42221512-00003 | 7 | Intravenous cannula, with introducer needle, radio-opaque, teflon/polyurethane, with double-tapered tip and stylet. 3-facetted J-point bevelled needle kLength: Not less than 33mm Size: 20G Colour: Pink Flow rate: 60-66 ml/min Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 327 UNIT: Box (100 per box) | 1 | B. BRAUN MEDICAL (PTY) LTD | 4.8 | Introcan FEP 20G x 1 1/4, 1,1 x 32mm | 4252110B | B. Braun | Germany | 42 | 50% | Euro | 100 | Yes | Each | Offered a length of 32mm instead of not less than 33mm | NOT COMPLY
42221512-00003 | 7 | Intravenous cannula, with introducer needle, radio-opaque, teflon/polyurethane, with double-tapered tip and stylet. 3-facetted J-point bevelled needle kLength: Not less than 33mm Size: 20G Colour: Pink Flow rate: 60-66 ml/min Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 327 UNIT: Box (100 per box) |  | LECHOBA MEDICAL TECHNOLOGIES (PTY) LTD | 4.98 | CHIRAFLEX | CHFX03020 | CHIRANA INJECTA A/S (PART OF CTI GROUP) | SLOVAK REPUBLIC | 21 DAYS | 75% | EURO | 100 | YES | EACH (PACKED 50/BOX) | Offered a length of 32mm instead of not less than 33mm | NOT COMPLY
42221512-00003 | 7 | Intravenous cannula, with introducer needle, radio-opaque, teflon/polyurethane, with double-tapered tip and stylet. 3-facetted J-point bevelled needle kLength: Not less than 33mm Size: 20G Colour: Pink Flow rate: 60-66 ml/min Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 327 UNIT: Box (100 per box) | Basic Line Offer | BIOCOR HOSPITAL SUPPLIES PTY LTD | 5.75 | Jelco Radiopaque | 4036 | Smith Medical | Italy | 21 | 85 | USD | 5 Boxes of 50 | Yes | N/A | No sample submitted for evaluation | NOT COMPLY
42221512-00003 | 7 | Intravenous cannula, with introducer needle, radio-opaque, teflon/polyurethane, with double-tapered tip and stylet. 3-facetted J-point bevelled needle kLength: Not less than 33mm Size: 20G Colour: Pink Flow rate: 60-66 ml/min Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 327 UNIT: Box (100 per box) |  | BAS MEDEXPRESS | 7.32 | Bicakcilap | 14616011 |  | Turkey | 1 - 3 days | 15% | N/A | As per client | Yes | Each | Offered a flow rate of 55 ml/min instead of Flow rate:60-66 ml/min. Offered a length of 32mm instead of not less than 33mm
42221512-00004 | 8 | Intravenous cannula, with introducer needle, radio-opaque, teflon/polyurethane, with double-tapered tip and stylet. 3-facetted J-point bevelled needle Length: Not less than 25mm Size: 22G Colour: Blue Flow rate: 36-40 ml/min Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 327 UNIT: Box (100 per box) | Only one offer submitted | Tara Technology | 195 | Wellmed / Pencath | TAR816 | Wellmed | India | 21 days | 85% | USD | 10 | Yes | 100/Box | COMPLY | COMPLY | Recommended | 100
42221512-00004 | 8 | Intravenous cannula, with introducer needle, radio-opaque, teflon/polyurethane, with double-tapered tip and stylet. 3-facetted J-point bevelled needle Length: Not less than 25mm Size: 22G Colour: Blue Flow rate: 36-40 ml/min Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 327 UNIT: Box (100 per box) |  | Chemical World cc | 1.96 | BIO ON | CIVNP22 | BIOMED | INDIA | 3 | 85 | USD | 1000 | YES | 1 | Offered a flow rate of 33 ml/min instead of Flow rate:36-40 ml/min. | COMPLY
42221512-00004 | 8 | Intravenous cannula, with introducer needle, radio-opaque, teflon/polyurethane, with double-tapered tip and stylet. 3-facetted J-point bevelled needle Length: Not less than 25mm Size: 22G Colour: Blue Flow rate: 36-40 ml/min Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 327 UNIT: Box (100 per box) | Basic Line offer | Becton Dickinson | 2 | BD Insyte 22G | 381221 | Becton Dickinson | San Augustine | 7 working days | 76% | USD | 50 units / 1 BOX | NO | 50 units in a box | offered a flow rate of 35ml/min 36-40 ml/min | NOT COMPLY
42221512-00004 | 8 | Intravenous cannula, with introducer needle, radio-opaque, teflon/polyurethane, with double-tapered tip and stylet. 3-facetted J-point bevelled needle Length: Not less than 25mm Size: 22G Colour: Blue Flow rate: 36-40 ml/min Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 327 UNIT: Box (100 per box) | OFFER 2 | SUPRA HEALTHCARE JOHANNESBURG PTY LTD | 2.3 | Vennula | IVC-22 | Romsons | INDIA | Initial 60, thereafter 14 | 84 | USD | 1 Box | YES | 100/BOX | offered a flow rate of 31ml/min 36-40 ml/min | NO TEST REPORTS
42221512-00004 | 8 | Intravenous cannula, with introducer needle, radio-opaque, teflon/polyurethane, with double-tapered tip and stylet. 3-facetted J-point bevelled needle Length: Not less than 25mm Size: 22G Colour: Blue Flow rate: 36-40 ml/min Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 327 UNIT: Box (100 per box) | Offer 1 of 1 | SSEM Mthembu Medical (Pty) Ltd | 2.47 | Polymed | PM121022 | Polymed | India | 56 | 85% | USD | 500 | Yes | 50/Box | COMPLY pending confirmation of latex free | COMPLY
42221512-00004 | 8 | Intravenous cannula, with introducer needle, radio-opaque, teflon/polyurethane, with double-tapered tip and stylet. 3-facetted J-point bevelled needle Length: Not less than 25mm Size: 22G Colour: Blue Flow rate: 36-40 ml/min Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 327 UNIT: Box (100 per box) |  | CRYSATL PIER TRADING 148cc | 2.76 | MEDIFLOW/SR CARE |  | GLOBAL MEDIKIT LIMITED C/O SR MEDICAL CARE | [6 empty] | EACH | COMPLY pending confirmation of latex free | COMPLY
42221512-00004 | 8 | Intravenous cannula, with introducer needle, radio-opaque, teflon/polyurethane, with double-tapered tip and stylet. 3-facetted J-point bevelled needle Length: Not less than 25mm Size: 22G Colour: Blue Flow rate: 36-40 ml/min Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 327 UNIT: Box (100 per box) | A | Quality Medical Supplies Pty Ltd | 2.76 | Vennula | TBA | Romsons International | India | 14 | 80% | USD | 100 Units | Yes | Box | offered a flow rate of 31ml/mininstead of 36-40 ml/min | NOT COMPLY
42221512-00004 | 8 | Intravenous cannula, with introducer needle, radio-opaque, teflon/polyurethane, with double-tapered tip and stylet. 3-facetted J-point bevelled needle Length: Not less than 25mm Size: 22G Colour: Blue Flow rate: 36-40 ml/min Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 327 UNIT: Box (100 per box) | OFFER 1 | SUPRA HEALTHCARE JOHANNESBURG PTY LTD | 2.78 | Healcath | IVC-22 | Harsoria | INDIA | Initial 60, thereafter 14 | 84 | USD | 1 Box | YES | 100/BOX | offered a flow rate of 33ml/min 36-40 ml/min | NOT COMPLY
42221512-00004 | 8 | Intravenous cannula, with introducer needle, radio-opaque, teflon/polyurethane, with double-tapered tip and stylet. 3-facetted J-point bevelled needle Length: Not less than 25mm Size: 22G Colour: Blue Flow rate: 36-40 ml/min Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 327 UNIT: Box (100 per box) | EACH | MSIMED CC | 3.45 | KANG KANG | CIV03 | ZHENJIANG | CHINA | 25 | 20% | USD | 20000 PCS | YES | EACH | offered a flow rate of 33ml/mininstead of 36-40 ml/min | NO TEST REPORTS
42221512-00004 | 8 | Intravenous cannula, with introducer needle, radio-opaque, teflon/polyurethane, with double-tapered tip and stylet. 3-facetted J-point bevelled needle Length: Not less than 25mm Size: 22G Colour: Blue Flow rate: 36-40 ml/min Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 327 UNIT: Box (100 per box) BOX (50 PER BOX) |  | JUMLA MEDICAL SUPPLIES (PTY) LTD | 4.01 | JELCO SERIVA | 8357 | Smiths Medical International | INDIA | 21 DAYS | 68.4% | BRITISH POUND | 500 | YES | NIL | COMPLY pending confirmation of latex free | COMPLY
42221512-00004 | 8 | Intravenous cannula, with introducer needle, radio-opaque, teflon/polyurethane, with double-tapered tip and stylet. 3-facetted J-point bevelled needle Length: Not less than 25mm Size: 22G Colour: Blue Flow rate: 36-40 ml/min Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 327 UNIT: Box (100 per box) | Alternative Offer | BIOCOR HOSPITAL SUPPLIES PTY LTD | 4.6 | Jelco Seriva | 8357 | Smith Medical | India | 21 | 85 | USD | 5 Boxes of Box | Yes | N/A | COMPLY | COMPLY
42221512-00004 | 8 | Intravenous cannula, with introducer needle, radio-opaque, teflon/polyurethane, with double-tapered tip and stylet. 3-facetted J-point bevelled needle Length: Not less than 25mm Size: 22G Colour: Blue Flow rate: 36-40 ml/min Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 327 UNIT: Box (100 per box) |  | ALMIKA TRADING PTY LTD | 4.63 | BD | 381223 | BD | USA | 14 | 0 | NO | 100 | YES | 1 | Offered a flow rate of 35 ml/min instead of Flow rate:36-40 ml/min. | NOT COMPLY
42221512-00004 | 8 | Intravenous cannula, with introducer needle, radio-opaque, teflon/polyurethane, with double-tapered tip and stylet. 3-facetted J-point bevelled needle Length: Not less than 25mm Size: 22G Colour: Blue Flow rate: 36-40 ml/min Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 327 UNIT: Box (100 per box) | BASIC LINE OFFER | A R FREIGHT PRIVATE LIMITED | 4.8 | JELCO SERIVA | 8357 | SMITHS MEDICAL | INDIA | 21 DAYS | 85% | USD | 100 | YES | N/A | COMPLY | COMPLY
42221512-00004 | 8 | Intravenous cannula, with introducer needle, radio-opaque, teflon/polyurethane, with double-tapered tip and stylet. 3-facetted J-point bevelled needle Length: Not less than 25mm Size: 22G Colour: Blue Flow rate: 36-40 ml/min Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 327 UNIT: Box (100 per box) | 1 | B. BRAUN MEDICAL (PTY) LTD | 4.8 | Introcan FEP 22G x 1", 0,9 x 25mm | 4252098B | B. Braun | Germany | 42 | 50% | Euro | 100 | Yes | Each | Offered a flow rate of 35 ml/min instead of Flow rate:36-40 ml/min. | NOT COMPLY
42221512-00004 | 8 | Intravenous cannula, with introducer needle, radio-opaque, teflon/polyurethane, with double-tapered tip and stylet. 3-facetted J-point bevelled needle Length: Not less than 25mm Size: 22G Colour: Blue Flow rate: 36-40 ml/min Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 327 UNIT: Box (100 per box) |  | LECHOBA MEDICAL TECHNOLOGIES (PTY) LTD | 4.98 | CHIRAFLEX | CHFX03022 | CHIRANA INJECTA A/S (PART OF CTI GROUP) | SLOVAK REPUBLIC | 21 DAYS | 75% | EURO | 100 | YES | EACH (PACKED 50/BOX) | COMPLY | NOT COMPLY
42221512-00004 | 8 | Intravenous cannula, with introducer needle, radio-opaque, teflon/polyurethane, with double-tapered tip and stylet. 3-facetted J-point bevelled needle Length: Not less than 25mm Size: 22G Colour: Blue Flow rate: 36-40 ml/min Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 327 UNIT: Box (100 per box) | Basic Line Offer | BIOCOR HOSPITAL SUPPLIES PTY LTD | 5.75 | Jelco Radiopaque | 4030 | Smith Medical | Italy | 21 | 85 | Usd | 5 Boxes of 50 | Yes | N/A | No sample submitted for evaluation | NOT COMPLY
42221512-00004 | 8 | Intravenous cannula, with introducer needle, radio-opaque, teflon/polyurethane, with double-tapered tip and stylet. 3-facetted J-point bevelled needle Length: Not less than 25mm Size: 22G Colour: Blue Flow rate: 36-40 ml/min Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 327 UNIT: Box (100 per box) |  | BAS MEDEXPRESS | 7.32 | Bicakcilap | 14622011 |  | Turkey | 1 - 3 days | 15% | N/A | As per client | Yes | Each | offered a flow rate of 32ml/min 36-40 ml/min
42221512-00005 | 9 | Intravenous cannula, with introducer needle, radio-opaque, teflon/polyurethane, with double-tapered tip and stylet. 3-facetted J-point bevelled needle Length: Not less than 19mm Size: 24G Colour: Yellow Flow rate: 22-26 ml/min Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 327 UNIT: Box (100 per box) | Basic Line offer | Becton Dickinson | 2 | BD Insyte 24G | 381212 | Becton Dickinson | San Augustine | 7 working days | 76% | USD | 50 units / 1 BOX | NO | 50 units in a box | COMPLY | NO TEST REPORT
42221512-00005 | 9 | Intravenous cannula, with introducer needle, radio-opaque, teflon/polyurethane, with double-tapered tip and stylet. 3-facetted J-point bevelled needle Length: Not less than 19mm Size: 24G Colour: Yellow Flow rate: 22-26 ml/min Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 327 UNIT: Box (100 per box) | Only one offer submitted | Tara Technology | 2.1 | Wellmed / Pencath | TAR817 | Wellmed | India | 21 days | 85% | USD | 10 | Yes | 100 per box | Offered a flow rate of 20ml/min instead 22-26 ml/min | NOT COMPLY
42221512-00005 | 9 | Intravenous cannula, with introducer needle, radio-opaque, teflon/polyurethane, with double-tapered tip and stylet. 3-facetted J-point bevelled needle Length: Not less than 19mm Size: 24G Colour: Yellow Flow rate: 22-26 ml/min Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 327 UNIT: Box (100 per box) |  | Chemical World cc | 2.14 | BIO ON | CIVNP24 | BIOMED | INDIA | 3 | 85 | USD | 1000 | YES | 1 | Offered a flow rate of 20ml/min instead 22-26 ml/min | NOT COMPLY
42221512-00005 | 9 | Intravenous cannula, with introducer needle, radio-opaque, teflon/polyurethane, with double-tapered tip and stylet. 3-facetted J-point bevelled needle Length: Not less than 19mm Size: 24G Colour: Yellow Flow rate: 22-26 ml/min Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 327 UNIT: Box (100 per box) | OFFER 2 | SUPRA HEALTHCARE JOHANNESBURG PTY LTD | 2.42 | Vennula | IVC-24 | Romsons | INDIA | Initial 60, thereafter 14 | 84 | USD | 1 Box | YES | 100/BOX | Offered a flow rate of 20ml/min instead 22-26 ml/min | NO TEST REPORTS
42221512-00005 | 9 | Intravenous cannula, with introducer needle, radio-opaque, teflon/polyurethane, with double-tapered tip and stylet. 3-facetted J-point bevelled needle Length: Not less than 19mm Size: 24G Colour: Yellow Flow rate: 22-26 ml/min Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 327 UNIT: Box (100 per box) | A | Quality Medical Supplies Pty Ltd | 2.76 | Vennula | TBA | Romsons International | India | 14 | 80% | USD | 100 Units | Yes | Box | Offered a flow rate of 20ml/min instead 22-26 ml/min | NOT COMPLY
42221512-00005 | 9 | Intravenous cannula, with introducer needle, radio-opaque, teflon/polyurethane, with double-tapered tip and stylet. 3-facetted J-point bevelled needle Length: Not less than 19mm Size: 24G Colour: Yellow Flow rate: 22-26 ml/min Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 327 UNIT: Box (100 per box) | Offer 1 of 1 | SSEM Mthembu Medical (Pty) Ltd | 275.32 | Polymed | PM121024 | Polymed | India | 56 | 85% | USD | 500 | Yes | 100/Box | COMPLY pending confirmation of latex free | COMPLY | Recommended | 95
42221512-00005 | 9 | Intravenous cannula, with introducer needle, radio-opaque, teflon/polyurethane, with double-tapered tip and stylet. 3-facetted J-point bevelled needle Length: Not less than 19mm Size: 24G Colour: Yellow Flow rate: 22-26 ml/min Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 327 UNIT: Box (100 per box) |  | CRYSATL PIER TRADING 148cc | 2.98 | MEDIFLOW/SR CARE |  | GLOBAL MEDIKIT LIMITED C/O SR MEDICAL CARE | [6 empty] | EACH | Offered a flow rate of 15ml/min instead 22-26 ml/min | NOT COMPLY
42221512-00005 | 9 | Intravenous cannula, with introducer needle, radio-opaque, teflon/polyurethane, with double-tapered tip and stylet. 3-facetted J-point bevelled needle Length: Not less than 19mm Size: 24G Colour: Yellow Flow rate: 22-26 ml/min Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 327 UNIT: Box (100 per box) | OFFER 1 | SUPRA HEALTHCARE JOHANNESBURG PTY LTD | 3.03 | Healcath | IVC-24 | Harsoria | INDIA | Initial 60, thereafter 14 | 84 | USD | 1 Box | YES | 100/BOX | Offered a flow rate of 18ml/min instead 22-26 ml/min | NOT COMPLY
42221512-00005 | 9 | Intravenous cannula, with introducer needle, radio-opaque, teflon/polyurethane, with double-tapered tip and stylet. 3-facetted J-point bevelled needle Length: Not less than 19mm Size: 24G Colour: Yellow Flow rate: 22-26 ml/min Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 327 UNIT: Box (100 per box) | EACH | MSIMED CC | 3.45 | KANG KANG | CIV04 | ZHENJIANG | CHINA | 25 | 20% | USD | 20000 PCS | YES | EACH | No sample submitted for evaluation | NO TEST REPORTS
42221512-00005 | 9 | Intravenous cannula, with introducer needle, radio-opaque, teflon/polyurethane, with double-tapered tip and stylet. 3-facetted J-point bevelled needle Length: Not less than 19mm Size: 24G Colour: Yellow Flow rate: 22-26 ml/min Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 327 UNIT: Box (100 per box) BOX (50 PER BOX) |  | JUMLA MEDICAL SUPPLIES (PTY) LTD | 4.01 | JELCO SERIVA | 8353 | Smiths Medical International | INDIA | 21 DAYS | 68.4% | BRITISH POUND | 500 | YES | NIL | COMPLY pending confirmation of latex free | COMPLY
42221512-00005 | 9 | Intravenous cannula, with introducer needle, radio-opaque, teflon/polyurethane, with double-tapered tip and stylet. 3-facetted J-point bevelled needle Length: Not less than 19mm Size: 24G Colour: Yellow Flow rate: 22-26 ml/min Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 327 UNIT: Box (100 per box) | Alternative Offer | BIOCOR HOSPITAL SUPPLIES PTY LTD | 4.6 | Jelco Seriva | 8353 | Smith Medical | India | 21 | 85 | USD | 5 Boxes 50 | Yes | N/A | COMPLY | COMPLY
42221512-00005 | 9 | Intravenous cannula, with introducer needle, radio-opaque, teflon/polyurethane, with double-tapered tip and stylet. 3-facetted J-point bevelled needle Length: Not less than 19mm Size: 24G Colour: Yellow Flow rate: 22-26 ml/min Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 327 UNIT: Box (100 per box) |  | ALMIKA TRADING PTY LTD | 4.63 | BD |  | BD | USA |  | 0 | NO | 100 | YES | 1 | COMPLY | NO TEST REPORTS
42221512-00005 | 9 | Intravenous cannula, with introducer needle, radio-opaque, teflon/polyurethane, with double-tapered tip and stylet. 3-facetted J-point bevelled needle Length: Not less than 19mm Size: 24G Colour: Yellow Flow rate: 22-26 ml/min Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 327 UNIT: Box (100 per box) | BASIC LINE OFFER | A R FREIGHT PRIVATE LIMITED | 4.8 | JELCO SERIVA | 8353 | SMITHS MEDICAL | INDIA | 21 DAYS | 85% | USD | 100 | YES | N/A | COMPLY | COMPLY
42221512-00005 | 9 | Intravenous cannula, with introducer needle, radio-opaque, teflon/polyurethane, with double-tapered tip and stylet. 3-facetted J-point bevelled needle Length: Not less than 19mm Size: 24G Colour: Yellow Flow rate: 22-26 ml/min Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 327 UNIT: Box (100 per box) | 1 | B. BRAUN MEDICAL (PTY) LTD | 4.8 | Introcan FEP 24G x 3/4, 0,7 x 19mm | 4252071B | B. Braun | Germany | 42 | 50% | Euro | 100 | Yes | Each | COMPLY | COMPLY
42221512-00005 | 9 | Intravenous cannula, with introducer needle, radio-opaque, teflon/polyurethane, with double-tapered tip and stylet. 3-facetted J-point bevelled needle Length: Not less than 19mm Size: 24G Colour: Yellow Flow rate: 22-26 ml/min Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 327 UNIT: Box (100 per box) |  | LECHOBA MEDICAL TECHNOLOGIES (PTY) LTD | 4.98 | TRO-VENOCATH SINE | 90083 | TROGE MEDICAL GmBH | GERMANY | 21 DAYS | 75% | USD | 100 | YES | EACH (PACKED 50/BOX) | COMPLY | NOT COMPLY
42221512-00005 | 9 | Intravenous cannula, with introducer needle, radio-opaque, teflon/polyurethane, with double-tapered tip and stylet. 3-facetted J-point bevelled needle Length: Not less than 19mm Size: 24G Colour: Yellow Flow rate: 22-26 ml/min Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 327 UNIT: Box (100 per box) | Basic Line Offer | BIOCOR HOSPITAL SUPPLIES PTY LTD | 5.75 | Jelco Radiopaque | 4033 | Smith Medical | Italy | 21 | 85 | USD | 5 Boxes of 50 | Yes | N/A | No sample submitted for evaluation | NOT COMPLY
42221512-00005 | 9 | Intravenous cannula, with introducer needle, radio-opaque, teflon/polyurethane, with double-tapered tip and stylet. 3-facetted J-point bevelled needle Length: Not less than 19mm Size: 24G Colour: Yellow Flow rate: 22-26 ml/min Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 327 UNIT: Box (100 per box) |  | BAS MEDEXPRESS | 7.32 | Bicakcilar | 146244011 |  | Turkey | 1-3 days | 15% | N/A | As per client | Yes | Box / 50 | Offered a flow rate of 32ml/min instead 22-26 ml/min Offered 22G instead of 24G
42221512-00007 | 10 | Intravenous cannula, with introducer needle, radio-opaque, teflon/polyurethane, with double-tapered tip and stylet. 3-facetted J-point bevelled needle Length: Not less than 19mm Size: 26G Colour: Yellow Flow rate: 13-17 ml/min Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 327 UNIT: Box (100 per box) |  | Chemical World cc | 2.42 | [10 empty] | No sample submitted for evaluation | NO TEST REPORTS
42221512-00007 | 10 | Intravenous cannula, with introducer needle, radio-opaque, teflon/polyurethane, with double-tapered tip and stylet. 3-facetted J-point bevelled needle Length: Not less than 19mm Size: 26G Colour: Violet Flow rate: 13-17 ml/min Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 327 UNIT: Box (100 per box) | Only one offer submitted | Tara Technology | 2.99 | Wellmed / Pencath | TAR818 | Wellmed | India | 21 days | 85% | USD | 10 | Yes | 100 per box | No sample submitted for evaluation. | NOT COMPLY
42221512-00007 | 10 | Intravenous cannula, with introducer needle, radio-opaque, teflon/polyurethane, with double-tapered tip and stylet. 3-facetted J-point bevelled needle Length: Not less than 19mm Size: 26G Colour: Violet Flow rate: 13-17 ml/min Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 327 UNIT: Box (100 per box) | Basic Line offer | CRYSATL PIER TRADING 148cc | 345 | MEDIFLOW/SR CARE | 391349 | GLOBAL MEDIKIT LIMITED C/O SR MEDICAL CARE | San Augustine | 7 working days | 76% | USD | 50 units / 1 BOX | NO | 100/Box | COMPLY | COMPLY | Recommended | 96.4
42221512-00007 | 10 | Intravenous cannula, with introducer needle, radio-opaque, teflon/polyurethane, with double-tapered tip and stylet. 3-facetted J-point bevelled needle Length: Not less than 19mm Size: 26G Colour: Violet Flow rate: 13-17 ml/min Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 327 UNIT: Box (100 per box) | OFFER 2 | SUPRA HEALTHCARE JOHANNESBURG PTY LTD | 3.22 | Micron | GS3033 | Romsons | INDIA | Initial 60, thereafter 14 | 84 | USD | 1 Box | YES | 100/BOX | COMPLY pending confirmation of latex free | NO TEST REPORTS
42221512-00007 | 10 | Intravenous cannula, with introducer needle, radio-opaque, teflon/polyurethane, with double-tapered tip and stylet. 3-facetted J-point bevelled needle Length: Not less than 19mm Size: 26G Colour: Violet Flow rate: 13-17 ml/min Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 327 UNIT: Box (100 per box) | OFFER 1 | SUPRA HEALTHCARE JOHANNESBURG PTY LTD | 3.28 | Healcath | 314061012 | Harsoria | INDIA | Initial 60, thereafter 14 | 84 | USD | 1 Box | YES | 100/BOX | Offered a flow rate of 10ml/min instead 13-17 ml/min | NOT COMPLY
42221512-00007 | 10 | Intravenous cannula, with introducer needle, radio-opaque, teflon/polyurethane, with double-tapered tip and stylet. 3-facetted J-point bevelled needle Length: Not less than 19mm Size: 26G Colour: Violet Flow rate: 13-17 ml/min Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 327 UNIT: Box (100 per box) | Offer 1 of 1 | SSEM Mthembu Medical (Pty) Ltd | 3.34 | Polymed | PM121026 | Polymed | India | 56 | 85% | USD | 500 | Yes | 50/Box | COMPLY pending confirmation of latex free | COMPLY
42221512-00007 | 10 | Intravenous cannula, with introducer needle, radio-opaque, teflon/polyurethane, with double-tapered tip and stylet. 3-facetted J-point bevelled needle Length: Not less than 19mm Size: 26G Colour: Violet Flow rate: 13-17 ml/min Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 327 UNIT: Box (100 per box) |  | CRYSATL PIER TRADING 148cc | 3.45 | MEDIFLOW/SR CARE |  | GLOBAL MEDIKIT LIMITED C/O SR MEDICAL CARE | [6 empty] | EACH | COMPLY pending confirmation of latex free | COMPLY
42221512-00007 | 10 | Intravenous cannula, with introducer needle, radio-opaque, teflon/polyurethane, with double-tapered tip and stylet. 3-facetted J-point bevelled needle Length: Not less than 19mm Size: 26G Colour: Violet Flow rate: 13-17 ml/min Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 327 UNIT: Box (100 per box) | A | Quality Medical Supplies Pty Ltd | 3.35 | Micron | TBA | Romsons International | India | 14 | 80% | USD | 100 Units | Yes | Box | COMPLY pending confirmation of latex free | NO TEST REPORTS
42221512-00007 | 10 | Intravenous cannula, with introducer needle, radio-opaque, teflon/polyurethane, with double-tapered tip and stylet. 3-facetted J-point bevelled needle Length: Not less than 19mm Size: 26G Colour: Violet Flow rate: 13-17 ml/min Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 327 UNIT: Box (100 per box) | BOX OF 100 | MSIMED CC | 3.5 | KANG KANG | CIV05 | ZHENJIANG | CHINA | 25 | 20% | USD | 20000 PCS | YES | BOX | No sample submitted for evaluation | NO TEST REPORTS
42221512-00007 | 10 | Intravenous cannula, with introducer needle, radio-opaque, teflon/polyurethane, with double-tapered tip and stylet. 3-facetted J-point bevelled needle Length: Not less than 19mm Size: 26G Colour: Violet Flow rate: 13-17 ml/min Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 327 UNIT: Box (100 per box) BOX (50 PER BOX) |  | JUMLA MEDICAL SUPPLIES (PTY) LTD | 4.01 | JELCO SERIVA | 8352 | Smiths Medical International | INDIA | 21 DAYS | 68.4% | BRITISH POUND | 500 | YES | NIL | COMPLY pending confirmation of latex free | COMPLY
42221512-00007 | 10 | Intravenous cannula, with introducer needle, radio-opaque, teflon/polyurethane, with double-tapered tip and stylet. 3-facetted J-point bevelled needle Length: Not less than 19mm Size: 26G Colour: Violet Flow rate: 13-17 ml/min Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 327 UNIT: Box (100 per box) | Basic Line Offer | BIOCOR HOSPITAL SUPPLIES PTY LTD | 4.6 | Jelco Seriva | 8352 | Smith Medical | India | 21 | 85 | USD | 5 Boxes of 50 | Yes | N/A | COMPLY | COMPLY
42221512-00007 | 10 | Intravenous cannula, with introducer needle, radio-opaque, teflon/polyurethane, with double-tapered tip and stylet. 3-facetted J-point bevelled needle Length: Not less than 19mm Size: 26G Colour: Violet Flow rate: 13-17 ml/min Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 327 UNIT: Box (100 per box) | BASIC LINE OFFER | A R FREIGHT PRIVATE LIMITED | 4.8 | JELCO SERIVA | 8352 | SMITHS MEDICAL | INDIA | 21 DAYS | 85% | USD | 100 | YES | N/A | COMPLY | COMPLY
42221512-00007 | 10 | Intravenous cannula, with introducer needle, radio-opaque, teflon/polyurethane, with double-tapered tip and stylet. 3-facetted J-point bevelled needle Length: Not less than 19mm Size: 26G Colour: Violet Flow rate: 13-17 ml/min Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 327 UNIT: Box (100 per box) |  | LECHOBA MEDICAL TECHNOLOGIES (PTY) LTD | 5.33 | CHIRAFLEX | CHFX03026 | CHIRANA INJECTA A/S (PART OF CTI GROUP) | SLOVAK REPUBLIC | 21 DAYS | 75% | EURO | 100 | YES | 50/BOX (PACKED 50/BOX) | COMPLY | COMPLY
42221512-00007 | 10 | Intravenous cannula, with introducer needle, radio-opaque, teflon/polyurethane, with double-tapered tip and stylet. 3-facetted J-point bevelled needle Length: Not less than 19mm Size: 26G Colour: Violet Flow rate: 13-17 ml/min Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 327 UNIT: Box (100 per box) |  | ALMIKA TRADING PTY LTD | 6.61 | BD | 391349 | BD | USA | 14 | 0 | NO | 100 | YES | 1 | COMPLY | COMPLY
42221512-00007 | 10 | Intravenous cannula, with introducer needle, radio-opaque, teflon/polyurethane, with double-tapered tip and stylet. 3-facetted J-point bevelled needle Length: Not less than 19mm Size: 26G Colour: Violet Flow rate: 13-17 ml/min Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 327 UNIT: Box (100 per box) |  | BAS MEDEXPRESS
42261508-00000 | 11 | Needle, biopsy, lightweight hub with a specimen retaining notch of not less than 20mm Sterile, Single use, Individually peel packed Size: 14G (1.9mm) x 75mm UNIT: Box (20 per box) |  | Chemical World cc | 1933.4 | MDL | PM14074 | MDL | ITALY | 3 | 85 | EURO | 20 | YES | 20 per box | COMPLY pending confirmation of latex free |  | Recommended | 100
42261508-00000 | 11 | Needle, biopsy, lightweight hub with a specimen retaining notch of not less than 20mm Sterile, Single use, Individually peel packed Size: 14G (1.9mm) x 75mm UNIT: Box (20 per box) |  | GIBELA TRADE AND INVEST 1001 CC | 99.0789 | TARA | NEE001 | Shandong Qiaopai Group Co., Ltd | CHINA | 21 | 85 | USD | 10 | YES | each | COMPLY pending confirmation of latex free
42261508-00000 | 11 | Needle, biopsy, lightweight hub with a specimen retaining notch of not less than 20mm Sterile, Single use, Individually peel packed Size: 14G (1.9mm) x 75mm UNIT: Box (20 per box) |  | UNITRADE 1032 CC | 115.83 | BIOSURE | 1 | SHANGHAI KINDLY ENTERPRISE DEVELOPMENT GROUP | ROC | 21 | 85 | USD | 100 | YES | 20 | No sample submitted for evaluation
42261508-00000 | 11 | Needle, biopsy, lightweight hub with a specimen retaining notch of not less than 20mm Sterile, Single use, Individually peel packed Size: 14G (1.9mm) x 75mm UNIT: Box (20 per box) | Basic Line Offer | BIOCOR HOSPITAL SUPPLIES PTY LTD | 132.55 | Technomed | TBN 14075 | Techno Med | South Africa | 21 | 0 | 0 | Box/50 | Yes | N/A | COMPLY pending confirmation of latex free
42261508-00000 | 11 | Needle, biopsy, lightweight hub with a specimen retaining notch of not less than 20mm Sterile, Single use, Individually peel packed Size: 14G (1.9mm) x 75mm UNIT: Box (20 per box) |  | SUPRA HEALTHCARE JOHANNESBURG PTY LTD | 166.57 | Fine-Cut | BIOPSY001 | EGEMEN | TURKEY | Initial 60, thereafter 14 | 84 | USD | 1 Box | YES |  | COMPLY pending confirmation of latex free
42261508-00000 | 11 | Needle, biopsy, lightweight hub with a specimen retaining notch of not less than 20mm Sterile, Single use, Individually peel packed Size: 14G (1.9mm) x 75mm UNIT: Box (20 per box) |  | Chemical World cc | 168.48 | MDL | PD01407 | MDL | ITALY | 3 | 85 | EURO | 20 | YES | 1 | No sample submitted for evaluation
42261508-00000 | 11 | Needle, biopsy, lightweight hub with a specimen retaining notch of not less than 20mm Sterile, Single use, Individually peel packed Size: 14G (1.9mm) x 75mm UNIT: Box (20 per box) BOX (10 PER BOX) |  | JUMLA MEDICAL SUPPLIES (PTY) LTD | 1959.95 | FINE-CUT | FC14075 | TIBBI MEDICAL | TURKEY | 21 DAYS | 70% | US DOLLAR | 1 BOX | YES | NIL | COMPLY pending confirmation of latex free
42261508-00001 | 12 | Needle, biopsy, lightweight hub with a specimen retaining notch of not less than 20mm Sterile, Single use, Individually peel packed Size: 14G (1.9mm) x 114 mm UNIT: Box (20 per box) |  | Chemical World cc | 1933.4 | MDL | PM1411 | MDL | ITALY | 3 | 85 | EURO | 20 | YES | 20 per box | COMPLY pending confirmation of latex free |  | Recommended | 100
42261508-00001 | 12 | Needle, biopsy, lightweight hub with a specimen retaining notch of not less than 20mm Sterile, Single use, Individually peel packed Size: 14G (1.9mm) x 114 mm UNIT: Box (20 per box) |  | GIBELA TRADE AND INVEST 1001 CC | 99.0789 | TARA | NEE002 | Shandong Qiaopai Group Co., Ltd | CHINA | 21 | 85 | USD | 10 | YES | each | COMPLY pending confirmation of latex free
42261508-00001 | 12 | Needle, biopsy, lightweight hub with a specimen retaining notch of not less than 20mm Sterile, Single use, Individually peel packed Size: 14G (1.9mm) x 114 mm UNIT: Box (20 per box) |  | UNITRADE 1032 CC | 115.83 | BIOSURE | 1 | SHANGHAI KINDLY ENTERPRISE DEVELOPMENT GROUP | ROC | 21 | 85 | USD | 100 | YES | 20 | No sample submitted for evaluation
42261508-00001 | 12 | Needle, biopsy, lightweight hub with a specimen retaining notch of not less than 20mm Sterile, Single use, Individually peel packed Size: 14G (1.9mm) x 114 mm UNIT: Box (20 per box) | Basic Line Offer | BIOCOR HOSPITAL SUPPLIES PTY LTD | 132.55 | Technomed | TBN 14115 | Techno Med | South Africa | 21 | 0 | 0 | Box/50 | Yes | N/A | COMPLY pending confirmation of latex free
42261508-00001 | 12 | Needle, biopsy, lightweight hub with a specimen retaining notch of not less than 20mm Sterile, Single use, Individually peel packed Size: 14G (1.9mm) x 114 mm UNIT: Box (20 per box) |  | SUPRA HEALTHCARE JOHANNESBURG PTY LTD | 166.57 | Fine-Cut | BIOPSY002 | EGEMEN | TURKEY | Initial 60, thereafter 14 | 84 | USD | 1 Box | YES |  | COMPLY pending confirmation of latex free
42261508-00001 | 12 | Needle, biopsy, lightweight hub with a specimen retaining notch of not less than 20mm Sterile, Single use, Individually peel packed Size: 14G (1.9mm) x 114 mm UNIT: Box (20 per box) |  | Chemical World cc | 168.48 | MDL | PDO1411 | MDL | ITALY | 3 | 85 | EURO | 20 | YES | 1 | No sample submitted for evaluation
42261508-00001 | 12 | Needle, biopsy, lightweight hub with a specimen retaining notch of not less than 20mm Sterile, Single use, Individually peel packed Size: 14G (1.9mm) x 114 mm UNIT: Box (20 per box) BOX (10 PER BOX) |  | JUMLA MEDICAL SUPPLIES (PTY) LTD | 1959.95 | FINE-CUT | FC14114 | TIBBI MEDICAL | TURKEY | 21 DAYS | 70% | US DOLLAR | 1 BOX | YES | NIL | COMPLY pending confirmation of latex free
42261508-00002 | 13 | Needle, biopsy, lightweight hub with a specimen retaining notch of not less than 20mm Sterile, Single use, Individually peel packed Size: 14G (1.9mm) x 152 mm UNIT: Box (20 per box) |  | Chemical World cc | 1933.4 | MDL | PM1415 | MDL | ITALY | 3 | 85 | EURO | 20 | YES | 20 per box | COMPLY pending confirmation of latex free |  | Recommended | 100
42261508-00002 | 13 | Needle, biopsy, lightweight hub with a specimen retaining notch of not less than 20mm Sterile, Single use, Individually peel packed Size: 14G (1.9mm) x 152 mm UNIT: Box (20 per box) |  | GIBELA TRADE AND INVEST 1001 CC | 99.0789 | TARA | NEE003 | Shandong Qiaopai Group Co., Ltd | CHINA | 21 | 85 | USD | 10 | YES | each | COMPLY pending confirmation of latex free
42261508-00002 | 13 | Needle, biopsy, lightweight hub with a specimen retaining notch of not less than 20mm Sterile, Single use, Individually peel packed Size: 14G (1.9mm) x 152 mm UNIT: Box (20 per box) |  | UNITRADE 1032 CC | 115.83 | BIOSURE | 1 | SHANGHAI KINDLY ENTERPRISE DEVELOPMENT GROUP | ROC | 21 | 85 | USD | 100 | YES | 20 | No sample submitted for evaluation
42261508-00002 | 13 | Needle, biopsy, lightweight hub with a specimen retaining notch of not less than 20mm Sterile, Single use, Individually peel packed Size: 14G (1.9mm) x 152 mm UNIT: Box (20 per box) | Basic Line Offer | BIOCOR HOSPITAL SUPPLIES PTY LTD | 132.55 | Technomed | TBN 14150 | Techno Med | South Africa | 21 | 0 | 0 | Box/50 | Yes | N/A | COMPLY pending confirmation of latex free
42261508-00002 | 13 | Needle, biopsy, lightweight hub with a specimen retaining notch of not less than 20mm Sterile, Single use, Individually peel packed Size: 14G (1.9mm) x 152 mm UNIT: Box (20 per box) |  | SUPRA HEALTHCARE JOHANNESBURG PTY LTD | 166.57 | Fine-Cut | BIOPSY003 | EGEMEN | TURKEY | Initial 60, thereafter 14 | 84 | USD | 1 Box | YES |  | COMPLY pending confirmation of latex free
42261508-00002 | 13 | Needle, biopsy, lightweight hub with a specimen retaining notch of not less than 20mm Sterile, Single use, Individually peel packed Size: 14G (1.9mm) x 152 mm UNIT: Box (20 per box) |  | Chemical World cc | 168.48 | MDL | PDO1415 | MDL | ITALY | 3 | 85 | EURO | 20 | YES | 1 | No sample submitted for evaluation
42261508-00002 | 13 | Needle, biopsy, lightweight hub with a specimen retaining notch of not less than 20mm Sterile, Single use, Individually peel packed Size: 14G (1.9mm) x 152 mm UNIT: Box (20 per box) BOX (10 PER BOX) |  | JUMLA MEDICAL SUPPLIES (PTY) LTD | 1959.95 | FINE-CUT | FC14152 | TIBBI MEDICAL | TURKEY | 21 DAYS | 70% | US DOLLAR | 1 BOX | YES | NIL | COMPLY pending confirmation of latex free
42261508-00003 | 14 | Needle, biopsy, bone marrow aspiration with adjustable needle stop Size: 14G x 30mm Colour code: Green Sterile, Single use, Individually peel packed UNIT: Box (10 per box) |  | SUPRA HEALTHCARE JOHANNESBURG PTY LTD | 114.48 | Stilo-can | BIOPSY006 | EGEMEN | TURKEY | Initial 60, thereafter 14 | 84 | USD | 1 Box | YES | 14G X40MM | Offered a length of 40mm instead of 30mm Offered 13G instead of 14G |  | Re-tender endusers to revise specification
42261508-00003 | 14 | Needle, biopsy, bone marrow aspiration with adjustable needle stop Size: 14G x 30mm Colour code: Green Sterile, Single use, Individually peel packed UNIT: Box (10 per box) |  | Chemical World cc | 120.75 | MDL | PIL1435100 | MDL | ITALY | 3 | 85 | EURO | 20 | NO.35MM | 1 | Offered a length of 35mm instead of 30mm |  | Re-tender endusers to revise specification
42261508-00003 | 14 | Needle, biopsy, bone marrow aspiration with adjustable needle stop Size: 14G x 30mm Colour code: Green Sterile, Single use, Individually peel packed UNIT: Box (10 per box) |  | ABCARE TECHNOLOGIES (PTY) LTD | 160 | ZAMAR | BM001 | ZAMAR DOO | CROATIA | 45 | 85 | EURO | 100 | YES | EACH | No sample submitted for evaluation |  | Re-tender endusers to revise specification
42261508-00003 | 14 | Needle, biopsy, bone marrow aspiration with adjustable needle stop Size: 14G x 30mm Colour code: Green Sterile, Single use, Individually peel packed UNIT: Box (10 per box) |  | JUMLA MEDICAL SUPPLIES (PTY) LTD | 1279.61 | STILO-CAN | SC1440 | TIBBI MEDICAL | TURKEY | 21 DAYS | 70% | US DOLLAR | 1 BOX | YES | NIL | Offered a length of 40mm instead of 30mm |  | Re-tender endusers to revise specification
42261508-00004 | 15 | Needle, biopsy, bone marrow aspiration with adjustable needle stop Size: 15G x 30mm Colour code: Blue Sterile, Single use, Individually peel packed UNIT: Box (10 per box) |  | SUPRA HEALTHCARE JOHANNESBURG PTY LTD | 114.48 | Fine-Cut | BIOPSY007 | EGEMEN | TURKEY | Initial 60, thereafter 14 | 84 | USD | 1 Box | YES | 15G X40MM | Offered a length of 40mm instead of 30mm |  | Re-tender endusers to revise specification
42261508-00004 | 15 | Needle, biopsy, bone marrow aspiration with adjustable needle stop Size: 15G x 30mm Colour code: Blue Sterile, Single use, Individually peel packed UNIT: Box (10 per box) |  | Chemical World cc | 120.75 | MDL | PIL1535100 | MDL | ITALY | 3 | 85 | EURO | 20 | NO.35MM | 1 | Offered a length of 35mm instead of 30mm |  | Re-tender endusers to revise specification
42261508-00004 | 15 | Needle, biopsy, bone marrow aspiration with adjustable needle stop Size: 15G x 30mm Colour code: Blue Sterile, Single use, Individually peel packed UNIT: Box (10 per box) |  | ABCARE TECHNOLOGIES (PTY) LTD | 160 | ZAMAR | BM002 | ZAMAR DOO | CROATIA | 45 | 85 | EURO | 100 | YES | EACH | Product information is reflected on a removable sticker. LOT number, shelf life and date of production of the sticker and outer package is not the same. Product information of a removable sticker indicates a size 21G and the product information of the outer package indicates size 28G |  | Re-tender endusers to revise specification
42261508-00004 | 15 | Needle, biopsy, bone marrow aspiration with adjustable needle stop Size: 15G x 30mm Colour code: Blue Sterile, Single use, Individually peel packed UNIT: Box (10 per box) |  | JUMLA MEDICAL SUPPLIES (PTY) LTD | 1279.61 | STILO-CAN | SC1540 | TIBBI MEDICAL | TURKEY | 21 DAYS | 70% | US DOLLAR | 1 BOX | YES | NIL | Offered a length of 40mm instead of 30mm |  | Re-tender endusers to revise specification
42142502-00021 | 16 | Needle, spinal, quincke bevel; luer lock. The needle must be colour coded at the stylet cap and the size must be clearly stated. Must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. Must be bevelled at 220degrees Needle size: 20G with a needle length of 40mm ±2mm excluding hub Needle and stylet must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 25) | BASIC LINE OFFER | A R FREIGHT PRIVATE LIMITED | 3.06 | SUMBOW | N/A | SUMBOW MEDICAL | CHINA | 60 DAYS | 85% | USD | 100 | YES | N/A | Offered a 22G X 40mm instead of 20G X 40mm
42142502-00021 | 16 | Needle, spinal, quincke bevel; luer lock. The needle must be colour coded at the stylet cap and the size must be clearly stated. Must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. Must be bevelled at 220degrees Needle size: 20G with a needle length of 40mm ±2mm excluding hub Needle and stylet must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 25) |  | Medi-Core Technologies | 3.15 | ULIFE | 42142502-00021 | JIAXING U-LIFE MEDICAL DEVICE TECHNOLOGY.,CO LTD | CHINA | 21 | 80 | USD | 1 box | YES | BOX | No sample submitted for evaluation
42142502-00021 | 16 | Needle, spinal, quincke bevel; luer lock. The needle must be colour coded at the stylet cap and the size must be clearly stated. Must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. Must be bevelled at 220degrees Needle size: 20G with a needle length of 40mm ±2mm excluding hub Needle and stylet must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 25) |  | Chemical World cc | 3.57 | MEDIJECT | NDSS2040 | RUNQIANG | CHINA | 3 | 85 | USD | 1000 | YES | 50 | Offered a needle length of 90mm instead of 40mm
42142502-00021 | 16 | Needle, spinal, quincke bevel; luer lock. The needle must be colour coded at the stylet cap and the size must be clearly stated. Must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. Must be bevelled at 220degrees Needle size: 20G with a needle length of 40mm ±2mm excluding hub Needle and stylet must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 25) | 1 | GROBIR MEDICAL SUPPLIERS | 4.73 | DURACAN | DQ.10.020.040 | PROMECON | CHINA | 21 | 80 | US DOLLAR | 25 | Yes | UOM AS SPECIFIED / MEETS SPECIFICATION | No sample submitted for evaluation.
42142502-00021 | 16 | Needle, spinal, quincke bevel; luer lock. The needle must be colour coded at the stylet cap and the size must be clearly stated. Must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. Must be bevelled at 220degrees Needle size: 20G with a needle length of 40mm ±2mm excluding hub Needle and stylet must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 25) | Basic Line Offer | BIOCOR HOSPITAL SUPPLIES PTY LTD | 4.88 | Medline | N/A | Yangzhong Medline Industry Co.Ltd | China | 21 | 85 | USD | 10 Boxes | Yes | N/A | No sample submitted for evaluation
42142502-00021 | 16 | Needle, spinal, quincke bevel; luer lock. The needle must be colour coded at the stylet cap and the size must be clearly stated. Must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. Must be bevelled at 220degrees Needle size: 20G with a needle length of 40mm ±2mm excluding hub Needle and stylet must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 25) | N/A | SAFARMEX (PTY) LTD | 4.98 | Saclin | To be confirmed | Zhejiang Mediunion Healthcare Group Co. Ltd | China | 3days | 80% | US$ | 10000 Boxes | Yes | Box of 25 | No sample submitted for evaluation
42142502-00021 | 16 |  |  | Zanotone (Pty) Ltd | 6.39 | Mekon |  | Mekon | [6 empty] | EACH | No sample submitted for evaluation
42142502-00021 | 16 | Needle, spinal, quincke bevel; luer lock. The needle must be colour coded at the stylet cap and the size must be clearly stated. Must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. Must be bevelled at 220degrees Needle size: 20G with a needle length of 40mm ±2mm excluding hub Needle and stylet must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 25) |  | SUPRA HEALTHCARE JOHANNESBURG PTY LTD | 185.92 | REVITAL | SPNQ20X40 | Revital Health | KENYA | Initial 60 days, thereafter 14 | 84 | USD | 1 Box | YES | 25/Box | COMPLY pending confirmation of latex free |  | Recommended | 96
42142502-00021 | 16 | Needle, spinal, quincke bevel; luer lock. The needle must be colour coded at the stylet cap and the size must be clearly stated. Must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. Must be bevelled at 220degrees Needle size: 20G with a needle length of 40mm ±2mm excluding hub Needle and stylet must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 25) |  | Viking Critical Care (Pty) Ltd | 46.65 | VYGON | 69F18349 | VYGON | FRANCE | EX. STOCK 30 - 60 DAYS | 80% | EURO | 1 | NO, OUR PRODUCT IS 38MM | 1 | No sample submitted for evaluation
42142502-00021 | 16 | Needle, spinal, quincke bevel; luer lock. The needle must be colour coded at the stylet cap and the size must be clearly stated. Must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. Must be bevelled at 220degrees Needle size: 20G with a needle length of 40mm ±2mm excluding hub Needle and stylet must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 25) |  | BAS MEDEXPRESS | 776.25 | Mercury | SN2790Q |  | India | 1 - 3 days | 15% | N/A | As per client | Yes | Box | No sample submitted for evaluation
42142502-00022 | 17 | Needle, spinal, quincke bevel; luer lock. The needle must be colour coded at the stylet cap and the size must be clearly stated. Must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. Must be bevelled at 220degrees Needle size: 22G with a needle length of 40mm ±2mm excluding hub Needle and stylet must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 25) | BASIC LINE OFFER | A R FREIGHT PRIVATE LIMITED | 3.06 | SUMBOW | N/A | SUMBOW MEDICAL | CHINA | 60 DAYS | 85% | USD | 100 | YES | N/A | No sample submitted for evaluation.
42142502-00022 | 17 | Needle, spinal, quincke bevel; luer lock. The needle must be colour coded at the stylet cap and the size must be clearly stated. Must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. Must be bevelled at 220degrees Needle size: 22G with a needle length of 40mm ±2mm excluding hub Needle and stylet must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 25) |  | Medi-Core Technologies | 82.82 | ULIFE | 42142502-00022 | JIAXING U-LIFE MEDICAL DEVICE TECHNOLOGY.,CO LTD | CHINA | 21 | 80 | USD | 1 box | YES | BOX OF 25 | COMPLY pending confirmation of latex free |  | Recommended | 100
42142502-00022 | 17 | Needle, spinal, quincke bevel; luer lock. The needle must be colour coded at the stylet cap and the size must be clearly stated. Must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. Must be bevelled at 220degrees Needle size: 22G with a needle length of 40mm ±2mm excluding hub Needle and stylet must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 25) | N/A | PROMED TECHNOLOGIES (PTY) LTD | 3.34 | U-LIFE | 562022 | JIAXING | CHINA | 14 DAYS | 85% | USD | 1 BOX | YES | BOX | COMPLY pending confirmation of latex free
42142502-00022 | 17 | Needle, spinal, quincke bevel; luer lock. The needle must be colour coded at the stylet cap and the size must be clearly stated. Must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. Must be bevelled at 220degrees Needle size: 22G with a needle length of 40mm ±2mm excluding hub Needle and stylet must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 25) |  | Chemical World cc | 3.39 | MEDIJECT | NDSS2240 | RUNQIANG | CHINA | 3 | 85 | USD | 1000 | YES | 50 | COMPLY pending confirmation of latex free
42142502-00022 | 17 | Needle, spinal, quincke bevel; luer lock. The needle must be colour coded at the stylet cap and the size must be clearly stated. Must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. Must be bevelled at 220degrees Needle size: 22G with a needle length of 40mm ±2mm excluding hub Needle and stylet must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 25) | 1 | GROBIR MEDICAL SUPPLIERS | 4.73 | DURACAN | DQ.10.022.040 | PROMECON | CHINA | 21 | 80 | US DOLLAR | 25 | Yes | UOM AS SPECIFIED / MEETS SPECIFICATION | No sample submitted for evaluation.
42142502-00022 | 17 | Needle, spinal, quincke bevel; luer lock. The needle must be colour coded at the stylet cap and the size must be clearly stated. Must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. Must be bevelled at 220degrees Needle size: 22G with a needle length of 40mm ±2mm excluding hub Needle and stylet must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 25) | Basic Line Offer | BIOCOR HOSPITAL SUPPLIES PTY LTD | 4.88 | Medline | N/A | Yangzhong Medline Industry Co.Ltd | China | 21 | 85 | USD | 10 Boxes | Yes | N/A | No sample submitted for evaluation
42142502-00022 | 17 | Needle, spinal, quincke bevel; luer lock. The needle must be colour coded at the stylet cap and the size must be clearly stated. Must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. Must be bevelled at 220degrees Needle size: 22G with a needle length of 40mm ±2mm excluding hub Needle and stylet must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 25) | N/A | SAFARMEX (PTY) LTD | 4.98 | Saclin | To be confirmed | Zhejiang Mediunion Healthcare Group Co. Ltd | China | 3days | 80% | US$ | 10000 Boxes | Yes | Box of 25 | No sample submitted for evaluation
42142502-00022 | 15 | Needle, spinal, quincke bevel; luer lock. The needle must be colour coded at the stylet cap and the size must be clearly stated. Must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. Must be bevelled at 220degrees Needle size: 22G with a needle length of 40mm ±2mm excluding hub Needle and stylet must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 25) |  | Zanotone (Pty) Ltd | 6.39 | Mekon |  | Mekon | [6 empty] | Each | No sample submitted for evaluation
42142502-00022 | 17 | Needle, spinal, quincke bevel; luer lock. The needle must be colour coded at the stylet cap and the size must be clearly stated. Must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. Must be bevelled at 220degrees Needle size: 22G with a needle length of 40mm ±2mm excluding hub Needle and stylet must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 25) | OFFER 1 | SUPRA HEALTHCARE JOHANNESBURG PTY LTD | 7.44 | REVITAL | SPNQ22X40 | Revital Health | KENYA | Initial 60 days, thereafter 14 | 84 | USD | 1 Box | YES |  | COMPLY pending confirmation of latex free
42142502-00022 | 17 | Needle, spinal, quincke bevel; luer lock. The needle must be colour coded at the stylet cap and the size must be clearly stated. Must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. Must be bevelled at 220degrees Needle size: 22G with a needle length of 40mm ±2mm excluding hub Needle and stylet must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 25) | Basic Line offer | Becton Dickinson | 9 | BD Quinke Spinal Needle 22G X 38mm | 405254 | Becton Dickinson | Fraga, Spain | 7 Working Days | 76% | USD | 1 BOX | YES | 25 units in a box | COMPLY pending confirmation of latex free
42142502-00022 | 17 | Needle, spinal, quincke bevel; luer lock. The needle must be colour coded at the stylet cap and the size must be clearly stated. Must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. Must be bevelled at 220degrees Needle size: 22G with a needle length of 40mm ±2mm excluding hub Needle and stylet must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 25) |  | ALMIKA TRADING PTY LTD | 10.58 | BD | 405254 | BD | USA | 14 | 0 | NO | 100 | YES | 1 | COMPLY pending confirmation of latex free
42142502-00022 | 17 | Needle, spinal, quincke bevel; luer lock. The needle must be colour coded at the stylet cap and the size must be clearly stated. Must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. Must be bevelled at 220degrees Needle size: 22G with a needle length of 40mm ±2mm excluding hub Needle and stylet must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 25) |  | LECHOBA MEDICAL TECHNOLOGIES (PTY) LTD | 12.99 | UNISIS UNIEVER | 1150371 | UNISIS CORPORATION | JAPAN | 21 DAYS | 75% | USD | 25 | YES | 25/BOX | No sample submitted for evaluation
42142502-00022 | 17 | Needle, spinal, quincke bevel; luer lock. The needle must be colour coded at the stylet cap and the size must be clearly stated. Must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. Must be bevelled at 220degrees Needle size: 22G with a needle length of 40mm ±2mm excluding hub Needle and stylet must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 25) | 1 | B. BRAUN MEDICAL (PTY) LTD | 18 | Spinocan 22G x 1 1/2" (0.7 x 40mm) | 4507401 | B. Braun | Germany | 42 | 50% | Euro | 25 | Yes | Box of 25 | COMPLY pending cpnfirmation of latex free
42142502-00022 | 17 | Needle, spinal, quincke bevel; luer lock. The needle must be colour coded at the stylet cap and the size must be clearly stated. Must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. Must be bevelled at 220degrees Needle size: 22G with a needle length of 40mm ±2mm excluding hub Needle and stylet must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 25) | OFFER 2 | SUPRA HEALTHCARE JOHANNESBURG PTY LTD | 22.99 | Egemen | SPNQ22X40 | Egemen | TURKEY | Initial 60 days, thereafter 14 | 84 | USD | 1 Box | YES | 22x40 | COMPLY pending confirmation of latex free
42142502-00022 | 17 | Needle, spinal, quincke bevel; luer lock. The needle must be colour coded at the stylet cap and the size must be clearly stated. Must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. Must be bevelled at 220degrees Needle size: 22G with a needle length of 40mm ±2mm excluding hub Needle and stylet must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 25) |  | Viking Critical Care (Pty) Ltd | 46.65 | VYGON | 69F18347 | VYGON | FRANCE | EX. STOCK 30 - 60 DAYS | 80% | EURO | 1 | NO, OUR PRODUCT IS 38MM | 1 | Offered a needle length of 38mm instead of 40mm
42142502-00022 | 17 | Needle, spinal, quincke bevel; luer lock. The needle must be colour coded at the stylet cap and the size must be clearly stated. Must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. Must be bevelled at 220degrees Needle size: 22G with a needle length of 40mm ±2mm excluding hub Needle and stylet must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 25) |  | JUMLA MEDICAL SUPPLIES (PTY) LTD | 550.85 | EGEMEN | TSPQ2240 | TIBBI MEDICAL | TURKEY | 21 DAYS | 70% | US DOLLAR | 2 BOXES | YES | NIl | COMPLY pending confirmation of latex free
42142502-00023 | 18 | Needle, spinal, quincke bevel; luer lock. The needle must be colour coded at the stylet cap and the size must be clearly stated. Must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. Must be bevelled at 220degrees Needle size: 22G with a needle length of 90mm excluding hub Needle and stylet must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 25) | BASIC LINE OFFER | A R FREIGHT PRIVATE LIMITED | 76.5 | SUMBOW | N/A | SUMBOW MEDICAL | CHINA | 60 DAYS | 85% | USD | 100 | YES | BOX OF 25 | COMPLY pending confirmation of latex free |  | Recommended | 100
42142502-00023 | 18 | Needle, spinal, quincke bevel; luer lock. The needle must be colour coded at the stylet cap and the size must be clearly stated. Must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. Must be bevelled at 220degrees Needle size: 22G with a needle length of 90mm excluding hub Needle and stylet must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 25) |  | Medi-Core Technologies | 3.32 | ULIFE | 42142502-00023 | JIAXING U-LIFE MEDICAL DEVICE TECHNOLOGY.,CO LTD | CHINA | 21 | 80 | USD | 1 box | YES |  | COMPLY pending confirmation of latex free
42142502-00023 | 18 | Needle, spinal, quincke bevel; luer lock. The needle must be colour coded at the stylet cap and the size must be clearly stated. Must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. Must be bevelled at 220degrees Needle size: 22G with a needle length of 90mm excluding hub Needle and stylet must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 25) | N/A | PROMED TECHNOLOGIES (PTY) LTD | 3.34 | U-LIFE | 502023 | JIAXING | CHINA | 14 DAYS | 85% | USD | 1 BOX | YES | BOX | COMPLY pending confirmation of latex free
42142502-00023 | 18 | Needle, spinal, quincke bevel; luer lock. The needle must be colour coded at the stylet cap and the size must be clearly stated. Must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. Must be bevelled at 220degrees Needle size: 22G with a needle length of 90mm excluding hub Needle and stylet must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 25) |  | Chemical World cc | 3.39 | MEDIJECT | NDSFG222 | RUNQIANG | CHINA | 3 | 85 | USD | 1000 | YES | 50 | COMPLY pending confirmation of latex free
42142502-00023 | 18 | Needle, spinal, quincke bevel; luer lock. The needle must be colour coded at the stylet cap and the size must be clearly stated. Must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. Must be bevelled at 220degrees Needle size: 22G with a needle length of 90mm excluding hub Needle and stylet must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 25) |  | UNITRADE 1032 CC | 4.37 | SPINOJEC | 1 | SHANGHAI KINDLY ENTERPRISE DEVELOPMENT GROUP | ROC | 21 | 85 | USD | 400 | YES | 25 | No brand name. No product code.
42142502-00023 | 18 | Needle, spinal, quincke bevel; luer lock. The needle must be colour coded at the stylet cap and the size must be clearly stated. Must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. Must be bevelled at 220degrees Needle size: 22G with a needle length of 90mm excluding hub Needle and stylet must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 25) | 1 | GROBIR MEDICAL SUPPLIERS | 4.73 | DURACAN | DQ.10.022.090 | PROMECON | CHINA | 21 | 80 | US DOLLAR | 25 | Yes | UOM AS SPECIFIED / MEETS SPECIFICATION | COMPLY pending confirmation of latex free
42142502-00023 | 18 | Needle, spinal, quincke bevel; luer lock. The needle must be colour coded at the stylet cap and the size must be clearly stated. Must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. Must be bevelled at 220degrees Needle size: 22G with a needle length of 90mm excluding hub Needle and stylet must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 25) | Basic Line Offer | BIOCOR HOSPITAL SUPPLIES PTY LTD | 4.88 | Medline | N/A | Yangzhong Medline Industry Co.Ltd | China | 60 | 85 | USD | 10 Boxes | Yes | N/A | COMPLY pending confirmation of latex free
42142502-00023 | 18 | Needle, spinal, quincke bevel; luer lock. The needle must be colour coded at the stylet cap and the size must be clearly stated. Must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. Must be bevelled at 220degrees Needle size: 22G with a needle length of 90mm excluding hub Needle and stylet must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 25) | N/A | SAFARMEX (PTY) LTD | 4.98 | Saclin | To be confirmed | Zhejiang Mediunion Healthcare Group Co. Ltd | China | 3days | 80% | US$ | 10000 pieces | Yes | Box of 25 | No sample submitted for evaluation
42142502-00023 | 18 | Needle, spinal, quincke bevel; luer lock. The needle must be colour coded at the stylet cap and the size must be clearly stated. Must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. Must be bevelled at 220degrees Needle size: 22G with a needle length of 90mm excluding hub Needle and stylet must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 25) |  | Zanotone (Pty) Ltd | 6.39 | Mekon |  | Mekon | [6 empty] | Each | Not tated to be a quinckie needle
42142502-00023 | 18 | Needle, spinal, quincke bevel; luer lock. The needle must be colour coded at the stylet cap and the size must be clearly stated. Must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. Must be bevelled at 220degrees Needle size: 22G with a needle length of 90mm excluding hub Needle and stylet must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 25) | OFFER 1 | SUPRA HEALTHCARE JOHANNESBURG PTY LTD | 7.44 | REVITAL | SPNQ22X40 | Revital Health | KENYA | Initial 60 days, thereafter 14 | 84 | USD | 1 Box | YES |  | COMPLY pending confirmation of latex free
42142502-00023 | 18 | Needle, spinal, quincke bevel; luer lock. The needle must be colour coded at the stylet cap and the size must be clearly stated. Must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. Must be bevelled at 220degrees Needle size: 22G with a needle length of 90mm excluding hub Needle and stylet must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 25) | Basic Line offer | Becton Dickinson | 9 | BD Quinke Spinal Needle 22G X 90mm | 405256 | Becton Dickinson | Fraga, Spain | 7 Working Days | 76% | USD | 1 BOX | YES | 26 units in a box | COMPLY pending confirmation of latex free
42142502-00023 | 18 | Needle, spinal, quincke bevel; luer lock. The needle must be colour coded at the stylet cap and the size must be clearly stated. Must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. Must be bevelled at 220degrees Needle size: 22G with a needle length of 90mm excluding hub Needle and stylet must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 25) |  | ALMIKA TRADING PTY LTD | 10.58 | BD | 405256 | BD | USA | 14 | 0 | NO | 100 | YES | 1 | COMPLY pending confirmation of latex free
42142502-00023 | 18 | Needle, spinal, quincke bevel; luer lock. The needle must be colour coded at the stylet cap and the size must be clearly stated. Must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. Must be bevelled at 220degrees Needle size: 22G with a needle length of 90mm excluding hub Needle and stylet must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 25) |  | LECHOBA MEDICAL TECHNOLOGIES (PTY) LTD | 12.99 | UNISIS UNIEVER | 1151261 | UNISIS CORPORATION | JAPAN | 21 DAYS | 75% | USD | 25 | YES | 25/BOX | No sample submitted for evaluation
42142502-00023 | 18 | Needle, spinal, quincke bevel; luer lock. The needle must be colour coded at the stylet cap and the size must be clearly stated. Must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. Must be bevelled at 220degrees Needle size: 22G with a needle length of 90mm excluding hub Needle and stylet must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 25) | BOX OF 50 | MSIMED CC | 13 | KANG KANG | SPN001 | ZHENJIANG | CHINA | 25 | 20% | USD | 20000PCS | YES | BOX | No sample submitted for evaluation
42142502-00023 | 18 | Needle, spinal, quincke bevel; luer lock. The needle must be colour coded at the stylet cap and the size must be clearly stated. Must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. Must be bevelled at 220degrees Needle size: 22G with a needle length of 90mm excluding hub Needle and stylet must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 25) | OFFER 2 | SUPRA HEALTHCARE JOHANNESBURG PTY LTD | 14.2 | Egemen | SPNQ22X40 | Egemen | TURKEY | Initial 60 days, thereafter 14 | 84 | USD | 1 Box | YES | 22x90 | COMPLY pending confirmation of latex free
42142502-00023 | 18 | Needle, spinal, quincke bevel; luer lock. The needle must be colour coded at the stylet cap and the size must be clearly stated. Must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. Must be bevelled at 220degrees Needle size: 22G with a needle length of 90mm excluding hub Needle and stylet must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 25) | 1 | B. BRAUN MEDICAL (PTY) LTD | 18 | Spinocan 22G x 3 1/2" (0.7 x 88mm) | 4507908 | B. Braun | Germany | 42 | 50% | Euro | 25 | Yes | Box of 25 | COMPLY pending cpnfirmation of latex free
42142502-00023 | 18 | Needle, spinal, quincke bevel; luer lock. The needle must be colour coded at the stylet cap and the size must be clearly stated. Must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. Must be bevelled at 220degrees Needle size: 22G with a needle length of 90mm excluding hub Needle and stylet must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 25) |  | Viking Critical Care (Pty) Ltd | 46.65 | VYGON | 69F18397 | VYGON | FRANCE | EX. STOCK 30 - 60 DAYS | 80% | EURO | 1 | YES | 1 | No sample submitted for evaluation
42142502-00023 | 18 | Needle, spinal, quincke bevel; luer lock. The needle must be colour coded at the stylet cap and the size must be clearly stated. Must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. Must be bevelled at 220degrees Needle size: 22G with a needle length of 90mm excluding hub Needle and stylet must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 25) |  | JUMLA MEDICAL SUPPLIES (PTY) LTD | 356.21 | EGEMEN | TSPQ2290 | TIBBI MEDICAL | TURKEY | 21 DAYS | 70% | US DOLLAR | 2 BOXES | YES | NIL | COMPLY pending confirmation of latex free
42142502-00023 | 18 | Needle, spinal, quincke bevel; luer lock. The needle must be colour coded at the stylet cap and the size must be clearly stated. Must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. Must be bevelled at 220degrees Needle size: 22G with a needle length of 90mm excluding hub Needle and stylet must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 25) |  | BAS MEDEXPRESS | 776.25 | Mercury | SN1690Q |  | India | 1 - 3 days | 15% | N/A | As per client | Yes | Box | No sample submitted for evaluation
42142502-00024 | 19 | Needle, spinal, quincke bevel; luer lock. The needle must be colour coded at the stylet cap and the size must be clearly stated. Must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. Must be bevelled at 220degrees Needle size: 22G extra-length with a needle length of 110mm excluding hub Needle and stylet must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 25) | BASIC LINE OFFER | A R FREIGHT PRIVATE LIMITED | 3.06 | SUMBOW | N/A | SUMBOW MEDICAL | CHINA | 60 DAYS | 85% | USD | 100 | YES | N/A | No sample submitted for evaluation.
42142502-00024 | 19 | Needle, spinal, quincke bevel; luer lock. The needle must be colour coded at the stylet cap and the size must be clearly stated. Must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. Must be bevelled at 220degrees Needle size: 22G extra-length with a needle length of 110mm excluding hub Needle and stylet must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 25) |  | Medi-Core Technologies | 3.32 | ULIFE | 42142502-00024 | JIAXING U-LIFE MEDICAL DEVICE TECHNOLOGY.,CO LTD | CHINA | 21 | 80 | USD | 1 box | YES |  | No sample submitted for evaluation
42142502-00024 | 19 | Needle, spinal, quincke bevel; luer lock. The needle must be colour coded at the stylet cap and the size must be clearly stated. Must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. Must be bevelled at 220degrees Needle size: 22G extra-length with a needle length of 110mm excluding hub Needle and stylet must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 25) | N/A | PROMED TECHNOLOGIES (PTY) LTD | 3.34 | U-LIFE | 502024 | JIAXING | CHINA | 14 DAYS | 85% | USD | 1 BOX | YES | BOX | No samples submitted for evaluation
42142502-00024 | 19 | Needle, spinal, quincke bevel; luer lock. The needle must be colour coded at the stylet cap and the size must be clearly stated. Must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. Must be bevelled at 220degrees Needle size: 22G extra-length with a needle length of 110mm excluding hub Needle and stylet must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 25) |  | UNITRADE 1032 CC | 4.37 | SPINOJEC | 1 | SHANGHAI KINDLY ENTERPRISE DEVELOPMENT GROUP | ROC | 21 | 85 | USD | 400 | YES | 25 | No sample submitted for evaluation
42142502-00024 | 19 | Needle, spinal, quincke bevel; luer lock. The needle must be colour coded at the stylet cap and the size must be clearly stated. Must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. Must be bevelled at 220degrees Needle size: 22G extra-length with a needle length of 110mm excluding hub Needle and stylet must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 25) | 1 | GROBIR MEDICAL SUPPLIERS | 4.73 | DURACAN | DQ.10.022.110 | PROMECON | CHINA | 21 | 80 | US DOLLAR | 25 | Yes | UOM AS SPECIFIED / MEETS SPECIFICATION | No sample submitted for evaluation.
42142502-00024 | 19 | Needle, spinal, quincke bevel; luer lock. The needle must be colour coded at the stylet cap and the size must be clearly stated. Must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. Must be bevelled at 220degrees Needle size: 22G extra-length with a needle length of 110mm excluding hub Needle and stylet must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 25) | Basic Line Offer | BIOCOR HOSPITAL SUPPLIES PTY LTD | 5.19 | Medline |  | Yangzhong Medline Industry Co.Ltd | China | 60 | 85 | USD | 10 Boxes | Yes | N/A | No sample submitted for evaluation
42142502-00024 | 19 | Needle, spinal, quincke bevel; luer lock. The needle must be colour coded at the stylet cap and the size must be clearly stated. Must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. Must be bevelled at 220degrees Needle size: 22G extra-length with a needle length of 110mm excluding hub Needle and stylet must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 25) | N/A | SAFARMEX (PTY) LTD | 5.38 | Saclin | To be confirmed | Zhejiang Mediunion Healthcare Group Co. Ltd | China | 3days | 80% | US$ | 10000 pieces | Yes | Box of 25 | No sample submitted for evaluation
42142502-00024 | 19 | Needle, spinal, quincke bevel; luer lock. The needle must be colour coded at the stylet cap and the size must be clearly stated. Must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. Must be bevelled at 220degrees Needle size: 22G extra-length with a needle length of 110mm excluding hub Needle and stylet must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 25) |  | Zanotone (Pty) Ltd | 6.39 | Mekon |  | Mekon | [6 empty] | EACH | No sample submitted for evaluation
42142502-00024 | 19 | Needle, spinal, quincke bevel; luer lock. The needle must be colour coded at the stylet cap and the size must be clearly stated. Must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. Must be bevelled at 220degrees Needle size: 22G extra-length with a needle length of 110mm excluding hub Needle and stylet must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 25) | Basic Line offer | Becton Dickinson | 7.2 | BD Quinke Spinal Needle 22G X 119mm | 405234 | Becton Dickinson | Fraga, Spain | 7 Working Days | 76% | USD | 1 BOX | YES | 27 units in a box | No sample submitted for evaluation.
42142502-00024 | 19 | Needle, spinal, quincke bevel; luer lock. The needle must be colour coded at the stylet cap and the size must be clearly stated. Must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. Must be bevelled at 220degrees Needle size: 22G extra-length with a needle length of 110mm excluding hub Needle and stylet must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 25) | OFFER 1 | SUPRA HEALTHCARE JOHANNESBURG PTY LTD | 185.92 | REVITAL | SPNQ22X110 | Revital Health | KENYA | Initial 60 days, thereafter 14 | 84 | USD | 1 Box | YES | 25/Box | COMPLY pending confirmation of latex free |  | Recommended | 96
42142502-00024 | 19 | Needle, spinal, quincke bevel; luer lock. The needle must be colour coded at the stylet cap and the size must be clearly stated. Must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. Must be bevelled at 220degrees Needle size: 22G extra-length with a needle length of 110mm excluding hub Needle and stylet must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 25) | BOX OF 50 | MSIMED CC | 13 | KANG KANG | SPN002 | ZHENJIANG | CHINA | 25 | 20% | USD | 20000PCS | YES | BOX | No sample submitted for evaluation
42142502-00024 | 19 | Needle, spinal, quincke bevel; luer lock. The needle must be colour coded at the stylet cap and the size must be clearly stated. Must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. Must be bevelled at 220degrees Needle size: 22G extra-length with a needle length of 110mm excluding hub Needle and stylet must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 25) |  | ALMIKA TRADING PTY LTD | 19.84 | BD | 405234 | BD | USA | 14 | 0 | NO | 100 | YES | 1 | No sample submitted for evaluation.
42142502-00024 | 19 | Needle, spinal, quincke bevel; luer lock. The needle must be colour coded at the stylet cap and the size must be clearly stated. Must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. Must be bevelled at 220degrees Needle size: 22G extra-length with a needle length of 110mm excluding hub Needle and stylet must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 25) | OFFER 2 | SUPRA HEALTHCARE JOHANNESBURG PTY LTD | 23.67 | Egemen | SPNQ22X120 | Egemen | TURKEY | Initial 60 days, thereafter 14 | 84 | USD | 1 Box | YES | 22GX120MM | COMPLY pending confirmation of latex free
42142502-00024 | 19 | Needle, spinal, quincke bevel; luer lock. The needle must be colour coded at the stylet cap and the size must be clearly stated. Must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. Must be bevelled at 220degrees Needle size: 22G extra-length with a needle length of 110mm excluding hub Needle and stylet must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 25) |  | Viking Critical Care (Pty) Ltd | 140.84 | VYGON | 69F518027 | VYGON | FRANCE | EX. STOCK 30 - 60 DAYS | 80% | EURO | 1 | YES | 1 | Offered a needle length of 120mm instead of 110mm
42142502-00024 | 19 | Needle, spinal, quincke bevel; luer lock. The needle must be colour coded at the stylet cap and the size must be clearly stated. Must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. Must be bevelled at 220degrees Needle size: 22G extra-length with a needle length of 110mm excluding hub Needle and stylet must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 25) |  | JUMLA MEDICAL SUPPLIES (PTY) LTD | 566.95 | EGEMEN | TSPQ22120 | TIBBI MEDICAL | TURKEY | 21 DAYS | 70% | US DOLLAR | 2 BOXES | YES | NIL | COMPLY pending confirmation of latex free
42142502-00024 | 19 | Needle, spinal, quincke bevel; luer lock. The needle must be colour coded at the stylet cap and the size must be clearly stated. Must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. Must be bevelled at 220degrees Needle size: 22G extra-length with a needle length of 110mm excluding hub Needle and stylet must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 25) |  | BAS MEDEXPRESS | 776.25 | Mercury | SN1890Q |  | India | 1 - 3 days | 15% | N/A | As per client | Yes | Box | No sample submitted for evaluation
42142502-00025 | 20 | Needle, spinal, quincke bevel; luer lock. The needle must be colour coded at the stylet cap and the size must be clearly stated. Must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. Must be bevelled at 220degrees Needle size: 23G with a needle length of 90mm excluding hub Needle and stylet must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 25) | BASIC LINE OFFER | A R FREIGHT PRIVATE LIMITED | 76.5 | SUMBOW | N/A | SUMBOW MEDICAL | CHINA | 60 DAYS | 85% | USD | 100 | YES | BOX OF 25 | COMPLY pending confirmation of latex free |  | Recommended | 100
42142502-00025 | 20 | Needle, spinal, quincke bevel; luer lock. The needle must be colour coded at the stylet cap and the size must be clearly stated. Must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. Must be bevelled at 220degrees Needle size: 23G with a needle length of 90mm excluding hub Needle and stylet must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 25) |  | Medi-Core Technologies | 3.32 | ULIFE | 42142502-00025 | JIAXING U-LIFE MEDICAL DEVICE TECHNOLOGY.,CO LTD | CHINA | 21 | 80 | USD | 1 box | YES | BOX | COMPLY pending confirmation of latex free |  |  |  | COMPLY pending confirmation of latex free | NO TEST REPORTS
42142502-00025 | 20 | Needle, spinal, quincke bevel; luer lock. The needle must be colour coded at the stylet cap and the size must be clearly stated. Must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. Must be bevelled at 220degrees Needle size: 23G with a needle length of 90mm excluding hub Needle and stylet must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 25) | N/A | PROMED TECHNOLOGIES (PTY) LTD | 3.34 | U-LIFE | 502025 | JIAXING | CHINA | 14 DAYS | 85% | USD | 1 BOX | YES | BOX | COMPLY pending confirmation of latex free |  |  |  | COMPLY pending confirmation of latex free | NO TEST REPORTS
42142502-00025 | 20 | Needle, spinal, quincke bevel; luer lock. The needle must be colour coded at the stylet cap and the size must be clearly stated. Must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. Must be bevelled at 220degrees Needle size: 23G with a needle length of 90mm excluding hub Needle and stylet must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 25) |  | Chemical World cc | 3.39 | MEDIJECT | NDSSFG23 | RUNQIANG | CHINA | 3 | 85 | USD | 1000 | YES | 50 | COMPLY pending confirmation of latex free
42142502-00025 | 20 | Needle, spinal, quincke bevel; luer lock. The needle must be colour coded at the stylet cap and the size must be clearly stated. Must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. Must be bevelled at 220degrees Needle size: 23G with a needle length of 90mm excluding hub Needle and stylet must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 25) | Basic Line Offer | BIOCOR HOSPITAL SUPPLIES PTY LTD | 4.15 | Medline | N/A | Yangzhong Medline Industry Co.Ltd | China | 21 | 85 | USD | 10 Boxes | Yes | N/A | COMPLY pending confirmation of latex free
42142502-00025 | 20 | Needle, spinal, quincke bevel; luer lock. The needle must be colour coded at the stylet cap and the size must be clearly stated. Must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. Must be bevelled at 220degrees Needle size: 23G with a needle length of 90mm excluding hub Needle and stylet must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 25) |  | UNITRADE 1032 CC | 4.37 | SPINOJEC | 1 | SHANGHAI KINDLY ENTERPRISE DEVELOPMENT GROUP | ROC | 21 | 85 | USD | 400 | YES | 25 | No sample submitted for evaluation |  |  |  | COMPLY pending confirmation of latex free | NO TEST REPORTS
42142502-00025 | 20 | Needle, spinal, quincke bevel; luer lock. The needle must be colour coded at the stylet cap and the size must be clearly stated. Must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. Must be bevelled at 220degrees Needle size: 23G with a needle length of 90mm excluding hub Needle and stylet must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 25) | 1 | GROBIR MEDICAL SUPPLIERS | 4.73 | DURACAN | DQ.10.023.090 | PROMECON | CHINA | 21 | 80 | US DOLLAR | 25 | Yes | UOM AS SPECIFIED / MEETS SPECIFICATION | No sample submitted for evaluation.
42142502-00025 | 20 | Needle, spinal, quincke bevel; luer lock. The needle must be colour coded at the stylet cap and the size must be clearly stated. Must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. Must be bevelled at 220degrees Needle size: 23G with a needle length of 90mm excluding hub Needle and stylet must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 25) | N/A | SAFARMEX (PTY) LTD | 4.98 | Saclin | To be confirmed | Zhejiang Mediunion Healthcare Group Co. Ltd | China | 3days | 80% | US$ | 10000 Pieces | Yes | 100 Pieces in a Box | No sample submitted for evaluation |  |  |  | No sample submitted for evaluation | NO TEST REPORTS
42142502-00025 | 20 | Needle, spinal, quincke bevel; luer lock. The needle must be colour coded at the stylet cap and the size must be clearly stated. Must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. Must be bevelled at 220degrees Needle size: 23G with a needle length of 90mm excluding hub Needle and stylet must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 25) |  | Zanotone (Pty) Ltd | 6.39 | Mekon |  | Mekon | [6 empty] | Each | COMPLY pending confirmation of latex free
42142502-00025 | 20 | Needle, spinal, quincke bevel; luer lock. The needle must be colour coded at the stylet cap and the size must be clearly stated. Must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. Must be bevelled at 220degrees Needle size: 23G with a needle length of 90mm excluding hub Needle and stylet must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 25) | OFFER 1 | SUPRA HEALTHCARE JOHANNESBURG PTY LTD | 7.44 | REVITAL | SPNQ23X90 | Revital Health | KENYA | Initial 60 days, thereafter 14 | 84 | USD | 1 Box | YES |  | COMPLY pending confirmation of latex free |  |  |  | No sample submitted for evaluation | NO TEST REPORTS
42142502-00025 | 20 | Needle, spinal, quincke bevel; luer lock. The needle must be colour coded at the stylet cap and the size must be clearly stated. Must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. Must be bevelled at 220degrees Needle size: 23G with a needle length of 90mm excluding hub Needle and stylet must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 25) | Basic Line offer | Becton Dickinson | 9 | BD Quinke Spinal Needle 23G X 90mm | 405240 | Becton Dickinson | Fraga, Spain | 7 Working Days | 76% | USD | 1 BOX | YES | 28 units in a box | No sample submitted for evaluation.
42142502-00025 | 20 | Needle, spinal, quincke bevel; luer lock. The needle must be colour coded at the stylet cap and the size must be clearly stated. Must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. Must be bevelled at 220degrees Needle size: 23G with a needle length of 90mm excluding hub Needle and stylet must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 25) |  | ALMIKA TRADING PTY LTD | 10.58 | BD | 405240 | BD | USA | 14 | 0 | NO | 100 | YES | 1 | No sample submitted for evaluation.
42142502-00025 | 20 | Needle, spinal, quincke bevel; luer lock. The needle must be colour coded at the stylet cap and the size must be clearly stated. Must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. Must be bevelled at 220degrees Needle size: 23G with a needle length of 90mm excluding hub Needle and stylet must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 25) |  | LECHOBA MEDICAL TECHNOLOGIES (PTY) LTD | 12.99 | UNISIS UNIEVER | 1160961 | UNISIS CORPORATION | JAPAN | 21 DAYS | 75% | USD | 25 | YES | 25/BOX | No sample submitted for evaluation
42142502-00025 | 20 | Needle, spinal, quincke bevel; luer lock. The needle must be colour coded at the stylet cap and the size must be clearly stated. Must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. Must be bevelled at 220degrees Needle size: 23G with a needle length of 90mm excluding hub Needle and stylet must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 25) | BOX OF 50 | MSIMED CC | 13.5 | KANG KANG | SPN003 | ZHENJIANG | CHINA | 25 | 20% | USD | 20000PCS | YES | BOX | No sample submitted for evaluation |  |  |  | No sample submitted for evaluation | NO TEST REPORTS
42142502-00025 | 20 | Needle, spinal, quincke bevel; luer lock. The needle must be colour coded at the stylet cap and the size must be clearly stated. Must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. Must be bevelled at 220degrees Needle size: 23G with a needle length of 90mm excluding hub Needle and stylet must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 25) | OFFER 2 | SUPRA HEALTHCARE JOHANNESBURG PTY LTD | 14.2 | Egemen | SPNQ23X90 | Egemen | TURKEY | Initial 60 days, thereafter 14 | 84 | USD | 1 Box | YES | 23GX90MM | COMPLY pending confirmation of latex free |  |  |  | COMPLY pending confirmation of latex free | NO TEST REPORTS
42142502-00025 | 20 | Needle, spinal, quincke bevel; luer lock. The needle must be colour coded at the stylet cap and the size must be clearly stated. Must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. Must be bevelled at 220degrees Needle size: 23G with a needle length of 90mm excluding hub Needle and stylet must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 25) |  | Viking Critical Care (Pty) Ltd | 46.95 | VYGON | 69F518076 | VYGON | FRANCE | EX. STOCK 30 - 60 DAYS | 80% | EURO | 1 | NO, OUR PRODUCT IS 75MM | 1 | No sample submitted for evaluation |  |  |  | No sample submitted for evaluation
42142502-00025 | 20 | Needle, spinal, quincke bevel; luer lock. The needle must be colour coded at the stylet cap and the size must be clearly stated. Must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. Must be bevelled at 220degrees Needle size: 23G with a needle length of 90mm excluding hub Needle and stylet must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 25) |  | JUMLA MEDICAL SUPPLIES (PTY) LTD | 340.12 | EGEMEN | TSPQ2390 | TIBBI MEDICAL | TURKEY | 21 DAYS | 70% | US DOLLAR | 2 BOXES | YES | NIL | COMPLY pending confirmation of latex free
42142502-00025 | 20 | Needle, spinal, quincke bevel; luer lock. The needle must be colour coded at the stylet cap and the size must be clearly stated. Must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. Must be bevelled at 220degrees Needle size: 23G with a needle length of 90mm excluding hub Needle and stylet must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 25) |  | BAS MEDEXPRESS | 776.25 | Mercury | SN2290Q |  | India | 1 - 3 days | 15% | N/A | As per client | Yes | Box | No sample submitted for evaluation
42142502-00026 | 21 | Needle, spinal, quincke bevel; luer lock. The needle must be colour coded at the stylet cap and the size must be clearly stated. Must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. Must be bevelled at 220degrees Needle size: 25G with a needle length of 25mm ±2mm excluding hub Needle and stylet must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 25) | BASIC LINE OFFER | A R FREIGHT PRIVATE LIMITED | 3.06 | SUMBOW | N/A | SUMBOW MEDICAL | CHINA | 60 DAYS | 85% | USD | 100 | YES | N/A | Offered a 26G X 90mm instead of 25G X 25mm |  |  |  | No sample submitted for evaluation
42142502-00026 | 21 | Needle, spinal, quincke bevel; luer lock. The needle must be colour coded at the stylet cap and the size must be clearly stated. Must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. Must be bevelled at 220degrees Needle size: 25G with a needle length of 25mm ±2mm excluding hub Needle and stylet must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 25) |  | Medi-Core Technologies | 3.32 | ULIFE | 42142502-00026 | JIAXING U-LIFE MEDICAL DEVICE TECHNOLOGY.,CO LTD | CHINA | 21 | 80 | USD | 1 box | YES | BOX | No sample submitted for evaluation |  |  |  | offered an eccentric instead of concentric nozzle offered a 24ml instead of 20ml | NO TEST REPORTS
42142502-00026 | 21 | Needle, spinal, quincke bevel; luer lock. The needle must be colour coded at the stylet cap and the size must be clearly stated. Must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. Must be bevelled at 220degrees Needle size: 25G with a needle length of 25mm ±2mm excluding hub Needle and stylet must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 25) | N/A | PROMED TECHNOLOGIES (PTY) LTD | 3.34 | U-LIFE | 502026 | JIAXING | CHINA | 14 DAYS | 85% | USD | 1 BOX | YES | BOX | No samples submitted for evaluation |  |  |  | No sample submitted for evaluation | NO TEST REPORTS
42142502-00026 | 21 | Needle, spinal, quincke bevel; luer lock. The needle must be colour coded at the stylet cap and the size must be clearly stated. Must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. Must be bevelled at 220degrees Needle size: 25G with a needle length of 25mm +-2mm excluding hub Needle and stylet must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 25) |  | Chemical World cc | 4.03 | [10 empty] | No sample submitted for evaluation |  |  |  | Offered an insulin syringe of 0.5ml instead of a two part luer 5ml syringe | NO TEST REPORTS
42142502-00026 | 21 | Needle, spinal, quincke bevel; luer lock. The needle must be colour coded at the stylet cap and the size must be clearly stated. Must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. Must be bevelled at 220degrees Needle size: 25G with a needle length of 25mm ±2mm excluding hub Needle and stylet must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 25) | Basic Line Offer | BIOCOR HOSPITAL SUPPLIES PTY LTD | 103.55 | Medline | N/A | Yangzhong Medline Industry Co.Ltd | China | 60 | 85 | USD | 10 Boxes | Yes | Box /25 | COMPLY pending confirmation of latex free |  | Recommended | 100 | No sample submitted for evaluation | NO TEST REPORTS
42142502-00026 | 21 | Needle, spinal, quincke bevel; luer lock. The needle must be colour coded at the stylet cap and the size must be clearly stated. Must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. Must be bevelled at 220degrees Needle size: 25G with a needle length of 25mm ±2mm excluding hub Needle and stylet must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 25) | 1 | GROBIR MEDICAL SUPPLIERS | 4.73 | DURACAN | DQ.10.025.025 | PROMECON | CHINA | 21 | 80 | US DOLLAR | 25 | Yes | UOM AS SPECIFIED / MEETS SPECIFICATION | No sample submitted for evaluation. |  |  |  | Offered a three part 12ML syringe instead of a two part 10ml syringe | NO TEST REPORTS
42142502-00026 | 21 | Needle, spinal, quincke bevel; luer lock. The needle must be colour coded at the stylet cap and the size must be clearly stated. Must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. Must be bevelled at 220degrees Needle size: 25G with a needle length of 25mm ±2mm excluding hub Needle and stylet must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 25) | N/A | SAFARMEX (PTY) LTD | 5.38 | Saclin | To be confirmed | Zhejiang Mediunion Healthcare Group Co. Ltd | China | 3days | 80% | US$ | 10000 Pieces | Yes | Box of 25 | No sample submitted for evaluation |  |  |  | offered an eccentric instead of concentric nozzle | NO TEST REPORTS
42142502-00026 | 21 | Needle, spinal, quincke bevel; luer lock. The needle must be colour coded at the stylet cap and the size must be clearly stated. Must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. Must be bevelled at 220degrees Needle size: 25G with a needle length of 25mm ±2mm excluding hub Needle and stylet must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 25) |  | Zanotone (Pty) Ltd | 6.39 | Mekon |  | Mekon | [6 empty] | EACH | Offered 26G instead of 25G
42142502-00026 | 21 | Needle, spinal, quincke bevel; luer lock. The needle must be colour coded at the stylet cap and the size must be clearly stated. Must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. Must be bevelled at 220degrees Needle size: 25G with a needle length of 25mm ±2mm excluding hub Needle and stylet must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 25) |  | SUPRA HEALTHCARE JOHANNESBURG PTY LTD | 7.44 | REVITAL | SPNQ25X25 | Revital Health | KENYA | Initial 60 days, thereafter 14 | 84 | USD | 1 Box | YES |  | Offered 26G instead of 25G |  |  |  | No sample submitted for evaluation | NO TEST REPORTS
42142502-00026 | 21 | Needle, spinal, quincke bevel; luer lock. The needle must be colour coded at the stylet cap and the size must be clearly stated. Must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. Must be bevelled at 220degrees Needle size: 25G with a needle length of 25mm ±2mm excluding hub Needle and stylet must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 25) | BOX OF 50 | MSIMED CC | 14.5 | KANG KANG | SPN004 | ZHENJIANG | CHINA | 25 | 20% | USD | 20000PCS | YES | BOX | No sample submitted for evaluation |  |  |  | Offered a three part instead of a two part. | NO TEST REPORTS
42142502-00026 | 21 | Needle, spinal, quincke bevel; luer lock. The needle must be colour coded at the stylet cap and the size must be clearly stated. Must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. Must be bevelled at 220degrees Needle size: 25G with a needle length of 25mm ±2mm excluding hub Needle and stylet must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 25) | Basic Line offer | Becton Dickinson | 18 | BD Quinke Spinal Needle 25G X 25mm | 405243 | Becton Dickinson | Fraga, Spain | 7 Working Days | 76% | USD | 1 BOX | YES | 29 units in a box | No sample submitted for evaluation. |  |  |  | Offered a three part instead of a two part. | NO TEST REPORTS
42142502-00026 | 21 | Needle, spinal, quincke bevel; luer lock. The needle must be colour coded at the stylet cap and the size must be clearly stated. Must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. Must be bevelled at 220degrees Needle size: 25G with a needle length of 25mm ±2mm excluding hub Needle and stylet must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 25) |  | ALMIKA TRADING PTY LTD | 19.84 | BD | 405243 | BD | USA | 14 | 0 | NO | 100 | YES | 1 | No sample submitted for evaluation. |  |  |  | Wrong sample, Offered a hypodermic needle 22G X 32mm instead of a blood collecting needle
42142502-00026 | 21 | Needle, spinal, quincke bevel; luer lock. The needle must be colour coded at the stylet cap and the size must be clearly stated. Must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. Must be bevelled at 220degrees Needle size: 25G with a needle length of 25mm ±2mm excluding hub Needle and stylet must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 25) |  | LECHOBA MEDICAL TECHNOLOGIES (PTY) LTD | 23.94 | UNISIS UNIEVER | 5240561 | UNISIS CORPORATION | JAPAN | 21 DAYS | 75% | USD | 25 | YES | 25/BOX | No sample submitted for evaluation |  |  |  | No sample submitted for evaluation | NO TEST REPORTS
42142502-00026 | 21 | Needle, spinal, quincke bevel; luer lock. The needle must be colour coded at the stylet cap and the size must be clearly stated. Must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. Must be bevelled at 220degrees Needle size: 25G with a needle length of 25mm ±2mm excluding hub Needle and stylet must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 25) |  | Viking Critical Care (Pty) Ltd | 71.63 | VYGON | 69F518036 | VYGON | FRANCE | EX. STOCK 30 - 60 DAYS | 80% | EURO | 1 | YES | 1 | Offered a needle length of 30mm instead of 25mm |  |  |  | No sample submitted for evaluation | NO TEST REPORTS
42142502-00026 | 21 | Needle, spinal, quincke bevel; luer lock. The needle must be colour coded at the stylet cap and the size must be clearly stated. Must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. Must be bevelled at 220degrees Needle size: 25G with a needle length of 25mm ±2mm excluding hub Needle and stylet must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 25) |  | JUMLA MEDICAL SUPPLIES (PTY) LTD | 340.12 | EGEMEN | TSPQ2690 | TIBBI MEDICAL | TURKEY | 21 DAYS | 70% | US DOLLAR | 2 BOXES | YES | NIL | No sample submitted for evaluation |  |  |  | No sample submitted for evaluation | NO TEST REPORTS
42142502-00026 | 21 | Needle, spinal, quincke bevel; luer lock. The needle must be colour coded at the stylet cap and the size must be clearly stated. Must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. Must be bevelled at 220degrees Needle size: 25G with a needle length of 25mm ±2mm excluding hub Needle and stylet must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 25) |  | BAS MEDEXPRESS | 776.25 | Mercury | SN2390Q |  | India | 1 - 3 days | 15% | N/A | As per client | Yes | Box | No sample submitted for evaluation |  |  |  | No sample submitted for evaluation | NO TEST REPORTS
42142502-00027 | 22 | Needle, spinal, quincke bevel; luer lock. The needle must be colour coded at the stylet cap and the size must be clearly stated. Must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. Must be bevelled at 220degrees Needle size: 25G with a needle length of 90mm excluding hub Needle and stylet must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 25) | BASIC LINE OFFER | A R FREIGHT PRIVATE LIMITED | 76.5 | SUMBOW | N/A | SUMBOW MEDICAL | CHINA | 60 DAYS | 85% | USD | 100 | YES | BOX OF 25 | COMPLY pending confirmation of latex free |  | Recommended | 100 | Item deleted | NO TEST REPORTS
42142502-00027 | 22 | Needle, spinal, quincke bevel; luer lock. The needle must be colour coded at the stylet cap and the size must be clearly stated. Must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. Must be bevelled at 220degrees Needle size: 25G with a needle length of 90mm excluding hub Needle and stylet must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 25) |  | Medi-Core Technologies | 3.32 | ULIFE | 42142502-00027 | JIAXING U-LIFE MEDICAL DEVICE TECHNOLOGY.,CO LTD | CHINA | 21 | 80 | USD | 1 box | YES | BOX | COMPLY pending confirmation of latex free |  |  |  | Offered 6ml instead Nominal capacity: 5ml Offered 38ml instead of 32ml The sample does not have a refractive safety mechanism | NO TEST REPORTS
42142502-00027 | 22 | Needle, spinal, quincke bevel; luer lock. The needle must be colour coded at the stylet cap and the size must be clearly stated. Must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. Must be bevelled at 220degrees Needle size: 25G with a needle length of 90mm excluding hub Needle and stylet must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 25) |  | PROMED TECHNOLOGIES (PTY) LTD | 3.34 | ULIFE |  | JIAXING | [6 empty] | 25 | COMPLY pending confirmation of latex free
42142502-00027 | 22 | Needle, spinal, quincke bevel; luer lock. The needle must be colour coded at the stylet cap and the size must be clearly stated. Must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. Must be bevelled at 220degrees Needle size: 25G with a needle length of 90mm excluding hub Needle and stylet must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 25) |  | Chemical World cc | 3.57 | MEDIJECT | NDSSFG25 | RUNQIANG | CHINA | 3 | 85 | USD | 1000 | YES | 50 | COMPLY pending confirmation of latex free |  |  |  | COMPLY pending confirmation of latex and DEHP free | NO TEST REPORTS
42142502-00027 | 22 | Needle, spinal, quincke bevel; luer lock. The needle must be colour coded at the stylet cap and the size must be clearly stated. Must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. Must be bevelled at 220degrees Needle size: 25G with a needle length of 90mm excluding hub Needle and stylet must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 25) | Basic Line Offer | BIOCOR HOSPITAL SUPPLIES PTY LTD | 4.15 | Medline | N/A | Yangzhong Medline Industry Co.Ltd | China | 60 | 85 | USD | 10 Boxes | Yes | N/A | COMPLY pending confirmation of latex free |  |  |  | No sample submitted for evaluation | NO TEST REPORTS
42142502-00027 | 22 | Needle, spinal, quincke bevel; luer lock. The needle must be colour coded at the stylet cap and the size must be clearly stated. Must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. Must be bevelled at 220degrees Needle size: 25G with a needle length of 90mm excluding hub Needle and stylet must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 25) |  | UNITRADE 1032 CC | 4.37 | SPINOJEC | 1 | SHANGHAI KINDLY ENTERPRISE DEVELOPMENT GROUP | ROC | 21 | 85 | USD | 400 | YES | 25 | No sample submitted for evaluation
42142502-00027 | 22 | Needle, spinal, quincke bevel; luer lock. The needle must be colour coded at the stylet cap and the size must be clearly stated. Must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. Must be bevelled at 220degrees Needle size: 25G with a needle length of 90mm excluding hub Needle and stylet must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 25) | 1 | GROBIR MEDICAL SUPPLIERS | 4.73 | DURACAN | DQ.10.025.090 | PROMECON | CHINA | 21 | 80 | US DOLLAR | 25 | Yes | UOM AS SPECIFIED / MEETS SPECIFICATION | COMPLY pending confirmation of latex free |  |  |  | Offered 29G instead of 30G
42142502-00027 | 22 | Needle, spinal, quincke bevel; luer lock. The needle must be colour coded at the stylet cap and the size must be clearly stated. Must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. Must be bevelled at 220degrees Needle size: 25G with a needle length of 90mm excluding hub Needle and stylet must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 25) | N/A | SAFARMEX (PTY) LTD | 4.98 | Saclin | To be confirmed | Zhejiang Mediunion Healthcare Group Co. Ltd | China | 3days | 80% | US$ | 10000 Pieces | Yes | Box of 25 | No sample submitted for evaluation |  |  |  | No sample submitted for evaluation | NO TEST REPORTS
42142502-00027 | 22 | Needle, spinal, quincke bevel; luer lock. The needle must be colour coded at the stylet cap and the size must be clearly stated. Must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. Must be bevelled at 220degrees Needle size: 25G with a needle length of 90mm excluding hub Needle and stylet must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 25) |  | Zanotone (Pty) Ltd | 6.39 | Mekon |  | Mekon | [6 empty] | Each | COMPLY pending confirmation of latex free
42142502-00027 | 22 | Needle, spinal, quincke bevel; luer lock. The needle must be colour coded at the stylet cap and the size must be clearly stated. Must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. Must be bevelled at 220degrees Needle size: 25G with a needle length of 90mm excluding hub Needle and stylet must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 25) | OFFER 1 | SUPRA HEALTHCARE JOHANNESBURG PTY LTD | 7.44 | REVITAL | SPNQ25X90 | Revital Health | KENYA | Initial 60 days, thereafter 14 | 84 | USD | 1 Box | YES |  | COMPLY pending confirmation of latex free
42142502-00027 | 22 | Needle, spinal, quincke bevel; luer lock. The needle must be colour coded at the stylet cap and the size must be clearly stated. Must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. Must be bevelled at 220degrees Needle size: 25G with a needle length of 90mm excluding hub Needle and stylet must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 25) | Basic Line offer | Becton Dickinson | 9 | BD Quinke Spinal Needle 25G X 90mm | 405257 | Becton Dickinson | Fraga, Spain | 7 Working Days | 76% | USD | 1 BOX | YES | 30 units in a box | COMPLY pending confirmation of latex free |  |  |  | Offered a 30G of 8mm instead of a 29G of 12,5 - 12,7 mm | NO TEST REPORTS
42142502-00027 | 22 | Needle, spinal, quincke bevel; luer lock. The needle must be colour coded at the stylet cap and the size must be clearly stated. Must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. Must be bevelled at 220degrees Needle size: 25G with a needle length of 90mm excluding hub Needle and stylet must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 25) |  | ALMIKA TRADING PTY LTD | 10.58 | BD | 405257 | BD | USA | 14 | 0 | NO | 100 | YES | 1 | COMPLY pending confirmation of latex free |  |  |  | No sample submitted for evaluation | NO TEST REPORTS
42142502-00027 | 22 | Needle, spinal, quincke bevel; luer lock. The needle must be colour coded at the stylet cap and the size must be clearly stated. Must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. Must be bevelled at 220degrees Needle size: 25G with a needle length of 90mm excluding hub Needle and stylet must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 25) |  | LECHOBA MEDICAL TECHNOLOGIES (PTY) LTD | 12.99 | UNISIS UNIEVER | 1182961 | UNISIS CORPORATION | JAPAN | 21 DAYS | 75% | USD | 25 | YES | 25/BOX | No sample submitted for evaluation |  |  |  | Wrong sample, offered hypodermic needle 22G instead of 1ml TB syringe | NO TEST REPORTS
42142502-00027 | 22 | Needle, spinal, quincke bevel; luer lock. The needle must be colour coded at the stylet cap and the size must be clearly stated. Must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. Must be bevelled at 220degrees Needle size: 25G with a needle length of 90mm excluding hub Needle and stylet must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 25) | BOX OF 50 | MSIMED CC | 13.2 | KANG KANG | SPN005 | ZHENJIANG | CHINA | 25 | 20% | USD | 20000PCS | YES | BOX | No sample submitted for evaluation |  |  |  | No sample submitted for evaluation | NO TEST REPORTS
42142502-00027 | 22 | Needle, spinal, quincke bevel; luer lock. The needle must be colour coded at the stylet cap and the size must be clearly stated. Must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. Must be bevelled at 220degrees Needle size: 25G with a needle length of 90mm excluding hub Needle and stylet must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 25) | OFFER 2 | SUPRA HEALTHCARE JOHANNESBURG PTY LTD | 14.2 | Egemen | SPNQ25X90 | Egemen | TURKEY | Initial 60 days, thereafter 14 | 84 | USD | 1 Box | YES |  | COMPLY pending confirmation of latex free
42142502-00027 | 22 | Needle, spinal, quincke bevel; luer lock. The needle must be colour coded at the stylet cap and the size must be clearly stated. Must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. Must be bevelled at 220degrees Needle size: 25G with a needle length of 90mm excluding hub Needle and stylet must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 25) | 1 | B. BRAUN MEDICAL (PTY) LTD | 18 | Spinocan 25G x 3 1/2" (0.53 x 88mm) | 4505905 | B. Braun | Germany | 42 | 50% | Euro | 25 | Yes | Box of 25 | COMPLY pending cpnfirmation of latex free
42142502-00027 | 22 | Needle, spinal, quincke bevel; luer lock. The needle must be colour coded at the stylet cap and the size must be clearly stated. Must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. Must be bevelled at 220degrees Needle size: 25G with a needle length of 90mm excluding hub Needle and stylet must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 25) |  | Viking Critical Care (Pty) Ltd | 77.21 | VYGON | 69F18005 | VYGON | FRANCE | EX. STOCK 30 - 60DAYS | 80% | EURO | 1 | YES | 1 | No sample submitted for evaluation
42142502-00027 | 22 | Needle, spinal, quincke bevel; luer lock. The needle must be colour coded at the stylet cap and the size must be clearly stated. Must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. Must be bevelled at 220degrees Needle size: 25G with a needle length of 90mm excluding hub Needle and stylet must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 25) |  | JUMLA MEDICAL SUPPLIES (PTY) LTD | 340.12 | EGEMEN | TSPQ2590 | TIBBI MEDICAL | TURKEY | 21 DAYS | 70% | US DOLLAR | 2 BOXES | YES | NIL | COMPLY pending confirmation of latex free |  |  |  | No sample submitted for evaluation | NO TEST REPORTS
42142502-00027 | 22 | Needle, spinal, quincke bevel; luer lock. The needle must be colour coded at the stylet cap and the size must be clearly stated. Must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. Must be bevelled at 220degrees Needle size: 25G with a needle length of 90mm excluding hub Needle and stylet must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 25) |  | BAS MEDEXPRESS | 776.25 | Mercury | SN2490Q |  | India | 1 - 3 days | 15% | N/A | As per client | Yes | Box | No sample submitted for evaluation |  |  |  | No sample submitted for evaluation | NO TEST REPORTS
42142502-00026 | 23 | Needle, spinal, quincke bevel; luer lock. The needle must be colour coded at the stylet cap and the size must be clearly stated. Must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. Must be bevelled at 220degrees Needle size: 26G with a needle length of 90mm excluding hub Needle and stylet must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 25) | BASIC LINE OFFER | A R FREIGHT PRIVATE LIMITED | 3.06 | SUMBOW | N/A | SUMBOW MEDICAL | CHINA | 60 DAYS | 85% | USD | 100 | YES | N/A | Offered a 26G X 90mm instead of 25G X 25mm |  | Re-tender endusers to revise specifications
42142502-00026 | 23 | Needle, spinal, quincke bevel; luer lock. The needle must be colour coded at the stylet cap and the size must be clearly stated. Must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. Must be bevelled at 220degrees Needle size: 26G with a needle length of 90mm excluding hub Needle and stylet must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 25) |  | Medi-Core Technologies | 3.32 | ULIFE | 42142502-00026 | JIAXING U-LIFE MEDICAL DEVICE TECHNOLOGY.,CO LTD | CHINA | 21 | 80 | USD | 1 box | YES | BOX | No sample submitted for evaluation |  | Re-tender endusers to revise specifications
42142502-00026 | 23 | Needle, spinal, quincke bevel; luer lock. The needle must be colour coded at the stylet cap and the size must be clearly stated. Must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. Must be bevelled at 220degrees Needle size: 26G with a needle length of 90mm excluding hub Needle and stylet must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 25) |  | Chemical World cc | 4.03 | [10 empty] | No sample submitted for evaluation |  | Re-tender endusers to revise specifications
42142502-00026 | 23 | Needle, spinal, quincke bevel; luer lock. The needle must be colour coded at the stylet cap and the size must be clearly stated. Must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. Must be bevelled at 220degrees Needle size: 26G with a needle length of 90mm excluding hub Needle and stylet must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 25) |  | UNITRADE 1032 CC | 4.37 | SPINOJEC | 1 | SHANGHAI KINDLY ENTERPRISE DEVELOPMENT GROUP | ROC | 21 | 85 | USD | 400 | YES | 25 | No sample submitted for evaluation |  | Re-tender endusers to revise specifications
42142502-00026 | 23 | Needle, spinal, quincke bevel; luer lock. The needle must be colour coded at the stylet cap and the size must be clearly stated. Must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. Must be bevelled at 220degrees Needle size: 26G with a needle length of 90mm excluding hub Needle and stylet must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 25) | 1 | GROBIR MEDICAL SUPPLIERS | 4.73 | DURACAN | DQ.10.026.090 | PROMECON | CHINA | 21 | 80 | US DOLLAR | 25 | Yes | UOM AS SPECIFIED / MEETS SPECIFICATION | No sample submitted for evaluation. |  | Re-tender endusers to revise specifications
42142502-00026 | 23 | Needle, spinal, quincke bevel; luer lock. The needle must be colour coded at the stylet cap and the size must be clearly stated. Must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. Must be bevelled at 220degrees Needle size: 26G with a needle length of 90mm excluding hub Needle and stylet must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 25) | N/A | SAFARMEX (PTY) LTD | 5.38 | Saclin | To be confirmed | Zhejiang Mediunion Healthcare Group Co. Ltd | China | 3days | 80% | US$ | 10000 Pieces | Yes | Box of 25 | No sample submitted for evaluation |  | Re-tender endusers to revise specifications
42142502-00026 | 23 | Needle, spinal, quincke bevel; luer lock. The needle must be colour coded at the stylet cap and the size must be clearly stated. Must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. Must be bevelled at 220degrees Needle size: 26G with a needle length of 90mm excluding hub Needle and stylet must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 25) | Basic Line offer | Becton Dickinson | 9 | BD Quinke Spinal Needle 26G X 90mm | 405258 | Becton Dickinson | Fraga, Spain | 7 Working Days | 76% | USD | 1 BOX | YES | 31 units in a box | No sample submitted for evaluation. |  | Re-tender endusers to revise specifications
42142502-00026 | 23 | Needle, spinal, quincke bevel; luer lock. The needle must be colour coded at the stylet cap and the size must be clearly stated. Must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. Must be bevelled at 220degrees Needle size: 26G with a needle length of 90mm excluding hub Needle and stylet must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 25) |  | ALMIKA TRADING PTY LTD | 10.58 | BD | 405258 | BD | USA | 14 | 0 | NO | 100 | YES | 1 | No sample submitted for evaluation. |  | Re-tender endusers to revise specifications
42142502-00026 | 23 | Needle, spinal, quincke bevel; luer lock. The needle must be colour coded at the stylet cap and the size must be clearly stated. Must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. Must be bevelled at 220degrees Needle size: 26G with a needle length of 90mm excluding hub Needle and stylet must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 25) | BOX OF 50 | MSIMED CC | 13.5 | KANG KANG | SPN006 | ZHENJIANG | CHINA | 25 | 20% | USD | 20000PCS | YES | BOX | No sample submitted for evaluation |  | Re-tender endusers to revise specifications
42142502-00026 | 23 | Needle, spinal, quincke bevel; luer lock. The needle must be colour coded at the stylet cap and the size must be clearly stated. Must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. Must be bevelled at 220degrees Needle size: 26G with a needle length of 90mm excluding hub Needle and stylet must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 25) | 1 | B. BRAUN MEDICAL (PTY) LTD | 18 | Spinocn 26G x 3 1/2" (0.47 x 88mm) | 4502906 | B. Braun | Germany | 42 | 50% | Euro | 25 | Yes | Box of 25 | Offered 28G instead of 25G |  | Re-tender endusers to revise specifications
42142502-00026 | 23 | Needle, spinal, quincke bevel; luer lock. The needle must be colour coded at the stylet cap and the size must be clearly stated. Must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. Must be bevelled at 220degrees Needle size: 26G with a needle length of 90mm excluding hub Needle and stylet must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 25) |  | LECHOBA MEDICAL TECHNOLOGIES (PTY) LTD | 23.94 | UNISIS UNIEVER | 5240561 | UNISIS CORPORATION | JAPAN | 21 DAYS | 75% | USD | 25 | YES | 25/BOX | No sample submitted for evaluation |  | Re-tender endusers to revise specifications
42142502-00026 | 23 | Needle, spinal, quincke bevel; luer lock. The needle must be colour coded at the stylet cap and the size must be clearly stated. Must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. Must be bevelled at 220degrees Needle size: 26G with a needle length of 90mm excluding hub Needle and stylet must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 25) |  | Viking Critical Care (Pty) Ltd | 47.55 | VYGON | 69F518035 | VYGON | FRANCE | EX. STOCK 30-60DAYS | 80% | EURO | 1 | YES | 1 | Offered a needle length of 30mm instead of 25mm |  | Re-tender endusers to revise specifications
42142502-00026 | 23 | Needle, spinal, quincke bevel; luer lock. The needle must be colour coded at the stylet cap and the size must be clearly stated. Must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. Must be bevelled at 220degrees Needle size: 26G with a needle length of 90mm excluding hub Needle and stylet must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 25) |  | JUMLA MEDICAL SUPPLIES (PTY) LTD | 340.12 | EGEMEN | TSPQ2690 | TIBBI MEDICAL | TURKEY | 21 DAYS | 70% | US DOLLAR | 2 BOXES | YES | NIL | No sample submitted for evaluation |  | Re-tender endusers to revise specifications
42142502-00026 | 23 | Needle, spinal, quincke bevel; luer lock. The needle must be colour coded at the stylet cap and the size must be clearly stated. Must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. Must be bevelled at 220degrees Needle size: 26G with a needle length of 90mm excluding hub Needle and stylet must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 25) |  | BAS MEDEXPRESS | 776.25 | Mercury | SN2390Q |  | India | 1 - 3 days | 15% | N/A | As per client | Yes | Box | No sample submitted for evaluation |  | Re-tender endusers to revise specifications
42142502-00028 | 24 | Needle, spinal, quincke bevel; luer lock. The needle must be colour coded at the stylet cap and the size must be clearly stated. Must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. Must be bevelled at 220degrees Needle size: 27G with a needle length of 90mm excluding hub Needle and stylet must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 25) | BASIC LINE OFFER | A R FREIGHT PRIVATE LIMITED | 76.5 | SUMBOW | N/A | SUMBOW MEDICAL | CHINA | 60 DAYS | 85% | USD | 100 | YES | BOX OF 25 | COMPLY pending confirmation of latex free |  | Recommended | 100
42142502-00028 | 24 | Needle, spinal, quincke bevel; luer lock. The needle must be colour coded at the stylet cap and the size must be clearly stated. Must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. Must be bevelled at 220degrees Needle size: 27G with a needle length of 90mm excluding hub Needle and stylet must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 25) |  | Medi-Core Technologies | 3.32 | ULIFE | 42142502-00028 | JIAXING U-LIFE MEDICAL DEVICE TECHNOLOGY.,CO LTD | CHINA |  | 80 | USD | 1 box | YES | BOX | COMPLY pending confirmation of latex free
42142502-00028 | 24 | Needle, spinal, quincke bevel; luer lock. The needle must be colour coded at the stylet cap and the size must be clearly stated. Must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. Must be bevelled at 220degrees Needle size: 27G with a needle length of 90mm excluding hub Needle and stylet must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 25) | N/A | PROMED TECHNOLOGIES (PTY) LTD | 3.34 | U-LIFE | 502028 | JIAXING | CHINA | 14 DAYS | 85% | USD | 1 BOX | YES | BOX | COMPLY pending confirmation of latex free
42142502-00028 | 24 | Needle, spinal, quincke bevel; luer lock. The needle must be colour coded at the stylet cap and the size must be clearly stated. Must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. Must be bevelled at 220degrees Needle size: 27G with a needle length of 90mm excluding hub Needle and stylet must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 25) |  | Chemical World cc | 4.03 | MEDIJECT | NDSSFG27 | RUNQIANG | CHINA | 3 | 85 | USD | 1000 | YES | 50 | No brand name. No product code.
42142502-00028 | 24 | Needle, spinal, quincke bevel; luer lock. The needle must be colour coded at the stylet cap and the size must be clearly stated. Must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. Must be bevelled at 220degrees Needle size: 27G with a needle length of 90mm excluding hub Needle and stylet must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 25) |  | UNITRADE 1032 CC | 4.37 | SPINOJEC | 1 | SHANGHAI KINDLY ENTERPRISE DEVELOPMENT GROUP | ROC | 21 | 85 | USD | 400 | YES | 25 | No sample submitted for evaluation
42142502-00028 | 24 | Needle, spinal, quincke bevel; luer lock. The needle must be colour coded at the stylet cap and the size must be clearly stated. Must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. Must be bevelled at 220degrees Needle size: 27G with a needle length of 90mm excluding hub Needle and stylet must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 25) | 1 | GROBIR MEDICAL SUPPLIERS | 4.73 | DURACAN | DQ.10.027.090 | PROMECON | CHINA | 21 | 80 | US DOLLAR | 25 | Yes | UOM AS SPECIFIED / MEETS SPECIFICATION | No sample submitted for evaluation.
42142502-00028 | 24 | Needle, spinal, quincke bevel; luer lock. The needle must be colour coded at the stylet cap and the size must be clearly stated. Must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. Must be bevelled at 220degrees Needle size: 27G with a needle length of 90mm excluding hub Needle and stylet must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 25) | N/A | SAFARMEX (PTY) LTD | 4.98 | Saclin | To be confirmed | Zhejiang Mediunion Healthcare Group Co. Ltd | China | 3days | 80% | US$ | 10000 Pieces | Yes | Box of 25 | No sample submitted for evaluation
42142502-00028 | 24 | Needle, spinal, quincke bevel; luer lock. The needle must be colour coded at the stylet cap and the size must be clearly stated. Must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. Must be bevelled at 220degrees Needle size: 27G with a needle length of 90mm excluding hub Needle and stylet must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 25) |  | Zanotone (Pty) Ltd | 6.39 | Mekon |  | Mekon | [6 empty] | EACH | No sample submitted for evaluation
42142502-00028 | 24 | Needle, spinal, quincke bevel; luer lock. The needle must be colour coded at the stylet cap and the size must be clearly stated. Must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. Must be bevelled at 220degrees Needle size: 27G with a needle length of 90mm excluding hub Needle and stylet must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 25) | OFFER 1 | SUPRA HEALTHCARE JOHANNESBURG PTY LTD | 7.44 | REVITAL | SPNQ27X90 | Revital Health | KENYA | Initial 60 days, thereafter 14 | 84 | USD | 1 Box | YES |  | COMPLY pending confirmation of latex free
42142502-00028 | 24 | Needle, spinal, quincke bevel; luer lock. The needle must be colour coded at the stylet cap and the size must be clearly stated. Must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. Must be bevelled at 220degrees Needle size: 27G with a needle length of 90mm excluding hub Needle and stylet must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 25) | Basic Line offer | Becton Dickinson | 9 | BD Quinke Spinal Needle 27G X 90mm | 405259 | Becton Dickinson | Fraga, Spain | 7 Working Days | 76% | USD | 1 BOX | YES | 25 Unita in a box | No sample submitted for evaluation.
42142502-00028 | 24 | Needle, spinal, quincke bevel; luer lock. The needle must be colour coded at the stylet cap and the size must be clearly stated. Must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. Must be bevelled at 220degrees Needle size: 27G with a needle length of 90mm excluding hub Needle and stylet must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 25) |  | ALMIKA TRADING PTY LTD | 10.58 | BD | 405259 | BD | USA | 14 | 0 | NO | 100 | YES | 1 | No sample submitted for evaluation.
42142502-00028 | 24 | Needle, spinal, quincke bevel; luer lock. The needle must be colour coded at the stylet cap and the size must be clearly stated. Must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. Must be bevelled at 220degrees Needle size: 27G with a needle length of 90mm excluding hub Needle and stylet must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 25) | BOX OF 50 | MSIMED CC | 13.5 | KANG KANG | SPN007 | ZHENJIANG | CHINA | 25 | 20% | USD | 20000PCS | YES | BOX | No sample submitted for evaluation
42142502-00028 | 24 | Needle, spinal, quincke bevel; luer lock. The needle must be colour coded at the stylet cap and the size must be clearly stated. Must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. Must be bevelled at 220degrees Needle size: 27G with a needle length of 90mm excluding hub Needle and stylet must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 25) | OFFER 2 | SUPRA HEALTHCARE JOHANNESBURG PTY LTD | 14.2 | Egemen | SPNQ27X90 | Egemen | TURKEY | Initial 60 days, thereafter 14 | 84 | USD | 1 Box | YES |  | COMPLY pending confirmation of latex free
42142502-00028 | 24 | Needle, spinal, quincke bevel; luer lock. The needle must be colour coded at the stylet cap and the size must be clearly stated. Must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. Must be bevelled at 220degrees Needle size: 27G with a needle length of 90mm excluding hub Needle and stylet must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 25) | 1 | B. BRAUN MEDICAL (PTY) LTD | 18 | Spinocan 27G x 3 1/2" (0.42 x 88mm) | 4503902 | B. Braun | Germany | 42 | 50% | Euro | 25 | Yes | Box of 25 | COMPLY pending cpnfirmation of latex free
42142502-00028 | 24 | Needle, spinal, quincke bevel; luer lock. The needle must be colour coded at the stylet cap and the size must be clearly stated. Must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. Must be bevelled at 220degrees Needle size: 27G with a needle length of 90mm excluding hub Needle and stylet must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 25) |  | LECHOBA MEDICAL TECHNOLOGIES (PTY) LTD | 27.97 | UNISIS UNIEVER | 5250561 | UNISIS CORPORATION | JAPAN | 21 DAYS | 75% | USD | 25 | YES | 25/BOX | No sample submitted for evaluation
42142502-00028 | 24 | Needle, spinal, quincke bevel; luer lock. The needle must be colour coded at the stylet cap and the size must be clearly stated. Must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. Must be bevelled at 220degrees Needle size: 27G with a needle length of 90mm excluding hub Needle and stylet must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 25) |  | Viking Critical Care (Pty) Ltd | 93.52 | VYGON | 69F518004 | VYGON | FRANCE | EX. STOCK 30-60DAYS | 80% | EURO | 1 | YES | 1 | COMPLY pending confirmation of latex free
42142502-00028 | 24 | Needle, spinal, quincke bevel; luer lock. The needle must be colour coded at the stylet cap and the size must be clearly stated. Must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. Must be bevelled at 220degrees Needle size: 27G with a needle length of 90mm excluding hub Needle and stylet must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 25) |  | JUMLA MEDICAL SUPPLIES (PTY) LTD | 340.12 | EGEMEN | TSPQ2790 | TIBBI MEDICAL | TURKEY | 21 DAYS | 70% | US DOLLAR | 2 BOXES | YES | NIL | COMPLY pending confirmation of latex free
42142502-00028 | 24 | Needle, spinal, quincke bevel; luer lock. The needle must be colour coded at the stylet cap and the size must be clearly stated. Must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. Must be bevelled at 220degrees Needle size: 27G with a needle length of 90mm excluding hub Needle and stylet must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 25) |  | BAS MEDEXPRESS | 776.25 | Mercury | SN2590Q |  | India | 1 - 3 days | 15% | N/A | As per client | Yes | Box | No sample submitted for evaluation
42142502-00029 | 25 | Needle, spinal, Pencil point bevel; luer lock with introducer. The needle must be colour coded at the stylet cap and the size must be clearly stated, must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. The needle must fit securely into the introducer. The introducer must be streamlined and easy to handle. The introducer size must not reduce the effective length of the cannula. The introducer hub must be transparent to asses for accidental CSF backflow. Needle size: 20G with a needle length of 90mm excluding hub Needle, stylet and introducers must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 50) |  | Medi-Core Technologies | 5.3 | ULIFE | 42142502-00029 | JIAXING U-LIFE MEDICAL DEVICE TECHNOLOGLY.,CO LTD | CHINA | 21 | 80 | USD | 1 box | YES | BOX | Re-tender endusers to revise specification |  | Re-tender endusers to revise specifications
42142502-00029 | 25 | Needle, spinal, Pencil point bevel; luer lock with introducer. The needle must be colour coded at the stylet cap and the size must be clearly stated, must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. The needle must fit securely into the introducer. The introducer must be streamlined and easy to handle. The introducer size must not reduce the effective length of the cannula. The introducer hub must be transparent to asses for accidental CSF backflow. Needle size: 20G with a needle length of 90mm excluding hub Needle, stylet and introducers must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 50) | N/A | PROMED TECHNOLOGIES (PTY) LTD | 5.34 | U-LIFE | 502029 | JIAXING | CHINA | 14 DAYS | 85% | USD | 1 BOX | YES | BOX | Re-tender end-users to revise specification |  | Re-tender endusers to revise specifications
42142502-00029 | 25 | Needle, spinal, Pencil point bevel; luer lock with introducer. The needle must be colour coded at the stylet cap and the size must be clearly stated, must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. The needle must fit securely into the introducer. The introducer must be streamlined and easy to handle. The introducer size must not reduce the effective length of the cannula. The introducer hub must be transparent to asses for accidental CSF backflow. Needle size: 20G with a needle length of 90mm excluding hub Needle, stylet and introducers must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 50) | BASIC LINE OFFER | A R FREIGHT PRIVATE LIMITED | 6.12 | SUMBOW | N/A | SUMBOW MEDICAL | CHINA | 60 DAYS | 85% | USD | 100 | YES | N/A | Re-tender end-users to revise specifications |  | Re-tender endusers to revise specifications
42142502-00029 | 25 | Needle, spinal, Pencil point bevel; luer lock with introducer. The needle must be colour coded at the stylet cap and the size must be clearly stated, must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. The needle must fit securely into the introducer. The introducer must be streamlined and easy to handle. The introducer size must not reduce the effective length of the cannula. The introducer hub must be transparent to asses for accidental CSF backflow. Needle size: 20G with a needle length of 90mm excluding hub Needle, stylet and introducers must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 50) |  | Chemical World cc | 7.85 | MEDIJECT | NDSPPI2090 | RUNQIANG | CHINA | 3 | 85 | USD | 1000 | YES | 50 | Re-tender end-users to revise specification |  | Re-tender endusers to revise specifications
42142502-00029 | 25 | Needle, spinal, Pencil point bevel; luer lock with introducer. The needle must be colour coded at the stylet cap and the size must be clearly stated, must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. The needle must fit securely into the introducer. The introducer must be streamlined and easy to handle. The introducer size must not reduce the effective length of the cannula. The introducer hub must be transparent to asses for accidental CSF backflow. Needle size: 20G with a needle length of 90mm excluding hub Needle, stylet and introducers must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 50) |  | UNITRADE 1032 CC | 7.94 | SPINOJEC PLUS | 1 | SHANGHAI KINDLY ENTERPRISE DEVELOPMENT GROUP | ROC | 21 | 85 | USD | 500 | YES | 50 | Re-tender end-users to revise specification |  | Re-tender endusers to revise specifications
42142502-00029 | 25 | Needle, spinal, Pencil point bevel; luer lock with introducer. The needle must be colour coded at the stylet cap and the size must be clearly stated, must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. The needle must fit securely into the introducer. The introducer must be streamlined and easy to handle. The introducer size must not reduce the effective length of the cannula. The introducer hub must be transparent to asses for accidental CSF backflow. Needle size: 20G with a needle length of 90mm excluding hub Needle, stylet and introducers must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 50) | N/A | SAFARMEX (PTY) LTD | 9.93 | Saclin | To be confirmed | Zhejiang Mediunion Healthcare Group Co. Ltd | China | 3days | 80% | US$ | 10000 Pieces | Yes | Box of 50 | Re-tender end-users to revise specification |  | Re-tender endusers to revise specifications
42142502-00029 | 25 | Needle, spinal, Pencil point bevel; luer lock with introducer. The needle must be colour coded at the stylet cap and the size must be clearly stated, must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. The needle must fit securely into the introducer. The introducer must be streamlined and easy to handle. The introducer size must not reduce the effective length of the cannula. The introducer hub must be transparent to asses for accidental CSF backflow. Needle size: 20G with a needle length of 90mm excluding hub Needle, stylet and introducers must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 50) | Basic Line Offer | BIOCOR HOSPITAL SUPPLIES PTY LTD | 12.6 | Medline | N/A | Medline | China | 60 | 85 | USD | 10 Boxes | Yes | N/A | Re-tender end-users to revise specification |  | Re-tender endusers to revise specifications
42142502-00029 | 25 | Needle, spinal, Pencil point bevel; luer lock with introducer. The needle must be colour coded at the stylet cap and the size must be clearly stated, must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. The needle must fit securely into the introducer. The introducer must be streamlined and easy to handle. The introducer size must not reduce the effective length of the cannula. The introducer hub must be transparent to asses for accidental CSF backflow. Needle size: 20G with a needle length of 90mm excluding hub Needle, stylet and introducers must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 50) | BOX OF 50 | MSIMED CC | 13.8 | KANG KANG | SPN008 | ZHENJIANG | CHINA | 25 | 20% | USD | 20000PCS | YES | BOX | Re-tender endusers to revise specification |  | Re-tender endusers to revise specifications
42142502-00029 | 25 | Needle, spinal, Pencil point bevel; luer lock with introducer. The needle must be colour coded at the stylet cap and the size must be clearly stated, must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. The needle must fit securely into the introducer. The introducer must be streamlined and easy to handle. The introducer size must not reduce the effective length of the cannula. The introducer hub must be transparent to asses for accidental CSF backflow. Needle size: 20G with a needle length of 90mm excluding hub Needle, stylet and introducers must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 50) | 1 | GROBIR MEDICAL SUPPLIERS | 13.95 | DURAGAURD | DG.11.020.090 | PROMECON | CHINA | 21 | 80 | US DOLLAR | 25 | Yes | UOM AS SPECIFIED / MEETS SPECIFICATION | Re-tender end-users to revise specification |  | Re-tender endusers to revise specifications
42142502-00029 | 25 | Needle, spinal, Pencil point bevel; luer lock with introducer. The needle must be colour coded at the stylet cap and the size must be clearly stated, must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. The needle must fit securely into the introducer. The introducer must be streamlined and easy to handle. The introducer size must not reduce the effective length of the cannula. The introducer hub must be transparent to asses for accidental CSF backflow. Needle size: 20G with a needle length of 90mm excluding hub Needle, stylet and introducers must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 50) | Alternative Offer | BIOCOR HOSPITAL SUPPLIES PTY LTD | 14.74 | Technomed | TCN 2090in | Techno Med | China | 21 | 0 | N/A | 5 Boxes of 50 | Yes | N/A | Re-tender end-users to revise specification |  | Re-tender endusers to revise specifications
42142502-00029 | 25 | Needle, spinal, Pencil point bevel; luer lock with introducer. The needle must be colour coded at the stylet cap and the size must be clearly stated, must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. The needle must fit securely into the introducer. The introducer must be streamlined and easy to handle. The introducer size must not reduce the effective length of the cannula. The introducer hub must be transparent to asses for accidental CSF backflow. Needle size: 20G with a needle length of 90mm excluding hub Needle, stylet and introducers must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 50) |  | Zanotone (Pty) Ltd | 16.66 | Mekon |  | Mekon | [6 empty] | EACH | Re-tender end-users to revise specification |  | Re-tender end-users to revise specification
42142502-00029 | 25 | Needle, spinal, Pencil point bevel; luer lock with introducer. The needle must be colour coded at the stylet cap and the size must be clearly stated, must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. The needle must fit securely into the introducer. The introducer must be streamlined and easy to handle. The introducer size must not reduce the effective length of the cannula. The introducer hub must be transparent to asses for accidental CSF backflow. Needle size: 20G with a needle length of 90mm excluding hub Needle, stylet and introducers must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 50) |  | SUPRA HEALTHCARE JOHANNESBURG PTY LTD | 26.37 | REVITAL | SPNPP20X90 | Revital Health | KENYA | Initial 60 days, thereafter 14 | 84 | USD | 1 Box | YES |  | Re-tender end-users to revise specification |  | Re-tender endusers to revise specifications
42142502-00029 | 25 | Needle, spinal, Pencil point bevel; luer lock with introducer. The needle must be colour coded at the stylet cap and the size must be clearly stated, must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. The needle must fit securely into the introducer. The introducer must be streamlined and easy to handle. The introducer size must not reduce the effective length of the cannula. The introducer hub must be transparent to asses for accidental CSF backflow. Needle size: 20G with a needle length of 90mm excluding hub Needle, stylet and introducers must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 50) |  | BAS MEDEXPRESS | 776.25 | Mercury | SN1690P |  | India | 1 - 3 days | 15% | N/A | As per client | Yes | Box | Re-tender endusers to revise specifications |  | Re-tender endusers to revise specifications
42142502-00057 | 26 | Needle, spinal, Pencil point bevel; luer lock with introducer.The needle must be colour coded at the stylet cap and the size must be clearly stated, must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. The needle must fit securely into the introducer. The introducer must be streamlined and easy to handle. The introducer size must not reduce the effective length of the cannula. The introducer hub must be transparent to asses for accidental CSF backflow. Needle size: 22G with a needle length of 90mm excluding hub Needle, stylet and introducers must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 50) |  | Medi-Core Technologies | 5.3 | ULIFE | 42142502-00057 | JIAXING U-LIFE MEDICAL DEVICE TECHNOLOGLY.,CO LTD | CHINA | 21 | 80 | USD | 1 box | YES | BOX | Re-tender endusers to revise specification |  | Re-tender endusers to revise specifications
42142502-00057 | 26 | Needle, spinal, Pencil point bevel; luer lock with introducer.The needle must be colour coded at the stylet cap and the size must be clearly stated, must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. The needle must fit securely into the introducer. The introducer must be streamlined and easy to handle. The introducer size must not reduce the effective length of the cannula. The introducer hub must be transparent to asses for accidental CSF backflow. Needle size: 22G with a needle length of 90mm excluding hub Needle, stylet and introducers must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 50) | N/A | PROMED TECHNOLOGIES (PTY) LTD | 5.34 | U-LIFE | 502057 | JIAXING | CHINA | 14 DAYS | 85% | USD | 1 BOX | YES | BOX | Re-tender end-users to revise specification |  | Re-tender endusers to revise specifications
42142502-00057 | 26 | Needle, spinal, Pencil point bevel; luer lock with introducer.The needle must be colour coded at the stylet cap and the size must be clearly stated, must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. The needle must fit securely into the introducer. The introducer must be streamlined and easy to handle. The introducer size must not reduce the effective length of the cannula. The introducer hub must be transparent to asses for accidental CSF backflow. Needle size: 22G with a needle length of 90mm excluding hub Needle, stylet and introducers must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 50) | BASIC LINE OFFER | A R FREIGHT PRIVATE LIMITED | 6.12 | SUMBOW | N/A | SUMBOW MEDICAL | CHINA | 60 DAYS | 85% | USD | 100 | YES | N/A | Re-tender end-users to revise specifications |  | Re-tender endusers to revise specifications
42142502-00057 | 26 | Needle, spinal, Pencil point bevel; luer lock with introducer.The needle must be colour coded at the stylet cap and the size must be clearly stated, must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. The needle must fit securely into the introducer. The introducer must be streamlined and easy to handle. The introducer size must not reduce the effective length of the cannula. The introducer hub must be transparent to asses for accidental CSF backflow. Needle size: 22G with a needle length of 90mm excluding hub Needle, stylet and introducers must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 50) | Basic Line offer | Becton Dickinson | 8.5 | BD Whitacre Spinal Needle 22G x 90mm | 401995 | Becton Dickinson | San Augustine | 1 week | 76% | USD | 1 BOX | YES | 25 units | Re-tender end-users to revise specifications |  | Re-tender endusers to revise specifications
42142502-00057 | 26 | Needle, spinal, Pencil point bevel; luer lock with introducer.The needle must be colour coded at the stylet cap and the size must be clearly stated, must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. The needle must fit securely into the introducer. The introducer must be streamlined and easy to handle. The introducer size must not reduce the effective length of the cannula. The introducer hub must be transparent to asses for accidental CSF backflow. Needle size: 22G with a needle length of 90mm excluding hub Needle, stylet and introducers must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 50) |  | Chemical World cc | 9.09 | MEDIJECT | NDSPPI22 | RUNQIANG | CHINA | 3 | 85 | USD | 1000 | YES | 50 | Re-tender end-users to revise specification |  | Re-tender endusers to revise specifications
42142502-00057 | 26 | Needle, spinal, Pencil point bevel; luer lock with introducer.The needle must be colour coded at the stylet cap and the size must be clearly stated, must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. The needle must fit securely into the introducer. The introducer must be streamlined and easy to handle. The introducer size must not reduce the effective length of the cannula. The introducer hub must be transparent to asses for accidental CSF backflow. Needle size: 22G with a needle length of 90mm excluding hub Needle, stylet and introducers must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 50) | N/A | SAFARMEX (PTY) LTD | 9.93 | Saclin | To be confirmed | Zhejiang Mediunion Healthcare Group Co. Ltd | China | 3days | 80% | US$ | 10000 Pieces | Yes | Box of 50 | Re-tender end-users to revise specification |  | Re-tender endusers to revise specifications
42142502-00057 | 26 | Needle, spinal, Pencil point bevel; luer lock with introducer.The needle must be colour coded at the stylet cap and the size must be clearly stated, must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. The needle must fit securely into the introducer. The introducer must be streamlined and easy to handle. The introducer size must not reduce the effective length of the cannula. The introducer hub must be transparent to asses for accidental CSF backflow. Needle size: 22G with a needle length of 90mm excluding hub Needle, stylet and introducers must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 50) | Basic Line Offer | BIOCOR HOSPITAL SUPPLIES PTY LTD | 12.6 | Medline | N/A | Yangzhong Medline Industry Co.Ltd | China | 60 | 85 | USD | 10 Boxes | Yes | N/A | Re-tender end-users to revise specification |  | Re-tender endusers to revise specifications
42142502-00057 | 26 | Needle, spinal, Pencil point bevel; luer lock with introducer.The needle must be colour coded at the stylet cap and the size must be clearly stated, must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. The needle must fit securely into the introducer. The introducer must be streamlined and easy to handle. The introducer size must not reduce the effective length of the cannula. The introducer hub must be transparent to asses for accidental CSF backflow. Needle size: 22G with a needle length of 90mm excluding hub Needle, stylet and introducers must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 50) | 1 | GROBIR MEDICAL SUPPLIERS | 13.95 | DURAGAURD | DG.11.022.090 | PROMECON | CHINA | 21 | 80 | US DOLLAR | 25 | Yes | UOM SOLD IN PACKS OF 25 CAN SUPPLY 2 BOXES TO MEET SPECIFICATION REQUIREMENT OF 50 UNITS / MEETS SPECIFICATION | Re-tender end-users to revise specification |  | Re-tender endusers to revise specifications
42142502-00057 | 26 | Needle, spinal, Pencil point bevel; luer lock with introducer.The needle must be colour coded at the stylet cap and the size must be clearly stated, must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. The needle must fit securely into the introducer. The introducer must be streamlined and easy to handle. The introducer size must not reduce the effective length of the cannula. The introducer hub must be transparent to asses for accidental CSF backflow. Needle size: 22G with a needle length of 90mm excluding hub Needle, stylet and introducers must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 50) | Alternative Offer | BIOCOR HOSPITAL SUPPLIES PTY LTD | 14.73 | Technomed | TCN2290in | Techno Med | South Africa | 21 | 0 | N/A | 5 Boxes of 50 | Yes | N/A | Re-tender end-users to revise specification |  | Re-tender endusers to revise specifications
42142502-00057 | 26 | Needle, spinal, Pencil point bevel; luer lock with introducer.The needle must be colour coded at the stylet cap and the size must be clearly stated, must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. The needle must fit securely into the introducer. The introducer must be streamlined and easy to handle. The introducer size must not reduce the effective length of the cannula. The introducer hub must be transparent to asses for accidental CSF backflow. Needle size: 22G with a needle length of 90mm excluding hub Needle, stylet and introducers must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 50) |  | Zanotone (Pty) Ltd | 16.66 | Mekon |  | Mekon | [6 empty] | EACH | Re-tender end-users to revise specification |  | Re-tender end-users to revise specification
42142502-00057 | 26 | Needle, spinal, Pencil point bevel; luer lock with introducer.The needle must be colour coded at the stylet cap and the size must be clearly stated, must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. The needle must fit securely into the introducer. The introducer must be streamlined and easy to handle. The introducer size must not reduce the effective length of the cannula. The introducer hub must be transparent to asses for accidental CSF backflow. Needle size: 22G with a needle length of 90mm excluding hub Needle, stylet and introducers must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 50) |  | ALMIKA TRADING PTY LTD | 19.84 | BD | 401995 | BD | USA | 14 | 0 | NO | 100 | YES | 1 | Re-tender end-users to revise specifications |  | Re-tender endusers to revise specifications
42142502-00057 | 26 | Needle, spinal, Pencil point bevel; luer lock with introducer.The needle must be colour coded at the stylet cap and the size must be clearly stated, must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. The needle must fit securely into the introducer. The introducer must be streamlined and easy to handle. The introducer size must not reduce the effective length of the cannula. The introducer hub must be transparent to asses for accidental CSF backflow. Needle size: 22G with a needle length of 90mm excluding hub Needle, stylet and introducers must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 50) | OFFER 1 | SUPRA HEALTHCARE JOHANNESBURG PTY LTD | 26.37 | REVITAL | SPNPP22X90 | Revital Health | KENYA | Initial 60 days, thereafter 14 | 84 | USD | 1 Box | YES |  | Re-tender end-users to revise specification |  | Re-tender endusers to revise specifications
42142502-00057 | 26 | Needle, spinal, Pencil point bevel; luer lock with introducer.The needle must be colour coded at the stylet cap and the size must be clearly stated, must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. The needle must fit securely into the introducer. The introducer must be streamlined and easy to handle. The introducer size must not reduce the effective length of the cannula. The introducer hub must be transparent to asses for accidental CSF backflow. Needle size: 22G with a needle length of 90mm excluding hub Needle, stylet and introducers must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 50) | OFFER 2 | SUPRA HEALTHCARE JOHANNESBURG PTY LTD | 31.44 | Egemen | SPNPP22X90 | Egemen | TURKEY | Initial 60 days, thereafter 14 | 84 | USD | 1 Box | YES |  | Re-tender end-users to revise specification |  | Re-tender endusers to revise specifications
42142502-00057 | 26 | Needle, spinal, Pencil point bevel; luer lock with introducer.The needle must be colour coded at the stylet cap and the size must be clearly stated, must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. The needle must fit securely into the introducer. The introducer must be streamlined and easy to handle. The introducer size must not reduce the effective length of the cannula. The introducer hub must be transparent to asses for accidental CSF backflow. Needle size: 22G with a needle length of 90mm excluding hub Needle, stylet and introducers must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 50) | 1 | B. BRAUN MEDICAL (PTY) LTD | 32 | Pencan 22G x 3 1/2" (0.783 x 88/90mm) | 4502035 | B. Braun | Germany | 42 | 50% | Euro | 50 | Yes | Box of 50 | Re-tender: End-users to revise specification |  | Re-tender endusers to revise specifications
42142502-00057 | 26 | Needle, spinal, Pencil point bevel; luer lock with introducer.The needle must be colour coded at the stylet cap and the size must be clearly stated, must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. The needle must fit securely into the introducer. The introducer must be streamlined and easy to handle. The introducer size must not reduce the effective length of the cannula. The introducer hub must be transparent to asses for accidental CSF backflow. Needle size: 22G with a needle length of 90mm excluding hub Needle, stylet and introducers must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 50) |  | LECHOBA MEDICAL TECHNOLOGIES (PTY) LTD | 42.86 | UNISIS UNIEVER | 2120471 | UNISIS CORPORATION | JAPAN | 21 DAYS | 75% | USD | 50 | YES | 25/BOX | Re-tender end-users to revise specification |  | Re-tender endusers to revise specifications
42142502-00057 | 26 | Needle, spinal, Pencil point bevel; luer lock with introducer.The needle must be colour coded at the stylet cap and the size must be clearly stated, must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. The needle must fit securely into the introducer. The introducer must be streamlined and easy to handle. The introducer size must not reduce the effective length of the cannula. The introducer hub must be transparent to asses for accidental CSF backflow. Needle size: 22G with a needle length of 90mm excluding hub Needle, stylet and introducers must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 50) BOX (BOX OF 25) |  | JUMLA MEDICAL SUPPLIES (PTY) LTD | 648.03 | EGEMEN | TSPP2290 | TIBBI MEDICAL | TURKEY | 21 DAYS | 70% | US DOLLAR | 2 BOXES | YES | NIL | Re-tender end-users to revise specification |  | Re-tender endusers to revise specifications
42142502-00030 | 27 | Needle, spinal, Pencil point bevel; luer lock with introducer. The needle must be colour coded at the stylet cap and the size must be clearly stated, must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. The needle must fit securely into the introducer. The introducer must be streamlined and easy to handle. The introducer size must not reduce the effective length of the cannula. The introducer hub must be transparent to asses for accidental CSF backflow. Needle size: 22G extra-length with a needle length of >110 to 120mm excluding hub. Needle, stylet and introducers must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 50) | BASIC LINE OFFER | A R FREIGHT PRIVATE LIMITED | 3.06 | SUMBOW | N/A | SUMBOW MEDICAL | CHINA | 60DAYS | 85% | USD | 100 | YES | N/A | Re-tender end-users to revise specifications |  | Re-tender endusers to revise specifications
42142502-00030 | 27 | Needle, spinal, Pencil point bevel; luer lock with introducer. The needle must be colour coded at the stylet cap and the size must be clearly stated, must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. The needle must fit securely into the introducer. The introducer must be streamlined and easy to handle. The introducer size must not reduce the effective length of the cannula. The introducer hub must be transparent to asses for accidental CSF backflow. Needle size: 22G extra-length with a needle length of >110 to 120mm excluding hub. Needle, stylet and introducers must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 50) |  | Medi-Core Technologies | 5.3 | ULIFE | 42142502-00030 | JIAXING U-LIFE MEDICAL DEVICE TECHNOLOGLY.,CO LTD | CHINA | 21 | 80 | USD | 1 box | YES | BOX | Re-tender endusers to revise specification |  | Re-tender endusers to revise specifications
42142502-00030 | 27 | Needle, spinal, Pencil point bevel; luer lock with introducer. The needle must be colour coded at the stylet cap and the size must be clearly stated, must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. The needle must fit securely into the introducer. The introducer must be streamlined and easy to handle. The introducer size must not reduce the effective length of the cannula. The introducer hub must be transparent to asses for accidental CSF backflow. Needle size: 22G extra-length with a needle length of >110 to 120mm excluding hub. Needle, stylet and introducers must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 50) | N/A | PROMED TECHNOLOGIES (PTY) LTD | 5.34 | U-LIFE | 502030 | JIAXING | CHINA | 14 DAYS | 85% | USD | 1 BOX | YES | BOX | Re-tender end-users to revise specification |  | Re-tender endusers to revise specifications
42142502-00030 | 27 | Needle, spinal, Pencil point bevel; luer lock with introducer. The needle must be colour coded at the stylet cap and the size must be clearly stated, must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. The needle must fit securely into the introducer. The introducer must be streamlined and easy to handle. The introducer size must not reduce the effective length of the cannula. The introducer hub must be transparent to asses for accidental CSF backflow. Needle size: 22G extra-length with a needle length of >110 to 120mm excluding hub. Needle, stylet and introducers must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 50) |  | UNITRADE 1032 CC | 7.94 | SPINOJEC PLUS | 1 | SHANGHAI KINDLY ENTERPRISE DEVELOPMENT GROUP | ROC | 21 | 85 | USD | 500 | YES | 50 | Re-tender end-users to revise specification |  | Re-tender endusers to revise specifications
42142502-00030 | 27 | Needle, spinal, Pencil point bevel; luer lock with introducer. The needle must be colour coded at the stylet cap and the size must be clearly stated, must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. The needle must fit securely into the introducer. The introducer must be streamlined and easy to handle. The introducer size must not reduce the effective length of the cannula. The introducer hub must be transparent to asses for accidental CSF backflow. Needle size: 22G extra-length with a needle length of >110 to 120mm excluding hub. Needle, stylet and introducers must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 50) | N/A | SAFARMEX (PTY) LTD | 10.76 | Saclin | To be confirmed | Zhejiang Mediunion Healthcare Group Co. Ltd | China | 3days | 80% | US$ | 10000 Pieces | Yes | Box of 50 | Re-tender end-users to revise specification |  | Re-tender endusers to revise specifications
42142502-00030 | 27 | Needle, spinal, Pencil point bevel; luer lock with introducer. The needle must be colour coded at the stylet cap and the size must be clearly stated, must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. The needle must fit securely into the introducer. The introducer must be streamlined and easy to handle. The introducer size must not reduce the effective length of the cannula. The introducer hub must be transparent to asses for accidental CSF backflow. Needle size: 22G extra-length with a needle length of >110 to 120mm excluding hub. Needle, stylet and introducers must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 50) | Basic Line Offer | BIOCOR HOSPITAL SUPPLIES PTY LTD | 13.82 | Medline | N/A | Yangzhong Medline Industry Co.Ltd | China | 60 | 85 | USD | 10 Boxes | Yes | N/A | Re-tender end-users to revise specification |  | Re-tender endusers to revise specifications
42142502-00030 | 27 | Needle, spinal, Pencil point bevel; luer lock with introducer. The needle must be colour coded at the stylet cap and the size must be clearly stated, must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. The needle must fit securely into the introducer. The introducer must be streamlined and easy to handle. The introducer size must not reduce the effective length of the cannula. The introducer hub must be transparent to asses for accidental CSF backflow. Needle size: 22G extra-length with a needle length of >110 to 120mm excluding hub. Needle, stylet and introducers must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 50) | 1 | GROBIR MEDICAL SUPPLIERS | 15.84 | DURAGAURD | DG.11.022.110 | PROMECON | CHINA | 21 | 80 | US DOLLAR | 25 | Yes | UOM SOLD IN PACKS OF 25 CAN SUPPLY 2 BOXES TO MEET SPECIFICATION REQUIREMENT OF 50 UNITS / MEETS SPECIFICATION | Re-tender end-users to revise specification |  | Re-tender endusers to revise specifications
42142502-00030 | 27 | Needle, spinal, Pencil point bevel; luer lock with introducer. The needle must be colour coded at the stylet cap and the size must be clearly stated, must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. The needle must fit securely into the introducer. The introducer must be streamlined and easy to handle. The introducer size must not reduce the effective length of the cannula. The introducer hub must be transparent to asses for accidental CSF backflow. Needle size: 22G extra-length with a needle length of >110 to 120mm excluding hub. Needle, stylet and introducers must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 50) | Alternative Offer | BIOCOR HOSPITAL SUPPLIES PTY LTD | 15.97 | Technomed | TCN 22120in | Techno Med | South Africa | 21 | 0 | N/A | 5 Boxes of 50 | Yes | N/A | Re-tender end-users to revise specification |  | Re-tender endusers to revise specifications
42142502-00030 | 27 |  |  | Zanotone (Pty) Ltd | 17.22 | Mekon |  | Mekon | [6 empty] | EACH | Re-tender end-users to revise specification |  | Re-tender end-users to revise specification
42142502-00030 | 27 | Needle, spinal, Pencil point bevel; luer lock with introducer. The needle must be colour coded at the stylet cap and the size must be clearly stated, must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. The needle must fit securely into the introducer. The introducer must be streamlined and easy to handle. The introducer size must not reduce the effective length of the cannula. The introducer hub must be transparent to asses for accidental CSF backflow. Needle size: 22G extra-length with a needle length of >110 to 120mm excluding hub. Needle, stylet and introducers must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 50) |  | SUPRA HEALTHCARE JOHANNESBURG PTY LTD | 26.37 | REVITAL | SPNPP22X120 | Revital Health | KENYA | INITIAL 60 DAYS, THEREAFTER 14 DAYS | 84 | USD | 1 Box | YES |  | Re-tender end-users to revise specification |  | Re-tender endusers to revise specifications
42142502-00030 | 27 | Needle, spinal, Pencil point bevel; luer lock with introducer. The needle must be colour coded at the stylet cap and the size must be clearly stated, must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. The needle must fit securely into the introducer. The introducer must be streamlined and easy to handle. The introducer size must not reduce the effective length of the cannula. The introducer hub must be transparent to asses for accidental CSF backflow. Needle size: 22G extra-length with a needle length of >110 to 120mm excluding hub. Needle, stylet and introducers must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 50) |  | BAS MEDEXPRESS | 776.25 | Mercury | SN1890P |  | India | 1 - 3 days | 15% | N/A | As per client | Yes | Box | Re-tender endusers to revise specifications |  | Re-tender endusers to revise specifications
42142502-00031 | 28 | Needle, spinal, Pencil point bevel; luer lock with introducer. The needle must be colour coded at the stylet cap and the size must be clearly stated, must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. The needle must fit securely into the introducer. The introducer must be streamlined and easy to handle. The introducer size must not reduce the effective length of the cannula. The introducer hub must be transparent to asses for accidental CSF backflow. Needle size: 23G with a needle length of 90mm excluding hub Needle, stylet and introducers must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 50) |  | Medi-Core Technologies | 5.3 | ULIFE | 42142501-00031 | JIAXING U-LIFE MEDICAL DEVICE TECHNOLOGLY.,CO LTD | CHINA | 21 | 80 | USD | 1 box | YES | BOX | Re-tender endusers to revise specification |  | Re-tender endusers to revise specifications
42142502-00031 | 28 | Needle, spinal, Pencil point bevel; luer lock with introducer. The needle must be colour coded at the stylet cap and the size must be clearly stated, must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. The needle must fit securely into the introducer. The introducer must be streamlined and easy to handle. The introducer size must not reduce the effective length of the cannula. The introducer hub must be transparent to asses for accidental CSF backflow. Needle size: 23G with a needle length of 90mm excluding hub Needle, stylet and introducers must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 50) | N/A | PROMED TECHNOLOGIES (PTY) LTD | 5.34 | U-LIFE | 502031 | JIAXING | CHINA | 14 DAYS | 85% | USD | 1 BOX | YES | BOX | Re-tender end-users to revise specification |  | Re-tender endusers to revise specifications
42142502-00031 | 28 | Needle, spinal, Pencil point bevel; luer lock with introducer. The needle must be colour coded at the stylet cap and the size must be clearly stated, must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. The needle must fit securely into the introducer. The introducer must be streamlined and easy to handle. The introducer size must not reduce the effective length of the cannula. The introducer hub must be transparent to asses for accidental CSF backflow. Needle size: 23G with a needle length of 90mm excluding hub Needle, stylet and introducers must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 50) | BASIC LINE OFFER | A R FREIGHT PRIVATE LIMITED | 6.12 | SUMBOW | N/A | SUMBOW MEDICAL | CHINA | 60 DAYS | 85% | USD | 100 | YES | N/A | Re-tender end-users to revise specifications |  | Re-tender endusers to revise specifications
42142502-00031 | 28 | Needle, spinal, Pencil point bevel; luer lock with introducer. The needle must be colour coded at the stylet cap and the size must be clearly stated, must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. The needle must fit securely into the introducer. The introducer must be streamlined and easy to handle. The introducer size must not reduce the effective length of the cannula. The introducer hub must be transparent to asses for accidental CSF backflow. Needle size: 23G with a needle length of 90mm excluding hub Needle, stylet and introducers must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 50) |  | UNITRADE 1032 CC | 7.94 | SPINOJEC PLUS | 1 | SHANGHAI KINDLY ENTERPRISE DEVELOPMENT GROUP | ROC | 21 | 85 | USD | 500 | YES | 50 | Re-tender end-users to revise specification |  | Re-tender endusers to revise specifications
42142502-00031 | 28 | Needle, spinal, Pencil point bevel; luer lock with introducer. The needle must be colour coded at the stylet cap and the size must be clearly stated, must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. The needle must fit securely into the introducer. The introducer must be streamlined and easy to handle. The introducer size must not reduce the effective length of the cannula. The introducer hub must be transparent to asses for accidental CSF backflow. Needle size: 23G with a needle length of 90mm excluding hub Needle, stylet and introducers must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 50) |  | Chemical World cc | 9.09 | MEDIJECT | NDSPPI23 | RUNQIANG | CHINA | 3 | 85 | USD | 1000 | YES | 50 | No sample submitted for evaluation |  | Re-tender endusers to revise specifications
42142502-00031 | 28 | Needle, spinal, Pencil point bevel; luer lock with introducer. The needle must be colour coded at the stylet cap and the size must be clearly stated, must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. The needle must fit securely into the introducer. The introducer must be streamlined and easy to handle. The introducer size must not reduce the effective length of the cannula. The introducer hub must be transparent to asses for accidental CSF backflow. Needle size: 23G with a needle length of 90mm excluding hub Needle, stylet and introducers must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 50) | N/A | SAFARMEX (PTY) LTD | 9.93 | Saclin | To be confirmed | Zhejiang Mediunion Healthcare Group Co. Ltd | China | 3days | 80% | US$ | 10000 Pieces | Yes | Box of 50 | Re-tender end-users to revise specification |  | Re-tender endusers to revise specifications
42142502-00031 | 28 | Needle, spinal, Pencil point bevel; luer lock with introducer. The needle must be colour coded at the stylet cap and the size must be clearly stated, must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. The needle must fit securely into the introducer. The introducer must be streamlined and easy to handle. The introducer size must not reduce the effective length of the cannula. The introducer hub must be transparent to asses for accidental CSF backflow. Needle size: 23G with a needle length of 90mm excluding hub Needle, stylet and introducers must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 50) | Basic Line Offer | BIOCOR HOSPITAL SUPPLIES PTY LTD | 12.6 | Medline | N/A | Yangzhong Medline Industry Co.Ltd | China | 60 | 85 | USD | 10 Boxes | Yes | N/A | Re-tender end-users to revise specification |  | Re-tender endusers to revise specifications
42142502-00031 | 28 | Needle, spinal, Pencil point bevel; luer lock with introducer. The needle must be colour coded at the stylet cap and the size must be clearly stated, must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. The needle must fit securely into the introducer. The introducer must be streamlined and easy to handle. The introducer size must not reduce the effective length of the cannula. The introducer hub must be transparent to asses for accidental CSF backflow. Needle size: 23G with a needle length of 90mm excluding hub Needle, stylet and introducers must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 50) | BOX OF 50 | MSIMED CC | 13.8 | KANG KANG | SPN009 | ZHENJIANG | CHINA | 25 | 20% | USD | 20000PCS | YES | BOX | Re-tender endusers to revise specification |  | Re-tender endusers to revise specifications
42142502-00031 | 28 | Needle, spinal, Pencil point bevel; luer lock with introducer. The needle must be colour coded at the stylet cap and the size must be clearly stated, must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. The needle must fit securely into the introducer. The introducer must be streamlined and easy to handle. The introducer size must not reduce the effective length of the cannula. The introducer hub must be transparent to asses for accidental CSF backflow. Needle size: 23G with a needle length of 90mm excluding hub Needle, stylet and introducers must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 50) | 1 | GROBIR MEDICAL SUPPLIERS | 13.95 | DURAGAURD | DG.11.023.090 | PROMECON | CHINA | 21 | 80 | US DOLLAR | 25 | Yes | UOM SOLD IN PACKS OF 25 CAN SUPPLY 2 BOXES TO MEET SPECIFICATION REQUIREMENT OF 50 UNITS / MEETS SPECIFICATION | Re-tender end-users to revise specification |  | Re-tender endusers to revise specifications
42142502-00031 | 28 | Needle, spinal, Pencil point bevel; luer lock with introducer. The needle must be colour coded at the stylet cap and the size must be clearly stated, must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. The needle must fit securely into the introducer. The introducer must be streamlined and easy to handle. The introducer size must not reduce the effective length of the cannula. The introducer hub must be transparent to asses for accidental CSF backflow. Needle size: 23G with a needle length of 90mm excluding hub Needle, stylet and introducers must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 50) | Alternative Offer | BIOCOR HOSPITAL SUPPLIES PTY LTD | 14.73 | Technomed | TCN2390in | Techno Med | South Africa | 21 | 0 | N/A | 5 Boxes of 50 | Yes | N/A | Re-tender end-users to revise specification |  | Re-tender endusers to revise specifications
42142502-00031 | 28 | Needle, spinal, Pencil point bevel; luer lock with introducer. The needle must be colour coded at the stylet cap and the size must be clearly stated, must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. The needle must fit securely into the introducer. The introducer must be streamlined and easy to handle. The introducer size must not reduce the effective length of the cannula. The introducer hub must be transparent to asses for accidental CSF backflow. Needle size: 23G with a needle length of 90mm excluding hub Needle, stylet and introducers must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 50) |  | Zanotone (Pty) Ltd | 16.66 | Mekon |  | Mekon | [6 empty] | EACH | Re-tender end-users to revise specification |  | Re-tender end-users to revise specification
42142502-00031 | 28 | Needle, spinal, Pencil point bevel; luer lock with introducer. The needle must be colour coded at the stylet cap and the size must be clearly stated, must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. The needle must fit securely into the introducer. The introducer must be streamlined and easy to handle. The introducer size must not reduce the effective length of the cannula. The introducer hub must be transparent to asses for accidental CSF backflow. Needle size: 23G with a needle length of 90mm excluding hub Needle, stylet and introducers must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 50) |  | SUPRA HEALTHCARE JOHANNESBURG PTY LTD | 26.37 | REVITAL | SPNPP23X90 | Revital Health | KENYA | INITIAL 60 DAYS, THEREAFTER 14 DAYS | 84 | USD | 1 Box | YES |  | Re-tender end-users to revise specification |  | Re-tender endusers to revise specifications
42142502-00031 | 28 | Needle, spinal, Pencil point bevel; luer lock with introducer. The needle must be colour coded at the stylet cap and the size must be clearly stated, must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. The needle must fit securely into the introducer. The introducer must be streamlined and easy to handle. The introducer size must not reduce the effective length of the cannula. The introducer hub must be transparent to asses for accidental CSF backflow. Needle size: 23G with a needle length of 90mm excluding hub Needle, stylet and introducers must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 50) |  | Viking Critical Care (Pty) Ltd | 166.47 | VYGON | 69F18195 | VYGON | FRANCE | EX. STOCK 30 - 60 DAYS | 80% | EURO | 1 | YES | 1 | Re-tender end-users to revise specification |  | Re-tender endusers to revise specifications
42142502-00031 | 28 | Needle, spinal, Pencil point bevel; luer lock with introducer. The needle must be colour coded at the stylet cap and the size must be clearly stated, must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. The needle must fit securely into the introducer. The introducer must be streamlined and easy to handle. The introducer size must not reduce the effective length of the cannula. The introducer hub must be transparent to asses for accidental CSF backflow. Needle size: 23G with a needle length of 90mm excluding hub Needle, stylet and introducers must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 50) BOX (BOX OF 25) |  | JUMLA MEDICAL SUPPLIES (PTY) LTD | 680.23 | EGEMEN | TSPP2390 | TIBBI MEDICAL | TURKEY | 21 DAYS | 70% | US DOLLAR | 2 BOXES | YES | NIL | Re-tender end-users to revise specification |  | Re-tender endusers to revise specifications
42142502-00031 | 28 | Needle, spinal, Pencil point bevel; luer lock with introducer. The needle must be colour coded at the stylet cap and the size must be clearly stated, must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. The needle must fit securely into the introducer. The introducer must be streamlined and easy to handle. The introducer size must not reduce the effective length of the cannula. The introducer hub must be transparent to asses for accidental CSF backflow. Needle size: 23G with a needle length of 90mm excluding hub Needle, stylet and introducers must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 50) |  | BAS MEDEXPRESS | 776.25 | Mercury | SN2290P |  | India | 1 - 3 days | 15% | N/A | As per client | Yes | Box | Re-tender endusers to revise specifications |  | Re-tender endusers to revise specifications
42142502-00032 | 29 | Needle, spinal, Pencil point bevel; luer lock with introducer. The needle must be colour coded at the stylet cap and the size must be clearly stated, must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. The needle must fit securely into the introducer. The introducer must be streamlined and easy to handle. The introducer size must not reduce the effective length of the cannula. The introducer hub must be transparent to asses for accidental CSF backflow. Needle size: 25G with a needle length of 90mm excluding hub. Introducer needle length of >36mm with an introducer hub length of <19mm. The effective length of introducer and stylet with cannula is >76mm Needle, stylet and introducers must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 50) |  | Medi-Core Technologies | 264.55 | ULIFE | 42142502-00032 | JIAXING U-LIFE MEDICAL DEVICE TECHNOLOGLY.,CO LTD | CHINA | 21 | 80 | USD | 1 box | YES | BOX OF 50 | COMPLY pending confirmation of introducer length |  | Recommended | 100
42142502-00032 | 29 | Needle, spinal, Pencil point bevel; luer lock with introducer. The needle must be colour coded at the stylet cap and the size must be clearly stated, must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. The needle must fit securely into the introducer. The introducer must be streamlined and easy to handle. The introducer size must not reduce the effective length of the cannula. The introducer hub must be transparent to asses for accidental CSF backflow. Needle size: 25G with a needle length of 90mm excluding hub. Introducer needle length of >36mm with an introducer hub length of <19mm. The effective length of introducer and stylet with cannula is >76mm Needle, stylet and introducers must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 50) | N/A | PROMED TECHNOLOGIES (PTY) LTD | 5.34 | U-LIFE | 502032 | JIAXING | CHINA | 14 DAYS | 85% | USD | 1 BOX | YES | BOX | COMPLY pending confirmation of latex free Accept with confirmation of the introducer length and introducer hub
42142502-00032 | 29 | Needle, spinal, Pencil point bevel; luer lock with introducer. The needle must be colour coded at the stylet cap and the size must be clearly stated, must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. The needle must fit securely into the introducer. The introducer must be streamlined and easy to handle. The introducer size must not reduce the effective length of the cannula. The introducer hub must be transparent to asses for accidental CSF backflow. Needle size: 25G with a needle length of 90mm excluding hub. Introducer needle length of >36mm with an introducer hub length of <19mm. The effective length of introducer and stylet with cannula is >76mm Needle, stylet and introducers must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 50) | BASIC LINE OFFER | A R FREIGHT PRIVATE LIMITED | 6.12 | SUMBOW | N/A | SUMBOW MEDICAL | CHINA | 60 DAYS | 85% | USD | 100 | YES | N/A | The needle does have an introducer needle
42142502-00032 | 29 | Needle, spinal, Pencil point bevel; luer lock with introducer. The needle must be colour coded at the stylet cap and the size must be clearly stated, must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. The needle must fit securely into the introducer. The introducer must be streamlined and easy to handle. The introducer size must not reduce the effective length of the cannula. The introducer hub must be transparent to asses for accidental CSF backflow. Needle size: 25G with a needle length of 90mm excluding hub. Introducer needle length of >36mm with an introducer hub length of <19mm. The effective length of introducer and stylet with cannula is >76mm Needle, stylet and introducers must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 50) |  | UNITRADE 1032 CC | 7.94 | SPINOJEC PLUS | 1 | SHANGHAI KINDLY ENTERPRISE DEVELOPMENT GROUP | ROC | 21 | 85 | USD | 500 | YES | 50 | No sample submitted for evaluation
42142502-00032 | 29 | Needle, spinal, Pencil point bevel; luer lock with introducer. The needle must be colour coded at the stylet cap and the size must be clearly stated, must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. The needle must fit securely into the introducer. The introducer must be streamlined and easy to handle. The introducer size must not reduce the effective length of the cannula. The introducer hub must be transparent to asses for accidental CSF backflow. Needle size: 25G with a needle length of 90mm excluding hub Needle, stylet and introducers must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 50) |  | Chemical World cc | 9.09 | [10 empty] | COMPLY pending confirmation of introducer length
42142502-00032 | 29 | Needle, spinal, Pencil point bevel; luer lock with introducer. The needle must be colour coded at the stylet cap and the size must be clearly stated, must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. The needle must fit securely into the introducer. The introducer must be streamlined and easy to handle. The introducer size must not reduce the effective length of the cannula. The introducer hub must be transparent to asses for accidental CSF backflow. Needle size: 25G with a needle length of 90mm excluding hub. Introducer needle length of >36mm with an introducer hub length of <19mm. The effective length of introducer and stylet with cannula is >76mm Needle, stylet and introducers must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 50) | N/A | SAFARMEX (PTY) LTD | 9.93 | Saclin | To be confirmed | Zhejiang Mediunion Healthcare Group Co. Ltd | China | 3days | 80% | US$ | 10000 Pieces | Yes | Box of 50 | No sample submitted for evaluation
42142502-00032 | 29 | Needle, spinal, Pencil point bevel; luer lock with introducer. The needle must be colour coded at the stylet cap and the size must be clearly stated, must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. The needle must fit securely into the introducer. The introducer must be streamlined and easy to handle. The introducer size must not reduce the effective length of the cannula. The introducer hub must be transparent to asses for accidental CSF backflow. Needle size: 25G with a needle length of 90mm excluding hub. Introducer needle length of >36mm with an introducer hub length of <19mm. The effective length of introducer and stylet with cannula is >76mm Needle, stylet and introducers must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 50) | Basic Line offer | Becton Dickinson | 12 | BD Whitacre Spinal Needle 25G X 90mm with Introducer | 405076 | Becton Dickinson | Fraga, Spain | 1 week | 76% | USD | 1 BOX | YES | 25 Units in a box | COMPLY pending confirmation of latex free
42142502-00032 | 29 | Needle, spinal, Pencil point bevel; luer lock with introducer. The needle must be colour coded at the stylet cap and the size must be clearly stated, must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. The needle must fit securely into the introducer. The introducer must be streamlined and easy to handle. The introducer size must not reduce the effective length of the cannula. The introducer hub must be transparent to asses for accidental CSF backflow. Needle size: 25G with a needle length of 90mm excluding hub. Introducer needle length of >36mm with an introducer hub length of <19mm. The effective length of introducer and stylet with cannula is >76mm Needle, stylet and introducers must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 50) | Basic Line Offer | BIOCOR HOSPITAL SUPPLIES PTY LTD | 12.6 | Medline | N/A | Medline | China | 60 | 85 | USD | 10 Boxes | Yes | N/A | COMPLY pending confirmation of introducer length
42142502-00032 | 29 | Needle, spinal, Pencil point bevel; luer lock with introducer. The needle must be colour coded at the stylet cap and the size must be clearly stated, must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. The needle must fit securely into the introducer. The introducer must be streamlined and easy to handle. The introducer size must not reduce the effective length of the cannula. The introducer hub must be transparent to asses for accidental CSF backflow. Needle size: 25G with a needle length of 90mm excluding hub. Introducer needle length of >36mm with an introducer hub length of <19mm. The effective length of introducer and stylet with cannula is >76mm Needle, stylet and introducers must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 50) | BOX OF 50 | MSIMED CC | 13.8 | KANG KANG | SPN011 | ZHENJIANG | CHINA | 25 | 20% | USD | 20000PCS | YES | BOX | No sample submitted for evaluation
42142502-00032 | 29 | Needle, spinal, Pencil point bevel; luer lock with introducer. The needle must be colour coded at the stylet cap and the size must be clearly stated, must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. The needle must fit securely into the introducer. The introducer must be streamlined and easy to handle. The introducer size must not reduce the effective length of the cannula. The introducer hub must be transparent to asses for accidental CSF backflow. Needle size: 25G with a needle length of 90mm excluding hub. Introducer needle length of >36mm with an introducer hub length of <19mm. The effective length of introducer and stylet with cannula is >76mm Needle, stylet and introducers must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 50) | 1 | GROBIR MEDICAL SUPPLIERS | 13.95 | DURAGAURD | DG.11.025.090 | PROMECON | CHINA | 21 | 80 | US DOLLAR | 25 | Yes | UOM SOLD IN PACKS OF 25 CAN SUPPLY 2 BOXES TO MEET SPECIFICATION REQUIREMENT OF 50 UNITS / MEETS SPECIFICATION | No sample submitted for evaluation.
42142502-00032 | 29 | Needle, spinal, Pencil point bevel; luer lock with introducer. The needle must be colour coded at the stylet cap and the size must be clearly stated, must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. The needle must fit securely into the introducer. The introducer must be streamlined and easy to handle. The introducer size must not reduce the effective length of the cannula. The introducer hub must be transparent to asses for accidental CSF backflow. Needle size: 25G with a needle length of 90mm excluding hub. Introducer needle length of >36mm with an introducer hub length of <19mm. The effective length of introducer and stylet with cannula is >76mm Needle, stylet and introducers must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 50) | Alternative Offer | BIOCOR HOSPITAL SUPPLIES PTY LTD | 14.73 | Technomed | TCN2590in | Techno Med | South Africa | 21 | 0 | N/A | 5 Boxes of 50 | Yes | N/A | The sample does not have an introducer
42142502-00032 | 29 | Needle, spinal, Pencil point bevel; luer lock with introducer. The needle must be colour coded at the stylet cap and the size must be clearly stated, must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. The needle must fit securely into the introducer. The introducer must be streamlined and easy to handle. The introducer size must not reduce the effective length of the cannula. The introducer hub must be transparent to asses for accidental CSF backflow. Needle size: 25G with a needle length of 90mm excluding hub. Introducer needle length of >36mm with an introducer hub length of <19mm. The effective length of introducer and stylet with cannula is >76mm Needle, stylet and introducers must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 50) |  | Zanotone (Pty) Ltd | 16.66 | Mekon |  | Mekon | [6 empty] | EaCH | COMPLY pending confirmation of latex free
42142502-00032 | 29 | Needle, spinal, Pencil point bevel; luer lock with introducer. The needle must be colour coded at the stylet cap and the size must be clearly stated, must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. The needle must fit securely into the introducer. The introducer must be streamlined and easy to handle. The introducer size must not reduce the effective length of the cannula. The introducer hub must be transparent to asses for accidental CSF backflow. Needle size: 25G with a needle length of 90mm excluding hub. Introducer needle length of >36mm with an introducer hub length of <19mm. The effective length of introducer and stylet with cannula is >76mm Needle, stylet and introducers must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 50) | OFFER 2 | SUPRA HEALTHCARE JOHANNESBURG PTY LTD | 26.37 | REVITAL | SPNPP25X90 | Revital Health | KENYA | Initial 60 days, thereafter 14 | 84 | USD | 1 Box | YES |  | COMPLY pending confirmation of latex free
42142502-00032 | 29 | Needle, spinal, Pencil point bevel; luer lock with introducer. The needle must be colour coded at the stylet cap and the size must be clearly stated, must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. The needle must fit securely into the introducer. The introducer must be streamlined and easy to handle. The introducer size must not reduce the effective length of the cannula. The introducer hub must be transparent to asses for accidental CSF backflow. Needle size: 25G with a needle length of 90mm excluding hub. Introducer needle length of >36mm with an introducer hub length of <19mm. The effective length of introducer and stylet with cannula is >76mm Needle, stylet and introducers must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 50) |  | ALMIKA TRADING PTY LTD | 26.45 | BD | 405076 | BD | USA | 14 | 0 | NO | 100 | YES | 1 | COMPLY pending confirmation of introducer length
42142502-00032 | 29 | Needle, spinal, Pencil point bevel; luer lock with introducer. The needle must be colour coded at the stylet cap and the size must be clearly stated, must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. The needle must fit securely into the introducer. The introducer must be streamlined and easy to handle. The introducer size must not reduce the effective length of the cannula. The introducer hub must be transparent to asses for accidental CSF backflow. Needle size: 25G with a needle length of 90mm excluding hub. Introducer needle length of >36mm with an introducer hub length of <19mm. The effective length of introducer and stylet with cannula is >76mm Needle, stylet and introducers must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 50) | OFFER 1 | SUPRA HEALTHCARE JOHANNESBURG PTY LTD | 31.44 | Egemen | SPNPP25X90 | Egemen | TURKEY | Initial 60 days, thereafter 14 | 84 | USD | 1 Box | YES |  | Offered an introducer length of 34mm instead of greater than 36mm
42142502-00032 | 29 | Needle, spinal, Pencil point bevel; luer lock with introducer. The needle must be colour coded at the stylet cap and the size must be clearly stated, must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. The needle must fit securely into the introducer. The introducer must be streamlined and easy to handle. The introducer size must not reduce the effective length of the cannula. The introducer hub must be transparent to asses for accidental CSF backflow. Needle size: 25G with a needle length of 90mm excluding hub. Introducer needle length of >36mm with an introducer hub length of <19mm. The effective length of introducer and stylet with cannula is >76mm Needle, stylet and introducers must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 50) | 1 | B. BRAUN MEDICAL (PTY) LTD | 32 | Pencan 25G x 3 1/2" (0.53 x 88mm) With Guide | 4502043 | B. Braun | Germany | 42 | 50% | Euro | 50 | Yes | Box of 50 | Offered an introducer length of 35mm instead of 36mm
42142502-00032 | 29 | Needle, spinal, Pencil point bevel; luer lock with introducer. The needle must be colour coded at the stylet cap and the size must be clearly stated, must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. The needle must fit securely into the introducer. The introducer must be streamlined and easy to handle. The introducer size must not reduce the effective length of the cannula. The introducer hub must be transparent to asses for accidental CSF backflow. Needle size: 25G with a needle length of 90mm excluding hub. Introducer needle length of >36mm with an introducer hub length of <19mm. The effective length of introducer and stylet with cannula is >76mm Needle, stylet and introducers must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 50) |  | LECHOBA MEDICAL TECHNOLOGIES (PTY) LTD | 48.86 | UNISIS UNIEVER | 5290761 | UNISIS CORPORATION | JAPAN | 21 DAYS | 75% | USD | 50 | YES | 25/BOX | No sample submitted for evaluation
42142502-00032 | 29 | Needle, spinal, Pencil point bevel; luer lock with introducer. The needle must be colour coded at the stylet cap and the size must be clearly stated, must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. The needle must fit securely into the introducer. The introducer must be streamlined and easy to handle. The introducer size must not reduce the effective length of the cannula. The introducer hub must be transparent to asses for accidental CSF backflow. Needle size: 25G with a needle length of 90mm excluding hub. Introducer needle length of >36mm with an introducer hub length of <19mm. The effective length of introducer and stylet with cannula is >76mm Needle, stylet and introducers must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 50) |  | Viking Critical Care (Pty) Ltd | 130.02 | VYGON | 69F18105 | VYGON | FRANCE | EX. STOCK 30 - 60 DAYS | 80% | EURO | 1 | YES | 1 | COMPLY pending confirmation of latex free
42142502-00032 | 29 | Needle, spinal, Pencil point bevel; luer lock with introducer. The needle must be colour coded at the stylet cap and the size must be clearly stated, must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. The needle must fit securely into the introducer. The introducer must be streamlined and easy to handle. The introducer size must not reduce the effective length of the cannula. The introducer hub must be transparent to asses for accidental CSF backflow. Needle size: 25G with a needle length of 90mm excluding hub. Introducer needle length of >36mm with an introducer hub length of <19mm. The effective length of introducer and stylet with cannula is >76mm Needle, stylet and introducers must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 50) BOX (BOX OF 25) |  | JUMLA MEDICAL SUPPLIES (PTY) LTD | 753.25 | EGEMEN | TSPPK2590 | TIBBI MEDICAL | TURKEY | 21 DAYS | 70% | US DOLLAR | 2 BOXES | YES | NIL | Offered an introducer length of 34mm instead of not less than 36mm
42142502-00032 | 29 | Needle, spinal, Pencil point bevel; luer lock with introducer. The needle must be colour coded at the stylet cap and the size must be clearly stated, must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. The needle must fit securely into the introducer. The introducer must be streamlined and easy to handle. The introducer size must not reduce the effective length of the cannula. The introducer hub must be transparent to asses for accidental CSF backflow. Needle size: 25G with a needle length of 90mm excluding hub. Introducer needle length of >36mm with an introducer hub length of <19mm. The effective length of introducer and stylet with cannula is >76mm Needle, stylet and introducers must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 50) |  | BAS MEDEXPRESS | 776.25 | Mercury | SN2390P |  | India | 1 - 3 days | 15% | N/A | As per client | Yes | Box | No sample submitted for evaluation
42142502-00033 | 30 | Needle, spinal, Pencil point bevel; luer lock with introducer. The needle must be colour coded at the stylet cap and the size must be clearly stated, must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. The needle must fit securely into the introducer. The introducer must be streamlined and easy to handle. The introducer size must not reduce the effective length of the cannula. The introducer hub must be transparent to asses for accidental CSF backflow. Needle size: 25G extra-length with a needle length of >110 to 120mm excluding hub. Introducer needle length of >35mm with an introducer hub length of <19mm. The effective length of introducer and stylet with cannula is ≥110mm Needle, stylet and introducers must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 50) |  | Medi-Core Technologies | 264.55 | ULIFE | 42142501-00033 | JIAXING U-LIFE MEDICAL DEVICE TECHNOLOGLY.,CO LTD | CHINA | 21 | 80 | USD | 1 box | YES | BOX OF 50 | COMPLY pending confirmation of introducer length |  | Recommended | 100
42142502-00033 | 30 | Needle, spinal, Pencil point bevel; luer lock with introducer. The needle must be colour coded at the stylet cap and the size must be clearly stated, must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. The needle must fit securely into the introducer. The introducer must be streamlined and easy to handle. The introducer size must not reduce the effective length of the cannula. The introducer hub must be transparent to asses for accidental CSF backflow. Needle size: 25G extra-length with a needle length of >110 to 120mm excluding hub. Introducer needle length of >35mm with an introducer hub length of <19mm. The effective length of introducer and stylet with cannula is ≥110mm Needle, stylet and introducers must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 50) | N/A | PROMED TECHNOLOGIES (PTY) LTD | 5.34 | U-LIFE | 502033 | JIAXING | CHINA | 14 DAYS | 85% | USD | 1 BOX | YES | BOX | COMPLY pending confirmation of latex free Accept with confirmation of the introducer length and introducer hub
42142502-00033 | 30 | Needle, spinal, Pencil point bevel; luer lock with introducer. The needle must be colour coded at the stylet cap and the size must be clearly stated, must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. The needle must fit securely into the introducer. The introducer must be streamlined and easy to handle. The introducer size must not reduce the effective length of the cannula. The introducer hub must be transparent to asses for accidental CSF backflow. Needle size: 25G extra-length with a needle length of >110 to 120mm excluding hub. Introducer needle length of >35mm with an introducer hub length of <19mm. The effective length of introducer and stylet with cannula is ≥110mm Needle, stylet and introducers must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 50) | BASIC LINE OFFER | A R FREIGHT PRIVATE LIMITED | 6.46 | SUMBOW | N/A | SUMBOW MEDICAL | CHINA | 60 DAYS | 85% | USD | 100 | YES | N/A | The needle does have an introducer needle
42142502-00033 | 30 | Needle, spinal, Pencil point bevel; luer lock with introducer. The needle must be colour coded at the stylet cap and the size must be clearly stated, must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. The needle must fit securely into the introducer. The introducer must be streamlined and easy to handle. The introducer size must not reduce the effective length of the cannula. The introducer hub must be transparent to asses for accidental CSF backflow. Needle size: 25G extra-length with a needle length of >110 to 120mm excluding hub. Introducer needle length of >35mm with an introducer hub length of <19mm. The effective length of introducer and stylet with cannula is ≥110mm Needle, stylet and introducers must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 50) |  | UNITRADE 1032 CC | 8.98 | SPINOJEC PLUS | 1 | SHANGHAI KINDLY ENTERPRISE DEVELOPMENT GROUP | ROC | 21 | 85 | USD | 500 | YES | 50 | No sample submitted for evaluation
42142502-00033 | 30 | Needle, spinal, Pencil point bevel; luer lock with introducer. The needle must be colour coded at the stylet cap and the size must be clearly stated, must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. The needle must fit securely into the introducer. The introducer must be streamlined and easy to handle. The introducer size must not reduce the effective length of the cannula. The introducer hub must be transparent to asses for accidental CSF backflow. Needle size: 25G extra-length with a needle length of >110 to 120mm excluding hub. Introducer needle length of >35mm with an introducer hub length of <19mm. The effective length of introducer and stylet with cannula is ≥110mm Needle, stylet and introducers must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 50) | N/A | SAFARMEX (PTY) LTD | 10.76 | Saclin | To be confirmed | Zhejiang Mediunion Healthcare Group Co. Ltd | China | 3days | 80% | US$ | 10000 Pieces | Yes | Box of 50 | No sample submitted for evaluation
42142502-00033 | 30 | Needle, spinal, Pencil point bevel; luer lock with introducer. The needle must be colour coded at the stylet cap and the size must be clearly stated, must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. The needle must fit securely into the introducer. The introducer must be streamlined and easy to handle. The introducer size must not reduce the effective length of the cannula. The introducer hub must be transparent to asses for accidental CSF backflow. Needle size: 25G extra-length with a needle length of >110 to 120mm excluding hub. Introducer needle length of >35mm with an introducer hub length of <19mm. The effective length of introducer and stylet with cannula is ≥110mm Needle, stylet and introducers must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 50) |  | Chemical World cc | 13.23 | MEDIJECT | NDSPP025 | RUNQIANG | CHINA | 3 | 85 | USD | 1000 | YES | 50 | No brand name. No product code. No introducer
42142502-00033 | 30 | Needle, spinal, Pencil point bevel; luer lock with introducer. The needle must be colour coded at the stylet cap and the size must be clearly stated, must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. The needle must fit securely into the introducer. The introducer must be streamlined and easy to handle. The introducer size must not reduce the effective length of the cannula. The introducer hub must be transparent to asses for accidental CSF backflow. Needle size: 25G extra-length with a needle length of >110 to 120mm excluding hub. Introducer needle length of >35mm with an introducer hub length of <19mm. The effective length of introducer and stylet with cannula is ≥110mm Needle, stylet and introducers must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 50) | Basic Line Offer | BIOCOR HOSPITAL SUPPLIES PTY LTD | 13.82 | Medline | N/A | Yangzhong Medline Industry Co.Ltd | China | 60 | 85 | USD | 10 Boxes | Yes | N/A | No sample submitted for evaluation
42142502-00033 | 30 | Needle, spinal, Pencil point bevel; luer lock with introducer. The needle must be colour coded at the stylet cap and the size must be clearly stated, must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. The needle must fit securely into the introducer. The introducer must be streamlined and easy to handle. The introducer size must not reduce the effective length of the cannula. The introducer hub must be transparent to asses for accidental CSF backflow. Needle size: 25G extra-length with a needle length of >110 to 120mm excluding hub. Introducer needle length of >35mm with an introducer hub length of <19mm. The effective length of introducer and stylet with cannula is ≥110mm Needle, stylet and introducers must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 50) | BOX OF 50 | MSIMED CC | 15 | KANG KANG | SPN012 | ZHENJIANG | CHINA | 25 | 20% | USD | 20000PCS | YES | BOX | No sample submitted for evaluation
42142502-00033 | 30 | Needle, spinal, Pencil point bevel; luer lock with introducer. The needle must be colour coded at the stylet cap and the size must be clearly stated, must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. The needle must fit securely into the introducer. The introducer must be streamlined and easy to handle. The introducer size must not reduce the effective length of the cannula. The introducer hub must be transparent to asses for accidental CSF backflow. Needle size: 25G extra-length with a needle length of >110 to 120mm excluding hub. Introducer needle length of >35mm with an introducer hub length of <19mm. The effective length of introducer and stylet with cannula is ≥110mm Needle, stylet and introducers must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 50) | 1 | GROBIR MEDICAL SUPPLIERS | 15.84 | DURAGAURD | DG.11.025.110 | PROMECON | CHINA | 21 | 80 | US DOLLAR | 25 | Yes | UOM SOLD IN PACKS OF 25 CAN SUPPLY 2 BOXES TO MEET SPECIFICATION REQUIREMENT OF 50 UNITS / MEETS SPECIFICATION | COMPLY pending confirmation of latex free
42142502-00033 | 30 | Needle, spinal, Pencil point bevel; luer lock with introducer. The needle must be colour coded at the stylet cap and the size must be clearly stated, must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. The needle must fit securely into the introducer. The introducer must be streamlined and easy to handle. The introducer size must not reduce the effective length of the cannula. The introducer hub must be transparent to asses for accidental CSF backflow. Needle size: 25G extra-length with a needle length of >110 to 120mm excluding hub. Introducer needle length of >35mm with an introducer hub length of <19mm. The effective length of introducer and stylet with cannula is ≥110mm Needle, stylet and introducers must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 50) | Alternative Offer | BIOCOR HOSPITAL SUPPLIES PTY LTD | 15.97 | Technomed | TCN25120in | Techno Med | South Africa | 21 | 0 | N/A | 5 Boxes of 50 | Yes | N/A | COMPLY pending confirmation of introducer length
42142502-00033 | 30 | Needle, spinal, Pencil point bevel; luer lock with introducer. The needle must be colour coded at the stylet cap and the size must be clearly stated, must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. The needle must fit securely into the introducer. The introducer must be streamlined and easy to handle. The introducer size must not reduce the effective length of the cannula. The introducer hub must be transparent to asses for accidental CSF backflow. Needle size: 25G extra-length with a needle length of >110 to 120mm excluding hub. Introducer needle length of >35mm with an introducer hub length of <19mm. The effective length of introducer and stylet with cannula is ≥110mm Needle, stylet and introducers must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 50) |  | Zanotone (Pty) Ltd | 17.22 | Mekon |  | Mekon | [6 empty] | EACH | No sample submitted for evaluation
42142502-00033 | 30 | Needle, spinal, Pencil point bevel; luer lock with introducer. The needle must be colour coded at the stylet cap and the size must be clearly stated, must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. The needle must fit securely into the introducer. The introducer must be streamlined and easy to handle. The introducer size must not reduce the effective length of the cannula. The introducer hub must be transparent to asses for accidental CSF backflow. Needle size: 25G extra-length with a needle length of >110 to 120mm excluding hub. Introducer needle length of >35mm with an introducer hub length of <19mm. The effective length of introducer and stylet with cannula is ≥110mm Needle, stylet and introducers must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 50) | Basic Line offer | Becton Dickinson | 17.5 | BD Whitacre Spinal Needle 25G X103mm with Introducer | 405112 | Becton Dickinson | Fraga, Spain | 1 week | 76% | USD | 1 BOX | YES | 25 Units in a box | Offered a needle length of 103mm instead of a needle length of greater than 110 to 120mm
42142502-00033 | 30 | Needle, spinal, Pencil point bevel; luer lock with introducer. The needle must be colour coded at the stylet cap and the size must be clearly stated, must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. The needle must fit securely into the introducer. The introducer must be streamlined and easy to handle. The introducer size must not reduce the effective length of the cannula. The introducer hub must be transparent to asses for accidental CSF backflow. Needle size: 25G extra-length with a needle length of >110 to 120mm excluding hub. Introducer needle length of >35mm with an introducer hub length of <19mm. The effective length of introducer and stylet with cannula is ≥110mm Needle, stylet and introducers must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 50) | OFFER 2 | SUPRA HEALTHCARE JOHANNESBURG PTY LTD | 26.37 | REVITAL | SPNPP25X120 | Revital Health | KENYA | Initial 60 days, thereafter 14 | 84 | USD | 1 Box | YES |  | No sample submitted for evaluation
42142502-00033 | 30 | Needle, spinal, Pencil point bevel; luer lock with introducer. The needle must be colour coded at the stylet cap and the size must be clearly stated, must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. The needle must fit securely into the introducer. The introducer must be streamlined and easy to handle. The introducer size must not reduce the effective length of the cannula. The introducer hub must be transparent to asses for accidental CSF backflow. Needle size: 25G extra-length with a needle length of >110 to 120mm excluding hub. Introducer needle length of >35mm with an introducer hub length of <19mm. The effective length of introducer and stylet with cannula is ≥110mm Needle, stylet and introducers must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 50) | 1 | B. BRAUN MEDICAL (PTY) LTD | 32 | Pencan 25G x 4" (0.53 x 103mm) With Guide | 4502116 | B. Braun | Germany | 42 | 50% | Euro | 50 | Yes | Box of 50 | Offered needle length of 103mm instead of a length of 110mm – 120mm
42142502-00033 | 30 | Needle, spinal, Pencil point bevel; luer lock with introducer. The needle must be colour coded at the stylet cap and the size must be clearly stated, must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. The needle must fit securely into the introducer. The introducer must be streamlined and easy to handle. The introducer size must not reduce the effective length of the cannula. The introducer hub must be transparent to asses for accidental CSF backflow. Needle size: 25G extra-length with a needle length of >110 to 120mm excluding hub. Introducer needle length of >35mm with an introducer hub length of <19mm. The effective length of introducer and stylet with cannula is ≥110mm Needle, stylet and introducers must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 50) | OFFER 1 | SUPRA HEALTHCARE JOHANNESBURG PTY LTD | 37.86 | Egemen | SPNPP25X120 | Egemen | TURKEY | Initial 60 days, thereafter 14 | 84 | USD | 1 Box | YES |  | COMPLY pending confirmation of latex free
42142502-00033 | 30 | Needle, spinal, Pencil point bevel; luer lock with introducer. The needle must be colour coded at the stylet cap and the size must be clearly stated, must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. The needle must fit securely into the introducer. The introducer must be streamlined and easy to handle. The introducer size must not reduce the effective length of the cannula. The introducer hub must be transparent to asses for accidental CSF backflow. Needle size: 25G extra-length with a needle length of >110 to 120mm excluding hub. Introducer needle length of >35mm with an introducer hub length of <19mm. The effective length of introducer and stylet with cannula is ≥110mm Needle, stylet and introducers must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 50) |  | ALMIKA TRADING PTY LTD | 39.68 | BD | 405112 | BD | USA | 14 | 0 | NO | 100 | YES | 1 | offered a needle length 103mm instead of 11omm – 120mm
42142502-00033 | 30 | Needle, spinal, Pencil point bevel; luer lock with introducer. The needle must be colour coded at the stylet cap and the size must be clearly stated, must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. The needle must fit securely into the introducer. The introducer must be streamlined and easy to handle. The introducer size must not reduce the effective length of the cannula. The introducer hub must be transparent to asses for accidental CSF backflow. Needle size: 25G extra-length with a needle length of >110 to 120mm excluding hub. Introducer needle length of >35mm with an introducer hub length of <19mm. The effective length of introducer and stylet with cannula is ≥110mm Needle, stylet and introducers must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 50) |  | LECHOBA MEDICAL TECHNOLOGIES (PTY) LTD | 52.82 | UNISIS UNIEVER | 5291771 | UNISIS CORPORATION | JAPAN | 21 DAYS | 75% | USD | 50 | YES | 25/BOX | No sample submitted for evaluation
42142502-00033 | 30 | Needle, spinal, Pencil point bevel; luer lock with introducer. The needle must be colour coded at the stylet cap and the size must be clearly stated, must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. The needle must fit securely into the introducer. The introducer must be streamlined and easy to handle. The introducer size must not reduce the effective length of the cannula. The introducer hub must be transparent to asses for accidental CSF backflow. Needle size: 25G extra-length with a needle length of >110 to 120mm excluding hub. Introducer needle length of >35mm with an introducer hub length of <19mm. The effective length of introducer and stylet with cannula is ≥110mm Needle, stylet and introducers must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 50) |  | Viking Critical Care (Pty) Ltd | 225.13 | VYGON | 69F18125 | VYGON | FRANCE | EX. STOCK 30 - 60 DAYS | 80% | EURO | 1 | YES | 1 | COMPLY pending confirmation of latex free
42142502-00033 | 30 | Needle, spinal, Pencil point bevel; luer lock with introducer. The needle must be colour coded at the stylet cap and the size must be clearly stated, must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. The needle must fit securely into the introducer. The introducer must be streamlined and easy to handle. The introducer size must not reduce the effective length of the cannula. The introducer hub must be transparent to asses for accidental CSF backflow. Needle size: 25G extra-length with a needle length of >110 to 120mm excluding hub. Introducer needle length of >35mm with an introducer hub length of <19mm. The effective length of introducer and stylet with cannula is ≥110mm Needle, stylet and introducers must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 50) |  | BAS MEDEXPRESS | 776.25 | Mercury | SN2490P |  | India | 1 - 3 days | 15% | N/A | As per client | Yes | Box | No sample submitted for evaluation
42142502-00033 | 30 | Needle, spinal, Pencil point bevel; luer lock with introducer. The needle must be colour coded at the stylet cap and the size must be clearly stated, must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. The needle must fit securely into the introducer. The introducer must be streamlined and easy to handle. The introducer size must not reduce the effective length of the cannula. The introducer hub must be transparent to asses for accidental CSF backflow. Needle size: 25G extra-length with a needle length of >110 to 120mm excluding hub. Introducer needle length of >35mm with an introducer hub length of <19mm. The effective length of introducer and stylet with cannula is ≥110mm Needle, stylet and introducers must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 50) BOX (BOX OF 25) |  | JUMLA MEDICAL SUPPLIES (PTY) LTD | 955.65 | EGEMEN | TSPPK25120 | TIBBI MEDICAL | TURKEY | 21 DAYS | 70% | US DOLLAR | 2 BOXES | YES | NIL | COMPLY pending confirmation of introducer length
42142502-00034 | 31 | Needle, spinal, Pencil point bevel; luer lock with introducer. The needle must be colour coded at the stylet cap and the size must be clearly stated, must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. The needle must fit securely into the introducer. The introducer must be streamlined and easy to handle. The introducer size must not reduce the effective length of the cannula. The introducer hub must be transparent to asses for accidental CSF backflow. Needle size: 26G with a needle length of 90mm excluding hub. Introducer needle length of >36mm with an introducer hub length of <19mm. The effective length of introducer and stylet with cannula is >76mm Needle, stylet and introducers must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 50) |  | Medi-Core Technologies | 264.55 | ULIFE | 42142502-00034 | JIAXING U-LIFE MEDICAL DEVICE TECHNOLOGLY.,CO LTD | CHINA | 21 | 80 | USD | 1 box | YES | BOX OF 50 | COMPLY pending confirmation of introducer length |  | Recommended | 100
42142502-00034 | 31 | Needle, spinal, Pencil point bevel; luer lock with introducer. The needle must be colour coded at the stylet cap and the size must be clearly stated, must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. The needle must fit securely into the introducer. The introducer must be streamlined and easy to handle. The introducer size must not reduce the effective length of the cannula. The introducer hub must be transparent to asses for accidental CSF backflow. Needle size: 26G with a needle length of 90mm excluding hub. Introducer needle length of >36mm with an introducer hub length of <19mm. The effective length of introducer and stylet with cannula is >76mm Needle, stylet and introducers must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 50) | N/A | PROMED TECHNOLOGIES (PTY) LTD | 5.34 | U-LIFE | 502034 | JIAXING | CHINA | 14 DAYS | 85% | USD | 1 BOX | YES | BOX | COMPLY pending confirmation of latex free Accept with confirmation of the introducer length and introducer hub
42142502-00034 | 31 | Needle, spinal, Pencil point bevel; luer lock with introducer. The needle must be colour coded at the stylet cap and the size must be clearly stated, must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. The needle must fit securely into the introducer. The introducer must be streamlined and easy to handle. The introducer size must not reduce the effective length of the cannula. The introducer hub must be transparent to asses for accidental CSF backflow. Needle size: 26G with a needle length of 90mm excluding hub. Introducer needle length of >36mm with an introducer hub length of <19mm. The effective length of introducer and stylet with cannula is >76mm Needle, stylet and introducers must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 50) | BASIC LINE OFFER | A R FREIGHT PRIVATE LIMITED | 6.12 | SUMBOW | N/A | SUMBOW MEDICAL | CHINA | 60 DAYS | 85% | USD | 100 | YES | N/A | The needle does have an introducer needle
42142502-00034 | 31 | Needle, spinal, Pencil point bevel; luer lock with introducer. The needle must be colour coded at the stylet cap and the size must be clearly stated, must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. The needle must fit securely into the introducer. The introducer must be streamlined and easy to handle. The introducer size must not reduce the effective length of the cannula. The introducer hub must be transparent to asses for accidental CSF backflow. Needle size: 26G with a needle length of 90mm excluding hub. Introducer needle length of >36mm with an introducer hub length of <19mm. The effective length of introducer and stylet with cannula is >76mm Needle, stylet and introducers must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 50) |  | UNITRADE 1032 CC | 7.94 | SPINOJEC PLUS | 1 | SHANGHAI KINDLY ENTERPRISE DEVELOPMENT GROUP | ROC | 21 | 85 | USD | 500 | YES | 50 | No sample submitted for evaluation
42142502-00034 | 31 | Needle, spinal, Pencil point bevel; luer lock with introducer. The needle must be colour coded at the stylet cap and the size must be clearly stated, must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. The needle must fit securely into the introducer. The introducer must be streamlined and easy to handle. The introducer size must not reduce the effective length of the cannula. The introducer hub must be transparent to asses for accidental CSF backflow. Needle size: 26G with a needle length of 90mm excluding hub. Introducer needle length of >36mm with an introducer hub length of <19mm. The effective length of introducer and stylet with cannula is >76mm Needle, stylet and introducers must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 50) |  | Chemical World cc | 9.09 | MEDIJECT | NDSPPI26 | RUNQIANG | CHINA | 3 | 85 | USD | 1000 | YES | 50 | The stylet is not stable and shakes or quiver and does not allow for easy re-insertion into the cannula.
42142502-00034 | 31 | Needle, spinal, Pencil point bevel; luer lock with introducer. The needle must be colour coded at the stylet cap and the size must be clearly stated, must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. The needle must fit securely into the introducer. The introducer must be streamlined and easy to handle. The introducer size must not reduce the effective length of the cannula. The introducer hub must be transparent to asses for accidental CSF backflow. Needle size: 26G with a needle length of 90mm excluding hub. Introducer needle length of >36mm with an introducer hub length of <19mm. The effective length of introducer and stylet with cannula is >76mm Needle, stylet and introducers must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 50) | N/A | SAFARMEX (PTY) LTD | 9.93 | Saclin | To be confirmed | Zhejiang Mediunion Healthcare Group Co. Ltd | China | 3days | 80% | US$ | 10000 Pieces | Yes | Box of 50 | No sample submitted for evaluation
42142502-00034 | 31 | Needle, spinal, Pencil point bevel; luer lock with introducer. The needle must be colour coded at the stylet cap and the size must be clearly stated, must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. The needle must fit securely into the introducer. The introducer must be streamlined and easy to handle. The introducer size must not reduce the effective length of the cannula. The introducer hub must be transparent to asses for accidental CSF backflow. Needle size: 26G with a needle length of 90mm excluding hub. Introducer needle length of >36mm with an introducer hub length of <19mm. The effective length of introducer and stylet with cannula is >76mm Needle, stylet and introducers must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 50) | Basic Line Offer | BIOCOR HOSPITAL SUPPLIES PTY LTD | 12.6 | Medline | N/A | Yangzhong Medline Industry Co.Ltd | China | 60 | 85 | USD | 10 Boxes | Yes | N/A | No sample submitted for evaluation
42142502-00034 | 31 | Needle, spinal, Pencil point bevel; luer lock with introducer. The needle must be colour coded at the stylet cap and the size must be clearly stated, must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. The needle must fit securely into the introducer. The introducer must be streamlined and easy to handle. The introducer size must not reduce the effective length of the cannula. The introducer hub must be transparent to asses for accidental CSF backflow. Needle size: 26G with a needle length of 90mm excluding hub. Introducer needle length of >36mm with an introducer hub length of <19mm. The effective length of introducer and stylet with cannula is >76mm Needle, stylet and introducers must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 50) | 1 | GROBIR MEDICAL SUPPLIERS | 13.95 | DURAGAURD | DG.11.026.090 | PROMECON | CHINA | 21 | 80 | US DOLLAR | 25 | Yes | UOM SOLD IN PACKS OF 25 CAN SUPPLY 2 BOXES TO MEET SPECIFICATION REQUIREMENT OF 50 UNITS / MEETS SPECIFICATION | Offered a quickie point instead of a pencil point.
42142502-00034 | 31 | Needle, spinal, Pencil point bevel; luer lock with introducer. The needle must be colour coded at the stylet cap and the size must be clearly stated, must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. The needle must fit securely into the introducer. The introducer must be streamlined and easy to handle. The introducer size must not reduce the effective length of the cannula. The introducer hub must be transparent to asses for accidental CSF backflow. Needle size: 26G with a needle length of 90mm excluding hub. Introducer needle length of >36mm with an introducer hub length of <19mm. The effective length of introducer and stylet with cannula is >76mm Needle, stylet and introducers must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 50) | Alternative Offer | BIOCOR HOSPITAL SUPPLIES PTY LTD | 14.73 | Technomed | TCN2690in | Techno Med | South Africa | 21 | 0 | N/A | 5 Boxes of 50 | Yes | N/A | COMPLY pending confirmation of introducer length
42142502-00034 | 31 | Needle, spinal, Pencil point bevel; luer lock with introducer. The needle must be colour coded at the stylet cap and the size must be clearly stated, must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. The needle must fit securely into the introducer. The introducer must be streamlined and easy to handle. The introducer size must not reduce the effective length of the cannula. The introducer hub must be transparent to asses for accidental CSF backflow. Needle size: 26G with a needle length of 90mm excluding hub. Introducer needle length of >36mm with an introducer hub length of <19mm. The effective length of introducer and stylet with cannula is >76mm Needle, stylet and introducers must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 50) | BOX OF 50 | MSIMED CC | 15 | KANG KANG | SPN013 | ZHENJIANG | CHINA | 25 | 20% | USD | 20000PCS | YES | BOX | No sample submitted for evaluation
42142502-00034 | 31 | Needle, spinal, Pencil point bevel; luer lock with introducer. The needle must be colour coded at the stylet cap and the size must be clearly stated, must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. The needle must fit securely into the introducer. The introducer must be streamlined and easy to handle. The introducer size must not reduce the effective length of the cannula. The introducer hub must be transparent to asses for accidental CSF backflow. Needle size: 26G with a needle length of 90mm excluding hub. Introducer needle length of >36mm with an introducer hub length of <19mm. The effective length of introducer and stylet with cannula is >76mm Needle, stylet and introducers must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 50) |  | Zanotone (Pty) Ltd | 16.66 | Mekon |  | Mekon | [6 empty] | Each | No sample submitted for evaluation
42142502-00034 | 31 | Needle, spinal, Pencil point bevel; luer lock with introducer. The needle must be colour coded at the stylet cap and the size must be clearly stated, must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. The needle must fit securely into the introducer. The introducer must be streamlined and easy to handle. The introducer size must not reduce the effective length of the cannula. The introducer hub must be transparent to asses for accidental CSF backflow. Needle size: 26G with a needle length of 90mm excluding hub. Introducer needle length of >36mm with an introducer hub length of <19mm. The effective length of introducer and stylet with cannula is >76mm Needle, stylet and introducers must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 50) | OFFER 1 | SUPRA HEALTHCARE JOHANNESBURG PTY LTD | 26.37 | REVITAL | SPNPP26X90 | Revital Health | KENYA | Initial 60 days, thereafter 14 | 84 | USD | 1 Box | YES |  | COMPLY pending confirmation of latex free
42142502-00034 | 31 | Needle, spinal, Pencil point bevel; luer lock with introducer. The needle must be colour coded at the stylet cap and the size must be clearly stated, must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. The needle must fit securely into the introducer. The introducer must be streamlined and easy to handle. The introducer size must not reduce the effective length of the cannula. The introducer hub must be transparent to asses for accidental CSF backflow. Needle size: 26G with a needle length of 90mm excluding hub. Introducer needle length of >36mm with an introducer hub length of <19mm. The effective length of introducer and stylet with cannula is >76mm Needle, stylet and introducers must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 50) | OFFER 2 | SUPRA HEALTHCARE JOHANNESBURG PTY LTD | 31.44 | Egemen | SPNPP26X90 | Egemen | TURKEY | Initial 60 days, thereafter 14 | 84 | USD | 1 Box | YES |  | Offered an introducer length of 34mm instead of greater than 36mm
42142502-00034 | 31 | Needle, spinal, Pencil point bevel; luer lock with introducer. The needle must be colour coded at the stylet cap and the size must be clearly stated, must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. The needle must fit securely into the introducer. The introducer must be streamlined and easy to handle. The introducer size must not reduce the effective length of the cannula. The introducer hub must be transparent to asses for accidental CSF backflow. Needle size: 26G with a needle length of 90mm excluding hub. Introducer needle length of >36mm with an introducer hub length of <19mm. The effective length of introducer and stylet with cannula is >76mm Needle, stylet and introducers must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 50) |  | LECHOBA MEDICAL TECHNOLOGIES (PTY) LTD | 49.81 | UNISIS UNIEVER | 5300661 | UNISIS CORPORATION | JAPAN | 21 DAYS | 75% | USD | 50 | YES | 25/BOX | No sample submitted for evaluation
42142502-00034 | 31 | Needle, spinal, Pencil point bevel; luer lock with introducer. The needle must be colour coded at the stylet cap and the size must be clearly stated, must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. The needle must fit securely into the introducer. The introducer must be streamlined and easy to handle. The introducer size must not reduce the effective length of the cannula. The introducer hub must be transparent to asses for accidental CSF backflow. Needle size: 26G with a needle length of 90mm excluding hub. Introducer needle length of >36mm with an introducer hub length of <19mm. The effective length of introducer and stylet with cannula is >76mm Needle, stylet and introducers must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 50) |  | Viking Critical Care (Pty) Ltd | 197.34 | VYGON | 69F18194 | VYGON | FRANCE | EX. STOCK 30 - 60 DAYS | 80% | EURO | 1 | YES | 1 | COMPLY pending confirmation of latex free
42142502-00034 | 31 | Needle, spinal, Pencil point bevel; luer lock with introducer. The needle must be colour coded at the stylet cap and the size must be clearly stated, must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. The needle must fit securely into the introducer. The introducer must be streamlined and easy to handle. The introducer size must not reduce the effective length of the cannula. The introducer hub must be transparent to asses for accidental CSF backflow. Needle size: 26G with a needle length of 90mm excluding hub. Introducer needle length of >36mm with an introducer hub length of <19mm. The effective length of introducer and stylet with cannula is >76mm Needle, stylet and introducers must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 50) BOX (BOX OF 25) |  | JUMLA MEDICAL SUPPLIES (PTY) LTD | 753.25 | EGEMEN | TSPPK2690 | TIBBI MEDICAL | TURKEY | 21 DAYS | 70% | US DOLLAR | 2 BOXES | YES | NIL | COMPLY pending confirmation of introducer length
42142502-00034 | 31 | Needle, spinal, Pencil point bevel; luer lock with introducer. The needle must be colour coded at the stylet cap and the size must be clearly stated, must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. The needle must fit securely into the introducer. The introducer must be streamlined and easy to handle. The introducer size must not reduce the effective length of the cannula. The introducer hub must be transparent to asses for accidental CSF backflow. Needle size: 26G with a needle length of 90mm excluding hub. Introducer needle length of >36mm with an introducer hub length of <19mm. The effective length of introducer and stylet with cannula is >76mm Needle, stylet and introducers must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 50) |  | BAS MEDEXPRESS | 776.25 | Mercury | SN2590P |  | India | 1 - 3 days | 15% | N/A | As per client | Yes | Box | No sample submitted for evaluation
42142502-00035 | 32 | Needle, spinal, Pencil point bevel; luer lock with introducer. The needle must be colour coded at the stylet cap and the size must be clearly stated, must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. The needle must fit securely into the introducer. The introducer must be streamlined and easy to handle. The introducer size must not reduce the effective length of the cannula. The introducer hub must be transparent to asses for accidental CSF backflow. Needle size: 26G extra-length with a needle length of >110 to 120mm excluding hub. Introducer needle length of >36mm with an introducer hub length of <19mm. The effective length of introducer and stylet with cannula is ≥110mm Needle, stylet and introducers must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 50) |  | Medi-Core Technologies | 264.55 | ULIFE | 42142502-00035 | JIAXING U-LIFE MEDICAL DEVICE TECHNOLOGLY.,CO LTD | CHINA | 21 | 80 | USD | 1 box | YES | BOX OF 50 | COMPLY pending confirmation of introducer length |  | Recommended | 100
42142502-00035 | 32 | Needle, spinal, Pencil point bevel; luer lock with introducer. The needle must be colour coded at the stylet cap and the size must be clearly stated, must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. The needle must fit securely into the introducer. The introducer must be streamlined and easy to handle. The introducer size must not reduce the effective length of the cannula. The introducer hub must be transparent to asses for accidental CSF backflow. Needle size: 26G extra-length with a needle length of >110 to 120mm excluding hub. Introducer needle length of >36mm with an introducer hub length of <19mm. The effective length of introducer and stylet with cannula is ≥110mm Needle, stylet and introducers must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 50) | N/A | PROMED TECHNOLOGIES (PTY) LTD | 5.34 | U-LIFE | 502035 | JIAXING | CHINA | 14 DAYS | 85% | USD | 1 BOX | YES | BOX | COMPLY pending confirmation of latex free Accept with confirmation of the introducer length and introducer hub
42142502-00035 | 32 | Needle, spinal, Pencil point bevel; luer lock with introducer. The needle must be colour coded at the stylet cap and the size must be clearly stated, must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. The needle must fit securely into the introducer. The introducer must be streamlined and easy to handle. The introducer size must not reduce the effective length of the cannula. The introducer hub must be transparent to asses for accidental CSF backflow. Needle size: 26G extra-length with a needle length of >110 to 120mm excluding hub. Introducer needle length of >36mm with an introducer hub length of <19mm. The effective length of introducer and stylet with cannula is ≥110mm Needle, stylet and introducers must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 50) |  | UNITRADE 1032 CC | 7.94 | SPINOJEC PLUS | 1 | SHANGHAI KINDLY ENTERPRISE DEVELOPMENT GROUP | ROC | 21 | 85 | USD | 500 | YES | 50 | No sample submitted for evaluation
42142502-00035 | 32 | Needle, spinal, Pencil point bevel; luer lock with introducer. The needle must be colour coded at the stylet cap and the size must be clearly stated, must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. The needle must fit securely into the introducer. The introducer must be streamlined and easy to handle. The introducer size must not reduce the effective length of the cannula. The introducer hub must be transparent to asses for accidental CSF backflow. Needle size: 26G extra-length with a needle length of >110 to 120mm excluding hub. Introducer needle length of >36mm with an introducer hub length of <19mm. The effective length of introducer and stylet with cannula is ≥110mm Needle, stylet and introducers must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 50) | N/A | SAFARMEX (PTY) LTD | 10.76 | Saclin | To be confirmed | Zhejiang Mediunion Healthcare Group Co. Ltd | China | 3days | 80% | US$ | 10000 Pieces | Yes | Box of 50 | No sample submitted for evaluation
42142502-00035 | 32 | Needle, spinal, Pencil point bevel; luer lock with introducer. The needle must be colour coded at the stylet cap and the size must be clearly stated, must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. The needle must fit securely into the introducer. The introducer must be streamlined and easy to handle. The introducer size must not reduce the effective length of the cannula. The introducer hub must be transparent to asses for accidental CSF backflow. Needle size: 26G extra-length with a needle length of >110 to 120mm excluding hub. Introducer needle length of >36mm with an introducer hub length of <19mm. The effective length of introducer and stylet with cannula is ≥110mm Needle, stylet and introducers must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 50) | Basic Line Offer | BIOCOR HOSPITAL SUPPLIES PTY LTD | 13.82 | Medline | N/A | Yangzhong Medline Industry Co.Ltd | China | 60 | 85 | USD | 10 Boxes | Yes | N/A | No sample submitted for evaluation
42142502-00035 | 32 | Needle, spinal, Pencil point bevel; luer lock with introducer. The needle must be colour coded at the stylet cap and the size must be clearly stated, must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. The needle must fit securely into the introducer. The introducer must be streamlined and easy to handle. The introducer size must not reduce the effective length of the cannula. The introducer hub must be transparent to asses for accidental CSF backflow. Needle size: 26G extra-length with a needle length of >110 to 120mm excluding hub. Introducer needle length of >36mm with an introducer hub length of <19mm. The effective length of introducer and stylet with cannula is ≥110mm Needle, stylet and introducers must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 50) | BOX OF 50 | MSIMED CC | 14.2 | KANG KANG | SPN014 | ZHENJIANG | CHINA | 25 | 20% | USD | 20000PCS | YES | BOX | No sample submitted for evaluation
42142502-00035 | 32 | Needle, spinal, Pencil point bevel; luer lock with introducer. The needle must be colour coded at the stylet cap and the size must be clearly stated, must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. The needle must fit securely into the introducer. The introducer must be streamlined and easy to handle. The introducer size must not reduce the effective length of the cannula. The introducer hub must be transparent to asses for accidental CSF backflow. Needle size: 26G extra-length with a needle length of >110 to 120mm excluding hub. Introducer needle length of >36mm with an introducer hub length of <19mm. The effective length of introducer and stylet with cannula is ≥110mm Needle, stylet and introducers must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 50) | 1 | GROBIR MEDICAL SUPPLIERS | 15.84 | DURAGAURD | DG.11.026.110 | PROMECON | CHINA | 21 | 80 | US DOLLAR | 25 | Yes | UOM SOLD IN PACKS OF 25 CAN SUPPLY 2 BOXES TO MEET SPECIFICATION REQUIREMENT OF 50 UNITS / MEETS SPECIFICATION | COMPLY pending confirmation of latex free
42142502-00035 | 32 | Needle, spinal, Pencil point bevel; luer lock with introducer. The needle must be colour coded at the stylet cap and the size must be clearly stated, must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. The needle must fit securely into the introducer. The introducer must be streamlined and easy to handle. The introducer size must not reduce the effective length of the cannula. The introducer hub must be transparent to asses for accidental CSF backflow. Needle size: 26G extra-length with a needle length of >110 to 120mm excluding hub. Introducer needle length of >36mm with an introducer hub length of <19mm. The effective length of introducer and stylet with cannula is ≥110mm Needle, stylet and introducers must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 50) | Alternative Offer | BIOCOR HOSPITAL SUPPLIES PTY LTD | 15.97 | Technomed | TCN26120in | Techno Med | South Africa | 21 | 0 | N/A | 5 Boxes of 50 | Yes | N/A | No sample submitted for evaluation
42142502-00035 | 32 | Needle, spinal, Pencil point bevel; luer lock with introducer. The needle must be colour coded at the stylet cap and the size must be clearly stated, must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. The needle must fit securely into the introducer. The introducer must be streamlined and easy to handle. The introducer size must not reduce the effective length of the cannula. The introducer hub must be transparent to asses for accidental CSF backflow. Needle size: 26G extra-length with a needle length of >110 to 120mm excluding hub. Introducer needle length of >36mm with an introducer hub length of <19mm. The effective length of introducer and stylet with cannula is ≥110mm Needle, stylet and introducers must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 50) |  | Zanotone (Pty) Ltd | 17.22 | Mekon |  | Mekon | [6 empty] | Each | No sample submitted for evaluation
42142502-00035 | 32 | Needle, spinal, Pencil point bevel; luer lock with introducer. The needle must be colour coded at the stylet cap and the size must be clearly stated, must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. The needle must fit securely into the introducer. The introducer must be streamlined and easy to handle. The introducer size must not reduce the effective length of the cannula. The introducer hub must be transparent to asses for accidental CSF backflow. Needle size: 26G extra-length with a needle length of >110 to 120mm excluding hub. Introducer needle length of >36mm with an introducer hub length of <19mm. The effective length of introducer and stylet with cannula is ≥110mm Needle, stylet and introducers must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 50) |  | SUPRA HEALTHCARE JOHANNESBURG PTY LTD | 26.37 | REVITAL | SPNPP26X110 | Revital Health | KENYA | Initial 60 days, thereafter 14 | 84 | USD | 1 Box | YES |  | No sample submitted for evaluation
42142502-00035 | 32 | Needle, spinal, Pencil point bevel; luer lock with introducer. The needle must be colour coded at the stylet cap and the size must be clearly stated, must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. The needle must fit securely into the introducer. The introducer must be streamlined and easy to handle. The introducer size must not reduce the effective length of the cannula. The introducer hub must be transparent to asses for accidental CSF backflow. Needle size: 26G extra-length with a needle length of >110 to 120mm excluding hub. Introducer needle length of >36mm with an introducer hub length of <19mm. The effective length of introducer and stylet with cannula is ≥110mm Needle, stylet and introducers must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 50) |  | Viking Critical Care (Pty) Ltd | 197.34 | VYGON | 69F18194 | VYGON | FRANCE | EX. STOCK 30 - 60 DAYS | 80% | EURO | 1 | NO, OUR PRODUCT IS 90MM | 1 | Offered a needle length of 90mm instead of the needle length greater 110mm to 120mm
42142502-00035 | 32 | Needle, spinal, Pencil point bevel; luer lock with introducer. The needle must be colour coded at the stylet cap and the size must be clearly stated, must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. The needle must fit securely into the introducer. The introducer must be streamlined and easy to handle. The introducer size must not reduce the effective length of the cannula. The introducer hub must be transparent to asses for accidental CSF backflow. Needle size: 26G extra-length with a needle length of >110 to 120mm excluding hub. Introducer needle length of >36mm with an introducer hub length of <19mm. The effective length of introducer and stylet with cannula is ≥110mm Needle, stylet and introducers must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 50) |  | BAS MEDEXPRESS | 776.25 | Mercury | SN2690P |  | India | 1 - 3 days | 15% | N/A | As per client | Yes | Box | No sample submitted for evaluation
42142502-00036 | 33 | Needle, spinal, Pencil point bevel; luer lock with introducer. The needle must be colour coded at the stylet cap and the size must be clearly stated, must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. The needle must fit securely into the introducer. The introducer must be streamlined and easy to handle. The introducer size must not reduce the effective length of the cannula. The introducer hub must be transparent to asses for accidental CSF backflow. Needle size: 27G with a needle length of 90mm excluding hub. Introducer needle length of >35mm with an introducer hub length of <19mm. The effective length of introducer and stylet with cannula is >76mm Needle, stylet and introducers must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 50) |  | Medi-Core Technologies | 264.55 | ULIFE | 42142502-00036 | JIAXING U-LIFE MEDICAL DEVICE TECHNOLOGLY.,CO LTD | CHINA | 21 | 80 | USD | 1 box | YES | BOX OF 50 | COMPLY pending confirmation of introducer length |  | Recommended | 100
42142502-00036 | 33 | Needle, spinal, Pencil point bevel; luer lock with introducer. The needle must be colour coded at the stylet cap and the size must be clearly stated, must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. The needle must fit securely into the introducer. The introducer must be streamlined and easy to handle. The introducer size must not reduce the effective length of the cannula. The introducer hub must be transparent to asses for accidental CSF backflow. Needle size: 27G with a needle length of 90mm excluding hub. Introducer needle length of >35mm with an introducer hub length of <19mm. The effective length of introducer and stylet with cannula is >76mm Needle, stylet and introducers must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 50) | N/A | PROMED TECHNOLOGIES (PTY) LTD | 5.34 | U-LIFE | 502036 | JIAXING | CHINA | 14 DAYS | 85% | USD | 1 BOX | YES | BOX | COMPLY pending confirmation of latex free Accept with confirmation of the introducer length and introducer hub
42142502-00036 | 33 | Needle, spinal, Pencil point bevel; luer lock with introducer. The needle must be colour coded at the stylet cap and the size must be clearly stated, must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. The needle must fit securely into the introducer. The introducer must be streamlined and easy to handle. The introducer size must not reduce the effective length of the cannula. The introducer hub must be transparent to asses for accidental CSF backflow. Needle size: 27G with a needle length of 90mm excluding hub. Introducer needle length of >35mm with an introducer hub length of <19mm. The effective length of introducer and stylet with cannula is >76mm Needle, stylet and introducers must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 50) |  | UNITRADE 1032 CC | 7.94 | SPINOJEC PLUS | 1 | SHANGHAI KINDLY ENTERPRISE DEVELOPMENT GROUP | ROC | 21 | 85 | USD | 500 | YES | 50 | No sample submitted for evaluation
42142502-00036 | 33 | Needle, spinal, Pencil point bevel; luer lock with introducer. The needle must be colour coded at the stylet cap and the size must be clearly stated, must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. The needle must fit securely into the introducer. The introducer must be streamlined and easy to handle. The introducer size must not reduce the effective length of the cannula. The introducer hub must be transparent to asses for accidental CSF backflow. Needle size: 27G with a needle length of 90mm excluding hub. Introducer needle length of >35mm with an introducer hub length of <19mm. The effective length of introducer and stylet with cannula is >76mm Needle, stylet and introducers must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 50) |  | Chemical World cc | 9.09 | MEDIJECT | NDSPPI27 | RUNQIANG | CHINA | 3 | 85 | USD | 1000 | YES | 50 | No brand name. No product code. No introducer
42142502-00036 | 33 | Needle, spinal, Pencil point bevel; luer lock with introducer. The needle must be colour coded at the stylet cap and the size must be clearly stated, must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. The needle must fit securely into the introducer. The introducer must be streamlined and easy to handle. The introducer size must not reduce the effective length of the cannula. The introducer hub must be transparent to asses for accidental CSF backflow. Needle size: 27G with a needle length of 90mm excluding hub. Introducer needle length of >35mm with an introducer hub length of <19mm. The effective length of introducer and stylet with cannula is >76mm Needle, stylet and introducers must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 50) | N/A | SAFARMEX (PTY) LTD | 9.93 | Saclin | To be confirmed | Zhejiang Mediunion Healthcare Group Co. Ltd | China | 3days | 80% | US$ | 10000 Pieces | Yes | Box of 50 | No sample submitted for evaluation
42142502-00036 | 33 | Needle, spinal, Pencil point bevel; luer lock with introducer. The needle must be colour coded at the stylet cap and the size must be clearly stated, must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. The needle must fit securely into the introducer. The introducer must be streamlined and easy to handle. The introducer size must not reduce the effective length of the cannula. The introducer hub must be transparent to asses for accidental CSF backflow. Needle size: 27G with a needle length of 90mm excluding hub. Introducer needle length of >35mm with an introducer hub length of <19mm. The effective length of introducer and stylet with cannula is >76mm Needle, stylet and introducers must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 50) | Basic Line offer | Becton Dickinson | 12 | BD Whitacre Spinal Needle 27G X 90mm with Introducer | 405075 | Becton Dickinson | Fraga, Spain | 1 week | 76% | USD | 1 BOX | YES | 25 Units in a box | COMPLY pending confirmation of latex free
42142502-00036 | 33 | Needle, spinal, Pencil point bevel; luer lock with introducer. The needle must be colour coded at the stylet cap and the size must be clearly stated, must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. The needle must fit securely into the introducer. The introducer must be streamlined and easy to handle. The introducer size must not reduce the effective length of the cannula. The introducer hub must be transparent to asses for accidental CSF backflow. Needle size: 27G with a needle length of 90mm excluding hub. Introducer needle length of >35mm with an introducer hub length of <19mm. The effective length of introducer and stylet with cannula is >76mm Needle, stylet and introducers must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 50) | Basic Line Offer | BIOCOR HOSPITAL SUPPLIES PTY LTD | 12.6 | Medline | N/A | Yangzhong Medline Industry Co.Ltd | China | 60 | 85 | USD | 10 Boxes | Yes | N/A | COMPLY pending confirmation of introducer length
42142502-00036 | 33 | Needle, spinal, Pencil point bevel; luer lock with introducer. The needle must be colour coded at the stylet cap and the size must be clearly stated, must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. The needle must fit securely into the introducer. The introducer must be streamlined and easy to handle. The introducer size must not reduce the effective length of the cannula. The introducer hub must be transparent to asses for accidental CSF backflow. Needle size: 27G with a needle length of 90mm excluding hub. Introducer needle length of >35mm with an introducer hub length of <19mm. The effective length of introducer and stylet with cannula is >76mm Needle, stylet and introducers must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 50) | 1 | GROBIR MEDICAL SUPPLIERS | 13.95 | DURAGAURD | DG.11.027.090 | PROMECON | CHINA | 21 | 80 | US DOLLAR | 25 | Yes | UOM SOLD IN PACKS OF 25 CAN SUPPLY 2 BOXES TO MEET SPECIFICATION REQUIREMENT OF 50 UNITS / MEETS SPECIFICATION | COMPLY pending confirmation of latex free
42142502-00036 | 33 | Needle, spinal, Pencil point bevel; luer lock with introducer. The needle must be colour coded at the stylet cap and the size must be clearly stated, must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. The needle must fit securely into the introducer. The introducer must be streamlined and easy to handle. The introducer size must not reduce the effective length of the cannula. The introducer hub must be transparent to asses for accidental CSF backflow. Needle size: 27G with a needle length of 90mm excluding hub. Introducer needle length of >35mm with an introducer hub length of <19mm. The effective length of introducer and stylet with cannula is >76mm Needle, stylet and introducers must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 50) | Alternative Offer | BIOCOR HOSPITAL SUPPLIES PTY LTD | 14.73 | Technomed | TCN2790in | Techno Med | South Africa | 21 | 0 | N/A | 5 Boxes of 50 | Yes | N/A | COMPLY pending confirmation of introducer length
42142502-00036 | 33 | Needle, spinal, Pencil point bevel; luer lock with introducer. The needle must be colour coded at the stylet cap and the size must be clearly stated, must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. The needle must fit securely into the introducer. The introducer must be streamlined and easy to handle. The introducer size must not reduce the effective length of the cannula. The introducer hub must be transparent to asses for accidental CSF backflow. Needle size: 27G with a needle length of 90mm excluding hub. Introducer needle length of >35mm with an introducer hub length of <19mm. The effective length of introducer and stylet with cannula is >76mm Needle, stylet and introducers must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 50) | BOX OF 50 | MSIMED CC | 15 | KANG KANG | SPN015 | ZHENJIANG | CHINA | 25 | 20% | USD | 20000PCS | YES | BOX | No sample submitted for evaluation
42142502-00036 | 33 | Needle, spinal, Pencil point bevel; luer lock with introducer. The needle must be colour coded at the stylet cap and the size must be clearly stated, must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. The needle must fit securely into the introducer. The introducer must be streamlined and easy to handle. The introducer size must not reduce the effective length of the cannula. The introducer hub must be transparent to asses for accidental CSF backflow. Needle size: 27G with a needle length of 90mm excluding hub. Introducer needle length of >35mm with an introducer hub length of <19mm. The effective length of introducer and stylet with cannula is >76mm Needle, stylet and introducers must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 50) |  | Zanotone (Pty) Ltd | 16.66 | Mekon |  | Mekon | [6 empty] | Each | COMPLY pending confirmation of latex free
42142502-00036 | 33 | Needle, spinal, Pencil point bevel; luer lock with introducer. The needle must be colour coded at the stylet cap and the size must be clearly stated, must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. The needle must fit securely into the introducer. The introducer must be streamlined and easy to handle. The introducer size must not reduce the effective length of the cannula. The introducer hub must be transparent to asses for accidental CSF backflow. Needle size: 27G with a needle length of 90mm excluding hub. Introducer needle length of >35mm with an introducer hub length of <19mm. The effective length of introducer and stylet with cannula is >76mm Needle, stylet and introducers must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 50) | OFFER 1 | SUPRA HEALTHCARE JOHANNESBURG PTY LTD | 26.37 | REVITAL | SPNPP27X90 | Revital Health | KENYA | Initial 60 days, thereafter 14 | 84 | USD | 1 Box | YES |  | COMPLY pending confirmation of latex free
42142502-00036 | 33 | Needle, spinal, Pencil point bevel; luer lock with introducer. The needle must be colour coded at the stylet cap and the size must be clearly stated, must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. The needle must fit securely into the introducer. The introducer must be streamlined and easy to handle. The introducer size must not reduce the effective length of the cannula. The introducer hub must be transparent to asses for accidental CSF backflow. Needle size: 27G with a needle length of 90mm excluding hub. Introducer needle length of >35mm with an introducer hub length of <19mm. The effective length of introducer and stylet with cannula is >76mm Needle, stylet and introducers must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 50) |  | ALMIKA TRADING PTY LTD | 26.45 | BD | 405075 | BD | USA | 14 | 0 | NO | 100 | YES | 1 | COMPLY pending confirmation of introducer length
42142502-00036 | 33 | Needle, spinal, Pencil point bevel; luer lock with introducer. The needle must be colour coded at the stylet cap and the size must be clearly stated, must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. The needle must fit securely into the introducer. The introducer must be streamlined and easy to handle. The introducer size must not reduce the effective length of the cannula. The introducer hub must be transparent to asses for accidental CSF backflow. Needle size: 27G with a needle length of 90mm excluding hub. Introducer needle length of >35mm with an introducer hub length of <19mm. The effective length of introducer and stylet with cannula is >76mm Needle, stylet and introducers must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 50) | OFFER 2 | SUPRA HEALTHCARE JOHANNESBURG PTY LTD | 31.44 | Egemen | SPNPP27X90 | Egemen | TURKEY | Initial 60 days, thereafter 14 | 84 | USD | 1 Box | YES |  | Offered an introducer length of 34mm instead of greater than 36mm
42142502-00036 | 33 | Needle, spinal, Pencil point bevel; luer lock with introducer. The needle must be colour coded at the stylet cap and the size must be clearly stated, must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. The needle must fit securely into the introducer. The introducer must be streamlined and easy to handle. The introducer size must not reduce the effective length of the cannula. The introducer hub must be transparent to asses for accidental CSF backflow. Needle size: 27G with a needle length of 90mm excluding hub. Introducer needle length of >35mm with an introducer hub length of <19mm. The effective length of introducer and stylet with cannula is >76mm Needle, stylet and introducers must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 50) | 1 | B. BRAUN MEDICAL (PTY) LTD | 32 | Pencan 27G x 3 1/2" (0.42 x 88mm) With Guide | 4502051 | B. Braun | Germany | 42 | 50% | Euro | 50 | Yes | Box of 50 | COMPLY pending cpnfirmation of latex free
42142502-00036 | 33 | Needle, spinal, Pencil point bevel; luer lock with introducer. The needle must be colour coded at the stylet cap and the size must be clearly stated, must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. The needle must fit securely into the introducer. The introducer must be streamlined and easy to handle. The introducer size must not reduce the effective length of the cannula. The introducer hub must be transparent to asses for accidental CSF backflow. Needle size: 27G with a needle length of 90mm excluding hub. Introducer needle length of >35mm with an introducer hub length of <19mm. The effective length of introducer and stylet with cannula is >76mm Needle, stylet and introducers must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 50) |  | Viking Critical Care (Pty) Ltd | 216.5 | VYGON | 69F18104 | VYGON | FRANCE | EX. STOCK 30 - 60 DAYS | 80% | EURO | 1 | YES | 1 | COMPLY pending confirmation of latex free
42142502-00036 | 33 | Needle, spinal, Pencil point bevel; luer lock with introducer. The needle must be colour coded at the stylet cap and the size must be clearly stated, must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. The needle must fit securely into the introducer. The introducer must be streamlined and easy to handle. The introducer size must not reduce the effective length of the cannula. The introducer hub must be transparent to asses for accidental CSF backflow. Needle size: 27G with a needle length of 90mm excluding hub. Introducer needle length of >35mm with an introducer hub length of <19mm. The effective length of introducer and stylet with cannula is >76mm Needle, stylet and introducers must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 50) BOX (BOX OF 25) |  | JUMLA MEDICAL SUPPLIES (PTY) LTD | 753.25 | EGEMEN | TSPPK2790 | TIBBI MEDICAL | TURKEY | 21 DAYS | 70% | US DOLLAR | 2 BOXES | YES | NIL | No sample submitted for evaluation
42142502-00036 | 33 | Needle, spinal, Pencil point bevel; luer lock with introducer. The needle must be colour coded at the stylet cap and the size must be clearly stated, must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. The needle must fit securely into the introducer. The introducer must be streamlined and easy to handle. The introducer size must not reduce the effective length of the cannula. The introducer hub must be transparent to asses for accidental CSF backflow. Needle size: 27G with a needle length of 90mm excluding hub. Introducer needle length of >35mm with an introducer hub length of <19mm. The effective length of introducer and stylet with cannula is >76mm Needle, stylet and introducers must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 50) |  | BAS MEDEXPRESS | 776.25 | Mercury | SN2790P |  | India | 1 - 3 days | 15% | N/A | As per client | Yes | Box | No sample submitted for evaluation
42142502-00036 | 33 | Needle, spinal, Pencil point bevel; luer lock with introducer. The needle must be colour coded at the stylet cap and the size must be clearly stated, must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. The needle must fit securely into the introducer. The introducer must be streamlined and easy to handle. The introducer size must not reduce the effective length of the cannula. The introducer hub must be transparent to asses for accidental CSF backflow. Needle size: 27G with a needle length of 90mm excluding hub. Introducer needle length of >35mm with an introducer hub length of <19mm. The effective length of introducer and stylet with cannula is >76mm Needle, stylet and introducers must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 50) |  | LECHOBA MEDICAL TECHNOLOGIES (PTY) LTD | 5183 | UNISIS UNIEVER | 5311371 | UNISIS CORPORATION | JAPAN | 21 DAYS | 75% | USD | 50 | YES | 25/BOX | No sample submitted for evaluation
42142502-00037 | 34 | Needle, spinal, Pencil point bevel; luer lock with introducer. The needle must be colour coded at the stylet cap and the size must be clearly stated, must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. The needle must fit securely into the introducer. The introducer must be streamlined and easy to handle. The introducer size must not reduce the effective length of the cannula. The introducer hub must be transparent to asses for accidental CSF backflow. Needle size: 27G extra-length with a needle length of >110 to 120mm excluding hub. Introducer needle length of >35mm with an introducer hub length of <19mm. The effective length of introducer and stylet with cannula is ≥110mm Needle, stylet and introducers must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 50) |  | LECHOBA MEDICAL TECHNOLOGIES (PTY) LTD | 2.24 | UNISIS UNIEVER | 5410061 | UNISIS CORPORATION | JAPAN | 21 DAYS | 75% | USD | 50 | YES | Each (PACKED 25/BOX) | No sample submitted for evaluation
42142502-00037 | 34 | Needle, spinal, Pencil point bevel; luer lock with introducer. The needle must be colour coded at the stylet cap and the size must be clearly stated, must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. The needle must fit securely into the introducer. The introducer must be streamlined and easy to handle. The introducer size must not reduce the effective length of the cannula. The introducer hub must be transparent to asses for accidental CSF backflow. Needle size: 27G extra-length with a needle length of >110 to 120mm excluding hub. Introducer needle length of >35mm with an introducer hub length of <19mm. The effective length of introducer and stylet with cannula is ≥110mm Needle, stylet and introducers must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 50) |  | Medi-Core Technologies | 264.55 | ULIFE | 42142502-00037 | JIAXING U-LIFE MEDICAL DEVICE TECHNOLOGLY.,CO LTD | CHINA | 21 | 80 | USD | 1 box | YES | BOX OF 50 | COMPLY pending confirmation of introducer length |  | Recommended | 100
42142502-00037 | 34 | Needle, spinal, Pencil point bevel; luer lock with introducer. The needle must be colour coded at the stylet cap and the size must be clearly stated, must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. The needle must fit securely into the introducer. The introducer must be streamlined and easy to handle. The introducer size must not reduce the effective length of the cannula. The introducer hub must be transparent to asses for accidental CSF backflow. Needle size: 27G extra-length with a needle length of >110 to 120mm excluding hub. Introducer needle length of >35mm with an introducer hub length of <19mm. The effective length of introducer and stylet with cannula is ≥110mm Needle, stylet and introducers must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 50) | N/A | PROMED TECHNOLOGIES (PTY) LTD | 5.34 | U-LIFE | 502037 | JIAXING | CHINA | 14 DAYS | 85% | USD | 1 BOX | YES | BOX | COMPLY pending confirmation of latex free Accept with confirmation of the introducer length and introducer hub
42142502-00037 | 34 | Needle, spinal, Pencil point bevel; luer lock with introducer. The needle must be colour coded at the stylet cap and the size must be clearly stated, must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. The needle must fit securely into the introducer. The introducer must be streamlined and easy to handle. The introducer size must not reduce the effective length of the cannula. The introducer hub must be transparent to asses for accidental CSF backflow. Needle size: 27G extra-length with a needle length of >110 to 120mm excluding hub. Introducer needle length of >35mm with an introducer hub length of <19mm. The effective length of introducer and stylet with cannula is ≥110mm Needle, stylet and introducers must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 50) | N/A | SAFARMEX (PTY) LTD | 10.76 | Saclin | To be confirmed | Zhejiang Mediunion Healthcare Group Co. Ltd | China | 3days | 80% | US$ | 10000 Pieces | Yes | Box of 50 | No sample submitted for evaluation
42142502-00037 | 34 | Needle, spinal, Pencil point bevel; luer lock with introducer. The needle must be colour coded at the stylet cap and the size must be clearly stated, must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. The needle must fit securely into the introducer. The introducer must be streamlined and easy to handle. The introducer size must not reduce the effective length of the cannula. The introducer hub must be transparent to asses for accidental CSF backflow. Needle size: 27G extra-length with a needle length of >110 to 120mm excluding hub. Introducer needle length of >35mm with an introducer hub length of <19mm. The effective length of introducer and stylet with cannula is ≥110mm Needle, stylet and introducers must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 50) |  | Chemical World cc | 13.23 | MEDIJECT | NDSPPI27 | RUNQIANG | CHINA | 3 | 85 | USD | 1000 | YES | 50 | No sample submitted for evaluation.
42142502-00037 | 34 | Needle, spinal, Pencil point bevel; luer lock with introducer. The needle must be colour coded at the stylet cap and the size must be clearly stated, must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. The needle must fit securely into the introducer. The introducer must be streamlined and easy to handle. The introducer size must not reduce the effective length of the cannula. The introducer hub must be transparent to asses for accidental CSF backflow. Needle size: 27G extra-length with a needle length of >110 to 120mm excluding hub. Introducer needle length of >35mm with an introducer hub length of <19mm. The effective length of introducer and stylet with cannula is ≥110mm Needle, stylet and introducers must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 50) | Basic Line Offer | BIOCOR HOSPITAL SUPPLIES PTY LTD | 13.81 | Medline | N/A | Yangzhong Medline Industry Co.Ltd | China | 60 | 85 | USD | 10 Boxes of 50 | Yes | N/A | No sample submitted for evaluation
42142502-00037 | 34 | Needle, spinal, Pencil point bevel; luer lock with introducer. The needle must be colour coded at the stylet cap and the size must be clearly stated, must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. The needle must fit securely into the introducer. The introducer must be streamlined and easy to handle. The introducer size must not reduce the effective length of the cannula. The introducer hub must be transparent to asses for accidental CSF backflow. Needle size: 27G extra-length with a needle length of >110 to 120mm excluding hub. Introducer needle length of >35mm with an introducer hub length of <19mm. The effective length of introducer and stylet with cannula is ≥110mm Needle, stylet and introducers must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 50) | EACH | MSIMED CC | 15 | KANG KANG | SPN016 | ZHENJIANG | CHINA | 25 | 20% | USD | 20000PCS | YES | EACH | No sample submitted for evaluation
42142502-00037 | 34 | Needle, spinal, Pencil point bevel; luer lock with introducer. The needle must be colour coded at the stylet cap and the size must be clearly stated, must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. The needle must fit securely into the introducer. The introducer must be streamlined and easy to handle. The introducer size must not reduce the effective length of the cannula. The introducer hub must be transparent to asses for accidental CSF backflow. Needle size: 27G extra-length with a needle length of >110 to 120mm excluding hub. Introducer needle length of >35mm with an introducer hub length of <19mm. The effective length of introducer and stylet with cannula is ≥110mm Needle, stylet and introducers must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 50) | 1 | GROBIR MEDICAL SUPPLIERS | 15.84 | DURAGAURD | DG.11.027.110 | PROMECON | CHINA | 21 | 80 | US DOLLAR | 25 | Yes | UOM SOLD IN PACKS OF 25 CAN SUPPLY 2 BOXES TO MEET SPECIFICATION REQUIREMENT OF 50 UNITS / MEETS SPECIFICATION | COMPLY pending confirmation of latex free
42142502-00037 | 34 | Needle, spinal, Pencil point bevel; luer lock with introducer. The needle must be colour coded at the stylet cap and the size must be clearly stated, must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. The needle must fit securely into the introducer. The introducer must be streamlined and easy to handle. The introducer size must not reduce the effective length of the cannula. The introducer hub must be transparent to asses for accidental CSF backflow. Needle size: 27G extra-length with a needle length of >110 to 120mm excluding hub. Introducer needle length of >35mm with an introducer hub length of <19mm. The effective length of introducer and stylet with cannula is ≥110mm Needle, stylet and introducers must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 50) | Alternative Offer | BIOCOR HOSPITAL SUPPLIES PTY LTD | 15.97 | Technomed | TCN27120in | Techno Med | South Africa | 21 | 0 | N/A | 5 Boxes of 50 | Yes | N/A | No sample submitted for evaluation
42142502-00037 | 34 | Needle, spinal, Pencil point bevel; luer lock with introducer. The needle must be colour coded at the stylet cap and the size must be clearly stated, must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. The needle must fit securely into the introducer. The introducer must be streamlined and easy to handle. The introducer size must not reduce the effective length of the cannula. The introducer hub must be transparent to asses for accidental CSF backflow. Needle size: 27G extra-length with a needle length of >110 to 120mm excluding hub. Introducer needle length of >35mm with an introducer hub length of <19mm. The effective length of introducer and stylet with cannula is ≥110mm Needle, stylet and introducers must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 50) |  | Zanotone (Pty) Ltd | 17.22 | Mekon |  | Mekon | [6 empty] | Each | No sample submitted for evaluation
42142502-00037 | 34 | Needle, spinal, Pencil point bevel; luer lock with introducer. The needle must be colour coded at the stylet cap and the size must be clearly stated, must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. The needle must fit securely into the introducer. The introducer must be streamlined and easy to handle. The introducer size must not reduce the effective length of the cannula. The introducer hub must be transparent to asses for accidental CSF backflow. Needle size: 27G extra-length with a needle length of >110 to 120mm excluding hub. Introducer needle length of >35mm with an introducer hub length of <19mm. The effective length of introducer and stylet with cannula is ≥110mm Needle, stylet and introducers must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 50) | Basic Line offer | Becton Dickinson | 17.5 | BD Whitacre Spinal Needle 27G X 103mm with Introducer | 405113 | Becton Dickinson | Fraga, Spain | 1 week | 76% | USD | 1 BOX | YES | 25 Units in a box | Offered a needle length of 103mm instead of a needle length of greater than 110 to 120mm
42142502-00037 | 34 | Needle, spinal, Pencil point bevel; luer lock with introducer. The needle must be colour coded at the stylet cap and the size must be clearly stated, must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. The needle must fit securely into the introducer. The introducer must be streamlined and easy to handle. The introducer size must not reduce the effective length of the cannula. The introducer hub must be transparent to asses for accidental CSF backflow. Needle size: 27G extra-length with a needle length of >110 to 120mm excluding hub. Introducer needle length of >35mm with an introducer hub length of <19mm. The effective length of introducer and stylet with cannula is ≥110mm Needle, stylet and introducers must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 50) | OFFER 1 | SUPRA HEALTHCARE JOHANNESBURG PTY LTD | 26.37 | REVITAL | SPNPP27X110 | Revital Health | KENYA | Initial 60 days, thereafter 14 | 84 | USD | 1 Box | YES |  | No sample submitted for evaluation
42142502-00037 | 34 | Needle, spinal, Pencil point bevel; luer lock with introducer. The needle must be colour coded at the stylet cap and the size must be clearly stated, must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. The needle must fit securely into the introducer. The introducer must be streamlined and easy to handle. The introducer size must not reduce the effective length of the cannula. The introducer hub must be transparent to asses for accidental CSF backflow. Needle size: 27G extra-length with a needle length of >110 to 120mm excluding hub. Introducer needle length of >35mm with an introducer hub length of <19mm. The effective length of introducer and stylet with cannula is ≥110mm Needle, stylet and introducers must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 50) | OFFER 2 | SUPRA HEALTHCARE JOHANNESBURG PTY LTD | 37.86 | Egemen | SPNPP27X120 | Egemen | TURKEY | Initial 60 days, thereafter 14 | 84 | USD | 1 Box | YES |  | Offered an introducer length of 34mm instead of greater than 36mm
42142502-00037 | 34 | Needle, spinal, Pencil point bevel; luer lock with introducer. The needle must be colour coded at the stylet cap and the size must be clearly stated, must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. The needle must fit securely into the introducer. The introducer must be streamlined and easy to handle. The introducer size must not reduce the effective length of the cannula. The introducer hub must be transparent to asses for accidental CSF backflow. Needle size: 27G extra-length with a needle length of >110 to 120mm excluding hub. Introducer needle length of >35mm with an introducer hub length of <19mm. The effective length of introducer and stylet with cannula is ≥110mm Needle, stylet and introducers must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 50) |  | ALMIKA TRADING PTY LTD | 39.68 | BD | 405113 | BD | USA | 14 | 0 | NO | 100 | YES | 1 | offered a needle length 103mm instead of 110mm – 120mm
42142502-00037 | 34 | Needle, spinal, Pencil point bevel; luer lock with introducer. The needle must be colour coded at the stylet cap and the size must be clearly stated, must have a clear/ transparent needle hub to allow for easy identification of the spinal fluid flashback. The flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and hub and must be clearly indicate the bevelled orientation. Have a stable stylet that does not shake or quiver to allow for easy re-insertion into the cannula. Must align preciously with the stylet point. The needle must fit securely into the introducer. The introducer must be streamlined and easy to handle. The introducer size must not reduce the effective length of the cannula. The introducer hub must be transparent to asses for accidental CSF backflow. Needle size: 27G extra-length with a needle length of >110 to 120mm excluding hub. Introducer needle length of >35mm with an introducer hub length of <19mm. The effective length of introducer and stylet with cannula is ≥110mm Needle, stylet and introducers must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size: Box (Box of 50) BOX (BOX OF 25) |  | JUMLA MEDICAL SUPPLIES (PTY) LTD | 955.65 | EGEMEN | TSPPK27120 | TIBBI MEDICAL | TURKEY | 21 DAYS | 70% | US DOLLAR | 2 BOXES | YES | NIL | COMPLY pending confirmation of introducer length
42142502-00000 | 35 | Needle, epidural, caudal: paediatric luer lock, keyed stylet with matched bevelled point Size: 22G (0.7mm) x 50mm Sterile, Single use, Individually peel packed. EACH | N/A | SAFARMEX (PTY) LTD | 5.04 | Saclin | To be confirmed | Zhejiang Mediunion Healthcare Group Co. Ltd | China | 3days | 80% | US$ |  | Yes | Box of 25 | No sample submitted for evaluation |  | Re-tender
42142502-00000 | 35 | Needle, epidural, caudal: paediatric luer lock, keyed stylet with matched bevelled point Size: 22G (0.7mm) x 50mm Sterile, Single use, Individually peel packed. EACH |  | Zanotone (Pty) Ltd | 11.11 | Mekon |  | Mekon | [6 empty] | Each | No sample submitted for evaluation |  | Re-tender
42142502-00001 | 36 | Needle, epidural, caudal: paediatric luer lock, keyed stylet with matched bevelled point Size: 25G (0.5mm) x 50mm Sterile, Single use, Individually peel packed. EACH | N/A | SAFARMEX (PTY) LTD | 5.04 | Saclin | To be confirmed | Zhejiang Mediunion Healthcare Group Co. Ltd | China | 3days | 80% | US$ |  | Yes | Box of 25 | No sample submitted for evaluation |  | Re-tender
42142502-00001 | 36 | Needle, epidural, caudal: paediatric luer lock, keyed stylet with matched bevelled point Size: 25G (0.5mm) x 50mm Sterile, Single use, Individually peel packed. EACH |  | Zanotone (Pty) Ltd | 11.11 | Mekon |  | Mekon | [6 empty] | Each | No sample submitted for evaluation |  | Re-tender
42142502-00001 | 36 | Needle, epidural, caudal: paediatric luer lock, keyed stylet with matched bevelled point Size: 25G (0.5mm) x 50mm Sterile, Single use, Individually peel packed. EACH BOX (BOX OF 25) |  | JUMLA MEDICAL SUPPLIES (PTY) LTD | 3199.02 | CAUDAL NDL | TKAU22 | TIBBI MEDICAL | TURKEY | 21 DAYS | 70% | US DOLLAR | 1 BOX | YES | NIL | No sample submitted for evaluation |  | Re-tender
42142502-00002 | 37 | Needle, epidural, caudal: paediatric luer lock, keyed stylet with matched bevelled point Size: 22G (0.7mm) x 35mm Sterile, Single use, Individually peel packed. EACH | N/A | PROMED TECHNOLOGIES (PTY) LTD | 4.2 | U-LIFE | 502002 | JIAXING | CHINA | 14 DAYS | 85% | USD | 1 BOX | YES | BOX | No samples submitted for evaluation
42142502-00002 | 37 | Needle, epidural, caudal: paediatric luer lock, keyed stylet with matched bevelled point Size: 22G (0.7mm) x 35mm Sterile, Single use, Individually peel packed. EACH | N/A | SAFARMEX (PTY) LTD | 6.39 | Saclin | To be confirmed | Zhejiang Mediunion Healthcare Group Co. Ltd | China | 3days | 80% | US$ |  | Yes | Box of 25 | No sample submitted for evaluation
42142502-00002 | 37 | Needle, epidural, caudal: paediatric luer lock, keyed stylet with matched bevelled point Size: 22G (0.7mm) x 35mm Sterile, Single use, Individually peel packed. EACH |  | Zanotone (Pty) Ltd | 11.11 | Mekon |  | Mekon | [6 empty] | Each | No sample submitted for evaluation
42142502-00002 | 37 | Needle, epidural, caudal: paediatric luer lock, keyed stylet with matched bevelled point Size: 22G (0.7mm) x 35mm Sterile, Single use, Individually peel packed. EACH |  | Medi-Core Technologies | 15.57 | ULIFE | 42142502-00002 | JIAXING U-LIFE MEDICAL DEVICE TECHNOLOGY.,CO LTD | CHINA | 21 | 80 | USD | 1 | YES | EACH | No sample submitted for evaluation
42142502-00002 | 37 | Needle, epidural, caudal: paediatric luer lock, keyed stylet with matched bevelled point Size: 22G (0.7mm) x 35mm Sterile, Single use, Individually peel packed. EACH | 1 | B. BRAUN MEDICAL (PTY) LTD | 35 | Epican Paed Caudal 22G x 1 3/8" (0.73 x 35mm) | 4502418 | B. Braun | Germany | 42 | 50% | Euro | Each | Yes | Each | COMPLY pending cpnfirmation of latex free |  | Recommended | 95
42142502-00002 | 37 | Needle, epidural, caudal: paediatric luer lock, keyed stylet with matched bevelled point Size: 22G (0.7mm) x 35mm Sterile, Single use, Individually peel packed. EACH |  | SUPRA HEALTHCARE JOHANNESBURG PTY LTD | 157.03 | Egemen | SPNQ22X35 | EGEMEN | TURKEY | Initial 60 days, thereafter 14 | 84 | USD | 10 Units | YES | 22x35 | COMPLY pending confirmation of latex free
42142502-00003 | 38 | Needle, epidural, caudal: paediatric luer lock, keyed stylet with matched bevelled point Size: 25G (0.5mm) x 30mm Sterile, Single use, Individually peel packed. EACH | N/A | SAFARMEX (PTY) LTD | 6.39 | Saclin | To be confirmed | Zhejiang Mediunion Healthcare Group Co. Ltd | China | 3days | 80% | US$ |  | Yes | Box of 25 | No sample submitted for evaluation
42142502-00003 | 38 | Needle, epidural, caudal: paediatric luer lock, keyed stylet with matched bevelled point Size: 25G (0.5mm) x 30mm Sterile, Single use, Individually peel packed. EACH |  | Zanotone (Pty) Ltd | 11.11 | Mekon |  | Mekon | [6 empty] | Each | No sample submitted for evaluation
42142502-00003 | 38 | Needle, epidural, caudal: paediatric luer lock, keyed stylet with matched bevelled point Size: 25G (0.5mm) x 30mm Sterile, Single use, Individually peel packed. EACH | 1 | B. BRAUN MEDICAL (PTY) LTD | 35 | Epican Paed Caudal 25G x 1 1/8" (0.53 x 30mm) | 4502400 | B. Braun | Germany | 42 | 50% | Euro | Each | Yes | Each | COMPLY pending cpnfirmation of latex free |  | Recommended | 95
42142502-00003 | 38 | Needle, epidural, caudal: paediatric luer lock, keyed stylet with matched bevelled point Size: 25G (0.5mm) x 30mm Sterile, Single use, Individually peel packed. EACH |  | SUPRA HEALTHCARE JOHANNESBURG PTY LTD | 157.03 | Egemen | SPNQ25X50 | EGEMEN | TURKEY | Initial 60 days, thereafter 14 | 84 | USD | 10 Units | YES | 25x40 | COMPLY pending confirmation of latex free
42142502-00003 | 38 | Needle, epidural, caudal: paediatric luer lock, keyed stylet with matched bevelled point Size: 25G (0.5mm) x 30mm Sterile, Single use, Individually peel packed. EACH BOX (BOX OF 25) |  | JUMLA MEDICAL SUPPLIES (PTY) LTD | 3199.02 | CAUDAL NDL | TKAU25 | TIBBI MEDICAL | TURKEY | 21 DAYS | 70% | US DOLLAR | 1 BOX | YES | NIL | COMPLY pending confirmation of latex free
42142502-00058 | 39 | Needle, epidural, paediatric, Tuohy bevel luer lock keyed stylet with matched bevelled point. The set must include: A Tuohy needle that: Is graduated in 1cm markings, has a precision fit stylet to prevent tissue coring and is marked to indicate bevel orientation and has a smooth inner bevel to prevent catheter shearing, clear ridged hub to aid with bevel positioning and a comfortable solid grip flange. Needle Size: 18G with a needle length of > 45 mm, < 50mm A loss of resistance syringe of 10ml with low friction An epidural catheter: Is kink resistant with good tensile strength, must have an atraumatic close ended tip with multiple side orifices, must be graduated with standard 1cm catheter markings from 5cm to 15cm to facilitate accurate catheter positioning and there must be double line at 10cm triple lines at 15cm and quadruple lines at 20cm. Epidural catheter size: 20G Catheter tip/catheter must be radio-opaque. A 20-micron filter that is flat with luer lock connectors. A securing device that connects the catheter securely to the filter – without risk of disconnection, compression of catheter or leaks. Needles, stylets and introducers must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free. Must be sterile and individually peel packed Pack size: Each |  | Medi-Core Technologies | 36.22 | ULIFE | 42142502-00058 | JIAXING U-LIFE MEDICAL DEVICE TECHNOLOGY CO.,LTD | CHINA | 21 | 80 | USD | 1 | YES | EACH | No sample submitted for evaluation |  | Re-tender
42142502-00058 | 39 | Needle, epidural, paediatric, Tuohy bevel luer lock keyed stylet with matched bevelled point. The set must include: A Tuohy needle that: Is graduated in 1cm markings, has a precision fit stylet to prevent tissue coring and is marked to indicate bevel orientation and has a smooth inner bevel to prevent catheter shearing, clear ridged hub to aid with bevel positioning and a comfortable solid grip flange. Needle Size: 18G with a needle length of > 45 mm, < 50mm A loss of resistance syringe of 10ml with low friction An epidural catheter: Is kink resistant with good tensile strength, must have an atraumatic close ended tip with multiple side orifices, must be graduated with standard 1cm catheter markings from 5cm to 15cm to facilitate accurate catheter positioning and there must be double line at 10cm triple lines at 15cm and quadruple lines at 20cm. Epidural catheter size: 20G Catheter tip/catheter must be radio-opaque. A 20-micron filter that is flat with luer lock connectors. A securing device that connects the catheter securely to the filter – without risk of disconnection, compression of catheter or leaks. Needles, stylets and introducers must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free. Must be sterile and individually peel packed Pack size: Each |  | SUPRA HEALTHCARE JOHANNESBURG PTY LTD | 224.84 | Portex | PORT18X45 | Babcock | United Kingdom | Initial 60 days, thereafter 14 | 84 | ZAR | 1 Pack | YES |  | Epidural catheter size: 18G instead of 20G Graduation markings start from zero instead of 5cm |  | Re-tender
42142502-00038 | 40 | Needle, epidural, paediatric, Tuohy bevel luer lock keyed stylet with matched bevelled point. The set must include: A Tuohy needle that: Is graduated in 1cm markings, has a precision fit stylet to prevent tissue coring and is marked to indicate bevel orientation and has a smooth inner bevel to prevent catheter shearing, clear ridged hub to aid with bevel positioning and a comfortable solid grip flange. Needle Size: 18G with a needle length of > 45 mm, < 50mm A loss of resistance syringe of 10ml with low friction An epidural catheter: Is kink resistant with good tensile strength, must have an atraumatic close ended tip with multiple side orifices, must be graduated with standard 1cm catheter markings from 5cm to 15cm to facilitate accurate catheter positioning and there must be double line at 10cm triple lines at 15cm and quadruple lines at 20cm. Epidural catheter size: 21G Catheter tip/catheter must be radio-opaque. A 20-micron filter that is flat with luer lock connectors. A securing device that connects the catheter securely to the filter – without risk of disconnection, compression of catheter or leaks. Needles, stylets and introducers must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free. Must be sterile and individually peel packed Pack size: Each |  | Medi-Core Technologies | 36.22 | ulife | 42142502-00038 | JIAXING U-LIFE MEDICAL DEVICE TECHNOLOGY CO.,LTD | CHINA | 21 | 80 | USD | 1 | YES | EACH | No sample submitted for evaluation |  | Re-tender
42142502-00038 | 40 | Needle, epidural, paediatric, Tuohy bevel luer lock keyed stylet with matched bevelled point. The set must include: A Tuohy needle that: Is graduated in 1cm markings, has a precision fit stylet to prevent tissue coring and is marked to indicate bevel orientation and has a smooth inner bevel to prevent catheter shearing, clear ridged hub to aid with bevel positioning and a comfortable solid grip flange. Needle Size: 18G with a needle length of > 45 mm, < 50mm A loss of resistance syringe of 10ml with low friction An epidural catheter: Is kink resistant with good tensile strength, must have an atraumatic close ended tip with multiple side orifices, must be graduated with standard 1cm catheter markings from 5cm to 15cm to facilitate accurate catheter positioning and there must be double line at 10cm triple lines at 15cm and quadruple lines at 20cm. Epidural catheter size: 21G Catheter tip/catheter must be radio-opaque. A 20-micron filter that is flat with luer lock connectors. A securing device that connects the catheter securely to the filter – without risk of disconnection, compression of catheter or leaks. Needles, stylets and introducers must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free. Must be sterile and individually peel packed Pack size: Each |  | SUPRA HEALTHCARE JOHANNESBURG PTY LTD | 224.84 | Portex | PORT18X45 | Babcock | United Kingdom | Initial 60 days, thereafter 14 | 84 | ZAR | 1 Pack | YES |  | Epidural catheter size: 18G instead of 21G Graduation markings start from zero instead of 5cm |  | Re-tender
42142502-00039 | 41 | Needle, epidural, paediatric, Tuohy bevel luer lock keyed stylet with matched bevelled point. The set must include: A Tuohy needle that: Is graduated in 1cm markings, has a precision fit stylet to prevent tissue coring and is marked to indicate bevel orientation and has a smooth inner bevel to prevent catheter shearing, clear ridged hub to aid with bevel positioning and a comfortable solid grip flange. Needle Size: 20G with a needle length of < 50mm A loss of resistance syringe of 10ml with low friction An epidural catheter: Is kink resistant with good tensile strength, must have an atraumatic close ended tip with multiple side orifices, must be graduated with standard 1cm catheter markings from 5cm to 15cm to facilitate accurate catheter positioning and there must be double line at 10cm triple lines at 15cm and quadruple lines at 20cm. Epidural catheter size: 24G Catheter tip/catheter must be radio-opaque. A 20-micron filter that is flat with luer lock connectors. A securing device that connects the catheter securely to the filter – without risk of disconnection, compression of catheter or leaks. Needles, stylets and introducers must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free. Must be sterile and individually peel packed Pack size: Each |  | Medi-Core Technologies | 36.22 | U-LIFE | 42142502-00039 | JIAXING U-LIFE MEDICAL DEVICE TECHNOLOGY CO.,LTD | CHINA | 21 | 80 | USD | 1 | YES | EACH | No sample submitted for evaluation |  | Re-tender
42142502-00039 | 41 | Needle, epidural, paediatric, Tuohy bevel luer lock keyed stylet with matched bevelled point. The set must include: A Tuohy needle that: Is graduated in 1cm markings, has a precision fit stylet to prevent tissue coring and is marked to indicate bevel orientation and has a smooth inner bevel to prevent catheter shearing, clear ridged hub to aid with bevel positioning and a comfortable solid grip flange. Needle Size: 20G with a needle length of < 50mm A loss of resistance syringe of 10ml with low friction An epidural catheter: Is kink resistant with good tensile strength, must have an atraumatic close ended tip with multiple side orifices, must be graduated with standard 1cm catheter markings from 5cm to 15cm to facilitate accurate catheter positioning and there must be double line at 10cm triple lines at 15cm and quadruple lines at 20cm. Epidural catheter size: 24G Catheter tip/catheter must be radio-opaque. A 20-micron filter that is flat with luer lock connectors. A securing device that connects the catheter securely to the filter – without risk of disconnection, compression of catheter or leaks. Needles, stylets and introducers must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free. Must be sterile and individually peel packed Pack size: Each |  | SUPRA HEALTHCARE JOHANNESBURG PTY LTD | 224.84 | Portex | PORT20X50 | Babcock | United Kingdom | Initial 60 days, thereafter 14 | 84 | ZAR | 1 Pack | YES |  | Package labelling indicates a needle size is 19G instead of 20G Epidural catheter size:19G instead of 24G Graduation markings start from zero instead of 5cm |  | Re-tender
42142502-00039 | 41 | Needle, epidural, paediatric, Tuohy bevel luer lock keyed stylet with matched bevelled point. The set must include: A Tuohy needle that: Is graduated in 1cm markings, has a precision fit stylet to prevent tissue coring and is marked to indicate bevel orientation and has a smooth inner bevel to prevent catheter shearing, clear ridged hub to aid with bevel positioning and a comfortable solid grip flange. Needle Size: 20G with a needle length of < 50mm A loss of resistance syringe of 10ml with low friction An epidural catheter: Is kink resistant with good tensile strength, must have an atraumatic close ended tip with multiple side orifices, must be graduated with standard 1cm catheter markings from 5cm to 15cm to facilitate accurate catheter positioning and there must be double line at 10cm triple lines at 15cm and quadruple lines at 20cm. Epidural catheter size: 24G Catheter tip/catheter must be radio-opaque. A 20-micron filter that is flat with luer lock connectors. A securing device that connects the catheter securely to the filter – without risk of disconnection, compression of catheter or leaks. Needles, stylets and introducers must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free. Must be sterile and individually peel packed Pack size: Each | 1 | B. BRAUN MEDICAL (PTY) LTD | 285 | Perifix One Paed Set 20 | 4512014C | B. Braun | Germany | 42 | 50% | Euro | Each | Yes | Each | Re-tender: End-users to revise specification |  | Re-tender
42142502-00040 | 42 | Needle, epidural, adult, Tuohy bevel luer lock keyed stylet with matched bevelled point. The set must include: A Tuohy needle that: Is graduated in 1cm markings, has a precision fit stylet to prevent tissue coring and is marked to indicate bevel orientation and has a smooth inner bevel to prevent catheter shearing, clear ridged hub to aid with bevel positioning and a comfortable solid grip flange. Needle Size: 16G with a needle length of 80 to 90mm. A loss of resistance syringe of 10ml with low friction An epidural catheter: Is kink resistant with good tensile strength, must have an atraumatic close ended tip with multiple side orifices, must be graduated with standard 1cm catheter markings from 5cm to 15cm to facilitate accurate catheter positioning and there must be double line at 10cm triple lines at 15cm and quadruple lines at 20cm. Epidural catheter size: 18G or 19G Catheter tip/catheter must be radio-opaque. A 20-micron filter that is flat with luer lock connectors. A securing device that connects the catheter securely to the filter – without risk of disconnection, compression of catheter or leaks. Needles, stylets and introducers must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free. Must be sterile and individually peel packed Pack size: Each |  | Medi-Core Technologies | 36.22 | U-LIFE | 42142502-00040 | JIAXING U-LIFE MEDICAL DEVICE TECHNOLOGY CO.,LTD | CHINA | 21 | 80 | USD | 1 | YES | EACH | The markings of the sample are graduated from zero instead from 5cm to 15cm the size of the epidural catherter is not stated the needle size is not indicated
42142502-00040 | 42 | Needle, epidural, adult, Tuohy bevel luer lock keyed stylet with matched bevelled point. The set must include: A Tuohy needle that: Is graduated in 1cm markings, has a precision fit stylet to prevent tissue coring and is marked to indicate bevel orientation and has a smooth inner bevel to prevent catheter shearing, clear ridged hub to aid with bevel positioning and a comfortable solid grip flange. Needle Size: 16G with a needle length of 80 to 90mm. A loss of resistance syringe of 10ml with low friction An epidural catheter: Is kink resistant with good tensile strength, must have an atraumatic close ended tip with multiple side orifices, must be graduated with standard 1cm catheter markings from 5cm to 15cm to facilitate accurate catheter positioning and there must be double line at 10cm triple lines at 15cm and quadruple lines at 20cm. Epidural catheter size: 18G or 19G Catheter tip/catheter must be radio-opaque. A 20-micron filter that is flat with luer lock connectors. A securing device that connects the catheter securely to the filter – without risk of disconnection, compression of catheter or leaks. Needles, stylets and introducers must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free. Must be sterile and individually peel packed Pack size: Each |  | CRYSATL PIER TRADING 148cc | 133.7 | EPIDURAL MINI PACK/SR CARE |  | GLOBAL MEDIKIT LIMITED C/O SR MEDICAL CARE | [6 empty] | Each | The micron filter is not stated. Epidural catheter size: 16G instead of 18/19G The markings starts from the tip of the catheter instead of from 5 cm.
42142502-00040 | 42 | Needle, epidural, adult, Tuohy bevel luer lock keyed stylet with matched bevelled point. The set must include: A Tuohy needle that: Is graduated in 1cm markings, has a precision fit stylet to prevent tissue coring and is marked to indicate bevel orientation and has a smooth inner bevel to prevent catheter shearing, clear ridged hub to aid with bevel positioning and a comfortable solid grip flange. Needle Size: 16G with a needle length of 80 to 90mm. A loss of resistance syringe of 10ml with low friction An epidural catheter: Is kink resistant with good tensile strength, must have an atraumatic close ended tip with multiple side orifices, must be graduated with standard 1cm catheter markings from 5cm to 15cm to facilitate accurate catheter positioning and there must be double line at 10cm triple lines at 15cm and quadruple lines at 20cm. Epidural catheter size: 18G or 19G Catheter tip/catheter must be radio-opaque. A 20-micron filter that is flat with luer lock connectors. A securing device that connects the catheter securely to the filter – without risk of disconnection, compression of catheter or leaks. Needles, stylets and introducers must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free. Must be sterile and individually peel packed Pack size: Each | OFFER 1 | SUPRA HEALTHCARE JOHANNESBURG PTY LTD | 172.3 | Epifix | EPI16X80 | EGEMEN | TURKEY | Initial 60 days, thereafter 14 | 84 | USD | 1 Box | YES |  | Epidural catheter size: 20G instead of 18G/19G
42142502-00040 | 42 | Needle, epidural, adult, Tuohy bevel luer lock keyed stylet with matched bevelled point. The set must include: A Tuohy needle that: Is graduated in 1cm markings, has a precision fit stylet to prevent tissue coring and is marked to indicate bevel orientation and has a smooth inner bevel to prevent catheter shearing, clear ridged hub to aid with bevel positioning and a comfortable solid grip flange. Needle Size: 16G with a needle length of 80 to 90mm. A loss of resistance syringe of 10ml with low friction An epidural catheter: Is kink resistant with good tensile strength, must have an atraumatic close ended tip with multiple side orifices, must be graduated with standard 1cm catheter markings from 5cm to 15cm to facilitate accurate catheter positioning and there must be double line at 10cm triple lines at 15cm and quadruple lines at 20cm. Epidural catheter size: 18G or 19G Catheter tip/catheter must be radio-opaque. A 20-micron filter that is flat with luer lock connectors. A securing device that connects the catheter securely to the filter – without risk of disconnection, compression of catheter or leaks. Needles, stylets and introducers must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free. Must be sterile and individually peel packed Pack size: Each | 1 | B. BRAUN MEDICAL (PTY) LTD | 230 | Perifix 431 | 4514319 | B. Braun | Germany | 42 | 50% | Euro | Each | Yes | Each | COMPLY pending confirmation of 1.05mm in terms of gauge |  | Recommended | 95
42142502-00040 | 42 | Needle, epidural, adult, Tuohy bevel luer lock keyed stylet with matched bevelled point. The set must include: A Tuohy needle that: Is graduated in 1cm markings, has a precision fit stylet to prevent tissue coring and is marked to indicate bevel orientation and has a smooth inner bevel to prevent catheter shearing, clear ridged hub to aid with bevel positioning and a comfortable solid grip flange. Needle Size: 16G with a needle length of 80 to 90mm. A loss of resistance syringe of 10ml with low friction An epidural catheter: Is kink resistant with good tensile strength, must have an atraumatic close ended tip with multiple side orifices, must be graduated with standard 1cm catheter markings from 5cm to 15cm to facilitate accurate catheter positioning and there must be double line at 10cm triple lines at 15cm and quadruple lines at 20cm. Epidural catheter size: 18G or 19G Catheter tip/catheter must be radio-opaque. A 20-micron filter that is flat with luer lock connectors. A securing device that connects the catheter securely to the filter – without risk of disconnection, compression of catheter or leaks. Needles, stylets and introducers must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free. Must be sterile and individually peel packed Pack size: Each | OFFER 2 | SUPRA HEALTHCARE JOHANNESBURG PTY LTD | 375.25 | Portex | PORT16X80 | Babcock | United Kingdom | Initial 60 days, thereafter 14 | 84 | ZAR | 1 Pack | YES |  | Epidural catheter size: 20G instead of 18G/19G
42142502-00040 | 42 | Needle, epidural, adult, Tuohy bevel luer lock keyed stylet with matched bevelled point. The set must include: A Tuohy needle that: Is graduated in 1cm markings, has a precision fit stylet to prevent tissue coring and is marked to indicate bevel orientation and has a smooth inner bevel to prevent catheter shearing, clear ridged hub to aid with bevel positioning and a comfortable solid grip flange. Needle Size: 16G with a needle length of 80 to 90mm. A loss of resistance syringe of 10ml with low friction An epidural catheter: Is kink resistant with good tensile strength, must have an atraumatic close ended tip with multiple side orifices, must be graduated with standard 1cm catheter markings from 5cm to 15cm to facilitate accurate catheter positioning and there must be double line at 10cm triple lines at 15cm and quadruple lines at 20cm. Epidural catheter size: 18G or 19G Catheter tip/catheter must be radio-opaque. A 20-micron filter that is flat with luer lock connectors. A securing device that connects the catheter securely to the filter – without risk of disconnection, compression of catheter or leaks. Needles, stylets and introducers must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free. Must be sterile and individually peel packed Pack size: Each BOX (B0X OF 25 PIECES) |  | JUMLA MEDICAL SUPPLIES (PTY) LTD | 3806.5 | EPIFIX-MININ | TEPIMA16903 | TIBBI MEDICAL | TURKEY | 21 DAYS | 70% | US DOLLAR | 1 BOX | YES | NIL | Offered an epidural catheter size of 20G Epidural catheter size: 18G or 19G
42142502-00041 | 43 | Needle, epidural, adult, Tuohy bevel luer lock keyed stylet with matched bevelled point. The set must include: A Tuohy needle that: Is graduated in 1cm markings, has a precision fit stylet to prevent tissue coring and is marked to indicate bevel orientation and has a smooth inner bevel to prevent catheter shearing, clear ridged hub to aid with bevel positioning and a comfortable solid grip flange. Needle Size: 17G with a needle length of 80 to 90mm. A loss of resistance syringe of 10ml with low friction An epidural catheter: Is kink resistant with good tensile strength, must have an atraumatic close ended tip with multiple side orifices, must be graduated with standard 1cm catheter markings from 5cm to 15cm to facilitate accurate catheter positioning and there must be double line at 10cm triple lines at 15cm and quadruple lines at 20cm. Epidural catheter size: 19G Catheter tip/catheter must be radio-opaque. A 20-micron filter that is flat with luer lock connectors. A securing device that connects the catheter securely to the filter – without risk of disconnection, compression of catheter or leaks. Needles, stylets and introducers must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free. Must be sterile and individually peel packed Pack size: Each |  | Medi-Core Technologies | 36.22 | ULIFE | 42142502-00041 | JIAXING U-LIFE MEDICAL DEVICE TECHNOLOGY CO.,LTD | CHINA | 21 | 80 | USD | 1 | YES | EACH | No sample submitted for evaluation
42142502-00042 | 44 | Needle, epidural, adult, Tuohy bevel luer lock keyed stylet with matched bevelled point. The set must include: A Tuohy needle that: Is graduated in 1cm markings, has a precision fit stylet to prevent tissue coring and is marked to indicate bevel orientation and has a smooth inner bevel to prevent catheter shearing, clear ridged hub to aid with bevel positioning and a comfortable solid grip flange. Needle Size: 18G with a needle length of 80 to 90mm. A loss of resistance syringe of 10ml with low friction An epidural catheter: Is kink resistant with good tensile strength, must have an atraumatic close ended tip with multiple side orifices, must be graduated with standard 1cm catheter markings from 5cm to 15cm to facilitate accurate catheter positioning and there must be double line at 10cm triple lines at 15cm and quadruple lines at 20cm. Epidural catheter size: 20G Catheter tip/catheter must be radio-opaque. A 20-micron filter that is flat with luer lock connectors. A securing device that connects the catheter securely to the filter – without risk of disconnection, compression of catheter or leaks. Needles, stylets and introducers must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free. Must be sterile and individually peel packed Pack size: Each |  | Medi-Core Technologies | 36.22 | ULIFE | 4214250-00042 | JIAXING U-LIFE MEDICAL DEVICE TECHNOLOGY CO.,LTD | CHINA | 21 | 80 | USD | 1 | YES | EACH | the markings of the sample are graduated from zero instead from 5cm to 15cm the size of the epidural catherter is not stated
42142502-00042 | 44 | Needle, epidural, adult, Tuohy bevel luer lock keyed stylet with matched bevelled point. The set must include: A Tuohy needle that: Is graduated in 1cm markings, has a precision fit stylet to prevent tissue coring and is marked to indicate bevel orientation and has a smooth inner bevel to prevent catheter shearing, clear ridged hub to aid with bevel positioning and a comfortable solid grip flange. Needle Size: 18G with a needle length of 80 to 90mm. A loss of resistance syringe of 10ml with low friction An epidural catheter: Is kink resistant with good tensile strength, must have an atraumatic close ended tip with multiple side orifices, must be graduated with standard 1cm catheter markings from 5cm to 15cm to facilitate accurate catheter positioning and there must be double line at 10cm triple lines at 15cm and quadruple lines at 20cm. Epidural catheter size: 20G Catheter tip/catheter must be radio-opaque. A 20-micron filter that is flat with luer lock connectors. A securing device that connects the catheter securely to the filter – without risk of disconnection, compression of catheter or leaks. Needles, stylets and introducers must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free. Must be sterile and individually peel packed Pack size: Each |  | CRYSATL PIER TRADING 148cc | 133.7 | EPIDURAL MINI PACK/SR CARE |  | GLOBAL MEDIKIT LIMITED C/O SR MEDICAL CARE | [6 empty] | EACH | The epidural catheter size is not stated. The micron filter is not stated. The markings starts from the tip of the catheter instead of from 5 cm.
42142502-00042 | 44 | Needle, epidural, adult, Tuohy bevel luer lock keyed stylet with matched bevelled point. The set must include: A Tuohy needle that: Is graduated in 1cm markings, has a precision fit stylet to prevent tissue coring and is marked to indicate bevel orientation and has a smooth inner bevel to prevent catheter shearing, clear ridged hub to aid with bevel positioning and a comfortable solid grip flange. Needle Size: 18G with a needle length of 80 to 90mm. A loss of resistance syringe of 10ml with low friction An epidural catheter: Is kink resistant with good tensile strength, must have an atraumatic close ended tip with multiple side orifices, must be graduated with standard 1cm catheter markings from 5cm to 15cm to facilitate accurate catheter positioning and there must be double line at 10cm triple lines at 15cm and quadruple lines at 20cm. Epidural catheter size: 20G Catheter tip/catheter must be radio-opaque. A 20-micron filter that is flat with luer lock connectors. A securing device that connects the catheter securely to the filter – without risk of disconnection, compression of catheter or leaks. Needles, stylets and introducers must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free. Must be sterile and individually peel packed Pack size: Each | OFFER 1 | SUPRA HEALTHCARE JOHANNESBURG PTY LTD | 172.3 | Epifix | EPI18X80 | EGEMEN | TURKEY | Initial 60 days, thereafter 14 | 84 | USD | 1 Box | YES | EACH | COMPLY pending confirmation of latex free |  | Recommended | 96
42142502-00042 | 44 | Needle, epidural, adult, Tuohy bevel luer lock keyed stylet with matched bevelled point. The set must include: A Tuohy needle that: Is graduated in 1cm markings, has a precision fit stylet to prevent tissue coring and is marked to indicate bevel orientation and has a smooth inner bevel to prevent catheter shearing, clear ridged hub to aid with bevel positioning and a comfortable solid grip flange. Needle Size: 18G with a needle length of 80 to 90mm. A loss of resistance syringe of 10ml with low friction An epidural catheter: Is kink resistant with good tensile strength, must have an atraumatic close ended tip with multiple side orifices, must be graduated with standard 1cm catheter markings from 5cm to 15cm to facilitate accurate catheter positioning and there must be double line at 10cm triple lines at 15cm and quadruple lines at 20cm. Epidural catheter size: 20G Catheter tip/catheter must be radio-opaque. A 20-micron filter that is flat with luer lock connectors. A securing device that connects the catheter securely to the filter – without risk of disconnection, compression of catheter or leaks. Needles, stylets and introducers must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free. Must be sterile and individually peel packed Pack size: Each |  | LECHOBA MEDICAL TECHNOLOGIES (PTY) LTD | 194.45 | UNISIS UNIEVER | 5142861 | UNISIS CORPORATION | JAPAN | 21 DAYS | 75% | USD | 1 | YES | Each (PACKED 10/BOX) | No sample submitted for evaluation
42142502-00042 | 44 | Needle, epidural, adult, Tuohy bevel luer lock keyed stylet with matched bevelled point. The set must include: A Tuohy needle that: Is graduated in 1cm markings, has a precision fit stylet to prevent tissue coring and is marked to indicate bevel orientation and has a smooth inner bevel to prevent catheter shearing, clear ridged hub to aid with bevel positioning and a comfortable solid grip flange. Needle Size: 18G with a needle length of 80 to 90mm. A loss of resistance syringe of 10ml with low friction An epidural catheter: Is kink resistant with good tensile strength, must have an atraumatic close ended tip with multiple side orifices, must be graduated with standard 1cm catheter markings from 5cm to 15cm to facilitate accurate catheter positioning and there must be double line at 10cm triple lines at 15cm and quadruple lines at 20cm. Epidural catheter size: 20G Catheter tip/catheter must be radio-opaque. A 20-micron filter that is flat with luer lock connectors. A securing device that connects the catheter securely to the filter – without risk of disconnection, compression of catheter or leaks. Needles, stylets and introducers must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free. Must be sterile and individually peel packed Pack size: Each | 1 | B. BRAUN MEDICAL (PTY) LTD | 265 | Perifix 901 | 4510305 | B. Braun | Germany | 42 | 50% | Euro | Each | Yes | Each
42142502-00042 | 44 | Needle, epidural, adult, Tuohy bevel luer lock keyed stylet with matched bevelled point. The set must include: A Tuohy needle that: Is graduated in 1cm markings, has a precision fit stylet to prevent tissue coring and is marked to indicate bevel orientation and has a smooth inner bevel to prevent catheter shearing, clear ridged hub to aid with bevel positioning and a comfortable solid grip flange. Needle Size: 18G with a needle length of 80 to 90mm. A loss of resistance syringe of 10ml with low friction An epidural catheter: Is kink resistant with good tensile strength, must have an atraumatic close ended tip with multiple side orifices, must be graduated with standard 1cm catheter markings from 5cm to 15cm to facilitate accurate catheter positioning and there must be double line at 10cm triple lines at 15cm and quadruple lines at 20cm. Epidural catheter size: 20G Catheter tip/catheter must be radio-opaque. A 20-micron filter that is flat with luer lock connectors. A securing device that connects the catheter securely to the filter – without risk of disconnection, compression of catheter or leaks. Needles, stylets and introducers must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free. Must be sterile and individually peel packed Pack size: Each | OFFER 2 | SUPRA HEALTHCARE JOHANNESBURG PTY LTD | 339.37 | Portex | PORT18X80 | Babcock | United Kingdom | Initial 60 days, thereafter 14 | 84 | ZAR | 1 Pack | YES |  | Epidural catheter size: 18G instead of 20G
42142502-00042 | 44 | Needle, epidural, adult, Tuohy bevel luer lock keyed stylet with matched bevelled point. The set must include: A Tuohy needle that: Is graduated in 1cm markings, has a precision fit stylet to prevent tissue coring and is marked to indicate bevel orientation and has a smooth inner bevel to prevent catheter shearing, clear ridged hub to aid with bevel positioning and a comfortable solid grip flange. Needle Size: 18G with a needle length of 80 to 90mm. A loss of resistance syringe of 10ml with low friction An epidural catheter: Is kink resistant with good tensile strength, must have an atraumatic close ended tip with multiple side orifices, must be graduated with standard 1cm catheter markings from 5cm to 15cm to facilitate accurate catheter positioning and there must be double line at 10cm triple lines at 15cm and quadruple lines at 20cm. Epidural catheter size: 20G Catheter tip/catheter must be radio-opaque. A 20-micron filter that is flat with luer lock connectors. A securing device that connects the catheter securely to the filter – without risk of disconnection, compression of catheter or leaks. Needles, stylets and introducers must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free. Must be sterile and individually peel packed Pack size: Each BOX (BOX OF 25) |  | JUMLA MEDICAL SUPPLIES (PTY) LTD | 3806.5 | EPIFIX-MINI | TEPIMA18903 | TIBBI MEDICAL | TURKEY | 21 DAYS | 70% | US DOLLAR | 1 BOX | YES | NIL | COMPLY pending confirmation of micro filter
42142509-00001 | 45 | The combined spinal epidural pack system must include: A Tuohy needle that must be graduated in 1cm markings must have a precision fit stylet to prevent tissue coring and is marked to indicate bevel orientation. Must have a smooth inner bevel to prevent catheter shearing clear ridged hub to aid with bevel positioning with a comfortable solid grip flange A loss of resistant syringe of low friction, 10ml. Must have a barrel that is graduate to indicate the degree of plunger advancement. Tuohy needle size: 17G with a length of 80mm to 90mm. A pencil point spinal needle that has a clear/ transparent needle hub to allow for easy identification of spinal fluid flashback and flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and the hub must be clearly indicating the bevel orientation. Must have a stylet that re-insert smoothly and does not hook on the cannula. Must have a stable stylet that does not shake and quiver to allow for easy re-insertion in the cannula. Is aligned precisely with the stylet point. Must have a needle through needle arrangement that does not protrude >15mm beyond Touhy needle. Pencil point spinal needle size: 26G with a must have a preferred locking mechanism. An Epidural catheter is kink resistant with good tensile strength. Must have an atraumatic close ended tip with multiple side orifices. Must be graduated with standard 1cm catheter markings from 5cm to 15cm to facilitate accurate catheter positioning. There must be double line at 10cm triple lines at 15cm and quadruple lines at 20cm. Epidural catheter size: 19G. A 20-micron filter that is flat with luer lock connectors. A securing device that connects the catheter securely to the filter- without risk of disconnection or compression of catheter or leaks. Needle, stylet and introducers must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free. Sterile and individually peel packed Pack size: Each | A | Quality Medical Supplies Pty Ltd | 112 | Romsons | TBA | Romsons International | India | 14 | 80% | USD | 50 Packs | Yes | Each | No sample submitted for evaluation |  | Re-tender
42142509-00001 | 45 | The combined spinal epidural pack system must include: A Tuohy needle that must be graduated in 1cm markings must have a precision fit stylet to prevent tissue coring and is marked to indicate bevel orientation. Must have a smooth inner bevel to prevent catheter shearing clear ridged hub to aid with bevel positioning with a comfortable solid grip flange A loss of resistant syringe of low friction, 10ml. Must have a barrel that is graduate to indicate the degree of plunger advancement. Tuohy needle size: 17G with a length of 80mm to 90mm. A pencil point spinal needle that has a clear/ transparent needle hub to allow for easy identification of spinal fluid flashback and flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and the hub must be clearly indicating the bevel orientation. Must have a stylet that re-insert smoothly and does not hook on the cannula. Must have a stable stylet that does not shake and quiver to allow for easy re-insertion in the cannula. Is aligned precisely with the stylet point. Must have a needle through needle arrangement that does not protrude >15mm beyond Touhy needle. Pencil point spinal needle size: 26G with a must have a preferred locking mechanism. An Epidural catheter is kink resistant with good tensile strength. Must have an atraumatic close ended tip with multiple side orifices. Must be graduated with standard 1cm catheter markings from 5cm to 15cm to facilitate accurate catheter positioning. There must be double line at 10cm triple lines at 15cm and quadruple lines at 20cm. Epidural catheter size: 19G. A 20-micron filter that is flat with luer lock connectors. A securing device that connects the catheter securely to the filter- without risk of disconnection or compression of catheter or leaks. Needle, stylet and introducers must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free. Sterile and individually peel packed Pack size: Each | OFFER 2 | SUPRA HEALTHCARE JOHANNESBURG PTY LTD | 185.38 | Portex | SECS18X90 | Babcock | United Kingdom | Initial 60 days, thereafter 14 | 84 | ZAR | 1 Unit | YES |  | Tuohy needle size: 16G instead of 17G Pencil point spinal needle size: 27G instead of 26G |  | Re-tender
42142509-00001 | 45 | The combined spinal epidural pack system must include: A Tuohy needle that must be graduated in 1cm markings must have a precision fit stylet to prevent tissue coring and is marked to indicate bevel orientation. Must have a smooth inner bevel to prevent catheter shearing clear ridged hub to aid with bevel positioning with a comfortable solid grip flange A loss of resistant syringe of low friction, 10ml. Must have a barrel that is graduate to indicate the degree of plunger advancement. Tuohy needle size: 17G with a length of 80mm to 90mm. A pencil point spinal needle that has a clear/ transparent needle hub to allow for easy identification of spinal fluid flashback and flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and the hub must be clearly indicating the bevel orientation. Must have a stylet that re-insert smoothly and does not hook on the cannula. Must have a stable stylet that does not shake and quiver to allow for easy re-insertion in the cannula. Is aligned precisely with the stylet point. Must have a needle through needle arrangement that does not protrude >15mm beyond Touhy needle. Pencil point spinal needle size: 26G with a must have a preferred locking mechanism. An Epidural catheter is kink resistant with good tensile strength. Must have an atraumatic close ended tip with multiple side orifices. Must be graduated with standard 1cm catheter markings from 5cm to 15cm to facilitate accurate catheter positioning. There must be double line at 10cm triple lines at 15cm and quadruple lines at 20cm. Epidural catheter size: 19G. A 20-micron filter that is flat with luer lock connectors. A securing device that connects the catheter securely to the filter- without risk of disconnection or compression of catheter or leaks. Needle, stylet and introducers must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free. Sterile and individually peel packed Pack size: Each | OFFER 1 | SUPRA HEALTHCARE JOHANNESBURG PTY LTD | 224.6 | Combifix | SECS18X90 | Egemen | TURKEY | Initial 60 days, thereafter 14 | 84 | USD | 1 Unit | YES |  | Tuohy needle size: 16G instead of 17G Pencil point spinal needle size: 27G instead of 26G |  | Re-tender
42142509-00001 | 45 | The combined spinal epidural pack system must include: A Tuohy needle that must be graduated in 1cm markings must have a precision fit stylet to prevent tissue coring and is marked to indicate bevel orientation. Must have a smooth inner bevel to prevent catheter shearing clear ridged hub to aid with bevel positioning with a comfortable solid grip flange A loss of resistant syringe of low friction, 10ml. Must have a barrel that is graduate to indicate the degree of plunger advancement. Tuohy needle size: 17G with a length of 80mm to 90mm. A pencil point spinal needle that has a clear/ transparent needle hub to allow for easy identification of spinal fluid flashback and flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and the hub must be clearly indicating the bevel orientation. Must have a stylet that re-insert smoothly and does not hook on the cannula. Must have a stable stylet that does not shake and quiver to allow for easy re-insertion in the cannula. Is aligned precisely with the stylet point. Must have a needle through needle arrangement that does not protrude >15mm beyond Touhy needle. Pencil point spinal needle size: 26G with a must have a preferred locking mechanism. An Epidural catheter is kink resistant with good tensile strength. Must have an atraumatic close ended tip with multiple side orifices. Must be graduated with standard 1cm catheter markings from 5cm to 15cm to facilitate accurate catheter positioning. There must be double line at 10cm triple lines at 15cm and quadruple lines at 20cm. Epidural catheter size: 19G. A 20-micron filter that is flat with luer lock connectors. A securing device that connects the catheter securely to the filter- without risk of disconnection or compression of catheter or leaks. Needle, stylet and introducers must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free. Sterile and individually peel packed Pack size: Each |  | JUMLA MEDICAL SUPPLIES (PTY) LTD | 226.77 | COMBIFIX | TSWS183 | TIBBI MEDICAL | TURKEY | 21 DAYS | 70% | US DOLLAR | 25 | NO | OFFERING 18G | Epidural catheter size: OF 20G instead The spinal needle of 27G instead of 26G Tuohy needle size of 19G instead of 17G |  | Re-tender
42142509-00001 | 45 | The combined spinal epidural pack system must include: A Tuohy needle that must be graduated in 1cm markings must have a precision fit stylet to prevent tissue coring and is marked to indicate bevel orientation. Must have a smooth inner bevel to prevent catheter shearing clear ridged hub to aid with bevel positioning with a comfortable solid grip flange A loss of resistant syringe of low friction, 10ml. Must have a barrel that is graduate to indicate the degree of plunger advancement. Tuohy needle size: 17G with a length of 80mm to 90mm. A pencil point spinal needle that has a clear/ transparent needle hub to allow for easy identification of spinal fluid flashback and flashback must be quick relative to the needle gauge. Must have a key/slot arrangement of the stylet and the hub must be clearly indicating the bevel orientation. Must have a stylet that re-insert smoothly and does not hook on the cannula. Must have a stable stylet that does not shake and quiver to allow for easy re-insertion in the cannula. Is aligned precisely with the stylet point. Must have a needle through needle arrangement that does not protrude >15mm beyond Touhy needle. Pencil point spinal needle size: 26G with a must have a preferred locking mechanism. An Epidural catheter is kink resistant with good tensile strength. Must have an atraumatic close ended tip with multiple side orifices. Must be graduated with standard 1cm catheter markings from 5cm to 15cm to facilitate accurate catheter positioning. There must be double line at 10cm triple lines at 15cm and quadruple lines at 20cm. Epidural catheter size: 19G. A 20-micron filter that is flat with luer lock connectors. A securing device that connects the catheter securely to the filter- without risk of disconnection or compression of catheter or leaks. Needle, stylet and introducers must be manufactured from good quality stainless steel. Plastic hub must be manufacture from medical grade plastic and all the components must be pyrogenic and latex free. Sterile and individually peel packed Pack size: Each | 1 | B. BRAUN MEDICAL (PTY) LTD | 515 | Espocan W. DS + Pencan 27G x 5 3/8" | 4556763 | B. Braun | Germany | 42 | 50% | Euro | Each | Yes | Each | Re-tender: End-users to revise specification |  | Re-tender
42142502-00044 | 46 | Insulated echogenic needle for regional anaesthesia. The set must include: Needle must have an ergonomic needle hub that is easy to grip, have an insulated shaft with a stimulating tip that allows for needle guidance using nerve stimulation and/ or ultrasound, have distance markers at 1cm intervals, be echogenic and show up well on ultrasound, be short bevelled at 20-30degree and allow for tactile feedback, have long flexible tubing that attaches to the syringe and have a long electrical lead with an end connector that is interchangeable with any nerve stimulator machine. Needle Size: 22G/23G with a needle length of 35mm Needles must be manufactured from good quality stainless steel and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size Each |  | Chemical World cc | 133.4 | MDL | AEQM | MDL | ITALY | 3 | 85 | EURO | 20 | 20 |  | Item cancelled:End-users to proqure using their delagated powers. |  | Item cancelled:End-users to proqure using their delagated powers.
42142502-00044 | 46 | Insulated echogenic needle for regional anaesthesia. The set must include: Needle must have an ergonomic needle hub that is easy to grip, have an insulated shaft with a stimulating tip that allows for needle guidance using nerve stimulation and/ or ultrasound, have distance markers at 1cm intervals, be echogenic and show up well on ultrasound, be short bevelled at 20-30degree and allow for tactile feedback, have long flexible tubing that attaches to the syringe and have a long electrical lead with an end connector that is interchangeable with any nerve stimulator machine. Needle Size: 22G/23G with a needle length of 35mm Needles must be manufactured from good quality stainless steel and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size Each |  | Viking Critical Care (Pty) Ltd | 152.86 | VYGON | 69F6194353 | VYGON | FRANCE | EX. STOCK 30 - 60 DAYS | 80% | EURO | 1 | YES | 1 | Item cancelled:End-users to proqure using their delagated powers. |  | Item cancelled:End-users to proqure using their delagated powers.
42142502-00045 | 47 | Insulated echogenic needle for regional anaesthesia. The set must include: Needle must have an ergonomic needle hub that is easy to grip, have an insulated shaft with a stimulating tip that allows for needle guidance using nerve stimulation and/ or ultrasound, have distance markers at 1cm intervals, be echogenic and show up well on ultrasound, be short bevelled at 20-30degree and allow for tactile feedback, have long flexible tubing that attaches to the syringe and have a long electrical lead with an end connector that is interchangeable with any nerve stimulator machine. Needle Size: 21G/22G with a needle length of 50mm Needles must be manufactured from good quality stainless steel and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size Each | 1 | B. BRAUN MEDICAL (PTY) LTD | 77.5 | Stimuplex A, 30°, 22G x 2", 0.70 x 50mm | 4894502 | B. Braun | Germany | 42 | 50% | Euro | Each | Yes | Each | Item cancelled:End-users to proqure using their delagated powers. |  | Item cancelled:End-users to proqure using their delagated powers.
42142502-00045 | 47 | Insulated echogenic needle for regional anaesthesia. The set must include: Needle must have an ergonomic needle hub that is easy to grip, have an insulated shaft with a stimulating tip that allows for needle guidance using nerve stimulation and/ or ultrasound, have distance markers at 1cm intervals, be echogenic and show up well on ultrasound, be short bevelled at 20-30degree and allow for tactile feedback, have long flexible tubing that attaches to the syringe and have a long electrical lead with an end connector that is interchangeable with any nerve stimulator machine. Needle Size: 21G/22G with a needle length of 50mm Needles must be manufactured from good quality stainless steel and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size Each |  | Chemical World cc | 133.4 | MDL | AEQM | MDL | ITALY | 3 | 85 | EURO | 20 |  |  | Item cancelled:End-users to proqure using their delagated powers. |  | Item cancelled:End-users to proqure using their delagated powers.
42142502-00045 | 47 | Insulated echogenic needle for regional anaesthesia. The set must include: Needle must have an ergonomic needle hub that is easy to grip, have an insulated shaft with a stimulating tip that allows for needle guidance using nerve stimulation and/ or ultrasound, have distance markers at 1cm intervals, be echogenic and show up well on ultrasound, be short bevelled at 20-30degree and allow for tactile feedback, have long flexible tubing that attaches to the syringe and have a long electrical lead with an end connector that is interchangeable with any nerve stimulator machine. Needle Size: 21G/22G with a needle length of 50mm Needles must be manufactured from good quality stainless steel and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size Each |  | Viking Critical Care (Pty) Ltd | 152.86 | VYGON | 69F6194503 | VYGON | FRANCE | EX. STOCK 30 - 60 DAYS | 80% | EURO | 1 | YES | 1 | Item cancelled:End-users to proqure using their delagated powers. |  | Item cancelled:End-users to proqure using their delagated powers.
42142502-00046 | 48 | Insulated echogenic needle for regional anaesthesia. The set must include: Needle must have an ergonomic needle hub that is easy to grip, have an insulated shaft with a stimulating tip that allows for needle guidance using nerve stimulation and/ or ultrasound, have distance markers at 1cm intervals, be echogenic and show up well on ultrasound, be short bevelled at 20-30degree and allow for tactile feedback, have long flexible tubing that attaches to the syringe and have a long electrical lead with an end connector that is interchangeable with any nerve stimulator machine. Needle Size: 21G/22G with a needle length of 90mm to 100mm Needles must be manufactured from good quality stainless steel and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size Each | 1 | B. BRAUN MEDICAL (PTY) LTD | 77.5 | Stimuplex A, 30°, 22G x 2", 0.80 x 100mm | 4894260 | B. Braun | Germany | 42 | 50% | Euro | Each | Yes | Each | Item cancelled:End-users to proqure using their delagated powers. |  | Item cancelled:End-users to proqure using their delagated powers.
42142502-00046 | 48 | Insulated echogenic needle for regional anaesthesia. The set must include: Needle must have an ergonomic needle hub that is easy to grip, have an insulated shaft with a stimulating tip that allows for needle guidance using nerve stimulation and/ or ultrasound, have distance markers at 1cm intervals, be echogenic and show up well on ultrasound, be short bevelled at 20-30degree and allow for tactile feedback, have long flexible tubing that attaches to the syringe and have a long electrical lead with an end connector that is interchangeable with any nerve stimulator machine. Needle Size: 21G/22G with a needle length of 90mm to 100mm Needles must be manufactured from good quality stainless steel and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size Each |  | Chemical World cc | 133.4 | MDL | AEQM | MDL | ITALY | 3 | 85 | EURO | 20 |  |  | Item cancelled:End-users to proqure using their delagated powers. |  | Item cancelled:End-users to proqure using their delagated powers.
42142502-00046 | 48 | Insulated echogenic needle for regional anaesthesia. The set must include: Needle must have an ergonomic needle hub that is easy to grip, have an insulated shaft with a stimulating tip that allows for needle guidance using nerve stimulation and/ or ultrasound, have distance markers at 1cm intervals, be echogenic and show up well on ultrasound, be short bevelled at 20-30degree and allow for tactile feedback, have long flexible tubing that attaches to the syringe and have a long electrical lead with an end connector that is interchangeable with any nerve stimulator machine. Needle Size: 21G/22G with a needle length of 90mm to 100mm Needles must be manufactured from good quality stainless steel and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size Each |  | Viking Critical Care (Pty) Ltd | 152.86 | VYGON | 69F6194103 | VYGON | FRANCE | EX. STOCK 30 - 60 DAYS | 80% | EURO | 1 | YES | 1 | Item cancelled:End-users to proqure using their delagated powers. |  | Item cancelled:End-users to proqure using their delagated powers.
42142502-00047 | 49 | Insulated echogenic needle for regional anaesthesia. The set must include: Needle must have an ergonomic needle hub that is easy to grip, have an insulated shaft with a stimulating tip that allows for needle guidance using nerve stimulation and/ or ultrasound, have distance markers at 1cm intervals, be echogenic and show up well on ultrasound, be short bevelled at 20-30degree and allow for tactile feedback, have long flexible tubing that attaches to the syringe and have a long electrical lead with an end connector that is interchangeable with any nerve stimulator machine. Needle Size: 20G with a needle length of 110mm to 120mm Needles must be manufactured from good quality stainless steel and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size Each |  | Chemical World cc | 133.4 | MDL | AEQM | MDL | ITALY | [6 empty] | Item cancelled:End-users to proqure using their delagated powers. |  | Item cancelled:End-users to proqure using their delagated powers.
42142502-00047 | 49 | Insulated echogenic needle for regional anaesthesia. The set must include: Needle must have an ergonomic needle hub that is easy to grip, have an insulated shaft with a stimulating tip that allows for needle guidance using nerve stimulation and/ or ultrasound, have distance markers at 1cm intervals, be echogenic and show up well on ultrasound, be short bevelled at 20-30degree and allow for tactile feedback, have long flexible tubing that attaches to the syringe and have a long electrical lead with an end connector that is interchangeable with any nerve stimulator machine. Needle Size: 20G with a needle length of 110mm to 120mm Needles must be manufactured from good quality stainless steel and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size Each |  | Viking Critical Care (Pty) Ltd | 229.32 | VYGON | 69F6194153 | VYGON | FRANCE | EX. STOCK 30 - 60 DAYS | 80% | EURO | 1 | YES | 1 | Item cancelled:End-users to proqure using their delagated powers. |  | Item cancelled:End-users to proqure using their delagated powers.
42142502-00048 | 50 | Insulated echogenic needle for regional anaesthesia. The set must include: Needle must have an ergonomic needle hub that is easy to grip, have an insulated shaft with a stimulating tip that allows for needle guidance using nerve stimulation and/ or ultrasound, have distance markers at 1cm intervals, be echogenic and show up well on ultrasound, be short bevelled at 20-30degree and allow for tactile feedback, have long flexible tubing that attaches to the syringe and have a long electrical lead with an end connector that is interchangeable with any nerve stimulator machine. Needle Size: 20G extra-length with a needle length of 150mm Needles must be manufactured from good quality stainless steel and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size Each | 1 | B. BRAUN MEDICAL (PTY) LTD | 77.5 | Stimuplex A, 30°, 20G x 6", 0.90 x 150mm | 4894278 | B. Braun | Germany | 42 | 50% | Euro | Each | Yes | Each |  |  | Item cancelled:End-users to proqure using their delagated powers.
42142502-00048 | 50 | Insulated echogenic needle for regional anaesthesia. The set must include: Needle must have an ergonomic needle hub that is easy to grip, have an insulated shaft with a stimulating tip that allows for needle guidance using nerve stimulation and/ or ultrasound, have distance markers at 1cm intervals, be echogenic and show up well on ultrasound, be short bevelled at 20-30degree and allow for tactile feedback, have long flexible tubing that attaches to the syringe and have a long electrical lead with an end connector that is interchangeable with any nerve stimulator machine. Needle Size: 20G extra-length with a needle length of 150mm Needles must be manufactured from good quality stainless steel and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size Each |  | Chemical World cc | 133.4 | MDL | AEQM | ITALY | 3 | 85 | EURO | 20 | [5 empty] | Item cancelled:End-users to proqure using their delagated powers.
42142502-00048 | 50 | Insulated echogenic needle for regional anaesthesia. The set must include: Needle must have an ergonomic needle hub that is easy to grip, have an insulated shaft with a stimulating tip that allows for needle guidance using nerve stimulation and/ or ultrasound, have distance markers at 1cm intervals, be echogenic and show up well on ultrasound, be short bevelled at 20-30degree and allow for tactile feedback, have long flexible tubing that attaches to the syringe and have a long electrical lead with an end connector that is interchangeable with any nerve stimulator machine. Needle Size: 20G extra-length with a needle length of 150mm Needles must be manufactured from good quality stainless steel and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size Each |  | Viking Critical Care (Pty) Ltd | 229.32 | VYGON | 69F6194153 | VYGON | FRANCE | EX. STOCK 30 - 60 DAYS | 80% | EURO | 1 | YES | 1 |  |  | Item cancelled:End-users to proqure using their delagated powers.
42142502-00049 | 51 | Needle echogenic, ultrasound localiser, nerve insulated 30º bevelled, needle tip and needle shaft must be 360º ultrasound visible the set must include: Needle must have an ergonomic needle hub that is easy to grip, have an insulated shaft that is 360º ultrasound visible, tip that allows for needle guidance using nerve stimulation and ultrasound, have distance markers at 1cm intervals, be echogenic and show up well on ultrasound, be short bevelled at 20-30degree and allow for tactile feedback, have long flexible tubing that attaches to the syringe and have a long electrical lead with an end connector that is interchangeable with any nerve stimulator machine. Needle size: 22G with a needle length of 50mm Needles must be manufactured from good quality stainless steel and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size Each |  | Chemical World cc | 133.4 | MDL | AEQM | ITALY | 3 | 85 | EURO | 20 |  |  |  | Item cancelled:End-users to proqure using their delagated powers. |  | Item cancelled:End-users to proqure using their delagated powers.
42142502-00049 | 51 | Needle echogenic, ultrasound localiser, nerve insulated 30º bevelled, needle tip and needle shaft must be 360º ultrasound visible the set must include: Needle must have an ergonomic needle hub that is easy to grip, have an insulated shaft that is 360º ultrasound visible, tip that allows for needle guidance using nerve stimulation and ultrasound, have distance markers at 1cm intervals, be echogenic and show up well on ultrasound, be short bevelled at 20-30degree and allow for tactile feedback, have long flexible tubing that attaches to the syringe and have a long electrical lead with an end connector that is interchangeable with any nerve stimulator machine. Needle size: 22G with a needle length of 50mm Needles must be manufactured from good quality stainless steel and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size Each | 1 | B. BRAUN MEDICAL (PTY) LTD | 135 | Stimuplex Ultra 360 30°, 22G x 50mm | 4892505-01 | B. Braun | Germany | 42 | 50% | Euro | Each | Yes | Each | Item cancelled:End-users to proqure using their delagated powers. |  | Item cancelled:End-users to proqure using their delagated powers.
42142502-00049 | 51 | Needle echogenic, ultrasound localiser, nerve insulated 30º bevelled, needle tip and needle shaft must be 360º ultrasound visible the set must include: Needle must have an ergonomic needle hub that is easy to grip, have an insulated shaft that is 360º ultrasound visible, tip that allows for needle guidance using nerve stimulation and ultrasound, have distance markers at 1cm intervals, be echogenic and show up well on ultrasound, be short bevelled at 20-30degree and allow for tactile feedback, have long flexible tubing that attaches to the syringe and have a long electrical lead with an end connector that is interchangeable with any nerve stimulator machine. Needle size: 22G with a needle length of 50mm Needles must be manufactured from good quality stainless steel and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size Each |  | Viking Critical Care (Pty) Ltd | 152.86 | VYGON | 69F6194503 | VYGON | FRANCE | EX. STOCK 30-60 DAYS | 80% | EURO | 1 | YES | 1 | Item cancelled:End-users to proqure using their delagated powers. |  | Item cancelled:End-users to proqure using their delagated powers.
42142502-00050 | 52 | Needle echogenic, ultrasound localiser, nerve insulated 30º bevelled, needle tip and needle shaft must be 360º ultrasound visible the set must include: Needle must have an ergonomic needle hub that is easy to grip, have an insulated shaft that is 360º ultrasound visible, tip that allows for needle guidance using nerve stimulation and ultrasound, have distance markers at 1cm intervals, be echogenic and show up well on ultrasound, be short bevelled at 20-30degree and allow for tactile feedback, have long flexible tubing that attaches to the syringe and have a long electrical lead with an end connector that is interchangeable with any nerve stimulator machine. Needle size: 22G with a needle length of 80mm Needles must be manufactured from good quality stainless steel and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size Each |  | Chemical World cc | 133.4 | MDL | AEQM | ITALY | 3 | 85 | EURO | 20 |  |  |  | Item cancelled:End-users to proqure using their delagated powers. |  | Item cancelled:End-users to proqure using their delagated powers.
42142502-00050 | 52 | Needle echogenic, ultrasound localiser, nerve insulated 30º bevelled, needle tip and needle shaft must be 360º ultrasound visible the set must include: Needle must have an ergonomic needle hub that is easy to grip, have an insulated shaft that is 360º ultrasound visible, tip that allows for needle guidance using nerve stimulation and ultrasound, have distance markers at 1cm intervals, be echogenic and show up well on ultrasound, be short bevelled at 20-30degree and allow for tactile feedback, have long flexible tubing that attaches to the syringe and have a long electrical lead with an end connector that is interchangeable with any nerve stimulator machine. Needle size: 22G with a needle length of 80mm Needles must be manufactured from good quality stainless steel and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size Each | 1 | B. BRAUN MEDICAL (PTY) LTD | 135 | Stimuplex Ultra 360 30°, 22G x 80mm | 4892508-01 | B. Braun | Germany | 42 | 50% | Euro | Each | Yes | Each | Item cancelled:End-users to proqure using their delagated powers. |  | Item cancelled:End-users to proqure using their delagated powers.
42142502-00050 | 52 | Needle echogenic, ultrasound localiser, nerve insulated 30º bevelled, needle tip and needle shaft must be 360º ultrasound visible the set must include: Needle must have an ergonomic needle hub that is easy to grip, have an insulated shaft that is 360º ultrasound visible, tip that allows for needle guidance using nerve stimulation and ultrasound, have distance markers at 1cm intervals, be echogenic and show up well on ultrasound, be short bevelled at 20-30degree and allow for tactile feedback, have long flexible tubing that attaches to the syringe and have a long electrical lead with an end connector that is interchangeable with any nerve stimulator machine. Needle size: 22G with a needle length of 80mm Needles must be manufactured from good quality stainless steel and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size Each |  | Viking Critical Care (Pty) Ltd | 152.86 | VYGON | 69F6194853 | VYGON | FRANCE | EX. STOCK 30-60 DAYS | 80% | EURO | 1 | NO, OUR PRODUCT IS 85MM | 1 | Item cancelled:End-users to proqure using their delagated powers. |  | Item cancelled:End-users to proqure using their delagated powers.
42142502-00043 | 53 | Needle echogenic, ultrasound localiser, nerve insulated 30º bevelled, needle tip and needle shaft must be 360º ultrasound visible the set must include: Needle must have an ergonomic needle hub that is easy to grip, have an insulated shaft that is 360º ultrasound visible, tip that allows for needle guidance using nerve stimulation and ultrasound, have distance markers at 1cm intervals, be echogenic and show up well on ultrasound, be short bevelled at 20-30degree and allow for tactile feedback, have long flexible tubing that attaches to the syringe and have a long electrical lead with an end connector that is interchangeable with any nerve stimulator machine. Needle size: 21G with a needle length of 100mm Needles must be manufactured from good quality stainless steel and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size Each |  | Chemical World cc | 133.34 |  |  | ITALY | 3 | 85 | EURO | 20 |  |  |  | Item cancelled:End-users to proqure using their delagated powers. |  | Item cancelled:End-users to proqure using their delagated powers.
42142502-00043 | 53 | Needle echogenic, ultrasound localiser, nerve insulated 30º bevelled, needle tip and needle shaft must be 360º ultrasound visible the set must include: Needle must have an ergonomic needle hub that is easy to grip, have an insulated shaft that is 360º ultrasound visible, tip that allows for needle guidance using nerve stimulation and ultrasound, have distance markers at 1cm intervals, be echogenic and show up well on ultrasound, be short bevelled at 20-30degree and allow for tactile feedback, have long flexible tubing that attaches to the syringe and have a long electrical lead with an end connector that is interchangeable with any nerve stimulator machine. Needle size: 21G with a needle length of 100mm Needles must be manufactured from good quality stainless steel and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size Each |  | Viking Critical Care (Pty) Ltd | 152.86 | VYGON | 69F6194103 | VYGON | FRANCE | EX. STOCK 30 - 60 DAYS | 80% | EURO | 1 | YES | 1 | Item cancelled:End-users to proqure using their delagated powers. |  | Item cancelled:End-users to proqure using their delagated powers.
42142502-00051 | 54 | Needle echogenic, ultrasound localiser, nerve insulated 30º bevelled, needle tip and needle shaft must be 360º ultrasound visible the set must include: Needle must have an ergonomic needle hub that is easy to grip, have an insulated shaft that is 360º ultrasound visible, tip that allows for needle guidance using nerve stimulation and ultrasound, have distance markers at 1cm intervals, be echogenic and show up well on ultrasound, be short bevelled at 20-30degree and allow for tactile feedback, have long flexible tubing that attaches to the syringe and have a long electrical lead with an end connector that is interchangeable with any nerve stimulator machine. Needle size: 20G with a needle length of 120mm Needles must be manufactured from good quality stainless steel and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size Each |  | Chemical World cc | 133.4 | MDL | AEQM | MDL | ITALY | [6 empty] | Item cancelled:End-users to proqure using their delagated powers. |  | Item cancelled:End-users to proqure using their delagated powers.
42142502-00051 | 54 | Needle echogenic, ultrasound localiser, nerve insulated 30º bevelled, needle tip and needle shaft must be 360º ultrasound visible the set must include: Needle must have an ergonomic needle hub that is easy to grip, have an insulated shaft that is 360º ultrasound visible, tip that allows for needle guidance using nerve stimulation and ultrasound, have distance markers at 1cm intervals, be echogenic and show up well on ultrasound, be short bevelled at 20-30degree and allow for tactile feedback, have long flexible tubing that attaches to the syringe and have a long electrical lead with an end connector that is interchangeable with any nerve stimulator machine. Needle size: 20G with a needle length of 120mm Needles must be manufactured from good quality stainless steel and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size Each | 1 | B. BRAUN MEDICAL (PTY) LTD | 135 | Stimuplex Ultra 360 30°, 20G x 150mm | 4892515-01 | B. Braun | Germany | 42 | 50% | Euro | Each | Yes | Each | Item cancelled:End-users to proqure using their delagated powers. |  | Item cancelled:End-users to proqure using their delagated powers.
42142502-00051 | 54 | Needle echogenic, ultrasound localiser, nerve insulated 30º bevelled, needle tip and needle shaft must be 360º ultrasound visible the set must include: Needle must have an ergonomic needle hub that is easy to grip, have an insulated shaft that is 360º ultrasound visible, tip that allows for needle guidance using nerve stimulation and ultrasound, have distance markers at 1cm intervals, be echogenic and show up well on ultrasound, be short bevelled at 20-30degree and allow for tactile feedback, have long flexible tubing that attaches to the syringe and have a long electrical lead with an end connector that is interchangeable with any nerve stimulator machine. Needle size: 20G with a needle length of 120mm Needles must be manufactured from good quality stainless steel and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size Each |  | Viking Critical Care (Pty) Ltd | 229.32 | VYGON | 69F6194123 | VYGON | FRANCE | EX. STOCK 30-60 DAYS | 80% | EURO | 1 | YES | 1 | Item cancelled:End-users to proqure using their delagated powers. |  | Item cancelled:End-users to proqure using their delagated powers.
42142502-00052 | 55 | Needle echogenic, ultrasound localiser, nerve insulated 30º bevelled, needle tip and needle shaft must be 360º ultrasound visible the set must include: Needle must have an ergonomic needle hub that is easy to grip, have an insulated shaft that is 360º ultrasound visible, tip that allows for needle guidance using ultrasound, have distance markers at 1cm intervals, be echogenic and show up well on ultrasound, be short bevelled at 20-30degree and allow for tactile feedback, have long flexible tubing that attaches to the syringe. Needle size: 22G with a needle length of 50mm Needles must be manufactured from good quality stainless steel and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size Each | 1 | B. BRAUN MEDICAL (PTY) LTD | 125 | Ultraplex 360 30°, 22G x 50mm | 4892605-01 | B. Braun | Germany | 42 | 50% | Euro | Each | Yes | Each | COMPLY pending confirmation of latex free |  | Recommended | 95
42142502-00052 | 55 | Needle echogenic, ultrasound localiser, nerve insulated 30º bevelled, needle tip and needle shaft must be 360º ultrasound visible the set must include: Needle must have an ergonomic needle hub that is easy to grip, have an insulated shaft that is 360º ultrasound visible, tip that allows for needle guidance using nerve stimulation and ultrasound, have distance markers at 1cm intervals, be echogenic and show up well on ultrasound, be short bevelled at 20-30degree and allow for tactile feedback, have long flexible tubing that attaches to the syringe and have a long electrical lead with an end connector that is interchangeable with any nerve stimulator machine. Needle size: 22G with a needle length of 50mm Needles must be manufactured from good quality stainless steel and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size Each |  | Chemical World cc | 133.4 | [10 empty] | No sample submitted for evaluation
42142502-00052 | 55 | Needle echogenic, ultrasound localiser, nerve insulated 30º bevelled, needle tip and needle shaft must be 360º ultrasound visible the set must include: Needle must have an ergonomic needle hub that is easy to grip, have an insulated shaft that is 360º ultrasound visible, tip that allows for needle guidance using ultrasound, have distance markers at 1cm intervals, be echogenic and show up well on ultrasound, be short bevelled at 20-30degree and allow for tactile feedback, have long flexible tubing that attaches to the syringe. Needle size: 22G with a needle length of 50mm Needles must be manufactured from good quality stainless steel and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size Each |  | Viking Critical Care (Pty) Ltd | 152.86 | VYGON | 69F6194503 | VYGON | FRANCE | EX. STOCK 30 - 60 DAYS | 80% | EURO | 1 | YES | 1 | COMPLY pending confirmation of latex free
42142502-00053 | 56 | Needle echogenic, ultrasound localiser, nerve insulated 30º bevelled, needle tip and needle shaft must be 360º ultrasound visible the set must include: Needle must have an ergonomic needle hub that is easy to grip, have an insulated shaft that is 360º ultrasound visible, tip that allows for needle guidance using ultrasound, have distance markers at 1cm intervals, be echogenic and show up well on ultrasound, be short bevelled at 20-30degree and allow for tactile feedback, have long flexible tubing that attaches to the syringe. Needle size: 22G with a needle length of 80mm Needles must be manufactured from good quality stainless steel and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size Each | 1 | B. BRAUN MEDICAL (PTY) LTD | 125 | Ultraplex 360 30°, 22G x 80mm | 4892608-01 | B. Braun | Germany | 42 | 50% | Euro | Each | Yes | Each | COMPLY pending confirmation of latex free |  | Recommended | 95
42142502-00053 | 56 | Needle echogenic, ultrasound localiser, nerve insulated 30º bevelled, needle tip and needle shaft must be 360º ultrasound visible the set must include: Needle must have an ergonomic needle hub that is easy to grip, have an insulated shaft that is 360º ultrasound visible, tip that allows for needle guidance using ultrasound, have distance markers at 1cm intervals, be echogenic and show up well on ultrasound, be short bevelled at 20-30degree and allow for tactile feedback, have long flexible tubing that attaches to the syringe. Needle size: 22G with a needle length of 80mm Needles must be manufactured from good quality stainless steel and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size Each |  | Chemical World cc | 133.4 | MDL | AEQM | MDL | ITALY | [6 empty] | No sample submitted for evaluation
42142502-00053 | 56 | Needle echogenic, ultrasound localiser, nerve insulated 30º bevelled, needle tip and needle shaft must be 360º ultrasound visible the set must include: Needle must have an ergonomic needle hub that is easy to grip, have an insulated shaft that is 360º ultrasound visible, tip that allows for needle guidance using ultrasound, have distance markers at 1cm intervals, be echogenic and show up well on ultrasound, be short bevelled at 20-30degree and allow for tactile feedback, have long flexible tubing that attaches to the syringe. Needle size: 22G with a needle length of 80mm Needles must be manufactured from good quality stainless steel and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size Each |  | Viking Critical Care (Pty) Ltd | 152.86 | VYGON | 69F6194853 | VYGON | FRANCE | EX. STOCK 30 - 60 DAYS | 80% | EURO | 1 | NO, OUR PRODUCT IS 85MM | 1 |  | Offered a 21G instead 22G Offered a needle length 85mm instead of 80mm
42142502-00054 | 57 | Needle echogenic, ultrasound localiser, nerve insulated 30º bevelled, needle tip and needle shaft must be 360º ultrasound visible the set must include: Needle must have an ergonomic needle hub that is easy to grip, have an insulated shaft that is 360º ultrasound visible, tip that allows for needle guidance using ultrasound, have distance markers at 1cm intervals, be echogenic and show up well on ultrasound, be short bevelled at 20-30degree and allow for tactile feedback, have long flexible tubing that attaches to the syringe. Needle size: 21G with a needle length of 110mm Needles must be manufactured from good quality stainless steel and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size Each |  | Chemical World cc | 133.4 | MDL | AEQM | MDL | ITALY | 3 | 85 | EURO | 20 |  |  | No sample submitted for evaluation |  | Re-tender
42142502-00054 | 57 | Needle echogenic, ultrasound localiser, nerve insulated 30º bevelled, needle tip and needle shaft must be 360º ultrasound visible the set must include: Needle must have an ergonomic needle hub that is easy to grip, have an insulated shaft that is 360º ultrasound visible, tip that allows for needle guidance using ultrasound, have distance markers at 1cm intervals, be echogenic and show up well on ultrasound, be short bevelled at 20-30degree and allow for tactile feedback, have long flexible tubing that attaches to the syringe. Needle size: 21G with a needle length of 110mm Needles must be manufactured from good quality stainless steel and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size Each |  | Viking Critical Care (Pty) Ltd | 152.86 | VYGON | 69F6194103 | VYGON | FRANCE | EX. STOCK 30 - 60 DAYS | 80% | EURO | 1 | NO, OUR PRODUCT IS 100MM | 1 | Offered a needle length 100mm instead of 110mm |  | Re-tender
42142502-00055 | 58 | Needle echogenic, ultrasound localiser, nerve insulated 30º bevelled, needle tip and needle shaft must be 360º ultrasound visible the set must include: Needle must have an ergonomic needle hub that is easy to grip, have an insulated shaft that is 360º ultrasound visible, tip that allows for needle guidance using ultrasound, have distance markers at 1cm intervals, be echogenic and show up well on ultrasound, be short bevelled at 20-30degree and allow for tactile feedback, have long flexible tubing that attaches to the syringe. Needle size: 21G with a needle length of 150mm Needles must be manufactured from good quality stainless steel and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size Each |  | Chemical World cc | 133.4 | MDL | AEQM | MDL | ITALY | 3 | 85 | EURO | 20 |  |  | No sample submitted for evaluation |  | Re-tender
42142502-00055 | 58 | Needle echogenic, ultrasound localiser, nerve insulated 30º bevelled, needle tip and needle shaft must be 360º ultrasound visible the set must include: Needle must have an ergonomic needle hub that is easy to grip, have an insulated shaft that is 360º ultrasound visible, tip that allows for needle guidance using ultrasound, have distance markers at 1cm intervals, be echogenic and show up well on ultrasound, be short bevelled at 20-30degree and allow for tactile feedback, have long flexible tubing that attaches to the syringe. Needle size: 21G with a needle length of 150mm Needles must be manufactured from good quality stainless steel and all the components must be pyrogenic and latex free Sterile and individually peel packed Pack size Each |  | Viking Critical Care (Pty) Ltd | 229.32 | VYGON | 69F6194153 | VYGON | FRANCE | EX. STOCK 30 - 60 DAYS | 80% | EURO | 1 | YES | 1 | Offered a 20G instead 21G |  | Re-tender
42142502-00056 | 59 | Vascular access echogenic needle. needle tip and needle shaft must be 360º ultrasound visible, Sizes: 18G Length 70mm Sterile and individually peel packed Pack size: Each |  | Chemical World cc | 80.5 | [10 empty] | No sample submitted for evaluation |  | Re-tender
42142502-00056 | 59 |  |  | Zanotone (Pty) Ltd | 4.72 | Mekon |  | Mekon | [6 empty] | Each | No sample submitted for evaluation |  | Re-tender
42142504-00000 | 60 | Needle, localisation, repositionable for non-palpable breast lesion. Size: 20G x 100mm Sterile, Single use, Individually peel packed EACH |  | ABCARE TECHNOLOGIES (PTY) LTD | 301 | ZAMAR | BN001 | ZAMAR DOO | CROATIA | 45 | 85 | EURO | 100 | YES | EACH | Product information is reflected on a removable sticker. LOT number, shelf life and date of production of the sticker and outer package is not the same. Product information of a removable sticker indicates a size 21G and the product information of the outer package indicates size 28G
42142504-00000 | 60 | Needle, localisation, repositionable for non-palpable breast lesion. Size: 20G x 100mm Sterile, Single use, Individually peel packed EACH |  | Chemical World cc | 363.4 | MDL | AEQM | MDL | ITALY | 3 | 85 | EURO | 20 |  | EACH | COMPLY pending confirmation of latex free |  | Recommended | 100
42142506-00000 | 61 | Needle, fill universal, with blunt tip. Unique 40º bevel needle to prevent needle stick injuries. Size: 18G x 40mm (1.5") Colour code: Red Hub and shield Sterile, Single Use, Individually peel packed Unit: Box (50 per box) |  | ALMIKA TRADING PTY LTD | 27 | BD | 305180 | BD | USA | 14 | 0 | NO | 100 | YES | 50 PER BOX | COMPLY |  | Recommended | 100
42142506-00000 | 61 | Needle, fill universal, with blunt tip. Unique 40º bevel needle to prevent needle stick injuries. Size: 18G x 40mm (1.5") Colour code: Red Hub and shield Sterile, Single Use, Individually peel packed Unit: Box (50 per box) |  | Zanotone (Pty) Ltd | 0.5 | Mekon |  | Mekon | [6 empty] | Each | No brand name No product code
42142506-00000 | 61 | Needle, fill universal, with blunt tip. Unique 40º bevel needle to prevent needle stick injuries. Size: 18G x 40mm (1.5") Colour code: Red Hub and shield Sterile, Single Use, Individually peel packed Unit: Box (50 per box) |  | LECHOBA MEDICAL TECHNOLOGIES (PTY) LTD | 0.7 | MEDOJECT BLUNT | CH18112BF | CHIRANA INJECTA A/S | SLOVAK REPUBLIC | 21 DAYS | 75% | EURO | 100 | YES | EACH (PACKED 100/BOX) | No sample submitted for evaluation
42142506-00000 | 61 | Needle, fill universal, with blunt tip. Unique 40º bevel needle to prevent needle stick injuries. Size: 18G x 40mm (1.5") Colour code: Red Hub and shield Sterile, Single Use, Individually peel packed Unit: Box (50 per box) | Basic Line offer | Becton Dickinson | 0.9 | BD™ Blunt Fill Needle 18G x 100 | 305180 | Becton Dickinson | Fraga, Spain | 1 week | 76% | USD | 1 BOX | NO | 100 Units in a box | COMPLY pending confirmation of latex free
42142506-00000 | 61 | Needle, fill universal, with blunt tip. Unique 40º bevel needle to prevent needle stick injuries. Size: 18G x 40mm (1.5") Colour code: Red Hub and shield Sterile, Single Use, Individually peel packed Unit: Box (50 per box) |  | ABCARE TECHNOLOGIES (PTY) LTD | 0.99 | MEDOJECT | BN002 | CHIRANA T INJEKTA | SLOVAKIA | 45 | 85 | EURO | 100 | YES | EACH | COMPLY
42142506-00000 | 61 | Needle, fill universal, with blunt tip. Unique 40º bevel needle to prevent needle stick injuries. Size: 18G x 40mm (1.5") Colour code: Red Hub and shield Sterile, Single Use, Individually peel packed Unit: Box (50 per box) | Basic Line Offer | BIOCOR HOSPITAL SUPPLIES PTY LTD | 1.97 | Medline | N/A | Yangzhong Medline Industry Co.Ltd | China | 60 | 85 | USD | 10 Boxes of 50 | Yes | N/A | COMPLY pending confirmation of latex free
42142506-00001 | 62 | Needle, fill universal, with blunt tip and 5-micron filter. Unique 40º bevel needle to prevent needle stick injuries. Size: 18G x 40mm x (1.5") Colour code: Red Hub and shield Sterile, Single Use, Individually peel packed Unit: Box (50 per box) |  | ALMIKA TRADING PTY LTD | 81 | BD | 305211 | BD | USA | 14 | 0 | NO | 100 | YES | 50 PER BOX | COMPLY |  | Recommended | 100
42142506-00001 | 62 | Needle, fill universal, with blunt tip and 5-micron filter. Unique 40º bevel needle to prevent needle stick injuries. Size: 18G x 40mm x (1.5") Colour code: Red Hub and shield Sterile, Single Use, Individually peel packed Unit: Box (50 per box) |  | Zanotone (Pty) Ltd | 1.94 | Mekon |  | Mekon | [6 empty] | Each | No brand name No product code
42142506-00001 | 62 | Needle, fill universal, with blunt tip and 5-micron filter. Unique 40º bevel needle to prevent needle stick injuries. Size: 18G x 40mm x (1.5") Colour code: Red Hub and shield Sterile, Single Use, Individually peel packed Unit: Box (50 per box) |  | ABCARE TECHNOLOGIES (PTY) LTD | 1.99 | MEDOJECT | BN003 | CHIRANA T INJEKTA | SLOVAKIA | 45 | 85 | EURO | 100 | YES | EACH | COMPLY
42142506-00001 | 62 | Needle, fill universal, with blunt tip and 5-micron filter. Unique 40º bevel needle to prevent needle stick injuries. Size: 18G x 40mm x (1.5") Colour code: Red Hub and shield Sterile, Single Use, Individually peel packed Unit: Box (50 per box) |  | LECHOBA MEDICAL TECHNOLOGIES (PTY) LTD | 2.3 | MEDOJECT BLUNT | CH18112BF-FILTER | CHIRANA INJECTA A/S | SLOVAK REPUBLIC | 21 DAYS | 75% | EURO | 100 | YES | EACH (PACKED 100/BOX) | No sample submitted for evaluation
42142506-00001 | 62 | Needle, fill universal, with blunt tip and 5-micron filter. Unique 40º bevel needle to prevent needle stick injuries. Size: 18G x 40mm x (1.5") Colour code: Red Hub and shield Sterile, Single Use, Individually peel packed Unit: Box (50 per box) | Basic Line offer | Becton Dickinson | 2.9 | BD™ Blunt Fill Needle with 5 Micron Filter x 100 | 305211 | Becton Dickinson | Fraga, Spain | 1 week | 76% | USD | 1 BOX | NO | 100 Units in a box | COMPLY pending confirmation of latex free
42142506-00001 | 62 | Needle, fill universal, with blunt tip and 5-micron filter. Unique 40º bevel needle to prevent needle stick injuries. Size: 18G x 40mm x (1.5") Colour code: Red Hub and shield Sterile, Single Use, Individually peel packed Unit: Box (50 per box) | 1 | Biologica Pharmaceuticals Pty Ltd | 4.07 | Sol- M | 110022F | Sol- Millennium | China | 4-6 Weeks | 100% | USD | 1000 Boxes | 100% | N/A
42142523-00015 | 63 | Needle, hypodermic, long point with plastic 6%, luer lock hub Size: 15G (1,8mm) x 40mm Colour code: Blue-grey Sterile, Single use, Individually packed To comply with the latest issue of SANS 1124-1 UNIT: Box (100 per box) | N/A | PROMED TECHNOLOGIES (PTY) LTD | 0.18 | U-LIFE | 523015 | JIAXING | CHINA | 14 DAYS | 85% | USD | 1 UNIT | YES | UNIT | No samples submitted for evaluation | NO TEST REPORTS
42142523-00015 | 63 | Needle, hypodermic, long point with plastic 6%, luer lock hub Size: 15G (1,8mm) x 40mm Colour code: Blue-grey Sterile, Single use, Individually packed To comply with the latest issue of SANS 1124-1 UNIT: Box (100 per box) |  | Chemical World cc | 0.1823 | SURGI PLUS | NDHD015 | TONGDA | CHINA | 3 | 85 | USD | 5000 | YES | EACH | COMPLY pending confirmation of latex free | NOT COMPLY
42142523-00015 | 63 | Needle, hypodermic, long point with plastic 6%, luer lock hub Size: 15G (1,8mm) x 40mm Colour code: Blue-grey Sterile, Single use, Individually packed To comply with the latest issue of SANS 1124-1 UNIT: Box (100 per box) |  | Medi-Core Technologies | 0.19 | ULIFE | 41242523-00015 | JIAXING U-LIFE MEDICAL DEVICE TECHNOLOGY CO.,LTD | CHINA | 21 | 80 | USD | 1 BOX | YES | BOX | No sample submitted for evaluation | NO TEST REPORTS
42142523-00015 | 63 | Needle, hypodermic, long point with plastic 6%, luer lock hub Size: 15G (1,8mm) x 40mm Colour code: Blue-grey Sterile, Single use, Individually packed To comply with the latest issue of SANS 1124-1 UNIT: Box (100 per box) | 1 | CENTRAL MEDICAL PTY LTD | 25.75 | LYSA | HDN1540 | Zhejian INI Medical Devices Co Ltd | China | 21 | 80 | China yuan | 100 boxes of 100 units | yes | box of 100 units | COMPLY pending confirmation of latex free | COMPLY | Recommended | 100
42142523-00015 | 63 | Needle, hypodermic, long point with plastic 6%, luer lock hub Size: 15G (1,8mm) x 40mm Colour code: Blue-grey Sterile, Single use, Individually packed To comply with the latest issue of SANS 1124-1 UNIT: Box (100 per box) | Basic Line Offer | BIOCOR HOSPITAL SUPPLIES PTY LTD | 0.27 | Triple -D | NDL-Care 15E | Jumla Medical | China | 21 | 0 | N/A | 5 Boxes | Yes | N/A | COMPLY pending confirmation of latex free | NO TEST REPORTS
42142523-00015 | 63 | Needle, hypodermic, long point with plastic 6%, luer lock hub Size: 15G (1,8mm) x 40mm Colour code: Blue-grey Sterile, Single use, Individually packed To comply with the latest issue of SANS 1124-1 UNIT: Box (100 per box) |  | GIBELA TRADE AND INVEST 1001 CC | 0.33434 | TARA | NEE004 | Shandong Qiaopai Group Co., Ltd | CHINA | 21 | 85 | USD | 10 | YES | each | No sample submitted for evaluation | NO TEST REPORTS
42142523-00015 | 63 | Needle, hypodermic, long point with plastic 6%, luer lock hub Size: 15G (1,8mm) x 40mm Colour code: Blue-grey Sterile, Single use, Individually packed To comply with the latest issue of SANS 1124-1 UNIT: Box (100 per box) |  | SUPRA HEALTHCARE JOHANNESBURG PTY LTD | 0.37 | Avacare | DN15 | HENAN AILE INDUSTRIAL CO.,LTD | CHINA | Initial 60 days, thereafter 14 | 84 | USD | 1 Box | YES | 100 pcs/box | COMPLY pending confirmation of latex free | NOT COMPLY
42142523-00015 | 63 | Needle, hypodermic, long point with plastic 6%, luer lock hub Size: 15G (1,8mm) x 40mm Colour code: Blue-grey Sterile, Single use, Individually packed To comply with the latest issue of SANS 1124-1 UNIT: Box (100 per box) |  | JUMLA MEDICAL SUPPLIES (PTY) LTD | 20.02 | NDL-CARE | NDL-CARE-15G | SUZHOU ACMED | CHINA | 21 DAYS | 72% | US DOLLAR | 10 BOXES | YES | NIL | COMPLY pending confirmation of latex free | NOT COMPLY
42142523-00014 | 64 | Needle, hypodermic, long point with plastic 6%, luer lock hub Size: 18G (1,2mm) x 40mm Colour code: Pink Sterile, Single use, individually peel packed To comply with the latest issue of SANS 1124-1 UNIT: Box (100 per box) |  | Isigidi Trading 95 cc | 0.15 | Isigidi | NE03 | Jiangsu Jichun | China | 14 days | 85% | USD | 100boxes | Yes | 100pcs/box | COMPLY pending confirmation of latex free | COMPLY
42142523-00014 | 64 | Needle, hypodermic, long point with plastic 6%, luer lock hub Size: 18G (1,2mm) x 40mm Colour code: Pink Sterile, Single use, individually peel packed To comply with the latest issue of SANS 1124-1 UNIT: Box (100 per box) | 1 | CENTRAL MEDICAL PTY LTD | 15.35 | LYSA | HDN1840 | Zhejian INI Medical Devices Co Ltd | China | 21 | 80 | China yuan | 100 boxes of 100 units | yes | box of 100 units | COMPLY pending confirmation of latex free | COMPLY | Recommended | 94
42142523-00014 | 64 | Needle, hypodermic, long point with plastic 6%, luer lock hub Size: 18G (1,2mm) x 40mm Colour code: Pink Sterile, Single use, individually peel packed To comply with the latest issue of SANS 1124-1 UNIT: Box (100 per box) | N/A | PROMED TECHNOLOGIES (PTY) LTD | 0.16 | U-LIFE | 523014 | JIAXING | CHINA | 14 DAYS | 85% | USD | 1 UNIT | YES | UNIT | COMPLY pending confirmation of latex free | COMPLY
42142523-00014 | 64 | Needle, hypodermic, long point with plastic 6%, luer lock hub Size: 18G (1,2mm) x 40mm Colour code: Pink Sterile, Single use, individually peel packed To comply with the latest issue of SANS 1124-1 UNIT: Box (100 per box) |  | Chemical World cc | 0.17 | SURGI PLUS | NDHD018 | TONGDA | CHINA | 3 | 85 | USD | 5000 | YES | EACH | COMPLY pending confirmation of latex free | NOT COMPLY
42142523-00014 | 64 | Needle, hypodermic, long point with plastic 6%, luer lock hub Size: 18G (1,2mm) x 40mm Colour code: Pink Sterile, Single use, individually peel packed To comply with the latest issue of SANS 1124-1 UNIT: Box (100 per box) |  | Medi-Core Technologies | 0.17 | ULIFE | 42142523-00014 | JIAXING U-LIFE MEDICAL DEVICE TECHNOLOGY CO.,LTD | CHINA | 21 | 80 | USD | 1 BOX | YES | BOX | COMPLY pending confirmation of latex free | NO TEST REPORTS
42142523-00014 | 64 | Needle, hypodermic, long point with plastic 6%, luer lock hub Size: 18G (1,2mm) x 40mm Colour code: Pink Sterile, Single use, individually peel packed To comply with the latest issue of SANS 1124-1 UNIT: Box (100 per box) |  | Mr First Aid (Pty) Ltd | 0.17 | Yue Kang |  | Yue Kang | China | 42 | 85 | USD | 10 | Yes |  | COMPLY subject to SABS conformation of packaging | NO TEST REPORTS
42142523-00014 | 64 | Needle, hypodermic, long point with plastic 6%, luer lock hub Size: 18G (1,2mm) x 40mm Colour code: Pink Sterile, Single use, individually peel packed To comply with the latest issue of SANS 1124-1 UNIT: Box (100 per box) |  | GIBELA TRADE AND INVEST 1001 CC | 0.17368 | TARA | NEE005 | Shandong Qiaopai Group Co., Ltd | CHINA | 21 | 85 | USD | 10 | YES | each | COMPLY pending confirmation of latex free | NOT COMPLY
42142523-00014 | 64 | Needle, hypodermic, long point with plastic 6%, luer lock hub Size: 18G (1,2mm) x 40mm Colour code: Pink Sterile, Single use, individually peel packed To comply with the latest issue of SANS 1124-1 UNIT: Box (100 per box) | 1 | Vital Life Medical Distributors (PTY) LTD | 0.19 | Healthease | NHN-18-1.5 | Neomedic | China | 14 days | 80% | US Dollar | 100 boxes | Yes | 100's per box | COMPLY pending confirmation of latex free | NO TEST REPORTS
42142523-00014 | 64 | Needle, hypodermic, long point with plastic 6%, luer lock hub Size: 18G (1,2mm) x 40mm Colour code: Pink Sterile, Single use, individually peel packed To comply with the latest issue of SANS 1124-1 UNIT: Box (100 per box) |  | SUPRA HEALTHCARE JOHANNESBURG PTY LTD | 0.2 | Avacare | DN18 | HENAN AILE INDUSTRIAL CO.,LTD | CHINA | Initial 60 days, thereafter 14 | 84 | USD | 1 Box | YES | 100 pcs/box | COMPLY pending confirmation of latex free | COMPLY
42142523-00014 | 64 | Needle, hypodermic, long point with plastic 6%, luer lock hub Size: 18G (1,2mm) x 40mm Colour code: Pink Sterile, Single use, individually peel packed To comply with the latest issue of SANS 1124-1 UNIT: Box (100 per box) | N/A | SAFARMEX (PTY) LTD | 0.23 | Saclin | SSYN-10016 | Zhejiang Mediunion Healthcare Group Co. Ltd | China | 3days | 80% | US$ | 30000 Boxes | Yes | Box of 100 | COMPLY pending confirmation of latex free | NO TEST REPORT
42142523-00014 | 64 | Needle, hypodermic, long point with plastic 6%, luer lock hub Size: 18G (1,2mm) x 40mm Colour code: Pink Sterile, Single use, individually peel packed To comply with the latest issue of SANS 1124-1 UNIT: Box (100 per box) | Basic Line Offer | BIOCOR HOSPITAL SUPPLIES PTY LTD | 0.24 | Triple -D | NDL-Care 1815 | Jumla Medical | South Africa | 21 | 0 | N/A | 5 Boxes of 100 | Yes | N/A | COMPLY pending confirmation of latex free | NO TEST REPORTS
42142523-00014 | 64 | Needle, hypodermic, long point with plastic 6%, luer lock hub Size: 18G (1,2mm) x 40mm Colour code: Pink Sterile, Single use, individually peel packed To comply with the latest issue of SANS 1124-1 UNIT: Box (100 per box) | Basic Line offer | Becton Dickinson | 0.25 | BD Microlance Needle, 18G X 40mm | 304622 | Becton Dickinson | Fraga, Spain | 1 week | 76% | USD | 1 BOX | YES | 100 Units in a box | COMPLY pending confirmation of latex free | COMPLY
42142523-00014 | 64 | Needle, hypodermic, long point with plastic 6%, luer lock hub Size: 18G (1,2mm) x 40mm Colour code: Pink Sterile, Single use, individually peel packed To comply with the latest issue of SANS 1124-1 UNIT: Box (100 per box) | 1 | B. BRAUN MEDICAL (PTY) LTD | 0.25 | Sterican G18 x 1 1/2", 1.20 x 40mm Pink | 4665120 | B. Braun | Germany | 42 | 50% | Euro | 100 | Yes | Each | COMPLY pending confirmation of latex free | NOT COMPLY
42142523-00014 | 64 | Needle, hypodermic, long point with plastic 6%, luer lock hub Size: 18G (1,2mm) x 40mm Colour code: Pink Sterile, Single use, individually peel packed To comply with the latest issue of SANS 1124-1 UNIT: Box (100 per box) |  | ALMIKA TRADING PTY LTD | 0.29 | BD | 304622 | BD | USA | 14 | 0 | NO | 100 | YES | 1 | COMPLY pending confirmation of latex free | COMPLY
42142523-00014 | 64 | Needle, hypodermic, long point with plastic 6%, luer lock hub Size: 18G (1,2mm) x 40mm Colour code: Pink Sterile, Single use, individually peel packed To comply with the latest issue of SANS 1124-1 UNIT: Box (100 per box) | 1 | Biologica Pharmaceuticals Pty Ltd | 0.38 | Sol- M | 11815 | Sol- Millennium | China | 4-6 Weeks | 100% | USD | 1000 Boxes | 100% | N/A
42142523-00014 | 64 | Needle, hypodermic, long point with plastic 6%, luer lock hub Size: 18G (1,2mm) x 40mm Colour code: Pink Sterile, Single use, individually peel packed To comply with the latest issue of SANS 1124-1 UNIT: Box (100 per box) |  | JUMLA MEDICAL SUPPLIES (PTY) LTD | 18.48 | NDL-CARE | NDL-CARE-1815 | SUZHOU ACMED | CHINA | 21 DAYS | 72% | US DOLLAR | 10 BOXES | YES | NIL | COMPLY pending confirmation of latex free | NOT COMPLY
42142523-00013 | 65 | Needle, hypodermic, long point with plastic 6%, luer lock hub Size: 19G (1,1mm) x 40mm Colour code: Cream Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 1124-1 UNIT: Box (100 per box) |  | Isigidi Trading 95 cc | 0.15 | Isigidi | NE05 | Jiangsu Jichun | China | 14 days | 85% | USD | 100boxes | Yes | 100pcs/box | COMPLY pending confirmation of latex free | COMPLY
42142523-00013 | 65 | Needle, hypodermic, long point with plastic 6%, luer lock hub Size: 19G (1,1mm) x 40mm Colour code: Cream Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 1124-1 UNIT: Box (100 per box) | 1 | CENTRAL MEDICAL PTY LTD | 0.16 | LYSA | HDN1940 | Zhejian INI Medical Devices Co Ltd | China | 21 | 80 | China yuan | 100 boxes of 100 units | yes | box of 100 units | COMPLY pending confirmation of latex free | NO TEST REPORTS
42142523-00013 | 65 | Needle, hypodermic, long point with plastic 6%, luer lock hub Size: 19G (1,1mm) x 40mm Colour code: Cream Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 1124-1 UNIT: Box (100 per box) |  | Chemical World cc | 16 | SURGI PLUS | NDHD019 | TONGDA | CHINA | 3 | 85 | USD | 5000 | YES | EACH | COMPLY pending confirmation of latex free | COMPLY | Recommended | 94
42142523-00013 | 65 | Needle, hypodermic, long point with plastic 6%, luer lock hub Size: 19G (1,1mm) x 40mm Colour code: Cream Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 1124-1 UNIT: Box (100 per box) |  | Medi-Core Technologies | 0.16 | ULIFE | 42142523-00013 | JIAXING U-LIFE MEDICAL DEVICE TECHNOLOGY CO.,LTD | CHINA | 21 | 80 | USD | 1 BOX | YES | BOX | No sample submitted for evaluation | NO TEST REPORTS
42142523-00013 | 65 | Needle, hypodermic, long point with plastic 6%, luer lock hub Size: 19G (1,1mm) x 40mm Colour code: Cream Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 1124-1 UNIT: Box (100 per box) | N/A | PROMED TECHNOLOGIES (PTY) LTD | 0.16 | U-LIFE | 523013 | JIAXING | CHINA | 14 DAYS | 85% | USD | 1 UNIT | YES | UNIT | No samples submitted for evaluation | NO TEST REPORTS
42142523-00013 | 65 | Needle, hypodermic, long point with plastic 6%, luer lock hub Size: 19G (1,1mm) x 40mm Colour code: Cream Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 1124-1 UNIT: Box (100 per box) |  | Mr First Aid (Pty) Ltd | 0.17 | Yue Kang |  | Yue Kang | China | 42 | 85 | USD | 10 | Yes |  | No sample submitted for evaluation | NO TEST REPORTS
42142523-00013 | 65 | Needle, hypodermic, long point with plastic 6%, luer lock hub Size: 19G (1,1mm) x 40mm Colour code: Cream Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 1124-1 UNIT: Box (100 per box) |  | GIBELA TRADE AND INVEST 1001 CC | 0.17368 | TARA | NEE006 | Shandong Qiaopai Group Co., Ltd | CHINA | 21 | 85 | USD | 10 | YES | each | COMPLY pending confirmation of latex free | NO TEST REPORTS
42142523-00013 | 65 | Needle, hypodermic, long point with plastic 6%, luer lock hub Size: 19G (1,1mm) x 40mm Colour code: Cream Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 1124-1 UNIT: Box (100 per box) | 1 | Vital Life Medical Distributors (PTY) LTD | 0.19 | Healthease | NHN-19-1.5 | Neomedic | China | 14 days | 80% | US Dollar | 100 boxes | Yes | 100's per box | COMPLY pending confirmation of latex free | NO TEST REPORTS
42142523-00013 | 65 | Needle, hypodermic, long point with plastic 6%, luer lock hub Size: 19G (1,1mm) x 40mm Colour code: Cream Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 1124-1 UNIT: Box (100 per box) | Basic Line offer | Becton Dickinson | 0.2 | BD Microlance Needle, 19G X 40mm | 301500 | Becton Dickinson | Fraga, Spain | 1 week | 76% | USD | 1 BOX | YES | 100 Units in a box | COMPLY pending confirmation of latex free | NO TEST REPORT
42142523-00013 | 65 | Needle, hypodermic, long point with plastic 6%, luer lock hub Size: 19G (1,1mm) x 40mm Colour code: Cream Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 1124-1 UNIT: Box (100 per box) |  | SUPRA HEALTHCARE JOHANNESBURG PTY LTD | 0.2 | Avacare | DN19 | HENAN AILE INDUSTRIAL CO.,LTD | CHINA | Initial 60 days, thereafter 14 | 84 | USD | 1 Box | YES | 100 pcs/box | COMPLY pending confirmation of latex free | COMPLY
42142523-00013 | 65 | Needle, hypodermic, long point with plastic 6%, luer lock hub Size: 19G (1,1mm) x 40mm Colour code: Cream Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 1124-1 UNIT: Box (100 per box) | EACH | MSIMED CC | 0.2145 | KANG KANG | HYP117 | ZHENJIANG | CHINA | 25 | 20% | USD | 20000PCS | YES | EACH | COMPLY pending confirmation of latex free | NO TEST REPORTS
42142523-00013 | 65 | Needle, hypodermic, long point with plastic 6%, luer lock hub Size: 19G (1,1mm) x 40mm Colour code: Cream Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 1124-1 UNIT: Box (100 per box) | N/A | SAFARMEX (PTY) LTD | 0.23 | Saclin | To be confirmed | Zhejiang Mediunion Healthcare Group Co. Ltd | China | 3days | 80% | US$ | 30000 Boxes | Yes | Box of 100 | No sample submitted for evaluation | NO TEST REPORT
42142523-00013 | 65 | Needle, hypodermic, long point with plastic 6%, luer lock hub Size: 19G (1,1mm) x 40mm Colour code: Cream Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 1124-1 UNIT: Box (100 per box) |  | ALMIKA TRADING PTY LTD | 0.24 | BD | 301500 | BD | USA | 14 | 0 | NO | 100 | YES | 1 | COMPLY pending confirmation of latex free | NO TEST REPORTS
42142523-00013 | 65 | Needle, hypodermic, long point with plastic 6%, luer lock hub Size: 19G (1,1mm) x 40mm Colour code: Cream Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 1124-1 UNIT: Box (100 per box) | Basic Line Offer | BIOCOR HOSPITAL SUPPLIES PTY LTD | 0.24 | Triple -D | NDL-Care 1915 | Jumla Medical | South Africa | 21 | 0 | N/A | 5 Boxes of 100 | Yes | N/A | COMPLY pending confirmation of latex free | NO TEST REPORTS
42142523-00013 | 65 | Needle, hypodermic, long point with plastic 6%, luer lock hub Size: 19G (1,1mm) x 40mm Colour code: Cream Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 1124-1 UNIT: Box (100 per box) | 1 | B. BRAUN MEDICAL (PTY) LTD | 0.25 | Sterican G19 x 1 1/2", 1.10 x 40mm Ivory | 4657799 | B. Braun | Germany | 42 | 50% | Euro | 100 | Yes | Each | COMPLY pending confirmation of latex free | NOT COMPLY
42142523-00013 | 65 | Needle, hypodermic, long point with plastic 6%, luer lock hub Size: 19G (1,1mm) x 40mm Colour code: Cream Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 1124-1 UNIT: Box (100 per box) | 1 | Biologica Pharmaceuticals Pty Ltd | 0.38 | Sol- M | 111915 | Sol- Millennium | China | 4-6 Weeks | 100% | USD | 1000 Boxes | 100% | N/A
42142523-00013 | 65 | Needle, hypodermic, long point with plastic 6%, luer lock hub Size: 19G (1,1mm) x 40mm Colour code: Cream Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 1124-1 UNIT: Box (100 per box) |  | JUMLA MEDICAL SUPPLIES (PTY) LTD | 18.48 | NDL-CARE | NDL-CARE-1915 | SUZHOU ACMED | CHINA | 21 DAYS | 72% | US DOLLAR | 10 BOXES | YES | NIL | COMPLY pending confirmation of latex free | NOT COMPLY
42142523-00012 | 66 | Needle, hypodermic, long point with plastic 6%, luer lock hub Size: 20G (0,9mm) x 40mm Colour code: Yellow Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 1124-1 UNIT: Box (100 per box) | 1 | CENTRAL MEDICAL PTY LTD | 13.65 | LYSA | HDN2040 | Zhejian INI Medical Devices Co Ltd | China | 21 | 80 | China yuan | 100 boxes of 100 units | yes | box of 100 units | COMPLY pending confirmation of latex free | COMPLY | Recommended | 100
42142523-00012 | 66 | Needle, hypodermic, long point with plastic 6%, luer lock hub Size: 20G (0,9mm) x 40mm Colour code: Yellow Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 1124-1 UNIT: Box (100 per box) |  | Chemical World cc | 0.15 | SURGI PLUS | NDHD020 | TONGDA | CHINA | 3 | 85 | USD | 5000 | YES | EACH | COMPLY pending confirmation of latex free | NOT COMPLY
42142523-00012 | 66 | Needle, hypodermic, long point with plastic 6%, luer lock hub Size: 20G (0,9mm) x 40mm Colour code: Yellow Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 1124-1 UNIT: Box (100 per box) |  | Isigidi Trading 95 cc | 0.15 | Isigidi | NE07 | Jiangsu Jichun | China | 14 days | 85% | USD | 100boxes | Yes | 100pcs/box | COMPLY pending confirmation of latex free | COMPLY
42142523-00012 | 66 | Needle, hypodermic, long point with plastic 6%, luer lock hub Size: 20G (0,9mm) x 40mm Colour code: Yellow Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 1124-1 UNIT: Box (100 per box) |  | GIBELA TRADE AND INVEST 1001 CC | 0.15371 | TARA | NEE007 | Shandong Qiaopai Group Co., Ltd | CHINA | 21 | 85 | USD | 10 | YES | each | COMPLY pending confirmation of latex free | NOT COMPLY
42142523-00012 | 66 | Needle, hypodermic, long point with plastic 6%, luer lock hub Size: 20G (0,9mm) x 40mm Colour code: Yellow Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 1124-1 UNIT: Box (100 per box) | Basic Line offer | Becton Dickinson | 0.16 | BD Microlance Needle, 20G X 40mm | 301300 | Becton Dickinson | Fraga, Spain | 1 week | 76% | USD | 1 BOX | YES | 100 Units in a box | COMPLY pending confirmation of latex free | COMPLY
42142523-00012 | 66 | Needle, hypodermic, long point with plastic 6%, luer lock hub Size: 20G (0,9mm) x 40mm Colour code: Yellow Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 1124-1 UNIT: Box (100 per box) |  | Medi-Core Technologies | 0.16 | ULIFE | 42142523-00012 | JIAXING U-LIFE MEDICAL DEVICE TECHNOLOGY CO.,LTD | CHINA | 21 | 80 | USD | 1 BOX | YES | BOX | COMPLY pending confirmation of latex free | NO TEST REPORTS
42142523-00012 | 66 | Needle, hypodermic, long point with plastic 6%, luer lock hub Size: 20G (0,9mm) x 40mm Colour code: Yellow Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 1124-1 UNIT: Box (100 per box) | N/A | PROMED TECHNOLOGIES (PTY) LTD | 0.16 | U-LIFE | 523012 | JIAXING | CHINA | 14 DAYS | 85% | USD | 1 UNIT | YES | UNIT | COMPLY pending confirmation of latex free | COMPLY
42142523-00012 | 66 | Needle, hypodermic, long point with plastic 6%, luer lock hub Size: 20G (0,9mm) x 40mm Colour code: Yellow Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 1124-1 UNIT: Box (100 per box) |  | Mr First Aid (Pty) Ltd | 0.17 | Yue Kang |  | Yue Kang | China | 42 | 85 | USD | 10 | Yes |  | COMPLY subject to SABS conformation of packaging. | NO TEST REPORTS
42142523-00012 | 66 | Needle, hypodermic, long point with plastic 6%, luer lock hub Size: 20G (0,9mm) x 40mm Colour code: Yellow Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 1124-1 UNIT: Box (100 per box) |  | ALMIKA TRADING PTY LTD | 0.19 | BD | 301300 | BD | USA | 14 | 0 | NO | 100 | YES | 1 | COMPLY pending confirmation of latex free | COMPLY
42142523-00012 | 66 | Needle, hypodermic, long point with plastic 6%, luer lock hub Size: 20G (0,9mm) x 40mm Colour code: Yellow Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 1124-1 UNIT: Box (100 per box) | 1 | Vital Life Medical Distributors (PTY) LTD | 0.19 | Healthease | NHN-20-1.5 | Neomedic | China | 14 days | 80% | US Dollar | 100 boxes | Yes | 100's per box | COMPLY pending confirmation of latex free | NO TEST REPORTS
42142523-00012 | 66 | Needle, hypodermic, long point with plastic 6%, luer lock hub Size: 20G (0,9mm) x 40mm Colour code: Yellow Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 1124-1 UNIT: Box (100 per box) |  | SUPRA HEALTHCARE JOHANNESBURG PTY LTD | 0.2 | Avacare | DN20 | HENAN AILE INDUSTRIAL CO.,LTD | CHINA | Initial 60 days, thereafter 14 | 84 | USD | 1 Box | YES | 100 pcs/box | COMPLY pending confirmation of latex free | COMPLY
42142523-00012 | 66 | Needle, hypodermic, long point with plastic 6%, luer lock hub Size: 20G (0,9mm) x 40mm Colour code: Yellow Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 1124-1 UNIT: Box (100 per box) | EACH | MSIMED CC | 0.2145 | KANG KANG | HYP116 | ZHENJIANG | CHINA | 25 | 20% | USD | 20000PCS | YES | EACH | COMPLY pending confirmation of latex free | NO TEST REPORTS
42142523-00012 | 66 | Needle, hypodermic, long point with plastic 6%, luer lock hub Size: 20G (0,9mm) x 40mm Colour code: Yellow Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 1124-1 UNIT: Box (100 per box) |  | LECHOBA MEDICAL TECHNOLOGIES (PTY) LTD | 0.22 | MEDOJECT | CH20112 | CHIRANA INJECTA A/S | SLOVAK REPUBLIC | 21 DAYS | 75% | EURO | 12000 | YES | EACH (PACKED 100/BOX) | COMPLY pending confirmation of latex free | COMPLY
42142523-00012 | 66 | Needle, hypodermic, long point with plastic 6%, luer lock hub Size: 20G (0,9mm) x 40mm Colour code: Yellow Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 1124-1 UNIT: Box (100 per box) | N/A | SAFARMEX (PTY) LTD | 0.22 | Saclin | SSYN-10017 | Zhejiang Mediunion Healthcare Group Co. Ltd | China | 3days | 80% | US$ | 30000 Boxes | Yes | Box of 100 | COMPLY pending confirmation of latex free | NOT COMPLY
42142523-00012 | 66 | Needle, hypodermic, long point with plastic 6%, luer lock hub Size: 20G (0,9mm) x 40mm Colour code: Yellow Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 1124-1 UNIT: Box (100 per box) | Basic Line Offer | BIOCOR HOSPITAL SUPPLIES PTY LTD | 0.24 | Triple -D | NDL-Care 2015 | Jumla Medical | South Africa | 21 | 0 | N/A | 5 Boxes of 100 | Yes | N/A | COMPLY pending confirmation of latex free | NO TEST REPORTS
42142523-00012 | 66 | Needle, hypodermic, long point with plastic 6%, luer lock hub Size: 20G (0,9mm) x 40mm Colour code: Yellow Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 1124-1 UNIT: Box (100 per box) | 1 | B. BRAUN MEDICAL (PTY) LTD | 0.25 | Sterican G20 x 1 1/2", 0.9 x 40mm Yellow | 4657519 | B. Braun | Germany | 42 | 50% | Euro | 100 | Yes | Each | COMPLY pending confirmation of latex free | NOT COMPLY
42142523-00012 | 66 | Needle, hypodermic, long point with plastic 6%, luer lock hub Size: 20G (0,9mm) x 40mm Colour code: Yellow Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 1124-1 UNIT: Box (100 per box) | 1 | Biologica Pharmaceuticals Pty Ltd | 0.34 | Sol- M | 112015 | Sol- Millennium | China | 4-6 Weeks | 100% | USD | 1000 Boxes | 100% | N/A
42142523-00012 | 66 | Needle, hypodermic, long point with plastic 6%, luer lock hub Size: 20G (0,9mm) x 40mm Colour code: Yellow Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 1124-1 UNIT: Box (100 per box) |  | JUMLA MEDICAL SUPPLIES (PTY) LTD | 18.48 | NDL-CARE | NDL-CARE-2015 | SUZHOU ACMED | CHINA | 21 DAYS | 72% | US DOLLAR | 10 BOXES | YES | NIL | COMPLY pending confirmation of latex free | NOT COMPLY
42142523-00011 | 67 | Needle, hypodermic, long point with plastic 6%, luer lock hub Size: 21G (0,8mm) x 40mm Colour code: Deep green Single use, sterile, individually peel packed To comply with the latest issue of SANS 1124-1 UNIT: Box (100 per box) |  | Isigidi Trading 95 cc | 0.13 | Isigidi | NE09 | Jiangsu Jichun | China | 14 days | 85% | USD | 100boxes | Yes | 100pcs/box | COMPLY pending confirmation of latex free | COMPLY
42142523-00011 | 67 | Needle, hypodermic, long point with plastic 6%, luer lock hub Size: 21G (0,8mm) x 40mm Colour code: Deep green Single use, sterile, individually peel packed To comply with the latest issue of SANS 1124-1 UNIT: Box (100 per box) | 1 | CENTRAL MEDICAL PTY LTD | 13.65 | LYSA | HDN2140 | Zhejian INI Medical Devices Co Ltd | China | 21 | 80 | China yuan | 100 boxes of 100 units | yes | box of 100 units | COMPLY pending confirmation of latex free | COMPLY | Recommended | 93.08
42142523-00011 | 67 | Needle, hypodermic, long point with plastic 6%, luer lock hub Size: 21G (0,8mm) x 40mm Colour code: Deep green Single use, sterile, individually peel packed To comply with the latest issue of SANS 1124-1 UNIT: Box (100 per box) |  | Medi-Core Technologies | 0.14 | ULIFE | 42142523-00011 | JIAXING U-LIFE MEDICAL DEVICE TECHNOLOGY CO.,LTD | CHINA | 21 | 80 | USD | 1 BOX | YES | BOX | COMPLY pending confirmation of latex free | NO TEST REPORTS
42142523-00011 | 67 | Needle, hypodermic, long point with plastic 6%, luer lock hub Size: 21G (0,8mm) x 40mm Colour code: Deep green Single use, sterile, individually peel packed To comply with the latest issue of SANS 1124-1 UNIT: Box (100 per box) | N/A | PROMED TECHNOLOGIES (PTY) LTD | 0.14 | U-LIFE | 523011 | JIAXING | CHINA | 14 DAYS | 85% | USD | 1 UNIT | YES | UNIT | COMPLY pending confirmation of latex free | COMPLY
42142523-00011 | 67 | Needle, hypodermic, long point with plastic 6%, luer lock hub Size: 21G (0,8mm) x 40mm Colour code: Deep green Single use, sterile, individually peel packed To comply with the latest issue of SANS 1124-1 UNIT: Box (100 per box) |  | Chemical World cc | 0.15 | SURGI PLUS | NDHD021 | TONGDA | CHINA | 3 | 85 | USD | 5000 | YES | EACH | COMPLY pending confirmation of latex free | COMPLY
42142523-00011 | 67 | Needle, hypodermic, long point with plastic 6%, luer lock hub Size: 21G (0,8mm) x 40mm Colour code: Deep green Single use, sterile, individually peel packed To comply with the latest issue of SANS 1124-1 UNIT: Box (100 per box) |  | GIBELA TRADE AND INVEST 1001 CC | 0.15371 | TARA | NEE008 | Shandong Qiaopai Group Co., Ltd | CHINA | 21 | 85 | USD | 10 | YES | each | COMPLY pending confirmation of latex free | NOT COMPLY
42142523-00011 | 67 | Needle, hypodermic, long point with plastic 6%, luer lock hub Size: 21G (0,8mm) x 40mm Colour code: Deep green Single use, sterile, individually peel packed To comply with the latest issue of SANS 1124-1 UNIT: Box (100 per box) | Basic Line offer | Becton Dickinson | 0.16 | BD Microlance Needle, 21G X 40mm | 304432 | Becton Dickinson | Fraga, Spain | 1 week | 76% | USD | 1 BOX | YES | 100 Units in a box | COMPLY pending confirmation of latex free | COMPLY
42142523-00011 | 67 | Needle, hypodermic, long point with plastic 6%, luer lock hub Size: 21G (0,8mm) x 40mm Colour code: Deep green Single use, sterile, individually peel packed To comply with the latest issue of SANS 1124-1 UNIT: Box (100 per box) |  | Mr First Aid (Pty) Ltd | 0.16 | Yue Kang |  | Yue Kang | China | 42 | 85 | USD | 10 | Yes |  | COMPLY subject to SABS conformation of packaging | NO TEST REPORTS
42142523-00011 | 67 | Needle, hypodermic, long point with plastic 6%, luer lock hub Size: 21G (0,8mm) x 40mm Colour code: Deep green Single use, sterile, individually peel packed To comply with the latest issue of SANS 1124-1 UNIT: Box (100 per box) |  | ALMIKA TRADING PTY LTD | 0.19 | BD | 304432 | BD | USA | 14 | 0 | NO | 100 | YES | 1 | COMPLY pending confirmation of latex free | COMPLY
42142523-00011 | 67 | Needle, hypodermic, long point with plastic 6%, luer lock hub Size: 21G (0,8mm) x 40mm Colour code: Deep green Single use, sterile, individually peel packed To comply with the latest issue of SANS 1124-1 UNIT: Box (100 per box) | 1 | Vital Life Medical Distributors (PTY) LTD | 0.19 | Healthease | NHN-21-1.5 | Neomedic | China | 14 days | 80% | US Dollar | 100 boxes | Yes | 100's per box | COMPLY pending confirmation of latex free | NO TEST REPORTS
42142523-00011 | 67 | Needle, hypodermic, long point with plastic 6%, luer lock hub Size: 21G (0,8mm) x 40mm Colour code: Deep green Single use, sterile, individually peel packed To comply with the latest issue of SANS 1124-1 UNIT: Box (100 per box) |  | SUPRA HEALTHCARE JOHANNESBURG PTY LTD | 0.2 | Avacare | DN21 | HENAN AILE INDUSTRIAL CO.,LTD | CHINA | Initial 60 days, thereafter 14 | 84 | USD | 1 Box | YES | 100 pcs/box | COMPLY pending confirmation of latex free | NOT COMPLY
42142523-00011 | 67 | Needle, hypodermic, long point with plastic 6%, luer lock hub Size: 21G (0,8mm) x 40mm Colour code: Deep green Single use, sterile, individually peel packed To comply with the latest issue of SANS 1124-1 UNIT: Box (100 per box) | EACH | MSIMED CC | 0.2125 | KANG KANG | HYP115 | ZHENJIANG | CHINA | 25 | 20% | USD | 20000PCS | YES | EACH | COMPLY pending confirmation of latex free | NO TEST REPORTS
42142523-00011 | 67 | Needle, hypodermic, long point with plastic 6%, luer lock hub Size: 21G (0,8mm) x 40mm Colour code: Deep green Single use, sterile, individually peel packed To comply with the latest issue of SANS 1124-1 UNIT: Box (100 per box) |  | LECHOBA MEDICAL TECHNOLOGIES (PTY) LTD | 0.22 | MEDOJECT | CH21112 | CHIRANA INJECTA A/S | SLOVAK REPUBLIC | 21 DAYS | 75% | EURO | 12000 | YES | EACH (PACKED 100/BOX) | COMPLY pending confirmation of latex free | NOT COMPLY
42142523-00011 | 67 | Needle, hypodermic, long point with plastic 6%, luer lock hub Size: 21G (0,8mm) x 40mm Colour code: Deep green Single use, sterile, individually peel packed To comply with the latest issue of SANS 1124-1 UNIT: Box (100 per box) | N/A | SAFARMEX (PTY) LTD | 0.22 | Saclin | SSYN-10018 | Zhejiang Mediunion Healthcare Group Co. Ltd | China | 3days | 80% | US$ | 30000 Boxes | Yes | Box of 100 | COMPLY pending confirmation of latex free | COMPLY
42142523-00011 | 67 | Needle, hypodermic, long point with plastic 6%, luer lock hub Size: 21G (0,8mm) x 40mm Colour code: Deep green Single use, sterile, individually peel packed To comply with the latest issue of SANS 1124-1 UNIT: Box (100 per box) | Basic Line Offer | BIOCOR HOSPITAL SUPPLIES PTY LTD | 0.23 | Triple -D | NDL-Care 2115 | Jumla Medical | South Africa | 21 | 0 | N/A | 5 Boxes of 100 | Yes | N/A | COMPLY pending confirmation of latex free | NO TEST REPORTS
42142523-00011 | 67 | Needle, hypodermic, long point with plastic 6%, luer lock hub Size: 21G (0,8mm) x 40mm Colour code: Deep green Single use, sterile, individually peel packed To comply with the latest issue of SANS 1124-1 UNIT: Box (100 per box) | 1 | B. BRAUN MEDICAL (PTY) LTD | 0.25 | Sterican G21 x 1 1/2", 0.80 x 40mm Green | 4657527 | B. Braun | Germany | 42 | 50% | Euro | 100 | Yes | Each | COMPLY pending confirmation of latex free | NOT COMPLY
42142523-00011 | 67 | Needle, hypodermic, long point with plastic 6%, luer lock hub Size: 21G (0,8mm) x 40mm Colour code: Deep green Single use, sterile, individually peel packed To comply with the latest issue of SANS 1124-1 UNIT: Box (100 per box) | 1 | Biologica Pharmaceuticals Pty Ltd | 0.34 | Sol- M | 112115 | Sol- Millennium | China | 4-6 Weeks | 100% | USD | 100 Boxes | 100% | N/A
42142523-00011 | 67 | Needle, hypodermic, long point with plastic 6%, luer lock hub Size: 21G (0,8mm) x 40mm Colour code: Deep green Single use, sterile, individually peel packed To comply with the latest issue of SANS 1124-1 UNIT: Box (100 per box) |  | JUMLA MEDICAL SUPPLIES (PTY) LTD | 17.18 | NDL-CARE | NDL-CARE-2115 | SUZHOU ACMED | CHINA | 21 DAYS | 72% | US DOLLAR | 10 BOXES | YES | NIL | COMPLY pending confirmation of latex free | NOT COMPLY
42142523-00010 | 68 | Needle, hypodermic, long point with plastic 6%, luer lock hub Size: 22G (0,7mm) x 32mm Colour code: Black Sterile, Single use, individually peel packed To comply with the latest issue of SANS 1124-1 UNIT: Box (100 per box) |  | Isigidi Trading 95 cc | 0.13 | Isigidi | NE11 | Jiangsu Jichun | China | 14 days | 85% | USD | 100boxes | Yes | 100pcs/box | COMPLY pending confirmation of latex free | COMPLY
42142523-00010 | 68 | Needle, hypodermic, long point with plastic 6%, luer lock hub Size: 22G (0,7mm) x 32mm Colour code: Black Sterile, Single use, individually peel packed To comply with the latest issue of SANS 1124-1 UNIT: Box (100 per box) | 1 | CENTRAL MEDICAL PTY LTD | 13.65 | LYSA | HDN2232 | Zhejian INI Medical Devices Co Ltd | China | 21 | 80 | China yuan | 100 boxes of 100 units | yes | box of 100 units | COMPLY pending confirmation of latex free | COMPLY | Recommended | 93.08
42142523-00010 | 68 | Needle, hypodermic, long point with plastic 6%, luer lock hub Size: 22G (0,7mm) x 32mm Colour code: Black Sterile, Single use, individually peel packed To comply with the latest issue of SANS 1124-1 UNIT: Box (100 per box) |  | Chemical World cc | 0.14 | SURGI PLUS | NDHD022 | TONGDA | CHINA | 3 | 85 | USD | 5000 | YES | EACH | COMPLY pending confirmation of latex free | COMPLY
42142523-00010 | 68 | Needle, hypodermic, long point with plastic 6%, luer lock hub Size: 22G (0,7mm) x 32mm Colour code: Black Sterile, Single use, individually peel packed To comply with the latest issue of SANS 1124-1 UNIT: Box (100 per box) |  | Medi-Core Technologies | 0.14 | ULIFE | 42142523-00010 | JIAXING U-LIFE MEDICAL DEVICE TECHNOLOGY CO.,LTD | CHINA | 21 | 80 | USD | 1 BOX | YES | BOX | No sample submitted for evaluation | NO TEST REPORTS
42142523-00010 | 68 | Needle, hypodermic, long point with plastic 6%, luer lock hub Size: 22G (0,7mm) x 32mm Colour code: Black Sterile, Single use, individually peel packed To comply with the latest issue of SANS 1124-1 UNIT: Box (100 per box) | N/A | PROMED TECHNOLOGIES (PTY) LTD | 0.14 | U-LIFE | 523010 | JIAXING | CHINA | 14 DAYS | 85% | USD | 1 UNIT | YES | UNIT | COMPLY pending confirmation of latex free | NOT COMPLY
42142523-00010 | 68 | Needle, hypodermic, long point with plastic 6%, luer lock hub Size: 22G (0,7mm) x 32mm Colour code: Black Sterile, Single use, individually peel packed To comply with the latest issue of SANS 1124-1 UNIT: Box (100 per box) | Basic Line offer | Becton Dickinson | 0.15 | BD Microlance Needle, 22G X 30mm | 300900 | Becton Dickinson | Fraga, Spain | 1 week | 76% | USD | 1 BOX | YES | 100 Units in a box | COMPLY pending confirmation of latex free | NO TEST REPORT
42142523-00010 | 68 | Needle, hypodermic, long point with plastic 6%, luer lock hub Size: 22G (0,7mm) x 32mm Colour code: Black Sterile, Single use, individually peel packed To comply with the latest issue of SANS 1124-1 UNIT: Box (100 per box) |  | GIBELA TRADE AND INVEST 1001 CC | 0.15034 | TARA | NEE009 | Shandong Qiaopai Group Co., Ltd | CHINA | 21 | 85 | USD | 10 | YES | each | COMPLY pending confirmation of latex free | NOT COMPLY
42142523-00010 | 68 | Needle, hypodermic, long point with plastic 6%, luer lock hub Size: 22G (0,7mm) x 32mm Colour code: Black Sterile, Single use, individually peel packed To comply with the latest issue of SANS 1124-1 UNIT: Box (100 per box) |  | Mr First Aid (Pty) Ltd | 0.16 | Yue Kang |  | Yue Kang | China | 42 | 85 | USD | 10 | Yes |  | COMPLY subject to SABS conformation of packaging | NO TEST REPORTS
42142523-00010 | 68 | Needle, hypodermic, long point with plastic 6%, luer lock hub Size: 22G (0,7mm) x 32mm Colour code: Black Sterile, Single use, individually peel packed To comply with the latest issue of SANS 1124-1 UNIT: Box (100 per box) |  | ALMIKA TRADING PTY LTD | 0.19 | BD | 30900 | BD | USA | 14 | 0 | NO | 100 | YES | 1 | COMPLY pending confirmation of latex free | NO TEST REPORTS
42142523-00010 | 68 | Needle, hypodermic, long point with plastic 6%, luer lock hub Size: 22G (0,7mm) x 32mm Colour code: Black Sterile, Single use, individually peel packed To comply with the latest issue of SANS 1124-1 UNIT: Box (100 per box) | 1 | Vital Life Medical Distributors (PTY) LTD | 0.19 | Healthease | NHN-22-1-1/4 | Neomedic | China | 14 days | 80% | US Dollar | 100 boxes | Yes | 100's per box | COMPLY pending confirmation of latex free | NO TEST REPORTS
42142523-00010 | 68 | Needle, hypodermic, long point with plastic 6%, luer lock hub Size: 22G (0,7mm) x 32mm Colour code: Black Sterile, Single use, individually peel packed To comply with the latest issue of SANS 1124-1 UNIT: Box (100 per box) |  | SUPRA HEALTHCARE JOHANNESBURG PTY LTD | 0.2 | Avacare | DN22 | HENAN AILE INDUSTRIAL CO.,LTD | CHINA | Initial 60 days, thereafter 14 | 84 | USD | 1 Box | YES | 100 pcs/box | COMPLY pending confirmation of latex free | NO TEST REPORTS
42142523-00010 | 68 | Needle, hypodermic, long point with plastic 6%, luer lock hub Size: 22G (0,7mm) x 32mm Colour code: Black Sterile, Single use, individually peel packed To comply with the latest issue of SANS 1124-1 UNIT: Box (100 per box) | N/A | SAFARMEX (PTY) LTD | 0.21 | Saclin | SSYN-10019 | Zhejiang Mediunion Healthcare Group Co. Ltd | China | 3days | 80% | US$ | 30000 Boxes | Yes | Box of 100 | COMPLY pending confirmation of latex free | COMPLY
42142523-00010 | 68 | Needle, hypodermic, long point with plastic 6%, luer lock hub Size: 22G (0,7mm) x 32mm Colour code: Black Sterile, Single use, individually peel packed To comply with the latest issue of SANS 1124-1 UNIT: Box (100 per box) | EACH | MSIMED CC | 0.2125 | KANG KANG | HYP114 | ZHENJIANG | CHINA | 25 | 20% | USD | 20000PCS | YES | EACH | offered a 32mm instead of 38mm | NO TEST REPORTS
42142523-00010 | 68 | Needle, hypodermic, long point with plastic 6%, luer lock hub Size: 22G (0,7mm) x 32mm Colour code: Black Sterile, Single use, individually peel packed To comply with the latest issue of SANS 1124-1 UNIT: Box (100 per box) | Basic Line Offer | BIOCOR HOSPITAL SUPPLIES PTY LTD | 0.23 | Triple -D | NDL-Care 22125 | Jumla Medical | South Africa | 21 | 0 | N/A | 5 Boxes of 100 | Yes | N/A | COMPLY pending confirmation of latex free | NO TEST REPORTS
42142523-00010 | 68 | Needle, hypodermic, long point with plastic 6%, luer lock hub Size: 22G (0,7mm) x 32mm Colour code: Black Sterile, Single use, individually peel packed To comply with the latest issue of SANS 1124-1 UNIT: Box (100 per box) | 1 | B. BRAUN MEDICAL (PTY) LTD | 0.25 | Sterican G22 x 1 1/4", 0.70 x 30mm Black | 4657624 | B. Braun | Germany | 42 | 50% | Euro | 100 | Yes | Each | COMPLY pending confirmation of latex free | NOT COMPLY
42142523-00010 | 68 | Needle, hypodermic, long point with plastic 6%, luer lock hub Size: 22G (0,7mm) x 32mm Colour code: Black Sterile, Single use, individually peel packed To comply with the latest issue of SANS 1124-1 UNIT: Box (100 per box) | 1 | Biologica Pharmaceuticals Pty Ltd | 0.34 | Sol- M | 1122125 | Sol- Millennium | China | 4-6 Weeks | 100% | USD | 1000 Boxes | 100% | N/A
42142523-00010 | 68 | Needle, hypodermic, long point with plastic 6%, luer lock hub Size: 22G (0,7mm) x 32mm Colour code: Black Sterile, Single use, individually peel packed To comply with the latest issue of SANS 1124-1 UNIT: Box (100 per box) |  | JUMLA MEDICAL SUPPLIES (PTY) LTD | 17.18 | NDL-CARE | NDL-CARE-22125 | SUZHOU ACMED | CHINA | 21 DAYS | 72% | US DOLLAR | 10 BOXES | YES | NIL | COMPLY pending confirmation of latex free | NO TEST REPORTS
42142523-00009 | 69 | Needle, hypodermic, long point with plastic 6%, luer lock hub Size: 23G (0,6mm) x 25mm Colour code: Deep blue Sterile, Single use, Individually packed To comply with the latest issue of SANS 1124-1 UNIT: Box (100 per box) |  | Isigidi Trading 95 cc | 0.13 | Isigidi | NE13 | Jiangsu Jichun | China | 14 days | 85% | USD | 100boxes | Yes | 100pcs/box | COMPLY pending confirmation of latex free | COMPLY
42142523-00009 | 69 | Needle, hypodermic, long point with plastic 6%, luer lock hub Size: 23G (0,6mm) x 25mm Colour code: Deep blue Sterile, Single use, Individually packed To comply with the latest issue of SANS 1124-1 UNIT: Box (100 per box) | 1 | CENTRAL MEDICAL PTY LTD | 13.65 | LYSA | HDN2325 | Zhejian INI Medical Devices Co Ltd | China | 21 | 80 | China yuan | 100 boxes of 100 units | yes | box of 100 units | COMPLY pending confirmation of latex free | COMPLY | Recommended | 93.08
42142523-00009 | 69 | Needle, hypodermic, long point with plastic 6%, luer lock hub Size: 23G (0,6mm) x 25mm Colour code: Deep blue Sterile, Single use, Individually packed To comply with the latest issue of SANS 1124-1 UNIT: Box (100 per box) |  | Chemical World cc | 0.14 | SURGI PLUS | NDHD023 | TONGDA | CHINA | 3 | 85 | USD | 5000 | YES | EACH | COMPLY pending confirmation of latex free | COMPLY
42142523-00009 | 69 | Needle, hypodermic, long point with plastic 6%, luer lock hub Size: 23G (0,6mm) x 25mm Colour code: Deep blue Sterile, Single use, Individually packed To comply with the latest issue of SANS 1124-1 UNIT: Box (100 per box) |  | Medi-Core Technologies | 0.14 | ULIFE | 42142523-00009 | JIAXING U-LIFE MEDICAL DEVICE TECHNOLOGY CO.,LTD | CHINA | 21 | 80 | USD | 1 BOX | YES | BOX | COMPLY pending confirmation of latex free | NO TEST REPORTS
42142523-00009 | 69 | Needle, hypodermic, long point with plastic 6%, luer lock hub Size: 23G (0,6mm) x 25mm Colour code: Deep blue Sterile, Single use, Individually packed To comply with the latest issue of SANS 1124-1 UNIT: Box (100 per box) | N/A | PROMED TECHNOLOGIES (PTY) LTD | 0.14 | U-LIFE |  | JIAXING | CHINA | 15 DAYS | 85% | USD | 2 UNIT |  |  | COMPLY pending confirmation of latex free | COMPLY
42142523-00009 | 69 | Needle, hypodermic, long point with plastic 6%, luer lock hub Size: 23G (0,6mm) x 25mm Colour code: Deep blue Sterile, Single use, Individually packed To comply with the latest issue of SANS 1124-1 UNIT: Box (100 per box) | Basic Line offer | Becton Dickinson | 0.15 | BD Microlance Needle, 23G X 25mm | 300800 | Becton Dickinson | Fraga, Spain | 1 week | 76% | USD | 1 BOX | YES | 100 Units in a box | COMPLY pending confirmation of latex free | NO TEST REPORT
42142523-00009 | 69 | Needle, hypodermic, long point with plastic 6%, luer lock hub Size: 23G (0,6mm) x 25mm Colour code: Deep blue Sterile, Single use, Individually packed To comply with the latest issue of SANS 1124-1 UNIT: Box (100 per box) |  | GIBELA TRADE AND INVEST 1001 CC | 0.15034 | TARA | NEE010 | Shandong Qiaopai Group Co., Ltd | CHINA | 21 | 85 | USD | 10 | YES | each | COMPLY pending confirmation of latex free | NOT COMPLY
42142523-00009 | 69 | Needle, hypodermic, long point with plastic 6%, luer lock hub Size: 23G (0,6mm) x 25mm Colour code: Deep blue Sterile, Single use, Individually packed To comply with the latest issue of SANS 1124-1 UNIT: Box (100 per box) |  | Mr First Aid (Pty) Ltd | 0.16 | Yue Kang |  | Yue Kang | China | 42 | 85 | USD | 10 | Yes |  | COMPLY subject to SABS conformation of packaging | NO TEST REPORTS
42142523-00009 | 69 | Needle, hypodermic, long point with plastic 6%, luer lock hub Size: 23G (0,6mm) x 25mm Colour code: Deep blue Sterile, Single use, Individually packed To comply with the latest issue of SANS 1124-1 UNIT: Box (100 per box) |  | ALMIKA TRADING PTY LTD | 0.19 | BD | 300800 | BD | USA | 14 | 0 | NO | 100 | YES | 1 | COMPLY pending confirmation of latex free | NO TEST REPORTS
42142523-00009 | 69 | Needle, hypodermic, long point with plastic 6%, luer lock hub Size: 23G (0,6mm) x 25mm Colour code: Deep blue Sterile, Single use, Individually packed To comply with the latest issue of SANS 1124-1 UNIT: Box (100 per box) | 1 | Vital Life Medical Distributors (PTY) LTD | 0.19 | Healthease | NHN-23-1 | Neomedic | China | 14 days | 80% | US Dollar | 100 boxes | Yes | 100's per box | COMPLY pending confirmation of latex free | NO TEST REPORTS
42142523-00009 | 69 | Needle, hypodermic, long point with plastic 6%, luer lock hub Size: 23G (0,6mm) x 25mm Colour code: Deep blue Sterile, Single use, Individually packed To comply with the latest issue of SANS 1124-1 UNIT: Box (100 per box) |  | SUPRA HEALTHCARE JOHANNESBURG PTY LTD | 0.2 | Avacare | DN23 | HENAN AILE INDUSTRIAL CO.,LTD | CHINA | Initial 60 days, thereafter 14 | 84 | USD | 1 Box | YES | 100 pcs/box | Offered 38mm instead of 25mm | NO TEST REPORTS
42142523-00009 | 69 | Needle, hypodermic, long point with plastic 6%, luer lock hub Size: 23G (0,6mm) x 25mm Colour code: Deep blue Sterile, Single use, Individually packed To comply with the latest issue of SANS 1124-1 UNIT: Box (100 per box) | N/A | SAFARMEX (PTY) LTD | 0.21 | Saclin | SSYN-10020 | Zhejiang Mediunion Healthcare Group Co. Ltd | China | 3days | 80% | US$ | 30000 Boxes | Yes | Box of 100 | COMPLY pending confirmation of latex free | COMPLY
42142523-00009 | 69 | Needle, hypodermic, long point with plastic 6%, luer lock hub Size: 23G (0,6mm) x 25mm Colour code: Deep blue Sterile, Single use, Individually packed To comply with the latest issue of SANS 1124-1 UNIT: Box (100 per box) | EACH | MSIMED CC | 0.2125 | KANG KANG | HYP113 | ZHENJIANG | CHINA | 25 | 20% | USD | 20000PCS | YES | EACH | Sample item number incorrectly labelled | NO TEST REPORTS
42142523-00009 | 69 | Needle, hypodermic, long point with plastic 6%, luer lock hub Size: 23G (0,6mm) x 25mm Colour code: Deep blue Sterile, Single use, Individually packed To comply with the latest issue of SANS 1124-1 UNIT: Box (100 per box) | Basic Line Offer | BIOCOR HOSPITAL SUPPLIES PTY LTD | 0.23 | Triple -D | NDL-Care 231 | Jumla Medical | South Africa | 60 | 0 | N/A | 5 Boxes of 100 | Yes | N/A | COMPLY pending confirmation of latex free | NO TEST REPORTS
42142523-00009 | 69 | Needle, hypodermic, long point with plastic 6%, luer lock hub Size: 23G (0,6mm) x 25mm Colour code: Deep blue Sterile, Single use, Individually packed To comply with the latest issue of SANS 1124-1 UNIT: Box (100 per box) | 1 | B. BRAUN MEDICAL (PTY) LTD | 0.25 | Sterican G23 x 1", 0.60 x 25mm Blue | 4657667 | B. Braun | Germany | 42 | 50% | Euro | 100 | Yes | Each | COMPLY pending confirmation of latex free | NOT COMPLY
42142523-00009 | 69 | Needle, hypodermic, long point with plastic 6%, luer lock hub Size: 23G (0,6mm) x 25mm Colour code: Deep blue Sterile, Single use, Individually packed To comply with the latest issue of SANS 1124-1 UNIT: Box (100 per box) | 1 | Biologica Pharmaceuticals Pty Ltd | 0.34 | Sol- M | 112310 | Sol- Millennium | China | 4-6 Weeks | 100% | USD | 1000 Boxes | 100% | N/A
42142523-00009 | 69 | Needle, hypodermic, long point with plastic 6%, luer lock hub Size: 23G (0,6mm) x 25mm Colour code: Deep blue Sterile, Single use, Individually packed To comply with the latest issue of SANS 1124-1 UNIT: Box (100 per box) |  | JUMLA MEDICAL SUPPLIES (PTY) LTD | 17.18 | NDL-CARE | NDL-CARE-231 | SUZHOU ACMED | CHINA | 21 DAYS | 72% | US DOLLAR | 10 BOXES | YES | NIL | COMPLY pending confirmation of latex free | NOT COMPLY
42142523-00008 | 70 | Needle, hypodermic, long point with plastic 6%, luer lock hub Size:25G (0,50mm) x 16mm Colour code: Orange Sterile, Single use, Individually packed To comply with the latest issue of SANS 1124-1 UNIT: Box (100 per box) |  | Chemical World cc | 0.13 | SURGI PLUS | NDHD025 | TONGDA | CHINA | 3 | 85 | USD | 5000 | YES | EACH | COMPLY pending confirmation of latex free | NOT COMPLY
42142523-00008 | 70 | Needle, hypodermic, long point with plastic 6%, luer lock hub Size:25G (0,50mm) x 16mm Colour code: Orange Sterile, Single use, Individually packed To comply with the latest issue of SANS 1124-1 UNIT: Box (100 per box) |  | Medi-Core Technologies | 0.13 | ULIFE | 42142523-00008 | JIAXING U-LIFE MEDICAL DEVICE TECHNOLOGY CO.,LTD | CHINA | 21 | 80 | USD | 1 BOX | YES | BOX | No sample submitted for evaluation | NO TEST REPORTS
42142523-00008 | 70 | Needle, hypodermic, long point with plastic 6%, luer lock hub Size:25G (0,50mm) x 16mm Colour code: Orange Sterile, Single use, Individually packed To comply with the latest issue of SANS 1124-1 UNIT: Box (100 per box) | 1 | CENTRAL MEDICAL PTY LTD | 13.65 | LYSA | HDN2516 | Zhejian INI Medical Devices Co Ltd | China | 21 | 80 | China yuan | 100 boxes of 100 units | yes | box of 100 units | COMPLY pending confirmation of latex free | COMPLY | Recommended | 93.08
42142523-00008 | 70 | Needle, hypodermic, long point with plastic 6%, luer lock hub Size:25G (0,50mm) x 16mm Colour code: Orange Sterile, Single use, Individually packed To comply with the latest issue of SANS 1124-1 UNIT: Box (100 per box) |  | Isigidi Trading 95 cc | 0.14 | Isigidi | NE25 | Jiangsu Jichun | China | 14 days | 85% | USD | 100boxes | Yes | 100pcs/box | COMPLY pending confirmation of latex free | COMPLY
42142523-00008 | 70 | Needle, hypodermic, long point with plastic 6%, luer lock hub Size:25G (0,50mm) x 16mm Colour code: Orange Sterile, Single use, Individually packed To comply with the latest issue of SANS 1124-1 UNIT: Box (100 per box) | N/A | PROMED TECHNOLOGIES (PTY) LTD | 0.14 | U-LIFE | 523008 | JIAXING | CHINA | 14 DAYS | 85% | USD | 1 UNIT | YES | UNIT | COMPLY pending confirmation of latex free | COMPLY
42142523-00008 | 70 | Needle, hypodermic, long point with plastic 6%, luer lock hub Size:25G (0,50mm) x 16mm Colour code: Orange Sterile, Single use, Individually packed To comply with the latest issue of SANS 1124-1 UNIT: Box (100 per box) | Basic Line offer | Becton Dickinson | 0.15 | BD Microlance Needle, 25G X 16mm | 300600 | Becton Dickinson | Fraga, Spain | 1 week | 76% | USD | 1 BOX | YES | 100 Units in a box | COMPLY pending confirmation of latex free | COMPLY
42142523-00008 | 70 | Needle, hypodermic, long point with plastic 6%, luer lock hub Size:25G (0,50mm) x 16mm Colour code: Orange Sterile, Single use, Individually packed To comply with the latest issue of SANS 1124-1 UNIT: Box (100 per box) |  | Mr First Aid (Pty) Ltd | 0.15 | Yue Kang |  | Yue Kang | China | 42 | 85 | USD | 10 | Yes |  | No sample submitted for evaluation. | NO TEST REPORTS
42142523-00008 | 70 | Needle, hypodermic, long point with plastic 6%, luer lock hub Size:25G (0,50mm) x 16mm Colour code: Orange Sterile, Single use, Individually packed To comply with the latest issue of SANS 1124-1 UNIT: Box (100 per box) |  | GIBELA TRADE AND INVEST 1001 CC | 0.15034 | TARA | NEE011 | Shandong Qiaopai Group Co., Ltd | CHINA | 21 | 85 | USD | 10 | YES | each | COMPLY pending confirmation of latex free | NOT COMPLY
42142523-00008 | 70 | Needle, hypodermic, long point with plastic 6%, luer lock hub Size:25G (0,50mm) x 16mm Colour code: Orange Sterile, Single use, Individually packed To comply with the latest issue of SANS 1124-1 UNIT: Box (100 per box) |  | ALMIKA TRADING PTY LTD | 0.19 | BD | 300600 | BD | USA | 14 | 0 | NO | 100 | YES | 1 | COMPLY pending confirmation of latex free | COMPLY
42142523-00008 | 70 | Needle, hypodermic, long point with plastic 6%, luer lock hub Size:25G (0,50mm) x 16mm Colour code: Orange Sterile, Single use, Individually packed To comply with the latest issue of SANS 1124-1 UNIT: Box (100 per box) | 1 | Vital Life Medical Distributors (PTY) LTD | 0.19 | Healthease | NHN-25-5/8 | Neomedic | China | 14 days | 80% | US Dollar | 100 boxes | Yes | 100's per box | COMPLY pending confirmation of latex free | NO TEST REPORTS
42142523-00008 | 70 | Needle, hypodermic, long point with plastic 6%, luer lock hub Size:25G (0,50mm) x 16mm Colour code: Orange Sterile, Single use, Individually packed To comply with the latest issue of SANS 1124-1 UNIT: Box (100 per box) |  | LECHOBA MEDICAL TECHNOLOGIES (PTY) LTD | 0.2 | MEDOJECT | CH25058 | CHIRANA INJECTA A/S | SLOVAK REPUBLIC | 21 DAYS | 75% | EURO | 12000 | YES | EACH (PACKED 100/BOX) | COMPLY pending confirmation of latex free | COMPLY
42142523-00008 | 70 | Needle, hypodermic, long point with plastic 6%, luer lock hub Size:25G (0,50mm) x 16mm Colour code: Orange Sterile, Single use, Individually packed To comply with the latest issue of SANS 1124-1 UNIT: Box (100 per box) |  | SUPRA HEALTHCARE JOHANNESBURG PTY LTD | 0.2 | Avacare | DN25 | HENAN AILE INDUSTRIAL CO.,LTD | CHINA | Initial 60 days, thereafter 14 | 84 | USD | 1 Box | YES | 100 pcs/box | COMPLY pending confirmation of latex free | NO TEST REPORTS
42142523-00008 | 70 | Needle, hypodermic, long point with plastic 6%, luer lock hub Size:25G (0,50mm) x 16mm Colour code: Orange Sterile, Single use, Individually packed To comply with the latest issue of SANS 1124-1 UNIT: Box (100 per box) | N/A | SAFARMEX (PTY) LTD | 0.21 | Saclin | SSYN-10021 | Zhejiang Mediunion Healthcare Group Co. Ltd | China | 3days | 80% | US$ | 30000 Boxes | Yes | Box of 100 | COMPLY pending confirmation of latex free | COMPLY
42142523-00008 | 70 | Needle, hypodermic, long point with plastic 6%, luer lock hub Size:25G (0,50mm) x 16mm Colour code: Orange Sterile, Single use, Individually packed To comply with the latest issue of SANS 1124-1 UNIT: Box (100 per box) | EACH | MSIMED CC | 0.22 | KANG KANG | HYP112 | ZHENJIANG | CHINA | 25 | 20% | USD | 20000PCS | YES | EACH | COMPLY pending confirmation of latex free | NO TEST REPORTS
42142523-00008 | 70 | Needle, hypodermic, long point with plastic 6%, luer lock hub Size:25G (0,50mm) x 16mm Colour code: Orange Sterile, Single use, Individually packed To comply with the latest issue of SANS 1124-1 UNIT: Box (100 per box) | Basic Line Offer | BIOCOR HOSPITAL SUPPLIES PTY LTD | 0.23 | Triple -D | NDL-Care 2558 | Jumla Medical | South Africa | 60 | 0 | N/A | 5 Boxes of 100 | Yes | N/A | COMPLY pending confirmation of latex free | NO TEST REPORTS
42142523-00008 | 70 | Needle, hypodermic, long point with plastic 6%, luer lock hub Size:25G (0,50mm) x 16mm Colour code: Orange Sterile, Single use, Individually packed To comply with the latest issue of SANS 1124-1 UNIT: Box (100 per box) | 1 | B. BRAUN MEDICAL (PTY) LTD | 0.25 | Sterican G25 x 5/8", 0.5 x 16mm Orange | 4657853 | B. Braun | Germany | 42 | 50% | Euro | 100 | Yes | Each | COMPLY pending confirmation of latex free | NOT COMPLY
42142523-00008 | 70 | Needle, hypodermic, long point with plastic 6%, luer lock hub Size:25G (0,50mm) x 16mm Colour code: Orange Sterile, Single use, Individually packed To comply with the latest issue of SANS 1124-1 UNIT: Box (100 per box) | 1 | Biologica Pharmaceuticals Pty Ltd | 0.34 | Sol- M | 112506 | Sol- Millennium | China | 4-6 Weeks | 100% | USD | 1000 Boxes | 100% | N/A
42142523-00008 | 70 | Needle, hypodermic, long point with plastic 6%, luer lock hub Size:25G (0,50mm) x 16mm Colour code: Orange Sterile, Single use, Individually packed To comply with the latest issue of SANS 1124-1 UNIT: Box (100 per box) |  | JUMLA MEDICAL SUPPLIES (PTY) LTD | 17.18 | NDL-CARE | NDL-CARE-2558 | SUZHOU ACMED | CHINA | 21 DAYS | 72% | US DOLLAR | 10 BOXES | YES | NIL | COMPLY pending confirmation of latex free | NOT COMPLY
42142523-00007 | 71 | Needle, hypodermic, long point with plastic 6%, luer lock hub Size: 26G (0.45mm) x 16mm Colour code: Brown Sterile, Single use, Individually packed To comply with the latest issue of SANS 1124-1 UNIT: Box (100 per box) |  | Chemical World cc | 0.13 | SURGI PLUS | NDHD026 | TONGDA | CHINA | 3 | 85 | USD | 5000 | YES | EACH | COMPLY pending confirmation of latex free | NOT COMPLY
42142523-00007 | 71 | Needle, hypodermic, long point with plastic 6%, luer lock hub Size: 26G (0.45mm) x 16mm Colour code: Brown Sterile, Single use, Individually packed To comply with the latest issue of SANS 1124-1 UNIT: Box (100 per box) |  | Medi-Core Technologies | 0.13 | ULIFE | 42142523-00007 | JIAXING U-LIFE MEDICAL DEVICE TECHNOLOGY CO.,LTD | CHINA | 21 | 80 | USD | 1 BOX | YES | BOX | COMPLY pending confirmation of latex free | NO TEST REPORTS
42142523-00007 | 71 | Needle, hypodermic, long point with plastic 6%, luer lock hub Size: 26G (0.45mm) x 16mm Colour code: Brown Sterile, Single use, Individually packed To comply with the latest issue of SANS 1124-1 UNIT: Box (100 per box) | 1 | CENTRAL MEDICAL PTY LTD | 13.65 | LYSA | HDN2616 | Zhejian INI Medical Devices Co Ltd | China | 21 | 80 | China yuan | 100 boxes of 100 units | yes | box of 100 units | COMPLY pending confirmation of latex free | COMPLY | Recommended | 93.08
42142523-00007 | 71 | Needle, hypodermic, long point with plastic 6%, luer lock hub Size: 26G (0.45mm) x 16mm Colour code: Brown Sterile, Single use, Individually packed To comply with the latest issue of SANS 1124-1 UNIT: Box (100 per box) |  | Isigidi Trading 95 cc | 0.14 | Isigidi | NE19 | Jiangsu Jichun | China | 14 days | 85% | USD | 100boxes | Yes | 100pcs/box | COMPLY pending confirmation of latex free | COMPLY
42142523-00007 | 71 | Needle, hypodermic, long point with plastic 6%, luer lock hub Size: 26G (0.45mm) x 16mm Colour code: Brown Sterile, Single use, Individually packed To comply with the latest issue of SANS 1124-1 UNIT: Box (100 per box) | N/A | PROMED TECHNOLOGIES (PTY) LTD | 0.14 | U-LIFE | 523007 | JIAXING | CHINA | 14 DAYS | 85% | USD | 1 UNIT | YES | UNIT | COMPLY pending confirmation of latex free | COMPLY
42142523-00007 | 71 | Needle, hypodermic, long point with plastic 6%, luer lock hub Size: 26G (0.45mm) x 16mm Colour code: Brown Sterile, Single use, Individually packed To comply with the latest issue of SANS 1124-1 UNIT: Box (100 per box) | Basic Line offer | Becton Dickinson | 0.15 | BD Microlance Needle, 26G X 13mm | 303800 | Becton Dickinson | Fraga, Spain | 1 week | 76% | USD | 1 BOX | YES | 100 Units in a box | Offered 13mm instead of 16mm | COMPLY
42142523-00007 | 71 | Needle, hypodermic, long point with plastic 6%, luer lock hub Size: 26G (0.45mm) x 16mm Colour code: Brown Sterile, Single use, Individually packed To comply with the latest issue of SANS 1124-1 UNIT: Box (100 per box) |  | Mr First Aid (Pty) Ltd | 0.15 | Yue Kang |  | Yue Kang | China | 42 | 85 | USD | 10 | Yes |  | No sample submitted for evaluation. | NO TEST REPORTS
42142523-00007 | 71 | Needle, hypodermic, long point with plastic 6%, luer lock hub Size: 26G (0.45mm) x 16mm Colour code: Brown Sterile, Single use, Individually packed To comply with the latest issue of SANS 1124-1 UNIT: Box (100 per box) |  | GIBELA TRADE AND INVEST 1001 CC | 0.15034 | TARA | NEE012 | Shandong Qiaopai Group Co., Ltd | CHINA | 21 | 85 | USD | 10 | YES | each | COMPLY pending confirmation of latex free | NOT COMPLY
42142523-00007 | 71 | Needle, hypodermic, long point with plastic 6%, luer lock hub Size: 26G (0.45mm) x 16mm Colour code: Brown Sterile, Single use, Individually packed To comply with the latest issue of SANS 1124-1 UNIT: Box (100 per box) |  | ALMIKA TRADING PTY LTD | 0.19 | BD | 303800 | BD | USA | 14 | 0 | NO | 100 | YES | 1 | offered a length of 13mm instead of 16mm | COMPLY
42142523-00007 | 71 | Needle, hypodermic, long point with plastic 6%, luer lock hub Size: 26G (0.45mm) x 16mm Colour code: Brown Sterile, Single use, Individually packed To comply with the latest issue of SANS 1124-1 UNIT: Box (100 per box) | 1 | Vital Life Medical Distributors (PTY) LTD | 0.19 | Healthease | NHN-26-1/2 | Neomedic | China | 14 days | 80% | US Dollar | 100 boxes | No | Needle 0.45mm x 13mm 100's per box | Offered 13mm instead of 16mm | NO TEST REPORTS
42142523-00007 | 71 | Needle, hypodermic, long point with plastic 6%, luer lock hub Size: 26G (0.45mm) x 16mm Colour code: Brown Sterile, Single use, Individually packed To comply with the latest issue of SANS 1124-1 UNIT: Box (100 per box) |  | SUPRA HEALTHCARE JOHANNESBURG PTY LTD | 0.2 | Avacare | DN26 | HENAN AILE INDUSTRIAL CO.,LTD | CHINA | Initial 60 days, thereafter 14 | 84 | USD | 1 Box | YES | 100 pcs/box | COMPLY pending confirmation of latex free | NO TEST REPORTS
42142523-00007 | 71 | Needle, hypodermic, long point with plastic 6%, luer lock hub Size: 26G (0.45mm) x 16mm Colour code: Brown Sterile, Single use, Individually packed To comply with the latest issue of SANS 1124-1 UNIT: Box (100 per box) | N/A | SAFARMEX (PTY) LTD | 0.21 | Saclin | SSYN-10022 | Zhejiang Mediunion Healthcare Group Co. Ltd | China | 3days | 80% | US$ | 30000 Boxes | Yes | Box of 100 | COMPLY pending confirmation of latex free | COMPLY
42142523-00007 | 71 | Needle, hypodermic, long point with plastic 6%, luer lock hub Size: 26G (0.45mm) x 16mm Colour code: Brown Sterile, Single use, Individually packed To comply with the latest issue of SANS 1124-1 UNIT: Box (100 per box) | EACH | MSIMED CC | 0.22 | KANG KANG | HYP111 | ZHENJIANG | CHINA | 25 | 20% | USD | 20000PCS | YES | EACH | offered a 13mm instead of 16mm | NO TEST REPORTS
42142523-00007 | 71 | Needle, hypodermic, long point with plastic 6%, luer lock hub Size: 26G (0.45mm) x 16mm Colour code: Brown Sterile, Single use, Individually packed To comply with the latest issue of SANS 1124-1 UNIT: Box (100 per box) | Basic Line Offer | BIOCOR HOSPITAL SUPPLIES PTY LTD | 0.23 | Triple -D | NDL-Care 2658 | Jumla Medical | South Africa | 21 | 0 | N/A | 5 Boxes of 100 | Yes | N/A | COMPLY pending confirmation of latex free | NO TEST REPORTS
42142523-00007 | 71 | Needle, hypodermic, long point with plastic 6%, luer lock hub Size: 26G (0.45mm) x 16mm Colour code: Brown Sterile, Single use, Individually packed To comply with the latest issue of SANS 1124-1 UNIT: Box (100 per box) | 1 | B. BRAUN MEDICAL (PTY) LTD | 0.25 | Sterican G26 x 1/2", 0.45 x 12mm Brown | 4665457 | B. Braun | Germany | 42 | 50% | Euro | 100 | Yes | Each | COMPLY pending confirmation of latex free | NOT COMPLY
42142523-00007 | 71 | Needle, hypodermic, long point with plastic 6%, luer lock hub Size: 26G (0.45mm) x 16mm Colour code: Brown Sterile, Single use, Individually packed To comply with the latest issue of SANS 1124-1 UNIT: Box (100 per box) |  | JUMLA MEDICAL SUPPLIES (PTY) LTD | 17.18 | NDL-CARE | NDL-CARE-2658 | SUZHOU ACMED | CHINA | 21 DAYS | 72% | US DOLLAR | 10 BOXES | YES | NIL | COMPLY pending confirmation of latex free
42142523-00006 | 72 | Needle, hypodermic, long point with plastic 6%, luer lock hub Size: 29G (0.33mm) x 8mm Colour code: Red Sterile, Single use, Individually packed To comply with the latest issue of SANS 1124-1 UNIT: Box (100 per box) |  | Chemical World cc | 0.18 | SURGI PLUS | NDHD029 | TONGDA | CHINA | 3 | 85 | USD | 5000 | YES | EACH | No sample submitted for evaluation | NO TEST REPORTS
42142523-00006 | 72 | Needle, hypodermic, long point with plastic 6%, luer lock hub Size: 29G (0.33mm) x 8mm Colour code: Red Sterile, Single use, Individually packed To comply with the latest issue of SANS 1124-1 UNIT: Box (100 per box) |  | Medi-Core Technologies | 0.2 | ULIFE | 42142523-00006 | JIAXING U-LIFE MEDICAL DEVICE TECHNOLOGY CO.,LTD | CHINA | 21 | 80 | USD | 1 BOX | YES | BOX | No sample submitted for evaluation | NO TEST REPORTS
42142523-00006 | 72 | Needle, hypodermic, long point with plastic 6%, luer lock hub Size: 29G (0.33mm) x 8mm Colour code: Red Sterile, Single use, Individually packed To comply with the latest issue of SANS 1124-1 UNIT: Box (100 per box) | N/A | PROMED TECHNOLOGIES (PTY) LTD | 0.2 | U-LIFE | 523006 | JIAXING | CHINA | 14 DAYS | 85% | USD | 1 BOX | YES | BOX | No samples submitted for evaluation | NO TEST REPORTS
42142523-00006 | 72 | Needle, hypodermic, long point with plastic 6%, luer lock hub Size: 29G (0.33mm) x 8mm Colour code: Red Sterile, Single use, Individually packed To comply with the latest issue of SANS 1124-1 UNIT: Box (100 per box) |  | Mr First Aid (Pty) Ltd | 0.21 | Yue Kang |  | Yue Kang | China | 42 | 85 | USD | 10 | Yes |  | No sample submitted for evaluation. | NO TEST REPORTS
42142523-00006 | 72 | Needle, hypodermic, long point with plastic 6%, luer lock hub Size: 29G (0.33mm) x 8mm Colour code: Red Sterile, Single use, Individually packed To comply with the latest issue of SANS 1124-1 UNIT: Box (100 per box) | Basic Line Offer | BIOCOR HOSPITAL SUPPLIES PTY LTD | 0.22 | Medline | N/A | Yangzhong Medline Industry Co.Ltd | China | 21 | 0 | N/A | 5 Boxes of 100 | Yes | N/A | No sample submitted for evaluation | NO TEST REPORTS
42142523-00006 | 72 | Needle, hypodermic, long point with plastic 6%, luer lock hub Size: 29G (0.33mm) x 8mm Colour code: Red Sterile, Single use, Individually packed To comply with the latest issue of SANS 1124-1 UNIT: Box (100 per box) |  | GIBELA TRADE AND INVEST 1001 CC | 0.22439 | TARA | NEE013 | Shandong Qiaopai Group Co., Ltd | CHINA | 21 | 85 | USD | 10 | YES | each | No sample submitted for evaluation | NO TEST REPORTS
42142523-00006 | 72 | Needle, hypodermic, long point with plastic 6%, luer lock hub Size: 29G (0.33mm) x 8mm Colour code: Red Sterile, Single use, Individually packed To comply with the latest issue of SANS 1124-1 UNIT: Box (100 per box) | 1 | CENTRAL MEDICAL PTY LTD | 0.26 | LYSA | HDN2908 | Zhejian INI Medical Devices Co Ltd | China | 21 | 80 | China yuan | 100 boxes of 100 units | yes | box of 100 units | COMPLY pending confirmation of latex free | NO TEST REPORTS
42142523-00006 | 72 | Needle, hypodermic, long point with plastic 6%, luer lock hub Size: 29G (0.33mm) x 8mm Colour code: Red Sterile, Single use, Individually packed To comply with the latest issue of SANS 1124-1 UNIT: Box (100 per box) |  | Isigidi Trading 95 cc | 25.3 | Isigidi | NE34 | Jiangsu Jichun | China | 14 days | 85% | USD | 100boxes | Yes | 100pcs/box | COMPLY pending confirmation of latex free | COMPLY | Recommended | 91
42142523-00006 | 72 | Needle, hypodermic, long point with plastic 6%, luer lock hub Size: 29G (0.33mm) x 8mm Colour code: Red Sterile, Single use, Individually packed To comply with the latest issue of SANS 1124-1 UNIT: Box (100 per box) | N/A | SAFARMEX (PTY) LTD | 0.3 | Saclin | To be confirmed | Zhejiang Mediunion Healthcare Group Co. Ltd | China | 3days | 80% | US$ | 30000 Boxes | Yes | Box of 100 | No sample submitted for evaluation | NO TEST REPORT
42142523-00006 | 72 | Needle, hypodermic, long point with plastic 6%, luer lock hub Size: 29G (0.33mm) x 8mm Colour code: Red Sterile, Single use, Individually packed To comply with the latest issue of SANS 1124-1 UNIT: Box (100 per box) |  | SUPRA HEALTHCARE JOHANNESBURG PTY LTD | 0.46 | Avacare | DN29 | HENAN AILE INDUSTRIAL CO.,LTD | CHINA | Initial 60 days, thereafter 14 | 84 | USD | 1 Box | YES | 100 pcs/box | COMPLY pending confirmation of latex free | NO TEST REPORTS
42142523-00018 | 73 | Needle, hypodermic, with inherent safety device to prevent needle stick injury, with plastic 6% luer lock hub Needle point: Long point Size: 21G (0,8mm) x 40mm Colour code: Deep green Sterile, Single use, Individually peel packed. To comply with the latest issue of SANS 1124-1 UNIT: Box (100 per box) |  | GIBELA TRADE AND INVEST 1001 CC | 0.7001 | TARA | NEE014 | Shandong Qiaopai Group Co., Ltd | CHINA | 21 | 85 | USD | 10 | YES | each | No sample submitted for evaluation | NO TEST REPORTS
42142523-00018 | 73 | Needle, hypodermic, with inherent safety device to prevent needle stick injury, with plastic 6% luer lock hub Needle point: Long point Size: 21G (0,8mm) x 40mm Colour code: Deep green Sterile, Single use, Individually peel packed. To comply with the latest issue of SANS 1124-1 UNIT: Box (100 per box) |  | Chemical World cc | 0.7475 | SURGI PLUS | NDHDS02115 | MEDLINE | CHINA | 3 | 85 | USD | 5000 | YES | EACH | Size and sterility is only indicated on the outer package. Size and sterility is not indicated on the immediate package. | NO TEST REPORTS
42142523-00018 | 73 | Needle, hypodermic, with inherent safety device to prevent needle stick injury, with plastic 6% luer lock hub Needle point: Long point Size: 21G (0,8mm) x 40mm Colour code: Deep green Sterile, Single use, Individually peel packed. To comply with the latest issue of SANS 1124-1 UNIT: Box (100 per box) |  | UNITRADE 1032 CC | 0.83 | PROJEC PLUS | 1 | SHANGHAI KINDLY ENTERPRISE DEVELOPMENT GROUP | ROC | 21 | 85 | USD | 2500 | YES | 100 | COMPLY pending confirmation of latex free | NOT COMPLY
42142523-00018 | 73 | Needle, hypodermic, with inherent safety device to prevent needle stick injury, with plastic 6% luer lock hub Needle point: Long point Size: 21G (0,8mm) x 40mm Colour code: Deep green Sterile, Single use, Individually peel packed. To comply with the latest issue of SANS 1124-1 UNIT: Box (100 per box) | 1 | CENTRAL MEDICAL PTY LTD | 0.85 | LYSA | HSN2140 | Zhejian INI Medical Devices Co Ltd | China | 21 | 80 | China yuan | 50 boxes of 100 units | yes | box of 100 units | COMPLY pending confirmation of latex free | NO TEST REPORTS
42142523-00018 | 73 | Needle, hypodermic, with inherent safety device to prevent needle stick injury, with plastic 6% luer lock hub Needle point: Long point Size: 21G (0,8mm) x 40mm Colour code: Deep green Sterile, Single use, Individually peel packed. To comply with the latest issue of SANS 1124-1 UNIT: Box (100 per box) | EACH | MSIMED CC | 0.85 | KANG KANG | SAF23 | ZHENJIANG | CHINA | 25 | 20% | USD | 20000PCS | YES | EACH | No sample submitted for evaluation | NO TEST REPORTS
42142523-00018 | 73 | Needle, hypodermic, with inherent safety device to prevent needle stick injury, with plastic 6% luer lock hub Needle point: Long point Size: 21G (0,8mm) x 40mm Colour code: Deep green Sterile, Single use, Individually peel packed. To comply with the latest issue of SANS 1124-1 UNIT: Box (100 per box) | Basic Line offer | Becton Dickinson | 110 | BD Eclipse™ 21G X 1 1/2" (38mm) | 305895 | Becton Dickinson | Fraga, Spain | 1 week | 76% | USD | 1 BOX | YES | 100 Units in a box | COMPLY pending confirmation of latex free | COMPLY | Recommended | 91
42142523-00018 | 73 | Needle, hypodermic, with inherent safety device to prevent needle stick injury, with plastic 6% luer lock hub Needle point: Long point Size: 21G (0,8mm) x 40mm Colour code: Deep green Sterile, Single use, Individually peel packed. To comply with the latest issue of SANS 1124-1 UNIT: Box (100 per box) |  | LECHOBA MEDICAL TECHNOLOGIES (PTY) LTD | 1.23 | MEDOJECT SAFETY | CHS21112 | CHIRANA INJECTA A/S | SLOVAK REPUBLIC | 21 DAYS | 75% | EURO | 300 | YES | EACH (PACKED 100/BOX) | COMPLY pending confirmation of latex free | NOT COMPLY
42142523-00018 | 73 | Needle, hypodermic, with inherent safety device to prevent needle stick injury, with plastic 6% luer lock hub Needle point: Long point Size: 21G (0,8mm) x 40mm Colour code: Deep green Sterile, Single use, Individually peel packed. To comply with the latest issue of SANS 1124-1 UNIT: Box (100 per box) | N/A | SAFARMEX (PTY) LTD | 1.24 | Saclin | To be confirmed | Zhejiang Mediunion Healthcare Group Co. Ltd | China | 3days | 80% | US$ | 2000 Boxes | Yes | Box of 100 | No sample submitted for evaluation | NO TEST REPORT
42142523-00018 | 73 | Needle, hypodermic, with inherent safety device to prevent needle stick injury, with plastic 6% luer lock hub Needle point: Long point Size: 21G (0,8mm) x 40mm Colour code: Deep green Sterile, Single use, Individually peel packed. To comply with the latest issue of SANS 1124-1 UNIT: Box (100 per box) |  | ALMIKA TRADING PTY LTD | 1.26 | BD | 305895 | BD | USA | 14 | 0 | NO | 100 | YES | 1 | COMPLY pending confirmation of latex free | COMPLY
42142523-00018 | 73 | Needle, hypodermic, with inherent safety device to prevent needle stick injury, with plastic 6% luer lock hub Needle point: Long point Size: 21G (0,8mm) x 40mm Colour code: Deep green Sterile, Single use, Individually peel packed. To comply with the latest issue of SANS 1124-1 UNIT: Box (100 per box) |  | SUPRA HEALTHCARE JOHANNESBURG PTY LTD | 1.26 | REVITAL | SN21X40 | Revital Health | KENYA | Initial 60 days, thereafter 14 | 84 | USD | 1 Box | YES | 100 pcs/box | COMPLY pending confirmation of latex free
42142523-00018 | 73 | Needle, hypodermic, with inherent safety device to prevent needle stick injury, with plastic 6% luer lock hub Needle point: Long point Size: 21G (0,8mm) x 40mm Colour code: Deep green Sterile, Single use, Individually peel packed. To comply with the latest issue of SANS 1124-1 UNIT: Box (100 per box) | 1 | B. BRAUN MEDICAL (PTY) LTD | 1.45 | Sterican Safety G21 x 1 1/2, 0.8 x 40mm | 4670045S-01 | B. Braun | Germany | 42 | 50% | Euro | 100 | Yes | Each | COMPLY pending confirmation of latex free | NO TEST REPORTS
42142523-00017 | 74 | Needle, hypodermic, with inherent safety device to prevent needle stick injury, with plastic 6% luer lock hub Needle point: Long point Size: 22G (0,7mm) x 32mm Colour code: Black Sterile, Single use, Individually peel packed. To comply with the latest issue of SANS 1124-1 UNIT: Box (100 per box) |  | GIBELA TRADE AND INVEST 1001 CC | 0.73376 | TARA | NEE015 | Shandong Qiaopai Group Co., Ltd | CHINA | 21 | 85 | USD | 10 | YES | each | No sample submitted for evaluation | NO TEST REPORTS
42142523-00017 | 74 | Needle, hypodermic, with inherent safety device to prevent needle stick injury, with plastic 6% luer lock hub Needle point: Long point Size: 22G (0,7mm) x 32mm Colour code: Black Sterile, Single use, Individually peel packed. To comply with the latest issue of SANS 1124-1 UNIT: Box (100 per box) |  | Chemical World cc | 0.7475 | SURGI PLUS | NDHDS02214 | MEDLINE | CHINA | 3 | 85 | USD | 5000 | YES | EACH | COMPLY pending confirmation of latex free | NO TEST REPORTS
42142523-00017 | 74 | Needle, hypodermic, with inherent safety device to prevent needle stick injury, with plastic 6% luer lock hub Needle point: Long point Size: 22G (0,7mm) x 32mm Colour code: Black Sterile, Single use, Individually peel packed. To comply with the latest issue of SANS 1124-1 UNIT: Box (100 per box) |  | UNITRADE 1032 CC | 0.83 | PROJEC PLUS | 1 | SHANGHAI KINDLY ENTERPRISE DEVELOPMENT GROUP | ROC | 21 | 85 | USD | 2500 | YES | 100 | No sample submitted for evaluation | NO TEST REPORTS
42142523-00017 | 74 | Needle, hypodermic, with inherent safety device to prevent needle stick injury, with plastic 6% luer lock hub Needle point: Long point Size: 22G (0,7mm) x 32mm Colour code: Black Sterile, Single use, Individually peel packed. To comply with the latest issue of SANS 1124-1 UNIT: Box (100 per box) | 1 | CENTRAL MEDICAL PTY LTD | 0.85 | LYSA | HSN2232 | Zhejian INI Medical Devices Co Ltd | China | 21 | 80 | China yuan | 50 boxes of 100 units | yes | box of 100 units | COMPLY pending confirmation of latex free | NO TEST REPORTS
42142523-00017 | 74 | Needle, hypodermic, with inherent safety device to prevent needle stick injury, with plastic 6% luer lock hub Needle point: Long point Size: 22G (0,7mm) x 32mm Colour code: Black Sterile, Single use, Individually peel packed. To comply with the latest issue of SANS 1124-1 UNIT: Box (100 per box) | EACH | MSIMED CC | 0.85 | KANG KANG | SAF22 | ZHENJIANG | CHINA | 25 | 20% | USD | 20000PCS | YES | EACH | Wrong sample, the sample does not include an inherent safety device to prevent needle stick injury | NO TEST REPORTS
42142523-00017 | 74 | Needle, hypodermic, with inherent safety device to prevent needle stick injury, with plastic 6% luer lock hub Needle point: Long point Size: 22G (0,7mm) x 32mm Colour code: Black Sterile, Single use, Individually peel packed. To comply with the latest issue of SANS 1124-1 UNIT: Box (100 per box) | Basic Line offer | Becton Dickinson | 110 | BD Eclipse™ 22G X (30mm) | 305887 | Becton Dickinson | Fraga, Spain | 1 week | 76% | USD | 1 BOX | YES | 100 Units in a box | COMPLY pending confirmation of latex free | COMPLY | Recommended | 91
42142523-00017 | 74 | Needle, hypodermic, with inherent safety device to prevent needle stick injury, with plastic 6% luer lock hub Needle point: Long point Size: 22G (0,7mm) x 32mm Colour code: Black Sterile, Single use, Individually peel packed. To comply with the latest issue of SANS 1124-1 UNIT: Box (100 per box) | N/A | SAFARMEX (PTY) LTD | 1.24 | Saclin | To be confirmed | Zhejiang Mediunion Healthcare Group Co. Ltd | China | 3days | 80% | US$ | 2000 Boxes | Yes | Box of 100 | No sample submitted for evaluation | NO TEST REPORT
42142523-00017 | 74 | Needle, hypodermic, with inherent safety device to prevent needle stick injury, with plastic 6% luer lock hub Needle point: Long point Size: 22G (0,7mm) x 32mm Colour code: Black Sterile, Single use, Individually peel packed. To comply with the latest issue of SANS 1124-1 UNIT: Box (100 per box) |  | ALMIKA TRADING PTY LTD | 1.26 | BD | 305887 | BD | USA | 14 | 0 | NO | 100 | YES | 1 | COMPLY pending confirmation of latex free | COMPLY
42142523-00017 | 74 | Needle, hypodermic, with inherent safety device to prevent needle stick injury, with plastic 6% luer lock hub Needle point: Long point Size: 22G (0,7mm) x 32mm Colour code: Black Sterile, Single use, Individually peel packed. To comply with the latest issue of SANS 1124-1 UNIT: Box (100 per box) |  | SUPRA HEALTHCARE JOHANNESBURG PTY LTD | 1.26 | REVITAL | SN22X32 | Revital Health | KENYA | Initial 60 days, thereafter 14 | 84 | USD | 1 Box | YES | 100 pcs/box | Offered 23G instead of 22G
42142523-00017 | 74 | Needle, hypodermic, with inherent safety device to prevent needle stick injury, with plastic 6% luer lock hub Needle point: Long point Size: 22G (0,7mm) x 32mm Colour code: Black Sterile, Single use, Individually peel packed. To comply with the latest issue of SANS 1124-1 UNIT: Box (100 per box) | 1 | B. BRAUN MEDICAL (PTY) LTD | 1.45 | Sterican Safety G22 x 1, 0.7 x 25mm | 4670030S-01 | B. Braun | Germany | 42 | 50% | Euro | 100 | Yes | Each | Offered a length of 25mm instead of 32mm | NO TEST REPORTS
42142523-00016 | 75 | Needle, hypodermic, with inherent safety device to prevent needle stick injury, with plastic 6% luer lock hub Needle point: Long point Size: 23G (0,6mm) x 25mm Colour code: Deep blue Sterile, Single use, Individually peel packed. To comply with the latest issue of SANS 1124-1 UNIT: Box (100 per box) |  | GIBELA TRADE AND INVEST 1001 CC | 0.73376 | TARA | NEE016 | Shandong Qiaopai Group Co., Ltd | CHINA | 21 | 85 | USD | 10 | YES | each | No sample submitted for evaluation | NO TEST REPORTS
42142523-00016 | 75 | Needle, hypodermic, with inherent safety device to prevent needle stick injury, with plastic 6% luer lock hub Needle point: Long point Size: 23G (0,6mm) x 25mm Colour code: Deep blue Sterile, Single use, Individually peel packed. To comply with the latest issue of SANS 1124-1 UNIT: Box (100 per box) |  | Chemical World cc | 0.7475 | SURGI PLUS | NDHDS0231 | MEDLINE | CHINA | 3 | 85 | USD | 5000 | YES | EACH | COMPLY pending confirmation of latex free | NO TEST REPORTS
42142523-00016 | 75 | Needle, hypodermic, with inherent safety device to prevent needle stick injury, with plastic 6% luer lock hub Needle point: Long point Size: 23G (0,6mm) x 25mm Colour code: Deep blue Sterile, Single use, Individually peel packed. To comply with the latest issue of SANS 1124-1 UNIT: Box (100 per box) |  | UNITRADE 1032 CC | 0.83 | PROJEC PLUS | 1 | SHANGHAI KINDLY ENTERPRISE DEVELOPMENT GROUP | ROC | 21 | 85 | USD | 2500 | YES | 100 | COMPLY pending confirmation of latex free | NOT COMPLY
42142523-00016 | 75 | Needle, hypodermic, with inherent safety device to prevent needle stick injury, with plastic 6% luer lock hub Needle point: Long point Size: 23G (0,6mm) x 25mm Colour code: Deep blue Sterile, Single use, Individually peel packed. To comply with the latest issue of SANS 1124-1 UNIT: Box (100 per box) | 1 | CENTRAL MEDICAL PTY LTD | 0.85 | LYSA | HSN2325 | Zhejian INI Medical Devices Co Ltd | China | 21 | 80 | China yuan | 50 boxes of 100 units | yes | box of 100 units | COMPLY pending confirmation of latex free | NO TEST REPORTS
42142523-00016 | 75 | Needle, hypodermic, with inherent safety device to prevent needle stick injury, with plastic 6% luer lock hub Needle point: Long point Size: 23G (0,6mm) x 25mm Colour code: Deep blue Sterile, Single use, Individually peel packed. To comply with the latest issue of SANS 1124-1 UNIT: Box (100 per box) | EACH | MSIMED CC | 0.85 | KANG KANG | SAF21 | ZHENJIANG | CHINA | 25 | 20% | USD | 20000PCS | YES | EACH | No sample submitted for evaluation | NO TEST REPORTS
42142523-00016 | 75 | Needle, hypodermic, with inherent safety device to prevent needle stick injury, with plastic 6% luer lock hub Needle point: Long point Size: 23G (0,6mm) x 25mm Colour code: Deep blue Sterile, Single use, Individually peel packed. To comply with the latest issue of SANS 1124-1 UNIT: Box (100 per box) | Basic Line offer | Becton Dickinson | 110 | BD Eclipse™ 23G X 25mm | 305892 | Becton Dickinson | Fraga, Spain | 1 week | 76% | USD | 1 BOX | YES | 100 Units in a box | COMPLY pending confirmation of latex free | COMPLY | Recommended | 91
42142523-00016 | 75 | Needle, hypodermic, with inherent safety device to prevent needle stick injury, with plastic 6% luer lock hub Needle point: Long point Size: 23G (0,6mm) x 25mm Colour code: Deep blue Sterile, Single use, Individually peel packed. To comply with the latest issue of SANS 1124-1 UNIT: Box (100 per box) | N/A | SAFARMEX (PTY) LTD | 1.24 | Saclin | To be confirmed | Zhejiang Mediunion Healthcare Group Co. Ltd | China | 3days | 80% | US$ | 2000 Boxes | Yes | Box of 100 | No sample submitted for evaluation | NO TEST REPORT
42142523-00016 | 75 | Needle, hypodermic, with inherent safety device to prevent needle stick injury, with plastic 6% luer lock hub Needle point: Long point Size: 23G (0,6mm) x 25mm Colour code: Deep blue Sterile, Single use, Individually peel packed. To comply with the latest issue of SANS 1124-1 UNIT: Box (100 per box) |  | ALMIKA TRADING PTY LTD | 1.26 | BD | 305892 | BD | USA | 14 | 0 | NO | 100 | YES | 1 | COMPLY pending confirmation of latex free | COMPLY
42142523-00016 | 75 | Needle, hypodermic, with inherent safety device to prevent needle stick injury, with plastic 6% luer lock hub Needle point: Long point Size: 23G (0,6mm) x 25mm Colour code: Deep blue Sterile, Single use, Individually peel packed. To comply with the latest issue of SANS 1124-1 UNIT: Box (100 per box) |  | SUPRA HEALTHCARE JOHANNESBURG PTY LTD | 1.26 | REVITAL | SN23X25 | Revital Health | KENYA | Initial 60 days, thereafter 14 | 84 | USD | 1 Box | YES | 100 pcs/box | COMPLY pending confirmation of latex free
42142523-00016 | 75 | Needle, hypodermic, with inherent safety device to prevent needle stick injury, with plastic 6% luer lock hub Needle point: Long point Size: 23G (0,6mm) x 25mm Colour code: Deep blue Sterile, Single use, Individually peel packed. To comply with the latest issue of SANS 1124-1 UNIT: Box (100 per box) | 1 | B. BRAUN MEDICAL (PTY) LTD | 1.45 | Sterican Safety G23 x 1, 0.6 x 25mm | 4670025S-01 | B. Braun | Germany | 42 | 50% | Euro | 100 | Yes | Each | COMPLY pending confirmation of latex free | NO TEST REPORTS
42142521-00000 | 76 | Needle, use for blood collecting. Needle type: multiple sample collecting needle, thin wall Patient side: The needle must be colour coded: green and needle size: 21G x 38mm In the middle of the needle it will contained screw-in tread which will help to attach the needle to the tube holder. Tube side: The needle on tube side must be covered with in a latex free elastic membrane to prevent the blood spilling when changing tubes. Both sides of the needle must be capped Sterile, individually peel packed Pack size: Box of 100 needles | 1 | B. BRAUN MEDICAL (PTY) LTD | 0.25 | Sterican G21 x 1 1/2", 0.80 x 40mm Green | 4657527 | B. Braun | Germany | 42 | 50% | Euro | 100 | Yes | Box of 100 | Offered a hypodermic needle instead of blood collecting needle
42142521-00000 | 76 | Needle, use for blood collecting. Needle type: multiple sample collecting needle, thin wall Patient side: The needle must be colour coded: green and needle size: 21G x 38mm In the middle of the needle it will contained screw-in tread which will help to attach the needle to the tube holder. Tube side: The needle on tube side must be covered with in a latex free elastic membrane to prevent the blood spilling when changing tubes. Both sides of the needle must be capped Sterile, individually peel packed Pack size: Box of 100 needles |  | GIBELA TRADE AND INVEST 1001 CC | 50.94 | TARA | NEE017 | Shandong Qiaopai Group Co., Ltd | CHINA | 21 | 85 | USD | 10 | YES | box 100 | COMPLY pending confirmation of latex free |  | Recommended | 100
42142521-00000 | 76 | Needle, use for blood collecting. Needle type: multiple sample collecting needle, thin wall Patient side: The needle must be colour coded: green and needle size: 21G x 38mm In the middle of the needle it will contained screw-in tread which will help to attach the needle to the tube holder. Tube side: The needle on tube side must be covered with in a latex free elastic membrane to prevent the blood spilling when changing tubes. Both sides of the needle must be capped Sterile, individually peel packed Pack size: Box of 100 needles |  | Medi-Core Technologies | 0.53 | ULIFE | 42142521-0000 | JIAXING U-LIFE MEDICAL DEVICE TECHNOLOGY.,CO LTD | CHINA | 21 | 80 | USD | 1 BOX | YES | BOX | No sample submitted for evaluation
42142521-00000 | 76 | Needle, use for blood collecting. Needle type: multiple sample collecting needle, thin wall Patient side: The needle must be colour coded: green and needle size: 21G x 38mm In the middle of the needle it will contained screw-in tread which will help to attach the needle to the tube holder. Tube side: The needle on tube side must be covered with in a latex free elastic membrane to prevent the blood spilling when changing tubes. Both sides of the needle must be capped Sterile, individually peel packed Pack size: Box of 100 needles |  | Chemical World cc | 0.61 | SURGI PLUS | BCND21 | XINLE | CHINA | 3 | 85 | USD | 5000 | YES | 100 | COMPLY pending confirmation of latex free
42142521-00000 | 76 |  |  | Zanotone (Pty) Ltd | 0.78 | Mekon |  | Mekon | [6 empty] | Each | No brand name No product code
42142521-00000 | 76 | Needle, use for blood collecting. Needle type: multiple sample collecting needle, thin wall Patient side: The needle must be colour coded: green and needle size: 21G x 38mm In the middle of the needle it will contained screw-in tread which will help to attach the needle to the tube holder. Tube side: The needle on tube side must be covered with in a latex free elastic membrane to prevent the blood spilling when changing tubes. Both sides of the needle must be capped Sterile, individually peel packed Pack size: Box of 100 needles | Basic Line offer | Becton Dickinson | 0.8 | BD VAC Precision Glide, 22G | 360213 | Becton Dickinson | United Kingdom | 7 Working Days | 76% | USD | 1 BOX | YES | 1 BOX = 100 UNITS | COMPLY pending confirmation of latex free
42142521-00000 | 76 | Needle, use for blood collecting. Needle type: multiple sample collecting needle, thin wall Patient side: The needle must be colour coded: green and needle size: 21G x 38mm In the middle of the needle it will contained screw-in tread which will help to attach the needle to the tube holder. Tube side: The needle on tube side must be covered with in a latex free elastic membrane to prevent the blood spilling when changing tubes. Both sides of the needle must be capped Sterile, individually peel packed Pack size: Box of 100 needles |  | FOCUS PRODUCTS (PTY) LTD | 0.83 | VACUCARE |  | VACUCARE | UNITED KINGDOM | 45 | 90 | POUND STERLING | 100 BOXES | YES | NONE | COMPLY pendicng confirmation of latex free
42142521-00000 | 76 | Needle, use for blood collecting. Needle type: multiple sample collecting needle, thin wall Patient side: The needle must be colour coded: green and needle size: 21G x 38mm In the middle of the needle it will contained screw-in tread which will help to attach the needle to the tube holder. Tube side: The needle on tube side must be covered with in a latex free elastic membrane to prevent the blood spilling when changing tubes. Both sides of the needle must be capped Sterile, individually peel packed Pack size: Box of 100 needles |  | SUPRA HEALTHCARE JOHANNESBURG PTY LTD | 0.96 | CD Newmed | BLDCOL21G | CD Newmed | CHINA | Initial 60 days, thereafter 14 days | 84 | USD | 1 Box | YES |  | COMPLY pending confirmation of latex free
42142521-00000 | 76 | Needle, use for blood collecting. Needle type: multiple sample collecting needle, thin wall Patient side: The needle must be colour coded: green and needle size: 21G x 38mm In the middle of the needle it will contained screw-in tread which will help to attach the needle to the tube holder. Tube side: The needle on tube side must be covered with in a latex free elastic membrane to prevent the blood spilling when changing tubes. Both sides of the needle must be capped Sterile, individually peel packed Pack size: Box of 100 needles | Offer 1 of 1 | SSEM Mthembu Medical (Pty) Ltd | 1.04 | Disera | DT250021 | Disera | Turkey | 56 | 85% | USD | 1 | Yes | 100/Box | COMPLY pending confirmation of latex free
42142521-00000 | 76 | Needle, use for blood collecting. Needle type: multiple sample collecting needle, thin wall Patient side: The needle must be colour coded: green and needle size: 21G x 38mm In the middle of the needle it will contained screw-in tread which will help to attach the needle to the tube holder. Tube side: The needle on tube side must be covered with in a latex free elastic membrane to prevent the blood spilling when changing tubes. Both sides of the needle must be capped Sterile, individually peel packed Pack size: Box of 100 needles |  | LECHOBA MEDICAL TECHNOLOGIES (PTY) LTD | 1.18 | CHIRAVAC | CHBCN21112 | CHIRANA INJECTA A/S | SLOVAK REPUBLIC | 21 DAYS | 75% | EURO | 1000 | YES | 100/BOX | COMPLY pending confirmation of latex free
42142521-00000 | 76 | Needle, use for blood collecting. Needle type: multiple sample collecting needle, thin wall Patient side: The needle must be colour coded: green and needle size: 21G x 38mm In the middle of the needle it will contained screw-in tread which will help to attach the needle to the tube holder. Tube side: The needle on tube side must be covered with in a latex free elastic membrane to prevent the blood spilling when changing tubes. Both sides of the needle must be capped Sterile, individually peel packed Pack size: Box of 100 needles | N/A | SAFARMEX (PTY) LTD | 1.26 | Saclin | To be confirmed | Zhejiang Mediunion Healthcare Group Co. Ltd | China | 3days | 80% | US$ | 50000 Pieces | Yes | Box of 100 Needles | No sample submitted for evaluation | NO TEST REPORT
42142521-00000 | 76 | Needle, use for blood collecting. Needle type: multiple sample collecting needle, thin wall Patient side: The needle must be colour coded: green and needle size: 21G x 38mm In the middle of the needle it will contained screw-in tread which will help to attach the needle to the tube holder. Tube side: The needle on tube side must be covered with in a latex free elastic membrane to prevent the blood spilling when changing tubes. Both sides of the needle must be capped Sterile, individually peel packed Pack size: Box of 100 needles | 1 | Biologica Pharmaceuticals Pty Ltd | 1.53 | Sol- M | 110201020005 | Sol- Millennium | China | 4- 6 Weeks | 100% | USD | 100 Boxes | 100% | N/A
42142521-00000 | 76 | Needle, use for blood collecting. Needle type: multiple sample collecting needle, thin wall Patient side: The needle must be colour coded: green and needle size: 21G x 38mm In the middle of the needle it will contained screw-in tread which will help to attach the needle to the tube holder. Tube side: The needle on tube side must be covered with in a latex free elastic membrane to prevent the blood spilling when changing tubes. Both sides of the needle must be capped Sterile, individually peel packed Pack size: Box of 100 needles | N/A | PROMED TECHNOLOGIES (PTY) LTD | 2.53 | U-LIFE | 521000 | JIAXING | CHINA | 14 DAYS | 85% | USD | 1 BOX | YES | BOX | No samples submitted for evaluation
42142521-00000 | 76 | Needle, use for blood collecting. Needle type: multiple sample collecting needle, thin wall Patient side: The needle must be colour coded: green and needle size: 21G x 38mm In the middle of the needle it will contained screw-in tread which will help to attach the needle to the tube holder. Tube side: The needle on tube side must be covered with in a latex free elastic membrane to prevent the blood spilling when changing tubes. Both sides of the needle must be capped Sterile, individually peel packed Pack size: Box of 100 needles |  | ABCARE TECHNOLOGIES (PTY) LTD | 75.99 | CHIRAVAC™ | BCN001 | CHIRANA T INJEKTA | SLOVAKIA | 45 | 85 | EURO | 100 | YES | BOX OF 100 | COMPLY
42142521-00001 | 77 | Needle, use for blood collecting. Needle type: multiple sample collecting needle, thin wall Patient side: The needle must be colour coded: black and needle size: 22G x 38mm In the middle of the needle it will contained screw-in tread which will help to attach the needle to the tube holder. Tube side: The needle on tube side must be covered with in a latex free elastic membrane to prevent the blood spilling when changing tubes. Both sides of the needle must be capped Sterile, individually peel packed Pack size: Box of 100 needles |  | GIBELA TRADE AND INVEST 1001 CC | 50.94 | TARA | NEE018 | Shandong Qiaopai Group Co., Ltd | CHINA | 21 | 85 | USD | 10 | YES | box 100 | COMPLY pending confirmation of latex free |  | Recommended | 100
42142521-00001 | 77 | Needle, use for blood collecting. Needle type: multiple sample collecting needle, thin wall Patient side: The needle must be colour coded: black and needle size: 22G x 38mm In the middle of the needle it will contained screw-in tread which will help to attach the needle to the tube holder. Tube side: The needle on tube side must be covered with in a latex free elastic membrane to prevent the blood spilling when changing tubes. Both sides of the needle must be capped Sterile, individually peel packed Pack size: Box of 100 needles |  | Medi-Core Technologies | 0.53 | UIFE | 42142521-00001 | JIAXING U-LIFE MEDICAL DEVICE TECHNOLOGY.,CO LTD | CHINA | 21 | 80 | USD | 1 BOX | YES | BOX | No sample submitted for evaluation
42142521-00001 | 77 | Needle, use for blood collecting. Needle type: multiple sample collecting needle, thin wall Patient side: The needle must be colour coded: black and needle size: 22G x 38mm In the middle of the needle it will contained screw-in tread which will help to attach the needle to the tube holder. Tube side: The needle on tube side must be covered with in a latex free elastic membrane to prevent the blood spilling when changing tubes. Both sides of the needle must be capped Sterile, individually peel packed Pack size: Box of 100 needles |  | Chemical World cc | 0.61 | SURGI PLUS | BCND22 | XINLE | CHINA | 3 | 85 | USD | 5000 | YES | 100 | COMPLY pending confirmation of latex free
42142521-00001 | 77 | Needle, use for blood collecting. Needle type: multiple sample collecting needle, thin wall Patient side: The needle must be colour coded: black and needle size: 22G x 38mm In the middle of the needle it will contained screw-in tread which will help to attach the needle to the tube holder. Tube side: The needle on tube side must be covered with in a latex free elastic membrane to prevent the blood spilling when changing tubes. Both sides of the needle must be capped Sterile, individually peel packed Pack size: Box of 100 needles |  | Zanotone (Pty) Ltd | 0.78 | Mekon |  | Mekon | [6 empty] | Each | No brand name No product codes
42142521-00001 | 77 | Needle, use for blood collecting. Needle type: multiple sample collecting needle, thin wall Patient side: The needle must be colour coded: black and needle size: 22G x 38mm In the middle of the needle it will contained screw-in tread which will help to attach the needle to the tube holder. Tube side: The needle on tube side must be covered with in a latex free elastic membrane to prevent the blood spilling when changing tubes. Both sides of the needle must be capped Sterile, individually peel packed Pack size: Box of 100 needles | Basic Line offer | Becton Dickinson | 0.81 | BD VAC Precision Glide, 21G | 360211 | Becton Dickinson | United Kingdom | 7 Working Days | 76% | USD | 1 BOX | YES | 1 BOX = 100 UNITS | COMPLY pending confirmation of latex free
42142521-00001 | 77 | Needle, use for blood collecting. Needle type: multiple sample collecting needle, thin wall Patient side: The needle must be colour coded: black and needle size: 22G x 38mm In the middle of the needle it will contained screw-in tread which will help to attach the needle to the tube holder. Tube side: The needle on tube side must be covered with in a latex free elastic membrane to prevent the blood spilling when changing tubes. Both sides of the needle must be capped Sterile, individually peel packed Pack size: Box of 100 needles |  | FOCUS PRODUCTS (PTY) LTD | 0.83 | VACUCARE |  | VACUCARE | UNITED KINGDOM | 45 | 90 | POUND STERLING | 100 BOXES | YES | NONE | No brand name. No product code.
42142521-00001 | 77 | Needle, use for blood collecting. Needle type: multiple sample collecting needle, thin wall Patient side: The needle must be colour coded: black and needle size: 22G x 38mm In the middle of the needle it will contained screw-in tread which will help to attach the needle to the tube holder. Tube side: The needle on tube side must be covered with in a latex free elastic membrane to prevent the blood spilling when changing tubes. Both sides of the needle must be capped Sterile, individually peel packed Pack size: Box of 100 needles |  | SUPRA HEALTHCARE JOHANNESBURG PTY LTD | 0.96 | CD Newmed | BLDCOL22G | CD Newmed | CHINA | Initial 60 days, thereafter 14 days | 84 | USD | 1 Box | YES |  | COMPLY pending confirmation of latex free
42142521-00001 | 77 | Needle, use for blood collecting. Needle type: multiple sample collecting needle, thin wall Patient side: The needle must be colour coded: black and needle size: 22G x 38mm In the middle of the needle it will contained screw-in tread which will help to attach the needle to the tube holder. Tube side: The needle on tube side must be covered with in a latex free elastic membrane to prevent the blood spilling when changing tubes. Both sides of the needle must be capped Sterile, individually peel packed Pack size: Box of 100 needles | Offer 1 of 1 | SSEM Mthembu Medical (Pty) Ltd | 1.04 | Disera | DT250022 | Disera | Turkey | 56 | 85% | USD | 1 | Yes | 100/Box | COMPLY pending confirmation of latex free
42142521-00001 | 77 | Needle, use for blood collecting. Needle type: multiple sample collecting needle, thin wall Patient side: The needle must be colour coded: black and needle size: 22G x 38mm In the middle of the needle it will contained screw-in tread which will help to attach the needle to the tube holder. Tube side: The needle on tube side must be covered with in a latex free elastic membrane to prevent the blood spilling when changing tubes. Both sides of the needle must be capped Sterile, individually peel packed Pack size: Box of 100 needles |  | LECHOBA MEDICAL TECHNOLOGIES (PTY) LTD | 1.18 | CHIRAVAC | CHBCN22112 | CHIRANA INJECTA A/S | SLOVAK REPUBLIC | 21 DAYS | 75% | EURO | 1000 | YES | 100/BOX | COMPLY pending confirmation of latex free
42142521-00001 | 77 | Needle, use for blood collecting. Needle type: multiple sample collecting needle, thin wall Patient side: The needle must be colour coded: black and needle size: 22G x 38mm In the middle of the needle it will contained screw-in tread which will help to attach the needle to the tube holder. Tube side: The needle on tube side must be covered with in a latex free elastic membrane to prevent the blood spilling when changing tubes. Both sides of the needle must be capped Sterile, individually peel packed Pack size: Box of 100 needles | N/A | SAFARMEX (PTY) LTD | 1.26 | Saclin | To be confirmed | Zhejiang Mediunion Healthcare Group Co. Ltd | China | 3days | 80% | US$ | 50000 Pieces | Yes | Box of 100 Needles | No sample submitted for evaluation
42142521-00001 | 77 | Needle, use for blood collecting. Needle type: multiple sample collecting needle, thin wall Patient side: The needle must be colour coded: black and needle size: 22G x 38mm In the middle of the needle it will contained screw-in tread which will help to attach the needle to the tube holder. Tube side: The needle on tube side must be covered with in a latex free elastic membrane to prevent the blood spilling when changing tubes. Both sides of the needle must be capped Sterile, individually peel packed Pack size: Box of 100 needles | 1 | Biologica Pharmaceuticals Pty Ltd | 1.53 | Sol- M | 110201020003 | Sol- Millennium | China | 4-6 Weeks | 100% | USD | 100 Boxes | 100% | N/A
42142521-00001 | 77 | Needle, use for blood collecting. Needle type: multiple sample collecting needle, thin wall Patient side: The needle must be colour coded: black and needle size: 22G x 38mm In the middle of the needle it will contained screw-in tread which will help to attach the needle to the tube holder. Tube side: The needle on tube side must be covered with in a latex free elastic membrane to prevent the blood spilling when changing tubes. Both sides of the needle must be capped Sterile, individually peel packed Pack size: Box of 100 needles | N/A | PROMED TECHNOLOGIES (PTY) LTD | 52.6 | U-LIFE | 521001 | JIAXING | CHINA | 14 DAYS | 85% | USD | 1 BOX | YES | BOX | No samples submitted for evaluation
42142521-00001 | 77 | Needle, use for blood collecting. Needle type: multiple sample collecting needle, thin wall Patient side: The needle must be colour coded: black and needle size: 22G x 38mm In the middle of the needle it will contained screw-in tread which will help to attach the needle to the tube holder. Tube side: The needle on tube side must be covered with in a latex free elastic membrane to prevent the blood spilling when changing tubes. Both sides of the needle must be capped Sterile, individually peel packed Pack size: Box of 100 needles |  | ABCARE TECHNOLOGIES (PTY) LTD | 75.99 | CHIRAVAC™ | BCN002 | CHIRANA T INJEKTA | SLOVAKIA | 45 | 85 | EURO | 100 | YES | BOX OF 100 | COMPLY
42142521-00002 | 78 | Needle, use for blood collecting inherent safety device. Needle must be a multiple sample collecting needle, thin wall. Patient side: The needle must be always position bevel-up towards the safety device which will eliminates the repositioning of the needle. The needle must be capped with is colour coded green on the patient side. Safety device must be an integral part of the device and must cover the needle completely. It must have a permanently, securely lock in place to eliminate hard-surface activation. It must easy operation, single-handed activation immediately after exiting the venepuncture site. In the middle of the needle it will contained screw-in tread which will help to attach the needle to the tube holder. Tube side: The needle must be covered with in a latex free elastic membrane to prevent the blood spilling when changing tubes. Needle size: 21G x 38mm Sterile, individually peel packed Box of 48 Unit: Box | Basic Line offer | Becton Dickinson | 68.25 | BD VAC ECLIPSE, GREEN, 21G | 368609 | Becton Dickinson | United Kingdom | 7 Working Days | 76% | USD | 1 BOX | NO | 1 BOX = 48 UNITS / 48 UNITS | COMPLY pending confirmation of latex free |  | Recommended | 91
42142521-00002 | 78 | Needle, use for blood collecting inherent safety device. Needle must be a multiple sample collecting needle, thin wall. Patient side: The needle must be always position bevel-up towards the safety device which will eliminates the repositioning of the needle. The needle must be capped with is colour coded green on the patient side. Safety device must be an integral part of the device and must cover the needle completely. It must have a permanently, securely lock in place to eliminate hard-surface activation. It must easy operation, single-handed activation immediately after exiting the venepuncture site. In the middle of the needle it will contained screw-in tread which will help to attach the needle to the tube holder. Tube side: The needle must be covered with in a latex free elastic membrane to prevent the blood spilling when changing tubes. Needle size: 21G x 38mm Sterile, individually peel packed Box of 100 Unit: Box |  | Chemical World cc | 1.04 | SURGI PLUS |  | XINLE | CHINA | 3 | 85 | USD | 5000 | YES | 100 | Wrong sample, offered a needle with a tube holder instead of a needle without a tube holder.
42142521-00002 | 78 | Needle, use for blood collecting inherent safety device. Needle must be a multiple sample collecting needle, thin wall. Patient side: The needle must be always position bevel-up towards the safety device which will eliminates the repositioning of the needle. The needle must be capped with is colour coded green on the patient side. Safety device must be an integral part of the device and must cover the needle completely. It must have a permanently, securely lock in place to eliminate hard-surface activation. It must easy operation, single-handed activation immediately after exiting the venepuncture site. In the middle of the needle it will contained screw-in tread which will help to attach the needle to the tube holder. Tube side: The needle must be covered with in a latex free elastic membrane to prevent the blood spilling when changing tubes. Needle size: 21G x 38mm Sterile, individually peel packed Box of 100 Unit: Box | N/A | SAFARMEX (PTY) LTD | 1.6 | Saclin | To be confirmed | Zhejiang Mediunion Healthcare Group Co. Ltd | China | 3days | 80% | US$ | 50000 Pieces | Yes | Box of 100 Needles | No sample submitted for evaluation
42142521-00002 | 78 | Needle, use for blood collecting inherent safety device. Needle must be a multiple sample collecting needle, thin wall. Patient side: The needle must be always position bevel-up towards the safety device which will eliminates the repositioning of the needle. The needle must be capped with is colour coded green on the patient side. Safety device must be an integral part of the device and must cover the needle completely. It must have a permanently, securely lock in place to eliminate hard-surface activation. It must easy operation, single-handed activation immediately after exiting the venepuncture site. In the middle of the needle it will contained screw-in tread which will help to attach the needle to the tube holder. Tube side: The needle must be covered with in a latex free elastic membrane to prevent the blood spilling when changing tubes. Needle size: 21G x 38mm Sterile, individually peel packed Box of 100 Unit: Box | Basic Line Offer | BIOCOR HOSPITAL SUPPLIES PTY LTD | 1.63 | Medline | N/A | Yangzhong Medline Industry Co.Ltd | China | 0.006 | N/A | USD | 5 Boxes | Yes | N/A | COMPLY pending confirmation of latex free
42142521-00002 | 78 | Needle, use for blood collecting inherent safety device. Needle must be a multiple sample collecting needle, thin wall. Patient side: The needle must be always position bevel-up towards the safety device which will eliminates the repositioning of the needle. The needle must be capped with is colour coded green on the patient side. Safety device must be an integral part of the device and must cover the needle completely. It must have a permanently, securely lock in place to eliminate hard-surface activation. It must easy operation, single-handed activation immediately after exiting the venepuncture site. In the middle of the needle it will contained screw-in tread which will help to attach the needle to the tube holder. Tube side: The needle must be covered with in a latex free elastic membrane to prevent the blood spilling when changing tubes. Needle size: 21G x 38mm Sterile, individually peel packed Box of 100 Unit: Box |  | Zanotone (Pty) Ltd | 2.22 | Mekon |  | Mekon | [6 empty] | Each | No brand name No product codes
42142521-00002 | 78 | Needle, use for blood collecting inherent safety device. Needle must be a multiple sample collecting needle, thin wall. Patient side: The needle must be always position bevel-up towards the safety device which will eliminates the repositioning of the needle. The needle must be capped with is colour coded green on the patient side. Safety device must be an integral part of the device and must cover the needle completely. It must have a permanently, securely lock in place to eliminate hard-surface activation. It must easy operation, single-handed activation immediately after exiting the venepuncture site. In the middle of the needle it will contained screw-in tread which will help to attach the needle to the tube holder. Tube side: The needle must be covered with in a latex free elastic membrane to prevent the blood spilling when changing tubes. Needle size: 21G x 38mm Sterile, individually peel packed Box of 100 Unit: Box |  | FOCUS PRODUCTS (PTY) LTD | 3.02 | GENERIC |  |  | CHINA | 45 | 90 | USD | 100 BOXES | YES | NONE | No brand name. No product code.
42142521-00002 | 78 | Needle, use for blood collecting inherent safety device. Needle must be a multiple sample collecting needle, thin wall. Patient side: The needle must be always position bevel-up towards the safety device which will eliminates the repositioning of the needle. The needle must be capped with is colour coded green on the patient side. Safety device must be an integral part of the device and must cover the needle completely. It must have a permanently, securely lock in place to eliminate hard-surface activation. It must easy operation, single-handed activation immediately after exiting the venepuncture site. In the middle of the needle it will contained screw-in tread which will help to attach the needle to the tube holder. Tube side: The needle must be covered with in a latex free elastic membrane to prevent the blood spilling when changing tubes. Needle size: 21G x 38mm Sterile, individually peel packed Box of 100 Unit: Box |  | SUPRA HEALTHCARE JOHANNESBURG PTY LTD | 3.45 | Avacare | BLDCOLS21G | HENAN AILE INDUSTRIAL CO.,LTD | CHINA | Initial 60 days, thereafter 14 days | 84 | USD | 1 Box | YES |  | No sample submitted for evaluation
42142521-00002 | 78 | Needle, use for blood collecting inherent safety device. Needle must be a multiple sample collecting needle, thin wall. Patient side: The needle must be always position bevel-up towards the safety device which will eliminates the repositioning of the needle. The needle must be capped with is colour coded green on the patient side. Safety device must be an integral part of the device and must cover the needle completely. It must have a permanently, securely lock in place to eliminate hard-surface activation. It must easy operation, single-handed activation immediately after exiting the venepuncture site. In the middle of the needle it will contained screw-in tread which will help to attach the needle to the tube holder. Tube side: The needle must be covered with in a latex free elastic membrane to prevent the blood spilling when changing tubes. Needle size: 21G x 38mm Sterile, individually peel packed Box of 100 Unit: Box | Offer 1 of 1 | SSEM Mthembu Medical (Pty) Ltd | 4.37 | Disera | DT24121190 | Disera | Turkey | 56 | 85% | USD | 1 | Yes | 100/Box | The package is tear off instead of a peel apart package
42142521-00003 | 79 | Needle, use for blood collecting inherent safety device. Needle must be a multiple sample collecting needle, thin wall. Patient side: The needle must be always position bevel-up towards the safety device which will eliminates the repositioning of the needle. The needle must be capped with is colour coded black on the patient side. Safety device must be an integral part of the device and must cover the needle completely. It must have a permanently, securely lock in place to eliminate hard-surface activation. It must easy operation, single-handed activation immediately after exiting the venepuncture site. In the middle of the needle it will contained screw-in tread which will help to attach the needle to the tube holder. Needle size: 22G x 38mm Sterile, individually peel packed Box of 100 Unit: Box | Basic Line offer | Becton Dickinson | 0.69 | BD VAC ECLIPSE, BLACK, 22G | 368610 | Becton Dickinson | United Kingdom | 7 Working Days | 76% | USD | 1 BOX | NO | 1 BOX = 48 UNITS / 48 UNITS | Offered a length of 32mm instead of 38mm
42142521-00003 | 79 | Needle, use for blood collecting inherent safety device. Needle must be a multiple sample collecting needle, thin wall. Patient side: The needle must be always position bevel-up towards the safety device which will eliminates the repositioning of the needle. The needle must be capped with is colour coded black on the patient side. Safety device must be an integral part of the device and must cover the needle completely. It must have a permanently, securely lock in place to eliminate hard-surface activation. It must easy operation, single-handed activation immediately after exiting the venepuncture site. In the middle of the needle it will contained screw-in tread which will help to attach the needle to the tube holder. Needle size: 22G x 38mm Sterile, individually peel packed Box of 100 Unit: Box |  | Chemical World cc | 1.04 | SURGI PLUS |  | XINLE | CHINA | 3 | 85 | USD | 5000 | YES | 100 | No sample submitted for evaluation
42142521-00003 | 79 | Needle, use for blood collecting inherent safety device. Needle must be a multiple sample collecting needle, thin wall. Patient side: The needle must be always position bevel-up towards the safety device which will eliminates the repositioning of the needle. The needle must be capped with is colour coded black on the patient side. Safety device must be an integral part of the device and must cover the needle completely. It must have a permanently, securely lock in place to eliminate hard-surface activation. It must easy operation, single-handed activation immediately after exiting the venepuncture site. In the middle of the needle it will contained screw-in tread which will help to attach the needle to the tube holder. Needle size: 22G x 38mm Sterile, individually peel packed Box of 100 Unit: Box | N/A | SAFARMEX (PTY) LTD | 1.6 | Saclin | To be confirmed | Zhejiang Mediunion Healthcare Group Co. Ltd | China | 3days | 80% | US$ | 50000 Pieces | Yes | Box of 100 Needles | No sample submitted for evaluation
42142521-00003 | 79 | Needle, use for blood collecting inherent safety device. Needle must be a multiple sample collecting needle, thin wall. Patient side: The needle must be always position bevel-up towards the safety device which will eliminates the repositioning of the needle. The needle must be capped with is colour coded black on the patient side. Safety device must be an integral part of the device and must cover the needle completely. It must have a permanently, securely lock in place to eliminate hard-surface activation. It must easy operation, single-handed activation immediately after exiting the venepuncture site. In the middle of the needle it will contained screw-in tread which will help to attach the needle to the tube holder. Needle size: 22G x 38mm Sterile, individually peel packed Box of 100 Unit: Box | Basic Line Offer | BIOCOR HOSPITAL SUPPLIES PTY LTD | 163 | Medline | N/A | Yangzhong Medline Industry Co.Ltd | China | 0.006 | 0.0085 | USD | 10 Boxes | Yes | Box/100 | COMPLY pending confirmation of latex free |  | Recommended | 100
42142521-00003 | 79 | Needle, use for blood collecting inherent safety device. Needle must be a multiple sample collecting needle, thin wall. Patient side: The needle must be always position bevel-up towards the safety device which will eliminates the repositioning of the needle. The needle must be capped with is colour coded black on the patient side. Safety device must be an integral part of the device and must cover the needle completely. It must have a permanently, securely lock in place to eliminate hard-surface activation. It must easy operation, single-handed activation immediately after exiting the venepuncture site. In the middle of the needle it will contained screw-in tread which will help to attach the needle to the tube holder. Needle size: 22G x 38mm Sterile, individually peel packed Box of 100 Unit: Box |  | Zanotone (Pty) Ltd | 2.22 | Mekon |  | Mekon | [6 empty] | Each | No brand name No product code
42142521-00003 | 79 | Needle, use for blood collecting inherent safety device. Needle must be a multiple sample collecting needle, thin wall. Patient side: The needle must be always position bevel-up towards the safety device which will eliminates the repositioning of the needle. The needle must be capped with is colour coded black on the patient side. Safety device must be an integral part of the device and must cover the needle completely. It must have a permanently, securely lock in place to eliminate hard-surface activation. It must easy operation, single-handed activation immediately after exiting the venepuncture site. In the middle of the needle it will contained screw-in tread which will help to attach the needle to the tube holder. Needle size: 22G x 38mm Sterile, individually peel packed Box of 100 Unit: Box |  | FOCUS PRODUCTS (PTY) LTD | 3.02 | GENERIC |  |  | CHINA | 45 | 90 | USD | 100 BOXES | YES | NONE | No sample submitted for evaluation
42142521-00003 | 79 | Needle, use for blood collecting inherent safety device. Needle must be a multiple sample collecting needle, thin wall. Patient side: The needle must be always position bevel-up towards the safety device which will eliminates the repositioning of the needle. The needle must be capped with is colour coded black on the patient side. Safety device must be an integral part of the device and must cover the needle completely. It must have a permanently, securely lock in place to eliminate hard-surface activation. It must easy operation, single-handed activation immediately after exiting the venepuncture site. In the middle of the needle it will contained screw-in tread which will help to attach the needle to the tube holder. Needle size: 22G x 38mm Sterile, individually peel packed Box of 100 Unit: Box |  | SUPRA HEALTHCARE JOHANNESBURG PTY LTD | 3.45 | Avacare | BLDCOLS23G | HENAN AILE INDUSTRIAL CO.,LTD | CHINA | Initial 60 days, thereafter 14 days | 84 | USD | 1 Box | YES |  | No sample submitted for evaluation
42142521-00003 | 79 | Needle, use for blood collecting inherent safety device. Needle must be a multiple sample collecting needle, thin wall. Patient side: The needle must be always position bevel-up towards the safety device which will eliminates the repositioning of the needle. The needle must be capped with is colour coded black on the patient side. Safety device must be an integral part of the device and must cover the needle completely. It must have a permanently, securely lock in place to eliminate hard-surface activation. It must easy operation, single-handed activation immediately after exiting the venepuncture site. In the middle of the needle it will contained screw-in tread which will help to attach the needle to the tube holder. Needle size: 22G x 38mm Sterile, individually peel packed Box of 100 Unit: Box | Offer 1 of 1 | SSEM Mthembu Medical (Pty) Ltd | 4.37 | Disera | DT24123190 | Disera | Turkey | 56 | 85% | USD | 1 | Yes | 100/Box | The package is tear off instead of a peel apart package Offered a needle size of 23G instead of 22G
42142505-00000 | 80 | Tube holder, used during multi blood collection. Device must be constructed from transparent plastic. Nozzle of the tube holder must be concentric which must contain screw-in tread in the inside of the nozzle where different sizes of multiple blood collection needles will fit. The tube holder must have a collar on the proximal end of the holder which will aloud easy changing of tubes. Non sterile, packed Box of 25 Unit: Box |  | GIBELA TRADE AND INVEST 1001 CC | 8.24 | TARA | TUB | Shandong Qiaopai Group Co., Ltd | CHINA | 21 | 85 | USD | 10 | YES | box 25 | COMPLY |  | Recommended | 100
42142505-00000 | 80 | Tube holder, used during multi blood collection. Device must be constructed from transparent plastic. Nozzle of the tube holder must be concentric which must contain screw-in tread in the inside of the nozzle where different sizes of multiple blood collection needles will fit. The tube holder must have a collar on the proximal end of the holder which will aloud easy changing of tubes. Non sterile, packed Box of 25 Unit: Box | Offer 1 of 1 | SSEM Mthembu Medical (Pty) Ltd | 0.37 | Disera | DT260001 | Disera | Turkey | 56 | 85% | USD | 1 | Yes | 25/Pack | COMPLY
42142505-00000 | 80 | Tube holder, used during multi blood collection. Device must be constructed from transparent plastic. Nozzle of the tube holder must be concentric which must contain screw-in tread in the inside of the nozzle where different sizes of multiple blood collection needles will fit. The tube holder must have a collar on the proximal end of the holder which will aloud easy changing of tubes. Non sterile, packed Box of 25 Unit: Box | N/A | PROMED TECHNOLOGIES (PTY) LTD | 0.43 | U-LIFE | 505000 | JIAXING | CHINA | 14 DAYS | 85% | ZAR | 1 BOX | YES | BOX | No samples submitted for evaluation
42142505-00000 | 80 | Tube holder, used during multi blood collection. Device must be constructed from transparent plastic. Nozzle of the tube holder must be concentric which must contain screw-in tread in the inside of the nozzle where different sizes of multiple blood collection needles will fit. The tube holder must have a collar on the proximal end of the holder which will aloud easy changing of tubes. Non sterile, packed Box of 25 Unit: Box |  | Medi-Core Technologies | 0.44 | ULIFE | 42142505-0000 | JIAXING U-LIFE MEDICAL DEVICE TECHNOLOGY CO.,LTD | CHINA | 21 | 80 | USD | 1 BOX | YES | BOX | No sample submitted for evaluation
42142505-00000 | 80 | Tube holder, used during multi blood collection. Device must be constructed from transparent plastic. Nozzle of the tube holder must be concentric which must contain screw-in tread in the inside of the nozzle where different sizes of multiple blood collection needles will fit. The tube holder must have a collar on the proximal end of the holder which will aloud easy changing of tubes. Non sterile, packed Box of 25 Unit: Box | N/A | SAFARMEX (PTY) LTD | 0.54 | Saclin | To be confirmed | Zhejiang Mediunion Healthcare Group Co. Ltd | China | 3days | 80% | US$ | 20000 Pieces | Yes | Box of 25 | No sample submitted for evaluation
42142505-00000 | 80 | Tube holder, used during multi blood collection. Device must be constructed from transparent plastic. Nozzle of the tube holder must be concentric which must contain screw-in tread in the inside of the nozzle where different sizes of multiple blood collection needles will fit. The tube holder must have a collar on the proximal end of the holder which will aloud easy changing of tubes. Non sterile, packed Box of 25 Unit: Box |  | SUPRA HEALTHCARE JOHANNESBURG PTY LTD | 0.58 | CD Newmed | HOLD001 | CD Newmed | CHINA | Initial 60 days, thereafter 14 days | 84 | USD | 10 Packs | NO | 25 Pieces/Bag | COMPLY
42142505-00000 | 80 | Tube holder, used during multi blood collection. Device must be constructed from transparent plastic. Nozzle of the tube holder must be concentric which must contain screw-in tread in the inside of the nozzle where different sizes of multiple blood collection needles will fit. The tube holder must have a collar on the proximal end of the holder which will aloud easy changing of tubes. Non sterile, packed Box of 25 Unit: Box |  | Zanotone (Pty) Ltd | 0.61 | Mekon |  | Mekon | [6 empty] | Each | COMPLY
42142505-00000 | 80 | Tube holder, used during multi blood collection. Device must be constructed from transparent plastic. Nozzle of the tube holder must be concentric which must contain screw-in tread in the inside of the nozzle where different sizes of multiple blood collection needles will fit. The tube holder must have a collar on the proximal end of the holder which will aloud easy changing of tubes. Non sterile, packed Box of 25 Unit: Box |  | UNITRADE 1032 CC | 0.61 | TUBEIN | 1 | SHANGHAI KINDLY ENTERPRISE DEVELOPMENT GROUP | ROC | 21 | 85 | USD | 1000 | YES | 25 | No brand name No product code
42142505-00000 | 80 | Tube holder, used during multi blood collection. Device must be constructed from transparent plastic. Nozzle of the tube holder must be concentric which must contain screw-in tread in the inside of the nozzle where different sizes of multiple blood collection needles will fit. The tube holder must have a collar on the proximal end of the holder which will aloud easy changing of tubes. Non sterile, packed Box of 25 Unit: Box |  | FOCUS PRODUCTS (PTY) LTD | 0.74 | GENERIC |  |  | CHINA | 45 | 90 | USD | 100 BOXES | YES | NONE | COMPLY
42142505-00000 | 80 | Tube holder, used during multi blood collection. Device must be constructed from transparent plastic. Nozzle of the tube holder must be concentric which must contain screw-in tread in the inside of the nozzle where different sizes of multiple blood collection needles will fit. The tube holder must have a collar on the proximal end of the holder which will aloud easy changing of tubes. Non sterile, packed Box of 25 Unit: Box | Basic Line Offer | BIOCOR HOSPITAL SUPPLIES PTY LTD | 0.75 | Medline | N/A | Yangzhong Medline Industry Co.Ltd | China | 60 | 85 | USD | 10 Boxes of 25 | Yes | N/A | COMPLY
42142505-00000 | 80 | Tube holder, used during multi blood collection. Device must be constructed from transparent plastic. Nozzle of the tube holder must be concentric which must contain screw-in tread in the inside of the nozzle where different sizes of multiple blood collection needles will fit. The tube holder must have a collar on the proximal end of the holder which will aloud easy changing of tubes. Non sterile, packed Box of 25 Unit: Box | WE OFFER UNIT OF MEASURE OF 100/PACK IN PLACE OF 25/PACK | LECHOBA MEDICAL TECHNOLOGIES (PTY) LTD | 0.76 | VACUMED | 45000 | FL MEDICAL | ITALY | 21 DAYS | 75% | EURO | 2500 | YES | 100/PACK | COMPLY
42142505-00000 | 80 | Tube holder, used during multi blood collection. Device must be constructed from transparent plastic. Nozzle of the tube holder must be concentric which must contain screw-in tread in the inside of the nozzle where different sizes of multiple blood collection needles will fit. The tube holder must have a collar on the proximal end of the holder which will aloud easy changing of tubes. Non sterile, packed Box of 25 Unit: Box | 1 | Biologica Pharmaceuticals Pty Ltd | 0.85 | Sol- M | 11020105001 | Sol- Millennium | China | 4-6 Weeks | 100% | USD | 1000 Boxes | 100% | N/A
42142505-00000 | 80 | Tube holder, used during multi blood collection. Device must be constructed from transparent plastic. Nozzle of the tube holder must be concentric which must contain screw-in tread in the inside of the nozzle where different sizes of multiple blood collection needles will fit. The tube holder must have a collar on the proximal end of the holder which will aloud easy changing of tubes. Non sterile, packed Box of 25 Unit: Box | Basic Line offer | Becton Dickinson | 2.91 | BD VAC 1-use Holder | 364815 | Becton Dickinson | United Kingdom | 7 Working Days | 76% | USD | 1 BOX | YES | 1 BOX = 250 UNITS | COMPLY
42142505-00000 | 80 | Tube holder, used during multi blood collection. Device must be constructed from transparent plastic. Nozzle of the tube holder must be concentric which must contain screw-in tread in the inside of the nozzle where different sizes of multiple blood collection needles will fit. The tube holder must have a collar on the proximal end of the holder which will aloud easy changing of tubes. Non sterile, packed Box of 25 Unit: Box |  | ABCARE TECHNOLOGIES (PTY) LTD | 19.99 | CHIRANA | TH001 | CHIRANA T INJEKTA | SLOVAKIA | 45 | 85 | USD | 100 | YES | PACK OF 25 | COMPLY
42142521-00004 | 81 | Needle, use for blood collecting inherent safety device. Needle must be a multiple sample collecting needle, thin wall. Patient side: The needle must be always position bevel-up towards the safety device which will eliminates the repositioning of the needle. The needle must be capped with is colour coded black on the patient side. Safety device must be an integral part of the device and must cover the needle completely. It must have a permanently, securely lock in place to eliminate hard-surface activation. It must easy operation, single-handed activation immediately after exiting the venepuncture site. In the middle of the needle it will contained screw-in tread which will help to attach the needle to the tube holder. Needle size: 21G x 38mm Sterile, individually peel packed Box of 100 Unit: Box |  | Chemical World cc | 1.61 | [10 empty] | No sample submitted for evaluation |  | Re-tender
42142521-00004 | 81 | Needle, use for blood collecting inherent safety device which is pre-attached to the tube holder. Needle must be a multiple sample collecting needle, thin wall. Patient side: The needle must be always position bevel-up towards the safety device which will eliminates the repositioning of the needle. The needle must be capped with is colour coded green on the patient side. The needle must be concentrically pre-attached to the tube holder for used during blood collection. Safety device must be an integral part of the device and must cover the needle completely. It must have a permanently, securely lock in place to eliminate hard-surface activation. It must easy operation, single-handed activation immediately after exiting the venepuncture site. The tube holder must be constructed from transparent plastic. Must have a collar on the proximal end of the holder which will aloud easy changing of tubes and reducing the risk of back-end needle stick injuries Tube side: The needle must be covered with in a latex free elastic membrane to prevent the blood spilling when changing tubes. Needle size: 21G x 38mm Sterile, individually peel packed Box of 100 Unit: Box |  | SUPRA HEALTHCARE JOHANNESBURG PTY LTD | 3.45 | Avacare | BLDCOLS21G | HENAN AILE INDUSTRIAL CO.,LTD | CHINA | Initial 60 days, thereafter 14 days | 84 | USD | 1 Box | YES |  | No sample submitted for evaluation |  | Re-tender
42142521-00004 | 81 | Needle, use for blood collecting inherent safety device which is pre-attached to the tube holder. Needle must be a multiple sample collecting needle, thin wall. Patient side: The needle must be always position bevel-up towards the safety device which will eliminates the repositioning of the needle. The needle must be capped with is colour coded green on the patient side. The needle must be concentrically pre-attached to the tube holder for used during blood collection. Safety device must be an integral part of the device and must cover the needle completely. It must have a permanently, securely lock in place to eliminate hard-surface activation. It must easy operation, single-handed activation immediately after exiting the venepuncture site. The tube holder must be constructed from transparent plastic. Must have a collar on the proximal end of the holder which will aloud easy changing of tubes and reducing the risk of back-end needle stick injuries Tube side: The needle must be covered with in a latex free elastic membrane to prevent the blood spilling when changing tubes. Needle size: 21G x 38mm Sterile, individually peel packed Box of 100 Unit: Box |  | Zanotone (Pty) Ltd | 3.47 | Mekon |  | Mekon | [6 empty] | Each | No brand name No product code |  | Re-tender
42142521-00004 | 81 | Needle, use for blood collecting inherent safety device which is pre-attached to the tube holder. Needle must be a multiple sample collecting needle, thin wall. Patient side: The needle must be always position bevel-up towards the safety device which will eliminates the repositioning of the needle. The needle must be capped with is colour coded green on the patient side. The needle must be concentrically pre-attached to the tube holder for used during blood collection. Safety device must be an integral part of the device and must cover the needle completely. It must have a permanently, securely lock in place to eliminate hard-surface activation. It must easy operation, single-handed activation immediately after exiting the venepuncture site. The tube holder must be constructed from transparent plastic. Must have a collar on the proximal end of the holder which will aloud easy changing of tubes and reducing the risk of back-end needle stick injuries Tube side: The needle must be covered with in a latex free elastic membrane to prevent the blood spilling when changing tubes. Needle size: 21G x 38mm Sterile, individually peel packed Box of 100 Unit: Box |  | FOCUS PRODUCTS (PTY) LTD | 3.91 | GENERIC |  |  | CHINA | 45 | 90 | USD | 100 BOXES | YES | NONE | No brand name. No product code. The sample does not have a needle |  | Re-tender
42142521-00004 | 81 | Needle, use for blood collecting inherent safety device which is pre-attached to the tube holder. Needle must be a multiple sample collecting needle, thin wall. Patient side: The needle must be always position bevel-up towards the safety device which will eliminates the repositioning of the needle. The needle must be capped with is colour coded green on the patient side. The needle must be concentrically pre-attached to the tube holder for used during blood collection. Safety device must be an integral part of the device and must cover the needle completely. It must have a permanently, securely lock in place to eliminate hard-surface activation. It must easy operation, single-handed activation immediately after exiting the venepuncture site. The tube holder must be constructed from transparent plastic. Must have a collar on the proximal end of the holder which will aloud easy changing of tubes and reducing the risk of back-end needle stick injuries Tube side: The needle must be covered with in a latex free elastic membrane to prevent the blood spilling when changing tubes. Needle size: 21G x 38mm Sterile, individually peel packed Box of 100 Unit: Box | WE OFFER A SAFETY BLOOD COLLECTION TUBE HOLDER WITH AN INHERENT NEEDLE PROTECTION DEVICE IN OF NEEDLE WITH INHERENT SAFETY DEVICE AND A BLOOD COLLECTION TUBE HOLDER. | LECHOBA MEDICAL TECHNOLOGIES (PTY) LTD | 4.42 | VACUMED | 45000 | FL MEDICAL | ITALY | 21 DAYS | 75% | EURO | 100 | NO. OFFER OPTION OF TUBE WITH SAFETY FEATURES WHICH CAN BE USED WITH STANDARD MULTI-SAMPLE NEEDLES. | 50/PACK | No pre-attached needle |  | Re-tender
42142521-00004 | 81 | Needle, use for blood collecting inherent safety device which is pre-attached to the tube holder. Needle must be a multiple sample collecting needle, thin wall. Patient side: The needle must be always position bevel-up towards the safety device which will eliminates the repositioning of the needle. The needle must be capped with is colour coded green on the patient side. The needle must be concentrically pre-attached to the tube holder for used during blood collection. Safety device must be an integral part of the device and must cover the needle completely. It must have a permanently, securely lock in place to eliminate hard-surface activation. It must easy operation, single-handed activation immediately after exiting the venepuncture site. The tube holder must be constructed from transparent plastic. Must have a collar on the proximal end of the holder which will aloud easy changing of tubes and reducing the risk of back-end needle stick injuries Tube side: The needle must be covered with in a latex free elastic membrane to prevent the blood spilling when changing tubes. Needle size: 21G x 38mm Sterile, individually peel packed Box of 100 Unit: Box | Basic Line offer | Becton Dickinson | 362.55 | BD VAC ECLIPSE PAH, GREEN, 21G | 368650 | Becton Dickinson | United Kingdom | 7 Working Days | 76% | USD | 1 BOX | NO | 1 BOX = 100 UNITS / LENGTH 32mm | Offered a length of 32mm instead of 38mm |  | Re-tender
42142521-00005 | 82 | Needle, use for blood collecting inherent safety device which is pre-attached to the tube holder. Needle must be a multiple sample collecting needle, thin wall. Patient side: The needle must be always position bevel-up towards the safety device which will eliminates the repositioning of the needle. The needle must be capped with is colour coded black on the patient side. The needle must be concentrically pre-attached to the tube holder for used during blood collection. Safety device must be an integral part of the device and must cover the needle completely. It must have a permanently, securely lock in place to eliminate hard-surface activation. It must easy operation, single-handed activation immediately after exiting the venepuncture site. The tube holder must be constructed from transparent plastic. Must have a collar on the proximal end of the holder which will aloud easy changing of tubes and reducing the risk of back-end needle stick injuries Tube side: The needle must be covered with in a latex free elastic membrane to prevent the blood spilling when changing tubes. Needle size: 22G x 38mm Sterile, individually peel packed Box of 100 Unit: Box |  | Chemical World cc | 1.61 | SURGI PLUS | XINLE | CHINA | 3 | 85 | USD | 5000 | YES | 100 |  | No sample submitted for evaluation
42142521-00005 | 82 | Needle, use for blood collecting inherent safety device which is pre-attached to the tube holder. Needle must be a multiple sample collecting needle, thin wall. Patient side: The needle must be always position bevel-up towards the safety device which will eliminates the repositioning of the needle. The needle must be capped with is colour coded black on the patient side. The needle must be concentrically pre-attached to the tube holder for used during blood collection. Safety device must be an integral part of the device and must cover the needle completely. It must have a permanently, securely lock in place to eliminate hard-surface activation. It must easy operation, single-handed activation immediately after exiting the venepuncture site. The tube holder must be constructed from transparent plastic. Must have a collar on the proximal end of the holder which will aloud easy changing of tubes and reducing the risk of back-end needle stick injuries Tube side: The needle must be covered with in a latex free elastic membrane to prevent the blood spilling when changing tubes. Needle size: 22G x 38mm Sterile, individually peel packed Box of 100 Unit: Box |  | SUPRA HEALTHCARE JOHANNESBURG PTY LTD | 3.45 | Avacare | BLDCOLS22G | HENAN AILE INDUSTRIAL CO.,LTD | CHINA | Initial 60 days, thereafter 14 days | 84 | USD | 1 Box | YES |  | No sample submitted for evaluation
42142521-00005 | 82 | Needle, use for blood collecting inherent safety device which is pre-attached to the tube holder. Needle must be a multiple sample collecting needle, thin wall. Patient side: The needle must be always position bevel-up towards the safety device which will eliminates the repositioning of the needle. The needle must be capped with is colour coded black on the patient side. The needle must be concentrically pre-attached to the tube holder for used during blood collection. Safety device must be an integral part of the device and must cover the needle completely. It must have a permanently, securely lock in place to eliminate hard-surface activation. It must easy operation, single-handed activation immediately after exiting the venepuncture site. The tube holder must be constructed from transparent plastic. Must have a collar on the proximal end of the holder which will aloud easy changing of tubes and reducing the risk of back-end needle stick injuries Tube side: The needle must be covered with in a latex free elastic membrane to prevent the blood spilling when changing tubes. Needle size: 22G x 38mm Sterile, individually peel packed Box of 100 Unit: Box |  | Zanotone (Pty) Ltd | 3.47 | Mekon |  | Mekon | [6 empty] | Each | No sample submitted for evaluation
42142521-00005 | 82 | Needle, use for blood collecting inherent safety device which is pre-attached to the tube holder. Needle must be a multiple sample collecting needle, thin wall. Patient side: The needle must be always position bevel-up towards the safety device which will eliminates the repositioning of the needle. The needle must be capped with is colour coded black on the patient side. The needle must be concentrically pre-attached to the tube holder for used during blood collection. Safety device must be an integral part of the device and must cover the needle completely. It must have a permanently, securely lock in place to eliminate hard-surface activation. It must easy operation, single-handed activation immediately after exiting the venepuncture site. The tube holder must be constructed from transparent plastic. Must have a collar on the proximal end of the holder which will aloud easy changing of tubes and reducing the risk of back-end needle stick injuries Tube side: The needle must be covered with in a latex free elastic membrane to prevent the blood spilling when changing tubes. Needle size: 22G x 38mm Sterile, individually peel packed Box of 100 Unit: Box | Basic Line offer | Becton Dickinson | 3.61 | BD VAC ECLIPSE, PAH, BLACK, 22G | 368651 | Becton Dickinson | United Kingdom | 7 Working Days | 76% | USD | 1 BOX | NO | 1 BOX = 100 UNITS / LENGTH 32mm | Offered a length of 32mm instead of 38mm
42142608-00000 | 83 | Syringe, two-part, plastic, latex free, non-pyrogenic and DEPH free. Type 2 (luer slip). Nominal capacity 2ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs UNIT: Box (Box of 100) | Basic Line offer | Becton Dickinson | 28 | BD 2ml Discardit | 300928 | Becton Dickinson | Fraga, Spain | 7 Working Days | 76% | USD | 1 BOX | YES | 100 Units in a box | COMPLY pending confirmation of latex and DEHP free | COMPLY | Recommended | 91
42142608-00000 | 83 | Syringe, two-part, plastic, latex free, non-pyrogenic and DEPH free. Type 2 (luer slip). Nominal capacity 2ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs UNIT: Box (Box of 100) |  | Isigidi Trading 95 cc | 0.31 | Isigidi | SY08 | Jiangsu Jichun Medical Devices Co. Ltd. | China | 14 days | 85% | USD | 30 boxes | Yes | 100pcs/box | The outer package indicates 2ml and immediate package indicates 2.5ml | COMPLY
42142608-00000 | 83 | Syringe, two-part, plastic, latex free, non-pyrogenic and DEPH free. Type 2 (luer slip). Nominal capacity 2ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs UNIT: Box (Box of 100) | 1 | B. BRAUN MEDICAL (PTY) LTD | 0.32 | Injekt 2ml | 4606027V | B. Braun | Germany | 42 | 50% | Euro | 100 | Yes | Each | offered a nominal capacity 3ml instead of 2ml | NO TEST REPORTS
42142608-00000 | 83 | Syringe, two-part, plastic, latex free, non-pyrogenic and DEPH free. Type 2 (luer slip). Nominal capacity 2ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs UNIT: Box (Box of 100) |  | ALMIKA TRADING PTY LTD | 0.33 | BD | 300928 | BD | USA | 14 | 0 | NO | 100 | YES | 1 | COMPLY pending confirmation of latex free | COMPLY
42142608-00000 | 83 | Syringe, two-part, plastic, latex free, non-pyrogenic and DEPH free. Type 2 (luer slip). Nominal capacity 2ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs UNIT: Box (Box of 100) |  | Chemical World cc | 0.35 | SURG PLUS | SYR202 | TONGDA | china | 3 | 85 | USD | 18000 | YES | EACH | COMPLY pnding confirmation of latex free | NO TEST REPORTS
42142608-00000 | 83 | Syringe, two-part, plastic, latex free, non-pyrogenic and DEPH free. Type 2 (luer slip). Nominal capacity 2ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs UNIT: Box (Box of 100) |  | Mr First Aid (Pty) Ltd | 0.38 | Yue Kang |  | Yue Kang | China | 42 | 85 | USD | 10 | Yes |  | COMPLY pending confirmation of latex and DEHP free | NO TEST REPORTS
42142608-00000 | 83 | Syringe, two-part, plastic, latex free, non-pyrogenic and DEPH free. Type 2 (luer slip). Nominal capacity 2ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs UNIT: Box (Box of 100) | N/A | SAFARMEX (PTY) LTD | 0.48 | Saclin | To be confirmed | Zhejiang Mediunion Healthcare Group Co. Ltd | China | 3days | 80% | US$ | 3000 Boxes | Yes | Box of 100 | No sample submitted for evaluation | NO TEST REPORT
42142608-00000 | 83 | Syringe, two-part, plastic, latex free, non-pyrogenic and DEPH free. Type 2 (luer slip). Nominal capacity 2ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs UNIT: Box (Box of 100) |  | GIBELA TRADE AND INVEST 1001 CC | 0.49686 | TARA | SYR001 | Shandong Qiaopai Group Co., Ltd | CHINA | 21 | 85 | USD | 10 | YES | each | COMPLY pending confirmation of latex and DEHP free | NOT COMPLY
42142608-00000 | 83 | Syringe, two-part, plastic, latex free, non-pyrogenic and DEPH free. Type 2 (luer slip). Nominal capacity 2ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs UNIT: Box (Box of 100) |  | SUPRA HEALTHCARE JOHANNESBURG PTY LTD | 0.54 | ABC | BS2ML2 | Revital Health | KENYA | Initial 60 days, thereafter 21 | 84 | USD | 10 Boxes | YES |  | COMPLY pending confirmation of latex and DEHP free | NO TEST REPORTS
42142608-00000 | 83 | Syringe, two-part, plastic, latex free, non-pyrogenic and DEPH free. Type 2 (luer slip). Nominal capacity 2ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs UNIT: Box (Box of 100) | 1 | Biologica Pharmaceuticals Pty Ltd | 0.64 | Sol- M | 313001 | Sol- Millennium | China | 4-6 Weeks | 100% | USD | 1000 Boxes | 100% | N/A
42142608-00004 | 84 | Syringe, two-part, plastic, latex free, non-pyrogenic and DEPH free. Type 3 (luer lock). Nominal capacity 2ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 100 UNIT: Box |  | Isigidi Trading 95 cc | 0.31 | Isigidi | SY31 | Jiangsu Jichun Medical Devices Co. Ltd. | China | 14 days | 85% | USD | 30 boxes | Yes | 100pcs/box | No sample submitted for evaluation | No test report | Re-tender
42142608-00004 | 84 | Syringe, two-part, plastic, latex free, non-pyrogenic and DEPH free. Type 2 (luer lock). Nominal capacity 2ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs UNIT: Box (Box of 100) |  | Chemical World cc | 0.33 | [10 empty] | No sample submitted for evaluation | NO TEST REPORTS | Re-tender
42142608-00004 | 84 | Syringe, two-part, plastic, latex free, non-pyrogenic and DEPH free. Type 3 (luer lock). Nominal capacity 2ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 100 UNIT: Box | N/A | PROMED TECHNOLOGIES (PTY) LTD | 0.35 | U-LIFE | 608004 | JIAXING | CHINA | 14 DAYS | 85% | USD | 1 UNIT | YES | BOX | No samples submitted for evaluation | NO TEST REPORTS | Re-tender
42142608-00004 | 84 | Syringe, two-part, plastic, latex free, non-pyrogenic and DEPH free. Type 3 (luer lock). Nominal capacity 2ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 100 UNIT: Box |  | Medi-Core Technologies | 0.36 | U-LIFE | 42142608-00004 | JIAXING U-LIFE MEDICAL DEVICE TECHNOLOGY CO.,LTD | CHINA | 21 | 80 | USD | 1 BOX | YES | BOX | No sample submitted for evaluation | NO TEST REPORTS | Re-tender
42142608-00004 | 84 | Syringe, two-part, plastic, latex free, non-pyrogenic and DEPH free. Type 3 (luer lock). Nominal capacity 2ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 100 UNIT: Box |  | Mr First Aid (Pty) Ltd | 0.4 | Yue Kang |  | Yue Kang | China | 42 | 85 | USD | 10 | Yes |  | Wrong sample, offered a three part instead of a two part syringe. Offered a luer slip instead luer lock. | NO TEST REPORTS | Re-tender
42142608-00004 | 84 | Syringe, two-part, plastic, latex free, non-pyrogenic and DEPH free. Type 3 (luer lock). Nominal capacity 2ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 100 UNIT: Box |  | SUPRA HEALTHCARE JOHANNESBURG PTY LTD | 0.47 | ABC | BS2ML2 | Revital Health | KENYA | Initial 60 days, thereafter 21 | 84 | USD | 10 Boxes | YES |  | No sample submitted for evaluation | NO TEST REPORTS | Re-tender
42142608-00004 | 84 | Syringe, two-part, plastic, latex free, non-pyrogenic and DEPH free. Type 3 (luer lock). Nominal capacity 2ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 100 UNIT: Box | Basic Line Offer | BIOCOR HOSPITAL SUPPLIES PTY LTD | 0.5 | Triple -D | SYR-Care 2ml | Jumla Medical | South Africa | 21 | 0 | N/A | 10 Boxes of 100 | Yes | N/A | No sample submitted for evaluation | NO TEST REPORTS | Re-tender
42142608-00004 | 84 | Syringe, two-part, plastic, latex free, non-pyrogenic and DEPH free. Type 3 (luer lock). Nominal capacity 2ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 100 UNIT: Box |  | GIBELA TRADE AND INVEST 1001 CC | 0.65706 | TARA | SYR002 | Shandong Qiaopai Group Co., Ltd | CHINA | 21 | 85 | USD | 10 | YES | each | No sample submitted for evaluation | NO TEST REPORTS | Re-tender
42142608-00021 | 85 | Syringe, two-part, plastic, latex free, non-pyrogenic and DEPH free. Type 2 (luer slip), with a concentric nozzle. Nominal capacity 5ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 100 UNIT: Box | Basic Line offer | Becton Dickinson | 0.35 | BD 5ml Discardit | 309050 | Becton Dickinson | Fraga, Spain | 7 Working Days | 76% | USD | 1 BOX | YES | 100 Units in a box | No sample submitted for evaluation. | NO TEST REPORT | Re-tender
42142608-00021 | 85 | Syringe, two-part, plastic, latex free, non-pyrogenic and DEPH free. Type 2 (luer slip), with a concentric nozzle. Nominal capacity 5ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 100 UNIT: Box | N/A | PROMED TECHNOLOGIES (PTY) LTD | 0.35 |  |  | JIAXING | CHINA | 15 DAYS | 85% | USD | 2 UNIT | YES | BOX | No samples submitted for evaluation | NO TEST REPORTS | Re-tender
42142608-00021 | 85 | Syringe, two-part, plastic, latex free, non-pyrogenic and DEPH free. Type 2 (luer slip), with a concentric nozzle. Nominal capacity 5ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 100 UNIT: Box |  | Chemical World cc | 0.38 | SURG PLUS | SYR205 | TONGDA | china | 3 | 85 | USD | 18000 | YES | 100 | COMPLY pending confirmation of latex free | NO TEST REPORTS | Re-tender
42142608-00021 | 85 | Syringe, two-part, plastic, latex free, non-pyrogenic and DEPH free. Type 2 (luer slip), with a concentric nozzle. Nominal capacity 5ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 100 UNIT: Box |  | Mr First Aid (Pty) Ltd | 0.4 | Yue Kang |  | Yue Kang | China | 42 | 85 | USD | 10 | Yes |  | COMPLY pending confirmation of latex and DEHP free | NO TEST REPORTS | Re-tender
42142608-00021 | 85 | Syringe, two-part, plastic, latex free, non-pyrogenic and DEPH free. Type 2 (luer slip), with a concentric nozzle. Nominal capacity 5ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 100 UNIT: Box | N/A | SAFARMEX (PTY) LTD | 0.48 | Saclin | To be confirmed | Zhejiang Mediunion Healthcare Group Co. Ltd | China | 3days | 80% | US$ | 3000 Boxes | Yes | Box of 100 | No sample submitted for evaluation | NO TEST REPORT | Re-tender
42142608-00021 | 85 | Syringe, two-part, plastic, latex free, non-pyrogenic and DEPH free. Type 2 (luer slip), with a concentric nozzle. Nominal capacity 5ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 100 UNIT: Box |  | SUPRA HEALTHCARE JOHANNESBURG PTY LTD | 0.62 | ABC | BS5ML2C | Revital Health | KENYA | Initial 60 days, thereafter 21 | 84 | USD | 10 Boxes | YES |  | COMPLY pending confirmation of latex and DEHP free | NO TEST REPORTS | Re-tender
42142608-00021 | 85 | Syringe, two-part, plastic, latex free, non-pyrogenic and DEPH free. Type 2 (luer slip), with a concentric nozzle. Nominal capacity 5ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 100 UNIT: Box |  | GIBELA TRADE AND INVEST 1001 CC | 0.66 | TARA | SYR003 | Shandong Qiaopai Group Co., Ltd | CHINA | 21 | 85 | USD | 10 | YES | box 100 | No sample submitted for evaluation | NO TEST REPORTS | Re-tender
42142608-00021 | 85 | Syringe, two-part, plastic, latex free, non-pyrogenic and DEPH free. Type 2 (luer slip), with a concentric nozzle. Nominal capacity 5ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 100 UNIT: Box | 1 | Biologica Pharmaceuticals Pty Ltd | 0.92 | Sol- M | 307005 | Sol- Millennium | China | 4-6 Weeks | 100% | USD | 1000 Boxes | 100% | N/A |  | NO TEST REPORTS | Re-tender
42142608-00022 | 86 | Syringe, two-part, plastic, latex free, non-pyrogenic and DEPH free. Type 2 (luer slip), with an eccentric nozzle. Nominal capacity 5ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 100 UNIT: Box |  | Isigidi Trading 95 cc | 34.96 | Isigidi | SY22 | Jiangsu Jichun Medical Devices Co. Ltd. | China | 14 days | 85% | USD | 20 boxes | Yes | 100pcs/box | COMPLY pending confirmation of latex and DEHP free | COMPLY | Recommended | 91
42142608-00022 | 86 | Syringe, two-part, plastic, latex free, non-pyrogenic and DEPH free. Type 2 (luer slip), with an eccentric nozzle. Nominal capacity 5ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 100 UNIT: Box |  | Chemical World cc | 0.38 | SURG PLUS | SYRE205 | TONGDA | china | 3 | 85 | USD | 18000 | YES | 100 | COMPLY pending confirmation of latex free | NO TEST REPORTS
42142608-00022 | 86 | Syringe, two-part, plastic, latex free, non-pyrogenic and DEPH free. Type 2 (luer slip), with an eccentric nozzle. Nominal capacity 5ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 100 UNIT: Box |  | Mr First Aid (Pty) Ltd | 0.4 | Yue Kang |  | Yue Kang | China | 42 | 85 | USD | 10 | Yes |  | COMPLY pending confirmation of latex and DEHP free | NO TEST REPORTS
42142608-00022 | 86 | Syringe, two-part, plastic, latex free, non-pyrogenic and DEPH free. Type 2 (luer slip), with an eccentric nozzle. Nominal capacity 5ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 100 UNIT: Box | 1 | B. BRAUN MEDICAL (PTY) LTD | 0.4 | Injekt 5ml | 4606051V | B. Braun | Germany | 42 | 50% | Euro | 100 | Yes | Box of 100 | COMPLY pending confirmation of latex and DEHP free | COMPLY
42142608-00022 | 86 | Syringe, two-part, plastic, latex free, non-pyrogenic and DEPH free. Type 2 (luer slip), with an eccentric nozzle. Nominal capacity 5ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 100 UNIT: Box |  | ALMIKA TRADING PTY LTD | 0.42 | BD | 309050 | BD | USA | 14 | 0 | NO | 100 | YES | 1 | COMPLY pending confirmation of latex free | COMPLY
42142608-00022 | 86 | Syringe, two-part, plastic, latex free, non-pyrogenic and DEPH free. Type 2 (luer slip), with an eccentric nozzle. Nominal capacity 5ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 100 UNIT: Box | 1 | Vital Life Medical Distributors (PTY) LTD | 0.42 | Healthease | NHS01-5ML | Neomedic | China | 14 days | 80% | US Dollar | 50 boxes | Yes | 100's per box | COMPLY pending confirmation of latex and DEHP free | NO TEST REPORTS
42142608-00022 | 86 | Syringe, two-part, plastic, latex free, non-pyrogenic and DEPH free. Type 2 (luer slip), with an eccentric nozzle. Nominal capacity 5ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 100 UNIT: Box | N/A | SAFARMEX (PTY) LTD | 0.48 | Saclin | To be confirmed | Zhejiang Mediunion Healthcare Group Co. Ltd | China | 3days | 80% | US$ | 3000 Boxes | Yes | Box of 100 | No sample submitted for evaluation | NO TEST REPORT
42142608-00022 | 86 | Syringe, two-part, plastic, latex free, non-pyrogenic and DEPH free. Type 2 (luer slip), with an eccentric nozzle. Nominal capacity 5ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 100 UNIT: Box |  | GIBELA TRADE AND INVEST 1001 CC | 0.66 | TARA | SYR004 | Shandong Qiaopai Group Co., Ltd | CHINA | 21 | 85 | USD | 10 | YES | box 100 | COMPLY pending confirmation and latex and DEHP free | NOT COMPLY
42142608-00022 | 86 | Syringe, two-part, plastic, latex free, non-pyrogenic and DEPH free. Type 2 (luer slip), with an eccentric nozzle. Nominal capacity 5ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 100 UNIT: Box |  | SUPRA HEALTHCARE JOHANNESBURG PTY LTD | 0.8 | Avacare | BS5ML2E | HENAN AILE INDUSTRIAL CO.,LTD | CHINA | Initial 60 days, thereafter 21 | 84 | USD | 10 Boxes | YES |  | No sample submitted for evaluation | NO TEST REPORTS
42142608-00022 | 86 | Syringe, two-part, plastic, latex free, non-pyrogenic and DEPH free. Type 2 (luer slip), with an eccentric nozzle. Nominal capacity 5ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 100 UNIT: Box | 1 | Biologica Pharmaceuticals Pty Ltd | 2.14 | Sol- M | 1800055T | Sol- Millennium | China | 4-6 Weeks | 100% | USD | 1000 Boxes | 100% | N/A
42142608-00023 | 87 | Syringe, two-part, plastic, latex free, non-pyrogenic and DEPH free. Type 3 (luer lock), with a concentric nozzle. Nominal capacity 5ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 100 UNIT: Box |  | Mr First Aid (Pty) Ltd | 0.4 | Yue Kang |  | Yue Kang | China | 42 | 85 | USD | 10 | Yes |  | No sample submitted for evaluation | NO TEST REPORTS | Re-tender
42142608-00023 | 87 | Syringe, two-part, plastic, latex free, non-pyrogenic and DEPH free. Type 3 (luer lock), with a concentric nozzle. Nominal capacity 5ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 100 UNIT: Box |  | SUPRA HEALTHCARE JOHANNESBURG PTY LTD | 0.62 | ABC | BS5ML2L/LC | Revital Health | KENYA | Initial 60 days, thereafter 21 | 84 | USD | 10 Boxes | YES |  | No sample submitted for evaluation | NO TEST REPORTS | Re-tender
42142608-00023 | 87 | Syringe, two-part, plastic, latex free, non-pyrogenic and DEPH free. Type 3 (luer lock), with a concentric nozzle. Nominal capacity 5ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 100 UNIT: Box |  | GIBELA TRADE AND INVEST 1001 CC | 0.66 | TARA | SYR005 | Shandong Qiaopai Group Co., Ltd | CHINA | 21 | 85 | USD | 10 | YES | box 100 | No sample submitted for evaluation | NO TEST REPORTS | Re-tender
42142608-00023 | 87 | Syringe, two-part, plastic, latex free, non-pyrogenic and DEPH free. Type 3 (luer lock), with a concentric nozzle. Nominal capacity 5ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 100 UNIT: Box | 1 | Biologica Pharmaceuticals Pty Ltd | 1.43 | Sol- M | 180005 | Sol- Millennium | China | 4-6 Weeks | 100% | USD | 1000 Boxes | 100% | N/A |  | NO TEST REPORTS | Re-tender
42142608-00024 | 88 | Syringe, two-part, plastic, latex free, non-pyrogenic and DEPH free. Type 3 (luer lock), with an eccentric nozzle. Nominal capacity 5ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 100 UNIT: Box |  | Mr First Aid (Pty) Ltd | 0.4 | Yue Kang |  | Yue Kang | China | 42 | 85 | USD | 10 | Yes |  | No sample submitted for evaluation | NO TEST REPORTS | Re-tender
42142608-00024 | 88 | Syringe, two-part, plastic, latex free, non-pyrogenic and DEPH free. Type 3 (luer lock), with an eccentric nozzle. Nominal capacity 5ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 100 UNIT: Box |  | GIBELA TRADE AND INVEST 1001 CC | 0.66 | TARA | SYR006 | Shandong Qiaopai Group Co., Ltd | CHINA | 21 | 85 | USD | 10 | YES | box 100 | No sample submitted for evaluation | NO TEST REPORTS | Re-tender
42142608-00024 | 88 | Syringe, two-part, plastic, latex free, non-pyrogenic and DEPH free. Type 3 (luer lock), with an eccentric nozzle. Nominal capacity 5ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 100 UNIT: Box | 1 | B. BRAUN MEDICAL (PTY) LTD | 0.7 | Omnifix 5ml Luer Lock | 4617053V | B. Braun | Germany | 42 | 50% | Euro | 100 | Yes | Box of 100 | offered a luer slip instead of luer lock | NO TEST REPORTS | Re-tender
42142608-00025 | 89 | Syringe, two-part, plastic, latex free, non-pyrogenic and DEPH free. Type 2 (luer slip), with a concentric nozzle. Nominal capacity 10ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 100 UNIT: Box | Basic Line offer | Becton Dickinson | 0.54 | BD 10ml Discardit | 309110 | Becton Dickinson | Fraga, Spain | 7 Working Days | 76% | USD | 1 BOX | YES | 100 units in a box | No sample submitted for evaluation. | NO TEST REPORT | Re-tender
42142608-00025 | 89 | Syringe, two-part, plastic, latex free, non-pyrogenic and DEPH free. Type 2 (luer slip), with a concentric nozzle. Nominal capacity 10ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 100 UNIT: Box |  | Mr First Aid (Pty) Ltd | 0.64 | Yue Kang |  | Yue Kang | China | 42 | 85 | USD | 10 | Yes |  | COMPLY pending confirmation of latex and DEHP free | NO TEST REPORTS | Re-tender
42142608-00025 | 89 | Syringe, two-part, plastic, latex free, non-pyrogenic and DEPH free. Type 2 (luer slip), with a concentric nozzle. Nominal capacity 10ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 100 UNIT: Box |  | GIBELA TRADE AND INVEST 1001 CC | 0.95 | TARA | SYR007 | Shandong Qiaopai Group Co., Ltd | CHINA | 21 | 85 | USD | 10 | YES | box 100 | No sample submitted for evaluation | NOT COMPLY | Re-tender
42142608-00025 | 89 | Syringe, two-part, plastic, latex free, non-pyrogenic and DEPH free. Type 2 (luer slip), with a concentric nozzle. Nominal capacity 10ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 100 UNIT: Box |  | SUPRA HEALTHCARE JOHANNESBURG PTY LTD | 1.17 | DCK | BS10ML2C | Revital Health | KENYA | Initial 60 days, thereafter 21 | 84 | USD | 10 Boxes | YES |  | COMPLY pending confirmation of latex and DEHP free | NO TEST REPORTS | Re-tender
42142608-00025 | 89 | Syringe, two-part, plastic, latex free, non-pyrogenic and DEPH free. Type 2 (luer slip), with a concentric nozzle. Nominal capacity 10ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 100 UNIT: Box | 1 | Biologica Pharmaceuticals Pty Ltd | 2.14 | Sol- M | 180010ST | Sol- Millennium | China | 4-6 Weeks | 100% | USD | 1000 Boxes | 100% | N/A |  | NO TEST REPORTS | Re-tender
42142608-00026 | 90 | Syringe, two-part, plastic, latex free, non-pyrogenic and DEPH free. Type 2 (luer slip), with an eccentric nozzle. Nominal capacity 10ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 100 UNIT: Box |  | GIBELA TRADE AND INVEST 1001 CC | 0.52 | TARA | SYR008 | Shandong Qiaopai Group Co., Ltd | CHINA | 21 | 85 | USD | 10 | YES | box 100 | COMPLY pending confirmation of latex and DEHP free | NOT COMPLY
42142608-00026 | 90 | Syringe, two-part, plastic, latex free, non-pyrogenic and DEPH free. Type 2 (luer slip), with an eccentric nozzle. Nominal capacity 10ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 100 UNIT: Box |  | Isigidi Trading 95 cc | 63.25 | Isigidi | SY23 | Jiangsu Jichun Medical Devices Co. Ltd. | China | 14 days | 85% | USD | 10boxes | Yes | 100pcs/box | COMPLY pending confirmation of latex and DEHP free | COMPLY | Recommended | 91
42142608-00026 | 90 | Syringe, two-part, plastic, latex free, non-pyrogenic and DEPH free. Type 2 (luer slip), with an eccentric nozzle. Nominal capacity 10ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 100 UNIT: Box |  | Mr First Aid (Pty) Ltd | 0.64 | Yue Kang |  | Yue Kang | China | 42 | 85 | USD | 10 | Yes |  | COMPLY pending confirmation of latex and DEHP free | NO TEST REPORTS
42142608-00026 | 90 | Syringe, two-part, plastic, latex free, non-pyrogenic and DEPH free. Type 2 (luer slip), with an eccentric nozzle. Nominal capacity 10ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 100 UNIT: Box |  | ALMIKA TRADING PTY LTD | 0.66 | BD | 309110 | BD | USA | 14 | 0 | NO | 100 | YES | 1 | COMPLY pending confirmation of latex free | COMPLY
42142608-00026 | 90 | Syringe, two-part, plastic, latex free, non-pyrogenic and DEPH free. Type 2 (luer slip), with an eccentric nozzle. Nominal capacity 10ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 100 UNIT: Box | 1 | B. BRAUN MEDICAL (PTY) LTD | 0.66 | Injekt 10ml | 4606108V | B. Braun | Germany | 42 | 50% | Euro | 100 | Yes | Box of 100 | COMPLY pending confirmation of latex and DEHP free | COMPLY
42142608-00026 | 90 | Syringe, two-part, plastic, latex free, non-pyrogenic and DEPH free. Type 2 (luer slip), with an eccentric nozzle. Nominal capacity 10ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 100 UNIT: Box |  | Chemical World cc | 0.69 | SURG PLUS | SYRE210 | TONGDA | china | 3 | 85 | USD | 32000 | YES | 100 | COMPLY pending confirmation of latex free | NO TEST REPORTS
42142608-00026 | 90 | Syringe, two-part, plastic, latex free, non-pyrogenic and DEPH free. Type 2 (luer slip), with an eccentric nozzle. Nominal capacity 10ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 100 UNIT: Box | 1 | Vital Life Medical Distributors (PTY) LTD | 0.69 | Healthease | NHS01-10ML | Neomedic | China | 14 days | 80% | US Dollar | 50 boxes | Yes | 100's per box | COMPLY pending confirmation of latex and DEHP free | NO TEST REPORTS
42142608-00026 | 90 | Syringe, two-part, plastic, latex free, non-pyrogenic and DEPH free. Type 2 (luer slip), with an eccentric nozzle. Nominal capacity 10ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 100 UNIT: Box | N/A | SAFARMEX (PTY) LTD | 0.87 | Saclin | To be confirmed | Zhejiang Mediunion Healthcare Group Co. Ltd | China | 3days | 80% | US$ | 3000 Boxes | Yes | Box of 100 | No sample submitted for evaluation | NO TEST REPORT
42142608-00026 | 90 | Syringe, two-part, plastic, latex free, non-pyrogenic and DEPH free. Type 2 (luer slip), with an eccentric nozzle. Nominal capacity 10ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 100 UNIT: Box |  | SUPRA HEALTHCARE JOHANNESBURG PTY LTD | 1.23 | Avacare | BS10ML2E | HENAN AILE INDUSTRIAL CO.,LTD | CHINA | Initial 60 days, thereafter 21 | 84 | USD | 10 Boxes | YES |  | No sample submitted for evaluation | NO TEST REPORTS
42142608-00026 | 90 | Syringe, two-part, plastic, latex free, non-pyrogenic and DEPH free. Type 2 (luer slip), with an eccentric nozzle. Nominal capacity 10ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 100 UNIT: Box | 1 | Biologica Pharmaceuticals Pty Ltd | 2.14 | Sol- M | 1800PET | Sol- Millennium | China | 4-6 Weeks | 100% | USD | 1000 Boxes | 100% | N/A
42142608-00027 | 91 | Syringe, two-part, plastic, latex free, non-pyrogenic and DEPH free. Type 3 (luer lock), with a concentric nozzle. Nominal capacity 10ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 100 UNIT: Box |  | Mr First Aid (Pty) Ltd | 0.64 | Yue Kang |  | Yue Kang | China | 42 | 85 | USD | 10 | Yes |  | No sample submitted for evaluation | NO TEST REPORTS | Re-tender
42142608-00027 | 91 | Syringe, two-part, plastic, latex free, non-pyrogenic and DEPH free. Type 3 (luer lock), with a concentric nozzle. Nominal capacity 10ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 100 UNIT: Box |  | GIBELA TRADE AND INVEST 1001 CC | 0.95 | TARA | SYR009 | Shandong Qiaopai Group Co., Ltd | CHINA | 21 | 85 | USD | 10 | YES | box 100 | No sample submitted for evaluation | NO TEST REPORTS | Re-tender
42142608-00027 | 91 | Syringe, two-part, plastic, latex free, non-pyrogenic and DEPH free. Type 3 (luer lock), with a concentric nozzle. Nominal capacity 10ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 100 UNIT: Box | 1 | Biologica Pharmaceuticals Pty Ltd | 2.21 | Sol- M | 180010 | Sol- Millennium | China | 4-6 Weeks | 100% | USD | 1000 Boxes | 100% | N/A |  | NO TEST REPORTS | Re-tender
42142608-00028 | 92 | Syringe, two-part, plastic, latex free, non-pyrogenic and DEPH free. Type 3 (luer lock), with an eccentric nozzle. Nominal capacity 10ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 100 UNIT: Box |  | Mr First Aid (Pty) Ltd | 0.64 | Yue Kang |  | Yue Kang | China | 42 | 85 | USD | 10 | Yes |  | No sample submitted for evaluation | NO TEST REPORTS | Re-tender
42142608-00028 | 92 | Syringe, two-part, plastic, latex free, non-pyrogenic and DEPH free. Type 3 (luer lock), with an eccentric nozzle. Nominal capacity 10ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 100 UNIT: Box |  | GIBELA TRADE AND INVEST 1001 CC | 0.95 | TARA | SYR010 | Shandong Qiaopai Group Co., Ltd | CHINA | 21 | 85 | USD | 10 | YES | box 100 | No sample submitted for evaluation | NO TEST REPORTS | Re-tender
42142608-00029 | 93 | Syringe, two-part, plastic, latex free, non-pyrogenic and DEPH free. Type 2 (luer slip), with a concentric nozzle. Nominal capacity 20ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 100 UNIT: Box |  | Mr First Aid (Pty) Ltd | 1 | Yue Kang |  | Yue Kang | China | 42 | 85 | USD | 10 | Yes | 50pc per box | COMPLY subject to SABS conformation of packaging | NO TEST REPORTS | Re-tender
42142608-00029 | 93 | Syringe, two-part, plastic, latex free, non-pyrogenic and DEPH free. Type 2 (luer slip), with a concentric nozzle. Nominal capacity 20ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 100 UNIT: Box |  | GIBELA TRADE AND INVEST 1001 CC | 1.35 | TARA | SYR011 | Shandong Qiaopai Group Co., Ltd | CHINA | 21 | 85 | USD | 10 | YES | box 100 | No sample submitted for evaluation | NO TEST REPORTS | Re-tender
42142608-00029 | 93 | Syringe, two-part, plastic, latex free, non-pyrogenic and DEPH free. Type 2 (luer slip), with a concentric nozzle. Nominal capacity 20ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 100 UNIT: Box | 1 | Biologica Pharmaceuticals Pty Ltd | 2.85 | Sol- M | 180020SJ | Sol- Millennium | China | 4-6 Weeks | 100% | USD | 1000 Boxes | 100% | N/A |  |  | Re-tender
42142608-00030 | 94 | Syringe, two-part, plastic, latex free, non-pyrogenic and DEPH free. Type 2 (luer slip), with an eccentric nozzle. Nominal capacity 20ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 100 UNIT: Box |  | Chemical World cc | 0.46 | SURG PLUS | SYRE220 | TONGDA | china | 3 | 85 | USD | 24000 | YES | 50 | COMPLY pending confirmation of latex free | NO TEST REPORTS
42142608-00030 | 94 | Syringe, two-part, plastic, latex free, non-pyrogenic and DEPH free. Type 2 (luer slip), with an eccentric nozzle. Nominal capacity 20ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 100 UNIT: Box |  | B. BRAUN MEDICAL (PTY) LTD | 95 | Injekt 20ml | 4606205V | B. Braun | China | 14 days | 85% | USD | 20boxes | Yes | BOX OF 100 | COMPLY pending confirmation of latex and DEHP free | COMPLY | Recommended | 91
42142608-00030 | 94 | Syringe, two-part, plastic, latex free, non-pyrogenic and DEPH free. Type 2 (luer slip), with an eccentric nozzle. Nominal capacity 20ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 100 UNIT: Box | Basic Line offer | Becton Dickinson | 0.87 | BD 20ml Discardit | 300296 | Becton Dickinson | Fraga, Spain | 7 Working Days | 76% | USD | 1 Box | Yes | 100 units in a box | COMPLY pending confirmation of latex and DEHP free | COMPLY
42142608-00030 | 94 | Syringe, two-part, plastic, latex free, non-pyrogenic and DEPH free. Type 2 (luer slip), with an eccentric nozzle. Nominal capacity 20ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 100 UNIT: Box | 1 | B. BRAUN MEDICAL (PTY) LTD | 0.95 | Injekt 20ml | 4606205V | B. Braun | Germany | 42 | 50% | Euro | 100 | Yes | Box of 100 | COMPLY pending confirmation of latex and DEHP free | COMPLY
42142608-00030 | 94 | Syringe, two-part, plastic, latex free, non-pyrogenic and DEPH free. Type 2 (luer slip), with an eccentric nozzle. Nominal capacity 20ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 100 UNIT: Box |  | Mr First Aid (Pty) Ltd | 1 | Yue Kang |  | Yue Kang | China | 42 | 85 | USD | 10 | Yes | 50pc per box | COMPLY subject to SABS conformation of packaging | NO TEST REPORTS
42142608-00030 | 94 | Syringe, two-part, plastic, latex free, non-pyrogenic and DEPH free. Type 2 (luer slip), with an eccentric nozzle. Nominal capacity 20ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 100 UNIT: Box |  | ALMIKA TRADING PTY LTD | 1.02 | BD | 300296 | BD | USA | 14 | 0 | NO | 100 | YES | 1 | COMPLY pending confirmation of latex free | COMPLY
42142608-00030 | 94 | Syringe, two-part, plastic, latex free, non-pyrogenic and DEPH free. Type 2 (luer slip), with an eccentric nozzle. Nominal capacity 20ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 100 UNIT: Box |  | GIBELA TRADE AND INVEST 1001 CC | 1.13 | TARA | SYR012 | Shandong Qiaopai Group Co., Ltd | CHINA | 21 | 85 | USD | 10 | YES | box 100 | COMPLY pending confirmation of latex and DEHP free | NOT COMPLY
42142608-00030 | 94 | Syringe, two-part, plastic, latex free, non-pyrogenic and DEPH free. Type 2 (luer slip), with an eccentric nozzle. Nominal capacity 20ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 100 UNIT: Box | 1 | Vital Life Medical Distributors (PTY) LTD | 1.27 | Healthease | NHS01-20ML | Neomedic | China | 14 days | 80% | US Dollar | 50 boxes | No | Pack size is boxes of 50's | COMPLY pending confirmation of latex and DEHP free | NO TEST REPORTS
42142608-00030 | 94 | Syringe, two-part, plastic, latex free, non-pyrogenic and DEPH free. Type 2 (luer slip), with an eccentric nozzle. Nominal capacity 20ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 100 UNIT: Box | N/A | SAFARMEX (PTY) LTD | 1.36 | Saclin | To be confirmed | Zhejiang Mediunion Healthcare Group Co. Ltd | China | 3days | 80% | US$ | 3000 Boxes | Yes | Box of 100 | No sample submitted for evaluation | NO TEST REPORT
42142608-00030 | 94 | Syringe, two-part, plastic, latex free, non-pyrogenic and DEPH free. Type 2 (luer slip), with an eccentric nozzle. Nominal capacity 20ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 100 UNIT: Box |  | SUPRA HEALTHCARE JOHANNESBURG PTY LTD | 1.94 | DCK | BS20ML2E | Revital Health | KENYA | Initial 60 days, thereafter 21 | 84 | USD | 10 Boxes | YES |  | COMPLY pending confirmation of latex and DEHP free | NO TEST REPORTS
42142608-00030 | 94 | Syringe, two-part, plastic, latex free, non-pyrogenic and DEPH free. Type 2 (luer slip), with an eccentric nozzle. Nominal capacity 20ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 100 UNIT: Box | 1 | Biologica Pharmaceuticals Pty Ltd | 2.85 | Sol- M | 180020ET | Sol- Millennium | China | 4-6 Weeks | 100% | USD | 1000 Boxes | 100% | N/A
42142608-00031 | 95 | Syringe, two-part, plastic, latex free, non-pyrogenic and DEPH free. Type 3 (luer lock), with a concentric nozzle. Nominal capacity 20ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 100 UNIT: Box |  | Mr First Aid (Pty) Ltd | 1 | Yue Kang |  | Yue Kang | China | 42 | 85 | USD | 10 | Yes | 50pc per box | No sample submitted for evaluation | NO TEST REPORTS | Re-tender
42142608-00031 | 95 | Syringe, two-part, plastic, latex free, non-pyrogenic and DEPH free. Type 3 (luer lock), with a concentric nozzle. Nominal capacity 20ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 100 UNIT: Box |  | GIBELA TRADE AND INVEST 1001 CC | 1.35 | TARA | SYR013 | Shandong Qiaopai Group Co., Ltd | CHINA | 21 | 85 | USD | 10 | YES | box 100 | No sample submitted for evaluation | NO TEST REPORTS | Re-tender
42142608-00032 | 96 | Syringe, two-part, plastic, latex free, non-pyrogenic and DEPH free. Type 3 (luer lock), with an eccentric nozzle. Nominal capacity 20ml Sterile, Single use, Individually peel packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 100 UNIT: Box |  | Mr First Aid (Pty) Ltd | 1 | Yue Kang |  | Yue Kang | China | 42 | 85 | USD | 10 | Yes | 50pc per box | No sample submitted for evaluation | NO TEST REPORTS | Re-tender
42142608-00032 | 96 | Syringe, two-part, plastic, latex free, non-pyrogenic and DEPH free. Type 3 (luer lock), with an eccentric nozzle. Nominal capacity 20ml Sterile, Single use, Individually peel packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 100 UNIT: Box |  | GIBELA TRADE AND INVEST 1001 CC | 1.35 | TARA | SYR014 | Shandong Qiaopai Group Co., Ltd | CHINA | 21 | 85 | USD | 10 | YES | box 100 | No sample submitted for evaluation | NO TEST REPORTS | Re-tender
42142608-00033 | 97 | Syringe, three-part, plastic, latex free, non-pyrogenic and DEPH free. Type 2 (luer slip). Nominal capacity 2ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs UNIT: Box (Box of 100) |  | Isigidi Trading 95 cc | 0.31 | Isigidi | SY31 | Jiangsu Jichun Medical Devices Co. Ltd. | China | 14 days | 85% | USD | 30boxes | Yes | 100pcs/box | Nominal capacity 2.5ml The outer package indicates 2ml and immediate package indicates 2.5ml | COMPLY
42142608-00033 | 97 | Syringe, three-part, plastic, latex free, non-pyrogenic and DEPH free. Type 2 (luer slip). Nominal capacity 2ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs UNIT: Box (Box of 100) |  | Chemical World cc | 0.32 | SURG PLUS | SYR002 | TONGDA | china | 3 | 85 | USD | 36000 | YES | 100 | COMPLY pending confirmation of latex free | NO TEST REPORTS
42142608-00033 | 97 | Syringe, three-part, plastic, latex free, non-pyrogenic and DEPH free. Type 2 (luer slip). Nominal capacity 2ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs UNIT: Box (Box of 100) |  | Medi-Core Technologies | 33.96 | ULIFE | 42142608-00033 | JIAXING ULIFE MEDICAL DEVICE TECHNOLOGY CO.,LTD | CHINA | 21 | 80 | USD | 1 box | YES | BOX OF 100 | COMPLY pending confirmation of latex free | COMPLY | Recommended | 81.29
42142608-00033 | 97 | Syringe, three-part, plastic, latex free, non-pyrogenic and DEPH free. Type 2 (luer slip). Nominal capacity 2ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs UNIT: Box (Box of 100) | Basic Line offer | Becton Dickinson | 0.35 | BD Emerald Syringe, 2ml | 307727 | Becton Dickinson | Fraga, Spain | 7 Working Days | 76% | USD | 1 box | YES | 100 units in a box | COMPLY pending confirmation of latex free and DEHP free | COMPLY
42142608-00033 | 97 | Syringe, three-part, plastic, latex free, non-pyrogenic and DEPH free. Type 2 (luer slip). Nominal capacity 2ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs UNIT: Box (Box of 100) | 1 | CENTRAL MEDICAL PTY LTD | 0.35 | LYSA | HSCS2000 | Zhejian INI Medical Devices Co Ltd | China | 21 | 80 | China yuan | 24 boxes of 100 units | yes | box of 100 units | Labeling of the outer/immediate package indicates Nominal capacity 2/3ml instead of 2ml | NO TEST REPORTS
42142608-00033 | 97 | Syringe, three-part, plastic, latex free, non-pyrogenic and DEPH free. Type 2 (luer slip). Nominal capacity 2ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs UNIT: Box (Box of 100) | N/A | PROMED TECHNOLOGIES (PTY) LTD | 0.35 | U-LIFE | 608033 | JIAXING | CHINA | 14 DAYS | 85% | USD | 1 BOX | YES | BOX | COMPLY pending confirmation of latex and DEHP free | COMPLY
42142608-00033 | 97 | Syringe, three-part, plastic, latex free, non-pyrogenic and DEPH free. Type 2 (luer slip). Nominal capacity 2ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs UNIT: Box (Box of 100) |  | Mr First Aid (Pty) Ltd | 0.4 | Yue Kang |  | Yue Kang | China | 42 | 85 | USD | 10 | Yes |  | COMPLY subject to SABS conformation of packaging | NO TEST REPORTS
42142608-00033 | 97 | Syringe, three-part, plastic, latex free, non-pyrogenic and DEPH free. Type 2 (luer slip). Nominal capacity 2ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs UNIT: Box (Box of 100) |  | ALMIKA TRADING PTY LTD | 0.42 | BD | 307727 | BD | USA | 14 | 0 | NO | 100 | YES | 1 | COMPLY pending confirmation of latex free | COMPLY
42142608-00033 | 97 | Syringe, three-part, plastic, latex free, non-pyrogenic and DEPH free. Type 2 (luer slip). Nominal capacity 2ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs UNIT: Box (Box of 100) |  | FOCUS PRODUCTS (PTY) LTD | 0.43 | HI-TECH |  | SR MEDICAL CARE | INDIA | 45 | 90 | USD | 100 BOXES | YES | NONE | Offered a tear off package instead of an individually peel apart package | NOT COMPLY
42142608-00033 | 97 | Syringe, three-part, plastic, latex free, non-pyrogenic and DEPH free. Type 2 (luer slip). Nominal capacity 2ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs UNIT: Box (Box of 100) |  | SUPRA HEALTHCARE JOHANNESBURG PTY LTD | 0.44 | Avacare | BS2ML | HENAN AILE INDUSTRIAL CO.,LTD | CHINA | Initial 60 days, thereafter 21 | 84 | USD | 10 Boxes | YES |  | COMPLY pending confirmation of latex and DEHP free | NOT COMPLY
42142608-00033 | 97 | Syringe, three-part, plastic, latex free, non-pyrogenic and DEPH free. Type 2 (luer slip). Nominal capacity 2ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs UNIT: Box (Box of 100) | 1 | Vital Life Medical Distributors (PTY) LTD | 0.44 | Healthease | NHPS01-5ML | Neomedic | China | 14 days | 80% | US Dollar | 50 boxes | Yes | 100's per box | No sample submitted for evaluation | NO TEST REPORTS
42142608-00033 | 97 | Syringe, three-part, plastic, latex free, non-pyrogenic and DEPH free. Type 2 (luer slip). Nominal capacity 2ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs UNIT: Box (Box of 100) | N/A | SAFARMEX (PTY) LTD | 0.48 | Saclin | SSYN-10024 | Zhejiang Mediunion Healthcare Group Co. Ltd | China | 3days | 80% | US$ | 3000 Boxes | Yes | Box of 100 | COMPLY pending confirmation of latex and DEHP free | COMPLY
42142608-00033 | 97 | Syringe, three-part, plastic, latex free, non-pyrogenic and DEPH free. Type 2 (luer slip). Nominal capacity 2ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs UNIT: Box (Box of 100) | Basic Line Offer | BIOCOR HOSPITAL SUPPLIES PTY LTD | 0.5 | Triple -D | SYR-Care 2ml | Jumla Medical | South Africa | 21 | 0 | N/A | 5 Boxes of 100 | Yes | N/A | COMPLY pending confirmation of latex and DEHP free | NO TEST REPORTS
42142608-00033 | 97 | Syringe, three-part, plastic, latex free, non-pyrogenic and DEPH free. Type 2 (luer slip). Nominal capacity 2ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs UNIT: Box (Box of 100) |  | GIBELA TRADE AND INVEST 1001 CC | 0.5 | TARA | SYR015 | Shandong Qiaopai Group Co., Ltd | CHINA | 21 | 85 | USD | 10 | YES | box 100 | offered a nominal capacity of 2.5ml instead 2ml | NOT COMPLY
42142608-00033 | 97 | Syringe, three-part, plastic, latex free, non-pyrogenic and DEPH free. Type 2 (luer slip). Nominal capacity 2ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs UNIT: Box (Box of 100) | 1 | B. BRAUN MEDICAL (PTY) LTD | 0.52 | Omnifix 2ml | 4606022V | B. Braun | Germany | 42 | 50% | Euro | 100 | Yes | Box of 100 | COMPLY pending confirmation of latex and DEHP free | COMPLY
42142608-00033 | 97 | Syringe, three-part, plastic, latex free, non-pyrogenic and DEPH free. Type 2 (luer slip). Nominal capacity 2ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs UNIT: Box (Box of 100) |  | JUMLA MEDICAL SUPPLIES (PTY) LTD | 37.54 | SYR-CARE | SYR-CARE-2ML | SUZHOU ACMED | CHINA | 21 DAYS | 72% | US DOLLAR | 10 BOXES | YES | NIL | COMPLY pending confirmation of latex and DEHP free | NO TEST REPORTS
42142608-00034 | 98 | Syringe, three-part, plastic, latex free, non-pyrogenic and DEPH free. Type 3 (luer lock). Nominal capacity 2ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 100 UNIT: Box |  | Medi-Core Technologies | 0.34 | ULIFE | 42142608-00034 | JIAXING ULIFE MEDICAL DEVICE TECHNOLOGY CO.,LTD | CHINA | 21 | 80 | USD | 1 box | YES | BOX | No sample submitted for evaluation | NO TEST REPORTS
42142608-00034 | 98 | Syringe, three-part, plastic, latex free, non-pyrogenic and DEPH free. Type 3 (luer lock). Nominal capacity 2ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 100 UNIT: Box | 1 | CENTRAL MEDICAL PTY LTD | 0.35 | LYSA | HSCL2000 | Zhejian INI Medical Devices Co Ltd | China | 21 | 80 | China yuan | 24 boxes of 100 units | yes | box of 100 units | Labeling of the outer/immediate package indicates Nominal capacity 2/3ml instead of 2ml | NO TEST REPORTS
42142608-00034 | 98 | Syringe, three-part, plastic, latex free, non-pyrogenic and DEPH free. Type 3 (luer lock). Nominal capacity 2ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 100 UNIT: Box | N/A | PROMED TECHNOLOGIES (PTY) LTD | 0.35 | U-LIFE | 608034 | JIAXING | CHINA | 14 DAYS | 85% | USD | 1 BOX | YES | BOX | No samples submitted for evaluation | NOT COMPLY
42142608-00034 | 98 | Syringe, three-part, plastic, latex free, non-pyrogenic and DEPH free. Type 3 (luer lock). Nominal capacity 2ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 100 UNIT: Box |  | Mr First Aid (Pty) Ltd | 0.4 | Yue Kang |  | Yue Kang | China | 42 | 85 | USD | 10 | Yes |  | COMPLY subject to SABS conformation of packaging | NO TEST REPORTS
42142608-00034 | 98 | Syringe, three-part, plastic, latex free, non-pyrogenic and DEPH free. Type 3 (luer lock). Nominal capacity 2ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 100 UNIT: Box |  | Isigidi Trading 95 cc | 46 | Isigidi | SY59 | Jiangsu Jichun Medical Devices Co. Ltd. | China | 14 days | 85% | USD | 30boxes | Yes | 100pcs/box | Nominal capacity 2.5ml The outer package indicates 2ml and immediate package indicates 2.5ml | COMPLY | Recommended | 91
42142608-00034 | 98 | Syringe, three-part, plastic, latex free, non-pyrogenic and DEPH free. Type 3 (luer lock). Nominal capacity 2ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 100 UNIT: Box | N/A | SAFARMEX (PTY) LTD | 0.48 | Saclin | To be confirmed | Zhejiang Mediunion Healthcare Group Co. Ltd | China | 3days | 80% | US$ | 3000 Boxes | Yes | Box of 100 | No sample submitted for evaluation | NO TEST REPORT
42142608-00034 | 98 | Syringe, three-part, plastic, latex free, non-pyrogenic and DEPH free. Type 3 (luer lock). Nominal capacity 2ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 100 UNIT: Box | Basic Line Offer | BIOCOR HOSPITAL SUPPLIES PTY LTD | 0.5 | Triple -D | SYR-Care 2ml | Jumla Medical | South Africa | 21 | 0 | N/A | 5 Boxes of 100 | Yes | N/A | COMPLY pending confirmation of latex and DEHP free | NO TEST REPORTS
42142608-00034 | 98 | Syringe, three-part, plastic, latex free, non-pyrogenic and DEPH free. Type 3 (luer lock). Nominal capacity 2ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 100 UNIT: Box |  | GIBELA TRADE AND INVEST 1001 CC | 0.5 | TARA | SYR016 | Shandong Qiaopai Group Co., Ltd | CHINA | 21 | 85 | USD | 10 | YES | box 100 | No sample submitted for evaluation | NO TEST REPORTS
42142608-00034 | 98 | Syringe, three-part, plastic, latex free, non-pyrogenic and DEPH free. Type 3 (luer lock). Nominal capacity 2ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 100 UNIT: Box | 1 | B. BRAUN MEDICAL (PTY) LTD | 0.54 | Omnifix 2ml Luer Lock | 4617029V | B. Braun | Germany | 42 | 50% | Euro | 100 | Yes | Box of 100 | COMPLY pending confirmation of latex free | COMPLY
42142608-00034 | 98 | Syringe, three-part, plastic, latex free, non-pyrogenic and DEPH free. Type 3 (luer lock). Nominal capacity 2ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 100 UNIT: Box |  | SUPRA HEALTHCARE JOHANNESBURG PTY LTD | 0.61 | Avacare | BS2ML/L2 | HENAN AILE INDUSTRIAL CO.,LTD | CHINA | Initial 60 days, thereafter 21 | 84 | USD | 10 Boxes | YES |  | Offered nominal capacity of 3ml instead of 2ml | NOT COMPLY
42142608-00034 | 98 | Syringe, three-part, plastic, latex free, non-pyrogenic and DEPH free. Type 3 (luer lock). Nominal capacity 2ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 100 UNIT: Box |  | ALMIKA TRADING PTY LTD | 0.86 | BD | 309658 | BD | USA | 14 | 0 | NO | 100 | YES | 1 |  | COMPLY
42142608-00034 | 98 | Syringe, three-part, plastic, latex free, non-pyrogenic and DEPH free. Type 3 (luer lock). Nominal capacity 2ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 100 UNIT: Box | Basic Line offer | Becton Dickinson | 1.5 | BD Plastipak™ Syringe 3ml Luer-Lok | 309658 | Becton Dickinson | Fraga, Spain | 7 Working Days | 76% | USD | 1 box | YES | 200 units in a box | Offered a 3mil instead of 2mil | COMPLY
42142608-00034 | 98 | Syringe, three-part, plastic, latex free, non-pyrogenic and DEPH free. Type 3 (luer lock). Nominal capacity 2ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 100 UNIT: Box |  | JUMLA MEDICAL SUPPLIES (PTY) LTD | 37.54 | SYR-CARE | SYR-CARE-2MLLL | SUZHOU ACMED | CHINA | 21 DAYS | 72% | US DOLLAR | 10 BOXES | YES | NIL | COMPLY pending confirmation of latex and DEHP free | NOT COMPLY
42142608-00035 | 99 | Syringe, three-part, plastic, latex free, non-pyrogenic and DEPH free. Type 2 (luer slip), with a concentric nozzle. Nominal capacity 5ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 100 UNIT: Box |  | Chemical World cc | 0.32 | SURG PLUS | SYR005 | TONGDA | china | 3 | 85 | USD | 36000 | YES | 100 | No sample for evaluation | NOT COMPLY
42142608-00035 | 99 | Syringe, three-part, plastic, latex free, non-pyrogenic and DEPH free. Type 2 (luer slip), with a concentric nozzle. Nominal capacity 5ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 100 UNIT: Box |  | Isigidi Trading 95 cc | 0.35 | Isigidi | SY62 | Jiangsu Jichun Medical Devices Co. Ltd. | China | 14 days | 85% | USD | 20boxes | Yes | 100pcs/box | COMPLY pending confirmation of latex and DEHP free | COMPLY
42142608-00035 | 99 | Syringe, three-part, plastic, latex free, non-pyrogenic and DEPH free. Type 2 (luer slip), with a concentric nozzle. Nominal capacity 5ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 100 UNIT: Box | 1 | CENTRAL MEDICAL PTY LTD | 35.5 | LYSA | HSCS5000 | Zhejian INI Medical Devices Co Ltd | China | 21 | 80 | China yuan | 24 boxes of 100 units | yes | box of 100 units | COMPLY pending confirmation of latex and DEHP free | COMPLY | Recommended | 99.1
42142608-00035 | 99 | Syringe, three-part, plastic, latex free, non-pyrogenic and DEPH free. Type 2 (luer slip), with a concentric nozzle. Nominal capacity 5ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 100 UNIT: Box | N/A | PROMED TECHNOLOGIES (PTY) LTD | 0.36 | U-LIFE | 608035 | JIAXING | CHINA | 14 DAYS | 85% | USD | 1 BOX | YES | BOX | COMPLY pending confirmation of latex and DEHP free | COMPLY
42142608-00035 | 99 | Syringe, three-part, plastic, latex free, non-pyrogenic and DEPH free. Type 2 (luer slip), with a concentric nozzle. Nominal capacity 5ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 100 UNIT: Box |  | Medi-Core Technologies | 0.38 | ULIFE | 42142608-00035 | JIAXING ULIFE MEDICAL DEVICE TECHNOLOGY CO.,LTD | CHINA | 21 | 80 | USD | 1 box | YES | BOX | COMPLY pending confirmation of latex free | COMPLY
42142608-00035 | 99 | Syringe, three-part, plastic, latex free, non-pyrogenic and DEPH free. Type 2 (luer slip), with a concentric nozzle. Nominal capacity 5ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 100 UNIT: Box |  | Mr First Aid (Pty) Ltd | 0.4 | Yue Kang |  | Yue Kang | China | 42 | 85 | USD | 10 | Yes |  | Packages of all samples were open which compromised sterility | NO TEST REPORTS
42142608-00035 | 99 | Syringe, three-part, plastic, latex free, non-pyrogenic and DEPH free. Type 2 (luer slip), with a concentric nozzle. Nominal capacity 5ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 100 UNIT: Box | 1 | Vital Life Medical Distributors (PTY) LTD | 0.44 | Healthease | NHPS01-5ML | Neomedic | China | 14 days | 80% | US Dollar | 50 boxes | Yes | 100's per box | offered a luer lock instead od luer slip | NO TEST REPORTS
42142608-00035 | 99 | Syringe, three-part, plastic, latex free, non-pyrogenic and DEPH free. Type 2 (luer slip), with a concentric nozzle. Nominal capacity 5ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 100 UNIT: Box | Basic Line offer | Becton Dickinson | 0.47 | BD Emerald Syringe, 5ml | 307731 | Becton Dickinson | Fraga, Spain | 7 Working Days | 76% | USD | 1 box | YES | 100 units in a box | COMPLY pending confirmation of latex and DEHP free | COMPLY
42142608-00035 | 99 | Syringe, three-part, plastic, latex free, non-pyrogenic and DEPH free. Type 2 (luer slip), with a concentric nozzle. Nominal capacity 5ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 100 UNIT: Box |  | FOCUS PRODUCTS (PTY) LTD | 0.49 | HI-TECH |  | SR MEDICAL CARE | INDIA | 45 | 90 | USD | 100 BOXES | YES | NONE | Offered a tear off package instead of an individually peel apart package | NOT COMPLY
42142608-00035 | 99 | Syringe, three-part, plastic, latex free, non-pyrogenic and DEPH free. Type 2 (luer slip), with a concentric nozzle. Nominal capacity 5ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 100 UNIT: Box |  | GIBELA TRADE AND INVEST 1001 CC | 0.52 | TARA | SYR017 | Shandong Qiaopai Group Co., Ltd | CHINA | 21 | 85 | USD | 10 | YES | box 100 | COMPLY pending confirmation of latex and DEHP free | NOT COMPLY
42142608-00035 | 99 | Syringe, three-part, plastic, latex free, non-pyrogenic and DEPH free. Type 2 (luer slip), with a concentric nozzle. Nominal capacity 5ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 100 UNIT: Box |  | SUPRA HEALTHCARE JOHANNESBURG PTY LTD | 0.52 | Avacare | BS5ML | HENAN AILE INDUSTRIAL CO.,LTD | CHINA | Initial 60 days, thereafter 21 | 84 | USD | 10 Boxes | YES |  | COMPLY pending confirmation of latex and DEHP free | NOT COMPLY
42142608-00035 | 99 | Syringe, three-part, plastic, latex free, non-pyrogenic and DEPH free. Type 2 (luer slip), with a concentric nozzle. Nominal capacity 5ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 100 UNIT: Box | Basic Line Offer | BIOCOR HOSPITAL SUPPLIES PTY LTD | 0.54 | Triple -D | SYR-Care 5ml | Jumla Medical | South Africa | 21 | 0 | N/A | 5 Boxes of 100 | Yes | N/A | COMPLY pending confirmation of latex and DEHP free | NO TEST REPORTS
42142608-00035 | 99 | Syringe, three-part, plastic, latex free, non-pyrogenic and DEPH free. Type 2 (luer slip), with a concentric nozzle. Nominal capacity 5ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 100 UNIT: Box | N/A | SAFARMEX (PTY) LTD | 0.55 | Saclin | SSYN-10025 | Zhejiang Mediunion Healthcare Group Co. Ltd | China | 3days | 80% | US$ | 3000 Boxes | Yes | Box of 100 | Offered Nominal capacity of 3ml instead 5ml | NOT COMPLY
42142608-00035 | 99 | Syringe, three-part, plastic, latex free, non-pyrogenic and DEPH free. Type 2 (luer slip), with a concentric nozzle. Nominal capacity 5ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 100 UNIT: Box |  | ALMIKA TRADING PTY LTD | 0.56 | BD | 307731 | BD | USA | 14 | 0 | NO | 100 | YES | 1 |  | COMPLY
42142608-00035 | 99 | Syringe, three-part, plastic, latex free, non-pyrogenic and DEPH free. Type 2 (luer slip), with a concentric nozzle. Nominal capacity 5ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 100 UNIT: Box |  | JUMLA MEDICAL SUPPLIES (PTY) LTD | 40.51 | SYR-CARE | SYR-CARE-5ML | SUZHOU ACMED | CHINA | 21 DAYS | 72% | US DOLLAR | 10 BOXES | YES | NIL | COMPLY pending confirmation of latex and DEHP free | NOT COMPLY
42142608-00036 | 100 | Syringe, three-part, plastic, latex free, non-pyrogenic and DEPH free. Type 2 (luer slip), with an eccentric nozzle. Nominal capacity 5ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed, Sterile, Single use, Individually peel packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 100 UNIT: Box |  | Isigidi Trading 95 cc | 34.96 | Isigidi | SY63 | Jiangsu Jichun Medical Devices Co. Ltd. | China | 14 days | 85% | USD | 20boxes | Yes | 100pcs/box | COMPLY pending confirmation of latex and DEHP free | COMPLY | Recommended | 91
42142608-00036 | 100 | Syringe, three-part, plastic, latex free, non-pyrogenic and DEPH free. Type 2 (luer slip), with an eccentric nozzle. Nominal capacity 5ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed, Sterile, Single use, Individually peel packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 100 UNIT: Box | N/A | PROMED TECHNOLOGIES (PTY) LTD | 0.36 | U-LIFE | 608036 | JIAXING | CHINA | 14 DAYS | 85% | USD | 1 BOX | YES | BOX | No samples submitted for evaluation | NO TEST REPORTS
42142608-00036 | 100 | Syringe, three-part, plastic, latex free, non-pyrogenic and DEPH free. Type 2 (luer slip), with an eccentric nozzle. Nominal capacity 5ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed, Sterile, Single use, Individually peel packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 100 UNIT: Box |  | Medi-Core Technologies | 0.38 | ULIFE | 42142608-00036 | JIAXING ULIFE MEDICAL DEVICE TECHNOLOGY CO.,LTD | CHINA | 21 | 80 | USD | 1 box | YES | BOX | No sample submitted for evaluation | NO TEST REPORTS
42142608-00036 | 100 | Syringe, three-part, plastic, latex free, non-pyrogenic and DEPH free. Type 2 (luer slip), with an eccentric nozzle. Nominal capacity 5ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed, Sterile, Single use, Individually peel packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 100 UNIT: Box |  | Mr First Aid (Pty) Ltd | 0.4 | Yue Kang |  | Yue Kang | China | 42 | 85 | USD | 10 | Yes |  | Packages of all samples were open which compromised sterility | NO TEST REPORTS
42142608-00036 | 100 | Syringe, three-part, plastic, latex free, non-pyrogenic and DEPH free. Type 2 (luer slip), with an eccentric nozzle. Nominal capacity 5ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed, Sterile, Single use, Individually peel packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 100 UNIT: Box | N/A | SAFARMEX (PTY) LTD | 0.55 | Saclin | SSYN-10026 | Zhejiang Mediunion Healthcare Group Co. Ltd | China | 3days | 80% | US$ | 3000 Boxes | Yes | Box of 100 | COMPLY pending confirmation of latex and DEHP free | COMPLY
42142608-00036 | 100 | Syringe, three-part, plastic, latex free, non-pyrogenic and DEPH free. Type 2 (luer slip), with an eccentric nozzle. Nominal capacity 5ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed, Sterile, Single use, Individually peel packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 100 UNIT: Box | 1 | B. BRAUN MEDICAL (PTY) LTD | 0.65 | Omnifix 5ml | 4616057V | B. Braun | Germany | 42 | 50% | Euro | 100 | Yes | Box of 100 | COMPLY pending confirmation of latex and DEHP free | COMPLY
42142608-00036 | 100 | Syringe, three-part, plastic, latex free, non-pyrogenic and DEPH free. Type 2 (luer slip), with an eccentric nozzle. Nominal capacity 5ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed, Sterile, Single use, Individually peel packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 100 UNIT: Box |  | GIBELA TRADE AND INVEST 1001 CC | 0.66 | TARA | SYR018 | Shandong Qiaopai Group Co., Ltd | CHINA | 21 | 85 | USD | 10 | YES | box 100 | No sample submitted for evaluation | NO TEST REPORTS
42142608-00036 | 100 | Syringe, three-part, plastic, latex free, non-pyrogenic and DEPH free. Type 2 (luer slip), with an eccentric nozzle. Nominal capacity 5ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed, Sterile, Single use, Individually peel packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 100 UNIT: Box |  | CRYSATL PIER TRADING 148cc | 0.72 | SRM-HYP0010 |  | GLOBAL MEDIKIT LIMITED C/O SR MEDICAL CARE | [6 empty] | BOX | No sample submitted for evaluation | NO TEST REPORTS
42142608-00037 | 101 | Syringe, three-part, plastic, latex free, non-pyrogenic and DEPH free. Type 3 (luer lock), with a concentric nozzle. Nominal capacity 5ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 100 UNIT: Box | 1 | CENTRAL MEDICAL PTY LTD | 0.36 | LYSA | HSCL5000 | Zhejian INI Medical Devices Co Ltd | China | 21 | 80 | China yuan | 24 boxes of 100 units | yes | box of 100 units | COMPLY pending confirmation of latex and DEHP free | NO TEST REPORTS
42142608-00037 | 101 | Syringe, three-part, plastic, latex free, non-pyrogenic and DEPH free. Type 3 (luer lock), with a concentric nozzle. Nominal capacity 5ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 100 UNIT: Box | N/A | PROMED TECHNOLOGIES (PTY) LTD | 0.36 | U-LIFE | 608037 | JIAXING | CHINA | 14 DAYS | 85% | USD | 1 BOX | YES | BOX | COMPLY pending confirmation of latex and DEHP free | NO TEST REPORTS
42142608-00037 | 101 | Syringe, three-part, plastic, latex free, non-pyrogenic and DEPH free. Type 3 (luer lock), with a concentric nozzle. Nominal capacity 5ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 100 UNIT: Box |  | Medi-Core Technologies | 0.38 | ULIFE | 42142608-00037 | JIAXING ULIFE MEDICAL DEVICE TECHNOLOGY CO.,LTD | CHINA | 21 | 80 | USD | 1 box | YES | BOX | No sample submitted for evaluation | NO TEST REPORTS
42142608-00037 | 101 | Syringe, three-part, plastic, latex free, non-pyrogenic and DEPH free. Type 3 (luer lock), with a concentric nozzle. Nominal capacity 5ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 100 UNIT: Box |  | Mr First Aid (Pty) Ltd | 0.4 | Yue Kang |  | Yue Kang | China | 42 | 85 | USD | 10 | Yes |  | COMPLY subject to SABS conformation of packaging | NO TEST REPORTS
42142608-00037 | 101 | Syringe, three-part, plastic, latex free, non-pyrogenic and DEPH free. Type 3 (luer lock), with a concentric nozzle. Nominal capacity 5ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 100 UNIT: Box | 1 | Vital Life Medical Distributors (PTY) LTD | 0.44 | Healthease | NHPSLL01-5ML | Neomedic | China | 14 days | 80% | US Dollar | 50 boxes | Yes | 100's per box | No sample submitted for evaluation | NO TEST REPORTS
42142608-00037 | 101 | Syringe, three-part, plastic, latex free, non-pyrogenic and DEPH free. Type 3 (luer lock), with a concentric nozzle. Nominal capacity 5ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 100 UNIT: Box | Basic Line Offer | BIOCOR HOSPITAL SUPPLIES PTY LTD | 0.54 | Triple -D | SYR-Care 5MILL | Jumla Medical | South Africa | 21 | 0 | N/A | 5 Boxes of 100 | Yes | N/A | COMPLY pending confirmation of latex and DEHP free | NO TEST REPORTS
42142608-00037 | 101 | Syringe, three-part, plastic, latex free, non-pyrogenic and DEPH free. Type 3 (luer lock), with a concentric nozzle. Nominal capacity 5ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 100 UNIT: Box | N/A | SAFARMEX (PTY) LTD | 0.55 | Saclin | To be confirmed | Zhejiang Mediunion Healthcare Group Co. Ltd | China | 3days | 80% | US$ | 3000 Boxes | Yes | Box of 100 | No sample submitted for evaluation | NO TEST REPORT
42142608-00037 | 101 | Syringe, three-part, plastic, latex free, non-pyrogenic and DEPH free. Type 3 (luer lock), with a concentric nozzle. Nominal capacity 5ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 100 UNIT: Box |  | Isigidi Trading 95 cc | 57.5 | Isigidi | SY64 | Jiangsu Jichun Medical Devices Co. Ltd. | China | 14 days | 85% | USD | 20boxes | Yes | 100pcs/box | COMPLY pending confirmation of latex and DEHP free | COMPLY | Recommended | 91
42142608-00037 | 101 | Syringe, three-part, plastic, latex free, non-pyrogenic and DEPH free. Type 3 (luer lock), with a concentric nozzle. Nominal capacity 5ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 100 UNIT: Box |  | SUPRA HEALTHCARE JOHANNESBURG PTY LTD | 0.62 | Avacare | BS5L/L | HENAN AILE INDUSTRIAL CO.,LTD | CHINA | Initial 60 days, thereafter 21 | 84 | USD | 10 Boxes | YES |  | COMPLY pending confirmation of latex and DEHP free | NOT COMPLY
42142608-00037 | 101 | Syringe, three-part, plastic, latex free, non-pyrogenic and DEPH free. Type 3 (luer lock), with a concentric nozzle. Nominal capacity 5ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 100 UNIT: Box |  | GIBELA TRADE AND INVEST 1001 CC | 0.66 | TARA | SYR019 | Shandong Qiaopai Group Co., Ltd | CHINA | 21 | 85 | USD | 10 | YES | box 100 | No sample submitted for evaluation | NO TEST REPORTS
42142608-00037 | 101 | Syringe, three-part, plastic, latex free, non-pyrogenic and DEPH free. Type 3 (luer lock), with a concentric nozzle. Nominal capacity 5ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 100 UNIT: Box |  | CRYSATL PIER TRADING 148cc | 0.74 | SRM-HYP0011 |  | GLOBAL MEDIKIT LIMITED C/O SR MEDICAL CARE | [6 empty] | BOX | No sample submitted for evaluation | NO TEST REPORTS
42142608-00037 | 101 | Syringe, three-part, plastic, latex free, non-pyrogenic and DEPH free. Type 3 (luer lock), with a concentric nozzle. Nominal capacity 5ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 100 UNIT: Box |  | ALMIKA TRADING PTY LTD | 1.26 | BD | 309649 | BD | USA | 14 | 0 | NO | 100 | YES | 1 |  | NO TEST REPORTS
42142608-00037 | 101 | Syringe, three-part, plastic, latex free, non-pyrogenic and DEPH free. Type 3 (luer lock), with a concentric nozzle. Nominal capacity 5ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 100 UNIT: Box | Basic Line offer | Becton Dickinson | 1.32 | BD Plastipak™ Syringe 5ml Luer Lok | 309649 | Becton Dickinson | Fraga, Spain | 7 Working Days | 76% | USD | 1 box | no | 125 units in a box | COMPLY pending confirmation of latex and DEHP free | NO TEST REPORTS
42142608-00037 | 101 | Syringe, three-part, plastic, latex free, non-pyrogenic and DEPH free. Type 3 (luer lock), with a concentric nozzle. Nominal capacity 5ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 100 UNIT: Box |  | JUMLA MEDICAL SUPPLIES (PTY) LTD | 40.51 | SYR-CARE | SYR-CARE-5MLLL | SUZHOU ACMED | CHINA | 21 DAYS | 72% | US DOLLAR | 10 BOXES | YES | NIL | COMPLY pending confirmation of latex and DEHP free | NOT COMPLY
42142608-00038 | 102 | Syringe, three-part, plastic, latex free, non-pyrogenic and DEPH free. Type 3 (luer lock), with an eccentric nozzle. Nominal capacity 5ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 100 UNIT: Box | N/A | PROMED TECHNOLOGIES (PTY) LTD | 0.36 | U-LIFE | 608038 | JIAXING | CHINA | 14 DAYS | 85% | USD | 1 BOX | YES | BOX | No samples submitted for evaluation | NO TEST REPORTS | Re-tender
42142608-00038 | 102 | Syringe, three-part, plastic, latex free, non-pyrogenic and DEPH free. Type 3 (luer lock), with an eccentric nozzle. Nominal capacity 5ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 100 UNIT: Box |  | Medi-Core Technologies | 0.38 | ULIFE | 42142608-00038 | JIAXING ULIFE MEDICAL DEVICE TECHNOLOGY CO.,LTD | CHINA | 21 | 80 | USD | 1 box | YES | BOX | No sample submitted for evaluation | NO TEST REPORTS | Re-tender
42142608-00038 | 102 | Syringe, three-part, plastic, latex free, non-pyrogenic and DEPH free. Type 3 (luer lock), with an eccentric nozzle. Nominal capacity 5ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 100 UNIT: Box |  | Mr First Aid (Pty) Ltd | 0.4 | Yue Kang |  | Yue Kang | China | 42 | 85 | USD | 10 | Yes |  | COMPLY subject to SABS conformation of packaging | NO TEST REPORTS | Re-tender
42142608-00038 | 102 | Syringe, three-part, plastic, latex free, non-pyrogenic and DEPH free. Type 3 (luer lock), with an eccentric nozzle. Nominal capacity 5ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 100 UNIT: Box |  | CRYSATL PIER TRADING 148cc | 0.74 | SRM-HYP0012 |  | GLOBAL MEDIKIT LIMITED C/O SR MEDICAL CARE | [6 empty] | BOX | No sample submitted for evaluation | NO TEST REPORTS | Re-tender
42142608-00039 | 103 | Syringe, three-part, plastic, latex free, non-pyrogenic and DEPH free. Type 2 (luer slip), with a concentric nozzle. Nominal capacity 10ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 100 UNIT: Box |  | Chemical World cc | 0.53 | SURG PLUS | SYR010 | TONGDA | china | 3 | 85 | USD | 32000 | YES | 100 | COMPLY pending confirmation of latex free | NOT COMPLY
42142608-00039 | 103 | Syringe, three-part, plastic, latex free, non-pyrogenic and DEPH free. Type 2 (luer slip), with a concentric nozzle. Nominal capacity 10ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 100 UNIT: Box |  | Medi-Core Technologies | 0.53 | ULIFE | 42142608-00039 | JIAXING ULIFE MEDICAL DEVICE TECHNOLOGY CO.,LTD | CHINA | 21 | 80 | USD | 1 box | YES | BOX | No sample submitted for evaluation | NO TEST REPORTS
42142608-00039 | 103 | Syringe, three-part, plastic, latex free, non-pyrogenic and DEPH free. Type 2 (luer slip), with a concentric nozzle. Nominal capacity 10ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 100 UNIT: Box | N/A | PROMED TECHNOLOGIES (PTY) LTD | 57 | U-LIFE | 608039 | JIAXING | CHINA | 14 DAYS | 85% | USD | 1 BOX | YES | BOX OF 100 | COMPLY pending confirmation of latex and DEHP free | COMPLY | Recommended | 100
42142608-00039 | 103 | Syringe, three-part, plastic, latex free, non-pyrogenic and DEPH free. Type 2 (luer slip), with a concentric nozzle. Nominal capacity 10ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 100 UNIT: Box | Basic Line offer | Becton Dickinson | 0.62 | BD Emerald Syringe, 10ml | 307736 | Becton Dickinson | Fraga, Spain | 7 Working Days | 76% | USD | 1 box | YES | 100 units in a box | COMPLY pending confirmation of latex and DEHP free | COMPLY
42142608-00039 | 103 | Syringe, three-part, plastic, latex free, non-pyrogenic and DEPH free. Type 2 (luer slip), with a concentric nozzle. Nominal capacity 10ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 100 UNIT: Box |  | Mr First Aid (Pty) Ltd | 0.63 | Yue Kang |  | Yue Kang | China | 42 | 85 | USD | 10 | Yes |  | COMPLY subject to SABS conformation of packaging | NO TEST REPORTS
42142608-00039 | 103 | Syringe, three-part, plastic, latex free, non-pyrogenic and DEPH free. Type 2 (luer slip), with a concentric nozzle. Nominal capacity 10ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 100 UNIT: Box |  | Isigidi Trading 95 cc | 0.64 | Isigidi | SY46 | Jiangsu Jichun Medical Devices Co. Ltd. | China | 14 days | 85% | USD | 10boxes | Yes | 100pcs/box | COMPLY pending confirmation of latex and DEHP free | COMPLY
42142608-00039 | 103 | Syringe, three-part, plastic, latex free, non-pyrogenic and DEPH free. Type 2 (luer slip), with a concentric nozzle. Nominal capacity 10ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 100 UNIT: Box | 1 | CENTRAL MEDICAL PTY LTD | 0.66 | LYSA | HSCS10000 | Zhejian INI Medical Devices Co Ltd | China | 21 | 80 | China yuan | 12 boxes of 100 units | yes | box of 100 units | COMPLY pending confirmation of latex and DEHP free | NO TEST REPORTS
42142608-00039 | 103 | Syringe, three-part, plastic, latex free, non-pyrogenic and DEPH free. Type 2 (luer slip), with a concentric nozzle. Nominal capacity 10ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 100 UNIT: Box |  | ALMIKA TRADING PTY LTD | 0.75 | BD | 307736 | BD | USA | 14 | 0 | NO | 100 | YES | 1 |  | COMPLY
42142608-00039 | 103 | Syringe, three-part, plastic, latex free, non-pyrogenic and DEPH free. Type 2 (luer slip), with a concentric nozzle. Nominal capacity 10ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 100 UNIT: Box |  | GIBELA TRADE AND INVEST 1001 CC | 0.77 | TARA | SYR020 | Shandong Qiaopai Group Co., Ltd | CHINA | 21 | 85 | USD | 10 | YES | box 100 | COMPLY pending confirmation of latex and DEHP free | NOT COMPLY
42142608-00039 | 103 | Syringe, three-part, plastic, latex free, non-pyrogenic and DEPH free. Type 2 (luer slip), with a concentric nozzle. Nominal capacity 10ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 100 UNIT: Box |  | FOCUS PRODUCTS (PTY) LTD | 0.82 | HI-TECH |  | SR MEDICAL CARE | INDIA | 45 | 90 | USD | 100 BOXES | YES | NONE | Offered a tear off package instead of an individually peel apart package | NOT COMPLY
42142608-00039 | 103 | Syringe, three-part, plastic, latex free, non-pyrogenic and DEPH free. Type 2 (luer slip), with a concentric nozzle. Nominal capacity 10ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 100 UNIT: Box | Basic Line Offer | BIOCOR HOSPITAL SUPPLIES PTY LTD | 0.84 | Triple -D | SYR-Care10ML | Jumla Medical | South Africa | 21 | 0 | N/A | 5 Boxes of 100 | Yes | N/A | COMPLY pending confirmation of latex and DEHP free | NO TEST REPORTS
42142608-00039 | 103 | Syringe, three-part, plastic, latex free, non-pyrogenic and DEPH free. Type 2 (luer slip), with a concentric nozzle. Nominal capacity 10ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 100 UNIT: Box | N/A | SAFARMEX (PTY) LTD | 0.89 | Saclin | To be confirmed | Zhejiang Mediunion Healthcare Group Co. Ltd | China | 3days | 80% | US$ | 3000 Boxes | Yes | Box of 100 | No sample submitted for evaluation | NO TEST REPORT
42142608-00039 | 103 | Syringe, three-part, plastic, latex free, non-pyrogenic and DEPH free. Type 2 (luer slip), with a concentric nozzle. Nominal capacity 10ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 100 UNIT: Box |  | SUPRA HEALTHCARE JOHANNESBURG PTY LTD | 0.91 | Avacare | BS10ML | HENAN AILE INDUSTRIAL CO.,LTD | CHINA | Initial 60 days, thereafter 21 | 84 | USD | 10 Boxes | YES |  | COMPLY pending confirmation of latex and DEHP free | NO TEST REPORTS
42142608-00039 | 103 | Syringe, three-part, plastic, latex free, non-pyrogenic and DEPH free. Type 2 (luer slip), with a concentric nozzle. Nominal capacity 10ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 100 UNIT: Box |  | CRYSATL PIER TRADING 148cc | 1.19 | MEDIPLUS SR CARE |  | GLOBAL MEDIKIT LIMITED C/O SR MEDICAL CARE | [6 empty] | BOX | No sample submitted for evaluation | NO TEST REPORTS
42142608-00039 | 103 | Syringe, three-part, plastic, latex free, non-pyrogenic and DEPH free. Type 2 (luer slip), with a concentric nozzle. Nominal capacity 10ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 100 UNIT: Box | 1 | Biologica Pharmaceuticals Pty Ltd | 1.43 | Sol- M | 308003 | Sol- Millennium | China | 4-6 Weeks | 100% | USD | 1000 Boxes | 100% | N/A |  | NO TEST REPORTS
42142608-00039 | 103 | Syringe, three-part, plastic, latex free, non-pyrogenic and DEPH free. Type 2 (luer slip), with a concentric nozzle. Nominal capacity 10ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 100 UNIT: Box |  | JUMLA MEDICAL SUPPLIES (PTY) LTD | 63.47 | SYR-CARE | SYR-CARE-10ML | SUZHOU ACMED | CHINA | 21 DAYS | 72% | US DOLLAR | 10 BOXES | YES | NIL | COMPLY pending confirmation of latex and DEHP free | NOT COMPLY
42142608-00040 | 104 | Syringe, three-part, plastic, latex free, non-pyrogenic and DEPH free. Type 2 (luer slip), with an eccentric nozzle. Nominal capacity 10ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 100 UNIT: Box |  | Chemical World cc | 0.53 | SURG PLUS | SYR010 | TONGDA | china | 3 | 85 | USD | 18000 | YES | EACH | COMPLY pending confirmation of latex free | NOT COMPLY
42142608-00040 | 104 | Syringe, three-part, plastic, latex free, non-pyrogenic and DEPH free. Type 2 (luer slip), with an eccentric nozzle. Nominal capacity 10ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 100 UNIT: Box |  | Medi-Core Technologies | 53 | ULIFE | 42142608-00040 | JIAXING ULIFE MEDICAL DEVICE TECHNOLOGY CO.,LTD | CHINA | 21 | 80 | USD | 1 box | YES | BOX OF 100 | COMPLY pending confirmation of latex and DEHP free | COMPLY | Recommended | 100
42142608-00040 | 104 | Syringe, three-part, plastic, latex free, non-pyrogenic and DEPH free. Type 2 (luer slip), with an eccentric nozzle. Nominal capacity 10ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 100 UNIT: Box | N/A | PROMED TECHNOLOGIES (PTY) LTD | 0.57 | U-LIFE | 608040 | JIAXING | CHINA | 14 DAYS | 85% | USD | 1 BOX | YES | BOX | COMPLY pending confirmation of latex and DEHP free | NO TEST REPORTS
42142608-00040 | 104 | Syringe, three-part, plastic, latex free, non-pyrogenic and DEPH free. Type 2 (luer slip), with an eccentric nozzle. Nominal capacity 10ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 100 UNIT: Box |  | Mr First Aid (Pty) Ltd | 0.63 | Yue Kang |  | Yue Kang | China | 42 | 85 | USD | 10 | Yes |  | Packages of all samples were open which compromised sterility | NO TEST REPORTS
42142608-00040 | 104 | Syringe, three-part, plastic, latex free, non-pyrogenic and DEPH free. Type 2 (luer slip), with an eccentric nozzle. Nominal capacity 10ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 100 UNIT: Box |  | Isigidi Trading 95 cc | 0.64 | Isigidi | SY47 | Jiangsu Jichun Medical Devices Co. Ltd. | China | 14 days | 85% | USD | 10boxes | Yes | 100pcs/box | COMPLY pending confirmation of latex and DEHP free | COMPLY
42142608-00040 | 104 | Syringe, three-part, plastic, latex free, non-pyrogenic and DEPH free. Type 2 (luer slip), with an eccentric nozzle. Nominal capacity 10ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 100 UNIT: Box | 1 | Vital Life Medical Distributors (PTY) LTD | 0.73 | Healthease | NHPS01-10ML | Neomedic | China | 14 days | 80% | US Dollar | 50 boxes | Yes | 100's per box | COMPLY pending confirmation of latex and DEHP free | NO TEST REPORTS
42142608-00040 | 104 | Syringe, three-part, plastic, latex free, non-pyrogenic and DEPH free. Type 2 (luer slip), with an eccentric nozzle. Nominal capacity 10ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 100 UNIT: Box |  | GIBELA TRADE AND INVEST 1001 CC | 0.87 | TARA | SYR021 | Shandong Qiaopai Group Co., Ltd | CHINA | 21 | 85 | USD | 10 | YES | box 100 | No sample submitted for evaluation | NO TEST REPORTS
42142608-00040 | 104 | Syringe, three-part, plastic, latex free, non-pyrogenic and DEPH free. Type 2 (luer slip), with an eccentric nozzle. Nominal capacity 10ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 100 UNIT: Box | N/A | SAFARMEX (PTY) LTD | 0.89 | Saclin | SSYN-10027 | Zhejiang Mediunion Healthcare Group Co. Ltd | China | 3days | 80% | US$ | 3000 Boxes | Yes | Box of 100 | COMPLY pending confirmation of latex and DEHP free | COMPLY
42142608-00040 | 104 | Syringe, three-part, plastic, latex free, non-pyrogenic and DEPH free. Type 2 (luer slip), with an eccentric nozzle. Nominal capacity 10ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 100 UNIT: Box | 1 | B. BRAUN MEDICAL (PTY) LTD | 0.9 | Omnifix 10ml | 4616103V | B. Braun | Germany | 42 | 50% | Euro | 100 | Yes | Box of 100 | COMPLY pending confirmation of latex and DEHP free | COMPLY
42142608-00040 | 104 | Syringe, three-part, plastic, latex free, non-pyrogenic and DEPH free. Type 2 (luer slip), with an eccentric nozzle. Nominal capacity 10ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 100 UNIT: Box |  | SUPRA HEALTHCARE JOHANNESBURG PTY LTD | 1.11 | Avacare | BS10ML/LE | HENAN AILE INDUSTRIAL CO.,LTD | CHINA | Initial 60 days, thereafter 21 | 84 | USD | 10 Boxes | YES |  | COMPLY pending confirmation of latex and DEHP free | NOT COMPLY
42142608-00040 | 104 | Syringe, three-part, plastic, latex free, non-pyrogenic and DEPH free. Type 2 (luer slip), with an eccentric nozzle. Nominal capacity 10ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 100 UNIT: Box |  | CRYSATL PIER TRADING 148cc | 1.19 | MEDIPLUS SR CARE | SRM-HYP0014 | GLOBAL MEDIKIT LIMITED C/O SR MEDICAL CARE | INDIA | 42 DAYS | 85% | USD | 100 000 PCS | YES | BOX | No sample submitted for evaluation | NO TEST REPORTS
42142608-00041 | 105 | Syringe, three-part, plastic, latex free, non-pyrogenic and DEPH free. Type 3 (luer lock), with a concentric nozzle. Nominal capacity 10ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 100 UNIT: Box | N/A | PROMED TECHNOLOGIES (PTY) LTD | 0.57 | U-LIFE | 608041 | JIAXING | CHINA | 14 DAYS | 85% | USD | 1 BOX | YES | BOX | No samples submitted for evaluation | NO TEST REPORTS
42142608-00041 | 105 | Syringe, three-part, plastic, latex free, non-pyrogenic and DEPH free. Type 3 (luer lock), with a concentric nozzle. Nominal capacity 10ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 100 UNIT: Box |  | Medi-Core Technologies | 0.59 | ULIFE | 42142608-00041 | JIAXING ULIFE MEDICAL DEVICE TECHNOLOGY CO.,LTD | CHINA | 21 | 80 | USD | 1 box | YES | BOX | No sample submitted for evaluation | NO TEST REPORTS
42142608-00041 | 105 | Syringe, three-part, plastic, latex free, non-pyrogenic and DEPH free. Type 3 (luer lock), with a concentric nozzle. Nominal capacity 10ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 100 UNIT: Box |  | Mr First Aid (Pty) Ltd | 0.63 | Yue Kang |  | Yue Kang | China | 42 | 85 | USD | 10 | Yes |  | COMPLY subject to SABS conformation of packaging | NO TEST REPORTS
42142608-00041 | 105 | Syringe, three-part, plastic, latex free, non-pyrogenic and DEPH free. Type 3 (luer lock), with a concentric nozzle. Nominal capacity 10ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 100 UNIT: Box | 1 | CENTRAL MEDICAL PTY LTD | 65.35 | LYSA | HSCL10000 | Zhejian INI Medical Devices Co Ltd | China | 21 | 80 | China yuan | 12 boxes of 100 units | yes | box of 100 units | COMPLY pending confirmation of latex and DEHP free | COMPLY | Recommended | 100
42142608-00041 | 105 | Syringe, three-part, plastic, latex free, non-pyrogenic and DEPH free. Type 3 (luer lock), with a concentric nozzle. Nominal capacity 10ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 100 UNIT: Box | 1 | Vital Life Medical Distributors (PTY) LTD | 0.73 | Healthease | NHPSLL01-10ML | Neomedic | China | 14 days | 80% | US Dollar | 50 boxes | Yes | 100's per box | No sample submitted for evaluation | NO TEST REPORTS
42142608-00041 | 105 | Syringe, three-part, plastic, latex free, non-pyrogenic and DEPH free. Type 3 (luer lock), with a concentric nozzle. Nominal capacity 10ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 100 UNIT: Box |  | GIBELA TRADE AND INVEST 1001 CC | 0.78 | TARA | SYR022 | Shandong Qiaopai Group Co., Ltd | CHINA | 21 | 85 | USD | 10 | YES | box 100 | No sample submitted for evaluation | NO TEST REPORTS
42142608-00041 | 105 | Syringe, three-part, plastic, latex free, non-pyrogenic and DEPH free. Type 3 (luer lock), with a concentric nozzle. Nominal capacity 10ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 100 UNIT: Box | Basic Line Offer | BIOCOR HOSPITAL SUPPLIES PTY LTD | 0.84 | Triple -D | SYR-Care 10MILL | Jumla Medical | South Africa | 21 | 0 | N/A | 5 Boxes of 100 | Yes | N/A | COMPLY pending confirmation of latex and DEHP free | NO TEST REPORTS
42142608-00041 | 105 | Syringe, three-part, plastic, latex free, non-pyrogenic and DEPH free. Type 3 (luer lock), with a concentric nozzle. Nominal capacity 10ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 100 UNIT: Box |  | Isigidi Trading 95 cc | 0.87 | Isigidi | SY48 | Jiangsu Jichun Medical Devices Co. Ltd. | China | 14 days | 85% | USD | 10boxes | Yes | 100pcs/box | COMPLY pending confirmation of latex and DEHP free | COMPLY
42142608-00041 | 105 | Syringe, three-part, plastic, latex free, non-pyrogenic and DEPH free. Type 3 (luer lock), with a concentric nozzle. Nominal capacity 10ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 100 UNIT: Box | 1 | B. BRAUN MEDICAL (PTY) LTD | 0.9 | Omnifix 10ml Luer Lock | 4617100V | B. Braun | Germany | 42 | 50% | Euro | 100 | Yes | Box of 100 | COMPLY pending confirmation of latex and DEHP free | COMPLY
42142608-00041 | 105 | Syringe, three-part, plastic, latex free, non-pyrogenic and DEPH free. Type 3 (luer lock), with a concentric nozzle. Nominal capacity 10ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 100 UNIT: Box |  | SUPRA HEALTHCARE JOHANNESBURG PTY LTD | 0.91 | Avacare | BS10ML/L2E | HENAN AILE INDUSTRIAL CO.,LTD | CHINA | Initial 60 days, thereafter 21 | 84 | USD | 10 Boxes | YES |  | COMPLY pending confirmation of latex and DEHP free | NOT COMPLY
42142608-00041 | 105 | Syringe, three-part, plastic, latex free, non-pyrogenic and DEPH free. Type 3 (luer lock), with a concentric nozzle. Nominal capacity 10ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 100 UNIT: Box |  | CRYSATL PIER TRADING 148cc | 1.21 | MEDIPLUS SR CARE | SRM-HYP0015 | GLOBAL MEDIKIT LIMITED C/O SR MEDICAL CARE | INDIA | 42 DAYS | 85% | USD | 100 000 PCS | YES | BOX | No sample submitted for evaluation | NO TEST REPORTS
42142608-00041 | 105 | Syringe, three-part, plastic, latex free, non-pyrogenic and DEPH free. Type 3 (luer lock), with a concentric nozzle. Nominal capacity 10ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 100 UNIT: Box | Basic Line offer | Becton Dickinson | 1.55 | BD Platipak Syringe 10ml Luer-Lok | 300912 | Becton Dickinson | Fraga, Spain | 7 Working Days | 76% | USD | 1 box | YES | 100 Units in a box | COMPLY pending confirmation of latex and DEHP free | NOT COMPLY
42142608-00041 | 105 | Syringe, three-part, plastic, latex free, non-pyrogenic and DEPH free. Type 3 (luer lock), with a concentric nozzle. Nominal capacity 10ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 100 UNIT: Box |  | ALMIKA TRADING PTY LTD | 1.85 | BD | 300912 | BD | USA | 14 | 0 | NO | 100 | YES | 1 |  | NOT COMPLY
42142608-00041 | 105 | Syringe, three-part, plastic, latex free, non-pyrogenic and DEPH free. Type 3 (luer lock), with a concentric nozzle. Nominal capacity 10ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 100 UNIT: Box |  | JUMLA MEDICAL SUPPLIES (PTY) LTD | 63.47 | SYR-CARE | SYR-CARE-10MLLL | SUZHOU ACMED | CHINA | 21 DAYS | 72% | US DOLLAR | 10 BOXES | YES | NIL | COMPLY pending confirmation of latex and DEHP free | NOT COMPLY
42142608-00042 | 106 | Syringe, three-part, plastic, latex free, non-pyrogenic and DEPH free. Type 3 (luer lock), with an eccentric nozzle. Nominal capacity 10ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off.Sterile, Sterile, Single use, Individually peel packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 100 UNIT: Box | N/A | PROMED TECHNOLOGIES (PTY) LTD | 0.57 | U-LIFE | 608042 | JIAXING | CHINA | 14 DAYS | 85% | USD | 1 BOX | YES | BOX | No samples submitted for evaluation | NO TEST REPORTS | Re-tender
42142608-00042 | 106 | Syringe, three-part, plastic, latex free, non-pyrogenic and DEPH free. Type 3 (luer lock), with an eccentric nozzle. Nominal capacity 10ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off.Sterile, Sterile, Single use, Individually peel packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 100 UNIT: Box |  | Medi-Core Technologies | 0.59 | ULIFE | 42142608-00042 | JIAXING ULIFE MEDICAL DEVICE TECHNOLOGY CO.,LTD | CHINA | 21 | 80 | USD | 1 box | YES | BOX | No sample submitted for evaluation | NO TEST REPORTS | Re-tender
42142608-00042 | 106 | Syringe, three-part, plastic, latex free, non-pyrogenic and DEPH free. Type 3 (luer lock), with an eccentric nozzle. Nominal capacity 10ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off.Sterile, Sterile, Single use, Individually peel packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 100 UNIT: Box |  | Mr First Aid (Pty) Ltd | 0.63 | Yue Kang |  | Yue Kang | China | 42 | 85 | USD | 10 | Yes |  | COMPLY subject to SABS conformation of packaging | NO TEST REPORTS | Re-tender
42142608-00042 | 106 | Syringe, three-part, plastic, latex free, non-pyrogenic and DEPH free. Type 3 (luer lock), with an eccentric nozzle. Nominal capacity 10ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off.Sterile, Sterile, Single use, Individually peel packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 100 UNIT: Box |  | GIBELA TRADE AND INVEST 1001 CC | 0.87 | TARA | SYR023 | Shandong Qiaopai Group Co., Ltd | CHINA | 21 | 85 | USD | 10 | YES | box 100 | No sample submitted for evaluation | NO TEST REPORTS | Re-tender
42142608-00042 | 106 | Syringe, three-part, plastic, latex free, non-pyrogenic and DEPH free. Type 3 (luer lock), with an eccentric nozzle. Nominal capacity 10ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off.Sterile, Sterile, Single use, Individually peel packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 100 UNIT: Box |  | CRYSATL PIER TRADING 148cc | 1.21 | MEDIPLUS SR CARE | SRM-HYP0016 | GLOBAL MEDIKIT LIMITED C/O SR MEDICAL CARE | INDIA | 42 DAYS | 85% | USD | 100 000 PCS | YES | BOX | No sample submitted for evaluation | NO TEST REPORTS | Re-tender
42142608-00042 | 106 | Syringe, three-part, plastic, latex free, non-pyrogenic and DEPH free. Type 3 (luer lock), with an eccentric nozzle. Nominal capacity 10ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off.Sterile, Sterile, Single use, Individually peel packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 100 UNIT: Box | Basic Line offer | Becton Dickinson | 2.4 | BD Plastipak™ 20ml Luer Lok | 300629 | Becton Dickinson | Fraga, Spain | 7 Working Days | 76% | USD | 1 box | YES | 120 units in a box | No sample submitted for evaluation. | NO TEST REPORTS | Re-tender
42142608-00043 | 107 | Syringe, three -part, plastic, latex free, non-pyrogenic and DEPH free. Type 2 (luer slip), with a concentric nozzle. Nominal capacity 20ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 100 UNIT: Box |  | PROMED TECHNOLOGIES (PTY) LTD | 102 | U-LIFE | 608043 | JIAXING | China | 14 days | 85% | USD | 20boxes | Yes | BOX OF 100 | COMPLY pending confirmation of latex and DEHP free | COMPLY | Recommended | 100
42142608-00043 | 107 | Syringe, three -part, plastic, latex free, non-pyrogenic and DEPH free. Type 2 (luer slip), with a concentric nozzle. Nominal capacity 20ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 100 UNIT: Box | Basic Line Offer | BIOCOR HOSPITAL SUPPLIES PTY LTD | 0.68 | SYR-Care | SYR-Care 20ML | Jumla Medical | South Africa | 21 | 0 | N/A | 5 Boxes of 100 | Yes | N/A | offered an ecentric instead concentric nozzle | NO TEST REPORTS
42142608-00043 | 107 | Syringe, three -part, plastic, latex free, non-pyrogenic and DEPH free. Type 2 (luer slip), with a concentric nozzle. Nominal capacity 20ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 100 UNIT: Box |  | Chemical World cc | 0.83 | SURG PLUS | SYRC020 | TONGDA | china | 3 | 85 | USD | 18000 | YES | EACH | No sample submitted for evaluation | NOT COMPLY
42142608-00043 | 107 | Syringe, three -part, plastic, latex free, non-pyrogenic and DEPH free. Type 2 (luer slip), with a concentric nozzle. Nominal capacity 20ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 100 UNIT: Box |  | Medi-Core Technologies | 0.99 | ULIFE | 42142608-00043 | JIAXING ULIFE MEDICAL DEVICE TECHNOLOGY CO.,LTD | CHINA | 21 | 80 | USD | 2 box | YES | PRICE PER 100-PACKED 50PCS/BOX | No sample submitted for evaluation | NO TEST REPORTS
42142608-00043 | 107 | Syringe, three -part, plastic, latex free, non-pyrogenic and DEPH free. Type 2 (luer slip), with a concentric nozzle. Nominal capacity 20ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 100 UNIT: Box | N/A | PROMED TECHNOLOGIES (PTY) LTD | 1.02 | U-LIFE | 608043 | JIAXING | CHINA | 14 DAYS | 85% | USD | 1 BOX | YES | BOX | COMPLY pending confirmation of latex and DEHP free | COMPLY
42142608-00043 | 107 | Syringe, three -part, plastic, latex free, non-pyrogenic and DEPH free. Type 2 (luer slip), with a concentric nozzle. Nominal capacity 20ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 100 UNIT: Box |  | Mr First Aid (Pty) Ltd | 1.04 | Yue Kang |  | Yue Kang | China | 42 | 85 | USD | 10 | Yes | 50pc per box | COMPLY pending confirmation of latex and DEHP free | NO TEST REPORTS
42142608-00043 | 107 | Syringe, three -part, plastic, latex free, non-pyrogenic and DEPH free. Type 2 (luer slip), with a concentric nozzle. Nominal capacity 20ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 100 UNIT: Box |  | GIBELA TRADE AND INVEST 1001 CC | 1.15 | TARA | SYR024 | Shandong Qiaopai Group Co., Ltd | CHINA | 21 | 85 | USD | 10 | YES | box 100 | No sample submitted for evaluation | NO TEST REPORTS
42142608-00043 | 107 | Syringe, three -part, plastic, latex free, non-pyrogenic and DEPH free. Type 2 (luer slip), with a concentric nozzle. Nominal capacity 20ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 100 UNIT: Box |  | SUPRA HEALTHCARE JOHANNESBURG PTY LTD | 1.35 | Avacare | BS20ML | HENAN AILE INDUSTRIAL CO.,LTD | CHINA | Initial 60 days, thereafter 21 | 84 | USD | 10 Boxes | YES | Box of 50 | No sample submitted for evaluation | NO TEST REPORTS
42142608-00043 | 107 | Syringe, three -part, plastic, latex free, non-pyrogenic and DEPH free. Type 2 (luer slip), with a concentric nozzle. Nominal capacity 20ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 100 UNIT: Box | 1 | Biologica Pharmaceuticals Pty Ltd | 2 | Sol- M | 309003 | Sol- Millennium | China | 4-6 Weeks | 100% | USD | 1000 Boxes | 100% | N/A
42142608-00043 | 107 | Syringe, three -part, plastic, latex free, non-pyrogenic and DEPH free. Type 2 (luer slip), with a concentric nozzle. Nominal capacity 20ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 100 BOX (BOX OF 50) UNIT: Box |  | JUMLA MEDICAL SUPPLIES (PTY) LTD | 51.31 | SYR-CARE | SYR-CARE-20ML | SUZHOU ACMED | CHINA | 21 DAYS | 72% | US DOLLAR | 10 BOXES | YES | NIL | COMPLY pending confirmation of latex and DEHP free | NO TEST REPORTS
42142608-00044 | 108 | Syringe, three-part, plastic, latex free, non-pyrogenic and DEPH free. Type 2 (luer slip), with an eccentric nozzle. Nominal capacity 20ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 100 UNIT: Box |  | Medi-Core Technologies | 99 | ULIFE | 42142608-00044 | JIAXING ULIFE MEDICAL DEVICE TECHNOLOGY CO.,LTD | China | 14 days | 85% | USD | 20boxes | Yes | BOX OF 100 | COMPLY pending confirmation of latex and DEHP free | COMPLY | Recommended | 100
42142608-00044 | 108 | Syringe, three-part, plastic, latex free, non-pyrogenic and DEPH free. Type 2 (luer slip), with an eccentric nozzle. Nominal capacity 20ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 100 UNIT: Box |  | Chemical World cc | 0.83 | SURG PLUS | SYR020 | TONGDA | china | 3 | 85 | USD | 18000 | YES | EACH | COMPLY pending confirmation of latex free | NOT COMPLY
42142608-00044 | 108 | Syringe, three-part, plastic, latex free, non-pyrogenic and DEPH free. Type 2 (luer slip), with an eccentric nozzle. Nominal capacity 20ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 100 UNIT: Box |  | Medi-Core Technologies | 0.99 | ULIFE | 42142608-00044 | JIAXING ULIFE MEDICAL DEVICE TECHNOLOGY CO.,LTD | CHINA | 21 | 80 | USD | 2 box | YES | PRICE PER 100-PACKED 50PCS/BOX | COMPLY pending confirmation of latex and DEHP free | COMPLY
42142608-00044 | 108 | Syringe, three-part, plastic, latex free, non-pyrogenic and DEPH free. Type 2 (luer slip), with an eccentric nozzle. Nominal capacity 20ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 100 UNIT: Box | N/A | PROMED TECHNOLOGIES (PTY) LTD | 1.02 | U-LIFE | 608044 | JIAXING | CHINA | 14 DAYS | 85% | USD | 1 BOX | YES | BOX | COMPLY pending confirmation of latex and DEHP free | NO TEST REPORTS
42142608-00044 | 108 | Syringe, three-part, plastic, latex free, non-pyrogenic and DEPH free. Type 2 (luer slip), with an eccentric nozzle. Nominal capacity 20ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 100 UNIT: Box |  | Mr First Aid (Pty) Ltd | 1.04 | Yue Kang |  | Yue Kang | China | 42 | 85 | USD | 10 | Yes | 50pc per box | COMPLY pending confirmation of latex and DEHP free | NO TEST REPORTS
42142608-00044 | 108 | Syringe, three-part, plastic, latex free, non-pyrogenic and DEPH free. Type 2 (luer slip), with an eccentric nozzle. Nominal capacity 20ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 100 UNIT: Box |  | GIBELA TRADE AND INVEST 1001 CC | 1.15 | TARA | SYR025 | Shandong Qiaopai Group Co., Ltd | CHINA | 21 | 85 | USD | 10 | YES | box 100 | COMPLY pending confirmation of latex and DEHP free | NOT COMPLY
42142608-00044 | 108 | Syringe, three-part, plastic, latex free, non-pyrogenic and DEPH free. Type 2 (luer slip), with an eccentric nozzle. Nominal capacity 20ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 100 UNIT: Box | 1 | Vital Life Medical Distributors (PTY) LTD | 1.2 | Healthease | NHPSLS-20ML | Neomedic | China | 14 days | 80% | US Dollar | 50 boxes | No | Pack size is boxes of 40's | COMPLY pending confirmation of latex and DEHP free | NO TEST REPORTS
42142608-00044 | 108 | Syringe, three-part, plastic, latex free, non-pyrogenic and DEPH free. Type 2 (luer slip), with an eccentric nozzle. Nominal capacity 20ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 100 UNIT: Box | 1 | B. BRAUN MEDICAL (PTY) LTD | 1.3 | Omnifx 20ml | 4616200V | B. Braun | Germany | 42 | 50% | Euro | 100 | Yes | Box of 100 | COMPLY pending confirmation of latex and DEHP free | COMPLY
42142608-00044 | 108 | Syringe, three-part, plastic, latex free, non-pyrogenic and DEPH free. Type 2 (luer slip), with an eccentric nozzle. Nominal capacity 20ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 100 UNIT: Box |  | SUPRA HEALTHCARE JOHANNESBURG PTY LTD | 1.38 | Avacare | BS20MLE | HENAN AILE INDUSTRIAL CO.,LTD | CHINA | Initial 60 days, thereafter 21 | 84 | USD | 10 Boxes | YES | Box of 50 | COMPLY pending confirmation of latex and DEHP free | NOT COMPLY
42142608-00044 | 108 | Syringe, three-part, plastic, latex free, non-pyrogenic and DEPH free. Type 2 (luer slip), with an eccentric nozzle. Nominal capacity 20ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 100 UNIT: Box | N/A | SAFARMEX (PTY) LTD | 1.45 | Saclin | SSYN-10028 | Zhejiang Mediunion Healthcare Group Co. Ltd | China | 3days | 80% | US$ | 3000 Boxes | Yes | Box of 50 | COMPLY pending confirmation of latex and DEHP free | COMPLY
42142608-00044 | 108 | Syringe, three-part, plastic, latex free, non-pyrogenic and DEPH free. Type 2 (luer slip), with an eccentric nozzle. Nominal capacity 20ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 100 UNIT: Box | Basic Line offer | Becton Dickinson | 2.22 | BD Plastipak™ 20ml | 300613 | Becton Dickinson | Fraga, Spain | 7 Working Days | 76% | USD | 1 box | YES | 120 units in a box | COMPLY pending confirmation of latex and DEHP free | COMPLY
42142608-00044 | 108 | Syringe, three-part, plastic, latex free, non-pyrogenic and DEPH free. Type 2 (luer slip), with an eccentric nozzle. Nominal capacity 20ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 100 UNIT: Box |  | ALMIKA TRADING PTY LTD | 2.25 | BD | 300613 | BD | USA | 14 | 0 | NO | 100 | YES | 1 |  | COMPLY
42142608-00045 | 109 | Syringe, three-part, plastic, latex free, non-pyrogenic and DEPH free. Type 3 (luer lock), with a concentric nozzle. Nominal capacity 20ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 100 UNIT: Box | Basic Line Offer | BIOCOR HOSPITAL SUPPLIES PTY LTD | 0.68 | Triple -D | SYR-Care 20MILL | Jumla Medical | South Africa | 21 | 0 | N/A | 5 Boxes of 100 | Yes | N/A | COMPLY pending confirmation of latex and DEHP free | NO TEST REPORTS
42142608-00045 | 109 | Syringe, three-part, plastic, latex free, non-pyrogenic and DEPH free. Type 3 (luer lock), with a concentric nozzle. Nominal capacity 20ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 100 UNIT: Box |  | Isigidi Trading 95 cc | 1.48 | Isigidi | SY58 | Jiangsu Jichun Medical Devices Co. Ltd. | China | 14 days | 85% | USD | 20boxes | Yes | 50pcs/box | COMPLY pending confirmation of latex and DEHP free | COMPLY
42142608-00045 | 109 | Syringe, three-part, plastic, latex free, non-pyrogenic and DEPH free. Type 3 (luer lock), with a concentric nozzle. Nominal capacity 20ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 100 UNIT: Box | N/A | PROMED TECHNOLOGIES (PTY) LTD | 1.02 | U-LIFE | 608045 | JIAXING | CHINA | 14 DAYS | 85% | USD | 1 BOX | YES | BOX | COMPLY pending confirmation of latex and DEHP free | NO TEST REPORTS
42142608-00045 | 109 | Syringe, three-part, plastic, latex free, non-pyrogenic and DEPH free. Type 3 (luer lock), with a concentric nozzle. Nominal capacity 20ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 100 UNIT: Box |  | Mr First Aid (Pty) Ltd | 1.04 | Yue Kang |  | Yue Kang | China | 42 | 85 | USD | 10 | Yes | 50pc per box | COMPLY pending confirmation of latex and DEHP free | NO TEST REPORTS
42142608-00045 | 109 | Syringe, three-part, plastic, latex free, non-pyrogenic and DEPH free. Type 3 (luer lock), with a concentric nozzle. Nominal capacity 20ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 100 UNIT: Box |  | Medi-Core Technologies | 1.05 | ULIFE | 42142608-00045 | JIAXING ULIFE MEDICAL DEVICE TECHNOLOGY CO.,LTD | CHINA | 21 | 80 | USD | 2 box | YES | PRICED PER 100-PACKED 50PCS/BOX | No sample submitted for evaluation | NO TEST REPORTS
42142608-00045 | 109 | Syringe, three-part, plastic, latex free, non-pyrogenic and DEPH free. Type 3 (luer lock), with a concentric nozzle. Nominal capacity 20ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 100 UNIT: Box |  | GIBELA TRADE AND INVEST 1001 CC | 1.15 | TARA | SYR026 | Shandong Qiaopai Group Co., Ltd | CHINA | 21 | 85 | USD | 10 | YES | box 100 | No sample submitted for evaluation | NO TEST REPORTS
42142608-00045 | 109 | Syringe, three-part, plastic, latex free, non-pyrogenic and DEPH free. Type 3 (luer lock), with a concentric nozzle. Nominal capacity 20ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 100 UNIT: Box | 1 | Vital Life Medical Distributors (PTY) LTD | 1.2 | Healthease | NHPSLL-20ML | Neomedic | China | 14 days | 80% | US Dollar | 50 boxes | No | Pack size is boxes of 40's | COMPLY pending confirmation of latex and DEHP free | NO TEST REPORTS
42142608-00045 | 109 | Syringe, three-part, plastic, latex free, non-pyrogenic and DEPH free. Type 3 (luer lock), with a concentric nozzle. Nominal capacity 20ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 100 UNIT: Box |  | SUPRA HEALTHCARE JOHANNESBURG PTY LTD | 1.42 | Avacare | BS20ML/LC | HENAN AILE INDUSTRIAL CO.,LTD | CHINA | Initial 60 days, thereafter 21 | 84 | USD | 10 Boxes | YES | Box of 50 | COMPLY pending confirmation of latex and DEHP free | NOT COMPLY
42142608-00045 | 109 | Syringe, three-part, plastic, latex free, non-pyrogenic and DEPH free. Type 3 (luer lock), with a concentric nozzle. Nominal capacity 20ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 100 UNIT: Box | N/A | SAFARMEX (PTY) LTD | 1.45 | Saclin | To be confirmed | Zhejiang Mediunion Healthcare Group Co. Ltd | China | 3days | 80% | US$ | 3000 Boxes | Yes | Box of 50 | No sample submitted for evaluation | NO TEST REPORT
42142608-00045 | 109 | Syringe, three-part, plastic, latex free, non-pyrogenic and DEPH free. Type 3 (luer lock), with a concentric nozzle. Nominal capacity 20ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 100 UNIT: Box | 1 | B. BRAUN MEDICAL (PTY) LTD | 1.5 | Omnifix 20ml Luer Lock | 4617207V | B. Braun | Germany | 42 | 50% | Euro | 100 | Yes | Box of 100 | COMPLY pending confirmation of latex and DEHP free | NOT COMPLY
42142608-00045 | 109 | Syringe, three-part, plastic, latex free, non-pyrogenic and DEPH free. Type 3 (luer lock), with a concentric nozzle. Nominal capacity 20ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 100 UNIT: Box |  | ALMIKA TRADING PTY LTD | 2.25 | BD | 300629 | BD | USA | 14 | 0 | NO | 100 | YES | 1 | No sample submitted for evaluation. | NO TEST REPORTS
42142608-00045 | 109 | Syringe, three-part, plastic, latex free, non-pyrogenic and DEPH free. Type 3 (luer lock), with a concentric nozzle. Nominal capacity 20ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 100 UNIT: Box | 1 | CENTRAL MEDICAL PTY LTD | 100 | LYSA | HSCL20000 | Zhejian INI Medical Devices Co Ltd | China | 21 | 80 | China yuan | 16 boxes of 50 units | yes | box of 50 units | COMPLY pending confirmation of latex and DEHP free | COMPLY | Recommended | 100
42142608-00045 | 109 | Syringe, three-part, plastic, latex free, non-pyrogenic and DEPH free. Type 3 (luer lock), with a concentric nozzle. Nominal capacity 20ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 100 UNIT: Box BOX (BOX OF 50) |  | JUMLA MEDICAL SUPPLIES (PTY) LTD | 51.31 | SYR-CARE | SYR-CARE-20MLLL | SUZHOU ACMED | CHINA | 21 DAYS | 72% | US DOLLAR | 10 BOXES | YES | NIL | COMPLY pending confirmation of latex and DEHP free | NOT COMPLY
42142608-00046 | 110 | Syringe, three-part, plastic, latex free, non-pyrogenic and DEPH free. Type 3 (luer lock), with an eccentric nozzle. Nominal capacity 20ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 100 UNIT: Box | N/A | PROMED TECHNOLOGIES (PTY) LTD | 1.02 | U-LIFE | 608046 | JIAXING | CHINA | 14 DAYS | 85% | USD | 1 BOX | YES | BOX | No samples submitted for evaluation | NO TEST REPORTS | Re-tender
42142608-00046 | 110 | Syringe, three-part, plastic, latex free, non-pyrogenic and DEPH free. Type 3 (luer lock), with an eccentric nozzle. Nominal capacity 20ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 100 UNIT: Box |  | Mr First Aid (Pty) Ltd | 1.04 | Yue Kang |  | Yue Kang | China | 42 | 85 | USD | 10 | Yes | 50pc per box | COMPLY pending confirmation of latex and DEHP free | NO TEST REPORTS | Re-tender
42142608-00046 | 110 | Syringe, three-part, plastic, latex free, non-pyrogenic and DEPH free. Type 3 (luer lock), with an eccentric nozzle. Nominal capacity 20ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 100 UNIT: Box |  | Medi-Core Technologies | 1.05 | ULIFE | 42142608-00046 | JIAXING ULIFE MEDICAL DEVICE TECHNOLOGY CO.,LTD | CHINA | 21 | 80 | USD | 2 box | YES | PRICED PER 100-PACKED 50PCS/BOX | No sample submitted for evaluation | NO TEST REPORTS | Re-tender
42142608-00046 | 110 | Syringe, three-part, plastic, latex free, non-pyrogenic and DEPH free. Type 3 (luer lock), with an eccentric nozzle. Nominal capacity 20ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 100 UNIT: Box |  | GIBELA TRADE AND INVEST 1001 CC | 1.14 | TARA | SYR027 | Shandong Qiaopai Group Co., Ltd | CHINA | 21 | 85 | USD | 10 | YES | box 100 | No sample submitted for evaluation | NO TEST REPORTS | Re-tender
42142608-00046 | 110 | Syringe, three-part, plastic, latex free, non-pyrogenic and DEPH free. Type 3 (luer lock), with an eccentric nozzle. Nominal capacity 20ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 100 UNIT: Box | BOX OF 100 | MSIMED CC | 1.31 | KANG KANG | SYR01 | ZHENJIANG | CHINA | 25 | 20% | USD | 20000PCS | YES | BOX | No sample submitted for evaluation | NO TEST REPORTS | Re-tender
42142608-00046 | 110 | Syringe, three-part, plastic, latex free, non-pyrogenic and DEPH free. Type 3 (luer lock), with an eccentric nozzle. Nominal capacity 20ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 100 UNIT: Box | 1 | Biologica Pharmaceuticals Pty Ltd | 2.92 | Sol- M | 180020 | Sol- Millennium | China | 4-6 Weeks | 100% | USD | 1000 Boxes | 100% | N/A |  | NO TEST REPORTS | Re-tender
42142608-00047 | 111 | Syringe, three-part, plastic, latex free, non-pyrogenic and DEPH free. Type 2 (luer slip), with a concentric nozzle Nominal capacity: 50ml (scale extended to 60ml), must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 25 UNIT: Box |  | GIBELA TRADE AND INVEST 1001 CC | 0.4 | TARA | SYR028 | Shandong Qiaopai Group Co., Ltd | CHINA | 21 | 85 | USD | 10 | YES | box 25 | No sample submitted for evaluation | NO TEST REPORTS | Re-tender
42142608-00047 | 111 | Syringe, three-part, plastic, latex free, non-pyrogenic and DEPH free. Type 2 (luer slip), with a concentric nozzle Nominal capacity: 50ml (scale extended to 60ml), must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 25 UNIT: Box |  | Chemical World cc | 2.44 | SURG PLUS | SYRC050 | TONGDA | china | 3 | 85 | USD | 1000 | YES | 25 | COMPLY pending confirmation of latex free | NO TEST REPORTS | Re-tender
42142608-00047 | 111 | Syringe, three-part, plastic, latex free, non-pyrogenic and DEPH free. Type 2 (luer slip), with a concentric nozzle Nominal capacity: 50ml (scale extended to 60ml), must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 25 UNIT: Box |  | Mr First Aid (Pty) Ltd | 2.6 | Yue Kang |  | Yue Kang | China | 42 | 85 | USD | 10 | Yes | 25pc per box | No sample submitted for evaluation | NO TEST REPORTS | Re-tender
42142608-00047 | 111 | Syringe, three-part, plastic, latex free, non-pyrogenic and DEPH free. Type 2 (luer slip), with a concentric nozzle Nominal capacity: 50ml (scale extended to 60ml), must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 25 UNIT: Box | N/A | PROMED TECHNOLOGIES (PTY) LTD | 2.78 | U-LIFE | 608047 | JIAXING | CHINA | 14 DAYS | 85% | USD | 1 BOX | YES | BOX | Scale does not extend to 60ml | NO TEST REPORTS | Re-tender
42142608-00047 | 111 | Syringe, three-part, plastic, latex free, non-pyrogenic and DEPH free. Type 2 (luer slip), with a concentric nozzle Nominal capacity: 50ml (scale extended to 60ml), must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 25 UNIT: Box |  | Medi-Core Technologies | 2.8 | ULIFE | 42142608-00047 | JIAXING ULIFE MEDICAL DEVICE TECHNOLOGY CO.,LTD | CHINA | 21 | 80 | USD | 1 box | YES | BOX | scale does not extend to 60ml | NOT COMPLY | Re-tender
42142608-00047 | 111 | Syringe, three-part, plastic, latex free, non-pyrogenic and DEPH free. Type 2 (luer slip), with a concentric nozzle Nominal capacity: 50ml (scale extended to 60ml), must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 25 UNIT: Box | BOX OF 20 | MSIMED CC | 3.64 | KANG KANG | SYR02 | ZHENJIANG | CHINA | 25 | 20% | USD | 20000PCS | YES | BOX | Wrong sample, offered an eccentric instead concentric nozzle | NO TEST REPORTS | Re-tender
42142608-00047 | 111 | Syringe, three-part, plastic, latex free, non-pyrogenic and DEPH free. Type 2 (luer slip), with a concentric nozzle Nominal capacity: 50ml (scale extended to 60ml), must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 25 UNIT: Box |  | SUPRA HEALTHCARE JOHANNESBURG PTY LTD | 3.88 | Avacare | BS50ML | HENAN AILE INDUSTRIAL CO.,LTD | CHINA | Initial 60 days, thereafter 21 | 84 | USD | 10 Boxes | YES | Box of 50 | No sample submitted for evaluation | NO TEST REPORTS | Re-tender
42142608-00047 | 111 | Syringe, three-part, plastic, latex free, non-pyrogenic and DEPH free. Type 2 (luer slip), with a concentric nozzle Nominal capacity: 50ml (scale extended to 60ml), must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 25 UNIT: Box | Basic Line offer | Becton Dickinson | 10.08 | BD Plastipak™ 50ml | 300866 | Becton Dickinson | Fraga, Spain | 7 Working Days | 76% | USD | 1 box | YES | 60 units in a box | No sample submitted for evaluation. | NO TEST REPORTS | Re-tender
42142608-00048 | 112 | Syringe, three-part, plastic, latex free, non-pyrogenic and DEPH free. Type 2 (luer slip), with an eccentric nozzle Nominal capacity: 50ml (scale extended to 60ml), must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 25 UNIT: Box |  | GIBELA TRADE AND INVEST 1001 CC | 0.4 | TARA | SYR029 | Shandong Qiaopai Group Co., Ltd | CHINA | 21 | 85 | USD | 10 | YES | box 25 | No sample submitted for evaluation | NO TEST REPORTS
42142608-00048 | 112 | Syringe, three-part, plastic, latex free, non-pyrogenic and DEPH free. Type 2 (luer slip), with an eccentric nozzle Nominal capacity: 50ml (scale extended to 60ml), must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 25 UNIT: Box |  | Chemical World cc | 2.44 | SURG PLUS | SYR050 | TONGDA | china | 3 | 85 | USD | 18000 | YES | EACH | COMPLY pending confirmation of latex free | NOT COMPLY
42142608-00048 | 112 | Syringe, three-part, plastic, latex free, non-pyrogenic and DEPH free. Type 2 (luer slip), with an eccentric nozzle Nominal capacity: 50ml (scale extended to 60ml), must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 25 UNIT: Box |  | Mr First Aid (Pty) Ltd | 2.6 | Yue Kang |  | Yue Kang | China | 42 | 85 | USD | 10 | Yes | 25pc per box | No sample submitted for evaluation | NO TEST REPORTS
42142608-00048 | 112 | Syringe, three-part, plastic, latex free, non-pyrogenic and DEPH free. Type 2 (luer slip), with an eccentric nozzle Nominal capacity: 50ml (scale extended to 60ml), must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 25 UNIT: Box |  | Isigidi Trading 95 cc | 69 | Isigidi | SY60 | Jiangsu Jichun Medical Devices Co. Ltd. | China | 14 days | 85% | USD | 12boxes | Yes | 25pcs/box | COMPLY pending confirmation of latex and DEHP free | COMPLY | Recommended | 91
42142608-00048 | 112 | Syringe, three-part, plastic, latex free, non-pyrogenic and DEPH free. Type 2 (luer slip), with an eccentric nozzle Nominal capacity: 50ml (scale extended to 60ml), must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 25 UNIT: Box | N/A | PROMED TECHNOLOGIES (PTY) LTD | 2.78 | U-LIFE | 608048 | JIAXING | CHINA | 14 DAYS | 85% | USD | 1 BOX | YES | BOX | No samples submitted for evaluation | NO TEST REPORTS
42142608-00048 | 112 | Syringe, three-part, plastic, latex free, non-pyrogenic and DEPH free. Type 2 (luer slip), with an eccentric nozzle Nominal capacity: 50ml (scale extended to 60ml), must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 25 UNIT: Box |  | Medi-Core Technologies | 2.8 | ULIFE | 42142608-00048 | JIAXING ULIFE MEDICAL DEVICE TECHNOLOGY CO.,LTD | CHINA | 21 | 80 | USD | 1 box | YES | BOX | No sample submitted for evaluation | NO TEST REPORTS
42142608-00048 | 112 | Syringe, three-part, plastic, latex free, non-pyrogenic and DEPH free. Type 2 (luer slip), with an eccentric nozzle Nominal capacity: 50ml (scale extended to 60ml), must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 25 UNIT: Box | 1 | CENTRAL MEDICAL PTY LTD | 3 | LYSA | HSES50000 | Zhejian INI Medical Devices Co Ltd | China | 21 | 80 | China yuan | 16 boxes of 25 units | yes | box of 25 units | COMPLY pending confirmation of latex and DEHP free | NO TEST REPORTS
42142608-00048 | 112 | Syringe, three-part, plastic, latex free, non-pyrogenic and DEPH free. Type 2 (luer slip), with an eccentric nozzle Nominal capacity: 50ml (scale extended to 60ml), must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 25 UNIT: Box | 1 | Vital Life Medical Distributors (PTY) LTD | 3.25 | Healthease | NHPSLS - 60ml | Neomedic | China | 14 days | 80% | US Dollar | 50 boxes | Yes | 20 per box | COMPLY pending confirmation of latex and DEHP free | NO TEST REPORTS
42142608-00048 | 112 | Syringe, three-part, plastic, latex free, non-pyrogenic and DEPH free. Type 2 (luer slip), with an eccentric nozzle Nominal capacity: 50ml (scale extended to 60ml), must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 25 UNIT: Box | Basic Line Offer | BIOCOR HOSPITAL SUPPLIES PTY LTD | 3.33 | Triple -D | SYR-Care 50MLS | Jumla Medical | South Africa | 21 | 0 | N/A | 5 Boxes of 25 | Yes | N/A | COMPLY pending confirmation of latex and DEHP free | NO TEST REPORTS
42142608-00048 | 112 | Syringe, three-part, plastic, latex free, non-pyrogenic and DEPH free. Type 2 (luer slip), with an eccentric nozzle Nominal capacity: 50ml (scale extended to 60ml), must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 25 UNIT: Box | N/A | SAFARMEX (PTY) LTD | 3.59 | Saclin | SSYN-10029 | Zhejiang Mediunion Healthcare Group Co. Ltd | China | 3days | 80% | US$ | 3000 Boxes | Yes | Box of 25 | COMPLY pending confirmation of latex and DEHP free | COMPLY
42142608-00048 | 112 | Syringe, three-part, plastic, latex free, non-pyrogenic and DEPH free. Type 2 (luer slip), with an eccentric nozzle Nominal capacity: 50ml (scale extended to 60ml), must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 25 UNIT: Box | BOX OF 20 | MSIMED CC | 3.64 | KANG KANG | SYR03 | ZHENJIANG | CHINA | 25 | 20% | USD | 20000PCS | YES | BOX | No sample submitted for evaluation | NO TEST REPORTS
42142608-00048 | 112 | Syringe, three-part, plastic, latex free, non-pyrogenic and DEPH free. Type 2 (luer slip), with an eccentric nozzle Nominal capacity: 50ml (scale extended to 60ml), must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 25 UNIT: Box |  | SUPRA HEALTHCARE JOHANNESBURG PTY LTD | 3.72 | Avacare | BS50MLE | HENAN AILE INDUSTRIAL CO.,LTD | CHINA | Initial 60 days, thereafter 21 | 84 | USD | 10 Boxes | YES | Box of 50 | COMPLY pending confirmation of latex and DEHP free | NO TEST REPORTS
42142608-00048 | 112 | Syringe, three-part, plastic, latex free, non-pyrogenic and DEPH free. Type 2 (luer slip), with an eccentric nozzle Nominal capacity: 50ml (scale extended to 60ml), must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 25 UNIT: Box |  | ALMIKA TRADING PTY LTD | 4.89 | BD | 300866 | BD | USA | 14 | 0 | NO | 100 | YES | 1 | COMPLY pending confirmation of latex free and DEHP free | NO TEST REPORTS
42142608-00048 | 112 | Syringe, three-part, plastic, latex free, non-pyrogenic and DEPH free. Type 2 (luer slip), with an eccentric nozzle Nominal capacity: 50ml (scale extended to 60ml), must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 25 UNIT: Box | 1 | B. BRAUN MEDICAL (PTY) LTD | 5.9 | Omnifix 50ml Luer Slip Syringe (L/F) | 4616502F | B. Braun | Germany | 42 | 50% | Euro | 25 | Yes | Box of 25 | COMPLY pending confirmation of latex and DEHP free | NOT COMPLY
42142608-00048 | 112 | Syringe, three-part, plastic, latex free, non-pyrogenic and DEPH free. Type 2 (luer slip), with an eccentric nozzle Nominal capacity: 50ml (scale extended to 60ml), must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 25 UNIT: Box | 1 | Biologica Pharmaceuticals Pty Ltd | 6.96 | Sol- M | 180060ET | Sol- Millennium | China | 4-6 Weeks | 100% | USD | 1000 Boxes | 100% | N/A
42142608-00048 | 112 | Syringe, three-part, plastic, latex free, non-pyrogenic and DEPH free. Type 2 (luer slip), with an eccentric nozzle Nominal capacity: 50ml (scale extended to 60ml), must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 25 UNIT: Box |  | JUMLA MEDICAL SUPPLIES (PTY) LTD | 62.78 | SYR-CARE | SYR-CARE-50MLS | SUZHOU ACMED | CHINA | 21 DAYS | 72% | US DOLLAR | 10 BOXES | YES | NIL | Wrong sample, offered a 20ml instead of 50ml The outer package is labelled 50ml and the immediate package or specimen is 20ml | NOT COMPLY
42142608-00049 | 113 | Syringe, three-part, plastic, latex free, non-pyrogenic and DEPH free. Type 3 (luer lock), with a concentric nozzle Nominal capacity: 50ml (scale extended to 60ml), must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 25 UNIT: Box |  | GIBELA TRADE AND INVEST 1001 CC | 0.4 | TARA | SYR030 | Shandong Qiaopai Group Co., Ltd | CHINA | 21 | 85 | USD | 10 | YES | box 25 | No sample submitted for evaluation | NO TEST REPORTS
42142608-00049 | 113 | Syringe, three-part, plastic, latex free, non-pyrogenic and DEPH free. Type 3 (luer lock), with a concentric nozzle Nominal capacity: 50ml (scale extended to 60ml), must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 25 UNIT: Box |  | Isigidi Trading 95 cc | 60 | Isigidi | SY37 | Jiangsu Jichun Medical Devices Co. Ltd. | China | 14 days | 85% | USD | 12boxes | Yes | 25pcs/box | COMPLY pending confirmation of latex and DEHP free | COMPLY | Recommended | 91
42142608-00049 | 113 | Syringe, three-part, plastic, latex free, non-pyrogenic and DEPH free. Type 3 (luer lock), with a concentric nozzle Nominal capacity: 50ml (scale extended to 60ml), must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 25 UNIT: Box |  | Chemical World cc | 2.49 | SURGI PLUS | SYRLL050 | TONGDA | china | 3 | 85 | USD | 1000 | YES | 25 | COMPLY pending confirmation of latex free | NOT COMPLY
42142608-00049 | 113 | Syringe, three-part, plastic, latex free, non-pyrogenic and DEPH free. Type 3 (luer lock), with a concentric nozzle Nominal capacity: 50ml (scale extended to 60ml), must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 25 UNIT: Box |  | Mr First Aid (Pty) Ltd | 2.6 | Yue Kang |  | Yue Kang | China | 42 | 85 | USD | 10 | Yes | 25pc per box | No sample submitted for evaluation | NO TEST REPORTS
42142608-00049 | 113 | Syringe, three-part, plastic, latex free, non-pyrogenic and DEPH free. Type 3 (luer lock), with a concentric nozzle Nominal capacity: 50ml (scale extended to 60ml), must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 25 UNIT: Box | N/A | PROMED TECHNOLOGIES (PTY) LTD | 2.78 | U-LIFE | 608049 | JIAXING | CHINA | 14 DAYS | 85% | USD | 1 BOX | YES | BOX | No samples submitted for evaluation | NO TEST REPORTS
42142608-00049 | 113 | Syringe, three-part, plastic, latex free, non-pyrogenic and DEPH free. Type 3 (luer lock), with a concentric nozzle Nominal capacity: 50ml (scale extended to 60ml), must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 25 UNIT: Box | 1 | CENTRAL MEDICAL PTY LTD | 2.85 | LYSA | HSCL50000 | Zhejian INI Medical Devices Co Ltd | China | 21 | 80 | China yuan | 16 boxes of 25 units | yes | box of 25 units | COMPLY pending confirmation of latex and DEHP free | COMPLY
42142608-00049 | 113 | Syringe, three-part, plastic, latex free, non-pyrogenic and DEPH free. Type 3 (luer lock), with a concentric nozzle Nominal capacity: 50ml (scale extended to 60ml), must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 25 UNIT: Box |  | Medi-Core Technologies | 2.85 | ULIFE | 42142608-00049 | JIAXING ULIFE MEDICAL DEVICE TECHNOLOGY CO.,LTD | CHINA | 21 | 80 | USD | 1 box | YES | BOX | COMPLY pending confirmation of latex and DEHP free | NO TEST REPORTS
42142608-00049 | 113 | Syringe, three-part, plastic, latex free, non-pyrogenic and DEPH free. Type 3 (luer lock), with a concentric nozzle Nominal capacity: 50ml (scale extended to 60ml), must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 25 UNIT: Box | 1 | Vital Life Medical Distributors (PTY) LTD | 3.25 | Healthease | NHPSLL-60ML | Neomedic | China | 14 days | 80% | US Dollar | 50 boxes | Yes | 20 per box | COMPLY pending confirmation of latex and DEHP free | NO TEST REPORTS
42142608-00049 | 113 | Syringe, three-part, plastic, latex free, non-pyrogenic and DEPH free. Type 3 (luer lock), with a concentric nozzle Nominal capacity: 50ml (scale extended to 60ml), must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 25 UNIT: Box | Basic Line Offer | BIOCOR HOSPITAL SUPPLIES PTY LTD | 3.38 | Triple -D | SYR-Care 50ML | Jumla Medical | South Africa | 21 | 0 | N/A | 5 Boxes of 25 | Yes | N/A | Scale does not extend to 60ml | NO TEST REPORTS
42142608-00049 | 113 | Syringe, three-part, plastic, latex free, non-pyrogenic and DEPH free. Type 3 (luer lock), with a concentric nozzle Nominal capacity: 50ml (scale extended to 60ml), must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 25 UNIT: Box | BOX OF 20 | MSIMED CC | 3.64 | KANG KANG | SYR04 | ZHENJIANG | CHINA | 25 | 20% | USD | 20000PCS | YES | BOX | COMPLY pending confirmation of latex and DEHP free | NO TEST REPORTS
42142608-00049 | 113 | Syringe, three-part, plastic, latex free, non-pyrogenic and DEPH free. Type 3 (luer lock), with a concentric nozzle Nominal capacity: 50ml (scale extended to 60ml), must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 25 UNIT: Box |  | SUPRA HEALTHCARE JOHANNESBURG PTY LTD | 3.72 | Avacare | BS50L/LC | HENAN AILE INDUSTRIAL CO.,LTD | CHINA | Initial 60 days, thereafter 21 | 84 | USD | 10 Boxes | YES | Box of 50 | COMPLY pending confirmation of latex and DEHP free | NO TEST REPORTS
42142608-00049 | 113 | Syringe, three-part, plastic, latex free, non-pyrogenic and DEPH free. Type 3 (luer lock), with a concentric nozzle Nominal capacity: 50ml (scale extended to 60ml), must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 25 UNIT: Box |  | ALMIKA TRADING PTY LTD | 4.63 | BD | 300865 | BD | USA | 14 | 0 | NO | 100 | YES | 1 | COMPLY pending confirmation of latex free and DEHP free | NO TEST REPORTS
42142608-00049 | 113 | Syringe, three-part, plastic, latex free, non-pyrogenic and DEPH free. Type 3 (luer lock), with a concentric nozzle Nominal capacity: 50ml (scale extended to 60ml), must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 25 UNIT: Box | 1 | Biologica Pharmaceuticals Pty Ltd | 7.13 | Sol- M | 180060 | Sol- Millennium | China | 4-6 Weeks | 100% | USD | 1000 Boxes | 100% | N/A
42142608-00049 | 113 | Syringe, three-part, plastic, latex free, non-pyrogenic and DEPH free. Type 3 (luer lock), with a concentric nozzle Nominal capacity: 50ml (scale extended to 60ml), must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 25 UNIT: Box | Basic Line offer | Becton Dickinson | 9.6 | BD Plastipak™ 50ml, Luer lok | 300865 | Becton Dickinson | Fraga, Spain | 7 Working Days | 78% | USD | 1 box | YES | 60 units in a box | COMPLY pending confirmation of latex and DEHP free | NOT COMPLY
42142608-00049 | 113 | Syringe, three-part, plastic, latex free, non-pyrogenic and DEPH free. Type 3 (luer lock), with a concentric nozzle Nominal capacity: 50ml (scale extended to 60ml), must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 25 UNIT: Box | 1 | Fresenius Kabi South Africa | 11.75 | Injectomat Luer Lock Syringe 50ml | FPD9000711 | Fresenius Kabi | France | ex stock (7 Days) or 6-8 weeks | 87.5% | Euro | 100 | Yes | Box | COMPLY pending confirmation of latex free | NO TEST REPORTS
42142608-00049 | 113 | Syringe, three-part, plastic, latex free, non-pyrogenic and DEPH free. Type 3 (luer lock), with a concentric nozzle Nominal capacity: 50ml (scale extended to 60ml), must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 25 UNIT: Box |  | JUMLA MEDICAL SUPPLIES (PTY) LTD | 63.54 | SYR-CARE | SYR-CARE-50MLL | SUZHOU ACMED | CHINA | 21 DAYS | 72% | US DOLLAR | 10 BOXES | YES | NIL | Offered a 50m and scale is not extended to 60ml | NO TEST REPORTS
42142608-00050 | 114 | Syringe, three-part, plastic, latex free, non-pyrogenic and DEPH free. Type 3 (luer lock), with an eccentric nozzle Nominal capacity: 50ml (scale extended to 60ml), must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 25 UNIT: Box |  | Mr First Aid (Pty) Ltd | 2.65 | Yue Kang |  | Yue Kang | China | 42 | 85 | USD | 10 | Yes | 25pc per box | Item to be deleted | NO TEST REPORTS | Item cancelled:End-users to proqure using their delagated powers.
42142608-00050 | 114 | Syringe, three-part, plastic, latex free, non-pyrogenic and DEPH free. Type 3 (luer lock), with an eccentric nozzle Nominal capacity: 50ml (scale extended to 60ml), must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 25 UNIT: Box | N/A | PROMED TECHNOLOGIES (PTY) LTD | 2.78 | U-LIFE | 608050 | JIAXING | CHINA | 14 DAYS | 85% | USD | 1 BOX | YES | BOX | Item deleted |  | Item cancelled:End-users to proqure using their delagated powers.
42142608-00050 | 114 | Syringe, three-part, plastic, latex free, non-pyrogenic and DEPH free. Type 3 (luer lock), with an eccentric nozzle Nominal capacity: 50ml (scale extended to 60ml), must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 25 UNIT: Box |  | Medi-Core Technologies | 2.85 | ULIFE | 42142608-00050 | JIAXING ULIFE MEDICAL DEVICE TECHNOLOGY CO.,LTD | CHINA | 21 | 80 | USD | 1 box | YES | BOX | Item deleted |  | Item cancelled:End-users to proqure using their delagated powers.
42142608-00050 | 114 | Syringe, three-part, plastic, latex free, non-pyrogenic and DEPH free. Type 3 (luer lock), with an eccentric nozzle Nominal capacity: 50ml (scale extended to 60ml), must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 25 UNIT: Box | BOX OF 20 | MSIMED CC | 3.64 | KANG KANG | SYR05 | ZHENJIANG | CHINA | 25 | 20% | USD | 20000PCS | YES | BOX | Item deleted | NO TEST REPORTS | Item cancelled:End-users to proqure using their delagated powers.
42142608-00050 | 114 | Syringe, three-part, plastic, latex free, non-pyrogenic and DEPH free. Type 3 (luer lock), with an eccentric nozzle Nominal capacity: 50ml (scale extended to 60ml), must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 25 UNIT: Box | 1 | B. BRAUN MEDICAL (PTY) LTD | 5.93 | Omnifix 50ml Luer Lock | 4617509F | B. Braun | Germany | 42 | 50% | Euro | 25 | Yes | Box of 25 | Item deleted. | NOT COMPLY | Item cancelled:End-users to proqure using their delagated powers.
42142608-00050 | 114 | Syringe, three-part, plastic, latex free, non-pyrogenic and DEPH free. Type 3 (luer lock), with an eccentric nozzle Nominal capacity: 50ml (scale extended to 60ml), must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 25 UNIT: Box | 1 | Biologica Pharmaceuticals Pty Ltd | 7.13 | Sol- M | 180060 | Sol- Millennium | China | 4-6 Weeks | 100% | USD | 1000 Boxes | 100% | N/A |  |  | Item cancelled:End-users to proqure using their delagated powers.
42142604-00000 | 115 | Syringe, three-part, plastic, latex free, non-pyrogenic and DEPH free. Type 4 (catheter tip) Nominal capacity: 50ml (scale extended to 60ml), must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Single use, sterile, individually packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 25 UNIT: Box |  | GIBELA TRADE AND INVEST 1001 CC | 2.56 | TARA | SYR031 | Shandong Qiaopai Group Co., Ltd | CHINA | 21 | 85 | USD | 10 | YES | box 25 | No sample submitted for evaluation | NO TEST REPORTS
42142604-00000 | 115 | Syringe, three-part, plastic, latex free, non-pyrogenic and DEPH free. Type 4 (catherter tip), with a concentric nozzle Nominal capacity: 50ml (scale extended to 60ml), must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Sterile, Single use, Individually peel packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 25 UNIT: Box |  | Chemical World cc | 2.58 | [10 empty] | COMPLY pending confirmation of latex free | NOT COMPLY
42142604-00000 | 115 | Syringe, three-part, plastic, latex free, non-pyrogenic and DEPH free. Type 4 (catheter tip) Nominal capacity: 50ml (scale extended to 60ml), must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Single use, sterile, individually packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 25 UNIT: Box |  | Medi-Core Technologies | 2.79 | ULIFE | 42142604-00000 | JIAXING U-LIFE MEDICAL DEVICE TECHNOLOGY CO.,LTD | CHINA | 21 | 80 | USD | 1 box | YES | BOX | scale does not extend to 60ml | NOT COMPLY
42142604-00000 | 115 | Syringe, three-part, plastic, latex free, non-pyrogenic and DEPH free. Type 4 (catheter tip) Nominal capacity: 50ml (scale extended to 60ml), must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Single use, sterile, individually packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 25 UNIT: Box | 1 | CENTRAL MEDICAL PTY LTD | 2.84 | LYSA | CTSS50000 | Zhejian INI Medical Devices Co Ltd | China | 21 | 80 | China yuan | 16 boxes of 25 units | yes | Box of 25 units | COMPLY pending confirmation of latex and DEHP free | NO TEST REPORTS
42142604-00000 | 115 | Syringe, three-part, plastic, latex free, non-pyrogenic and DEPH free. Type 4 (catheter tip) Nominal capacity: 50ml (scale extended to 60ml), must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Single use, sterile, individually packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 25 UNIT: Box | 1 | Vital Life Medical Distributors (PTY) LTD | 3.25 | Healthease | NHPSCT-60ML | Neomedic | China | 14 days | 80% | US Dollar | 50 Boxes | No | Pack size is boxes of 20's | COMPLY pending confirmation of latex and DEHP free | NO TEST REPORTS
42142604-00000 | 115 | Syringe, three-part, plastic, latex free, non-pyrogenic and DEPH free. Type 4 (catheter tip) Nominal capacity: 50ml (scale extended to 60ml), must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Single use, sterile, individually packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 25 UNIT: Box | BOX OF 20 | MSIMED CC | 3.64 | KANG KANG | SYR06 | ZHENJIANG | CHINA | 25 | 20% | USD | 20000PCS | YES | BOX | No sample submitted for evaluation | NO TEST REPORTS
42142604-00000 | 115 | Syringe, three-part, plastic, latex free, non-pyrogenic and DEPH free. Type 4 (catheter tip) Nominal capacity: 50ml (scale extended to 60ml), must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Single use, sterile, individually packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 25 UNIT: Box |  | SUPRA HEALTHCARE JOHANNESBURG PTY LTD | 3.83 | Avacare | BS50CT | HENAN AILE INDUSTRIAL CO.,LTD | CHINA | Initial 60 days, thereafter 21 | 84 | USD | 10 Boxes | YES |  | COMPLY pending confirmation of latex and DEHP free | NO TEST REPORTS
42142604-00000 | 115 | Syringe, three-part, plastic, latex free, non-pyrogenic and DEPH free. Type 4 (catheter tip) Nominal capacity: 50ml (scale extended to 60ml), must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Single use, sterile, individually packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 25 UNIT: Box |  | Isigidi Trading 95 cc | 110.4 | Isigidi | SY61 | Jiangsu Jichun Medical Devices Co. Ltd. | China | 14 days | 85% | USD | 12boxes | Yes | 25pcs/box | COMPLY pending confirmation of latex and DEHP free | COMPLY | Recommended | 91
42142604-00000 | 115 | Syringe, three-part, plastic, latex free, non-pyrogenic and DEPH free. Type 4 (catheter tip) Nominal capacity: 50ml (scale extended to 60ml), must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Single use, sterile, individually packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 25 UNIT: Box |  | ALMIKA TRADING PTY LTD | 5.03 | BD | 300867 | BD | USA | 14 | 0 | NO | 100 | YES | 1 | COMPLY pending confirmation of latex free and DEHP free | NO TEST REPORTS
42142604-00000 | 115 | Syringe, three-part, plastic, latex free, non-pyrogenic and DEPH free. Type 4 (catheter tip) Nominal capacity: 50ml (scale extended to 60ml), must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Single use, sterile, individually packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 25 UNIT: Box | Basic Line offer | Becton Dickinson | 10.08 | BD Plastipak™ 50ml Catheter Tip | 300867 | Becton Dickinson | Fraga, Spain | 7 Working Days | 78% | USD | 1 box | YES | 60 units in a box | COMPLY pending confirmation of latex and DEHP free | NO TEST REPORTS
42142604-00000 | 115 | Syringe, three-part, plastic, latex free, non-pyrogenic and DEPH free. Type 4 (catheter tip) Nominal capacity: 50ml (scale extended to 60ml), must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. Single use, sterile, individually packed, To comply with the latest issue of SANS 1124-2 Boxed in suitable shelf packs Box of 25 UNIT: Box |  | JUMLA MEDICAL SUPPLIES (PTY) LTD | 64.14 | SYR-CARE | SYR-CARE-50MLC | SUZHOU ACMED | CHINA | 21 DAYS | 72% | US DOLLAR | 10 BOXES | YES | NIL | COMPLY pending confirmation of latex and DEHP free | NOT COMPLY
42142609-00022 | 116 | Syringe, insulin, plastic, latex free, non-pyrogenic and DEPH free. Type 4 (with a fixed needle). Nominal capacity: 0,5 ml, 100 units per 1ml, must have clear graduation markings that are not easily removed or rubbed off. Needle: 29G (0,33 mm) x 12,5 - 12,7 mm Single use, sterile path To comply with the latest issue of SANS 1166 Packed in units of 10 syringes in a plastic packet 10 Packets in a box. UNIT: Box | N/A | PROMED TECHNOLOGIES (PTY) LTD | 0.68 | U-LIFE | 609022 | JIAXING | CHINA | 14 DAYS | 85% | USD | 1 BOX | YES | BOX | COMPLY pending confirmation of latex and DEHP free | NO TEST REPORTS
42142609-00022 | 116 | Syringe, insulin, plastic, latex free, non-pyrogenic and DEPH free. Type 4 (with a fixed needle). Nominal capacity: 0,5 ml, 100 units per 1ml, must have clear graduation markings that are not easily removed or rubbed off. Needle: 29G (0,33 mm) x 12,5 - 12,7 mm Single use, sterile path To comply with the latest issue of SANS 1166 Packed in units of 10 syringes in a plastic packet 10 Packets in a box. UNIT: Box |  | SUPRA HEALTHCARE JOHANNESBURG PTY LTD | 0.86 | REVITAL | IS0529P | Revital Health | KENYA | Initial 60 days, thereafter 21 | 84 | USD | 10 Boxes | YES |  | No sample submitted for evaluation | NO TEST REPORTS
42142609-00022 | 116 | Syringe, insulin, plastic, latex free, non-pyrogenic and DEPH free. Type 4 (with a fixed needle). Nominal capacity: 0,5 ml, 100 units per 1ml, must have clear graduation markings that are not easily removed or rubbed off. Needle: 29G (0,33 mm) x 12,5 - 12,7 mm Single use, sterile path To comply with the latest issue of SANS 1166 Packed in units of 10 syringes in a plastic packet 10 Packets in a box. UNIT: Box | N/A | SAFARMEX (PTY) LTD | 0.93 | Saclin | To be confirmed | Zhejiang Mediunion Healthcare Group Co. Ltd | China | 3days | 80% | US$ | 3000 Boxes | Yes | 10 Packets in a Box | No sample submitted for evaluation | NO TEST REPORT
42142609-00022 | 116 | Syringe, insulin, plastic, latex free, non-pyrogenic and DEPH free. Type 4 (with a fixed needle). Nominal capacity: 0,5 ml, 100 units per 1ml, must have clear graduation markings that are not easily removed or rubbed off. Needle: 29G (0,33 mm) x 12,5 - 12,7 mm Single use, sterile path To comply with the latest issue of SANS 1166 Packed in units of 10 syringes in a plastic packet 10 Packets in a box. UNIT: Box |  | Isigidi Trading 95 cc | 63.25 | Isigidi | SY40 | Jiangsu Jichun Medical Devices Co. Ltd. | China | 14 days | 85% | USD | 30boxes | Yes | 100pcs/box | COMPLY pending confirmation of latex and DEHP free | COMPLY (ACCEPTED WITH A DEVIATION OF 13MM) | Recommended | 91
42142609-00022 | 116 | Syringe, insulin, plastic, latex free, non-pyrogenic and DEPH free. Type 4 (with a fixed needle). Nominal capacity: 0,5 ml, 100 units per 1ml, must have clear graduation markings that are not easily removed or rubbed off. Needle: 29G (0,33 mm) x 12,5 - 12,7 mm Single use, sterile path To comply with the latest issue of SANS 1166 Packed in units of 10 syringes in a plastic packet 10 Packets in a box. UNIT: Box | Basic Line offer | Becton Dickinson | 6.4 | BD Microfine Plus | 320921 | Becton Dickinson | USA | 7 Working Days | 76 | USD | 1 BOX | Yes | 1 Box -100 Units | COMPLY pending confirmation of latex and DEHP free | NO TEST REPORTS
42142609-00022 | 116 | Syringe, insulin, plastic, latex free, non-pyrogenic and DEPH free. Type 4 (with a fixed needle). Nominal capacity: 0,5 ml, 100 units per 1ml, must have clear graduation markings that are not easily removed or rubbed off. Needle: 29G (0,33 mm) x 12,5 - 12,7 mm Single use, sterile path To comply with the latest issue of SANS 1166 Packed in units of 10 syringes in a plastic packet 10 Packets in a box. UNIT: Box | 1 | CENTRAL MEDICAL PTY LTD | 6.98 | LYSA | INSL052913 | Zhejian INI Medical Devices Co Ltd | China | 21 | 80 | China yuan | 20 boxes of 10 packets of 10 units | yes | Box of 10 packets of 10 unit | COMPLY pending confirmation of latex and DEHP free | COMPLY
42142609-00022 | 116 | Syringe, insulin, plastic, latex free, non-pyrogenic and DEPH free. Type 4 (with a fixed needle). Nominal capacity: 0,5 ml, 100 units per 1ml, must have clear graduation markings that are not easily removed or rubbed off. Needle: 29G (0,33 mm) x 12,5 - 12,7 mm Single use, sterile path To comply with the latest issue of SANS 1166 Packed in units of 10 syringes in a plastic packet 10 Packets in a box. UNIT: Box |  | GIBELA TRADE AND INVEST 1001 CC | 17.43 | TARA | SYR032 | Shandong Qiaopai Group Co., Ltd | CHINA | 21 | 85 | USD | 10 | YES | box 100 | No sample submitted for evaluation | NO TEST REPORTS
42142609-00023 | 117 | Syringe, insulin, plastic, latex free, non-pyrogenic and DEPH free. Type 4 (with a fixed needle). Nominal capacity: 0,5 ml, 100 units per 1 ml, must have clear graduation markings that are not easily removed or rubbed off. Needle: 29G (0,33 mm) x 12,5 - 12,7 mm Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 1166 Packed in box of 100. UNIT: Box |  | Isigidi Trading 95 cc | 0.64 | Isigidi | SY45 | Jiangsu Jichun Medical Devices Co. Ltd. | China | 14 days | 85% | USD | 30boxes | Yes | 100pcs/box | COMPLY pending confirmation of latex and DEHP free | COMPLY (ACCEPTED WITH A DEVIATION OF 13MM)
42142609-00023 | 117 | Syringe, insulin, plastic, latex free, non-pyrogenic and DEPH free. Type 4 (with a fixed needle). Nominal capacity: 0,5 ml, 100 units per 1 ml, must have clear graduation markings that are not easily removed or rubbed off. Needle: 29G (0,33 mm) x 12,5 - 12,7 mm Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 1166 Packed in box of 100. UNIT: Box | N/A | PROMED TECHNOLOGIES (PTY) LTD | 68 | U-LIFE | 609023 | JIAXING | CHINA | 14 DAYS | 85% | USD | 1 BOX | YES | BOX OF 100 | COMPLY pending confirmation of latex and DEHP free | COMPLY | Recommended | 93.70
42142609-00023 | 117 | Syringe, insulin, plastic, latex free, non-pyrogenic and DEPH free. Type 4 (with a fixed needle). Nominal capacity: 0,5 ml, 100 units per 1 ml, must have clear graduation markings that are not easily removed or rubbed off. Needle: 29G (0,33 mm) x 12,5 - 12,7 mm Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 1166 Packed in box of 100. UNIT: Box | 1 | CENTRAL MEDICAL PTY LTD | 0.69 | LYSA | INSL052913 | Zhejian INI Medical Devices Co Ltd | China | 21 | 80 | China yuan | 20 boxes of 100 units | yes | Box of 100 units | Offered a length of 13mm | NO TEST REPORTS
42142609-00023 | 117 | Syringe, insulin, plastic, latex free, non-pyrogenic and DEPH free. Type 4 (with a fixed needle). Nominal capacity: 0,5 ml, 100 units per 1 ml, must have clear graduation markings that are not easily removed or rubbed off. Needle: 29G (0,33 mm) x 12,5 - 12,7 mm Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 1166 Packed in box of 100. UNIT: Box |  | Chemical World cc | 0.69 | SURG PLUS | SYRI05 | TONGDA | china | 3 | 85 | USD | 2000 | YES | 100 | COMPLY pending confirmation of latex free | NOT COMPLY
42142609-00023 | 117 | Syringe, insulin, plastic, latex free, non-pyrogenic and DEPH free. Type 4 (with a fixed needle). Nominal capacity: 0,5 ml, 100 units per 1 ml, must have clear graduation markings that are not easily removed or rubbed off. Needle: 29G (0,33 mm) x 12,5 - 12,7 mm Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 1166 Packed in box of 100. UNIT: Box | 1 | Vital Life Medical Distributors (PTY) LTD | 0.81 | Healthease | NHIS-0.5ML-29-0.5 | Neomedic | China | 14 days | 80% | US Dollar | 50 Boxes | No | Needle 0.33mm x 13mm 100's per box | No sample submitted for evaluation | NO TEST REPORTS
42142609-00023 | 117 | Syringe, insulin, plastic, latex free, non-pyrogenic and DEPH free. Type 4 (with a fixed needle). Nominal capacity: 0,5 ml, 100 units per 1 ml, must have clear graduation markings that are not easily removed or rubbed off. Needle: 29G (0,33 mm) x 12,5 - 12,7 mm Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 1166 Packed in box of 100. UNIT: Box | Basic Line offer | Becton Dickinson | 0.82 | BD Microfine Plus | 324892 | BECTON DICKINSON | USA | 7 Working Days | 76% | USD | 1 BOX | Yes | 1 Box =100 Units | COMPLY pending confirmation of latex and DEHP free | NO TEST REPORTS
42142609-00023 | 117 | Syringe, insulin, plastic, latex free, non-pyrogenic and DEPH free. Type 4 (with a fixed needle). Nominal capacity: 0,5 ml, 100 units per 1 ml, must have clear graduation markings that are not easily removed or rubbed off. Needle: 29G (0,33 mm) x 12,5 - 12,7 mm Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 1166 Packed in box of 100. UNIT: Box |  | SUPRA HEALTHCARE JOHANNESBURG PTY LTD | 0.82 | REVITAL | IS0529B | Revital Health | KENYA | Initial 60 days, thereafter 21 | 84 | USD | 10 Boxes | YES |  | wrong sample, offered 5ml hypodermic syringe instead of an insulin syringe | NO TEST REPORTS
42142609-00023 | 117 | Syringe, insulin, plastic, latex free, non-pyrogenic and DEPH free. Type 4 (with a fixed needle). Nominal capacity: 0,5 ml, 100 units per 1 ml, must have clear graduation markings that are not easily removed or rubbed off. Needle: 29G (0,33 mm) x 12,5 - 12,7 mm Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 1166 Packed in box of 100. UNIT: Box | N/A | SAFARMEX (PTY) LTD | 0.9 | Saclin | To be confirmed | Zhejiang Mediunion Healthcare Group Co. Ltd | China | 3days | 80% | US$ | 3000 Boxes | Yes | Packed in a Box of 100 | No sample submitted for evaluation | NO TEST REPORT
42142609-00023 | 117 | Syringe, insulin, plastic, latex free, non-pyrogenic and DEPH free. Type 4 (with a fixed needle). Nominal capacity: 0,5 ml, 100 units per 1 ml, must have clear graduation markings that are not easily removed or rubbed off. Needle: 29G (0,33 mm) x 12,5 - 12,7 mm Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 1166 Packed in box of 100. UNIT: Box |  | GIBELA TRADE AND INVEST 1001 CC | 1.75 | TARA | SYR033 | Shandong Qiaopai Group Co., Ltd | CHINA | 21 | 85 | USD | 10 | YES | box 100 | No sample submitted for evaluation | NO TEST REPORTS
42142609-00023 | 117 | Syringe, insulin, plastic, latex free, non-pyrogenic and DEPH free. Type 4 (with a fixed needle). Nominal capacity: 0,5 ml, 100 units per 1 ml, must have clear graduation markings that are not easily removed or rubbed off. Needle: 29G (0,33 mm) x 12,5 - 12,7 mm Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 1166 Packed in box of 100. UNIT: Box | 1 | Biologica Pharmaceuticals Pty Ltd | 1.83 | Sol- M | 1652912B | Sol- Millennium | China | 4-6 Weeks | 100% | USD | 1000 Boxes | 100% | N/A |  | NO TEST REPORTS
42142609-00024 | 118 | Syringe, insulin, plastic, latex free, non-pyrogenic and DEPH free. Type 4 (with a fixed needle). Nominal capacity: 1 ml, 100 units per ml, must have clear graduation markings that are not easily removed or rubbed off. Needle: 29G (0,33 mm) x 12,5 - 12,7 mm Single use, sterile path To comply with the latest issue of SANS 1166 Packed in units of 10 syringes in a plastic packet 10 Packets in a box. UNIT: Box | N/A | PROMED TECHNOLOGIES (PTY) LTD | 0.68 | U-LIFE | 609024 | JIAXING | CHINA | 14 DAYS | 85% | USD | 1 BOX | YES | BOX | COMPLY pending confirmation of latex and DEHP free | NO TEST REPORTS
42142609-00024 | 118 | Syringe, insulin, plastic, latex free, non-pyrogenic and DEPH free. Type 4 (with a fixed needle). Nominal capacity: 1 ml, 100 units per ml, must have clear graduation markings that are not easily removed or rubbed off. Needle: 29G (0,33 mm) x 12,5 - 12,7 mm Single use, sterile path To comply with the latest issue of SANS 1166 Packed in units of 10 syringes in a plastic packet 10 Packets in a box. UNIT: Box |  | SUPRA HEALTHCARE JOHANNESBURG PTY LTD | 0.86 | REVITAL | IS0129P | Revital Health | KENYA | Initial 60 days, thereafter 21 | 84 | USD | 10 Boxes | YES |  | No sample submitted for evaluation | NO TEST REPORTS
42142609-00024 | 118 | Syringe, insulin, plastic, latex free, non-pyrogenic and DEPH free. Type 4 (with a fixed needle). Nominal capacity: 1 ml, 100 units per ml, must have clear graduation markings that are not easily removed or rubbed off. Needle: 29G (0,33 mm) x 12,5 - 12,7 mm Single use, sterile path To comply with the latest issue of SANS 1166 Packed in units of 10 syringes in a plastic packet 10 Packets in a box. UNIT: Box | N/A | SAFARMEX (PTY) LTD | 0.93 | Saclin | To be confirmed | Zhejiang Mediunion Healthcare Group Co. Ltd | China | 3days | 80% | US$ | 3000 Boxes | Yes | 10 Packets in a Box | No sample submitted for evaluation | NO TEST REPORT
42142609-00024 | 118 | Syringe, insulin, plastic, latex free, non-pyrogenic and DEPH free. Type 4 (with a fixed needle). Nominal capacity: 1 ml, 100 units per ml, must have clear graduation markings that are not easily removed or rubbed off. Needle: 29G (0,33 mm) x 12,5 - 12,7 mm Single use, sterile path To comply with the latest issue of SANS 1166 Packed in units of 10 syringes in a plastic packet 10 Packets in a box. UNIT: Box | 1 | Biologica Pharmaceuticals Pty Ltd | 1.83 | Sol- M | 1612912B | Sol- Millennium | China | 4-6 Weeks | 100% | USD | 1000 Boxes | 100% | N/A |  | NO TEST REPORTS
42142609-00024 | 118 | Syringe, insulin, plastic, latex free, non-pyrogenic and DEPH free. Type 4 (with a fixed needle). Nominal capacity: 1 ml, 100 units per ml, must have clear graduation markings that are not easily removed or rubbed off. Needle: 29G (0,33 mm) x 12,5 - 12,7 mm Single use, sterile path To comply with the latest issue of SANS 1166 Packed in units of 10 syringes in a plastic packet 10 Packets in a box. UNIT: Box |  | Isigidi Trading 95 cc | 63.25 | Isigidi | SY54 | Jiangsu Jichun Medical Devices Co. Ltd. | China | 14 days | 85% | USD | 30boxes | Yes | 100pcs/box | COMPLY pending confirmation of latex and DEHP free | COMPLY (ACCEPTED WITH A DEVIATION OF 13MM) | Recommended | 91
42142609-00024 | 118 | Syringe, insulin, plastic, latex free, non-pyrogenic and DEPH free. Type 4 (with a fixed needle). Nominal capacity: 1 ml, 100 units per ml, must have clear graduation markings that are not easily removed or rubbed off. Needle: 29G (0,33 mm) x 12,5 - 12,7 mm Single use, sterile path To comply with the latest issue of SANS 1166 Packed in units of 10 syringes in a plastic packet 10 Packets in a box. UNIT: Box | Basic Line offer | Becton Dickinson | 6.4 | BD Microfine Plus | 320909 | BECTON DICKINSON | USA | 7 Working Days | 76% | USD | 1 BOX | Yes | 1 Box =100 Units | COMPLY pending confirmation of latex and DEHP free | NO TEST REPORTS
42142609-00024 | 118 | Syringe, insulin, plastic, latex free, non-pyrogenic and DEPH free. Type 4 (with a fixed needle). Nominal capacity: 1 ml, 100 units per ml, must have clear graduation markings that are not easily removed or rubbed off. Needle: 29G (0,33 mm) x 12,5 - 12,7 mm Single use, sterile path To comply with the latest issue of SANS 1166 Packed in units of 10 syringes in a plastic packet 10 Packets in a box. UNIT: Box | 1 | CENTRAL MEDICAL PTY LTD | 6.98 | LYSA | INSL102913 | Zhejian INI Medical Devices Co Ltd | China | 21 | 80 | China yuan | 20 boxes of 10 packets of 10 units | yes | Box of 10 packets of 10 unit | COMPLY pending confirmation of latex and DEHP free | COMPLY
42142609-00024 | 118 | Syringe, insulin, plastic, latex free, non-pyrogenic and DEPH free. Type 4 (with a fixed needle). Nominal capacity: 1 ml, 100 units per ml, must have clear graduation markings that are not easily removed or rubbed off. Needle: 29G (0,33 mm) x 12,5 - 12,7 mm Single use, sterile path To comply with the latest issue of SANS 1166 Packed in units of 10 syringes in a plastic packet 10 Packets in a box. UNIT: Box |  | GIBELA TRADE AND INVEST 1001 CC | 8.7 | TARA | SYR034 | Shandong Qiaopai Group Co., Ltd | CHINA | 21 | 85 | USD | 10 | YES | box 100 | No sample submitted for evaluation | NO TEST REPORTS
42142609-00024 | 118 | Syringe, insulin, plastic, latex free, non-pyrogenic and DEPH free. Type 4 (with a fixed needle). Nominal capacity: 1 ml, 100 units per ml, must have clear graduation markings that are not easily removed or rubbed off. Needle: 29G (0,33 mm) x 12,5 - 12,7 mm Single use, sterile path To comply with the latest issue of SANS 1166 Packed in units of 10 syringes in a plastic packet 10 Packets in a box. UNIT: Box |  | LECHOBA MEDICAL TECHNOLOGIES (PTY) LTD | 8.92 | CHIRANA | CHINS0129PB | CHIRANA INJECTA A/S | SLOVAK REPUBLIC | 21 DAYS | 75% | EURO | 12000 | YES | 100/BOX (10 PCS X10 PACKS) | No sample submitted for evaluation | NO TEST REPORTS
42142609-00025 | 119 | Syringe, insulin, plastic, latex free, non-pyrogenic and DEPH free. Type 4 (with a fixed needle). Nominal capacity: 1 ml, 100 units per 1ml, must have clear graduation markings that are not easily removed or rubbed off. Needle: 29G (0,33 mm) x 12,5 - 12,7 mm Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 1166 Packed in box of 100. UNIT: Box |  | Isigidi Trading 95 cc | 0.64 | Isigidi | SY53 | Jiangsu Jichun Medical Devices Co. Ltd. | China | 14 days | 85% | USD | 30boxes | Yes | 100pcs/box | COMPLY pending confirmation of latex and DEHP free | COMPLY
42142609-00025 | 119 | Syringe, insulin, plastic, latex free, non-pyrogenic and DEPH free. Type 4 (with a fixed needle). Nominal capacity: 1 ml, 100 units per 1ml, must have clear graduation markings that are not easily removed or rubbed off. Needle: 29G (0,33 mm) x 12,5 - 12,7 mm Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 1166 Packed in box of 100. UNIT: Box | N/A | PROMED TECHNOLOGIES (PTY) LTD | 68 | U-LIFE | 609025 | JIAXING | CHINA | 14 DAYS | 85% | USD | 1 BOX | YES | BOX OF 100 | COMPLY pending confirmation of latex and DEHP free | COMPLY | Recommended | 94.38
42142609-00025 | 119 | Syringe, insulin, plastic, latex free, non-pyrogenic and DEPH free. Type 4 (with a fixed needle). Nominal capacity: 1 ml, 100 units per 1ml, must have clear graduation markings that are not easily removed or rubbed off. Needle: 29G (0,33 mm) x 12,5 - 12,7 mm Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 1166 Packed in box of 100. UNIT: Box | 1 | CENTRAL MEDICAL PTY LTD | 0.69 | LYSA | INSL102913 | Zhejian INI Medical Devices Co Ltd | China | 21 | 80 | China yuan | 20 boxes of 100 units | yes | Box of 100 units | Offered a length of 13mm | NO TEST REPORTS
42142609-00025 | 119 | Syringe, insulin, plastic, latex free, non-pyrogenic and DEPH free. Type 4 (with a fixed needle). Nominal capacity: 1 ml, 100 units per 1ml, must have clear graduation markings that are not easily removed or rubbed off. Needle: 29G (0,33 mm) x 12,5 - 12,7 mm Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 1166 Packed in box of 100. UNIT: Box |  | Chemical World cc | 0.71 | SURG PLUS | SYRI29 | TONGDA | china | 3 | 85 | USD | 20000 | YES | 100 | COMPLY pending confirmation of latex free
42142609-00025 | 119 | Syringe, insulin, plastic, latex free, non-pyrogenic and DEPH free. Type 4 (with a fixed needle). Nominal capacity: 1 ml, 100 units per 1ml, must have clear graduation markings that are not easily removed or rubbed off. Needle: 29G (0,33 mm) x 12,5 - 12,7 mm Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 1166 Packed in box of 100. UNIT: Box | 1 | Vital Life Medical Distributors (PTY) LTD | 0.81 | Healthease | NHIS-1ML-29-0.5 | Neomedic | China | 14 days | 80% | US Dollar | 50 Boxes | No | Needle 0.33 x 13mm 100's per box | No sample submitted for evaluation | NO TEST REPORTS
42142609-00025 | 119 | Syringe, insulin, plastic, latex free, non-pyrogenic and DEPH free. Type 4 (with a fixed needle). Nominal capacity: 1 ml, 100 units per 1ml, must have clear graduation markings that are not easily removed or rubbed off. Needle: 29G (0,33 mm) x 12,5 - 12,7 mm Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 1166 Packed in box of 100. UNIT: Box | Basic Line offer | Becton Dickinson | 0.82 | BD Microfine Plus | 324891 | BECTON DICKINSON | USA | 7 Working Days | 76% | USD | 1 BOX | Yes | 1 Box =100 Units | COMPLY pending confirmation of latex and DEHP free | NO TEST REPORTS
42142609-00025 | 119 | Syringe, insulin, plastic, latex free, non-pyrogenic and DEPH free. Type 4 (with a fixed needle). Nominal capacity: 1 ml, 100 units per 1ml, must have clear graduation markings that are not easily removed or rubbed off. Needle: 29G (0,33 mm) x 12,5 - 12,7 mm Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 1166 Packed in box of 100. UNIT: Box |  | SUPRA HEALTHCARE JOHANNESBURG PTY LTD | 0.82 | REVITAL | IS0129B | Revital Health | KENYA | Initial 60 days, thereafter 21 | 84 | USD | 10 Boxes | YES |  | No sample submitted for evaluation | NO TEST REPORTS
42142609-00025 | 119 | Syringe, insulin, plastic, latex free, non-pyrogenic and DEPH free. Type 4 (with a fixed needle). Nominal capacity: 1 ml, 100 units per 1ml, must have clear graduation markings that are not easily removed or rubbed off. Needle: 29G (0,33 mm) x 12,5 - 12,7 mm Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 1166 Packed in box of 100. UNIT: Box |  | GIBELA TRADE AND INVEST 1001 CC | 0.87 | TARA | SYR035 | Shandong Qiaopai Group Co., Ltd | CHINA | 21 | 85 | USD | 10 | YES | box 100 | No sample submitted for evaluation | NO TEST REPORTS
42142609-00025 | 119 | Syringe, insulin, plastic, latex free, non-pyrogenic and DEPH free. Type 4 (with a fixed needle). Nominal capacity: 1 ml, 100 units per 1ml, must have clear graduation markings that are not easily removed or rubbed off. Needle: 29G (0,33 mm) x 12,5 - 12,7 mm Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 1166 Packed in box of 100. UNIT: Box | N/A | SAFARMEX (PTY) LTD | 0.9 | Saclin | To be confirmed | Zhejiang Mediunion Healthcare Group Co. Ltd | China | 3days | 80% | US$ | 3000 Boxes | Yes | Packed in a Box of 100 | No sample submitted for evaluation | NO TEST REPORT
42142609-00025 | 119 | Syringe, insulin, plastic, latex free, non-pyrogenic and DEPH free. Type 4 (with a fixed needle). Nominal capacity: 1 ml, 100 units per 1ml, must have clear graduation markings that are not easily removed or rubbed off. Needle: 29G (0,33 mm) x 12,5 - 12,7 mm Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 1166 Packed in box of 100. UNIT: Box |  | LECHOBA MEDICAL TECHNOLOGIES (PTY) LTD | 1 | CHIRANA | CHINS0129 | CHIRANA INJECTA A/S | SLOVAK REPUBLIC | 21 DAYS | 75% | EURO | 3000 | YES | 100/BOX | COMPLY pending confirmattion of latex and DEHP free | COMPLY
42142609-00025 | 119 | Syringe, insulin, plastic, latex free, non-pyrogenic and DEPH free. Type 4 (with a fixed needle). Nominal capacity: 1 ml, 100 units per 1ml, must have clear graduation markings that are not easily removed or rubbed off. Needle: 29G (0,33 mm) x 12,5 - 12,7 mm Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 1166 Packed in box of 100. UNIT: Box | 1 | Biologica Pharmaceuticals Pty Ltd | 1.83 | Sol- M | 1612912B | Sol- Millennium | China | 4-6 Weeks | 100% | USD | 1000 Boxes | 100% | N/A
42142609-00001 | 120 | Syringe, insulin, plastic, latex free, non-pyrogenic and DEPH free. Type 4 (with a fixed needle). Nominal capacity: 0,5 ml, 100 units per ml, must have clear graduation markings that are not easily removed or rubbed off. Needle: 30G x 8 mm Single use, sterile path To comply with the latest issue of SANS 1166 Packed in units of 10 syringes in a plastic packet 10 Packets in a box. UNIT: Box |  | Medi-Core Technologies | 0.65 | ULIFE | 42142609-00001 | JIAXING U-LIFE MEDICAL DEVICE TECHNOLOGY CO.,LTD | CHINA | 21 | 80 | USD | 1 box | YES | BOX | COMPLY pending confirmation of latex and DEHP free | NOT COMPLY
42142609-00001 | 120 | Syringe, insulin, plastic, latex free, non-pyrogenic and DEPH free. Type 4 (with a fixed needle). Nominal capacity: 0,5 ml, 100 units per ml, must have clear graduation markings that are not easily removed or rubbed off. Needle: 30G x 8 mm Single use, sterile path To comply with the latest issue of SANS 1166 Packed in units of 10 syringes in a plastic packet 10 Packets in a box. UNIT: Box | N/A | PROMED TECHNOLOGIES (PTY) LTD | 0.68 | U-LIFE | 609001 | JIAXING | CHINA | 14 DAYS | 85% | USD | 1 BOX | YES | BOX | COMPLY pending confirmation of latex and DEHP free | NO TEST REPORTS
42142609-00001 | 120 | Syringe, insulin, plastic, latex free, non-pyrogenic and DEPH free. Type 4 (with a fixed needle). Nominal capacity: 0,5 ml, 100 units per ml, must have clear graduation markings that are not easily removed or rubbed off. Needle: 30G x 8 mm Single use, sterile path To comply with the latest issue of SANS 1166 Packed in units of 10 syringes in a plastic packet 10 Packets in a box. UNIT: Box | N/A | SAFARMEX (PTY) LTD | 0.93 | Saclin | To be confirmed | Zhejiang Mediunion Healthcare Group Co. Ltd | China | 3days | 80% | US$ | 3000 Boxes | Yes | 10 Packets in a Box | No sample submitted for evaluation | NO TEST REPORT
42142609-00001 | 120 | Syringe, insulin, plastic, latex free, non-pyrogenic and DEPH free. Type 4 (with a fixed needle). Nominal capacity: 0,5 ml, 100 units per ml, must have clear graduation markings that are not easily removed or rubbed off. Needle: 30G x 8 mm Single use, sterile path To comply with the latest issue of SANS 1166 Packed in units of 10 syringes in a plastic packet 10 Packets in a box. UNIT: Box | PKT OF 10 | MSIMED CC | 0.95 | KANG KANG | INS1 | ZHENJIANG | CHINA | 25 | 20% | USD | 20000PCS | YES | BOX | No sample submitted for evaluation | NO TEST REPORTS
42142609-00001 | 120 | Syringe, insulin, plastic, latex free, non-pyrogenic and DEPH free. Type 4 (with a fixed needle). Nominal capacity: 0,5 ml, 100 units per ml, must have clear graduation markings that are not easily removed or rubbed off. Needle: 30G x 8 mm Single use, sterile path To comply with the latest issue of SANS 1166 Packed in units of 10 syringes in a plastic packet 10 Packets in a box. UNIT: Box |  | SUPRA HEALTHCARE JOHANNESBURG PTY LTD | 0.97 | REVITAL | IS0530P | Revital Health | KENYA | Initial 60 days, thereafter 21 | 84 | USD | 10 Boxes | YES |  | No sample submitted for evaluation | NO TEST REPORTS
42142609-00001 | 120 | Syringe, insulin, plastic, latex free, non-pyrogenic and DEPH free. Type 4 (with a fixed needle). Nominal capacity: 0,5 ml, 100 units per ml, must have clear graduation markings that are not easily removed or rubbed off. Needle: 30G x 8 mm Single use, sterile path To comply with the latest issue of SANS 1166 Packed in units of 10 syringes in a plastic packet 10 Packets in a box. UNIT: Box | 1 | B. BRAUN MEDICAL (PTY) LTD | 1.57 | Omnican 50 0.5ml INS. U100 30G x 8mm IND | 9151117S | B. Braun | Germany | 42 | 50% | Euro | 10 | Yes | Box of 10 | offered an individually peel apart package instead packed in units of 10 syringes in a plastic packet | NOT COMPLY
42142609-00001 | 120 | Syringe, insulin, plastic, latex free, non-pyrogenic and DEPH free. Type 4 (with a fixed needle). Nominal capacity: 0,5 ml, 100 units per ml, must have clear graduation markings that are not easily removed or rubbed off. Needle: 30G x 8 mm Single use, sterile path To comply with the latest issue of SANS 1166 Packed in units of 10 syringes in a plastic packet 10 Packets in a box. UNIT: Box | 1 | Biologica Pharmaceuticals Pty Ltd | 1.83 | Sol- M | 1653051B | Sol- Millennium | China | 4-6 Weeks | 100% | USD | 1000 Boxes | 100% | N/A |  | NO TEST REPORTS
42142609-00001 | 120 | Syringe, insulin, plastic, latex free, non-pyrogenic and DEPH free. Type 4 (with a fixed needle). Nominal capacity: 0,5 ml, 100 units per ml, must have clear graduation markings that are not easily removed or rubbed off. Needle: 30G x 8 mm Single use, sterile path To comply with the latest issue of SANS 1166 Packed in units of 10 syringes in a plastic packet 10 Packets in a box. UNIT: Box |  | Isigidi Trading 95 cc | 63.25 | Isigidi | SY43 | Jiangsu Jichun Medical Devices Co. Ltd. | China | 14 days | 85% | USD | 30boxes | Yes | 100pcs/box | COMPLY pending confirmation of latex and DEHP free | COMPLY | Recommended | 91
42142609-00001 | 120 | Syringe, insulin, plastic, latex free, non-pyrogenic and DEPH free. Type 4 (with a fixed needle). Nominal capacity: 0,5 ml, 100 units per ml, must have clear graduation markings that are not easily removed or rubbed off. Needle: 30G x 8 mm Single use, sterile path To comply with the latest issue of SANS 1166 Packed in units of 10 syringes in a plastic packet 10 Packets in a box. UNIT: Box | Basic Line offer | Becton Dickinson | 6.4 | BD Microfine Plus | 320930 | BECTON DICKINSON | USA | 7 Working Days | 76% | USD | 1 BOX | Yes | 1 Box =100 Units | COMPLY pending confirmation of latex and DEHP free | NO TEST REPORTS
42142609-00001 | 120 | Syringe, insulin, plastic, latex free, non-pyrogenic and DEPH free. Type 4 (with a fixed needle). Nominal capacity: 0,5 ml, 100 units per ml, must have clear graduation markings that are not easily removed or rubbed off. Needle: 30G x 8 mm Single use, sterile path To comply with the latest issue of SANS 1166 Packed in units of 10 syringes in a plastic packet 10 Packets in a box. UNIT: Box | 1 | CENTRAL MEDICAL PTY LTD | 7.11 | LYSA | INSL053008 | Zhejian INI Medical Devices Co Ltd | China | 21 | 80 | China yuan | 20 boxes of 10 packets of 10 units | yes | Box of 10 packets of 10 unit | COMPLY pending confirmation of latex and DEHP free | NO TEST REPORTS
42142609-00001 | 120 | Syringe, insulin, plastic, latex free, non-pyrogenic and DEPH free. Type 4 (with a fixed needle). Nominal capacity: 0,5 ml, 100 units per ml, must have clear graduation markings that are not easily removed or rubbed off. Needle: 30G x 8 mm Single use, sterile path To comply with the latest issue of SANS 1166 Packed in units of 10 syringes in a plastic packet 10 Packets in a box. UNIT: Box |  | GIBELA TRADE AND INVEST 1001 CC | 8.56 | TARA | SYR036 | Shandong Qiaopai Group Co., Ltd | CHINA | 21 | 85 | USD | 10 | YES | box 100 | No sample submitted for evaluation | NO TEST REPORTS
42142609-00001 | 120 | Syringe, insulin, plastic, latex free, non-pyrogenic and DEPH free. Type 4 (with a fixed needle). Nominal capacity: 0,5 ml, 100 units per ml, must have clear graduation markings that are not easily removed or rubbed off. Needle: 30G x 8 mm Single use, sterile path To comply with the latest issue of SANS 1166 Packed in units of 10 syringes in a plastic packet 10 Packets in a box. UNIT: Box |  | LECHOBA MEDICAL TECHNOLOGIES (PTY) LTD | 8.83 | CHIRANA | CHINS005308PB | CHIRANA INJECTA A/S | SLOVAK REPUBLIC | 21 DAYS | 75% | EURO | 3000 | YES | 100/BOX (10 PCS X10 PACKS) | COMPLY pending confirmattion of latex and DEHP free | NOT COMPLY
42142609-00002 | 121 | Syringe, insulin, plastic, latex free, non-pyrogenic and DEPH free. Type 4 (with a fixed needle). Nominal capacity: 0,5 ml, 100 units per 1 ml, must have clear graduation markings that are not easily removed or rubbed off. Needle: 30G x 8 mm Sterile, Single use, Individually peel packed Packed in box of 100. UNIT: Box |  | Isigidi Trading 95 cc | 63.25 | Isigidi | SY42 | Jiangsu Jichun Medical Devices Co. Ltd. | China | 14 days | 85% | USD | 30boxes | Yes | 100pcs/box | COMPLY pending confirmation of latex and DEHP free | COMPLY | Recommended | 91
42142609-00002 | 121 | Syringe, insulin, plastic, latex free, non-pyrogenic and DEPH free. Type 4 (with a fixed needle). Nominal capacity: 0,5 ml, 100 units per 1 ml, must have clear graduation markings that are not easily removed or rubbed off. Needle: 30G x 8 mm Sterile, Single use, Individually peel packed Packed in box of 100. UNIT: Box |  | Medi-Core Technologies | 0.67 | ULIFE | 42142609-00002 | JIAXING U-LIFE MEDICAL DEVICE TECHNOLOGY CO.,LTD | CHINA | 21 | 80 | USD | 1 box | YES | BOX | No sample submitted for evaluation | NO TEST REPORTS
42142609-00002 | 121 | Syringe, insulin, plastic, latex free, non-pyrogenic and DEPH free. Type 4 (with a fixed needle). Nominal capacity: 0,5 ml, 100 units per 1 ml, must have clear graduation markings that are not easily removed or rubbed off. Needle: 30G x 8 mm Sterile, Single use, Individually peel packed Packed in box of 100. UNIT: Box | N/A | PROMED TECHNOLOGIES (PTY) LTD | 0.68 | U-LIFE | 609002 | JIAXING | CHINA | 14 DAYS | 85% | USD | 1 BOX | YES | BOX | No samples submitted for evaluation | NO TEST REPORTS
42142609-00002 | 121 | Syringe, insulin, plastic, latex free, non-pyrogenic and DEPH free. Type 4 (with a fixed needle). Nominal capacity: 0,5 ml, 100 units per 1 ml, must have clear graduation markings that are not easily removed or rubbed off. Needle: 30G x 8 mm Sterile, Single use, Individually peel packed Packed in box of 100. UNIT: Box | 1 | CENTRAL MEDICAL PTY LTD | 0.7 | LYSA | INSL053008 | Zhejian INI Medical Devices Co Ltd | China | 21 | 80 | China yuan | 20 boxes of 100 units | yes | Box of 100 units | No sample submitted for evaluation | NO TEST REPORTS
42142609-00002 | 121 | Syringe, insulin, plastic, latex free, non-pyrogenic and DEPH free. Type 4 (with a fixed needle). Nominal capacity: 0,5 ml, 100 units per 1 ml, must have clear graduation markings that are not easily removed or rubbed off. Needle: 30G x 8 mm Sterile, Single use, Individually peel packed Packed in box of 100. UNIT: Box | Basic Line offer | Becton Dickinson | 0.82 | BD Microfine Plus | 324893 | BECTON DICKINSON | USA | 7 Working Days | 76% | USD | 1 BOX | Yes | 1 Box =100 Units | COMPLY pending confirmation of latex and DEHP free | NO TEST REPORT
42142609-00002 | 121 | Syringe, insulin, plastic, latex free, non-pyrogenic and DEPH free. Type 4 (with a fixed needle). Nominal capacity: 0,5 ml, 100 units per 1 ml, must have clear graduation markings that are not easily removed or rubbed off. Needle: 30G x 8 mm Sterile, Single use, Individually peel packed Packed in box of 100. UNIT: Box |  | GIBELA TRADE AND INVEST 1001 CC | 0.84 | TARA | SYR037 | Shandong Qiaopai Group Co., Ltd | CHINA | 21 | 85 | USD | 10 | YES | box 100 | No sample submitted for evaluation | NO TEST REPORTS
42142609-00002 | 121 | Syringe, insulin, plastic, latex free, non-pyrogenic and DEPH free. Type 4 (with a fixed needle). Nominal capacity: 0,5 ml, 100 units per 1 ml, must have clear graduation markings that are not easily removed or rubbed off. Needle: 30G x 8 mm Sterile, Single use, Individually peel packed Packed in box of 100. UNIT: Box | N/A | SAFARMEX (PTY) LTD | 0.9 | Saclin | To be confirmed | Zhejiang Mediunion Healthcare Group Co. Ltd | China | 3days | 80% | US$ | 3000 Boxes | Yes | Packed in a Box of 100 | No sample submitted for evaluation | NO TEST REPORT
42142609-00002 | 121 | Syringe, insulin, plastic, latex free, non-pyrogenic and DEPH free. Type 4 (with a fixed needle). Nominal capacity: 0,5 ml, 100 units per 1 ml, must have clear graduation markings that are not easily removed or rubbed off. Needle: 30G x 8 mm Sterile, Single use, Individually peel packed Packed in box of 100. UNIT: Box | PKT OF 10 | MSIMED CC | 0.95 | KANG KANG | INS2 | ZHENJIANG | CHINA | 25 | 20% | USD | 20000PCS | YES | BOX | No sample submitted for evaluation | NO TEST REPORTS
42142609-00002 | 121 | Syringe, insulin, plastic, latex free, non-pyrogenic and DEPH free. Type 4 (with a fixed needle). Nominal capacity: 0,5 ml, 100 units per 1 ml, must have clear graduation markings that are not easily removed or rubbed off. Needle: 30G x 8 mm Sterile, Single use, Individually peel packed Packed in box of 100. UNIT: Box |  | SUPRA HEALTHCARE JOHANNESBURG PTY LTD | 0.97 | REVITAL | IS0530B | Revital Health | KENYA | Initial 60 days, thereafter 21 | 84 | USD | 10 Boxes | YES |  | No sample submitted for evaluation | NO TEST REPORTS
42142609-00002 | 121 | Syringe, insulin, plastic, latex free, non-pyrogenic and DEPH free. Type 4 (with a fixed needle). Nominal capacity: 0,5 ml, 100 units per 1 ml, must have clear graduation markings that are not easily removed or rubbed off. Needle: 30G x 8 mm Sterile, Single use, Individually peel packed Packed in box of 100. UNIT: Box | Basic Line Offer | BIOCOR HOSPITAL SUPPLIES PTY LTD | 1.04 | Medline | N/A | Yangzhong Medline Industry Co.Ltd | China | 21 | 85 | USD | 5 Boxes of 100 | Yes | N/A | Offered a 12mm instead of 8mm | NO TEST REPORTS
42142609-00002 | 121 | Syringe, insulin, plastic, latex free, non-pyrogenic and DEPH free. Type 4 (with a fixed needle). Nominal capacity: 0,5 ml, 100 units per 1 ml, must have clear graduation markings that are not easily removed or rubbed off. Needle: 30G x 8 mm Sterile, Single use, Individually peel packed Packed in box of 100. UNIT: Box | 1 | Biologica Pharmaceuticals Pty Ltd | 1.83 | Sol- M | 1653051B | Sol- Millennium | China | 4-6 Weeks | 100% | USD | 1000 Boxes | 100% | N/A
42142609-00002 | 121 | Syringe, insulin, plastic, latex free, non-pyrogenic and DEPH free. Type 4 (with a fixed needle). Nominal capacity: 0,5 ml, 100 units per 1 ml, must have clear graduation markings that are not easily removed or rubbed off. Needle: 30G x 8 mm Sterile, Single use, Individually peel packed Packed in box of 100. UNIT: Box |  | JUMLA MEDICAL SUPPLIES (PTY) LTD | 80.2 | SYR-CARE | SYR-CARE-0.5M30 | SUZHOU ACMED | CHINA | 21 DAYS | 72% | US DOLLAR | 10 BOXES | NO | 12MM NEEDLE | Offered 12mm instead of 8mm | NO TEST REPORTS
42142609-00003 | 122 | Syringe, insulin, plastic, latex free, non-pyrogenic and DEPH free. Type 4 (with a fixed needle). Nominal capacity: 1 ml, 100 units per 1ml, must have clear graduation markings that are not easily removed or rubbed off. Needle: 30G x 8 mm Single use, sterile path To comply with the latest issue of SANS 1166 Packed in units of 10 syringes in a plastic packet 10 Packets in a box. UNIT: Box |  | Medi-Core Technologies | 0.65 | ULIFE | 42142609-00003 | JIAXING U-LIFE MEDICAL DEVICE TECHNOLOGY CO.,LTD | CHINA | 21 | 80 | USD | 1 box | YES | BOX | COMPLY pending confirmation of latex and DEHP free | NOT COMPLY
42142609-00003 | 122 | Syringe, insulin, plastic, latex free, non-pyrogenic and DEPH free. Type 4 (with a fixed needle). Nominal capacity: 1 ml, 100 units per 1ml, must have clear graduation markings that are not easily removed or rubbed off. Needle: 30G x 8 mm Single use, sterile path To comply with the latest issue of SANS 1166 Packed in units of 10 syringes in a plastic packet 10 Packets in a box. UNIT: Box | N/A | PROMED TECHNOLOGIES (PTY) LTD | 0.68 | U-LIFE | 609003 | JIAXING | CHINA | 14 DAYS | 85% | USD | 1 BOX | YES | BOX | COMPLY pending confirmation of latex and DEHP free | NO TEST REPORTS
42142609-00003 | 122 | Syringe, insulin, plastic, latex free, non-pyrogenic and DEPH free. Type 4 (with a fixed needle). Nominal capacity: 1 ml, 100 units per 1ml, must have clear graduation markings that are not easily removed or rubbed off. Needle: 30G x 8 mm Single use, sterile path To comply with the latest issue of SANS 1166 Packed in units of 10 syringes in a plastic packet 10 Packets in a box. UNIT: Box | N/A | SAFARMEX (PTY) LTD | 0.93 | Saclin | To be confirmed | Zhejiang Mediunion Healthcare Group Co. Ltd | China | 3days | 80% | US$ | 3000 Boxes | Yes | 10 Packets in a Box | No sample submitted for evaluation | NO TEST REPORT
42142609-00003 | 122 | Syringe, insulin, plastic, latex free, non-pyrogenic and DEPH free. Type 4 (with a fixed needle). Nominal capacity: 1 ml, 100 units per 1ml, must have clear graduation markings that are not easily removed or rubbed off. Needle: 30G x 8 mm Single use, sterile path To comply with the latest issue of SANS 1166 Packed in units of 10 syringes in a plastic packet 10 Packets in a box. UNIT: Box | PKT OF 10 | MSIMED CC | 0.95 | KANG KANG | INS3 | ZHENJIANG | CHINA | 25 | 20% | USD | 20000PCS | YES | BOX | No sample submitted for evaluation | NO TEST REPORTS
42142609-00003 | 122 | Syringe, insulin, plastic, latex free, non-pyrogenic and DEPH free. Type 4 (with a fixed needle). Nominal capacity: 1 ml, 100 units per 1ml, must have clear graduation markings that are not easily removed or rubbed off. Needle: 30G x 8 mm Single use, sterile path To comply with the latest issue of SANS 1166 Packed in units of 10 syringes in a plastic packet 10 Packets in a box. UNIT: Box |  | SUPRA HEALTHCARE JOHANNESBURG PTY LTD | 0.97 | REVITAL | IS130P | Revital Health | KENYA | Initial 60 days, thereafter 21 | 84 | USD | 10 Boxes | YES |  | No sample submitted for evaluation | NO TEST REPORTS
42142609-00003 | 122 | Syringe, insulin, plastic, latex free, non-pyrogenic and DEPH free. Type 4 (with a fixed needle). Nominal capacity: 1 ml, 100 units per 1ml, must have clear graduation markings that are not easily removed or rubbed off. Needle: 30G x 8 mm Single use, sterile path To comply with the latest issue of SANS 1166 Packed in units of 10 syringes in a plastic packet 10 Packets in a box. UNIT: Box |  | Isigidi Trading 95 cc | 63.25 | Isigidi | SY49 | Jiangsu Jichun Medical Devices Co. Ltd. | China | 14 days | 85% | USD | 30boxes | Yes | 100pcs/box | COMPLY pending confirmation of latex and DEHP free | COMPLY | Recommended | 91
42142609-00003 | 122 | Syringe, insulin, plastic, latex free, non-pyrogenic and DEPH free. Type 4 (with a fixed needle). Nominal capacity: 1 ml, 100 units per 1ml, must have clear graduation markings that are not easily removed or rubbed off. Needle: 30G x 8 mm Single use, sterile path To comply with the latest issue of SANS 1166 Packed in units of 10 syringes in a plastic packet 10 Packets in a box. UNIT: Box | Basic Line offer | Becton Dickinson | 6.4 | BD Microfine Plus | 320929 | BECTON DICKINSON | USA | 7 Working Days | 76% | USD | 1 BOX | Yes | 1 Box =100 Units | COMPLY pending confirmation of latex and DEHP free | NO TEST REPORTS
42142609-00003 | 122 | Syringe, insulin, plastic, latex free, non-pyrogenic and DEPH free. Type 4 (with a fixed needle). Nominal capacity: 1 ml, 100 units per 1ml, must have clear graduation markings that are not easily removed or rubbed off. Needle: 30G x 8 mm Single use, sterile path To comply with the latest issue of SANS 1166 Packed in units of 10 syringes in a plastic packet 10 Packets in a box. UNIT: Box |  | GIBELA TRADE AND INVEST 1001 CC | 8.22 | TARA | SYR038 | Shandong Qiaopai Group Co., Ltd | CHINA | 21 | 85 | USD | 10 | YES | box100 | No sample submitted for evaluation | NO TEST REPORTS
42142609-00003 | 122 | Syringe, insulin, plastic, latex free, non-pyrogenic and DEPH free. Type 4 (with a fixed needle). Nominal capacity: 1 ml, 100 units per 1ml, must have clear graduation markings that are not easily removed or rubbed off. Needle: 30G x 8 mm Single use, sterile path To comply with the latest issue of SANS 1166 Packed in units of 10 syringes in a plastic packet 10 Packets in a box. UNIT: Box |  | LECHOBA MEDICAL TECHNOLOGIES (PTY) LTD | 8.92 | CHIRANA | CHINS01308PB | CHIRANA INJECTA A/S | SLOVAK REPUBLIC | 21 DAYS | 75% | EURO | 3000 | YES | 100/BOX (10 PCS X10 PACKS) | COMPLY pending confirmattion of latex and DEHP free | NOT COMPLY
42142609-00004 | 123 | Syringe, insulin, plastic, latex free, non-pyrogenic and DEPH free. Type 4 (with a fixed needle). Nominal capacity: 1 ml, 100 units per 1ml, must have clear graduation markings that are not easily removed or rubbed off. Needle: 30G x 8 mm Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 1166 Packed in box of 100. UNIT: Box |  | Isigidi Trading 95 cc | 63.25 | Isigidi | SY50 | Jiangsu Jichun Medical Devices Co. Ltd. | China | 14 days | 85% | USD | 30boxes | Yes | 100pcs/box | COMPLY pending confirmation of latex and DEHP free | COMPLY | Recommended | 91
42142609-00004 | 123 | Syringe, insulin, plastic, latex free, non-pyrogenic and DEPH free. Type 4 (with a fixed needle). Nominal capacity: 1 ml, 100 units per 1ml, must have clear graduation markings that are not easily removed or rubbed off. Needle: 30G x 8 mm Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 1166 Packed in box of 100. UNIT: Box |  | Medi-Core Technologies | 0.67 | ULIFE | 42142609-00004 | JIAXING U-LIFE MEDICAL DEVICE TECHNOLOGY CO.,LTD | CHINA | 21 | 80 | USD | 1 box | YES | BOX | No sample submitted for evaluation | NOT COMPLY
42142609-00004 | 123 | Syringe, insulin, plastic, latex free, non-pyrogenic and DEPH free. Type 4 (with a fixed needle). Nominal capacity: 1 ml, 100 units per 1ml, must have clear graduation markings that are not easily removed or rubbed off. Needle: 30G x 8 mm Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 1166 Packed in box of 100. UNIT: Box | N/A | PROMED TECHNOLOGIES (PTY) LTD | 0.68 | U-LIFE | 609004 | JIAXING | CHINA | 14 DAYS | 85% | USD | 1 BOX | YES | BOX | No samples submitted for evaluation | NO TEST REPORTS
42142609-00004 | 123 | Syringe, insulin, plastic, latex free, non-pyrogenic and DEPH free. Type 4 (with a fixed needle). Nominal capacity: 1 ml, 100 units per 1ml, must have clear graduation markings that are not easily removed or rubbed off. Needle: 30G x 8 mm Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 1166 Packed in box of 100. UNIT: Box |  | GIBELA TRADE AND INVEST 1001 CC | 0.81 | TARA | SYR039 | Shandong Qiaopai Group Co., Ltd | CHINA | 21 | 85 | USD | 10 | YES | box 100 | No sample submitted for evaluation | NO TEST REPORTS
42142609-00004 | 123 | Syringe, insulin, plastic, latex free, non-pyrogenic and DEPH free. Type 4 (with a fixed needle). Nominal capacity: 1 ml, 100 units per 1ml, must have clear graduation markings that are not easily removed or rubbed off. Needle: 30G x 8 mm Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 1166 Packed in box of 100. UNIT: Box | Basic Line offer | Becton Dickinson | 0.82 | BD Microfine Plus | 324899 | BECTON DICKINSON | USA | 7 Working Days | 76% | USD | 1 BOX | Yes | 1 Box =100 Units | COMPLY pending confirmation of latex and DEHP free | NO TEST REPORTS
42142609-00004 | 123 | Syringe, insulin, plastic, latex free, non-pyrogenic and DEPH free. Type 4 (with a fixed needle). Nominal capacity: 1 ml, 100 units per 1ml, must have clear graduation markings that are not easily removed or rubbed off. Needle: 30G x 8 mm Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 1166 Packed in box of 100. UNIT: Box | N/A | SAFARMEX (PTY) LTD | 0.9 | Saclin | To be confirmed | Zhejiang Mediunion Healthcare Group Co. Ltd | China | 3days | 80% | US$ | 3000 Boxes | Yes | Packed in a Box of 100 | No sample submitted for evaluation | NO TEST REPORT
42142609-00004 | 123 | Syringe, insulin, plastic, latex free, non-pyrogenic and DEPH free. Type 4 (with a fixed needle). Nominal capacity: 1 ml, 100 units per 1ml, must have clear graduation markings that are not easily removed or rubbed off. Needle: 30G x 8 mm Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 1166 Packed in box of 100. UNIT: Box | PKT OF 10 | MSIMED CC | 0.95 | KANG KANG | INS4 | ZHENJIANG | CHINA | 25 | 20% | USD | 20000PCS | YES | BOX | No sample submitted for evaluation | NO TEST REPORTS
42142609-00004 | 123 | Syringe, insulin, plastic, latex free, non-pyrogenic and DEPH free. Type 4 (with a fixed needle). Nominal capacity: 1 ml, 100 units per 1ml, must have clear graduation markings that are not easily removed or rubbed off. Needle: 30G x 8 mm Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 1166 Packed in box of 100. UNIT: Box |  | SUPRA HEALTHCARE JOHANNESBURG PTY LTD | 0.97 | REVITAL | IS130B | Revital Health | KENYA | Initial 60 days, thereafter 21 | 84 | USD | 10 Boxes | YES |  | No sample submitted for evaluation | NO TEST REPORTS
42142609-00004 | 123 | Syringe, insulin, plastic, latex free, non-pyrogenic and DEPH free. Type 4 (with a fixed needle). Nominal capacity: 1 ml, 100 units per 1ml, must have clear graduation markings that are not easily removed or rubbed off. Needle: 30G x 8 mm Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 1166 Packed in box of 100. UNIT: Box |  | JUMLA MEDICAL SUPPLIES (PTY) LTD | 74.81 | SYR-CARE | SYR-CARE-1ML31G | SUZHOU ACMED | CHINA | 21 DAYS | 72% | US DOLLAR | 10 BOXES | NO | 12MM NEEDLE | Offered 12mm instead of 8mm | NO TEST REPORTS
42142609-00005 | 124 | Syringe, insulin, plastic, latex free, non-pyrogenic and DEPH free. Type 4 (with a fixed needle). Nominal capacity: 0,5 ml, 100 units per ml, must have clear graduation markings that are not easily removed or rubbed off. Needle: 31G x 6 mm Single use, sterile path To comply with the latest issue of SANS 1166 Packed in units of 10 syringes in a plastic packet 10 Packets in a box. UNIT: Box |  | Medi-Core Technologies | 0.65 | ULIFE | 42142609-00005 | JIAXING U-LIFE MEDICAL DEVICE TECHNOLOGY CO.,LTD | CHINA | 21 | 80 | USD | 1 box | YES | BOX | No sample submitted for evaluation | NO TEST REPORTS | Re-tender
42142609-00005 | 124 | Syringe, insulin, plastic, latex free, non-pyrogenic and DEPH free. Type 4 (with a fixed needle). Nominal capacity: 0,5 ml, 100 units per ml, must have clear graduation markings that are not easily removed or rubbed off. Needle: 31G x 6 mm Single use, sterile path To comply with the latest issue of SANS 1166 Packed in units of 10 syringes in a plastic packet 10 Packets in a box. UNIT: Box | N/A | PROMED TECHNOLOGIES (PTY) LTD | 0.68 | U-LIFE | 609005 | JIAXING | CHINA | 14 DAYS | 85% | USD | 1 BOX | YES | BOX | No samples submitted for evaluation | NO TEST REPORTS | Re-tender
42142609-00005 | 124 | Syringe, insulin, plastic, latex free, non-pyrogenic and DEPH free. Type 4 (with a fixed needle). Nominal capacity: 0,5 ml, 100 units per ml, must have clear graduation markings that are not easily removed or rubbed off. Needle: 31G x 6 mm Single use, sterile path To comply with the latest issue of SANS 1166 Packed in units of 10 syringes in a plastic packet 10 Packets in a box. UNIT: Box | N/A | SAFARMEX (PTY) LTD | 0.93 | Saclin | To be confirmed | Zhejiang Mediunion Healthcare Group Co. Ltd | China | 3days | 80% | US$ | 3000 Boxes | Yes | 10 Pasckets in a Box | No sample submitted for evaluation | NO TEST REPORT | Re-tender
42142609-00005 | 124 | Syringe, insulin, plastic, latex free, non-pyrogenic and DEPH free. Type 4 (with a fixed needle). Nominal capacity: 0,5 ml, 100 units per ml, must have clear graduation markings that are not easily removed or rubbed off. Needle: 31G x 6 mm Single use, sterile path To comply with the latest issue of SANS 1166 Packed in units of 10 syringes in a plastic packet 10 Packets in a box. UNIT: Box |  | SUPRA HEALTHCARE JOHANNESBURG PTY LTD | 0.97 | REVITAL | IS0531P | Revital Health | KENYA | Initial 60 days, thereafter 21 | 84 | USD | 10 Boxes | YES |  | No sample submitted for evaluation | NO TEST REPORTS | Re-tender
42142609-00005 | 124 | Syringe, insulin, plastic, latex free, non-pyrogenic and DEPH free. Type 4 (with a fixed needle). Nominal capacity: 0,5 ml, 100 units per ml, must have clear graduation markings that are not easily removed or rubbed off. Needle: 31G x 6 mm Single use, sterile path To comply with the latest issue of SANS 1166 Packed in units of 10 syringes in a plastic packet 10 Packets in a box. UNIT: Box |  | Isigidi Trading 95 cc | 6.33 | Isigidi | SY41 | Jiangsu Jichun Medical Devices Co. Ltd. | China | 14 days | 85% | USD | 30boxes | Yes | 100pcs/box | COMPLY pending confirmation of latex and DEHP free | NOT COMPLY (NEEDLE LENGTH IS 5.2MM) | Re-tender
42142609-00005 | 124 | Syringe, insulin, plastic, latex free, non-pyrogenic and DEPH free. Type 4 (with a fixed needle). Nominal capacity: 0,5 ml, 100 units per ml, must have clear graduation markings that are not easily removed or rubbed off. Needle: 31G x 6 mm Single use, sterile path To comply with the latest issue of SANS 1166 Packed in units of 10 syringes in a plastic packet 10 Packets in a box. UNIT: Box | Basic Line offer | Becton Dickinson | 6.4 | BD Microfine Plus | 324904 | BECTON DICKINSON | USA | 7 Working Days | 76% | USD | 1 BOX | Yes | 1 Box =100 Units | COMPLY pending confirmation of latex and DEHP free | NO TEST REPORTS | Re-tender
42142609-00005 | 124 | Syringe, insulin, plastic, latex free, non-pyrogenic and DEPH free. Type 4 (with a fixed needle). Nominal capacity: 0,5 ml, 100 units per ml, must have clear graduation markings that are not easily removed or rubbed off. Needle: 31G x 6 mm Single use, sterile path To comply with the latest issue of SANS 1166 Packed in units of 10 syringes in a plastic packet 10 Packets in a box. UNIT: Box |  | GIBELA TRADE AND INVEST 1001 CC | 8.56 | TARA | SYR040 | Shandong Qiaopai Group Co., Ltd | CHINA | 21 | 85 | USD | 10 | YES | box 100 | No sample submitted for evaluation | NO TEST REPORTS | Re-tender
42142609-00006 | 125 | Syringe, insulin, plastic, latex free, non-pyrogenic and DEPH free. Type 4 (with a fixed needle). Nominal capacity: 0,5 ml, 100 units per 1 ml, must have clear graduation markings that are not easily removed or rubbed off. Needle: 31G x 6 mm Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 1166 Packed in box of 100. UNIT: Box |  | Isigidi Trading 95 cc | 0.64 | Isigidi | SY44 | Jiangsu Jichun Medical Devices Co. Ltd. | China | 14 days | 85% | USD | 30boxes | Yes | 100pcs/box | COMPLY pending confirmation of latex and DEHP free | NOT COMPLY (NEEDLE LENGTH IS 5.2MM) | Re-tender
42142609-00006 | 125 | Syringe, insulin, plastic, latex free, non-pyrogenic and DEPH free. Type 4 (with a fixed needle). Nominal capacity: 0,5 ml, 100 units per 1 ml, must have clear graduation markings that are not easily removed or rubbed off. Needle: 31G x 6 mm Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 1166 Packed in box of 100. UNIT: Box |  | Medi-Core Technologies | 0.67 | ULIFE | 42142609-00006 | JIAXING U-LIFE MEDICAL DEVICE TECHNOLOGY CO.,LTD | CHINA | 21 | 80 | USD | 1 box | YES | BOX | No sample submitted for evaluation | NO TEST REPORTS | Re-tender
42142609-00006 | 125 | Syringe, insulin, plastic, latex free, non-pyrogenic and DEPH free. Type 4 (with a fixed needle). Nominal capacity: 0,5 ml, 100 units per 1 ml, must have clear graduation markings that are not easily removed or rubbed off. Needle: 31G x 6 mm Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 1166 Packed in box of 100. UNIT: Box | N/A | PROMED TECHNOLOGIES (PTY) LTD | 0.68 | U-LIFE | 609006 | JIAXING | CHINA | 14 DAYS | 85% | USD | 1 BOX | YES | BOX | No samples submitted for evaluation | NO TEST REPORTS | Re-tender
42142609-00006 | 125 | Syringe, insulin, plastic, latex free, non-pyrogenic and DEPH free. Type 4 (with a fixed needle). Nominal capacity: 0,5 ml, 100 units per 1 ml, must have clear graduation markings that are not easily removed or rubbed off. Needle: 31G x 6 mm Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 1166 Packed in box of 100. UNIT: Box | 1 | CENTRAL MEDICAL PTY LTD | 0.77 | LYSA | INSL053106 | Zhejian INI Medical Devices Co Ltd | China | 21 | 80 | China yuan | 20 boxes of 100 units | yes | Box of 100 units | COMPLY pending confirmation of latex and DEHP free | NO TEST REPORTS | Re-tender
42142609-00006 | 125 | Syringe, insulin, plastic, latex free, non-pyrogenic and DEPH free. Type 4 (with a fixed needle). Nominal capacity: 0,5 ml, 100 units per 1 ml, must have clear graduation markings that are not easily removed or rubbed off. Needle: 31G x 6 mm Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 1166 Packed in box of 100. UNIT: Box | Basic Line offer | Becton Dickinson | 0.82 | BD Microfine Plus | 326677 | BECTON DICKINSON | USA | 7 Working Days | 76% | USD | 1 BOX | Yes | 1 Box =100 Units | COMPLY pending confirmation of latex and DEHP free | NO TEST REPORTS | Re-tender
42142609-00006 | 125 | Syringe, insulin, plastic, latex free, non-pyrogenic and DEPH free. Type 4 (with a fixed needle). Nominal capacity: 0,5 ml, 100 units per 1 ml, must have clear graduation markings that are not easily removed or rubbed off. Needle: 31G x 6 mm Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 1166 Packed in box of 100. UNIT: Box |  | GIBELA TRADE AND INVEST 1001 CC | 0.84 | TARA | SYR041 | Shandong Qiaopai Group Co., Ltd | CHINA | 21 | 85 | USD | 10 | YES | box 100 | No sample submitted for evaluation | NO TEST REPORTS | Re-tender
42142609-00006 | 125 | Syringe, insulin, plastic, latex free, non-pyrogenic and DEPH free. Type 4 (with a fixed needle). Nominal capacity: 0,5 ml, 100 units per 1 ml, must have clear graduation markings that are not easily removed or rubbed off. Needle: 31G x 6 mm Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 1166 Packed in box of 100. UNIT: Box | N/A | SAFARMEX (PTY) LTD | 0.9 | Saclin | To be confirmed | Zhejiang Mediunion Healthcare Group Co. Ltd | China | 3days | 80% | US$ | 3000 Boxes | Yes | Packed in a Box of 100 | No sample submitted for evaluation | NO TEST REPORT | Re-tender
42142609-00006 | 125 | Syringe, insulin, plastic, latex free, non-pyrogenic and DEPH free. Type 4 (with a fixed needle). Nominal capacity: 0,5 ml, 100 units per 1 ml, must have clear graduation markings that are not easily removed or rubbed off. Needle: 31G x 6 mm Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 1166 Packed in box of 100. UNIT: Box |  | SUPRA HEALTHCARE JOHANNESBURG PTY LTD | 0.97 | REVITAL | IS0531B | Revital Health | KENYA | Initial 60 days, thereafter 21 | 84 | USD | 10 Boxes | YES |  | No sample submitted for evaluation | NO TEST REPORTS | Re-tender
42142609-00007 | 126 | Syringe, insulin, plastic, latex free, non-pyrogenic and DEPH free. Type 4 (with a fixed needle). Nominal capacity: 1 ml, 100 units per 1ml, must have clear graduation markings that are not easily removed or rubbed off. Needle: 31G x 6 mm Single use, sterile path To comply with the latest issue of SANS 1166 Packed in units of 10 syringes in a plastic packet 10 Packets in a box. UNIT: Box |  | Medi-Core Technologies | 0.65 | ULIFE | 42142609-00007 | JIAXING U-LIFE MEDICAL DEVICE TECHNOLOGY CO.,LTD | CHINA | 21 | 80 | USD | 1 box | YES | BOX | COMPLY pending confirmation of latex and DEHP free | NO TEST REPORTS
42142609-00007 | 126 | Syringe, insulin, plastic, latex free, non-pyrogenic and DEPH free. Type 4 (with a fixed needle). Nominal capacity: 1 ml, 100 units per 1ml, must have clear graduation markings that are not easily removed or rubbed off. Needle: 31G x 6 mm Single use, sterile path To comply with the latest issue of SANS 1166 Packed in units of 10 syringes in a plastic packet 10 Packets in a box. UNIT: Box | N/A | PROMED TECHNOLOGIES (PTY) LTD | 0.68 | U-LIFE | 609007 | JIAXING | CHINA | 14 DAYS | 85% | USD | 1 BOX | YES | BOX | COMPLY pending confirmation of latex and DEHP free | NO TEST REPORTS
42142609-00007 | 126 | Syringe, insulin, plastic, latex free, non-pyrogenic and DEPH free. Type 4 (with a fixed needle). Nominal capacity: 1 ml, 100 units per 1ml, must have clear graduation markings that are not easily removed or rubbed off. Needle: 31G x 6 mm Single use, sterile path To comply with the latest issue of SANS 1166 Packed in units of 10 syringes in a plastic packet 10 Packets in a box. UNIT: Box | N/A | SAFARMEX (PTY) LTD | 0.93 | Saclin | To be confirmed | Zhejiang Mediunion Healthcare Group Co. Ltd | China | 3days | 80% | US$ | 3000 Boxes | Yes | 10 Pasckets in a Box | No sample submitted for evaluation | NO TEST REPORT
42142609-00007 | 126 | Syringe, insulin, plastic, latex free, non-pyrogenic and DEPH free. Type 4 (with a fixed needle). Nominal capacity: 1 ml, 100 units per 1ml, must have clear graduation markings that are not easily removed or rubbed off. Needle: 31G x 6 mm Single use, sterile path To comply with the latest issue of SANS 1166 Packed in units of 10 syringes in a plastic packet 10 Packets in a box. UNIT: Box |  | SUPRA HEALTHCARE JOHANNESBURG PTY LTD | 0.97 | REVITAL | IS131P | Revital Health | KENYA | Initial 60 days, thereafter 21 | 84 | USD | 10 Boxes | YES |  | No sample submitted for evaluation | NO TEST REPORTS
42142609-00007 | 126 | Syringe, insulin, plastic, latex free, non-pyrogenic and DEPH free. Type 4 (with a fixed needle). Nominal capacity: 1 ml, 100 units per 1ml, must have clear graduation markings that are not easily removed or rubbed off. Needle: 31G x 6 mm Single use, sterile path To comply with the latest issue of SANS 1166 Packed in units of 10 syringes in a plastic packet 10 Packets in a box. UNIT: Box |  | Isigidi Trading 95 cc | 63.25 | Isigidi | SY51 | Jiangsu Jichun Medical Devices Co. Ltd. | China | 14 days | 85% | USD | 30boxes | Yes | 100pcs/box | COMPLY pending confirmation of latex and DEHP free | COMPLY(ACCEPTED WITH DEVIATION NEEDLE LENGTH5.5MM) | Recommended | 91
42142609-00007 | 126 | Syringe, insulin, plastic, latex free, non-pyrogenic and DEPH free. Type 4 (with a fixed needle). Nominal capacity: 1 ml, 100 units per 1ml, must have clear graduation markings that are not easily removed or rubbed off. Needle: 31G x 6 mm Single use, sterile path To comply with the latest issue of SANS 1166 Packed in units of 10 syringes in a plastic packet 10 Packets in a box. UNIT: Box | Basic Line offer | Becton Dickinson | 6.4 | BD Microfine Plus | 324905 | BECTON DICKINSON | USA | 7 Working Days | 76% | USD | 1 BOX | Yes | 1 Box =100 Units | COMPLY pending confirmation of latex and DEHP free | NO TEST REPORTS
42142609-00007 | 126 | Syringe, insulin, plastic, latex free, non-pyrogenic and DEPH free. Type 4 (with a fixed needle). Nominal capacity: 1 ml, 100 units per 1ml, must have clear graduation markings that are not easily removed or rubbed off. Needle: 31G x 6 mm Single use, sterile path To comply with the latest issue of SANS 1166 Packed in units of 10 syringes in a plastic packet 10 Packets in a box. UNIT: Box | 1 | CENTRAL MEDICAL PTY LTD | 7.78 | LYSA | INSL103106 | Zhejian INI Medical Devices Co Ltd | China | 21 | 80 | China yuan | 20 boxes of 10 packets of 10 units | yes | Box of 10 packets of 10 unit | COMPLY pending confirmation of latex and DEHP free | NO TEST REPORTS
42142609-00007 | 126 | Syringe, insulin, plastic, latex free, non-pyrogenic and DEPH free. Type 4 (with a fixed needle). Nominal capacity: 1 ml, 100 units per 1ml, must have clear graduation markings that are not easily removed or rubbed off. Needle: 31G x 6 mm Single use, sterile path To comply with the latest issue of SANS 1166 Packed in units of 10 syringes in a plastic packet 10 Packets in a box. UNIT: Box |  | GIBELA TRADE AND INVEST 1001 CC | 8.34 | TARA | SYR042 | Shandong Qiaopai Group Co., Ltd | CHINA | 21 | 85 | USD | 10 | YES | box 100 | No sample submitted for evaluation | NO TEST REPORTS
42142609-00008 | 127 | Syringe, insulin, plastic, latex free, non-pyrogenic and DEPH free. Type 4 (with a fixed needle). Nominal capacity: 1 ml, 100 units per 1ml, must have clear graduation markings that are not easily removed or rubbed off. Needle: 31G x 6 mm Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 1166 Box of 100. UNIT: Box |  | Isigidi Trading 95 cc | 63.25 | Isigidi | SY52 | Jiangsu Jichun Medical Devices Co. Ltd. | China | 14 days | 85% | USD | 30boxes | Yes | 100pcs/box | COMPLY pending confirmation of latex and DEHP free | COMPLY(ACCEPTED WITH DEVIATION NEEDLE LENGTH5.5MM) | Recommended | 91
42142609-00008 | 127 | Syringe, insulin, plastic, latex free, non-pyrogenic and DEPH free. Type 4 (with a fixed needle). Nominal capacity: 1 ml, 100 units per 1ml, must have clear graduation markings that are not easily removed or rubbed off. Needle: 31G x 6 mm Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 1166 Box of 100. UNIT: Box |  | Medi-Core Technologies | 0.67 | ULIFE | 42142609-00008 | JIAXING U-LIFE MEDICAL DEVICE TECHNOLOGY CO.,LTD | CHINA | 21 | 80 | USD | 1 box | YES | BOX | No sample submitted for evaluation | NO TEST REPORTS
42142609-00008 | 127 | Syringe, insulin, plastic, latex free, non-pyrogenic and DEPH free. Type 4 (with a fixed needle). Nominal capacity: 1 ml, 100 units per 1ml, must have clear graduation markings that are not easily removed or rubbed off. Needle: 31G x 6 mm Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 1166 Box of 100. UNIT: Box | N/A | PROMED TECHNOLOGIES (PTY) LTD | 0.68 | U-LIFE | 609008 | JIAXING | CHINA | 14 DAYS | 85% | USD | 1BOX | YES | BOX | No samples submitted for evaluation | NO TEST REPORTS
42142609-00008 | 127 | Syringe, insulin, plastic, latex free, non-pyrogenic and DEPH free. Type 4 (with a fixed needle). Nominal capacity: 1 ml, 100 units per 1ml, must have clear graduation markings that are not easily removed or rubbed off. Needle: 31G x 6 mm Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 1166 Box of 100. UNIT: Box | 1 | CENTRAL MEDICAL PTY LTD | 0.77 | LYSA | INSL103106 | Zhejian INI Medical Devices Co Ltd | China | 21 | 80 | China yuan | 20 boxes of 100 units | yes | Box of 100 units | COMPLY pending confirmation of latex and DEHP free | NO TEST REPORTS
42142609-00008 | 127 | Syringe, insulin, plastic, latex free, non-pyrogenic and DEPH free. Type 4 (with a fixed needle). Nominal capacity: 1 ml, 100 units per 1ml, must have clear graduation markings that are not easily removed or rubbed off. Needle: 31G x 6 mm Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 1166 Box of 100. UNIT: Box | Basic Line offer | Becton Dickinson | 0.82 | BD Microfine Plus | 326676 | BECTON DICKINSON | USA | 7 Working Days | 76% | USD | 1 BOX | Yes | 1 Box =100 Units | COMPLY pending confirmation of latex and DEHP free | NO TEST REPORTS
42142609-00008 | 127 | Syringe, insulin, plastic, latex free, non-pyrogenic and DEPH free. Type 4 (with a fixed needle). Nominal capacity: 1 ml, 100 units per 1ml, must have clear graduation markings that are not easily removed or rubbed off. Needle: 31G x 6 mm Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 1166 Box of 100. UNIT: Box |  | GIBELA TRADE AND INVEST 1001 CC | 0.83 | TARA | SYR043 | Shandong Qiaopai Group Co., Ltd | CHINA | 21 | 85 | USD | 10 | YES | box 100 | No sample submitted for evaluation | NO TEST REPORTS
42142609-00008 | 127 | Syringe, insulin, plastic, latex free, non-pyrogenic and DEPH free. Type 4 (with a fixed needle). Nominal capacity: 1 ml, 100 units per 1ml, must have clear graduation markings that are not easily removed or rubbed off. Needle: 31G x 6 mm Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 1166 Box of 100. UNIT: Box | N/A | SAFARMEX (PTY) LTD | 0.9 | Saclin | To be confirmed | Zhejiang Mediunion Healthcare Group Co. Ltd | China | 3days | 80% | US$ | 3000 Boxes | Yes | Box of 100 | No sample submitted for evaluation | NO TEST REPORT
42142609-00008 | 127 | Syringe, insulin, plastic, latex free, non-pyrogenic and DEPH free. Type 4 (with a fixed needle). Nominal capacity: 1 ml, 100 units per 1ml, must have clear graduation markings that are not easily removed or rubbed off. Needle: 31G x 6 mm Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 1166 Box of 100. UNIT: Box |  | SUPRA HEALTHCARE JOHANNESBURG PTY LTD | 0.97 | REVITAL | IS131B | Revital Health | KENYA | Initial 60 days, thereafter 21 | 84 | USD | 10 Boxes | YES |  | No sample submitted for evaluation | NO TEST REPORTS
42142609-00009 | 128 | Syringe, plastic, latex free, non-pyrogenic and DEPH free. Concentric nozzle, luer slip, Type 1. Nominal capacity: 1 ml, must have clear graduation markings that are not easily removed or rubbed off. To be labelled "FOR TUBERCULIN" With Fixed Needle 25G (0.5mm) x 16mm - for intradermal use Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 1124-1 and 1124-2 Box of 100. UNIT: Box |  | Isigidi Trading 95 cc | 48.3 | Isigidi | SY55 | Jiangsu Jichun Medical Devices Co. Ltd. | China | 14 days | 85% | USD | 30boxes | Yes | 100pcs/box | offered a removable needle instead of a fixed needle | COMPLY | Recommended | 91
42142609-00009 | 128 | Syringe, plastic, latex free, non-pyrogenic and DEPH free. Concentric nozzle, luer slip, Type 1. Nominal capacity: 1 ml, must have clear graduation markings that are not easily removed or rubbed off. To be labelled "FOR TUBERCULIN" With Fixed Needle 25G (0.5mm) x 16mm - for intradermal use Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 1124-1 and 1124-2 Box of 100. UNIT: Box | N/A | SAFARMEX (PTY) LTD | 0.56 | Saclin | To be confirmed | Zhejiang Mediunion Healthcare Group Co. Ltd | China | 3days | 80% | US$ | 3000 Boxes | Yes | Box of 100 | No sample submitted for evaluation | NO TEST REPORT
42142609-00009 | 128 | Syringe, plastic, latex free, non-pyrogenic and DEPH free. Concentric nozzle, luer slip, Type 1. Nominal capacity: 1 ml, must have clear graduation markings that are not easily removed or rubbed off. To be labelled "FOR TUBERCULIN" With Fixed Needle 25G (0.5mm) x 16mm - for intradermal use Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 1124-1 and 1124-2 Box of 100. UNIT: Box |  | GIBELA TRADE AND INVEST 1001 CC | 0.58 | TARA | SYR044 | Shandong Qiaopai Group Co., Ltd | CHINA | 21 | 85 | USD | 10 | YES | box 100 | No sample submitted for evaluation | NO TEST REPORTS
42142609-00009 | 128 | Syringe, plastic, latex free, non-pyrogenic and DEPH free. Concentric nozzle, luer slip, Type 1. Nominal capacity: 1 ml, must have clear graduation markings that are not easily removed or rubbed off. To be labelled "FOR TUBERCULIN" With Fixed Needle 25G (0.5mm) x 16mm - for intradermal use Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 1124-1 and 1124-2 Box of 100. UNIT: Box | 1 | B. BRAUN MEDICAL (PTY) LTD | 0.65 | Omnifix 1ml Fine Dosage Duo 25G x 5/8" | 9161465V | B. Braun | Germany | 42 | 50% | Euro | 100 | Yes | Box of 100 | offered a removable needle instead of a fixed needle | NOT COMPLY
42142609-00009 | 128 | Syringe, plastic, latex free, non-pyrogenic and DEPH free. Concentric nozzle, luer slip, Type 1. Nominal capacity: 1 ml, must have clear graduation markings that are not easily removed or rubbed off. To be labelled "FOR TUBERCULIN" With Fixed Needle 25G (0.5mm) x 16mm - for intradermal use Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 1124-1 and 1124-2 Box of 100. UNIT: Box | 1 | CENTRAL MEDICAL PTY LTD | 0.72 | LYSA | TUB102516 | Zhejian INI Medical Devices Co Ltd | China | 21 | 80 | China yuan | 20 boxes of 100 units | yes | Box of 100 units | COMPLY pending confirmation of latex and DEHP free | NO TEST REPORTS
42142609-00009 | 128 | Syringe, plastic, latex free, non-pyrogenic and DEPH free. Concentric nozzle, luer slip, Type 1. Nominal capacity: 1 ml, must have clear graduation markings that are not easily removed or rubbed off. To be labelled "FOR TUBERCULIN" With Fixed Needle 25G (0.5mm) x 16mm - for intradermal use Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 1124-1 and 1124-2 Box of 100. UNIT: Box |  | ALMIKA TRADING PTY LTD | 0.99 | BD | 303175 | BD | USA | 14 | 0 | NO | 100 | YES | 1 | No sample submitted for evaluation. | NO TEST REPORTS
42142609-00009 | 128 | Syringe, plastic, latex free, non-pyrogenic and DEPH free. Concentric nozzle, luer slip, Type 1. Nominal capacity: 1 ml, must have clear graduation markings that are not easily removed or rubbed off. To be labelled "FOR TUBERCULIN" With Fixed Needle 25G (0.5mm) x 16mm - for intradermal use Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 1124-1 and 1124-2 Box of 100. UNIT: Box | Basic Line offer | Becton Dickinson | 1.02 | BD Plastipak™ 1ml, with 25G Needle X 16mm | 303175 | Becton Dickinson | Fraga, Spain | 7 Working Days | 76% | USD | 1 Box | NO | 120 units in a box | Package / sample is not marked for TB. Offered a removable needle instead of a fixed needle. | NO TEST REPORTS
42142609-00009 | 128 | Syringe, plastic, latex free, non-pyrogenic and DEPH free. Concentric nozzle, luer slip, Type 1. Nominal capacity: 1 ml, must have clear graduation markings that are not easily removed or rubbed off. To be labelled "FOR TUBERCULIN" With Fixed Needle 25G (0.5mm) x 16mm - for intradermal use Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 1124-1 and 1124-2 Box of 100. UNIT: Box |  | SUPRA HEALTHCARE JOHANNESBURG PTY LTD | 1.22 | AVACARE | 1S125B | HENAN AILE INDUSTRIAL CO.,LTD | CHINA | Initial 60 days, thereafter 21 | 84 | USD | 10 Boxes | YES |  | No sample submitted for evaluation | NO TEST REPORTS
42142609-00009 | 128 | Syringe, plastic, latex free, non-pyrogenic and DEPH free. Concentric nozzle, luer slip, Type 1. Nominal capacity: 1 ml, must have clear graduation markings that are not easily removed or rubbed off. To be labelled "FOR TUBERCULIN" With Fixed Needle 25G (0.5mm) x 16mm - for intradermal use Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 1124-1 and 1124-2 Box of 100. UNIT: Box | 1 | Biologica Pharmaceuticals Pty Ltd | 3.39 | Sol- M | 10018IM | Sol- Millennium | China | 4-6 Weeks | 100% | USD | 1000 Boxes | 100% | N/A
42142609-00010 | 129 | Syringe, plastic, latex free, non-pyrogenic and DEPH free. Concentric nozzle, luer slip, Type 1. Nominal capacity: 1 ml, must have clear graduation markings that are not easily removed or rubbed off. To be labelled "FOR TUBERCULIN" With Fixed Needle 25G (0.5mm) x 16mm - for intradermal use With automated retraction or similar safety mechanism Sterile, Single use, individually peel packed To comply with the latest issue of SANS 1124-1 and 1124-2 Box of 100. UNIT: Box |  | GIBELA TRADE AND INVEST 1001 CC | 1.75 | TARA | SYR045 | Shandong Qiaopai Group Co., Ltd | CHINA | 21 | 85 | USD | 10 | YES | box 100 | No sample submitted for evaluation | NO TEST REPORTS | Re-tender
42142609-00010 | 129 | Syringe, plastic, latex free, non-pyrogenic and DEPH free. Concentric nozzle, luer slip, Type 1. Nominal capacity: 1 ml, must have clear graduation markings that are not easily removed or rubbed off. To be labelled "FOR TUBERCULIN" With Fixed Needle 25G (0.5mm) x 16mm - for intradermal use With automated retraction or similar safety mechanism Sterile, Single use, individually peel packed To comply with the latest issue of SANS 1124-1 and 1124-2 Box of 100. UNIT: Box | 1 | Biologica Pharmaceuticals Pty Ltd | 3.39 | Sol- M | 10018IM | Sol- Millennium | China | 4-6 Weeks | 100% | USD | 1000 Boxes | 100% | N/A |  |  | Re-tender
42142609-00073 | 130 | Syringe, plastic, latex free, non-pyrogenic and DEPH free. Concentric nozzle, luer slip, Type 1. Nominal capacity: 1 ml, must have clear graduation markings that are not easily removed or rubbed off. To be labelled "FOR TUBERCULIN" With Fixed Needle 27G (0.4mm) x 12.7mm - for intradermal use With automated retraction or similar safety mechanism Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 1124-1 and 2 Box of 100 UNIT: Box |  | GIBELA TRADE AND INVEST 1001 CC | 1.75 | TARA | SYR046 | Shandong Qiaopai Group Co., Ltd | CHINA | 21 | 85 | USD | 10 | YES | box 100 | No sample submitted for evaluation | NO TEST REPORTS | Re-tender
42142608-00019 | 131 | Syringe, plastic, latex free, non-pyrogenic and DEPH free. Concentric nozzle, luer slip, Type 1. Nominal capacity: 1 ml, must have clear graduation markings that are not easily removed or rubbed off. To be labelled "FOR TUBERCULIN" Without Needle Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 1124-2 Box of 100. UNIT: Box |  | Isigidi Trading 95 cc | 40.25 | Isigidi | SY07 | Jiangsu Jichun Medical Devices Co. Ltd. | China | 14 days | 85% | USD | 30boxes | Yes | 100pcs/box | COMPLY pending confirmation of latex and DEHP free | COMPLY | Recommended | 91
42142608-00019 | 131 | Syringe, plastic, latex free, non-pyrogenic and DEPH free. Concentric nozzle, luer slip, Type 1. Nominal capacity: 1 ml, must have clear graduation markings that are not easily removed or rubbed off. To be labelled "FOR TUBERCULIN" Without Needle Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 1124-2 Box of 100. UNIT: Box | 1 | CENTRAL MEDICAL PTY LTD | 0.43 | LYSA | HSCS1000 | Zhejian INI Medical Devices Co Ltd | China | 21 | 80 | China yuan | 20 Boxes of 100 units | Yes | Box of 100 | COMPLY pending confirmation of latex and DEHP free | NO TEST REPORTS
42142608-00019 | 131 | Syringe, plastic, latex free, non-pyrogenic and DEPH free. Concentric nozzle, luer slip, Type 1. Nominal capacity: 1 ml, must have clear graduation markings that are not easily removed or rubbed off. To be labelled "FOR TUBERCULIN" Without Needle Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 1124-2 Box of 100. UNIT: Box | Basic Line Offer | BIOCOR HOSPITAL SUPPLIES PTY LTD | 0.54 | Triple -D | SYR-Care 1ML TB | Jumla Medical | South Africa | 21 | 0 | N/A | 10 Boxes of 100 | Yes | N/A | COMPLY pending confirmation of latex and DEHP free | NO TEST REPORTS
42142608-00019 | 131 | Syringe, plastic, latex free, non-pyrogenic and DEPH free. Concentric nozzle, luer slip, Type 1. Nominal capacity: 1 ml, must have clear graduation markings that are not easily removed or rubbed off. To be labelled "FOR TUBERCULIN" Without Needle Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 1124-2 Box of 100. UNIT: Box |  | GIBELA TRADE AND INVEST 1001 CC | 0.58533 | TARA | SYR047 | Shandong Qiaopai Group Co., Ltd | CHINA | 21 | 85 | USD | 10 | YES | each | No sample submitted for evaluation | NO TEST REPORTS
42142608-00019 | 131 | Syringe, plastic, latex free, non-pyrogenic and DEPH free. Concentric nozzle, luer slip, Type 1. Nominal capacity: 1 ml, must have clear graduation markings that are not easily removed or rubbed off. To be labelled "FOR TUBERCULIN" Without Needle Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 1124-2 Box of 100. UNIT: Box |  | ALMIKA TRADING PTY LTD | 0.93 | BD | 303172 | BD | USA | 14 | 0 | NO | 100 | YES | 1 | not labelled "for tuberculin" | NO TEST REPORTS
42142608-00019 | 131 | Syringe, plastic, latex free, non-pyrogenic and DEPH free. Concentric nozzle, luer slip, Type 1. Nominal capacity: 1 ml, must have clear graduation markings that are not easily removed or rubbed off. To be labelled "FOR TUBERCULIN" Without Needle Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 1124-2 Box of 100. UNIT: Box | Basic Line offer | Becton Dickinson | 0.94 | BD Plastipak™ 1ml, TB Syringe | 303172 | Becton Dickinson | Fraga, Spain | 7 Working Days | 76% | USD | 1 box | YES | 120 units in a box | The product is not labelled for TB | NO TEST REPORTS
42142608-00019 | 131 | Syringe, plastic, latex free, non-pyrogenic and DEPH free. Concentric nozzle, luer slip, Type 1. Nominal capacity: 1 ml, must have clear graduation markings that are not easily removed or rubbed off. To be labelled "FOR TUBERCULIN" Without Needle Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 1124-2 Box of 100. UNIT: Box |  | JUMLA MEDICAL SUPPLIES (PTY) LTD | 41.04 | SYR-CARE | SYR-CARE-1MLTB | SUZHOU ACMED | CHINA | 21 DAYS | 72% | US DOLLAR | 10 BOXES | YES | NIL | COMPLY pending confirmation of latex and DEHP free | NOT COMPLY
42142608-00019 | 131 | Syringe, plastic, latex free, non-pyrogenic and DEPH free. Concentric nozzle, luer slip, Type 1. Nominal capacity: 1 ml, must have clear graduation markings that are not easily removed or rubbed off. To be labelled "FOR TUBERCULIN" Without Needle Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 1124-2 Box of 100. UNIT: Box |  | SUPRA HEALTHCARE JOHANNESBURG PTY LTD | 54.65 | Avacare | BS1ML | HENAN AILE INDUSTRIAL CO.,LTD | CHINA | Initial 60 days, thereafter 21 | 84 | USD | 10 Boxes | YES |  | not labelled for“ fortuberculin" | NO TEST REPORTS
42142609-00000 | 132 | Syringe, plastic, latex free, non-pyrogenic and DEPH free, with fixed needle for intradermal BCG vaccine. Nominal capacity: 0,5 ml, must have clear graduation markings that are not easily removed or rubbed off. Needle: 27G (0,4mm) x 10 mm Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 1124-1 and 1124-2 Packed in units of 10 syringes in a plastic packet 10 Packets in a box. UNIT: Box |  | Medi-Core Technologies | 0.67 | ULIFE | 42142609-00000 | JIAXING U-LIFE MEDICAL DEVICE TECHNOLOGY CO.,LTD | CHINA | 21 | 80 | USD | 1 | YES | EACH | Re-tender endusers to revise specification |  | Re-tender endusers to revise specification
42142609-00000 | 132 | Syringe, plastic, latex free, non-pyrogenic and DEPH free, with fixed needle for intradermal BCG vaccine. Nominal capacity: 0,5 ml, must have clear graduation markings that are not easily removed or rubbed off. Needle: 27G (0,4mm) x 10 mm Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 1124-1 and 1124-2 Packed in units of 10 syringes in a plastic packet 10 Packets in a box. UNIT: Box |  | PROMED TECHNOLOGIES (PTY) LTD | 0.68 | U-LIFE | 609000 |  | CHINA | 14 DAYS | 85% | USD | 1 UNIT | YES | UNIT | Re-tender end-users to revise specification |  | Re-tender endusers to revise specification
42142609-00000 | 132 | Syringe, plastic, latex free, non-pyrogenic and DEPH free, with fixed needle for intradermal BCG vaccine. Nominal capacity: 0,5 ml, must have clear graduation markings that are not easily removed or rubbed off. Needle: 27G (0,4mm) x 10 mm Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 1124-1 and 1124-2 Packed in units of 10 syringes in a plastic packet 10 Packets in a box. UNIT: Box |  | Chemical World cc | 0.87 | SURG PLUS | SYRBCG | TONGDA | CHINA | [6 empty] | Re-tender end-users to revise specification |  | Re-tender endusers to revise specification
42142609-00000 | 132 | Syringe, plastic, latex free, non-pyrogenic and DEPH free, with fixed needle for intradermal BCG vaccine. Nominal capacity: 0,5 ml, must have clear graduation markings that are not easily removed or rubbed off. Needle: 27G (0,4mm) x 10 mm Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 1124-1 and 1124-2 Packed in units of 10 syringes in a plastic packet 10 Packets in a box. UNIT: Box |  | SUPRA HEALTHCARE JOHANNESBURG PTY LTD | 1.1 | Avacare | BCG27 | HENAN AILE INDUSTRIAL CO.,LTD | CHINA | Initial 60 days, thereafter 21 | 84 | USD | 10 Boxes | YES |  | Re-tender end-users to revise specification |  | Re-tender endusers to revise specification
42142609-00000 | 132 | Syringe, plastic, latex free, non-pyrogenic and DEPH free, with fixed needle for intradermal BCG vaccine. Nominal capacity: 0,5 ml, must have clear graduation markings that are not easily removed or rubbed off. Needle: 27G (0,4mm) x 10 mm Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 1124-1 and 1124-2 Packed in units of 10 syringes in a plastic packet 10 Packets in a box. UNIT: Box |  | ALMIKA TRADING PTY LTD | 1.12 | BD | 300860 | BD | USA | 14 | 0 | NO | 100 | YES | 1 | Re-tender: End-users to revise specifications |  | Re-tender endusers to revise specification
42142609-00000 | 132 | Syringe, plastic, latex free, non-pyrogenic and DEPH free, with fixed needle for intradermal BCG vaccine. Nominal capacity: 0,5 ml, must have clear graduation markings that are not easily removed or rubbed off. Needle: 27G (0,4mm) x 10 mm Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 1124-1 and 1124-2 Packed in units of 10 syringes in a plastic packet 10 Packets in a box. UNIT: Box |  | GIBELA TRADE AND INVEST 1001 CC | 5.15 | TARA | SYR048 | Shandong Qiaopai Group Co., Ltd | CHINA | 21 | 85 | USD | 10 | YES | box 100 | Re-tender endusers to revise specifications |  | Re-tender endusers to revise specification
42142609-00000 | 132 | Syringe, plastic, latex free, non-pyrogenic and DEPH free, with fixed needle for intradermal BCG vaccine. Nominal capacity: 0,5 ml, must have clear graduation markings that are not easily removed or rubbed off. Needle: 27G (0,4mm) x 10 mm Sterile, Single use, Individually peel packed To comply with the latest issue of SANS 1124-1 and 1124-2 Packed in units of 10 syringes in a plastic packet 10 Packets in a box. UNIT: Box | Basic Line offer | Becton Dickinson | 9.5 | BD BCG Syringe 0.5ml, 27G x 10mm | 300860 | Becton Dickinson | Fraga, Spain | 7 Working Days | 76% | USD | 1 box | YES | 100 units in a box | Re-tender end-users to revise specifications |  | Re-tender endusers to revise specification
42142616-00000 | 133 | Syringe, aspiration, arterial blood collection Nominal capacity: 1ml, must have clear graduation markings that are not easily removed or rubbed off. Luer Slip Without Needle Contains Calcium balanced Heparin Unit: Box (Box of 100) |  | Chemical World cc | 362 | SURG PLUS |  | CAPRICORN | china | 3 | 85 | USD | 18000 | YES | Box of 100 | COMPLY pending confirmation of latex free |  | Recommended | 100
42142616-00000 | 133 | Syringe, aspiration, arterial blood collection Nominal capacity: 1ml, must have clear graduation markings that are not easily removed or rubbed off. Luer Slip Without Needle Contains Calcium balanced Heparin Unit: Box (Box of 100) | Basic Line offer | Becton Dickinson | 4.36 | BD VAC A-line, 1ml | 364356 | Becton Dickinson | United Kingdom | 7 Working Days | 76% | USD | 1 BOX | YES | 1 BOX = 100 UNITS | COMPLY
42142616-00000 | 133 | Syringe, aspiration, arterial blood collection Nominal capacity: 1ml, must have clear graduation markings that are not easily removed or rubbed off. Luer Slip Without Needle Contains Calcium balanced Heparin Unit: Box (Box of 100) |  | GIBELA TRADE AND INVEST 1001 CC | 45.43 | TARA | SYR049 | Shandong Qiaopai Group Co., Ltd | CHINA | 21 | 85 | USD | 10 | YES | each | No sample submitted for evaluation | NO TEST REPORTS
42142616-00001 | 134 | Syringe, aspiration, arterial blood collection Nominal capacity: 3ml, must have clear graduation markings that are not easily removed or rubbed off. Luer Slip Without Needle Contains Calcium balanced Heparin UNIT: Box (Box of 100) |  | Chemical World cc | 305 | SURG PLUS |  | CAPRICORN | RSA | 3 | 85 | USD | 18000 | YES | Box of 100 | COMPLY pending confirmation of latex free |  | Recommended | 100
42142616-00001 | 134 | Syringe, aspiration, arterial blood collection Nominal capacity: 3ml, must have clear graduation markings that are not easily removed or rubbed off. Luer Slip Without Needle Contains Calcium balanced Heparin UNIT: Box (Box of 100) | Basic Line offer | Becton Dickinson | 4.14 | BD VAC A-line, 3ml | 364376 | Becton Dickinson | United Kingdom | 7 Working Days | 76% | USD | 1 BOX | YES | 1 BOX = 100 UNITS | COMPLY
42142616-00001 | 134 | Syringe, aspiration, arterial blood collection Nominal capacity: 3ml, must have clear graduation markings that are not easily removed or rubbed off. Luer Slip Without Needle Contains Calcium balanced Heparin UNIT: Box (Box of 100) |  | GIBELA TRADE AND INVEST 1001 CC | 45.43 | TARA | SYR050 | Shandong Qiaopai Group Co., Ltd | CHINA | 21 | 85 | USD | 10 | YES | each | No sample submitted for evaluation
42142616-00002 | 135 | Syringe, aspiration, arterial blood collection Nominal capacity: 1ml, must have clear graduation markings that are not easily removed or rubbed off. Luer Slip With Needle, 23G Contains Calcium balanced Heparin UNIT: Box (Box of 100) |  | Chemical World cc | 362 | SURG PLUS |  | CAPRICORN | RSA | 3 | 85 | USD | 18000 | YES | Box of 100 | COMPLY pending confirmation of latex free |  | Recommended | 100
42142616-00002 | 135 | Syringe, aspiration, arterial blood collection Nominal capacity: 1ml, must have clear graduation markings that are not easily removed or rubbed off. Luer Slip With Needle, 23G Contains Calcium balanced Heparin UNIT: Box (Box of 100) | Basic Line offer | Becton Dickinson | 7.25 | BD VAC Preset, 1ml | 364413 | Becton Dickinson | United Kingdom | 7 Working Days | 76% | USD | 1 BOX | YES | 1 BOX = 100 UNITS | COMPLY
42142616-00002 | 135 | Syringe, aspiration, arterial blood collection Nominal capacity: 1ml, must have clear graduation markings that are not easily removed or rubbed off. Luer Slip With Needle, 23G Contains Calcium balanced Heparin UNIT: Box (Box of 100) |  | SUPRA HEALTHCARE JOHANNESBURG PTY LTD | 27 | Avacare | BSASP01 | HENAN AILE INDUSTRIAL CO.,LTD | CHINA | Initial 60 days, thereafter 21 | 84 | USD | 1 Box | YES | 50 pcs/box | No sample submitted for evaluation
42142616-00002 | 135 | Syringe, aspiration, arterial blood collection Nominal capacity: 1ml, must have clear graduation markings that are not easily removed or rubbed off. Luer Slip With Needle, 23G Contains Calcium balanced Heparin UNIT: Box (Box of 100) |  | GIBELA TRADE AND INVEST 1001 CC | 45.43 | TARA | SYR051 | Shandong Qiaopai Group Co., Ltd | CHINA | 21 | 85 | USD | 10 | YES | each | No sample submitted for evaluation
42142616-00003 | 136 | Syringe, aspiration, arterial blood collection Nominal capacity: 3ml, must have clear graduation markings that are not easily removed or rubbed off. Luer Slip With Needle, 22G With auto retractable mechanism or similar safety mechanism Contains Calcium balanced Heparin UNIT: Box (Box of 25) |  | SUPRA HEALTHCARE JOHANNESBURG PTY LTD | 27 | Avacare | BSASP03 | HENAN AILE INDUSTRIAL CO.,LTD | CHINA | Initial 60 days, thereafter 21 | 84 | USD | 1 Box | YES | 50 pcs/box | No sample submitted for evaluation |  | Re-tender
42142616-00003 | 136 | Syringe, aspiration, arterial blood collection Nominal capacity: 3ml, must have clear graduation markings that are not easily removed or rubbed off. Luer Slip With Needle, 22G With auto retractable mechanism or similar safety mechanism Contains Calcium balanced Heparin UNIT: Box (Box of 25) | Basic Line offer | Becton Dickinson | 28.44 | BD VAC Preset, 3ml | 364390 | Becton Dickinson | United Kingdom | 7 Working Days | 76% | USD | 1 BOX | YES | 1 BOX = 100 UNITS | Offered a luer lock instead of a luer slip |  | Re-tender
42142616-00003 | 136 | Syringe, aspiration, arterial blood collection Nominal capacity: 3ml, must have clear graduation markings that are not easily removed or rubbed off. Luer Slip With Needle, 22G With auto retractable mechanism or similar safety mechanism Contains Calcium balanced Heparin UNIT: Box (Box of 25) |  | GIBELA TRADE AND INVEST 1001 CC | 45.43 | TARA | SYR052 | Shandong Qiaopai Group Co., Ltd | CHINA | 21 | 85 | USD | 10 | YES | each | No sample submitted for evaluation |  | Re-tender
42142602-00000 | 137 | Syringe, bulb, for irrigation, latex free, non-pyrogenic and DEPH free, with rubber or plastic bulb and catheter tip, length of neck of bulb: 15mm. Nominal capacity: Not less than 50ml, must have clear graduation markings that are not easily removed or rubbed off. Sterile, Single use, Individually peel packed UNIT: Box (Box of 25) |  | Chemical World cc | 199.75 | SURG PLUS |  | CAPRICORN | RSA | 3 | 85 | USD | 18000 | YES | Box of 25 | COMPLY pending confirmation of latex and DEHP free |  | Recommended | 100
42142602-00000 | 137 | Syringe, bulb, for irrigation, latex free, non-pyrogenic and DEPH free, with rubber or plastic bulb and catheter tip, length of neck of bulb: 15mm. Nominal capacity: Not less than 50ml, must have clear graduation markings that are not easily removed or rubbed off. Sterile, Single use, Individually peel packed UNIT: Box (Box of 25) | BASIC LINE OFFER | MALACHITE MEDICAL (PTY) LTD | 8.02 | PRIONGUARD | SYB00010 | CHANGZHOU CHUANGJIA MEDICAL APPLIANCE CO., LTD | CHINA | 21 | 74.27 | USD | EACH | YES | 60 PER CASE | COMPLY pending confirmation of latex and DEHP free
42142602-00000 | 137 | Syringe, bulb, for irrigation, latex free, non-pyrogenic and DEPH free, with rubber or plastic bulb and catheter tip, length of neck of bulb: 15mm. Nominal capacity: Not less than 50ml, must have clear graduation markings that are not easily removed or rubbed off. Sterile, Single use, Individually peel packed UNIT: Box (Box of 25) |  | SUPRA HEALTHCARE JOHANNESBURG PTY LTD | 8.24 | Avacare | BSBULB | Suprahealthcare Pty Ltd | South africa | 14 days | 65 | USD | 1 Box | YES |  | COMPLY pending confirmation of latex and DEHP free
42142602-00000 | 137 | Syringe, bulb, for irrigation, latex free, non-pyrogenic and DEPH free, with rubber or plastic bulb and catheter tip, length of neck of bulb: 15mm. Nominal capacity: Not less than 50ml, must have clear graduation markings that are not easily removed or rubbed off. Sterile, Single use, Individually peel packed UNIT: Box (Box of 25) | Basic Line Offer | BIOCOR HOSPITAL SUPPLIES PTY LTD | 9.96 | Technomed | TSB 600 | Techno Med | South Africa | 21 | 0 | N/A | 2 Boxes of 50 | Yes | N/A | COMPLY pending confirmation of latex and DEHP free
42142602-00000 | 137 | Syringe, bulb, for irrigation, latex free, non-pyrogenic and DEPH free, with rubber or plastic bulb and catheter tip, length of neck of bulb: 15mm. Nominal capacity: Not less than 50ml, must have clear graduation markings that are not easily removed or rubbed off. Sterile, Single use, Individually peel packed UNIT: Box (Box of 25) |  | UNITRADE 1032 CC | 16.49 | IRRISURE | 1 | UNITRADE 1032 CC | RSA | 21 | 0 | N/A | 100 | YES | 1 | No brand name No product code
42142609-00011 | 138 | Syringe, plastic, latex free, non-pyrogenic and DEPH free. Nominal capacity: 0.5ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. With attachable needle 20G(0,9mm) x 40mm, colour code: yellow and must have an automated retractable or similar safety mechanism Sterile, Single use, individually peel packed To comply with the latest issue of SANS 1124-1 and 1124-2 UNIT: Box |  | GIBELA TRADE AND INVEST 1001 CC | 1.75 | TARA | SYR053 | Shandong Qiaopai Group Co., Ltd | CHINA | 21 | 85 | USD | 10 | YES | box 100 | No sample submitted for evaluation | NO TEST REPORTS | Re-tender
42142609-00012 | 139 | Syringe, plastic, latex free, non-pyrogenic and DEPH free. Nominal capacity: 0.5ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. With attachable needle 21G(0,8mm) x 40mm, colour code: deep green and must have an automated retractable or similar safety mechanism Sterile, Single use, individually peel packed To comply with the latest issue of SANS 1124-1 and 1124-2 UNIT: Box |  | GIBELA TRADE AND INVEST 1001 CC | 1.75 | TARA | SYR054 | Shandong Qiaopai Group Co., Ltd | CHINA | 21 | 85 | USD | 10 | YES | box 100 | No sample submitted for evaluation | NO TEST REPORTS | Re-tender
42142609-00013 | 140 | Syringe, plastic, latex free, non-pyrogenic and DEPH free. Nominal capacity: 0.5 ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. With attachable needle 22G(0,7mm) x 32mm, colour code: black and must have an automated retractable or similar safety mechanism Sterile, Single use, individually peel packed To comply with the latest issue of SANS 1124-1 and 1124-2 UNIT: Box |  | GIBELA TRADE AND INVEST 1001 CC | 1.75 | TARA | SYR055 | Shandong Qiaopai Group Co., Ltd | CHINA | 21 | 85 | USD | 10 | YES | box 100 | No sample submitted for evaluation | NO TEST REPORTS | Re-tender
42142609-00014 | 141 | Syringe, plastic, latex free, non-pyrogenic and DEPH free. Nominal capacity: 0.5ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. With attachable needle 23G(0,6mm) x 25mm,colour code: deep blue and must have an automated retractable or similar safety mechanism Sterile, Single use, individually peel packed To comply with the latest issue of SANS 1124-1 and 1124-2 UNIT: Box | N/A | SAFARMEX (PTY) LTD | 1.08 | Saclin | To be confirmed | Zhejiang Mediunion Healthcare Group Co. Ltd | China | 3days | 80% | US$ | 3000 Boxes | Yes | 100 Pieces in a Box | No sample submitted for evaluation | NO TEST REPORT | Re-tender
42142609-00014 | 141 | Syringe, plastic, latex free, non-pyrogenic and DEPH free. Nominal capacity: 0.5ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. With attachable needle 23G(0,6mm) x 25mm,colour code: deep blue and must have an automated retractable or similar safety mechanism Sterile, Single use, individually peel packed To comply with the latest issue of SANS 1124-1 and 1124-2 UNIT: Box |  | GIBELA TRADE AND INVEST 1001 CC | 1.75 | TARA | SYR056 | Shandong Qiaopai Group Co., Ltd | CHINA | 21 | 85 | USD | 10 | YES | box 100 | No sample submitted for evaluation | NO TEST REPORTS | Re-tender
42142609-00015 | 142 | Syringe, plastic, latex free, non-pyrogenic and DEPH free. Nominal capacity: 0.5ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. With attachable needle 24G(0.55mm) x 20mm, colour code: medium purple and must have an automated retractable or similar safety mechanism Sterile, Single use, individually peel packed To comply with the latest issue of SANS 1124-1 and 1124-2 UNIT: Box | N/A | SAFARMEX (PTY) LTD | 1.08 | Saclin | To be confirmed | Zhejiang Mediunion Healthcare Group Co. Ltd | China | 3days | 80% | US$ | 3000 Boxes | Yes | 100 Pieces in a Box | No sample submitted for evaluation | NO TEST REPORT | Re-tender
42142609-00015 | 142 | Syringe, plastic, latex free, non-pyrogenic and DEPH free. Nominal capacity: 0.5ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. With attachable needle 24G(0.55mm) x 20mm, colour code: medium purple and must have an automated retractable or similar safety mechanism Sterile, Single use, individually peel packed To comply with the latest issue of SANS 1124-1 and 1124-2 UNIT: Box |  | GIBELA TRADE AND INVEST 1001 CC | 1.75 | TARA | SYR057 | Shandong Qiaopai Group Co., Ltd | CHINA | 21 | 85 | USD | 10 | YES | box 100 | No sample submitted for evaluation | NO TEST REPORTS | Re-tender
42142609-00016 | 143 | Syringe, plastic, latex free, non-pyrogenic and DEPH free. Nominal capacity: 0.5ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. With attachable needle 25G(0,50mm) x 16mm, colour code: orange and must have an automated retractable or similar safety mechanism Sterile, Single use, individually peel packed To comply with the latest issue of SANS 1124-1 and 1124-2 UNIT: Box | N/A | SAFARMEX (PTY) LTD | 1.08 | Saclin | To be confirmed | Zhejiang Mediunion Healthcare Group Co. Ltd | China | 3days | 80% | US$ | 3000 Boxes | Yes | 100 Pieces in a Box | No sample submitted for evaluation | NO TEST REPORT | Re-tender
42142609-00016 | 143 | Syringe, plastic, latex free, non-pyrogenic and DEPH free. Nominal capacity: 0.5ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. With attachable needle 25G(0,50mm) x 16mm, colour code: orange and must have an automated retractable or similar safety mechanism Sterile, Single use, individually peel packed To comply with the latest issue of SANS 1124-1 and 1124-2 UNIT: Box |  | GIBELA TRADE AND INVEST 1001 CC | 1.75 | TARA | SYR058 | Shandong Qiaopai Group Co., Ltd | CHINA | 21 | 85 | USD | 10 | YES | box 100 | No sample submitted for evaluation | NO TEST REPORTS | Re-tender
42142609-00017 | 144 | Syringe, plastic, latex free, non-pyrogenic and DEPH free. Nominal capacity: 0.5ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. With attachable needle 26G(0.45mm) x 16mm, colour code: brown an automated retractable or similar safety mechanism Sterile, Single use, individually peel packed To comply with the latest issue of SANS 1124-1 and 1124-2 UNIT: Box | N/A | SAFARMEX (PTY) LTD | 1.08 | Saclin | To be confirmed | Zhejiang Mediunion Healthcare Group Co. Ltd | China | 3days | 80% | US$ | 3000 Boxes | Yes | 100 Pieces in a Box | No sample submitted for evaluation | NO TEST REPORT | Re-tender
42142609-00017 | 144 | Syringe, plastic, latex free, non-pyrogenic and DEPH free. Nominal capacity: 0.5ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. With attachable needle 26G(0.45mm) x 16mm, colour code: brown an automated retractable or similar safety mechanism Sterile, Single use, individually peel packed To comply with the latest issue of SANS 1124-1 and 1124-2 UNIT: Box |  | GIBELA TRADE AND INVEST 1001 CC | 1.75 | TARA | SYR059 | Shandong Qiaopai Group Co., Ltd | CHINA | 21 | 85 | USD | 10 | YES | box 100 | No sample submitted for evaluation | NO TEST REPORTS | Re-tender
42142609-00018 | 145 | Syringe, plastic, latex free, non-pyrogenic and DEPH free. Nominal capacity: 0.5ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. With attachable needle 27G(0.4mm) x 8mm, colour code: medium grey and must have an automated retractable or similar safety mechanism Sterile, Single use, individually peel packed To comply with the latest issue of SANS 1124-1 and 1124-2 UNIT: Box | N/A | SAFARMEX (PTY) LTD | 1.08 | Saclin | To be confirmed | Zhejiang Mediunion Healthcare Group Co. Ltd | China | 3days | 80% | US$ | 3000 Boxes | Yes | 100 Pieces in a Box | No sample submitted for evaluation | NO TEST REPORT | Re-tender
42142609-00018 | 145 | Syringe, plastic, latex free, non-pyrogenic and DEPH free. Nominal capacity: 0.5ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. With attachable needle 27G(0.4mm) x 8mm, colour code: medium grey and must have an automated retractable or similar safety mechanism Sterile, Single use, individually peel packed To comply with the latest issue of SANS 1124-1 and 1124-2 UNIT: Box |  | SUPRA HEALTHCARE JOHANNESBURG PTY LTD | 2.82 | Avacare | 1S0527BS | HENAN AILE INDUSTRIAL CO.,LTD | CHINA | Initial 60 days, thereafter 21 | 84 | USD | 10 Boxes | YES |  | No sample submitted for evaluation | NO TEST REPORTS | Re-tender
42142609-00019 | 146 | Syringe, plastic, latex free, non-pyrogenic and DEPH free. Nominal capacity: 0.5ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. With attachable needle 29G(0.33mm) x 8mm, colour code: red and must have an automated retractable or similar safety mechanism Sterile, Single use, individually peel packed To comply with the latest issue of SANS 1124-1 and 1124-2 UNIT: Box | N/A | SAFARMEX (PTY) LTD | 1.9 | Saclin | To be confirmed | Zhejiang Mediunion Healthcare Group Co. Ltd | China | 3days | 80% | US$ | 3000 Boxes | Yes | 100 Pieces in a Box | No sample submitted for evaluation | NO TEST REPORT | Re-tender
42142609-00020 | 147 | Syringe, plastic, latex free, non-pyrogenic and DEPH free. Nominal capacity: 0.5ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. With attachable needle 30G(0.3mm) x 8mm, colour code: yellow and must have an automated retractable or similar safety mechanism Sterile, Single use, individually peel packed To comply with the latest issue of SANS 1124-1 and 1124-2 UNIT: Box | N/A | SAFARMEX (PTY) LTD | 1.9 | Saclin | To be confirmed | Zhejiang Mediunion Healthcare Group Co. Ltd | China | 3days | 80% | US$ | 3000 Boxes | Yes | 100 Pieces in a Box | No sample submitted for evaluation | NO TEST REPORT | Re-tender
42142609-00021 | 148 | Syringe, plastic, latex free, non-pyrogenic and DEPH free. Nominal capacity: 1ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. With attachable needle 20G(0,9mm) x 40mm, colour code: yellow and must have an automated retractable or similar safety mechanism Sterile, Single use, individually peel packed To comply with the latest issue of SANS 1124-1 and 1124-2 UNIT: Box |  | SUPRA HEALTHCARE JOHANNESBURG PTY LTD | 1.03 | REVITAL | 1S120BS | Revital Health | KENYA | Initial 60 days, thereafter 21 | 84 | USD | 10 Boxes | YES |  | No automated retractable or similar safety mechanism | NO TEST REPORTS | Re-tender
42142609-00021 | 148 | Syringe, plastic, latex free, non-pyrogenic and DEPH free. Nominal capacity: 1ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. With attachable needle 20G(0,9mm) x 40mm, colour code: yellow and must have an automated retractable or similar safety mechanism Sterile, Single use, individually peel packed To comply with the latest issue of SANS 1124-1 and 1124-2 UNIT: Box |  | GIBELA TRADE AND INVEST 1001 CC | 1.75 | TARA | SYR060 | Shandong Qiaopai Group Co., Ltd | CHINA | 21 | 85 | USD | 10 | YES | box 100 | No sample submitted for evaluation | NO TEST REPORTS | Re-tender
42142609-00026 | 149 | Syringe, plastic, latex free, non-pyrogenic and DEPH free. Nominal capacity: 1ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. With attachable needle 21G(0,8mm) x 40mm, colour code: deep green and must have an automated retractable or similar safety mechanism Sterile, Single use, individually peel packed To comply with the latest issue of SANS 1124-1 and 1124-2 UNIT: Box |  | SUPRA HEALTHCARE JOHANNESBURG PTY LTD | 1.06 | REVITAL | BSS0121 | Revital Health | KENYA | Initial 60 days, thereafter 21 | 84 | USD | 10 Boxes | YES | 100/bOX | No automated retractable or similar safety mechanism | NO TEST REPORTS | Re-tender
42142609-00026 | 149 | Syringe, plastic, latex free, non-pyrogenic and DEPH free. Nominal capacity: 1ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. With attachable needle 21G(0,8mm) x 40mm, colour code: deep green and must have an automated retractable or similar safety mechanism Sterile, Single use, individually peel packed To comply with the latest issue of SANS 1124-1 and 1124-2 UNIT: Box |  | GIBELA TRADE AND INVEST 1001 CC | 1.75 | TARA | SYR061 | Shandong Qiaopai Group Co., Ltd | CHINA | 21 | 85 | USD | 10 | YES | box 100 | No sample submitted for evaluation | NO TEST REPORTS | Re-tender
42142609-00027 | 150 | Syringe, plastic, latex free, non-pyrogenic and DEPH free. Nominal capacity: 1ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. With attachable needle 22G(0,7mm) x 32mm, colour code: black and must have an automated retractable or similar safety mechanism Sterile, Single use, individually peel packed To comply with the latest issue of SANS 1124-1 and 1124-2 UNIT: Box |  | SUPRA HEALTHCARE JOHANNESBURG PTY LTD | 1.06 | REVITAL | BSS0122 | Revital Health | KENYA | Initial 60 days, thereafter 21 | 84 | USD | 10 Boxes | YES | 100/bOX | No automated retractable or similar safety mechanism | NO TEST REPORTS | Re-tender
42142609-00027 | 150 | Syringe, plastic, latex free, non-pyrogenic and DEPH free. Nominal capacity: 1ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. With attachable needle 22G(0,7mm) x 32mm, colour code: black and must have an automated retractable or similar safety mechanism Sterile, Single use, individually peel packed To comply with the latest issue of SANS 1124-1 and 1124-2 UNIT: Box | N/A | PROMED TECHNOLOGIES (PTY) LTD | 1.67 | U-LIFE | 502027 | JIAXING | CHINA | 14 DAYS | 85% | USD | 1 BOX | YES | BOX | No samples submitted for evaluation | NO TEST REPORTS | Re-tender
42142609-00027 | 150 | Syringe, plastic, latex free, non-pyrogenic and DEPH free. Nominal capacity: 1ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. With attachable needle 22G(0,7mm) x 32mm, colour code: black and must have an automated retractable or similar safety mechanism Sterile, Single use, individually peel packed To comply with the latest issue of SANS 1124-1 and 1124-2 UNIT: Box |  | GIBELA TRADE AND INVEST 1001 CC | 1.75 | TARA | SYR062 | Shandong Qiaopai Group Co., Ltd | CHINA | 21 | 85 | USD | 10 | YES | box 100 | No sample submitted for evaluation | NO TEST REPORTS | Re-tender
42142609-00028 | 151 | Syringe, plastic, latex free, non-pyrogenic and DEPH free. Nominal capacity: 1ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. With attachable needle 23G(0,6mm) x 25mm,colour code: deep blue and must have an automated retractable or similar safety mechanism Sterile, Single use, individually peel packed To comply with the latest issue of SANS 1124-1 and 1124-2 UNIT: Box |  | SUPRA HEALTHCARE JOHANNESBURG PTY LTD | 1.06 | REVITAL | BSS0123 | Revital Health | KENYA | Initial 60 days, thereafter 21 | 84 | USD | 10 Boxes | YES | 100/bOX | No automated retractable or similar safety mechanism | NO TEST REPORTS | Re-tender
42142609-00028 | 151 | Syringe, plastic, latex free, non-pyrogenic and DEPH free. Nominal capacity: 1ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. With attachable needle 23G(0,6mm) x 25mm,colour code: deep blue and must have an automated retractable or similar safety mechanism Sterile, Single use, individually peel packed To comply with the latest issue of SANS 1124-1 and 1124-2 UNIT: Box | N/A | PROMED TECHNOLOGIES (PTY) LTD | 1.67 | U-LIFE |  | JIAXING | CHINA | 15 DAYS | 85% | USD | 2 BOX | YES | BOX | No samples submitted for evaluation | NO TEST REPORTS | Re-tender
42142609-00028 | 151 | Syringe, plastic, latex free, non-pyrogenic and DEPH free. Nominal capacity: 1ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. With attachable needle 23G(0,6mm) x 25mm,colour code: deep blue and must have an automated retractable or similar safety mechanism Sterile, Single use, individually peel packed To comply with the latest issue of SANS 1124-1 and 1124-2 UNIT: Box |  | GIBELA TRADE AND INVEST 1001 CC | 1.75 | TARA | SYR063 | Shandong Qiaopai Group Co., Ltd | CHINA | 21 | 85 | USD | 10 | YES | box 100 | No sample submitted for evaluation | NO TEST REPORTS | Re-tender
42142609-00029 | 152 | Syringe, plastic, latex free, non-pyrogenic and DEPH free. Nominal capacity: 1ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. With attachable needle 24G(0.55mm) x 20mm, colour code: medium purple and must have an automated retractable or similar safety mechanism Sterile, Single use, individually peel packed To comply with the latest issue of SANS 1124-1 and 1124-2 UNIT: Box |  | SUPRA HEALTHCARE JOHANNESBURG PTY LTD | 1.06 | REVITAL | BSS0124 | Revital Health | KENYA | Initial 60 days, thereafter 21 | 84 | USD | 10 Boxes | YES | 100/bOX | No automated retractable or similar safety mechanism | NO TEST REPORTS | Re-tender
42142609-00029 | 152 | Syringe, plastic, latex free, non-pyrogenic and DEPH free. Nominal capacity: 1ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. With attachable needle 24G(0.55mm) x 20mm, colour code: medium purple and must have an automated retractable or similar safety mechanism Sterile, Single use, individually peel packed To comply with the latest issue of SANS 1124-1 and 1124-2 UNIT: Box |  | GIBELA TRADE AND INVEST 1001 CC | 1.75 | TARA | SYR064 | Shandong Qiaopai Group Co., Ltd | CHINA | 21 | 85 | USD | 10 | YES | box 100 | No sample submitted for evaluation | NO TEST REPORTS | Re-tender
42142609-00029 | 152 | Syringe, plastic, latex free, non-pyrogenic and DEPH free. Nominal capacity: 1ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. With attachable needle 24G(0.55mm) x 20mm, colour code: medium purple and must have an automated retractable or similar safety mechanism Sterile, Single use, individually peel packed To comply with the latest issue of SANS 1124-1 and 1124-2 UNIT: Box | N/A | SAFARMEX (PTY) LTD | 2.55 | Saclin | To be confirmed | Zhejiang Mediunion Healthcare Group Co. Ltd | China | 3days | 80% | US$ | 3000 Boxes | Yes | Box of 100 | No sample submitted for evaluation | NO TEST REPORT | Re-tender
42142609-00030 | 153 | Syringe, plastic, latex free, non-pyrogenic and DEPH free. Nominal capacity: 1ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. With attachable needle 25G(0,50mm) x 16mm, colour code: orange and must have an automated retractable or similar safety mechanism Sterile, Single use, individually peel packed To comply with the latest issue of SANS 1124-1 and 1124-2 UNIT: Box |  | SUPRA HEALTHCARE JOHANNESBURG PTY LTD | 1.06 | REVITAL | BSS0125 | Revital Health | KENYA | Initial 60 days, thereafter 21 | 84 | USD | 10 Boxes | YES | 100/bOX | No automated retractable or similar safety mechanism | NO TEST REPORTS | Re-tender
42142609-00030 | 153 | Syringe, plastic, latex free, non-pyrogenic and DEPH free. Nominal capacity: 1ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. With attachable needle 25G(0,50mm) x 16mm, colour code: orange and must have an automated retractable or similar safety mechanism Sterile, Single use, individually peel packed To comply with the latest issue of SANS 1124-1 and 1124-2 UNIT: Box |  | GIBELA TRADE AND INVEST 1001 CC | 1.75 | TARA | SYR065 | Shandong Qiaopai Group Co., Ltd | CHINA | 21 | 85 | USD | 10 | YES | box 100 | No sample submitted for evaluation | NO TEST REPORTS | Re-tender
42142609-00030 | 153 | Syringe, plastic, latex free, non-pyrogenic and DEPH free. Nominal capacity: 1ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. With attachable needle 25G(0,50mm) x 16mm, colour code: orange and must have an automated retractable or similar safety mechanism Sterile, Single use, individually peel packed To comply with the latest issue of SANS 1124-1 and 1124-2 UNIT: Box | N/A | SAFARMEX (PTY) LTD | 2.55 | Saclin | To be confirmed | Zhejiang Mediunion Healthcare Group Co. Ltd | China | 3days | 80% | US$ | 3000 Boxes | Yes | Box of 100 | No sample submitted for evaluation | NO TEST REPORT | Re-tender
42142609-00030 | 153 | Syringe, plastic, latex free, non-pyrogenic and DEPH free. Nominal capacity: 1ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. With attachable needle 25G(0,50mm) x 16mm, colour code: orange and must have an automated retractable or similar safety mechanism Sterile, Single use, individually peel packed To comply with the latest issue of SANS 1124-1 and 1124-2 UNIT: Box |  | Chemical World cc | 209 | WEALY | SYRS | WEALY | china | 3 | 85 | USD | 18000 | YES | 100 | COMPLY pending confirmation of latex and DEHP free | NOT COMPLY | Re-tender
42142609-00031 | 154 | Syringe, plastic, latex free, non-pyrogenic and DEPH free. Nominal capacity: 1ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. With attachable needle 26G(0.45mm) x 16mm, colour code: brown an automated retractable or similar safety mechanism Sterile, Single use, individually peel packed To comply with the latest issue of SANS 1124-1 and 1124-2 UNIT: Box |  | SUPRA HEALTHCARE JOHANNESBURG PTY LTD | 1.06 | REVITAL | BSS0126 | Revital Health | KENYA | Initial 60 days, thereafter 21 | 84 | USD | 10 Boxes | YES | 100/bOX | No automated retractable or similar safety mechanism | NO TEST REPORTS | Re-tender
42142609-00031 | 154 | Syringe, plastic, latex free, non-pyrogenic and DEPH free. Nominal capacity: 1ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. With attachable needle 26G(0.45mm) x 16mm, colour code: brown an automated retractable or similar safety mechanism Sterile, Single use, individually peel packed To comply with the latest issue of SANS 1124-1 and 1124-2 UNIT: Box |  | GIBELA TRADE AND INVEST 1001 CC | 1.75 | TARA | SYR066 | Shandong Qiaopai Group Co., Ltd | CHINA | 21 | 85 | USD | 10 | YES | box 100 | No sample submitted for evaluation | NO TEST REPORTS | Re-tender
42142609-00031 | 154 | Syringe, plastic, latex free, non-pyrogenic and DEPH free. Nominal capacity: 1ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. With attachable needle 26G(0.45mm) x 16mm, colour code: brown an automated retractable or similar safety mechanism Sterile, Single use, individually peel packed To comply with the latest issue of SANS 1124-1 and 1124-2 UNIT: Box | N/A | SAFARMEX (PTY) LTD | 2.55 | Saclin | To be confirmed | Zhejiang Mediunion Healthcare Group Co. Ltd | China | 3days | 80% | US$ | 3000 Boxes | Yes | Box of 100 | No sample submitted for evaluation | NO TEST REPORT | Re-tender
42142609-00032 | 155 | Syringe, plastic, latex free, non-pyrogenic and DEPH free. Nominal capacity: 1ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. With attachable needle 27G(0.4mm) x 8mm, colour code: medium grey and must have an automated retractable or similar safety mechanism Sterile, Single use, individually peel packed To comply with the latest issue of SANS 1124-1 and 1124-2 UNIT: Box |  | SUPRA HEALTHCARE JOHANNESBURG PTY LTD | 1.06 | REVITAL | BSS0127 | Revital Health | KENYA | Initial 60 days, thereafter 21 | 84 | USD | 10 Boxes | YES | 100/bOX | No automated retractable or similar safety mechanism | NO TEST REPORTS | Re-tender
42142609-00032 | 155 | Syringe, plastic, latex free, non-pyrogenic and DEPH free. Nominal capacity: 1ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. With attachable needle 27G(0.4mm) x 8mm, colour code: medium grey and must have an automated retractable or similar safety mechanism Sterile, Single use, individually peel packed To comply with the latest issue of SANS 1124-1 and 1124-2 UNIT: Box | N/A | SAFARMEX (PTY) LTD | 2.55 | Saclin | To be confirmed | Zhejiang Mediunion Healthcare Group Co. Ltd | China | 3days | 80% | US$ | 3000 Boxes | Yes | Box of 100 | No sample submitted for evaluation | NO TEST REPORT | Re-tender
42142609-00033 | 156 | Syringe, plastic, latex free, non-pyrogenic and DEPH free. Nominal capacity: 1ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. With attachable needle 29G(0.33mm) x 8mm, colour code: red and must have an automated retractable or similar safety mechanism Sterile, Single use, individually peel packed To comply with the latest issue of SANS 1124-1 and 1124-2 UNIT: Box |  | GIBELA TRADE AND INVEST 1001 CC | 0.66 | TARA | SYR067 | Shandong Qiaopai Group Co., Ltd | CHINA | 21 | 85 | USD | 10 | YES | box 100 | No sample submitted for evaluation | NO TEST REPORTS | Re-tender
42142609-00033 | 156 | Syringe, plastic, latex free, non-pyrogenic and DEPH free. Nominal capacity: 1ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. With attachable needle 29G(0.33mm) x 8mm, colour code: red and must have an automated retractable or similar safety mechanism Sterile, Single use, individually peel packed To comply with the latest issue of SANS 1124-1 and 1124-2 UNIT: Box |  | SUPRA HEALTHCARE JOHANNESBURG PTY LTD | 1.06 | REVITAL | BSS0129 | Revital Health | KENYA | Initial 60 days, thereafter 21 | 84 | USD | 10 Boxes | YES | 100/bOX | No automated retractable or similar safety mechanism | NO TEST REPORTS | Re-tender
42142609-00033 | 156 | Syringe, plastic, latex free, non-pyrogenic and DEPH free. Nominal capacity: 1ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. With attachable needle 29G(0.33mm) x 8mm, colour code: red and must have an automated retractable or similar safety mechanism Sterile, Single use, individually peel packed To comply with the latest issue of SANS 1124-1 and 1124-2 UNIT: Box | N/A | SAFARMEX (PTY) LTD | 1.9 | Saclin | To be confirmed | Zhejiang Mediunion Healthcare Group Co. Ltd | China | 3days | 80% | US$ | 3000 Boxes | Yes | Box of 100 | No sample submitted for evaluation | NO TEST REPORT | Re-tender
42142609-00033 | 156 | Syringe, plastic, latex free, non-pyrogenic and DEPH free. Nominal capacity: 1ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. With attachable needle 29G(0.33mm) x 8mm, colour code: red and must have an automated retractable or similar safety mechanism Sterile, Single use, individually peel packed To comply with the latest issue of SANS 1124-1 and 1124-2 UNIT: Box |  | Chemical World cc | 2.14 | WEALY | SYRS | WEALY | china | 3 | 85 | USD | 18000 | YES | 100 | length of 12mm instead of 8mm | NO TEST REPORTS | Re-tender
42142609-00034 | 157 | Syringe, plastic, latex free, non-pyrogenic and DEPH free. Nominal capacity: 1ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. With attachable needle 30G(0.3mm) x 8mm, colour code: yellow and must have an automated retractable or similar safety mechanism Sterile, Single use, individually peel packed To comply with the latest issue of SANS 1124-1 and 1124-2 UNIT: Box |  | GIBELA TRADE AND INVEST 1001 CC | 0.66 | TARA | SYR068 | Shandong Qiaopai Group Co., Ltd | CHINA | 21 | 85 | USD | 10 | YES | box 100 | No sample submitted for evaluation | NO TEST REPORTS | Re-tender
42142609-00034 | 157 | Syringe, plastic, latex free, non-pyrogenic and DEPH free. Nominal capacity: 1ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. With attachable needle 30G(0.3mm) x 8mm, colour code: yellow and must have an automated retractable or similar safety mechanism Sterile, Single use, individually peel packed To comply with the latest issue of SANS 1124-1 and 1124-2 UNIT: Box |  | SUPRA HEALTHCARE JOHANNESBURG PTY LTD | 1.06 | REVITAL | BSS0130 | Revital Health | KENYA | Initial 60 days, thereafter 21 | 84 | USD | 10 Boxes | YES | 100/bOX | No automated retractable or similar safety mechanism | NO TEST REPORTS | Re-tender
42142609-00034 | 157 | Syringe, plastic, latex free, non-pyrogenic and DEPH free. Nominal capacity: 1ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. With attachable needle 30G(0.3mm) x 8mm, colour code: yellow and must have an automated retractable or similar safety mechanism Sterile, Single use, individually peel packed To comply with the latest issue of SANS 1124-1 and 1124-2 UNIT: Box | N/A | SAFARMEX (PTY) LTD | 1.9 | Saclin | To be confirmed | Zhejiang Mediunion Healthcare Group Co. Ltd | China | 3days | 80% | US$ | 3000 Boxes | Yes | Box of 100 | No sample submitted for evaluation | NO TEST REPORT | Re-tender
42142609-00035 | 158 | Syringe, plastic, latex free, non-pyrogenic and DEPH free. Nominal capacity: 3ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. With attachable needle 20G(0,9mm) x 40mm, colour code: yellow and must have an automated retractable or similar safety mechanism Sterile, Single use, individually peel packed To comply with the latest issue of SANS 1124-1 and 1124-2 UNIT: Box |  | GIBELA TRADE AND INVEST 1001 CC | 1.71 | TARA | SYR069 | Shandong Qiaopai Group Co., Ltd | CHINA | 21 | 85 | USD | 10 | YES | box 100 | No sample submitted for evaluation | NO TEST REPORTS | Re-tender
42142609-00035 | 158 | Syringe, plastic, latex free, non-pyrogenic and DEPH free. Nominal capacity: 3ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. With attachable needle 20G(0,9mm) x 40mm, colour code: yellow and must have an automated retractable or similar safety mechanism Sterile, Single use, individually peel packed To comply with the latest issue of SANS 1124-1 and 1124-2 UNIT: Box | N/A | SAFARMEX (PTY) LTD | 2.49 | Saclin | To be confirmed | Zhejiang Mediunion Healthcare Group Co. Ltd | China | 3days | 80% | US$ | 3000 Pieces | Yes | Box of 100 | No sample submitted for evaluation | NO TEST REPORT | Re-tender
42142609-00035 | 158 | Syringe, plastic, latex free, non-pyrogenic and DEPH free. Nominal capacity: 3ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. With attachable needle 20G(0,9mm) x 40mm, colour code: yellow and must have an automated retractable or similar safety mechanism Sterile, Single use, individually peel packed To comply with the latest issue of SANS 1124-1 and 1124-2 UNIT: Box |  | SUPRA HEALTHCARE JOHANNESBURG PTY LTD | 2.97 | Avacare | BSS0320 | HENAN AILE INDUSTRIAL CO.,LTD | CHINA | Initial 60 days, thereafter 21 | 84 | USD | 10 Boxes | YES | 100 pcs/box | No sample submitted for evaluation | NO TEST REPORTS | Re-tender
42142609-00036 | 159 | Syringe, plastic, latex free, non-pyrogenic and DEPH free. Nominal capacity: 3ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. With attachable needle 21G(0,8mm) x 40mm, colour code: deep green and must have an automated retractable or similar safety mechanism Sterile, Single use, individually peel packed To comply with the latest issue of SANS 1124-1 and 1124-2 UNIT: Box |  | GIBELA TRADE AND INVEST 1001 CC | 1.71 | TARA | SYR070 | Shandong Qiaopai Group Co., Ltd | CHINA | 21 | 85 | USD | 10 | YES | box 100 | No sample submitted for evaluation | NO TEST REPORTS | Re-tender
42142609-00036 | 159 | Syringe, plastic, latex free, non-pyrogenic and DEPH free. Nominal capacity: 3ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. With attachable needle 21G(0,8mm) x 40mm, colour code: deep green and must have an automated retractable or similar safety mechanism Sterile, Single use, individually peel packed To comply with the latest issue of SANS 1124-1 and 1124-2 UNIT: Box |  | Chemical World cc | 2.24 | WEAL+F204:O204Y |  |  | china | 3 | 85 | USD | 18000 | YES | 100 | COMPLY pending confirmation of latex and DEHP free | NOT COMPLY | Re-tender
42142609-00036 | 159 | Syringe, plastic, latex free, non-pyrogenic and DEPH free. Nominal capacity: 3ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. With attachable needle 21G(0,8mm) x 40mm, colour code: deep green and must have an automated retractable or similar safety mechanism Sterile, Single use, individually peel packed To comply with the latest issue of SANS 1124-1 and 1124-2 UNIT: Box | N/A | SAFARMEX (PTY) LTD | 2.49 | Saclin | To be confirmed | Zhejiang Mediunion Healthcare Group Co. Ltd | China | 3days | 80% | US$ | 3000 Pieces | Yes | Box of 100 | No sample submitted for evaluation | NO TEST REPORT | Re-tender
42142609-00036 | 159 | Syringe, plastic, latex free, non-pyrogenic and DEPH free. Nominal capacity: 3ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. With attachable needle 21G(0,8mm) x 40mm, colour code: deep green and must have an automated retractable or similar safety mechanism Sterile, Single use, individually peel packed To comply with the latest issue of SANS 1124-1 and 1124-2 UNIT: Box |  | SUPRA HEALTHCARE JOHANNESBURG PTY LTD | 2.97 | Avacare | BSS0321 | HENAN AILE INDUSTRIAL CO.,LTD | CHINA | Initial 60 days, thereafter 21 | 84 | USD | 10 Boxes | YES | 100 pcs/box | Offered nominal capacity of 2ml instead of 3ml | NO TEST REPORTS | Re-tender
42142609-00037 | 160 | Syringe, plastic, latex free, non-pyrogenic and DEPH free. Nominal capacity: 3ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. With attachable needle 22G(0,7mm) x 32mm, colour code: black and must have an automated retractable or similar safety mechanism Sterile, Single use, individually peel packed To comply with the latest issue of SANS 1124-1 and 1124-2 UNIT: Box | Basic Line offer | Becton Dickinson | 0.85 | 3ml 22G x 25mm Emerald Saf T | 305974 | Becton Dickinson | Brazil | 7 Working Days | 76% | USD | 1 box | YES | 100 units in a box | No sample submitted for evaluation. | NO TEST REPORTS | Re-tender
42142609-00037 | 160 | Syringe, plastic, latex free, non-pyrogenic and DEPH free. Nominal capacity: 3ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. With attachable needle 22G(0,7mm) x 32mm, colour code: black and must have an automated retractable or similar safety mechanism Sterile, Single use, individually peel packed To comply with the latest issue of SANS 1124-1 and 1124-2 UNIT: Box |  | ALMIKA TRADING PTY LTD | 1.02 | BD | 305974 | BD | USA | 14 | 0 | NO | 100 | YES | 1 | Re-tender: End-users to revise specifications | NO TEST REPORTS | Re-tender
42142609-00037 | 160 | Syringe, plastic, latex free, non-pyrogenic and DEPH free. Nominal capacity: 3ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. With attachable needle 22G(0,7mm) x 32mm, colour code: black and must have an automated retractable or similar safety mechanism Sterile, Single use, individually peel packed To comply with the latest issue of SANS 1124-1 and 1124-2 UNIT: Box |  | GIBELA TRADE AND INVEST 1001 CC | 1.71 | TARA | SYR071 | Shandong Qiaopai Group Co., Ltd | CHINA | 21 | 85 | USD | 10 | YES | box 100 | No sample submitted for evaluation | NO TEST REPORTS | Re-tender
42142609-00037 | 160 | Syringe, plastic, latex free, non-pyrogenic and DEPH free. Nominal capacity: 3ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. With attachable needle 22G(0,7mm) x 32mm, colour code: black and must have an automated retractable or similar safety mechanism Sterile, Single use, individually peel packed To comply with the latest issue of SANS 1124-1 and 1124-2 UNIT: Box | N/A | SAFARMEX (PTY) LTD | 2.49 | Saclin | To be confirmed | Zhejiang Mediunion Healthcare Group Co. Ltd | China | 3days | 80% | US$ | 3000 Boxes | Yes | Box of 100 | No sample submitted for evaluation | NO TEST REPORT | Re-tender
42142609-00037 | 160 | Syringe, plastic, latex free, non-pyrogenic and DEPH free. Nominal capacity: 3ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. With attachable needle 22G(0,7mm) x 32mm, colour code: black and must have an automated retractable or similar safety mechanism Sterile, Single use, individually peel packed To comply with the latest issue of SANS 1124-1 and 1124-2 UNIT: Box |  | SUPRA HEALTHCARE JOHANNESBURG PTY LTD | 2.97 | Avacare | BSS0322 | HENAN AILE INDUSTRIAL CO.,LTD | CHINA | Initial 60 days, thereafter 21 | 84 | USD | 10 Boxes | YES | 100 pcs/box | No sample submitted for evaluation | NO TEST REPORTS | Re-tender
42142609-00038 | 161 | Syringe, plastic, latex free, non-pyrogenic and DEPH free. Nominal capacity: 3ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. With attachable needle 23G(0,6mm) x 25mm,colour code: deep blue and must have an automated retractable or similar safety mechanism Sterile, Single use, individually peel packed To comply with the latest issue of SANS 1124-1 and 1124-2 UNIT: Box | Basic Line offer | Becton Dickinson | 0.85 | 3ml 23G x 25mm Emerald Saf T | 305976 | Becton Dickinson | Brazil | 7 Working Days | 76% | USD | 1 box | YES | 100 units in a box | No sample submitted for evaluation. | NO TEST REPORTS | Re-tender
42142609-00038 | 161 | Syringe, plastic, latex free, non-pyrogenic and DEPH free. Nominal capacity: 3ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. With attachable needle 23G(0,6mm) x 25mm,colour code: deep blue and must have an automated retractable or similar safety mechanism Sterile, Single use, individually peel packed To comply with the latest issue of SANS 1124-1 and 1124-2 UNIT: Box |  | ALMIKA TRADING PTY LTD | 1.02 | BD | 305976 | BD | USA | 14 | 0 | NO | 100 | YES | 1 | Re-tender: End-users to revise specifications | NO TEST REPORTS | Re-tender
42142609-00038 | 161 | Syringe, plastic, latex free, non-pyrogenic and DEPH free. Nominal capacity: 3ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. With attachable needle 23G(0,6mm) x 25mm,colour code: deep blue and must have an automated retractable or similar safety mechanism Sterile, Single use, individually peel packed To comply with the latest issue of SANS 1124-1 and 1124-2 UNIT: Box |  | GIBELA TRADE AND INVEST 1001 CC | 1.71 | TARA | SYR072 | Shandong Qiaopai Group Co., Ltd | CHINA | 21 | 85 | USD | 10 | YES | box 100 | No sample submitted for evaluation | NO TEST REPORTS | Re-tender
42142609-00038 | 161 | Syringe, plastic, latex free, non-pyrogenic and DEPH free. Nominal capacity: 3ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. With attachable needle 23G(0,6mm) x 25mm,colour code: deep blue and must have an automated retractable or similar safety mechanism Sterile, Single use, individually peel packed To comply with the latest issue of SANS 1124-1 and 1124-2 UNIT: Box | N/A | SAFARMEX (PTY) LTD | 2.49 | Saclin | To be confirmed | Zhejiang Mediunion Healthcare Group Co. Ltd | China | 3days | 80% | US$ | 3000 Boxes | Yes | Box of 100 | No sample submitted for evaluation | NO TEST REPORT | Re-tender
42142609-00038 | 161 | Syringe, plastic, latex free, non-pyrogenic and DEPH free. Nominal capacity: 3ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. With attachable needle 23G(0,6mm) x 25mm,colour code: deep blue and must have an automated retractable or similar safety mechanism Sterile, Single use, individually peel packed To comply with the latest issue of SANS 1124-1 and 1124-2 UNIT: Box |  | SUPRA HEALTHCARE JOHANNESBURG PTY LTD | 2.97 | Avacare | BSS0323 | HENAN AILE INDUSTRIAL CO.,LTD | CHINA | Initial 60 days, thereafter 21 | 84 | USD | 10 Boxes | YES | 100 pcs/box | No sample submitted for evaluation | NO TEST REPORTS | Re-tender
42142609-00038 | 161 | Syringe, plastic, latex free, non-pyrogenic and DEPH free. Nominal capacity: 3ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. With attachable needle 23G(0,6mm) x 25mm,colour code: deep blue and must have an automated retractable or similar safety mechanism Sterile, Single use, individually peel packed To comply with the latest issue of SANS 1124-1 and 1124-2 UNIT: Box | 1 | Biologica Pharmaceuticals Pty Ltd | 4.07 | Sol- M | 100077IM | Sol- Millennium | China | 4-6 Weeks | 100% | USD | 1000 Boxes | 100% | N/A |  | NO TEST REPORTS | Re-tender
42142609-00039 | 162 | Syringe, plastic, latex free, non-pyrogenic and DEPH free. Nominal capacity: 3ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. With attachable needle 24G(0.55mm) x 20mm, colour code: medium purple and must have an automated retractable or similar safety mechanism Sterile, Single use, individually peel packed To comply with the latest issue of SANS 1124-1 and 1124-2 UNIT: Box |  | GIBELA TRADE AND INVEST 1001 CC | 1.71 | TARA | SYR073 | Shandong Qiaopai Group Co., Ltd | CHINA | 21 | 85 | USD | 10 | YES | box 100 | No sample submitted for evaluation | NO TEST REPORTS | Re-tender
42142609-00039 | 162 | Syringe, plastic, latex free, non-pyrogenic and DEPH free. Nominal capacity: 3ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. With attachable needle 24G(0.55mm) x 20mm, colour code: medium purple and must have an automated retractable or similar safety mechanism Sterile, Single use, individually peel packed To comply with the latest issue of SANS 1124-1 and 1124-2 UNIT: Box | N/A | SAFARMEX (PTY) LTD | 2.49 | Saclin | To be confirmed | Zhejiang Mediunion Healthcare Group Co. Ltd | China | 3days | 80% | US$ | 3000 Boxes | Yes | Box of 100 | No sample submitted for evaluation | NO TEST REPORT | Re-tender
42142609-00039 | 162 | Syringe, plastic, latex free, non-pyrogenic and DEPH free. Nominal capacity: 3ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. With attachable needle 24G(0.55mm) x 20mm, colour code: medium purple and must have an automated retractable or similar safety mechanism Sterile, Single use, individually peel packed To comply with the latest issue of SANS 1124-1 and 1124-2 UNIT: Box |  | SUPRA HEALTHCARE JOHANNESBURG PTY LTD | 2.97 | Avacare | BSS0324 | HENAN AILE INDUSTRIAL CO.,LTD | CHINA | Initial 60 days, thereafter 21 | 84 | USD | 10 Boxes | YES | 100 pcs/box | No sample submitted for evaluation | NO TEST REPORTS | Re-tender
42142609-00040 | 163 | Syringe, plastic, latex free, non-pyrogenic and DEPH free. Nominal capacity: 3ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. With attachable needle 25G(0,50mm) x 16mm, colour code: orange and must have an automated retractable or similar safety mechanism Sterile, Single use, individually peel packed To comply with the latest issue of SANS 1124-1 and 1124-2 UNIT: Box |  | GIBELA TRADE AND INVEST 1001 CC | 1.71 | TARA | SYR074 | Shandong Qiaopai Group Co., Ltd | CHINA | 21 | 85 | USD | 10 | YES | box 100 | No sample submitted for evaluation | NO TEST REPORTS | Re-tender
42142609-00040 | 163 | Syringe, plastic, latex free, non-pyrogenic and DEPH free. Nominal capacity: 3ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. With attachable needle 25G(0,50mm) x 16mm, colour code: orange and must have an automated retractable or similar safety mechanism Sterile, Single use, individually peel packed To comply with the latest issue of SANS 1124-1 and 1124-2 UNIT: Box | N/A | SAFARMEX (PTY) LTD | 2.49 | Saclin | To be confirmed | Zhejiang Mediunion Healthcare Group Co. Ltd | China | 3days | 80% | US$ | 3000 Boxes | Yes | Box of 100 | No sample submitted for evaluation | NO TEST REPORT | Re-tender
42142609-00040 | 163 | Syringe, plastic, latex free, non-pyrogenic and DEPH free. Nominal capacity: 3ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. With attachable needle 25G(0,50mm) x 16mm, colour code: orange and must have an automated retractable or similar safety mechanism Sterile, Single use, individually peel packed To comply with the latest issue of SANS 1124-1 and 1124-2 UNIT: Box |  | SUPRA HEALTHCARE JOHANNESBURG PTY LTD | 2.97 | Avacare | BSS0325 | HENAN AILE INDUSTRIAL CO.,LTD | CHINA | Initial 60 days, thereafter 21 | 84 | USD | 10 Boxes | YES | 100 pcs/box | No sample submitted for evaluation | NO TEST REPORTS | Re-tender
42142609-00041 | 164 | Syringe, plastic, latex free, non-pyrogenic and DEPH free. Nominal capacity: 3ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. With attachable needle 26G(0.45mm) x 16mm, colour code: brown an automated retractable or similar safety mechanism Sterile, Single use, individually peel packed To comply with the latest issue of SANS 1124-1 and 1124-2 UNIT: Box |  | GIBELA TRADE AND INVEST 1001 CC | 1.71 | TARA | SYR075 | Shandong Qiaopai Group Co., Ltd | CHINA | 21 | 85 | USD | 10 | YES | box 100 | No sample submitted for evaluation | NO TEST REPORTS | Re-tender
42142609-00041 | 164 | Syringe, plastic, latex free, non-pyrogenic and DEPH free. Nominal capacity: 3ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. With attachable needle 26G(0.45mm) x 16mm, colour code: brown an automated retractable or similar safety mechanism Sterile, Single use, individually peel packed To comply with the latest issue of SANS 1124-1 and 1124-2 UNIT: Box | N/A | SAFARMEX (PTY) LTD | 2.49 | Saclin | To be confirmed | Zhejiang Mediunion Healthcare Group Co. Ltd | China | 3days | 80% | US$ | 3000 Boxes | Yes | Box of 100 | No sample submitted for evaluation | NO TEST REPORT | Re-tender
42142609-00041 | 164 | Syringe, plastic, latex free, non-pyrogenic and DEPH free. Nominal capacity: 3ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. With attachable needle 26G(0.45mm) x 16mm, colour code: brown an automated retractable or similar safety mechanism Sterile, Single use, individually peel packed To comply with the latest issue of SANS 1124-1 and 1124-2 UNIT: Box |  | SUPRA HEALTHCARE JOHANNESBURG PTY LTD | 2.97 | Avacare | BSS0326 | HENAN AILE INDUSTRIAL CO.,LTD | CHINA | Initial 60 days, thereafter 21 | 84 | USD | 10 Boxes | YES | 100 pcs/box | No sample submitted for evaluation | NO TEST REPORTS | Re-tender
42142609-00042 | 165 | Syringe, plastic, latex free, non-pyrogenic and DEPH free. Nominal capacity: 3ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. With attachable needle 27G(0.4mm) x 8mm, colour code: medium grey and must have an automated retractable or similar safety mechanism Sterile, Single use, individually peel packed To comply with the latest issue of SANS 1124-1 and 1124-2 UNIT: Box | N/A | SAFARMEX (PTY) LTD | 2.49 | Saclin | To be confirmed | Zhejiang Mediunion Healthcare Group Co. Ltd | China | 3days | 80% | US$ | 3000 Boxes | Yes | Box of 100 | No sample submitted for evaluation | NO TEST REPORT | Re-tender
42142609-00042 | 165 | Syringe, plastic, latex free, non-pyrogenic and DEPH free. Nominal capacity: 3ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. With attachable needle 27G(0.4mm) x 8mm, colour code: medium grey and must have an automated retractable or similar safety mechanism Sterile, Single use, individually peel packed To comply with the latest issue of SANS 1124-1 and 1124-2 UNIT: Box |  | SUPRA HEALTHCARE JOHANNESBURG PTY LTD | 2.97 | Avacare | BSS0327 | HENAN AILE INDUSTRIAL CO.,LTD | CHINA | Initial 60 days, thereafter 21 | 84 | USD | 10 Boxes | YES | 100 pcs/box | No sample submitted for evaluation | NO TEST REPORTS | Re-tender
42142609-00044 | 167 | Syringe, plastic, latex free, non-pyrogenic and DEPH free. Nominal capacity: 3ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. With attachable needle 30G(0.3mm) x 8mm, colour code: yellow and must have an automated retractable or similar safety mechanism Sterile, Single use, individually peel packed To comply with the latest issue of SANS 1124-1 and 1124-2 UNIT: Box |  | GIBELA TRADE AND INVEST 1001 CC | 1.71 | TARA | SYR076 | Shandong Qiaopai Group Co., Ltd | CHINA | 21 | 85 | USD | 10 | YES | box 100 | No sample submitted for evaluation | NO TEST REPORTS | Re-tender
42142609-00045 | 168 | Syringe, plastic, latex free, non-pyrogenic and DEPH free. Nominal capacity: 5ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. With attachable needle 20G(0,9mm) x 40mm, colour code: yellow and must have an automated retractable or similar safety mechanism Sterile, Single use, individually peel packed To comply with the latest issue of SANS 1124-1 and 1124-2 UNIT: Box |  | SUPRA HEALTHCARE JOHANNESBURG PTY LTD | 1.15 | REVITAL | BSS0320 | Revital Health | KENYA | Initial 60 days, thereafter 21 | 84 | USD | 10 Boxes | YES | 100/bOX | No automated retractable or similar safety mechanism | NO TEST REPORTS | Re-tender
42142609-00045 | 168 | Syringe, plastic, latex free, non-pyrogenic and DEPH free. Nominal capacity: 5ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. With attachable needle 20G(0,9mm) x 40mm, colour code: yellow and must have an automated retractable or similar safety mechanism Sterile, Single use, individually peel packed To comply with the latest issue of SANS 1124-1 and 1124-2 UNIT: Box |  | GIBELA TRADE AND INVEST 1001 CC | 1.75 | TARA | SYR077 | Shandong Qiaopai Group Co., Ltd | CHINA | 21 | 85 | USD | 10 | YES | box 100 | No sample submitted for evaluation | NO TEST REPORTS | Re-tender
42142609-00045 | 168 | Syringe, plastic, latex free, non-pyrogenic and DEPH free. Nominal capacity: 5ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. With attachable needle 20G(0,9mm) x 40mm, colour code: yellow and must have an automated retractable or similar safety mechanism Sterile, Single use, individually peel packed To comply with the latest issue of SANS 1124-1 and 1124-2 UNIT: Box | N/A | SAFARMEX (PTY) LTD | 2.49 | Saclin | To be confirmed | Zhejiang Mediunion Healthcare Group Co. Ltd | China | 3days | 80% | US$ | 3000 Boxes | Yes | Box of 100 | No sample submitted for evaluation | NO TEST REPORT | Re-tender
42142609-00046 | 169 | Syringe, plastic, latex free, non-pyrogenic and DEPH free. Nominal capacity: 5ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. With attachable needle 21G(0,8mm) x 40mm, colour code: deep green and must have an automated retractable or similar safety mechanism Sterile, Single use, individually peel packed To comply with the latest issue of SANS 1124-1 and 1124-2 UNIT: Box |  | SUPRA HEALTHCARE JOHANNESBURG PTY LTD | 1.15 | REVITAL | BSS0521 | Revital Health | KENYA | Initial 60 days, thereafter 21 | 84 | USD | 10 Boxes | YES | 100/bOX | No automated retractable or similar safety mechanism | NO TEST REPORTS | Re-tender
42142609-00046 | 169 | Syringe, plastic, latex free, non-pyrogenic and DEPH free. Nominal capacity: 5ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. With attachable needle 21G(0,8mm) x 40mm, colour code: deep green and must have an automated retractable or similar safety mechanism Sterile, Single use, individually peel packed To comply with the latest issue of SANS 1124-1 and 1124-2 UNIT: Box |  | GIBELA TRADE AND INVEST 1001 CC | 1.75 | TARA | SYR078 | Shandong Qiaopai Group Co., Ltd | CHINA | 21 | 85 | USD | 10 | YES | box 100 | No sample submitted for evaluation | NO TEST REPORTS | Re-tender
42142609-00046 | 169 | Syringe, plastic, latex free, non-pyrogenic and DEPH free. Nominal capacity: 5ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. With attachable needle 21G(0,8mm) x 40mm, colour code: deep green and must have an automated retractable or similar safety mechanism Sterile, Single use, individually peel packed To comply with the latest issue of SANS 1124-1 and 1124-2 UNIT: Box | N/A | SAFARMEX (PTY) LTD | 2.49 | Saclin | To be confirmed | Zhejiang Mediunion Healthcare Group Co. Ltd | China | 3days | 80% | US$ | 3000 Boxes | Yes | Box of 100 | No sample submitted for evaluation | NO TEST REPORT | Re-tender
42142609-00047 | 170 | Syringe, plastic, latex free, non-pyrogenic and DEPH free. Nominal capacity: 5ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. With attachable needle 22G(0,7mm) x 32mm, colour code: black and must have an automated retractable or similar safety mechanism Sterile, Single use, individually peel packed To comply with the latest issue of SANS 1124-1 and 1124-2 UNIT: Box | Basic Line offer | Becton Dickinson | 1.15 | 5ml 22G x 25mm Emerald Saf T | 305982 | Becton Dickinson | Brazil | 7 Working Days | 76% | USD | 1 Box | YES | 100 units in a box | No sample submitted for evaluation. | NO TEST REPORTS | Re-tender
42142609-00047 | 170 | Syringe, plastic, latex free, non-pyrogenic and DEPH free. Nominal capacity: 5ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. With attachable needle 22G(0,7mm) x 32mm, colour code: black and must have an automated retractable or similar safety mechanism Sterile, Single use, individually peel packed To comply with the latest issue of SANS 1124-1 and 1124-2 UNIT: Box |  | SUPRA HEALTHCARE JOHANNESBURG PTY LTD | 1.15 | REVITAL | BSS0522 | Revital Health | KENYA | Initial 60 days, thereafter 21 | 84 | USD | 10 Boxes | YES | 100/bOX | No automated retractable or similar safety mechanism | NO TEST REPORTS | Re-tender
42142609-00047 | 170 | Syringe, plastic, latex free, non-pyrogenic and DEPH free. Nominal capacity: 5ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. With attachable needle 22G(0,7mm) x 32mm, colour code: black and must have an automated retractable or similar safety mechanism Sterile, Single use, individually peel packed To comply with the latest issue of SANS 1124-1 and 1124-2 UNIT: Box |  | ALMIKA TRADING PTY LTD | 1.32 | BD | 305982 | BD | USA | 14 | 0 | NO | 100 | YES | 1 | Re-tender: End-users to revise specifications | NO TEST REPORTS | Re-tender
42142609-00047 | 170 | Syringe, plastic, latex free, non-pyrogenic and DEPH free. Nominal capacity: 5ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. With attachable needle 22G(0,7mm) x 32mm, colour code: black and must have an automated retractable or similar safety mechanism Sterile, Single use, individually peel packed To comply with the latest issue of SANS 1124-1 and 1124-2 UNIT: Box |  | GIBELA TRADE AND INVEST 1001 CC | 1.75 | TARA | SYR079 | Shandong Qiaopai Group Co., Ltd | CHINA | 21 | 85 | USD | 10 | YES | box 100 | No sample submitted for evaluation | NO TEST REPORTS | Re-tender
42142609-00047 | 170 | Syringe, plastic, latex free, non-pyrogenic and DEPH free. Nominal capacity: 5ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. With attachable needle 22G(0,7mm) x 32mm, colour code: black and must have an automated retractable or similar safety mechanism Sterile, Single use, individually peel packed To comply with the latest issue of SANS 1124-1 and 1124-2 UNIT: Box | N/A | SAFARMEX (PTY) LTD | 2.49 | Saclin | To be confirmed | Zhejiang Mediunion Healthcare Group Co. Ltd | China | 3days | 80% | US$ | 3000 Boxes | Yes | Box of 100 | No sample submitted for evaluation | NO TEST REPORT | Re-tender
42142609-00047 | 170 | Syringe, plastic, latex free, non-pyrogenic and DEPH free. Nominal capacity: 5ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. With attachable needle 22G(0,7mm) x 32mm, colour code: black and must have an automated retractable or similar safety mechanism Sterile, Single use, individually peel packed To comply with the latest issue of SANS 1124-1 and 1124-2 UNIT: Box | 1 | Biologica Pharmaceuticals Pty Ltd | 4.41 | Sol- M | 140076IM | Sol- Millennium | China | 4-6 Weeks | 100% | USD | 1000 Boxes | 100% | N/A |  | NO TEST REPORTS | Re-tender
42142609-00048 | 171 | Syringe, plastic, latex free, non-pyrogenic and DEPH free. Nominal capacity: 5ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. With attachable needle 23G(0,6mm) x 25mm,colour code: deep blue and must have an automated retractable or similar safety mechanism Sterile, Single use, individually peel packed To comply with the latest issue of SANS 1124-1 and 1124-2 UNIT: Box |  | SUPRA HEALTHCARE JOHANNESBURG PTY LTD | 1.15 | REVITAL | BSS0523 | Revital Health | KENYA | Initial 60 days, thereafter 21 | 84 | USD | 10 Boxes | YES | 100/bOX | No automated retractable or similar safety mechanism | NO TEST REPORTS | Re-tender
42142609-00048 | 171 | Syringe, plastic, latex free, non-pyrogenic and DEPH free. Nominal capacity: 5ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. With attachable needle 23G(0,6mm) x 25mm,colour code: deep blue and must have an automated retractable or similar safety mechanism Sterile, Single use, individually peel packed To comply with the latest issue of SANS 1124-1 and 1124-2 UNIT: Box |  | GIBELA TRADE AND INVEST 1001 CC | 1.75 | TARA | SYR080 | Shandong Qiaopai Group Co., Ltd | CHINA | 21 | 85 | USD | 10 | YES | box 100 | No sample submitted for evaluation | NO TEST REPORTS | Re-tender
42142609-00048 | 171 | Syringe, plastic, latex free, non-pyrogenic and DEPH free. Nominal capacity: 5ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. With attachable needle 23G(0,6mm) x 25mm,colour code: deep blue and must have an automated retractable or similar safety mechanism Sterile, Single use, individually peel packed To comply with the latest issue of SANS 1124-1 and 1124-2 UNIT: Box |  | Chemical World cc | 2.32 | WEALY | SYRS | WEALY | china | 3 | 85 | USD | 18000 | YES | 100 | COMPLY pending confirmation of latex and DEHP free | NOT COMPLY | Re-tender
42142609-00048 | 171 | Syringe, plastic, latex free, non-pyrogenic and DEPH free. Nominal capacity: 5ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. With attachable needle 23G(0,6mm) x 25mm,colour code: deep blue and must have an automated retractable or similar safety mechanism Sterile, Single use, individually peel packed To comply with the latest issue of SANS 1124-1 and 1124-2 UNIT: Box | N/A | SAFARMEX (PTY) LTD | 2.49 | Saclin | To be confirmed | Zhejiang Mediunion Healthcare Group Co. Ltd | China | 3days | 80% | US$ | 3000 Boxes | Yes | Box of 100 | No sample submitted for evaluation | NO TEST REPORT | Re-tender
42142609-00048 | 171 | Syringe, plastic, latex free, non-pyrogenic and DEPH free. Nominal capacity: 5ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. With attachable needle 23G(0,6mm) x 25mm,colour code: deep blue and must have an automated retractable or similar safety mechanism Sterile, Single use, individually peel packed To comply with the latest issue of SANS 1124-1 and 1124-2 UNIT: Box | 1 | Biologica Pharmaceuticals Pty Ltd | 4.41 | Sol- M | 140077IM | Sol- Millennium | China | 4-6 Weeks | 100% | USD | 1000 Boxes | 100% | N/A |  | NO TEST REPORTS | Re-tender
42142609-00049 | 172 | Syringe, plastic, latex free, non-pyrogenic and DEPH free. Nominal capacity: 5ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. With attachable needle 24G(0.55mm) x 20mm, colour code: medium purple and must have an automated retractable or similar safety mechanism Sterile, Single use, individually peel packed To comply with the latest issue of SANS 1124-1 and 1124-2 UNIT: Box |  | SUPRA HEALTHCARE JOHANNESBURG PTY LTD | 1.15 | REVITAL | BSS0524 | Revital Health | KENYA | Initial 60 days, thereafter 21 | 84 | USD | 10 Boxes | YES | 100/bOX | No automated retractable or similar safety mechanism | NO TEST REPORTS | Re-tender
42142609-00049 | 172 | Syringe, plastic, latex free, non-pyrogenic and DEPH free. Nominal capacity: 5ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. With attachable needle 24G(0.55mm) x 20mm, colour code: medium purple and must have an automated retractable or similar safety mechanism Sterile, Single use, individually peel packed To comply with the latest issue of SANS 1124-1 and 1124-2 UNIT: Box |  | GIBELA TRADE AND INVEST 1001 CC | 1.75 | TARA | SYR081 | Shandong Qiaopai Group Co., Ltd | CHINA | 21 | 85 | USD | 10 | YES | box 100 | No sample submitted for evaluation | NO TEST REPORTS | Re-tender
42142609-00049 | 172 | Syringe, plastic, latex free, non-pyrogenic and DEPH free. Nominal capacity: 5ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. With attachable needle 24G(0.55mm) x 20mm, colour code: medium purple and must have an automated retractable or similar safety mechanism Sterile, Single use, individually peel packed To comply with the latest issue of SANS 1124-1 and 1124-2 UNIT: Box | N/A | SAFARMEX (PTY) LTD | 2.49 | Saclin | To be confirmed | Zhejiang Mediunion Healthcare Group Co. Ltd | China | 3days | 80% | US$ | 3000 Boxes | Yes | Box of 100 | No sample submitted for evaluation | NO TEST REPORT | Re-tender
42142609-00050 | 173 | Syringe, plastic, latex free, non-pyrogenic and DEPH free. Nominal capacity: 5ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. With attachable needle 25G(0,50mm) x 16mm, colour code: orange and must have an automated retractable or similar safety mechanism Sterile, Single use, individually peel packed To comply with the latest issue of SANS 1124-1 and 1124-2 UNIT: Box |  | SUPRA HEALTHCARE JOHANNESBURG PTY LTD | 1.15 | REVITAL | BSS0525 | Revital Health | KENYA | Initial 60 days, thereafter 21 | 84 | USD | 10 Boxes | YES | 100/bOX | No automated retractable or similar safety mechanism | NO TEST REPORTS | Re-tender
42142609-00050 | 173 | Syringe, plastic, latex free, non-pyrogenic and DEPH free. Nominal capacity: 5ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. With attachable needle 25G(0,50mm) x 16mm, colour code: orange and must have an automated retractable or similar safety mechanism Sterile, Single use, individually peel packed To comply with the latest issue of SANS 1124-1 and 1124-2 UNIT: Box |  | GIBELA TRADE AND INVEST 1001 CC | 1.75 | TARA | SYR082 | Shandong Qiaopai Group Co., Ltd | CHINA | 21 | 85 | USD | 10 | YES | box 100 | No sample submitted for evaluation | NO TEST REPORTS | Re-tender
42142609-00050 | 173 | Syringe, plastic, latex free, non-pyrogenic and DEPH free. Nominal capacity: 5ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. With attachable needle 25G(0,50mm) x 16mm, colour code: orange and must have an automated retractable or similar safety mechanism Sterile, Single use, individually peel packed To comply with the latest issue of SANS 1124-1 and 1124-2 UNIT: Box | N/A | SAFARMEX (PTY) LTD | 2.49 | Saclin | To be confirmed | Zhejiang Mediunion Healthcare Group Co. Ltd | China | 3days | 80% | US$ | 3000 Boxes | Yes | Box of 100 | No sample submitted for evaluation | NO TEST REPORT | Re-tender
42142609-00051 | 174 | Syringe, plastic, latex free, non-pyrogenic and DEPH free. Nominal capacity: 5ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. With attachable needle 26G(0.45mm) x 16mm, colour code: brown an automated retractable or similar safety mechanism Sterile, Single use, individually peel packed To comply with the latest issue of SANS 1124-1 and 1124-2 UNIT: Box |  | SUPRA HEALTHCARE JOHANNESBURG PTY LTD | 1.15 | REVITAL | BSS0526 | Revital Health | KENYA | Initial 60 days, thereafter 21 | 84 | USD | 10 Boxes | YES | 100/bOX | No automated retractable or similar safety mechanism | NO TEST REPORTS | Re-tender
42142609-00051 | 174 | Syringe, plastic, latex free, non-pyrogenic and DEPH free. Nominal capacity: 5ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. With attachable needle 26G(0.45mm) x 16mm, colour code: brown an automated retractable or similar safety mechanism Sterile, Single use, individually peel packed To comply with the latest issue of SANS 1124-1 and 1124-2 UNIT: Box |  | GIBELA TRADE AND INVEST 1001 CC | 1.75 | TARA | SYR083 | Shandong Qiaopai Group Co., Ltd | CHINA | 21 | 85 | USD | 10 | YES | box 100 | No sample submitted for evaluation | NO TEST REPORTS | Re-tender
42142609-00051 | 174 | Syringe, plastic, latex free, non-pyrogenic and DEPH free. Nominal capacity: 5ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. With attachable needle 26G(0.45mm) x 16mm, colour code: brown an automated retractable or similar safety mechanism Sterile, Single use, individually peel packed To comply with the latest issue of SANS 1124-1 and 1124-2 UNIT: Box | N/A | SAFARMEX (PTY) LTD | 2.49 | Saclin | To be confirmed | Zhejiang Mediunion Healthcare Group Co. Ltd | China | 3days | 80% | US$ | 3000 Boxes | Yes | Box of 100 | No sample submitted for evaluation | NO TEST REPORT | Re-tender
42142609-00052 | 175 | Syringe, plastic, latex free, non-pyrogenic and DEPH free. Nominal capacity: 5ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. With attachable needle 27G(0.4mm) x 8mm, colour code: medium grey and must have an automated retractable or similar safety mechanism Sterile, Single use, individually peel packed To comply with the latest issue of SANS 1124-1 and 1124-2 UNIT: Box |  | SUPRA HEALTHCARE JOHANNESBURG PTY LTD | 1.15 | REVITAL | BSS0527 | Revital Health | KENYA | Initial 60 days, thereafter 21 | 84 | USD | 10 Boxes | YES | 100/bOX | No automated retractable or similar safety mechanism | NO TEST REPORTS | Re-tender
42142609-00052 | 175 | Syringe, plastic, latex free, non-pyrogenic and DEPH free. Nominal capacity: 5ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. With attachable needle 27G(0.4mm) x 8mm, colour code: medium grey and must have an automated retractable or similar safety mechanism Sterile, Single use, individually peel packed To comply with the latest issue of SANS 1124-1 and 1124-2 UNIT: Box | N/A | SAFARMEX (PTY) LTD | 2.49 | Saclin | To be confirmed | Zhejiang Mediunion Healthcare Group Co. Ltd | China | 3days | 80% | US$ | 3000 Boxes | Yes | Box of 100 | No sample submitted for evaluation | NO TEST REPORT | Re-tender
42142609-00053 | 176 | Syringe, plastic, latex free, non-pyrogenic and DEPH free. Nominal capacity: 5ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. With attachable needle 29G(0.33mm) x 8mm, colour code: red and must have an automated retractable or similar safety mechanism Sterile, Single use, individually peel packed To comply with the latest issue of SANS 1124-1 and 1124-2 UNIT: Box |  | SUPRA HEALTHCARE JOHANNESBURG PTY LTD | 1.15 | REVITAL | BSS0529 | Revital Health | KENYA | Initial 60 days, thereafter 21 | 84 | USD | 10 Boxes | YES | 100/bOX | No automated retractable or similar safety mechanism | NO TEST REPORTS | Re-tender
42142609-00054 | 177 | Syringe, plastic, latex free, non-pyrogenic and DEPH free. Nominal capacity: 5ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. With attachable needle 30G(0.3mm) x 8mm, colour code: yellow and must have an automated retractable or similar safety mechanism Sterile, Single use, individually peel packed To comply with the latest issue of SANS 1124-1 and 1124-2 UNIT: Box |  | SUPRA HEALTHCARE JOHANNESBURG PTY LTD | 1.15 | REVITAL | BSS0529L | Revital Health | KENYA | Initial 60 days, thereafter 21 | 84 | USD | 10 Boxes | YES | 100/bOX | No automated retractable or similar safety mechanism | NO TEST REPORTS | Re-tender
42142609-00055 | 178 | Syringe, plastic, latex free, non-pyrogenic and DEPH free. Nominal capacity: 10ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. With attachable needle 20G(0,9mm) x 40mm, colour code: yellow and must have an automated retractable or similar safety mechanism Sterile, Single use, individually peel packed To comply with the latest issue of SANS 1124-1 and 1124-2 UNIT: Box |  | SUPRA HEALTHCARE JOHANNESBURG PTY LTD | 1.98 | REVITAL | BS1020M | Revital Health | KENYA | Initial 60 days, thereafter 21 | 84 | USD | 10 Boxes | YES | 100/bOX | No automated retractable or similar safety mechanism | NO TEST REPORTS | Re-tender
42142609-00055 | 178 | Syringe, plastic, latex free, non-pyrogenic and DEPH free. Nominal capacity: 10ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. With attachable needle 20G(0,9mm) x 40mm, colour code: yellow and must have an automated retractable or similar safety mechanism Sterile, Single use, individually peel packed To comply with the latest issue of SANS 1124-1 and 1124-2 UNIT: Box |  | GIBELA TRADE AND INVEST 1001 CC | 2.14 | TARA | SYR084 | Shandong Qiaopai Group Co., Ltd | CHINA | 21 | 85 | USD | 10 | YES | box 100 | No sample submitted for evaluation | NO TEST REPORTS | Re-tender
42142609-00055 | 178 | Syringe, plastic, latex free, non-pyrogenic and DEPH free. Nominal capacity: 10ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. With attachable needle 20G(0,9mm) x 40mm, colour code: yellow and must have an automated retractable or similar safety mechanism Sterile, Single use, individually peel packed To comply with the latest issue of SANS 1124-1 and 1124-2 UNIT: Box |  | Chemical World cc | 2.44 | WEALY | SYRS | WEALY | china | 3 | 85 | USD | 18000 | YES | 100 | COMPLY pending confirmation of latex and DEHP free | NOT COMPLY | Re-tender
42142609-00055 | 178 | Syringe, plastic, latex free, non-pyrogenic and DEPH free. Nominal capacity: 10ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. With attachable needle 20G(0,9mm) x 40mm, colour code: yellow and must have an automated retractable or similar safety mechanism Sterile, Single use, individually peel packed To comply with the latest issue of SANS 1124-1 and 1124-2 UNIT: Box | N/A | SAFARMEX (PTY) LTD | 2.86 | Saclin | To be confirmed | Zhejiang Mediunion Healthcare Group Co. Ltd | China | 3days | 80% | US$ | 3000 Boxes | Yes | Box of 100 | No sample submitted for evaluation | NO TEST REPORT | Re-tender
42142609-00056 | 179 | Syringe, plastic, latex free, non-pyrogenic and DEPH free. Nominal capacity: 10ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. With attachable needle 21G(0,8mm) x 40mm, colour code: deep green and must have an automated retractable or similar safety mechanism Sterile, Single use, individually peel packed To comply with the latest issue of SANS 1124-1 and 1124-2 UNIT: Box |  | SUPRA HEALTHCARE JOHANNESBURG PTY LTD | 1.98 | REVITAL | BS1021M | Revital Health | KENYA | Initial 60 days, thereafter 21 | 84 | USD | 10 Boxes | YES | 100/bOX | No automated retractable or similar safety mechanism | NO TEST REPORTS | Re-tender
42142609-00056 | 179 | Syringe, plastic, latex free, non-pyrogenic and DEPH free. Nominal capacity: 10ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. With attachable needle 21G(0,8mm) x 40mm, colour code: deep green and must have an automated retractable or similar safety mechanism Sterile, Single use, individually peel packed To comply with the latest issue of SANS 1124-1 and 1124-2 UNIT: Box |  | GIBELA TRADE AND INVEST 1001 CC | 2.14 | TARA | SYR085 | Shandong Qiaopai Group Co., Ltd | CHINA | 21 | 85 | USD | 10 | YES | box 100 | No sample submitted for evaluation | NO TEST REPORTS | Re-tender
42142609-00056 | 179 | Syringe, plastic, latex free, non-pyrogenic and DEPH free. Nominal capacity: 10ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. With attachable needle 21G(0,8mm) x 40mm, colour code: deep green and must have an automated retractable or similar safety mechanism Sterile, Single use, individually peel packed To comply with the latest issue of SANS 1124-1 and 1124-2 UNIT: Box | N/A | SAFARMEX (PTY) LTD | 2.86 | Saclin | To be confirmed | Zhejiang Mediunion Healthcare Group Co. Ltd | China | 3days | 80% | US$ | 3000 Boxes | Yes | Box of 100 | No sample submitted for evaluation | NO TEST REPORT | Re-tender
42142609-00056 | 179 | Syringe, plastic, latex free, non-pyrogenic and DEPH free. Nominal capacity: 10ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. With attachable needle 21G(0,8mm) x 40mm, colour code: deep green and must have an automated retractable or similar safety mechanism Sterile, Single use, individually peel packed To comply with the latest issue of SANS 1124-1 and 1124-2 UNIT: Box | 1 | Biologica Pharmaceuticals Pty Ltd | 5.09 | Sol- M | 160072IM | Sol- Millennium | China | 4-6 Weeks | 100% | USD | 1000 Boxes | 100% | N/A |  | NO TEST REPORTS | Re-tender
42142609-00057 | 180 | Syringe, plastic, latex free, non-pyrogenic and DEPH free. Nominal capacity: 10ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. With attachable needle 22G(0,7mm) x 32mm, colour code: black and must have an automated retractable or similar safety mechanism Sterile, Single use, individually peel packed To comply with the latest issue of SANS 1124-1 and 1124-2 UNIT: Box | Basic Line offer | Becton Dickinson | 1.15 | 10ml 22G x 25mm Emerald Saf T | 305986 | Becton Dickinson | Brazil | 7 Working Days | 76% | USD | 1 box | YES | 100 units in a box | No sample submitted for evaluation. | NO TEST REPORTS | Re-tender
42142609-00057 | 180 | Syringe, plastic, latex free, non-pyrogenic and DEPH free. Nominal capacity: 10ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. With attachable needle 22G(0,7mm) x 32mm, colour code: black and must have an automated retractable or similar safety mechanism Sterile, Single use, individually peel packed To comply with the latest issue of SANS 1124-1 and 1124-2 UNIT: Box |  | ALMIKA TRADING PTY LTD | 1.32 | BD | 305986 | BD | USA | 14 | 0 | NO | 100 | YES | 1 | Re-tender: End-users to revise specifications | NO TEST REPORTS | Re-tender
42142609-00057 | 180 | Syringe, plastic, latex free, non-pyrogenic and DEPH free. Nominal capacity: 10ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. With attachable needle 22G(0,7mm) x 32mm, colour code: black and must have an automated retractable or similar safety mechanism Sterile, Single use, individually peel packed To comply with the latest issue of SANS 1124-1 and 1124-2 UNIT: Box |  | SUPRA HEALTHCARE JOHANNESBURG PTY LTD | 1.98 | REVITAL | BS1022M | Revital Health | KENYA | Initial 60 days, thereafter 21 | 84 | USD | 10 Boxes | YES | 100/bOX | No automated retractable or similar safety mechanism | NO TEST REPORTS | Re-tender
42142609-00057 | 180 | Syringe, plastic, latex free, non-pyrogenic and DEPH free. Nominal capacity: 10ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. With attachable needle 22G(0,7mm) x 32mm, colour code: black and must have an automated retractable or similar safety mechanism Sterile, Single use, individually peel packed To comply with the latest issue of SANS 1124-1 and 1124-2 UNIT: Box |  | GIBELA TRADE AND INVEST 1001 CC | 2.14 | TARA | SYR086 | Shandong Qiaopai Group Co., Ltd | CHINA | 21 | 85 | USD | 10 | YES | box 100 | No sample submitted for evaluation | NO TEST REPORTS | Re-tender
42142609-00057 | 180 | Syringe, plastic, latex free, non-pyrogenic and DEPH free. Nominal capacity: 10ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. With attachable needle 22G(0,7mm) x 32mm, colour code: black and must have an automated retractable or similar safety mechanism Sterile, Single use, individually peel packed To comply with the latest issue of SANS 1124-1 and 1124-2 UNIT: Box | N/A | SAFARMEX (PTY) LTD | 2.86 | Saclin | To be confirmed | Zhejiang Mediunion Healthcare Group Co. Ltd | China | 3days | 80% | US$ | 3000 Boxes | Yes | Box of 100 | No sample submitted for evaluation | NO TEST REPORT | Re-tender
42142609-00058 | 181 | Syringe, plastic, latex free, non-pyrogenic and DEPH free. Nominal capacity: 10ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. With attachable needle 23G(0,6mm) x 25mm,colour code: deep blue and must have an automated retractable or similar safety mechanism Sterile, Single use, individually peel packed To comply with the latest issue of SANS 1124-1 and 1124-2 UNIT: Box |  | SUPRA HEALTHCARE JOHANNESBURG PTY LTD | 1.98 | REVITAL | BS1023M | Revital Health | KENYA | Initial 60 days, thereafter 21 | 84 | USD | 10 Boxes | YES | 100/bOX | No automated retractable or similar safety mechanism | NO TEST REPORTS | Re-tender
42142609-00058 | 181 | Syringe, plastic, latex free, non-pyrogenic and DEPH free. Nominal capacity: 10ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. With attachable needle 23G(0,6mm) x 25mm,colour code: deep blue and must have an automated retractable or similar safety mechanism Sterile, Single use, individually peel packed To comply with the latest issue of SANS 1124-1 and 1124-2 UNIT: Box |  | GIBELA TRADE AND INVEST 1001 CC | 2.14 | TARA | SYR087 | Shandong Qiaopai Group Co., Ltd | CHINA | 21 | 85 | USD | 10 | YES | box 100 | No sample submitted for evaluation | NO TEST REPORTS | Re-tender
42142609-00058 | 181 | Syringe, plastic, latex free, non-pyrogenic and DEPH free. Nominal capacity: 10ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. With attachable needle 23G(0,6mm) x 25mm,colour code: deep blue and must have an automated retractable or similar safety mechanism Sterile, Single use, individually peel packed To comply with the latest issue of SANS 1124-1 and 1124-2 UNIT: Box | N/A | SAFARMEX (PTY) LTD | 2.86 | Saclin | To be confirmed | Zhejiang Mediunion Healthcare Group Co. Ltd | China | 3days | 80% | US$ | 3000 Boxes | Yes | Box of 100 | No sample submitted for evaluation | NO TEST REPORT | Re-tender
42142609-00059 | 182 | Syringe, plastic, latex free, non-pyrogenic and DEPH free. Nominal capacity: 10ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. With attachable needle 24G(0.55mm) x 20mm, colour code: medium purple and must have an automated retractable or similar safety mechanism Sterile, Single use, individually peel packed To comply with the latest issue of SANS 1124-1 and 1124-2 UNIT: Box |  | SUPRA HEALTHCARE JOHANNESBURG PTY LTD | 1.98 | REVITAL | BS1024M | Revital Health | KENYA | Initial 60 days, thereafter 21 | 84 | USD | 10 Boxes | YES | 100/bOX | No automated retractable or similar safety mechanism | NO TEST REPORTS | Re-tender
42142609-00059 | 182 | Syringe, plastic, latex free, non-pyrogenic and DEPH free. Nominal capacity: 10ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. With attachable needle 24G(0.55mm) x 20mm, colour code: medium purple and must have an automated retractable or similar safety mechanism Sterile, Single use, individually peel packed To comply with the latest issue of SANS 1124-1 and 1124-2 UNIT: Box |  | GIBELA TRADE AND INVEST 1001 CC | 2.14 | TARA | SYR088 | Shandong Qiaopai Group Co., Ltd | CHINA | 21 | 85 | USD | 10 | YES | box 100 | No sample submitted for evaluation | NO TEST REPORTS | Re-tender
42142609-00059 | 182 | Syringe, plastic, latex free, non-pyrogenic and DEPH free. Nominal capacity: 10ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. With attachable needle 24G(0.55mm) x 20mm, colour code: medium purple and must have an automated retractable or similar safety mechanism Sterile, Single use, individually peel packed To comply with the latest issue of SANS 1124-1 and 1124-2 UNIT: Box | N/A | SAFARMEX (PTY) LTD | 2.86 | Saclin | To be confirmed | Zhejiang Mediunion Healthcare Group Co. Ltd | China | 3days | 80% | US$ | 3000 Boxes | Yes | Box of 100 | No sample submitted for evaluation | NO TEST REPORT | Re-tender
42142609-00060 | 183 | Syringe, plastic, latex free, non-pyrogenic and DEPH free. Nominal capacity: 10ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. With attachable needle 25G(0,50mm) x 16mm, colour code: orange and must have an automated retractable or similar safety mechanism Sterile, Single use, individually peel packed To comply with the latest issue of SANS 1124-1 and 1124-2 UNIT: Box |  | SUPRA HEALTHCARE JOHANNESBURG PTY LTD | 1.98 | REVITAL | BS1025M | Revital Health | KENYA | Initial 60 days, thereafter 21 | 84 | USD | 10 Boxes | YES | 100/bOX | No automated retractable or similar safety mechanism | NO TEST REPORTS | Re-tender
42142609-00060 | 183 | Syringe, plastic, latex free, non-pyrogenic and DEPH free. Nominal capacity: 10ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. With attachable needle 25G(0,50mm) x 16mm, colour code: orange and must have an automated retractable or similar safety mechanism Sterile, Single use, individually peel packed To comply with the latest issue of SANS 1124-1 and 1124-2 UNIT: Box |  | GIBELA TRADE AND INVEST 1001 CC | 2.14 | TARA | SYR089 | Shandong Qiaopai Group Co., Ltd | CHINA | 21 | 85 | USD | 10 | YES | box 100 | No sample submitted for evaluation | NO TEST REPORTS | Re-tender
42142609-00060 | 183 | Syringe, plastic, latex free, non-pyrogenic and DEPH free. Nominal capacity: 10ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. With attachable needle 25G(0,50mm) x 16mm, colour code: orange and must have an automated retractable or similar safety mechanism Sterile, Single use, individually peel packed To comply with the latest issue of SANS 1124-1 and 1124-2 UNIT: Box | N/A | SAFARMEX (PTY) LTD | 2.86 | Saclin | To be confirmed | Zhejiang Mediunion Healthcare Group Co. Ltd | China | 3days | 80% | US$ | 3000 Boxes | Yes | Box of 100 | No sample submitted for evaluation | NO TEST REPORT | Re-tender
42142609-00061 | 184 | Syringe, plastic, latex free, non-pyrogenic and DEPH free. Nominal capacity: 10ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. With attachable needle 26G(0.45mm) x 16mm, colour code: brown an automated retractable or similar safety mechanism Sterile, Single use, individually peel packed To comply with the latest issue of SANS 1124-1 and 1124-2 UNIT: Box |  | SUPRA HEALTHCARE JOHANNESBURG PTY LTD | 1.98 | REVITAL | BS1026M | Revital Health | KENYA | Initial 60 days, thereafter 21 | 84 | USD | 10 Boxes | YES | 100/bOX | No automated retractable or similar safety mechanism | NO TEST REPORTS | Re-tender
42142609-00061 | 184 | Syringe, plastic, latex free, non-pyrogenic and DEPH free. Nominal capacity: 10ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. With attachable needle 26G(0.45mm) x 16mm, colour code: brown an automated retractable or similar safety mechanism Sterile, Single use, individually peel packed To comply with the latest issue of SANS 1124-1 and 1124-2 UNIT: Box |  | GIBELA TRADE AND INVEST 1001 CC | 2.14 | TARA | SYR090 | Shandong Qiaopai Group Co., Ltd | CHINA | 21 | 85 | USD | 10 | YES | box 100 | No sample submitted for evaluation | NO TEST REPORTS | Re-tender
42142609-00061 | 184 | Syringe, plastic, latex free, non-pyrogenic and DEPH free. Nominal capacity: 10ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. With attachable needle 26G(0.45mm) x 16mm, colour code: brown an automated retractable or similar safety mechanism Sterile, Single use, individually peel packed To comply with the latest issue of SANS 1124-1 and 1124-2 UNIT: Box | N/A | SAFARMEX (PTY) LTD | 2.86 | Saclin | To be confirmed | Zhejiang Mediunion Healthcare Group Co. Ltd | China | 3days | 80% | US$ | 3000 Boxes | Yes | Box of 100 | No sample submitted for evaluation | NO TEST REPORT | Re-tender
42142609-00062 | 185 | Syringe, plastic, latex free, non-pyrogenic and DEPH free. Nominal capacity: 10ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. With attachable needle 27G(0.4mm) x 8mm, colour code: medium grey and must have an automated retractable or similar safety mechanism Sterile, Single use, individually peel packed To comply with the latest issue of SANS 1124-1 and 1124-2 UNIT: Box |  | SUPRA HEALTHCARE JOHANNESBURG PTY LTD | 1.98 | REVITAL | BS1027M | Revital Health | KENYA | Initial 60 days, thereafter 21 | 84 | USD | 10 Boxes | YES | 100/bOX | No automated retractable or similar safety mechanism | NO TEST REPORTS | Re-tender
42142609-00062 | 185 | Syringe, plastic, latex free, non-pyrogenic and DEPH free. Nominal capacity: 10ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. With attachable needle 27G(0.4mm) x 8mm, colour code: medium grey and must have an automated retractable or similar safety mechanism Sterile, Single use, individually peel packed To comply with the latest issue of SANS 1124-1 and 1124-2 UNIT: Box | N/A | SAFARMEX (PTY) LTD | 2.86 | Saclin | To be confirmed | Zhejiang Mediunion Healthcare Group Co. Ltd | China | 3days | 80% | US$ | 3000 Boxes | Yes | Box of 100 | No sample submitted for evaluation | NO TEST REPORT | Re-tender
42142609-00063 | 186 | Syringe, plastic, latex free, non-pyrogenic and DEPH free. Nominal capacity: 10ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. With attachable needle 29G(0.33mm) x 8mm, colour code: red and must have an automated retractable or similar safety mechanism Sterile, Single use, individually peel packed To comply with the latest issue of SANS 1124-1 and 1124-2 UNIT: Box |  | SUPRA HEALTHCARE JOHANNESBURG PTY LTD | 1.98 | REVITAL | BS1029M | Revital Health | KENYA | Initial 60 days, thereafter 21 | 84 | USD | 10 Boxes | YES | 100/bOX | No automated retractable or similar safety mechanism | NO TEST REPORTS | Re-tender
42142609-00063 | 186 | Syringe, plastic, latex free, non-pyrogenic and DEPH free. Nominal capacity: 10ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. With attachable needle 29G(0.33mm) x 8mm, colour code: red and must have an automated retractable or similar safety mechanism Sterile, Single use, individually peel packed To comply with the latest issue of SANS 1124-1 and 1124-2 UNIT: Box | N/A | SAFARMEX (PTY) LTD | 2.86 | Saclin | To be confirmed | Zhejiang Mediunion Healthcare Group Co. Ltd | China | 3days | 80% | US$ | 3000 Boxes | Yes | Box of 100 | No sample submitted for evaluation | NO TEST REPORT | Re-tender
42142609-00064 | 187 | Syringe, plastic, latex free, non-pyrogenic and DEPH free. Nominal capacity: 10ml, must have clear graduation markings that are not easily removed with spirits and allow for permanent marker markings without rubbing off. With attachable needle 30G(0.3mm) x 8mm, colour code: yellow and must have an automated retractable or similar safety mechanism Sterile, Single use, individually peel packed To comply with the latest issue of SANS 1124-1 and 1124-2 UNIT: Box |  | SUPRA HEALTHCARE JOHANNESBURG PTY LTD | 1.98 | REVITAL | BS1030M | Revital Health | KENYA | Initial 60 days, thereafter 21 | 84 | USD | 10 Boxes | YES | 100/bOX | No automated retractable or similar safety mechanism | NO TEST REPORTS | Re-tender
Note on F171: Johnny Ramoroka:R266,5 is it a price 100 or 50 in a box
Note on A256: Johnny Ramoroka: confirm pricing with the bidder it it is R180
Note on A317: Johnny Ramoroka: Verify the sample of BD if it complies.
Note on F376: Johnny Ramoroka: r1071,5 is it price of 25 of 50 in a box
Note on A541: Johnny Ramoroka: check sample evaluation report
Note on F810: Johnny Ramoroka: UOM 252,60 and pricing schedule is 52,60
Note on F1287: Johnny Ramoroka: R45,43 is for each
Note on F1290: Johnny Ramoroka: R45,43 is for each
Note on F1294: Johnny Ramoroka: R45,43 is for each
Note on F1297: Johnny Ramoroka: R45,43 is for each