## TECHNICAL SPEC
ANNEXURE A: TECHNICAL SPECIFICATION REQUIREMENTS RT31-1-2023: SUPPLY AND DELIVERY OF MEDICAL ADMINISTRATION SETS TO THE STATE FOR A PERIOD ENDING 31 MAY 2027
Line # | Item Number | Description | Compliant with specification (Yes/No) | Indicate the Deviation if applicable
1 | RT31-01-008 | ADMINISTRATION SET standard, for paediatric use,Latex Free, With two needle Y-injection sites, one proximal and one distal Accurate flow regulating roller clamp between the two Y-sites. The tube must be non-kinking at any site and all temperatures. The drip chamber must be fitted with a drop needle to ensure uniformity of drop size. Drip volume: 60 drops/ml. Distal luer lock male adaptor. Retractable Rotating male luer slip lock adaptor Total length: 150cm-290cm Latex free: Latex free to be stated on package Sterile fluid pathway. Individually packed, sterile, tamper proof. To comply with the latest issue of SANS 1775. Compliance certificate to be submitted with the bid. EACH
2 | RT31-01-018 | ADMINISTRATION SET standard, for paediatric use, Latex and DEHP free, With two needle Y-injection sites, one proximal and one distal. Accurate flow regulating roller clamp between the two Y-sites. The tube must be non-kinking at any site and all temperatures. The drip chamber must be fitted with a drop needle to ensure uniformity of drop size. Drip volume: 60 drops/ml. Distal luer lock male adaptor. Retractable Rotating male luer slip lock adaptor Total length: 150cm-290cm Latex and DEHP free: Latex and DEHP free to be stated on package Sterile fluid pathway. Individually packed, sterile, tamper proof. To comply with the latest issue of SANS 1775. Compliance certificate to be submitted with the bid. EACH
3 | RT31-02-001 | BLOOD ADMINISTRATION SET: Blood and other blood products for adult use, Latex free , With single chamber and filter. Drip chamber Lenght: minimum of 7cm. Filter: 180-220 micron nylon. Length of tubing: Total length of 150cm-290cm The tube must be non-kinking at any site and all temperatures. Drip Chamber: Vented Drip volume: 10 drops /ml Distal luer lock male adaptor. Retractable Rotating male luer slip lock adaptor Accurate flow regulating roller clamp. Sterile fluid pathway Latex free: Latex free to be stated on package Individually packed, sterile, tamper proof. To comply with the latest issue of SANS 1775 Compliance certificate to be submitted with the bid. EACH
4 | RT31-02-002 | BLOOD ADMINISTRATION SET: Blood and other blood products for adult use, Latex free, With single chamber and filter. Drip chamber Lenght: minimum of 7cm. Filter: 180-220 micron nylon. Length of tubing: Total length 150cm-290cm The tube must be non-kinking at any site and all temperatures. Drip Chamber: non-vented Drip volume: 10 drops /ml Distal luer lock male adaptor. Retractable Rotating male luer slip lock adaptor Accurate flow regulating roller clamp. Sterile fluid pathway Latex free: Latex free to be stated on package Individually packed, sterile, tamper proof. To comply with the latest issue of SANS 1775 Compliance certificate to be submitted with the bid. EACH
5 | RT31-02-003 | BLOOD ADMINISTRATION SET: Blood and other blood products for adult use, Latex free, With one needle Y-injection site with a non-return valve for continuous flow. With single chamber and filter. Drip chamber lenght: minimum of 7cm. Filter: 180-220 micron nylon. Length of tubing: Total length 150cm-290cm. The tube must be non-kinking at any site and all temperatures. Drip Chamber: Vented Drip volume: 10 drops /ml Distal luer lock male adaptor. Retractable Rotating male luer slip lock adaptor Accurate flow regulating roller clamp. Sterile fluid pathway Latex free: Latex free to be stated on package Individually packed, sterile, tamper proof. To comply with the latest issue of SANS 1775 Compliance certificate to be submitted with the bid. EACH
6 | RT31-02-004 | BLOOD ADMINISTRATION SET: Blood and other blood products for adult use, With one needle Y-injection site with a non-return valve for continuous flow. With single chamber and filter. Drip chamber lenght: minimum of 7cm. Filter: 180-220 micron nylon. Length of tubing: Total length 150cm-290cm The tube must be non-kinking at any site and all temperatures. Drip Chamber: Non-vented Drip volume: 10 drops /ml Distal luer lock male adaptor. Retractable Rotating male luer slip lock adaptor Accurate flow regulating roller clamp. Sterile fluid pathway Latex free: Latex free to be stated on package Individually packed, sterile, tamper proof. To comply with the latest issue of SANS 1775 Compliance certificate to be submitted with the bid. EACH
7 | RT31-02-005 | BLOOD ADMINISTRATION SET: Blood and other blood products for adult use, Latex free, With double chamber Filter: 180-220 micron nylon fitted in the lower chamber Length of tubing: Total length of 150cm-180cm. The tube must be non-kinking at any site and all temperatures. Drip volume: 10 drops /ml Distal luer lock male adaptor. Retractable Rotating male luer slip lock adaptor Accurate flow regulating roller clamp. Sterile fluid pathway Latex free: Latex free to be stated on package Individually packed, sterile, tamper proof. To comply with the latest issue of SANS 1775 Compliance certificate to be submitted with the bid. EACH
8 | RT31-02-006 | BLOOD ADMINISTRATION SET: Blood and other blood products for adult use, Latex free, With one needle Y-injection site with a non-return valve for continuous flow. With double chamber Filter: 180-220 micron nylon fitted in the lower chamber. Length of tubing: Total length 150cm-290cm. The tube must be non-kinking at any site and all temperatures. Drip volume: 10 drops /ml Distal luer lock male adaptor. Retractable Rotating male luer slip lock adaptor Accurate flow regulating roller clamp. Sterile fluid pathway Latex free: Latex free to be stated on package Individually packed, sterile, tamper proof. To comply with the latest issue of SANS 1775 Compliance certificate to be submitted with the bid. EACH
9 | RT31-02-007 | BLOOD ADMINISTRATION SET: Blood and other blood products for adult use, Latex free, With one needle free Y-injection site with a non-return valve for continuous flow. With double chamber Filter: 180-220 micron nylon fitted in the lower chamber. Length of tubing: Total length of 150cm-290cm. The tube must be non-kinking at any site and all temperatures. Drip volume: 10 drops /ml Distal luer lock male adaptor. Retractable Rotating male luer slip lock adaptor Accurate flow regulating roller clamp. Sterile fluid pathway Latex free: Latex free to be stated on package Individually packed, sterile, tamper proof. To comply with the latest issue of SANS 1775 Compliance certificate to be submitted with the bid. EACH
10 | RT31-02-009 | BLOOD ADMINISTRATION SET: Blood and other blood products for adult use, Latex and DEHP free, With single chamber and filter. Drip chamber lenght: minimum of 7cm. Filter: 180-220 micron nylon. Length of tubing: Total length 150cm-290cm. The tube must be non-kinking at any site and all temperatures. Drip Chamber: Vented Drip volume: 10 drops /ml Distal luer lock male adaptor. Retractable Rotating male luer slip lock adaptor Accurate flow regulating roller clamp. Sterile fluid pathway Latex and DEHP free: Latex and DEHP free to be stated on package Individually packed, sterile, tamper proof. To comply with the latest issue of SANS 1775 Compliance certificate to be submitted with the bid. EACH
11 | RT31-02-010 | BLOOD ADMINISTRATION SET: Blood and other blood products for adult use, Latex and DEHP free, With single chamber and filter. Drip chamber lenght: minimum of 6-7cm. Filter: 180-220 micron nylon. Length of tubing: Total length 150cm-290cm. The tube must be non-kinking at any site and all temperatures. Drip Chamber: Non-Vented Drip volume: 10 drops /ml Distal luer lock male adaptor. Retractable Rotating male luer slip lock adaptor Accurate flow regulating roller clamp. Sterile fluid pathway Latex and DEHP free: Latex and DEHP free to be stated on package Individually packed, sterile, tamper proof. To comply with the latest issue of SANS 1775 Compliance certificate to be submitted with the bid EACH
12 | RT31-02-011 | BLOOD ADMINISTRATION SET: Blood and other blood products for paediatric use, Latex and DEHP free, With single chamber and filter. Drip chamber Length: minimum of 6-7cm. Filter: 180-220 micron nylon. Length of tubing: Total length 150cm-290cm. The tube must be non-kinking at any site and all temperatures. Drip Chamber: Vented Drip volume: 20 drops /ml Distal luer lock male adaptor. Retractable Rotating male luer slip lock adaptor Accurate flow regulating roller clamp. Sterile fluid pathway Latex and DEHP free: Latex and DEHP free to be stated on package Individually packed, sterile, tamper proof. To comply with the latest issue of SANS 1775 Compliance certificate to be submitted with the bid. EACH
13 | RT31-02-013 | BLOOD ADMINISTRATION SET: Blood and other blood products for paediatric use, With one needle Y-injection site with a non-return valve for continuous flow. With single chamber and filter. Drip chamber lenght: minimum of 6-7cm. Filter: 180-220 micron nylon. Length of tubing: Total length 150cm-290cm. The tube must be non-kinking at any site and all temperatures. Drip Chamber: Vented Drip volume: 20 drops /ml Distal luer lock male adaptor. Retractable Rotating male luer slip lock adaptor Accurate flow regulating roller clamp. Sterile fluid pathway Latex and DEHP free: Latex and DEHP free to be stated on package Individually packed, sterile, tamper proof. To comply with the latest issue of SANS 1775 Compliance certificate to be submitted with the bid. EACH
14 | RT31-02-014 | BLOOD ADMINISTRATION SET: Blood and other blood products for paediatric use, Latex and DEHP free, With one needle Y-injection site with a non-return valve for continuous flow. With single chamber and filter. Drip chamber lenght: minimum of 6-7cm. Filter: 180-220 micron nylon. Length of tubing: Total length 150cm-290cm. The tube must be non-kinking at any site and all temperatures. Drip Chamber: Non-Vented Drip volume: 20 drops /ml Distal luer lock male adaptor. Retractable Rotating male luer slip lock adaptor Accurate flow regulating roller clamp. Sterile fluid pathway Latex and DEHP free: Latex and DEHP free to be stated on package Individually packed, sterile, tamper proof. To comply with the latest issue of SANS 1775 Compliance certificate to be submitted with the bid. EACH
15 | RT31-02-015 | BLOOD ADMINISTRATION SET: Blood and other blood products for adult use, Latex and DEHP free, without any Y-injection site With double chamber Filter: 180-220 micron nylon fitted in the lower chamber. Length of tubing: Total length 150cm-290cm. The tube must be non-kinking at any site and all temperatures. Drip Chamber: Vented Drip volume: 10 drops /ml Distal luer lock male adaptor. Retractable Rotating male luer slip lock adaptor Accurate flow regulating roller clamp. Sterile fluid pathway Latex and DEHP free: Latex and DEHP free to be stated on package Individually packed, sterile, tamper proof. To comply with the latest issue of SANS 1775 Compliance certificate to be submitted with the bid. EACH
16 | RT31-02-016 | BLOOD ADMINISTRATION SET: Blood and other blood products for paediatric use, Latex and DEHP free, without any Y-injection site With double chamber Filter: 180-220 micron nylon fitted in the lower chamber. Length of tubing: Total length 150cm-290cm The tube must be non-kinking at any site and all temperatures. Drip Chamber: Vented Drip volume: 20 drops /ml Distal luer lock male adaptor. Retractable Rotating male luer slip lock adaptor Accurate flow regulating roller clamp. Sterile fluid pathway Latex and DEHP free: Latex and DEHP free to be stated on package Individually packed, sterile, tamper proof. To comply with the latest issue of SANS 1775 Compliance certificate to be submitted with the bid. EACH
17 | RT31-02-017 | BLOOD ADMINISTRATION SET: Blood and other blood products for paediatric use, Latex and DEHP free, With one needle Y-injection site with a non-return valve for continuous flow. With double chamber Filter: 180-220 micron nylon fitted in the lower chamber. Length of tubing: Total length 150cm-290cm. The tube must be non-kinking at any site and all temperatures. Drip Chamber: Vented Drip volume: 20 drops /ml Distal luer lock male adaptor. Retractable Rotating male luer slip lock adaptor Accurate flow regulating roller clamp. Sterile fluid pathway Latex and DEHP free: Latex and DEHP free to be stated on package Individually packed, sterile, tamper proof. To comply with the latest issue of SANS 1775 Compliance certificate to be submitted with the bid. EACH
18 | RT31-02-018 | BLOOD ADMINISTRATION SET: Blood and other blood products for paediatric use, Latex and DEHP free, With one needle free Y-injection site with a non-return valve for continuous flow. With double chamber Filter: 180-220 micron nylon fitted in the lower chamber. Length of tubing: Total length 150cm-290cm. The tube must be non-kinking at any site and all temperatures. Drip Chamber: Vented Drip volume: 20 drops /ml Distal luer lock male adaptor . Retractable Rotating male luer slip lock adaptor Accurate flow regulating roller clamp. Sterile fluid pathway Latex and DEHP free: Latex and DEHP free to be stated on package Individually packed, sterile, tamper proof. To comply with the latest issue of SANS 1775 Compliance certificate to be submitted with the bid. EACH
19 | RT31-03-003 | ADMINISTRATION SET, IN‐LINE, CONTROLLED VOLUME UNIT, Latex free, With a 150ml calibrated burette. For use with blood solutions or parenteral nutrition of precise volume infusion. Proximal end: · Spike · Tube, length of 15cm-20cm · Single clamp. Burette: · Volume 150ml – calibrated from 0 – 150ml with graduated markings at each 1ml. · Hanger · Must have a bacteria retentive air inlet · Needle Injection Port Distal end: · Open tube, length of 5 -10cm. Sterile fluid pathway Latex free: Latex free to be stated on package. Individually packed, sterile, tamper proof. To comply with the latest issue of SANS 1775 Compliance certificate to be submitted with the bid. EACH
20 | RT31-03-004 | ADMINISTRATION SET, IN‐LINE, CONTROLLED VOLUME UNIT, Latex free, With a 150ml calibrated burette. For use with blood solutions or parenteral nutrition of precise volume infusion. Proximal end: · Spike, air vented with bacterial protection · Tube length: +/- 15cm · Single clamp. Burette: · Volume 150ml – calibrated from 0 – 150ml with graduated markings at each 1ml. · Hanger · Must have a bacterial protection air inlet · Needle Injection Port · Automatic shut off valve to prevent air entry if the burette chamber gets empty. Distal end: · Must start with a chamber: · The drip chamber must be fitted with a drop needle to ensure uniformity of drop size. · Drip volume: 60 drops/ml. · Tube length: ±150cm · One injection Y-site · Accurate flow regulating roller clamp between the chamber and Y-site. · The tube must be non-kinking at any site and all temperatures. · Distal luer lock male adaptor . Sterile fluid pathway Latex free: Latex free to be stated on package. Individually packed, sterile, tamper proof. To comply with the latest issue of SANS 1775 Compliance certificate to be submitted with the bid. EACH
21 | RT31-03-005 | ADMINISTRATION SET, IN‐LINE, CONTROLLED VOLUME UNIT, Latex free With a 150ml calibrated burette. For use with blood solutions or parenteral nutrition of precise volume infusion. Proximal end: · Spike, air vented with bacterial protection · Tube length: +/- 15cm · Single clamp. Burette: · Volume 150ml – calibrated from 0 – 150ml with graduated markings at each 1ml. · Hanger · Must have a bacterial protection air inlet · Needle Free Injection Port · Automatic shut off valve to prevent air entry if the burette chamber gets empty. Distal end: · Must start with a chamber: · The drip chamber must be fitted with a drop needle to ensure uniformity of drop size. · Drip volume: 60 drops/ml. · Tube length: 150cm-180cm · One needle free injection Y-site · Accurate flow regulating roller clamp between the chamber and Y-site. · The tube must be non-kinking at any site and all temperatures. · Distal luer lock male adaptor Sterile fluid pathway Latex free: Latex free to be stated on package. Individually packed, sterile, tamper proof. To comply with the latest issue of SANS 1775 Compliance certificate to be submitted with the bid. EACH
22 | RT31-03-006 | ADMINISTRATION SET, IN‐LINE, CONTROLLED VOLUME UNIT, Latex and DEHP free With a 150ml calibrated burette. For use with blood solutions or parenteral nutrition of precise volume infusion. Proximal end: · Spike, air vented with bacterial protection · Tube, length of +/- 15cm · Single clamp. Burette: · Volume 150ml – calibrated from 0 – 150ml with graduated markings at each 1ml. · Hanger · Must have a bacterial protection air inlet · Needle Free Injection Port · Automatic shut off valve to prevent air entry if the burette chamber gets empty. Distal end: · Must start with a chamber: · The drip chamber must be fitted with a drop needle to ensure uniformity of drop size. · Drip volume: 60 drops/ml. · Tube length: minimum of 150cm-180cm · One needle free injection Y-site · Accurate flow regulating roller clamp between the chamber and Y-site. · The tube must be non-kinking at any site and all temperatures. · Distal luer lock male adaptor Sterile fluid pathway Latex and DEHP free: Latex and DEHP free to be stated on package. Individually packed, sterile, tamper proof. To comply with the latest issue of SANS 1775 Compliance certificate to be submitted with the bid. EACH
23 | RT31-04-001 | INTRAVENOUS CONTROL FLOW DEVICE, Proximal end: · Female connector on a short tube (Length of the tube 0,65 to 0,75m) Flow control device: · Scale: Visible control range: off/0 - 250ml/hour · Able to compensate for head height variations Distal end: · Short tube (Length of the tube 0,65 to 0,75m) ending with male adaptor to be the luer lock type Latex free: Latex free to be stated on package. Sterile fluid pathway Individually packed, sterile, tamper proof. To comply with the latest issue of SANS 1775 Compliance certificate to be submitted with the bid. EACH
24 | RT31-04-002 | INTRAVENOUS CONTROL FLOW DEVICE, Latex free, Proximal end: · Female connector on a short tube (Length of the tube 0,65 to 0,75m) Flow control device: · Scale: Visible control range: off/0 - 250ml/hour · Able to compensate for head height variations Distal end: · Tube length: 0,65 to 0,75m · One needle injection Y-site · The tube must be non-kinking at any site and all temperatures. · Distal luer lock male adaptor Sterile fluid pathway Latex free: Latex free to be stated on package. Individually packed, sterile, tamper proof. To comply with the latest issue of SANS 1775 Compliance certificate to be submitted with the bid. EACH
25 | RT31-04-003 | INTRAVENOUS CONTROL FLOW DEVICE,Latex and DEHP free, Proximal end: · Female connector on a short tube (Length of the tube 0,65 to 0,75m) Flow control device: · Scale: Visible control range: off/0 - 250ml/hour · Able to compensate for head height variations Distal end: · Tube length: 0,65 to 0,75m · One needle free injection Y-site · The tube must be non-kinking at any site and all temperatures. · Distal luer lock male adaptor Sterile fluid pathway Latex and DEHP free: Latex and DEHP free to be stated on package. Individually packed, sterile, tamper proof. To comply with the latest issue of SANS 1775 Compliance certificate to be submitted with the bid. EACH
26 | RT31-06-001 | EXTENSION SETS: Latex Free, Suitable for use with administration set and an infusion pump. Minimal residual volume Tube length: ± 1m Inner diameter (ID) of tubing: 1 to 2mm Supplied with one male and one female luer lock connector The tube must be non-kinking at any site and all temperatures. PVC material Sterile fluid pathway Latex free: Latex free to be stated on package. Individually packed, sterile, tamper proof. To comply with the latest issue of SANS 1775 Compliance certificate to be submitted with the bid. EACH
27 | RT31-06-003 | EXTENSION SETS: Latex Free, Suitable for use with administration set and an infusion pump. Minimal residual volume Tube length: ±2m Inner diameter (ID) of tubing: 1 to 2mm Supplied with one male and one female luer lock connector The tube must be non-kinking at any site and all temperatures. PVC material Sterile fluid pathway Latex free: Latex free to be stated on package. Individually packed, sterile, tamper proof. To comply with the latest issue of SANS 1775 Compliance certificate to be submitted with the bid. EACH
28 | RT31-06-004 | EXTENSION SETS: Latex Free, Suitable for use with administration set and an infusion pump. Minimal residual volume Spiral tube length: ± 1m ID of tubing: 1 to 2mm Supplied with one male and one female luer lock connector The tube must be non-kinking at any site and all temperatures. PVC material Sterile fluid pathway Latex free: Latex free to be stated on package. Individually packed, sterile, tamper proof. To comply with the latest issue of SANS 1775 Compliance certificate to be submitted with the bid. EACH
29 | RT31-06-005 | EXTENSION SETS: Latex Free, Suitable for use with administration set and an infusion pump. Minimal residual volume Spiral tube length: ± 2m ID of tubing: 1 to 2mm Supplied with one male and one female luer lock connector The tube must be non-kinking at any site and all temperatures. PVC material Sterile fluid pathway Latex free: Latex free to be stated on package. Individually packed, sterile, tamper proof. To comply with the latest issue of SANS 1775 Compliance certificate to be submitted with the bid. EACH
30 | RT31-06-006 | EXTENSION SETS: Latex and DEHP free, Suitable for use with administration set and an infusion pump. Minimal residual volume Tube length: ± 1m Inner diameter (ID) of tubing: 1 to 2mm Supplied with one male and one female luer lock connector The tube must be non-kinking at any site and all temperatures. Sterile fluid pathway Latex and DEHP free: Latex and DEHP free to be stated on package. Individually packed, sterile, tamper proof. To comply with the latest issue of SANS 1775 Compliance certificate to be submitted with the bid. EACH
31 | RT31-06-007 | EXTENSION SETS: Latex and DEHP free, Suitable for use with administration set and an infusion pump. Minimal residual volume Tube length: ± 1.5m Inner diameter (ID) of tubing: 1 to 2mm Supplied with one male and one female luer lock connector The tube must be non-kinking at any site and all temperatures. Sterile fluid pathway Latex and DEHP free: Latex and DEHP free to be stated on package. Individually packed, sterile, tamper proof. To comply with the latest issue of SANS 1775 Compliance certificate to be submitted with the bid. EACH
32 | RT31-06-008 | EXTENSION SETS: Latex and DEHP Free, Suitable for use with administration set and an infusion pump. Minimal residual volume Tube length: ± 2m Inner diameter (ID) of tubing: 0,8 to 2mm Supplied with one male and one female luer lock connector The tube must be non-kinking at any site and all temperatures. Sterile fluid pathway Latex and DEHP free: Latex and DEHP free to be stated on package. Individually packed, sterile, tamper proof. To comply with the latest issue of SANS 1775 Compliance certificate to be submitted with the bid. EACH
33 | RT31-06-009 | EXTENSION SETS: Latex and DEHP Free, Suitable for use with administration set and an infusion pump. Light protection / Opaque Minimal residual volume Tube length: ± 1m Inner diameter (ID) of tubing: 1 to 2mm Supplied with one male and one female luer lock connector The tube must be non-kinking at any site and all temperatures. Sterile fluid pathway Latex and DEHP free: Latex and DEHP free to be stated on package. Individually packed, sterile, tamper proof. To comply with the latest issue of SANS 1775Compliance certificate to be submitted with the bid. EACH
34 | RT31-06-010 | EXTENSION SETS: Latex and DEHP Free, Suitable for use with administration set and an infusion pump. Light protection / Opaque Minimal residual volume Tube length: ± 1.5m Inner diameter (ID) of tubing: 1 to 2mm Supplied with one male and one female luer lock connector The tube must be non-kinking at any site and all temperatures. Sterile fluid pathway Latex and DEHP free: Latex and DEHP free to be stated on package. Individually packed, sterile, tamper proof. To comply with the latest issue of SANS 1775Compliance certificate to be submitted with the bid. EACH
35 | RT31-06-011 | EXTENSION SETS: Latex and DEHP Free, Suitable for use with administration set and an infusion pump. Light protection / Opaque Minimal residual volume Tube length: ± 2m Inner diameter (ID) of tubing: 1 to 2mm Supplied with one male and one female luer lock connector The tube must be non-kinking at any site and all temperatures. Sterile fluid pathway Latex and DEHP free: Latex and DEHP free to be stated on package. Individually packed, sterile, tamper proof. To comply with the latest issue of SANS 1775Compliance certificate to be submitted with the bid. EACH
36 | RT31-06-012 | EXTENSION SETS: Latex and DEHP Free, Suitable for use with administration set and an infusion pump. Minimal residual volume Spiral tube length: ± 1m ID of tubing: 1 to 2mm Supplied with one male and one female luer lock connector The tube must be non-kinking at any site and all temperatures. Sterile fluid pathway Latex and DEHP free: Latex and DEHP free to be stated on package. Individually packed, sterile, tamper proof. To comply with the latest issue of SANS 1775 Compliance certificate to be submitted with the bid. EACH
37 | RT31-06-013 | EXTENSION SETS: Latex and DEHP Free, Suitable for use with administration set and an infusion pump. Minimal residual volume Spiral tube length: ± 2m ID of tubing: 1 to 2mm Supplied with one male and one female luer lock connector The tube must be non-kinking at any site and all temperatures. Sterile fluid pathway Latex and DEHP free: Latex and DEHP free to be stated on package. Individually packed, sterile, tamper proof. To comply with the latest issue of SANS 1775 Compliance certificate to be submitted with the bid. EACH
38 | RT31-06-014 | EXTENSION SETS: High Capacity Latex Free, Suitable for infusion of viscous solutions. Minimal residual volume Tube length: ± 1m Inner diameter (ID) of tubing: 3mm-5mm Supplied with one male and one female luer lock connector The tube must be non-kinking at any site and all temperatures. PVC material Sterile fluid pathway Latex free: Latex free to be stated on package. Individually packed, sterile, tamper proof. To comply with the latest issue of SANS 1775 Compliance certificate to be submitted with the bid. EACH
39 | RT31-06-015 | EXTENSION SETS: High Capacity Latex Free, For anaesthetic use Minimal residual volume Tube length: ±50cm Inner diameter (ID) of tubing: 3mm-5mm Supplied with one male and one female luer lock connector The tube must be non-kinking at any site and all temperatures. PVC material Sterile fluid pathway Latex free: Latex free to be stated on package. Individually packed, sterile, tamper proof. To comply with the latest issue of SANS 1775 Compliance certificate to be submitted with the bid. EACH
40 | RT31-06-016 | EXTENSION SETS: High Capacity Latex Free, For anaesthetic use Minimal residual volume Tube length: ± 100cm Inner diameter (ID) of tubing: 3mm-5mm Supplied with one male and one female luer lock connector The tube must be non-kinking at any site and all temperatures. Sterile fluid pathway Latex free: Latex free to be stated on package. Individually packed, sterile, tamper proof. To comply with the latest issue of SANS 1775 Compliance certificate to be submitted with the bid. EACH
41 | RT31-06-017 | EXTENSION SETS: High Capacity Latex and DEHP Free, For anaesthetic use Minimal residual volume Tube length: ± 100cm Inner diameter (ID) of tubing: 3mm-5mm Supplied with one male and one female luer lock connector The tube must be non-kinking at any site and all temperatures. Sterile fluid pathway Latex free: Latex free to be stated on package. Individually packed, sterile, tamper proof. To comply with the latest issue of SANS 1775 Compliance certificate to be submitted with the bid. EACH
42 | RT31-06-018 | EXTENSION SETS: High Capacity Latex and DEHP Free, For anaesthetic use Minimal residual volume Tube length: ±50cm Inner diameter (ID) of tubing: 3mm-5mm Supplied with one male and one female luer lock connector The tube must be non-kinking at any site and all temperatures. PVC material Sterile fluid pathway Latex and DEHP free: Latex and DEHP free to be stated on package. Individually packed, sterile, tamper proof. To comply with the latest issue of SANS 1775 Compliance certificate to be submitted with the bid. EACH
43 | RT31-06-019 | EXTENSION SETS: High Capacity Latex and DEHP Free, For anaesthetic use Minimal residual volume Tube length: ± 100cm Inner diameter (ID) of tubing: 3mm-5mm Supplied with one male and one female luer lock connector The tube must be non-kinking at any site and all temperatures. Sterile fluid pathway Latex and DEHP free: Latex and DEHP free to be stated on package. Individually packed, sterile, tamper proof. To comply with the latest issue of SANS 1775 Compliance certificate to be submitted with the bid. EACH
44 | RT31-06-020 | EXTENSION SETS: T- Connector Latexfree For use with infusions and side drips in neonates and infants in high care, theatre and intensive care units Length of tubing between connectors: ±130mm Flushing volume: ± 0,3ml (maximum) Supplied with one male luer slip connector with rubber cap (injection port) and one female luer lock connector with plastic screw cap. Slide clamp between the male connector and female connector Sterile fluid pathway Latex free: Latex free to be stated on package Individually packed, sterile, tamper proof. To comply with the latest issue of SANS 1775 Compliance certificate to be submitted with the bid. EACH
45 | RT31-06-021 | EXTENSION SETS: T- Connector Latex and DEHP Free, For use with infusions and side drips in neonates and infants in high care, theatre and intensive care units Length of tubing between connectors: ±130mm Flushing volume: ± 0,3ml (maximum) Supplied with one male luer slip connector with rubber cap (injection port) and one female luer lock connector with plastic screw cap. Slide clamp between the male connector and female connector Sterile fluid pathway Latex and DEHP free: Latex and DEHP free to be stated on package Individually packed, sterile, tamper proof. To comply with the latest issue of SANS 1775 Compliance certificate to be submitted with the bid. EACH
46 | RT31-06-023 | EXTENSION SETS: Double Lumen Latex Free, Suitable to be used with infusion pump. Minimal priming volume Proximal end female luer lock connector attached on a short tubing without a sliding clamp Distal end branches into two tubes 10 – 15 cm each with a sliding clamp. The clamps must be of two different colours Each of above tubes ends with a non-return, anti-siphon valve on a needle free connector. Sterile fluid pathway Latex free: Latex free to be stated on package Individually packed, sterile, tamper proof. To comply with the latest issue of SANS 1775 Compliance certificate to be submitted with the bid. EACH
47 | RT31-06-024 | EXTENSION SETS: Double Lumen Latex and DEHP Free, Suitable to be used with infusion pump. Minimal priming volume Proximal end female luer lock connector attached on a short tubing without a clamp Distal end branches into two tubes 10 – 15 cm each with a clamp. The clamps must be of two different colours Each of above tubes ends with a non-return, anti-siphon valve on a needle free connector. Sterile fluid pathway Latex and DEHP free: Latex and DEHP free to be stated on package Individually packed, sterile, tamper proof. To comply with the latest issue of SANS 1775 Compliance certificate to be submitted with the bid. EACH
48 | RT31-06-026 | TRIPLE LUMEN EXTENSION SET (Trifurcation Set), Latex and DEHP fre Suitable to be used with infusion pump. Minimal priming volume Proximal end female luer lock connector attached on a short tubing with a clamp Distal end branches into three tubes 10 – 15 cm each with a clamp. The clamps must be of three different colours Each of above tubes ends with a non-return, anti-siphon valve on a needle free connector. Sterile fluid pathway Latex and DEHP free: Latex and DEHP free to be stated on package Individually packed, sterile, tamper proof. To comply with the latest issue of SANS 1775 Compliance certificate to be submitted with the bid. EACH
49 | RT31-06-027 | EXTENSION SETS: Latex , Suitable for use with administration set and an infusion pump. Minimal residual volume Tube length: ± 1m Inner diameter (ID) of tubing: 1 to 2mm Supplied with one male and one female luer lock connector The tube must be non-kinking at any site and all temperatures. PVC material Sterile fluid pathway Latex : Latex to be stated on package. Individually packed, sterile, tamper proof. To comply with the latest issue of SANS 1775 Compliance certificate to be submitted with the bid. EACH
50 | RT31-06-028 | EXTENSION SETS: Latex , Suitable for use with administration set and an infusion pump. Minimal residual volume Tube length: ±2m Inner diameter (ID) of tubing: 1 to 2mm Supplied with one male and one female luer lock connector The tube must be non-kinking at any site and all temperatures. PVC material Sterile fluid pathway Latex: Latex to be stated on package. Individually packed, sterile, tamper proof. To comply with the latest issue of SANS 1775 Compliance certificate to be submitted with the bid. EACH
51 | RT31-06-029 | EXTENSION SETS: T- Connector Latexfree For use with infusions and side drips in neonates and infants in high care, theatre and intensive care units Length of tubing between connectors: 120mm - 160mm Flushing volume: ± 0,3ml (maximum) Supplied with one male luer slip connector with rubber cap (needle free injection port) and one female luer lock connector with plastic screw cap. Slide clamp between the male connector and female connector Sterile fluid pathway Latex free: Latex free to be stated on package Individually packed, sterile, tamper proof. To comply with the latest issue of SANS 1775 Compliance certificate to be submitted with the bid. EACH
52 | RT31-06-030 | EXTENSION SETS: T- Connector Latex and DEHP Free, For use with infusions and side drips in neonates and infants in high care, theatre and intensive care units Length of tubing between connectors: ±130mm Flushing volume: ± 0,3ml (maximum) Supplied with one male luer slip connector with rubber cap (needle free injection port) and one female luer lock connector with plastic screw cap. Slide clamp between the male connector and female connector Sterile fluid pathway Latex and DEHP free: Latex and DEHP free to be stated on package Individually packed, sterile, tamper proof. To comply with the latest issue of SANS 1775 Compliance certificate to be submitted with the bid. EACH
53 | RT31-07-003 | INFUSION SET, CONTINUOUS, WINGED, Latex free- Needle size: 1mm (19G) x 20mm Extension tube: 20-37,5cm Adaptor on extension tube must be female luer with a suitable cover Wings: Flexible colour coded to international standards Latex free: Latex free to be stated on package Individually packed, sterile, tamper proof. To comply with the latest issue of SANS 305 Compliance certificate to be submitted with the bid. EACH
54 | RT31-07-004 | INFUSION SET, CONTINUOUS, WINGED, Latex free- Needle size: 0,5mm (25G) x 20mm Extension tube: 20-37,5cm Adaptor on extension tube must be female luer with a suitable cover Wings: Flexible colour coded to international standards Latex free: Latex free to be stated on package Individually packed, sterile, tamper proof. To comply with the latest issue of SANS 305 Compliance certificate to be submitted with the bid. EACH
55 | RT31-07-005 | SAFETY WINGED INFUSION SET, CONTINUOUS, WINGED, Latex free- Needle size: 1mm (19G) x 20mm Extension tube: 20-37,5cm Adaptor on extension tube must be female luer with a suitable cover Wings: Flexible colour coded to international standards Latex free: Latex free to be stated on package Individually packed, sterile, tamper proof. To comply with the latest issue of SANS 305 Compliance certificate to be submitted with the bid. EACH
56 | RT31-07-006 | SAFETY WINGED INFUSION SET, CONTINUOUS, WINGED, Latex free- Needle size: 0,5mm (25G) x 20mm Extension tube: 20-37,5cm Adaptor on extension tube must be female luer with a suitable cover Wings: Flexible colour coded to international standards Latex free: Latex free to be stated on package Individually packed, sterile, tamper proof. To comply with the latest issue of SANS 305 Compliance certificate to be submitted with the bid. EACH
57 | RT31-09-001 | MONITORING SET, VENOUS PRESSURE, Latex free - Administration set: · With one needle Y-injection site. · Accurate flow regulating roller clamp. · The tube must be non-kinking at any site and all temperatures. · Drip volume: 20 drops/ml. · With a 180-220 micron filter, housed in the top chamber. · Distal luer lock male adaptor. Retractable Rotating male luer slip lock adaptor · Total length: ± 150cm. Central venous pressure measurement: · Three way stopcock. · The above 3-way stopcock must be connected on a tube with a flow indicator and the tube must end with an air vented bacterial protection. · 1 x graduated manometer tape Sterile fluid path way Latex free: Latex free to be stated on package Individually packed, sterile, tamper proof. To comply with the latest issue of SANS 1775 Compliance certificate to be submitted with the bid. EACH
58 | RT31-09-002 | MONITORING SET, VENOUS PRESSURE, Latex free- Administration set: · With needle free Y-injection site. · Accurate flow regulating roller clamp. · The tube must be non-kinking at any site and all temperatures. · Drip volume: 20 drops/ml. · With a 180-220 micron filter, housed in the top chamber. · Distal luer lock male adaptor. Retractable Rotating male luer slip lock adaptor · Total length: ± 150cm. Central venous pressure measurement: · Three way stopcock. · The above 3-way stopcock must be connected on a tube with a flow indicator and the tube must end with an air vented bacterial protection. · 1 x graduated manometer tape Sterile fluid path way Latex free: Latex free to be stated on package Individually packed, sterile, tamper proof. To comply with the latest issue of SANS 1775 Compliance certificate to be submitted with the bid. EACH
59 | RT31-12-001 | IRRIGATION SET (TUR Irrigation Set), Latex free- For urology use Proximal end: · 2 x spike connections, each connected on ± 50cm tube, with one on/off clamp on each · The 50 cm tubes are joined in a Y-connection to ends distal in one tube. · The tubes must have an inner diameter (ID) of large bore tubing for delivering high flow rates, · The tubing must be non-kinking at any site and all temperatures. Drip volume: 20 drops/ml. Distal luer lock male adaptor. Retractable Rotating male luer slip lock adaptor Total length: ±150cm. Latex free: Latex free to be stated on package. Sterile fluid pathway. Individually packed, sterile, tamper proof. To comply with the latest issue of SANS 1775. Compliance certificate to be submitted with the bid. EACH
60 | RT31-12-002 | IRRIGATION SET , Latex free- For urology use Proximal end: • 1 x spike connections, each connected on ± 50cm tube, with one on/off clamp • The tubes must have an inner diameter (ID) of large bore tubing for delivering high flow rates, • The tubing must be non-kinking at any site and all temperatures. Drip volume: 20 drops/ml. Distal luer lock male adaptor. Retractable Rotating male luer slip lock adaptor Total length: ±150cm. Latex free: Latex free to be stated on package. Sterile fluid pathway. Individually packed, sterile, tamper proof. To comply with the latest issue of SANS 1775. Compliance certificate to be submitted with the bid. EACH