## ANNEXURE A TECHNICAL SPECIFICAT
BID NUMBER: RT42-2025
DESCRIPTION: SUPPLY AND DELIVERY OF BANDAGES AND DRESSING TO THE STATE FOR THE PERIOD OF 60 MONTHS
ITEM NUMBER | ITEM DESCRIPTION | LOCAL PRODUCATION AND CONTENT % | NUMBER OF SAMPLES TO BE SUBMITTED FOR VISUAL SCREENING | CONFORM WITH SPECIFICAITON? YES OR NO | IF NO, PLEASE STATE THE DEVIATION
RT42-01-001 | Dressing, burn, first aid non-adherent, each Sterile dressing for emergency or primary application to list 1st, 2nd, 3rd degree burns and scalds Size: 10cm x 10cm Properties of dressing: water based, absorb and dissipate heat, reduce burn progression, limit skin damage, inhibit fluid loss Properties of Gel: colourless to white, non-oily, contain 90 - 98% water, non-toxic, hypo-allergenic, low irritant, even to eyes, viscosity: 15 000 to 30 000 cPs pH: 5.5 to 7.0 Each | 100% | 3 samples in an original packaging
RT42-01-002 | Dressing, burn, first aid, non-adherent, each Sterile dressing for emergency or primary application to list 1st, 2nd, 3rd degree burns and scalds Size: 20cm x 20cm Properties of dressing: water based, absorb and dissipate heat reduce burn progression, limit skin damage, inhibit fluid loss Properties of Gel: colourless to white, non-oily, contain 90-98% water, non-toxic, hypo-allergenic, low irritant, even to eyes, bacteriostatic viscosity: 15 000 - 30 000 cPs pH: 5.5 to 7.0 Label to state bacteriostatic Each | 100% | 3 samples in an original packaging
RT42-01-003 | Dressing, burn, first aid, non-adherent, each Sterile dressing for emergency or primary application to list 1st, 2nd, 3rd degree burns and scalds Size: 60cm x 40cm Properties of dressing: water based, absorb and dissipate heat, reduce burn progression, limit skin damage, inhibit fluid loss Properties of Gel: colourless to white, non-oily, contain 90-98% water, non-toxic, hypo-allergenic, low irritant, even to eyes, viscosity: 15 000 - 30 000 cPs pH: 5.5 to 7.0 Each | 100% | 3 samples in an original packaging
RT42-01-004 | Dressing, burn, first aid, multiple kit, Sterile dressing for pre-medical application to 1st, 2nd and 3rd degree burns and scalds Kit = 6 dressing: 2 x 100mm x 100mm 2 x 200mm x 200mm 2 x 600mm x 400mm Properties of dressing: water based, absorb and dissipate heat, reduce burn progression, limit skin damage, inhibit fluid loss Properties of Gel: colourless to white, non-oily, contain 96% water, non-toxic, hypo-allergenic, low irritant, even to eyes, viscosity: 15 000 - 30 000 cPs pH: 5.5 to 7.0 Kit | 100% | 3 samples in an original packaging
RT42-01-005 | Dressing, burn, occlusive, hydrogel sheet, non-adherent, each Sterile dressing for application to 1st, 2nd degree burns and low to moderate draining wounds Size: 150cm x 225cm Properties of dressing: * water based absorb and dissipate heat reduce burn progression * limit skin damage * inhibit fluid loss Properties of Gel: * Transparent, colourless to white * Non-oily * Contain 90-98% water * Non-toxic, hypo allergenic * low irritant even to eyes Each | 100% | 3 samples in an original packaging
RT42-01-006 | Dressing, burn, first aid non-adherent, each Sterile dressing for emergency or primary application to list 1st, 2nd, 3rd degree burns and scalds Size: 5cm x15cm Properties of dressing: water based, absorb and dissipate heat, reduce burn progression, limit skin damage, inhibit fluid loss Properties of Gel: colourless to white, non-oily, contain 90-98% water, non-toxic, hypo-allergenic, low irritant, even to eyes viscosity: 15 000 to 30 000 cPs pH: 5.5 to 7.0 Each | 100% | 3 samples in an original packaging
RT42-01-007 | Dressing, burn, first aid non-adherent, each Sterile dressing for emergency or primary application to list 1st, 2nd, 3rd degree burns and scalds Size: 10cm x 40cm Properties of dressing: water based, absorb and dissipate heat, reduce burn progression, limit skin damage, inhibit fluid loss Properties of Gel: colourless to white, non-oily, contain 90-98% water, non-toxic, hypo-allergenic, low irritant, even to eyes, viscosity: 15 000 to 30 000 cPs pH: 5.5 to 7.0 Each | 100% | 3 samples in an original packaging
RT42-01-008 | Dressing, burn, first aid non-adherent, each Sterile dressing for emergency or primary application to list 1st, 2nd, 3rd degree burns and scalds Size: 20cm x 45cm Properties of dressing: water based, absorb and dissipate heat, reduce burn progression, limit skin damage, inhibit fluid loss Properties of Gel: colourless to white, non-oily, contain 90-98% water, non-toxic, hypo-allergenic, low irritant, even to eyes, viscosity: 15 000 to 30 000 cPs pH: 5.5 to 7.0 Each | 100% | 3 samples in an original packaging
RT42-01-009 | Dressing, burn, first aid non-adherent, each Sterile dressing for emergency or primary application to list 1st, 2nd, 3rd degree burns and scalds Size: 20cm x 55cm (hand dressing) Properties of dressing: water based, absorb and dissipate heat, reduce burn progression, limit skin damage, inhibit fluid loss Properties of Gel: colourless to white, non-oily, contain 90-98% water, non-toxic, hypo-allergenic, low irritant, even to eyes, viscosity: 15 000 to 30 000 cPs pH: 5.5 to 7.0 Each | 100% | 3 samples in an original packaging
RT42-01-010 | Dressing, burn, first aid non-adherent, each Sterile dressing for emergency or primary application to list 1st, 2nd, 3rd degree burns and scalds Size: 30cm x 40cm (face mask) Properties of dressing: water based, absorb and dissipate heat, reduce burn progression, limit skin damage, inhibit fluid loss Properties of Gel: colourless to white, non-oily, contain 90-98% water, non-toxic, hypo-allergenic, low irritant, even to eyes, viscosity: 15 000 to 30 000 cPs pH: 5.5 to 7.0 Each | 100% | 3 samples in an original packaging
RT42-01-011 | Dressing, burn, first aid non-adherent, each Sterile dressing, wrap, for emergency or primary application to list 1st, 2nd, 3rd degree burns and scalds Size: ±1m x 1m Properties of dressing: water based, absorb and dissipate heat, reduce burn progression, limit skin damage, inhibit fluid loss Properties of Gel: colourless to white, non-oily, contain 90-98% water, non-toxic, hypo-allergenic, low irritant, even to eyes, viscosity: 15 000 to 30 000 cPs pH: 5.5 to 7.0 Each | 100% | 3 samples in an original packaging
RT42-01-012 | Dressing, burn, first aid non-adherent, each Sterile dressing, wrap, for emergency or primary application to list 1st, 2nd, 3rd degree burns and scalds Size: ±244cm x 183cm Properties of dressing: water based, absorb and dissipate heat, reduce burn progression, limit skin damage, inhibit fluid loss Properties of Gel: colourless to white, non-oily, contain 90-98% water, non-toxic, hypo-allergenic, low irritant, even to eyes, viscosity: 15 000 to 30 000 cPs pH: 5.5 to 7.0 Each | 100% | 3 samples in an original packaging
RT42-01-013 | Dressing, burn, first aid non-adherent, each Sterile dressing for emergency or primary application to list 1st, 2nd, 3rd degree burns and scalds Size: 50cm x 1m Properties of dressing: water based, absorb and dissipate heat, reduce burn progression, limit skin damage, inhibit fluid loss Properties of Gel: colourless to white, non-oily, contain 96% water, non-toxic, hypo-allergenic, low irritant, even to eyes, viscosity: 15 000 to 30 000 cPs pH: 5.5 to 7.0 Each | 100% | 3 samples in an original packaging
RT42-01-014 | Dressing, standard, first aid, sterile Consisting of an absorbent cotton wool pad to comply with the latest issue of SANS 228, enclosed in absorbent woven gauze to comply with the latest issue of SANS 446, attached to a conforming crepe or knitted type bandage to comply with the latest issue of SANS 1427 Compliance certificate to be submitted with bid Size: 50mm x 75mm pad 50mm x 1m bandage Sterile, individually peeled packed Kit | Not applicable | 3 samples in an original packaging
RT42-01-015 | Dressing, standard, first aid, sterile Consisting of an absorbent cotton wool pad to comply with the latest issue of SANS 228, enclosed in absorbent woven gauze to comply with the latest issue of of SANS 446, attached to a conforming crepe or knitted type bandage to comply with the latest issue of SANS 1427 Compliance certificate to be submitted with bid Size: 75mm x 100mm pad; 75mm x 2m bandage Sterile, individually peeled packed Kit | Not applicable | 3 samples in an original packaging
RT42-01-016 | Dressing, standard, first aid, sterile Consisting of an absorbent cotton wool pad to comply with the latest issue of SANS 228, enclosed in absorbent woven gauze o comply with the latest issue of of SANS 446, attached to a conforming crepe or knitted type bandage to comply with the latest issue of SANS 1427 Compliance certificate to be submitted with bid Size: 150mm x 200mm pad, 100mm x 4.5m bandage Sterile, individually peeled packed Kit | Not applicable | 3 samples in an original packaging
RT42-01-017 | Dressing, standard, first aid, sterile, Consisting of an absorbent cotton wool pad to comply with the latest issue of SANS 228, enclosed in absorbent woven gauze to comply with the latest issue of of SANS 446, attached to a conforming crepe or knitted type bandage to comply with the latest issue of SANS 1427 Compliance certificate to be submitted with bid Size: 200mm x 300mm pad, 200mm x 4.5m bandage Sterile, individually peeled packed Kit | Not applicable | 3 samples in an original packaging
RT42-01-018 | Dressing, burn, occlusive, hydrogel sheet, non-adherent, each Sterile dressing for application to 1st, 2nd degree burns and low to moderate draining wounds Size: 95mm x 95mm Properties of dressing: * water based absorb and dissipate heat * reduce burn progression * limit skin damage * inhibit fluid loss Properties of Gel: * transparent, colourless to white * non-oily * contain 90% water * non-toxic, hypo-allergenic * low irritant, even to eyes *bacteriostatic Each | 100% | 3 samples in an original packaging
RT42-01-019 | Dressing for burn wounds. Dressing must have open fingers which is an oversized glove. Dressing must have an antimicrobial silver contact wound barrier. Dressing should be able to release a continues flow of silver ions to wound. Dressing should be able to stayed for up to 7 days in place Latex Free individually wrapped Size: Extra Small Each | Not applicable | 3 samples in an original packaging
RT42-01-020 | Dressing for burn wounds. Dressing must have open fingers which is an oversized glove. Dressing must have an antimicrobial silver contact wound barrier. Dressing should be able to release a continues flow of silver ions to wound. Dressing should be able to stayed for up to 7 days in place Latex Free individually wrapped Size: Small Each | Not applicable | 3 samples in an original packaging
RT42-01-021 | Dressing for burn wounds. Dressing must have open fingers which is an oversized glove. Dressing must have an antimicrobial silver contact wound barrier. Dressing should be able to release a continues flow of silver ions to wound. Dressing should be able to stayed for up to 7 days in place Latex Free individually wrapped Size: Medium Each | Not applicable | 3 samples in an original packaging
RT42-01-022 | Dressing for burn wounds. Dressing must have open fingers which is an oversized glove. Dressing must have an antimicrobial silver contact wound barrier. Dressing should be able to release a continues flow of silver ions to wound. Dressing should be able to stayed for up to 7 days in place Latex Free individually wrapped Size: Large Each | Not applicable | 3 samples in an original packaging
RT42-01-023 | Dressing for burn wounds. Dressing must have open fingers which is an oversized glove. Dressing must have an antimicrobial silver contact wound barrier. Dressing should be able to release a continues flow of silver ions to wound. Dressing should be able to stayed for up to 7 days in place Latex Free individually wrapped Size: X-large Each | Not applicable | 3 samples in an original packaging
RT42-01-024 | Dressing strip, fabric, elastic adhesive with a soft non-adherent, absorbent wound padding , B.P. Size: 40 mm x 1 m Each | Not applicable | 3 samples in an original packaging
RT42-01-025 | Dressing strip, fabric, elastic adhesive with a soft non-adherent, absorbent wound padding , B.P. Size: 65 mm x 1 m Each | Not applicable | 3 samples in an original packaging
RT42-01-026 | Dressing, absorbent gauze, type 1, bleached, 100 dressings Non sterile, autoclavable Size: 225mm x 225mm x 16ply To comply with the latest issue of SANS 446 Compliance certificate to be submitted with bid Packet | 100% | 3 samples in an original packaging
RT42-01-027 | Dressing, paraffin gauze, B.P., 10 dressings in a tray Size: 100mm x 100mm Tray | 100% | 3 samples in an original packaging
RT42-01-028 | Dressing, paraffin gauze, B.P., in a tray Size: 100mm x 7m One dressing per tray Tray | 100% | 3 samples in an original packaging
RT42-01-029 | Dressing, paraffin gauze, B.P., 10 dressings Size: 100mm x 100mm Sterile individually peel packed Box | 100% | 3 samples in an original packaging
RT42-01-030 | Dressing, paraffin gauze B.P. 10 dressings Size: 100mm x 400mm Sterile, individuallly peel packed Box | 100% | 3 samples in an original packaging
RT42-01-031 | Dressing, ribbon gauze, cotton fabric with antimicrobial hydrophobic coating (e.g. dialkyl carbamoyl chloride) that physically binds and removes bacteria, size 20mm x 500mm, packaging: roll, individually wrapped, sterile. Roll | Not applicable | 3 samples in an original packaging
RT42-01-032 | Dressing, ribbon gauze, cotton fabric with antimicrobial hydrophobic coating (e.g. dialkyl carbamoyl chloride) that physically binds and removes bacteria, size 50mm x 2000mm, packaging: roll, individually wrapped, sterile. Roll | Not applicable | 3 samples in an original packaging
RT42-01-033 | Dressing Gauze, Impregnated with Honey, non-adherent.. Sterile, individually peel packed Size of Dressing: ±50mm x 50mm Each | Not applicable | 3 samples in an original packaging
RT42-01-034 | Dressing, Gauze, Impregnated with Honey, non-adherent.. Sterile, individually peel packed Size of Dressing: ±100mm x 100mm Each | Not applicable | 3 samples in an original packaging
RT42-01-035 | Dressing, gauze, non-woven, with haemostatic properties trauma pad Size: 100mm X 100mm Individually packed Sterile Box of 10 dressings | Not applicable | 3 samples in an original packaging
RT42-01-036 | Dressing, tracheostomy, absorbent, non-adherent, waterproof Size: 100mm x 100mm Dressing to be: easy to apply and removable Sterile, individually peel packed Must be clearly marked "FOR TRACHEOSTOMY" Each | Not applicable | 3 samples in an original packaging
RT42-01-037 | Dressing, tulle gras, 10 dressings Impregnated with chlorhexidine 0.5% m/m or other bacterial properties Size: 100mm x 100mm Sterile, individually peel packed Box | Not applicable | 3 samples in an original packaging
RT42-01-038 | Dressing, tulle grass Impregnated with chlorhexidine 0.5% m/m or other bacterial properties Size: 100mm x 400mm Sterile, individually peel packed Each | Not applicable | 3 samples in an original packaging
RT42-01-039 | Dressing, film, transparent, barrier material, For intravenous catheter placement Barrier material to be: tear resistant, conform to irregular surfaces, antistatic, permiable to oxygen and water vapour and waterproof Adhesive: low-irritant hypo-allergenic pressure sensitive, to adhere firmly so as to keep catheter in position covered with protective silicone treated release backing Packing: Must include instructions for use Sterile, individually peel packed with description on outer pack Size: 60mm x 70mm Each | 100% | 3 samples in an original packaging
RT42-01-040 | Dressing, film, transparent, barrier material, For intravenous catheter placement Barrier material to be: tear resistant, conform to irregular surfaces, antistatic, permiable to oxygen and water vapour and waterproof Adhesive: low-irritant hypo-allergenic pressure sensitive, to adhere firmly so as to keep catheter in position covered with protective silicone treated release backing Packing: Must include instructions for use Sterile, individually peel packed with description on outer pack Size: 100mm x 120mm Each | 100% | 3 samples in an original packaging
RT42-01-041 | Dressing, film, transparent, barrier material, For intravenous catheter placement Barrier material to be: tear resistant, conform to irregular surfaces, antistatic, permiable to oxygen and water vapour and waterproof Adhesive: low-irritant hypo-allergenic pressure sensitive, to adhere firmly so as to keep catheter in position covered with protective silicone treated release backing Packing: Must include instructions for use Sterile, individually peel packed with description on outer pack Size: 100mm x 200mm Each | 100% | 3 samples in an original packaging
RT42-01-042 | Dressing, film, transparent, barrier material, For intravenous catheter placement Barrier material to be: tear resistant, conform to irregular surfaces, antistatic, permiable to oxygen and water vapour and waterproof Adhesive: low-irritant hypo-allergenic pressure sensitive, to adhere firmly so as to keep catheter in position covered with protective silicone treated release backing Packing: Must include instructions for use Sterile, individually peel packed with description on outer pack Size: 100mm x 250mm Each | 100% | 3 samples in an original packaging
RT42-01-043 | Dressing, film, transparent, barrier material, For intravenous catheter placement Barrier material to be: tear resistant, conform to irregular surfaces, antistatic, permiable to oxygen and water vapour and waterproof Adhesive: low-irritant hypo-allergenic pressure sensitive, to adhere firmly so as to keep catheter in position covered with protective silicone treated release backing Packing: Must include instructions for use Sterile, individually peel packed with description on outer pack Size: 150mm x 200mm Each | 100% | 3 samples in an original packaging
RT42-01-044 | Dressing, film, transparent, barrier material, For intravenous catheter placement Barrier material to be: tear resistant, conform to irregular surfaces, antistatic, permiable to oxygen and water vapour and waterproof Adhesive: low-irritant hypo-allergenic pressure sensitive, to adhere firmly so as to keep catheter in position covered with protective silicone treated release backing Packing: Must include instructions for use Sterile, individually peel packed with description on outer pack Size: 200mm x 300mm Each | 100% | 3 samples in an original packaging
RT42-01-045 | Dressing, film, transparent, Island Dressing, Central Island to be: highly absorbent, non-adherent pad Barrier material to be: transparent, matt finished and tear resistant conformable to irregular surfaces antistatic permiable to oxygen and water vapour waterproof Adhesive: low-irritant hypo-allergenic pressure sensitive, to adhere firmly so as to keep catheter in position covered with protective silicone treated release backing two tape handles to facilitate easy tear off after positioning Packing: must include instructions for use on outer pack Sterile, individually peel packed with description on outer pack Size Dressing: 50mm x 70mm Pad Size: 50mm x 35mm Each | 100% | 3 samples in an original packaging
RT42-01-046 | Dressing, film, transparent, Island Dressing, Central Island to be: highly absorbent, non-adherent pad Barrier material to be: transparent, matt finished and tear resistant conformable to irregular surfaces antistatic permiable to oxygen and water vapour waterproof Adhesive: low-irritant hypo-allergenic pressure sensitive, to adhere firmly so as to keep catheter in position covered with protective silicone treated release backing two tape handles to facilitate easy tear off after positioning Packing: must include instructions for use on outer pack Sterile, individually peel packed with description on outer pack Size Dressing: 83mm x 60mm Pad Size: 50mm x 35mm Each | 100% | 3 samples in an original packaging
RT42-01-047 | Dressing, film, transparent, Island Dressing, Central Island to be: highly absorbent, non-adherent pad Barrier material to be: transparent, matt finished and tear resistant conformable to irregular surfaces antistatic permiable to oxygen and water vapour waterproof Adhesive: low-irritant hypo-allergenic pressure sensitive, to adhere firmly so as to keep catheter in position covered with protective silicone treated release backing two tape handles to facilitate easy tear off after positioning Packing: must include instructions for use on outer pack Sterile, individually peel packed with description on outer pack Size Dressing: 83mm x 240mm Pad Size: 45mm x 170mm Each | 100% | 3 samples in an original packaging
RT42-01-048 | Dressing, film, transparent, Island Dressing, Central Island to be: highly absorbent, non-adherent pad Barrier material to be: transparent, matt finished and tear resistant conformable to irregular surfaces antistatic permiable to oxygen and water vapour waterproof Adhesive: low-irritant hypo-allergenic pressure sensitive, to adhere firmly so as to keep catheter in position covered with protective silicone treated release backing two tape handles to facilitate easy tear off after positioning Packing: must include instructions for use on outer pack Sterile, individually peel packed with description on outer pack Size Dressing: 100mm x 60mm Pad Size: 50mm x 35mm Each | 100% | 3 samples in an original packaging
RT42-01-049 | Dressing, film, transparent, Island Dressing, Central Island to be: highly absorbent, non-adherent pad Barrier material to be: transparent, matt finished and tear resistant conformable to irregular surfaces antistatic permiable to oxygen and water vapour waterproof Adhesive: low-irritant hypo-allergenic pressure sensitive, to adhere firmly so as to keep catheter in position covered with protective silicone treated release backing two tape handles to facilitate easy tear off after positioning Packing: must include instructions for use on outer pack Sterile, individually peel packed with description on outer pack Size Dressing: 100mm x 120mm Pad Size: 45mm x 75mm Each | 100% | 3 samples in an original packaging
RT42-01-050 | Dressing, film, transparent, Island Dressing, Central Island to be: highly absorbent, non-adherent pad Barrier material to be: transparent, matt finished and tear resistant conformable to irregular surfaces antistatic permiable to oxygen and water vapour waterproof Adhesive: low-irritant hypo-allergenic pressure sensitive, to adhere firmly so as to keep catheter in position covered with protective silicone treated release backing two tape handles to facilitate easy tear off after positioning Packing: must include instructions for use on outer pack Sterile, individually peel packed with description on outer pack Size Dressing: 100mm x 200mm Pad Size: 45mm x 150mm Each | 100% | 3 samples in an original packaging
RT42-01-051 | Dressing, film, transparent, Island Dressing, Central Island to be: highly absorbent, non-adherent pad Barrier material to be: transparent, matt finished and tear resistant conformable to irregular surfaces antistatic permiable to oxygen and water vapour waterproof Adhesive: low-irritant hypo-allergenic pressure sensitive, to adhere firmly so as to keep catheter in position covered with protective silicone treated release backing two tape handles to facilitate easy tear off after positioning Packing: must include instructions for use on outer pack Sterile, individually peel packed with description on outer pack Size Dressing: 100mm x 240mm Pad Size: 45mm x 170mm Each | 100% | 3 samples in an original packaging
RT42-01-052 | Dressing, film, transparent, Island Dressing, Central Island to be: highly absorbent, non-adherent pad Barrier material to be: transparent, matt finished and tear resistant conformable to irregular surfaces antistatic permiable to oxygen and water vapour waterproof Adhesive: low-irritant hypo-allergenic press covered with protective silicone treated release backing two tape handles to facilitate easy tear off after positioning Packing: must include instructions for use on outer pack Sterile, individually peel packed with description on outer pack Size Dressing: 100mm x 350mm Pad Size: 45mm x 270mm Each | 100% | 3 samples in an original packaging
RT42-01-053 | Dressing, transparent film, post-operative/wound, adhesive, hypoallergenic, allows visibility of incision site, Sterile with a central see through absorbent pad, on-adherent pad will not disturb staples, sutures or surgical glue, Sterile, individually peel packed with description peel packed with description on outer pack Size Dressing: 80mm x 100mm Each | 100% | 3 samples in an original packaging
RT42-01-054 | Dressing, transparent film, post-operative/wound, adhesive, hypoallergenic, allows visibility of incision site, Sterile with a central see through absorbent pad, on-adherent pad will not disturb staples, sutures or surgical glue, Sterile, individually peel packed with description peel packed with description on outer pack Size Dressing: 100mm x 300mm Each | Not applicable | 3 samples in an original packaging
RT42-01-055 | Dressing, transparent film, post-operative/wound, adhesive, hypoallergenic, allows visibility of incision site, Sterile with a central see through absorbent pad, on-adherent pad will not disturb staples, sutures or surgical glue, Sterile, individually peel packed with description peel packed with description on outer pack Size Dressing: 150mm x 100mm Each | 100% | 3 samples in an original packaging
RT42-01-056 | Dressing, transparent film, post-operative/wound, adhesive, hypoallergenic, allows visibility of incision site, Sterile with a central see through absorbent pad, on-adherent pad will not disturb staples, sutures or surgical glue, Sterile, individually peel packed with description peel packed with description on outer pack Size Dressing: 200mm x 100mm Each | 100% | 3 samples in an original packaging
RT42-01-057 | Dressing, transparent film, post-operative/wound, adhesive, hypoallergenic, allows visibility of incision site, Sterile with a central see through absorbent pad, on-adherent pad will not disturb staples, sutures or surgical glue, Sterile, individually peel packed with description peel packed with description on outer pack Size Dressing: 250mm x 100mm Each | 100% | 3 samples in an original packaging
RT42-01-058 | Dressing, transparent film, post-operative/wound, adhesive, hypoallergenic, allows visibility of incision site, Sterile with a central see through absorbent pad, on-adherent pad will not disturb staples, sutures or surgical glue, Sterile, individually peel packed with description peel packed with description on outer pack Size Dressing: 350mm x 100mm Each | 100% | 3 samples in an original packaging
RT42-01-059 | Dressing, post-operative/wound, conformable, adhesive, hypoallergenic, Sterile with a central highly absorbent non-adherent non-striKe through pad must include instructions for use on outer pack Sterile, individually peel packed with description on outer pack Size Dressing: 83mm x 60mm Pad Size: 50mm x 35mm Each | 100% | 3 samples in an original packaging
RT42-01-060 | Dressing, post-operative/wound, conformable, adhesive, hypoallergenic, Sterile with a central highly absorbent non-adherent non-striKe through pad must include instructions for use on outer pack Sterile, individually peel packed with description on outer pack Size Dressing: 83mm x 120mm Pad Size: 45mm x 75mm Each | 100% | 3 samples in an original packaging
RT42-01-061 | Dressing, post-operative/wound, conformable, adhesive, hypoallergenic, Sterile with a central highly absorbent non-adherent non-striKe through pad must include instructions for use on outer pack Sterile, individually peel packed with description on outer pack Size Dressing: 83mm x 180mm Pad Size: 45mm x 120mm Each | 100% | 3 samples in an original packaging
RT42-01-062 | Dressing, post-operative/wound, conformable, adhesive, hypoallergenic, Sterile with a central highly absorbent non-adherent non-striKe through pad Must include instructions for use on outer pack Sterile, individually peel packed with description on outer pack Size Dressing: 83mm x 240mm Pad Size: 45mm x 170mm Each | 100% | 3 samples in an original packaging
RT42-01-063 | Dressing, post-operative/wound, conformable, adhesive, hypoallergenic, Sterile with a central highly absorbent non-adherent non-striKe through pad Must include instructions for use on outer pack Sterile, individually peel packed with description on outer pack Size Dressing: 83mm x 415mm Pad Size: 45mm x 345mm Each | 100% | 3 samples in an original packaging
RT42-01-064 | Dressing, post-operative/wound, conformable, adhesive, hypoallergenic, Sterile with a central highly absorbent non-adherent non-striKe through pad Must include instructions for use on outer pack Sterile, individually peel packed with description on outer pack Size Dressing: 100mm x 300mm Pad Size: 45mm x 250mm Each | 100% | 3 samples in an original packaging
RT42-01-065 | Dressing Wound, absorbent, self-adherent Highly absorbent centre layer to accommodate large volume exudate and odour, layer of dressing to be waterproof. Outer layer waterproof, tapered borders to prevent leakage. Non-adherent wound contact layer. Sterile, individually peel packed Size of Dressing: 75mm x 75mm Each | Not applicable | 3 samples in an original packaging
RT42-01-066 | Dressing Wound, absorbent, self-adherent Highly absorbent centre layer to accommodate large volume exudate and odour, layer of dressing to be waterproof. Outer layer waterproof, tapered borders to prevent leakage. Non-adherent wound contact layer.. Sterile, individually peel packed Size of Dressing: 100mm x 100mm Each | Not applicable | 3 samples in an original packaging
RT42-01-067 | Dressing Wound, absorbent, self-adherent Highly absorbent centre layer to accommodate large volume exudate and odour, layer of dressing to be waterproof. Outer layer waterproof, tapered borders to prevent leakage. Non-adherent wound contact layer. Sterile, individually peel packed Size of Dressing: 125mm x 125mm Each | Not applicable | 3 samples in an original packaging
RT42-01-068 | Dressing Wound, absorbent, self-adherent Highly absorbent centre layer to accommodate large volume exudate and odour, layer of dressing to be waterproof. Outer layer waterproof, tapered borders to prevent leakage. Non-adherent wound contact layer. Sterile, individually peel packed Size of Dressing: 150mm x 150mm Each | Not applicable | 3 samples in an original packaging
RT42-01-069 | Dressing Wound, absorbent, self-adherent Highly absorbent centre layer to accommodate large volume exudate and odour, layer of dressing to be waterproof. Outer layer waterproof, tapered borders to prevent leakage. Non-adherent wound contact layer. Sterile, individually peel packed Size of Dressing:180mm x 180mm Each | Not applicable | 3 samples in an original packaging
RT42-01-070 | Dressing Wound, absorbent, self-adherent, Heel Highly absorbent centre layer to accommodate large volume exudate and odour, layer of dressing to be waterproof. Outer layer waterproof, tapered boders to prevent leakage. Non-adherent wound contact layer. Sterile, individually peel packed Size of Dressing: Each | Not applicable | 3 samples in an original packaging
RT42-01-071 | Dressing Wound, absorbent, self-adherent, Sacral Highly absorbent centre layer to accommodate large volume exudate and odour, layer of dressing to be waterproof. Outer layer waterproof, tapered borders to prevent leakage. Non-adherent wound contact layer. Sterile, individually peel packed Size: 170mm x 170mm Each | Not applicable | 3 samples in an original packaging
RT42-01-072 | Dressing, self-adhesive, highly absorbent polyurethane foam core, acetate fabric with antimicrobial hydrophobic coating that physically binds and removes bacteria, size 75mm x 75mm, packaging individually wrapped, sterile. Each | 100% | 3 samples in an original packaging
RT42-01-073 | Dressing, self-adhesive, highly absorbent polyurethane foam core, acetate fabric with antimicrobial hydrophobic coating that physically binds and removes bacteria, size 125mm x 125mm, packaging individually wrapped, sterile. Each | 100% | 3 samples in an original packaging
RT42-01-074 | Dressing, self-adhesive, highly absorbent polyurethane foam core, acetate fabric with antimicrobial hydrophobic coating that physically binds and removes bacteria, size 150mm x 150mm, packaging individually wrapped, sterile. Each | 100% | 3 samples in an original packaging
RT42-01-075 | Dressing, self-adhesive, highly absorbent polyurethane foam core, acetate fabric with antimicrobial hydrophobic coating that physically binds and removes bacteria, size 175mm x 175mm, packaging individually wrapped, sterile. Each | 100% | 3 samples in an original packaging
RT42-01-076 | Dressing, self-adhesive, highly absorbent polyurethane foam core, acetate fabric with antimicrobial hydrophobic coating that physically binds and removes bacteria, size 225 x 225mm, packaging individually wrapped, sterile. Each | 100% | 3 samples in an original packaging
RT42-01-077 | Dressing, self-adhesive, absorbent hydropolymer core, acetate fabric with antimicrobial hydrophobic coating that physically binds and removes bacteria, size 190mm x 190mm, packaging individually wrapped, sterile. Each | 100% | 3 samples in an original packaging
RT42-01-078 | Dressing, self-adhesive, absorbent hydropolymer core, acetate fabric with antimicrobial hydrophobic coating that physically binds and removes bacteria, size 140mm x 240mm, packaging individually wrapped, sterile. Each | 100% | 3 samples in an original packaging
RT42-01-079 | Dressing Wound, absorbent, with adhesive border Highly absorbent centre layer to accommodate large volume exudate and odour, polypropylene/polyethylene wound contact layer,non-adherent. Outerlayer of dressing to be waterproof. Sterile, individually peel packed Size Dressing: 150mm x 150mm Each | Not applicable | 3 samples in an original packaging
RT42-01-080 | Dressing Wound, absorbent, with adhesive border Highly absorbent centre layer to accommodate large volume exudate and odour, polypropylene/polyethylene wound contact layer,non-adherent. Outerlayer of dressing to be waterproof. Sterile, individually peel packed Size Dressing: 200mm x 200mm Each | Not applicable | 3 samples in an original packaging
RT42-01-081 | Dressing, wound Non-adherent dressing polypropylene/polythylene wound contact layer Size: 100mm x 100mm x 4ply Sterile, individually peel packed Each | Not applicable | 3 samples in an original packaging
RT42-01-082 | Dressing wound, non-adherent, Size: ±105mm x, 105mm, absorbent, contains activated charcoal, sterile,individually peel packed. Each | Not applicable | 3 samples in an original packaging
RT42-01-083 | Dressing, wound, non-adherent, fenestrated of highly absorbant hydrophyllic antimicrobial foam is impregnated with polyhexamethylene biguanide Individually wrapped Sterile Size: 100mm x ±100mm Box of 5 dressings | Not applicable | 3 samples in an original packaging
RT42-01-084 | Dressing, wound, non-adherent, fenestrated of highly absorbant hydrophyllic antimicrobial foam is impregnated with polyhexamethylene biguanide Individually wrapped Sterile Size: 100mm x ±200mm Box of 5 dressings | Not applicable | 3 samples in an original packaging
RT42-01-085 | Dressing, wound, non-adherent, fenestrated of highly absorbant hydrophyllic antimicrobial foam is impregnated with polyhexamethylene biguanide Individually wrapped Sterile Size: 200mm x ±200mm Box of 5 dressings | Not applicable | 3 samples in an original packaging
RT42-01-086 | Dressing, wound, non-adherent, fenestrated of highly absorbant hydrophyllic antimicrobial foam is impregnated with polyhexamethylene biguanide Individually wrapped Sterile Size: 50mm x ±50mm Box of 5 dressings | Not applicable | 3 samples in an original packaging
RT42-01-087 | Dressing Wound, Povidone-Iodine impregnated, non-adherent Sterile, individually peel packed Size of Dressing: 50mm x 50mm Each | Not applicable | 3 samples in an original packaging
RT42-01-088 | Dressing Wound, Povidone-Iodine impregnated, non-adherent Sterile, individually peel packed Size of Dressing: 100mm x 100mm Each | 100% | 3 samples in an original packaging
RT42-01-089 | Dressing Wound, Honey, non-adherent. Sterile, individually peel packed Size of Dressing: 100mm x 100mm Each | Not applicable | 3 samples in an original packaging
RT42-01-090 | Dressing Wound, Honey, non-adherent.. Sterile, individually peel packed Size of Dressing: ±100mm x 200mm Each | Not applicable | 3 samples in an original packaging
RT42-01-091 | Dressing Wound, Alginate, Impregnated with Honey, non-adherent.. Sterile, individually peel packed Size of Dressing: 50mm x 50mm Each | Not applicable | 3 samples in an original packaging
RT42-01-092 | Dressing Wound, Alginate, Impregnated with Honey, non-adherent.. Sterile, individually peel packed Size of Dressing: 100mm x 100mm Each | Not applicable | 3 samples in an original packaging
RT42-01-093 | Dressing, Alginate Absorbent, non-adherent, sterile, Sterile, individually peel packed Size of Dressing: 60mm x 60mm Each | Not applicable | 3 samples in an original packaging
RT42-01-094 | Dressing, Alginate Absorbent, non-adherent, sterile, Sterile, individually peel packed Size of Dressing: 100mm x 100mm Each | Not applicable | 3 samples in an original packaging
RT42-01-095 | Dressing, Alginate Absorbent, non-adherent, sterile, Sterile, individually peel packed Size of Dressing: 100mm x 200mm Each | Not applicable | 3 samples in an original packaging
RT42-01-096 | Dressing, Alginate Absorbent, non-adherent, sterile, Sterile, individually peel packed Size of Dressing: 150mm x 150mm Each | Not applicable | 3 samples in an original packaging
RT42-01-097 | Dressing, Alginate Absorbent, non-adherent, sterile, Sterile, individually peel packed Size of Dressing: 150mm x 250mm Each | Not applicable | 3 samples in an original packaging
RT42-01-098 | Dressing Wound, , super absorbent capillary action, woven Absorbent, layered dressing to assist in the removal of exudate and dead tissue, be non-adherent to wound surface, sterile. . Should have description on outer pack Sterile, individually peel packed Size Dressing: 50mm x 50mm Each | 100% | 3 samples in an original packaging
RT42-01-099 | Dressing Wound, super absorbent capillary action, non-woven Absorbent, layered dressing to assist in the removal of exudate and dead tissue, be non-adherent to wound surface, sterile. . Should have description on outer pack Sterile, individually peel packed To comply with the latest issue of SANS 8840 Compliance certificate to be submitted with bid Size Dressing: 50mm x 50mm Each | 100% | 3 samples in an original packaging
RT42-01-100 | Dressing Wound ,super absorbent capillary action, woven Absorbent, layered dressing to assist in the removal of exudate and dead tissue, be non-adherent to wound surface, sterile. . Should have description on outer pack Sterile, individually peel packed Size Dressing: 75mm x 75mm Each | 100% | 3 samples in an original packaging
RT42-01-101 | Dressing Wound,super absorbent capillary action non-woven Absorbent, layered dressing to assist in the removal of exudate and dead tissue, be non-adherent to wound surface, sterile. . Should have description on outer pack Sterile, individually peel packed To comply with the latest issue of SANS 8840 Compliance certificate to be submitted with bid Size Dressing: 75mm x 75mm Each | 100% | 3 samples in an original packaging
RT42-01-102 | Dressing Wound, super absorbent capillary action, woven Absorbent, layered dressing to assist in the removal of exudate and dead tissue, be non-adherent to wound surface, sterile. . Should have description on outer pack Sterile, individually peel packed Size Dressing: 100mm x 100mm Each | 100% | 3 samples in an original packaging
RT42-01-103 | Dressing Wound, super absorbent capillary action, non-woven Absorbent, layered dressing to assist in the removal of exudate and dead tissue, be non-adherent to wound surface, sterile. . Should have description on outer pack Sterile, individually peel packed To comply with the latest issue of SANS 8840 Compliance certificate to be submitted with bid Size Dressing: 100mm x 100mm Each | 100% | 3 samples in an original packaging
RT42-01-104 | Dressing Wound,super absorbent capillary action, woven Absorbent, layered dressing to assist in the removal of exudate and dead tissue, be non-adherent to wound surface, sterile. . Should have description on outer pack Sterile, individually peel packed Size Dressing: 150mm x 150mm Each | 100% | 3 samples in an original packaging
RT42-01-105 | Dressing Wound, super absorbent capillary action,, non-woven Absorbent, layered dressing to assist in the removal of exudate and dead tissue, be non-adherent to wound surface, sterile. . Should have description on outer pack Sterile, individually peel packed To comply with the latest issue of SANS 8840 Compliance certificate to be submitted with bid Size Dressing: 150mm x 150mm Each | 100% | 3 samples in an original packaging
RT42-01-106 | Dressing Wound, super absorbent capillary action,, woven Absorbent, layered dressing to assist in the removal of exudate and dead tissue, be non-adherent to wound surface, sterile. . Should have description on outer pack Sterile, individually peel packed Size Dressing: 150mm x 200mm Each | 100% | 3 samples in an original packaging
RT42-01-107 | Dressing Wound,super absorbent capillary action,, non-woven Absorbent, layered dressing to assist in the removal of exudate and dead tissue, be non-adherent to wound surface, sterile. . Should have description on outer pack Sterile, individually peel packed To comply with the latest issue of SANS 8840 Compliance certificate to be submitted with bid Size Dressing: 150mm x 200mm Each | 100% | 3 samples in an original packaging
RT42-01-108 | Dressing Wound, super absorbent capillary action,, woven Absorbent, layered dressing to assist in the removal of exudate and dead tissue, be non-adherent to wound surface, sterile. Should have description on outer pack Sterile, individually peel packed Size Dressing: 100mm x 1m Each | 100% | 3 samples in an original packaging
RT42-01-109 | Dressing Wound,super absorbent capillary action,y, non-woven Absorbent, layered dressing to assist in the removal of exudate and dead tissue, be non-adherent to wound surface, sterile. Should have description on outer pack Sterile, individually peel packed To comply with the latest issue of SANS 8840 Compliance certificate to be submitted with bid Size Dressing: 100mm x 1m Each | 100% | 3 samples in an original packaging
RT42-01-110 | Dressing Wound, super absorbent capillary action,, woven Absorbent, layered dressing to assist in the removal of exudate and dead tissue, be non-adherent to wound surface, sterile. . Should have description on outer pack Sterile, individually peel packed Size Dressing: 200mm x 200mm Each | 100% | 3 samples in an original packaging
RT42-01-111 | Dressing Wound,super absorbent capillary action, non-woven Absorbent, layered dressing to assist in the removal of exudate and dead tissue, be non-adherent to wound surface, sterile. . Should have description on outer pack Sterile, individually peel packed To comply with the latest issue of SANS 8840 Compliance certificate to be submitted with bid Size Dressing: 200mm x 200mm Each | 100% | 3 samples in an original packaging
RT42-01-112 | Dressing wound, super absorbent capillary action, woven, stretch, absorbent, layered dressing to assist in the removal of exudate and dead tissue, be non-adherent to wound surface, sterile, should have description on outer pack, sterile, individually peel packed. Size: 200mm x 1m Each | 100% | 3 samples in an original packaging
RT42-01-113 | Dressing wound,super absorbent capillary action, non-woven, stretch, absorbent, layered dressing to assist in the removal of exudate and dead tissue, be non-adherent to wound surface, sterile, should have description on outer pack, sterile, individually peel packed. To comply with the latest issue of SANS 8840 Compliance certificate to be submitted with bid Size: 200mm x 1m Each | 100% | 3 samples in an original packaging
RT42-01-114 | Dressing Wound,super absorbent capillary action,woven, Tracheostomy Absorbent, layered dressing to assist in the removal of exudate and dead tissue, be non-adherent to wound surface, sterile. Should have description on outer pack Sterile, individually peel packed Size Dressing :100mm x100mm Each | 100% | 3 samples in an original packaging
RT42-01-115 | Dressing Wound, super absorbent capillary action, non-woven, Tracheostomy Absorbent, layered dressing to assist in the removal of exudate and dead tissue, be non-adherent to wound surface, sterile. Should have description on outer pack Sterile, individually peel packed To comply with the latest issue of SANS 8840 Compliance certificate to be submitted with bid Size Dressing :100mm x100mm Each | 100% | 3 samples in an original packaging
RT42-01-116 | Dressing Wound, super absorbent capillary action, woven , Cavity Absorbent, layered dressing to assist in the removal of exudate and dead tissue, be non-adherent to wound surface, sterile. . Should have description on outer pack Sterile, individually peel packed Size Dressing: 100mm x 100mm Each | 100% | 3 samples in an original packaging
RT42-01-117 | Dressing Wound, super absorbent capillary action, non-woven , Cavity Absorbent, layered dressing to assist in the removal of exudate and dead tissue, be non-adherent to wound surface, sterile. . Should have description on outer pack Sterile, individually peel packed To comply with the latest issue of SANS 8840 Compliance certificate to be submitted with bid Size Dressing: 100mm x 100mm Each | 100% | 3 samples in an original packaging
RT42-01-118 | Dressing wound, imprenaged with antimicrobial/antibacterial absorbent properties, must be have an open-pored sterile, freeze-dried matrix, must contains ionically bound silver, which provides antimicrobial protection against bacteria and infection. The absorbent collagen sheets manage wound moisture which must converts to soft gel sheet that stays in intimate contact with the wound bed, Three-dimensional structure helps support cellular growth Dressing must be composed of collagen and oxidised regenerated cellulose Size: 123m² Sterile, Individually wrapped Box of 10 dressings | Not applicable | 3 samples in an original packaging
RT42-01-119 | Dressing wound, imprenaged with antimicrobial/antibacterial absorbent properties, must be have an open-pored sterile, freeze-dried matrix, must contains ionically bound silver, which provides antimicrobial protection against bacteria and infection. The absorbent collagen sheets manage wound moisture which must converts to soft gel sheet that stays in intimate contact with the wound bed, Three-dimensional structure helps support cellular growth Dressing must be composed of collagen and oxidised regenerated cellulose Size: 28m² Sterile, Individually wrapped Box of 10 dressings | Not applicable | 3 samples in an original packaging
RT42-01-120 | Dressing, wound, construct of a gelling fibre, must have an absorbent wound covers, which combined, manage wound drainage. Dressing must be able to remove of dead skin, damage and infected tissue. Dressing must have an antimicrobial action with to clean. Dressing must be able to easy removal removes easily and reduce risk of skin damage. Dressing must be able to be used on treating of medium to highly fluid exudating wounds Individually wrapped Sterile Size: 10cmx10cm, Box of 10 dressings, | Not applicable | 3 samples in an original packaging
RT42-01-121 | Dressing, wound, construct of a gelling fibre, must have an absorbent wound covers, which combined, manage wound drainage. Dressing must be able to remove of dead skin, damage and infected tissue. Dressing must have an antimicrobial action with to clean. Dressing must be able to easy removal removes easily and reduce risk of skin damage. Dressing must be able to be used on treating of medium to highly fluid exudating wounds Individually wrapped Sterile Size: 15cmx15cm, Box of 10 dressings, | Not applicable | 3 samples in an original packaging
RT42-01-122 | Dressing, wound, construct of a gelling fibre, must have an absorbent wound covers, which combined, manage wound drainage. Dressing must be able to remove of dead skin, damage and infected tissue. Dressing must have an antimicrobial action with to clean. Dressing must be able to easy removal removes easily and reduce risk of skin damage. Dressing must be able to be used on treating of medium to highly fluid exudating wounds Individually wrapped Sterile Size: 20cmx30cm, Box of 10 dressings, | Not applicable | 3 samples in an original packaging
RT42-01-123 | Dressing, wound, construct of a gelling fibre, must have an absorbent wound covers, which combined, manage wound drainage. Dressing must be able to remove of dead skin, damage and infected tissue. Dressing must have an antimicrobial action with to clean. Dressing must be able to easy removal removes easily and reduce risk of skin damage. Dressing must be able to be used on treating of medium to highly fluid exudating wounds Individually wrapped Sterile Size: 2cmx45cm, Box of 5 dressings, | Not applicable | 3 samples in an original packaging
RT42-01-124 | Dressing wound, hydrokinetic, super absorbent, layered mechanically modified cellulose and gel forming polymerdressing to assist in the removal of exudate, bacteria and dead tissue, be non-adherent to wound surface, sterile, individually peel packed, size dressing: 7.5cm x 7.5cm. Each | Not applicable | 3 samples in an original packaging
RT42-01-125 | Dressing wound, hydrokinetic, super absorbent, layered mechanically modified cellulose and gel forming polymerdressing to assist in the removal of exudate, bacteria and dead tissue, be non-adherent to wound surface, sterile, individually peel packed, size dressing:12cm x 5cm. Each | Not applicable | 3 samples in an original packaging
RT42-01-126 | Dressing wound, hydrokinetic, super absorbent, layered mechanically modified cellulose and gel forming polymerdressing to assist in the removal of exudate, bacteria and dead tissue, be non-adherent to wound surface, sterile, individually peel packed, size dressing:10cm x 10cm. Each | Not applicable | 3 samples in an original packaging
RT42-01-127 | Dressing wound, hydrokinetic, super absorbent, layered mechanically modified cellulose and gel forming polymerdressing to assist in the removal of exudate, bacteria and dead tissue, be non-adherent to wound surface, sterile, individually peel packed, size dressing: 20cm x 10cm. Each | Not applicable | 3 samples in an original packaging
RT42-01-128 | Dressing wound, hydrokinetic, super absorbent, layered mechanically modified cellulose and gel forming polymerdressing to assist in the removal of exudate, bacteria and dead tissue, be non-adherent to wound surface, sterile, individually peel packed, size dressing:15cm x 15cm Each | Not applicable | 3 samples in an original packaging
RT42-01-129 | Dressing wound, hydrokinetic, super absorbent, layered mechanically modified cellulose and gel forming polymerdressing to assist in the removal of exudate, bacteria and dead tissue, be non-adherent to wound surface, sterile, individually peel packed, size dressing: 20cm x 20cm. Each | Not applicable | 3 samples in an original packaging
RT42-01-130 | Dressing wound, hydrokinetic, super absorbent, layered mechanically modified cellulose and gel forming polymerdressing to assist in the removal of exudate, bacteria and dead tissue, be non-adherent to wound surface, sterile, individually peel packed, size dressing:30cm x 20cm Each | Not applicable | 3 samples in an original packaging
RT42-01-131 | Dressing wound, hydrokinetic, super absorbent, layered mechanically modified cellulose and gel forming polymerdressing to assist in the removal of exudate, bacteria and dead tissue, be non-adherent to wound surface, sterile, individually peel packed, size dressing: round 8cm diameter. Each | Not applicable | 3 samples in an original packaging
RT42-01-132 | Dressing wound, hydrokinetic, super absorbent, layered mechanically modified cellulose and gel forming polymerdressing to assist in the removal of exudate, bacteria and dead tissue, be non-adherent to wound surface, sterile, individually peel packed, size dressing: round 4cm diameter. Each | Not applicable | 3 samples in an original packaging
RT42-01-133 | Dressing wound, self-adherent, hydrokinetic, super absorbent, layered mechanically modified cellulose and gel forming polymerdressing to assist in the removal of exudate, bacteria and dead tissue, be non-adherent to wound surface, sterile. individually peel packed. size dressing: 10cm x 10cm. Each | Not applicable | 3 samples in an original packaging
RT42-01-134 | Dressing wound, self-adherent, hydrokinetic, super absorbent, layered mechanically modified cellulose and gel forming polymerdressing to assist in the removal of exudate, bacteria and dead tissue, be non-adherent to wound surface, sterile. individually peel packed. size dressing: 15cm x 15cm. Each | Not applicable | 3 samples in an original packaging
RT42-01-135 | Dressing wound, self-adherent, hydrokinetic, super absorbent, layered mechanically modified cellulose and gel forming polymerdressing to assist in the removal of exudate, bacteria and dead tissue, be non-adherent to wound surface, sterile. individually peel packed. size dressing: 15cm x 25cm. Each | Not applicable | 3 samples in an original packaging
RT42-01-136 | Dressing wound, negative pressure wound therapy system. Kit must include: *Dressing consist of a patented four layer dressing or sponge with a size: 10cm x 20cm *Pocket sized pump battery powered pump which must last for a maximum of 7 days Individually packed Single use Sterile Kit | 100% | 3 samples in an original packaging
RT42-01-137 | Dressing wound, negative pressure wound therapy system. Kit must include: *Dressing consist of a patented four layer dressing or sponge with a size: 10cm x 30cm *Pocket sized pump battery powered pump which must last for a maximum of 7 days Individually packed Single use Sterile Kit | 100% | 3 samples in an original packaging
RT42-01-138 | Dressing wound,negative pressure wound therapy system. Kit must include: *Dressing consist of a patented four layer dressing or sponge with a size: 10cm x 40cm *Pocket sized pump battery powered pump which must last for a maximum of 7 days Individually packed Single use Sterile Kit | 100% | 3 samples in an original packaging
RT42-01-139 | Dressing wound, Single use negative pressure wound therapy system. Kit must include: *Dressing consist of a patented four layer dressing or sponge with a size: 15cm x 15cm *Pocket sized pump battery powered pump which must last for a maximum of 7 days Individually packed Single use Sterile Kit | 100% | 3 samples in an original packaging
RT42-01-140 | Dressing wound, Single use negative pressure wound therapy system. Kit must include: *Dressing consist of a patented four layer dressing or sponge with a size: 15cm x 20cm *Pocket sized pump battery powered pump which must last for a maximum of 7 days Individually packed Single use Sterile Kit | 100% | 3 samples in an original packaging
RT42-01-141 | Dressing wound, Single use negative pressure wound therapy system. Kit must include: *Dressing consist of a patented four layer dressing or sponge with a size: 15cm x 30cm *Pocket sized pump battery powered pump which must last for a maximum of 7 days Individually packed Single use Sterile Kit | 100% | 3 samples in an original packaging
RT42-01-142 | Dressing wound, Single use negative pressure wound therapy system. Kit must include: *Dressing consist of a patented four layer dressing or sponge with a size: 20cm x 20cm *Pocket sized pump battery powered pump which must last for a maximum of 7 days Individually packed Single use Sterile Kit | 100% | 3 samples in an original packaging
RT42-01-143 | Dressing wound, Single use negative pressure wound therapy system. Kit must include: *Dressing consist of a patented four layer dressing or sponge with a size: 25cm x 25cm *Pocket sized pump battery powered pump which must last for a maximum of 7 days Individually packed Single use Sterile Kit | 100% | 3 samples in an original packaging
RT42-01-144 | Negative pressure wound therapy system Equipment specification for placement • To treat chronic and acute wounds with high levels of exudate • Portable • Lightweight • Easy to use • Free air displacement • Battery/power operation • Multiple safety alarms system • With a patient lock-out feature • Negative pressure cycle: Continuous or intermittent • Analogue display • Low sound level The machines must be on loan. All sizes for the consumables to be purchased must be compatible with the machines. | 100% | 1 sample in an original packaging
RT42-01-145 | Negative pressure wound therapy system Small Dressing kit Sterile, Peel packed Each This item must be compatible with item RT42-01-144 and be the same brand offered on RT42-01-144 by the same bidder. | 100% | 3 samples in an original packaging
RT42-01-146 | Negative pressure wound therapy system Medium Dressing kit Sterile, Peel packed Each This item must be compatible with item RT42-01-144 and be the same brand offered on RT42-01-144 by the same bidder. | 100% | 3 samples in an original packaging
RT42-01-147 | Negative pressure wound therapy system Large Dressing kit Sterile, Peel packed Each This item must be compatible with item RT42-01-144 and be the same brand offered on RT42-01-144 by the same bidder. | 100% | 3 samples in an original packaging
RT42-01-148 | Negative pressure wound therapy system X-Large Dressing kit Sterile, Peel packed Each This item must be compatible with item RT42-01-144 and be the same brand offered on RT42-01-144 by the same bidder. | 100% | 3 samples in an original packaging
RT42-01-149 | Negative pressure wound therapy system XX-Large Dressing kit Sterile, Peel packed Each This item must be compatible with item RT42-01-144 and be the same brand offered on RT42-01-144 by the same bidder. | 100% | 3 samples in an original packaging
RT42-01-150 | Negative pressure wound therapy system - canister Size: 500ml Each This item must be compatible with item RT42-01-144 and be the same brand offered on RT42-01-144 by the same bidder. | Not applicable | 3 samples in an original packaging
RT42-01-151 | Negative pressure wound therapy system - canister Size: 1000ml Each This item must be compatible with item RT42-01-144 and be the same brand offered on RT42-01-144 by the same bidder. | Not applicable | 3 samples in an original packaging
RT42-01-152 | Dressing, Wound Polymer dressing, Highly absorbent Sterile non-adherent, moisture-vapour permeable Waterproof polyurethane. Dressing contains silver complex Must include instructions for use on outer pack Sterile, individually peel packed Size Dressing: 50mm x 70mm Each | Not applicable | 3 samples in an original packaging
RT42-01-153 | Dressing, Wound Polymer dressing, Highly absorbent Sterile non-adherent, moisture-vapour permeable Waterproof polyurethane. Dressing contains silver complex Must include instructions for use on outer pack Sterile, individually peel packed Size Dressing: 100mm x 100mm Each | Not applicable | 3 samples in an original packaging
RT42-01-154 | Dressing,Wound, Polymer dressing, Highly absorbent Sterile non-adherent, moisture-vapour permeable Waterproof polyurethane. Dressing contains silver complex Must include instructions for use on outer pack Sterile, individually peel packed Size Dressing: 100mm x 120mm Each | Not applicable | 3 samples in an original packaging
RT42-01-155 | Dressing, Wound Polymer dressing, Highly absorbent Sterile non-adherent, moisture-vapour permeable Waterproof polyurethane. Dressing contains silver complex Must include instructions for use on outer pack Sterile, individually peel packed Size Dressing: 100mm x 200mm Each | Not applicable | 3 samples in an original packaging
RT42-01-156 | Dressing, Wound Polymer dressing, Highly absorbent Sterile non-adherent, moisture-vapour permeable Waterproof polyurethane. Dressing contains silver complex Must include instructions for use on outer pack Sterile, individually peel packed Size Dressing: 150mm x 150mm Each | Not applicable | 3 samples in an original packaging
RT42-01-157 | Dressing, Wound, Polymer dressing, Highly absorbent Sterile non-adherent, moisture-vapour permeable Waterproof polyurethane. Dressing contains silver complex Must include instructions for use on outer pack Sterile, individually peel packed Size Dressing: 150mm x 200mm Each | Not applicable | 3 samples in an original packaging
RT42-01-158 | Dressing, Wound Polymer dressing, Highly absorbent Sterile non-adherent, moisture-vapour permeable Waterproof polyurethane. Dressing contains silver complex Must include instructions for use on outer pack Sterile, individually peel packed Size Dressing: 200mm x 200mm Each | Not applicable | 3 samples in an original packaging
RT42-01-159 | Dressing, Wound Polymer dressing, Highly absorbent Sterile non-adherent, moisture-vapour permeable Waterproof polyurethane. with ibuprofen Dressing release Ibuprofen into wound, released at 0.5g/cm². Sterile, individually peel packed Size Dressing: 100mm x 100mm Each | Not applicable | 3 samples in an original packaging
RT42-01-160 | Dressing, Wound Polymer dressing, Highly absorbent Sterile non-adherent, moisture-vapour permeable Waterproof polyurethane. with ibuprofen Dressing release Ibuprofen into wound, released at 0.5g/cm². Sterile, individually peel packed Size Dressing: 100mm x 200mm Each | Not applicable | 3 samples in an original packaging
RT42-01-161 | Dressing, Wound Polymer dressing, Highly absorbent Sterile non-adherent, moisture-vapour permeable Waterproof polyurethane with ibuprofen Dressing release Ibuprofen into wound, released at 0.5g/cm². Sterile, individually peel packed Size Dressing: 150mm x 150mm Each | Not applicable | 3 samples in an original packaging
RT42-01-162 | Dressing, Wound Polymer dressing, Highly absorbent Sterile non-adherent, moisture-vapour permeable Waterproof polyurethane with ibuprofen Dressing release Ibuprofen into wound, released at 0.5g/cm². Sterile, individually peel packed Size Dressing: 200mm x 200mm Each | Not applicable | 3 samples in an original packaging
RT42-01-163 | Dressing Wound, Foam, absorbent, non-adherent Highly absorbent centre layer to accommodate large volume exudate and odour, polypropylene/polyethylene wound contact layer of dressing to be water proof backing layer. Non-adherent wound contact layer. Sterile, individually peel packed Size Dressing: 75mm x75mm Each | 100% | 3 samples in an original packaging
RT42-01-164 | Dressing Wound, Foam, absorbent, non-adherent Highly absorbent centre layer to accommodate large volume exudate and odour, polypropylene/polyethylene wound contact layer of dressing to be water proof backing layer. Non-adherent wound contact layer. Sterile, individually peel packed Size Dressing: 100mm x 100mm Each | 100% | 3 samples in an original packaging
RT42-01-165 | Dressing Wound, Foam, absorbent, non-adherent Highly absorbent centre layer to accommodate large volume exudate and odour, polypropylene/polyethylene wound contact layer of dressing to be water proof backing layer Non-adherent wound contact layer. Sterile, individually peel packed Size Dressing: 100mm x 200mm Each | 100% | 3 samples in an original packaging
RT42-01-166 | Dressing Wound, Foam, absorbent, non-adherent Highly absorbent centre layer to accommodate large volume exudate and odour, polypropylene/polyethylene wound contact layer of dressing to be water proof backing layer. Non-adherent wound contact layer. Sterile, individually peel packed Size Dressing: 150mm x 150mm Each | 100% | 3 samples in an original packaging
RT42-01-167 | Dressing Wound, Foam, absorbent, non-adherent Highly absorbent centre layer to accommodate large volume exudate and odour, polypropylene/polyethylene wound contact layer of dressing to be water proof backing layer. Non-adherent wound contact layer. Sterile, individually peel packed Size Dressing: 200mm x 200mm Each | 100% | 3 samples in an original packaging
RT42-01-168 | Dressing Wound, Foam, absorbent, non-adherent, Tracheostomy Highly absorbent centre layer to accommodate large volume exudate and odour, polypropylene/polyethylene wound contact layer of dressing to be water proof backing layer. Non-adherent wound contact layer. Sterile, individually peel packed Size Dressing: 100mm x 100mm Each | 100% | 3 samples in an original packaging
RT42-01-169 | Dressing Wound, Foam, absorbent, non-adherent, Heel Highly absorbent centre layer to accommodate large volume exudate and odour, polypropylene/polyethylene wound contact layer of dressing to be water proof backing layer. Non-adherent wound contact layer. Sterile, individually peel packed Each | 100% | 3 samples in an original packaging
RT42-01-170 | Dressing Wound, Foam, absorbent, self-adherent Highly absorbent centre layer to accommodate large volume exudate and odour, polypropylene/polyethylene wound contact layer impregnated with silver, non-adherent and sterile. Outerlayer to be water proof. Non-adherent wound contact layer. Sterile, individually peel packed Size Dressing: 75mm x 75mm Each | 100% | 3 samples in an original packaging
RT42-01-171 | Dressing Wound, Foam, absorbent, self-adherent Highly absorbent centre layer to accommodate large volume exudate and odour, polypropylene/polyethylene wound contact layer impregnated with silver, non-adherent and sterile. Outerlayer to be water proof. Non-adherent wound contact layer. Sterile, individually peel packed Size Dressing: 100mm x 100mm Each | 100% | 3 samples in an original packaging
RT42-01-172 | Dressing Wound, Foam, absorbent, self-adherent Highly absorbent centre layer to accommodate large volume exudate and odour, polypropylene/polyethylene wound contact layer impregnated with silver, non-adherent and sterile. Outerlayer to be water proof. Non-adherent wound contact layer. Sterile, individually peel packed Size Dressing: 150mm x 150mm Each | 100% | 3 samples in an original packaging
RT42-01-173 | Dressing Wound, Foam, absorbent, self-adherent Highly absorbent centre layer to accommodate large volume exudate and odour, polypropylene/polyethylene wound contact layer impregnated with silver, non-adherent and sterile. Outerlayer to be water proof. Non-adherent wound contact layer. Sterile, individually peel packed Size Dressing: 225mm x 225mm Each | 100% | 3 samples in an original packaging
RT42-01-174 | Dressing Wound, Foam, absorbent, self-adherent , Heel Highly absorbent centre layer to accommodate large volume exudate and odour, polypropylene/polyethylene wound contact layer impregnated with silver, non-adherent and sterile. Outerlayer to be water proof. Non-adherent wound contact layer. Sterile, individually peel packed Size Dressing: 100mm x130mm Each | Not applicable | 3 samples in an original packaging
RT42-01-175 | Dressing Wound, Foam, absorbent, self-adherent , Sacral Highly absorbent centre layer to accommodate large volume exudate and odour, polypropylene/polyethylene wound contact layer impregnated with silver, non-adherent and sterile. Outerlayer to be water proof. Non-adherent wound contact layer. Sterile, individually peel packed Size Dressing: 220mm x220mm Each | Not applicable | 3 samples in an original packaging
RT42-01-176 | Dressing, hydrocolloid Sodium Carboxymethyl cellulose hydrocolloid layer bonded to a polyurethane film, transparent for visual inspection, flexible dressing. Sterile, individually peel packed Size of Dressing: 40mm x 60mm Each | Not applicable | 3 samples in an original packaging
RT42-01-177 | Dressing, hydrocolloid Sodium Carboxymethyl cellulose hydrocolloid layer bonded to a polyurethane film, transparent for visual inspection, flexible dressing. Sterile, individually peel packed Size of Dressing: 100mm x 100mm Each | 100% | 3 samples in an original packaging
RT42-01-178 | Dressing, hydrocolloid Sodium Carboxymethyl cellulose hydrocolloid layer bonded to a polyurethane film, transparent for visual inspection, flexible dressing. Sterile, individually peel packed Size of Dressing: 150mm x 150mm Each | Not applicable | 3 samples in an original packaging
RT42-01-179 | Dressing, hydrocolloid Sodium Carboxymethyl cellulose hydrocolloid layer bonded to a polyurethane film, transparent for visual inspection, flexible dressing. Sterile, individually peel packed Size of Dressing: 200mm x 200mm Each | 100% | 3 samples in an original packaging
RT42-01-180 | Dressing, hydrocolloid, Sacral dressing Sodium Carboxymethyl cellulose hydrocolloid layer bonded to a polyurethane film, transparent for visual inspection, flexible dressing. Sterile, individually peel packed Size of Dressing: 180mm x 150mm Each | Not applicable | 3 samples in an original packaging
RT42-01-181 | Dressing, hydrocolloid, Sacral dressing Sodium Carboxymethyl cellulose hydrocolloid layer bonded to a polyurethane film, transparent for visual inspection, flexible dressing. Sterile, individually peel packed Size of Dressing: 200mm x 230mm Each | Not applicable | 3 samples in an original packaging
RT42-01-182 | Dressing, hydrocolloid, with border Sodium Carboxymethyl cellulose, pectin, gelatin Adhesive to be: hydrocolloid layer bonded to a polyurethane foam layer for protection, flexible, hypo-allergenic. Sterile, individually peel packed Size of Dressing: 100mm x 100mm Each | Not applicable | 3 samples in an original packaging
RT42-01-183 | Dressing, hydrocolloid, with border Sodium Carboxymethyl cellulose, pectin, gelatin Adhesive to be: hydrocolloid layer bonded to a polyurethane foam layer for protection, flexible, hypo-allergenic. Sterile, individually peel packed Size of Dressing: 150mm x 150mm Each | Not applicable | 3 samples in an original packaging
RT42-01-184 | Dressing, hydrocolloid, with border Sodium Carboxymethyl cellulose, pectin, gelatin Adhesive to be: hydrocolloid layer bonded to a polyurethane foam layer for protection, flexible, hypo-allergenic. Sterile, individually peel packed Size of Dressing: 200mm x 200mm Each | Not applicable | 3 samples in an original packaging
RT42-01-185 | Dressing, hydrocolloid, without borders Transparent, Sodium Carboxymethyl cellulose, pectin, gelatin, thin, hydrocolloid layer bonded to a polyurethane film, transparent for visual inspection, flexible dressing. Sterile, individually peel packed Size of Dressing: 50mm x 50mm Each | Not applicable | 3 samples in an original packaging
RT42-01-186 | Dressing, hydrocolloid, without borders Transparent, Sodium Carboxymethyl cellulose, pectin, gelatin, thin, hydrocolloid layer bonded to a polyurethane film, transparent for visual inspection, flexible dressing. Sterile, individually peel packed Size of Dressing: 100mm x 100mm Each | 100% | 3 samples in an original packaging
RT42-01-187 | Dressing, hydrocolloid, without borders Transparent, Sodium Carboxymethyl cellulose, pectin, gelatin, thin, hydrocolloid layer bonded to a polyurethane film, transparent for visual inspection, flexible dressing. Sterile, individually peel packed Size of Dressing: 150mm x 150mm Each | Not applicable | 3 samples in an original packaging
RT42-01-188 | Dressing, hydrocolloid, without borders Transparent, Sodium Carboxymethyl cellulose, pectin, gelatin, thin, hydrocolloid layer bonded to a polyurethane film, transparent for visual inspection, flexible dressing. Sterile, individually peel packed Size of Dressing: 200mm x 200mm Each | 100% | 3 samples in an original packaging
RT42-01-189 | Dressing, Silver Meshed Non-woven, porous, contains Silver Sulphate, sterile. Sterile, individually peel packed Size of Dressing: 100mm x 120mm Each | Not applicable | 3 samples in an original packaging
RT42-01-190 | Dressing, Silver Meshed Non-woven, porous, contains Silver Sulphate , sterile. Sterile, individually peel packed Size of Dressing: 100mm x 200mm Each | Not applicable | 3 samples in an original packaging
RT42-01-191 | Dressing, Silver Meshed Non-woven, porous, contains Silver Sulphate, sterile. Sterile, individually peel packed Size of Dressing: 150mm x 200mm Each | Not applicable | 3 samples in an original packaging
RT42-01-192 | Dressing, Silver Meshed Non-woven, porous, contains Silver Sulphate, sterile. Sterile, individually peel packed Size of Dressing: 200mm x 200mm Each | Not applicable | 3 samples in an original packaging
RT42-01-193 | Dressing, Silver Meshed Non-woven, porous, contains Silver Sulphate, sterile. Sterile, individually peel packed Size of Dressing: 400mm x 400mm Each | Not applicable | 3 samples in an original packaging
RT42-01-194 | Dressing, keliod and hypertropic scars, Size: 100mm x180mm transparent medical grade silicone rubber sheet, soft adhesive ,washable and reusable. Reinforced with silicone membrane backing, conformable. Sterile, indivually peel- packed. Each | Not applicable | 3 samples in an original packaging
RT42-01-195 | Dressing, keliod and hypertropic scars, Size: 120mm x150mm transparent medical grade silicone rubber sheet, soft adhesive ,washable and reusable. Reinforced with silicone membrane backing, conformable. Sterile, indivually peel- packed. Each | Not applicable | 3 samples in an original packaging
RT42-01-196 | Dressing, keliod and hypertropic scars, Size: 60mm X 120mm transparent medical grade silicone rubber sheet, soft adhesive ,washable and reusable. Reinforced with silicone membrane backing, conformable. Sterile, indivually peel- packed. Each | Not applicable | 3 samples in an original packaging
RT42-01-197 | Dressing, incorporating ionic silver and calcium alginate with polyurethrane foam, adhesive Size: 60mm x 60mm, sterile, individually peel pack, Each | Not applicable | 3 samples in an original packaging
RT42-01-198 | Dressing, incorporating ionic silver and calcium alginate with polyurethrane foam,adhesive Size: 100mm x 100m , sterile, individually peel pack, Each | Not applicable | 3 samples in an original packaging
RT42-01-199 | Dressing, incorporating ionic silver and calcium alginate with polyurethrane foam, adhesive Size: 150mm x 150mm , sterile, individually peel pack, Each | Not applicable | 3 samples in an original packaging
RT42-01-200 | Dressing, incorporating ionic silver and calcium alginate with polyurethrane foam,adhesive Size: 200mm x 200mm , sterile, individually peel pack Each | Not applicable | 3 samples in an original packaging
RT42-01-201 | Self adhering foam designed to distribute body weight and prevent skin break down. Size:12mm thick, (216 mm x 279mm) Each | Not applicable | 3 samples in an original packaging
RT42-01-202 | Dressing, antimicrobial Size: 85mm x 115mm . Semipermeable polyurethrane see through pad with outer adhesive border, contains Chlorhexidine Gluconate Sterile Individually Peel Packed, Each | Not applicable | 3 samples in an original packaging
RT42-01-203 | Dressing, antimicrobial, Size: 100mm x 120mm. Semipermeable polyurethrane see through pad with outer adhesive border, contains Chlorhexidine Gluconate .Sterile Individually Peel Packed, Each | Not applicable | 3 samples in an original packaging
RT42-01-204 | Dressing, antimicrobial , Size: 100mm x 155mm . Semipermeable polyurethrane see through pad with outer adhesive border, contains Chlorhexidine Gluconate Sterile Individually Peel Packed.. Each | Not applicable | 3 samples in an original packaging
RT42-01-205 | Dressing Paediatric , IV Dressing, Transparent polyurethane film, water proof, moisture vapour permeable, unique frame delivery system, softcloth reinforcement for added secural, latex free, hypoallergenic, additional medipore cloth for added secural, pre-cut deep notch for catheter anchorage, transparent window for easy monitoring, single use dressing for 3-5 day wear time. . Size 5cm x 5.7cm. Each | 100% | 3 samples in an original packaging
RT42-01-206 | Dressing , IV for peripheral Line, Transparent polyurethane film, water proof, moisture vapour permeable, unique frame delivery system, softcloth reinforcement for added secural, latex free, hypoallergenic, additional medipore cloth for added secural, pre-cut deep notch for catheter anchorage, transparent window for easy monitoring, single use dressing for 3-5 day wear time. . Size 70mm x 85mm. Each | 100% | 3 samples in an original packaging
RT42-01-207 | Dressing , IV for CVC Line, Transparent polyurethane film, water proof, moisture vapour permeable, unique frame delivery system, softcloth reinforcement for added secural, latex free, hypoallergenic, additional medipore cloth for added secural, pre-cut deep notch for catheter anchorage, transparent window for easy monitoring, single use dressing for 3-5 day wear time. Size 85mm x 105mm. Each | 100% | 3 samples in an original packaging
RT42-01-208 | Dressing , Universal IV . Transparent polyurethane film, water proof, moisture vapour permeable, unique frame delivery system, softcloth reinforcement for added secural, latex free, hypoallergenic, additional medipore cloth for added secural, pre-cut deep notch for catheter anchorage, transparent window for easy monitoring, single use dressing for 3-5 day wear time. Size 8.9cm x 11.5cm. Each | 100% | 3 samples in an original packaging
RT42-01-209 | Dressing, Acrylic:Size: 76mm x 95mm, allows for moist wound healing promoting autolytic debridement , absorbent, clear, unique construction - dressing made up of two layers of transparent film sandwitched between a layer of acrylic polymer, self adherent, latex free, sterile,, moisture vapour permeable, hypoallergenic, individually packed, 5 dressings per box. | Not applicable | 3 samples in an original packaging
RT42-01-210 | Dressing, Acrylic:Size: 150mm x 150mm, allows for moist wound healing promoting autolytic debridement , absorbent, clear, unique construction - dressing made up of two layers of transparent film sandwitched between a layer of acrylic polymer, self adherent, latex free, sterile,, moisture vapour permeable, hypoallergenic, individually packed, 5 dressings per box. | Not applicable | 3 samples in an original packaging
RT42-01-211 | Dressing, Acrylic:Size: 170mm x 190mm, allows for moist wound healing promoting autolytic debridement , absorbent, clear, unique construction - dressing made up of two layers of transparent film sandwitched between a layer of acrylic polymer, self adherent, latex free, sterile,, moisture vapour permeable, hypoallergenic, individually packed, 5 dressings per box. | Not applicable | 3 samples in an original packaging
RT42-01-212 | Dressing, Acrylic:Size: 200mm x 200mm, allows for moist wound healing promoting autolytic debridement , absorbent, clear, unique construction - dressing made up of two layers of transparent film sandwitched between a layer of acrylic polymer, self adherent, latex free, sterile,, moisture vapour permeable, hypoallergenic, individually packed, 5 dressings per box. | Not applicable | 3 samples in an original packaging
RT42-01-213 | Dressing, Acrylic:Size: 110mm x 130mm, allows for moist wound healing promoting autolytic debridement , absorbent, clear, unique construction - dressing made up of two layers of transparent film sandwitched between a layer of acrylic polymer, self adherent, latex free, sterile, transparent film to be barrier to contamination. moisture vapour permeable, hypoallergenic, individually packed, 5 dressings per box. | Not applicable | 3 samples in an original packaging
RT42-01-214 | Dressing, occlusive, adhesive, Size: 110mm x 120mm, transparent polyurethrane-based synthetic, hydrophilic layer adheres without adhesive, sterile, Indivicually peel packed. Each | Not applicable | 3 samples in an original packaging
RT42-01-215 | Dressing, occlusive, adhesive, Size: 110mm x 330mm, transparent polyurethrane-based synthetic, hydrophilic layer adheres w/out adhesive, sterile, Indivicually peel packed. Each | Not applicable | 3 samples in an original packaging
RT42-01-216 | Dressing, occlusive, adhesive, Size: 160mm x 2100mm, transparent polyurethrane-based synthetic, hydrophilic layer adheres without adhesive, sterile, Indivicually peel packed Each | Not applicable | 3 samples in an original packaging
RT42-01-217 | Dressing Pack: Surgical: Tray must be packed as follows 2x hand towels 1x red plastic bag 1x Pair of latex free gloves, size Medium, non powdered 2x Disposable soft drape 90cmx90cm Cleansing Swabs 30mm - 1 bundle of 5 balls 10 X Plain gauze 100mm x 100mm 8ply 2x Disposable forceps Sterile, peel packed , Comply with the latest SANS 446 Each | 100% | 3 samples in an original packaging
RT42-01-218 | Dressing Pack: Maternity Tray must be packed as follows: 2x hand towels 1x red plastic bag 1x Pair of latex free gloves, size Medium, non powdered Cleansing Swabs 65mm - 1 bundle of 5 balls 10 X Plain gauze 100mm x 100mm 8ply 2 x Maternity sanitary towels loops Sterile, peel packed , Comply with the latest SANS 446 Each | 100% | 3 samples in an original packaging
RT42-01-219 | Dressing Surgical set: tray ,swabbing Tray must be packed as follows: 2x paper towels, 1x red plastic (280mm x400mm), 2x powder free latex gloves (medium), 1x protection sheet, 10x cotton wool balls 2g, 1x sanitary towel looped, 2x gallipots (75ml). Sterile, peel packed , Comply with the latest SANS 228 Compliance certificate to be submitted with bid Each | 100% | 3 samples in an original packaging
RT42-01-220 | Dressing tray sterile. Tray must be packed as follows, 2x paper towels, 1x red plastic (280mm x 400mm), 2x powder free latex gloves (medium), 1x protection sheet, 5 x gauze swabs (10mm x 100mm x 8 ply), 2 x gallipots (75ml), 1 x forcepts (plastic).Sterile, peel packed , Comply with the latest SANS 446 Each | 100% | 3 samples in an original packaging
RT42-01-221 | Dressing finger and toe, small, polyamide, Polyethersulfoae (PES); elastane/lycra, Each. | Not applicable | 3 samples in an original packaging
RT42-01-222 | Dressing finger and toe, medium, polyamide, Polyethersulfoae (PES); elastane/lycra, Each. | Not applicable | 3 samples in an original packaging
RT42-01-223 | Dressing finger and toe, large, polyamide, Polyethersulfoae (PES); elastane/lycra, Each. | Not applicable | 3 samples in an original packaging
RT42-01-224 | Dressing finger and toe, x-large, polyamide, Polyethersulfoae (PES); elastane/lycra, Each. | Not applicable | 3 samples in an original packaging
RT42-01-225 | Dressing finger and toe, Size: 750mm x 1500mm, polyamide,Polyethersulfoae ( PES); elastane/lycra, Each, roll. | Not applicable | 3 samples in an original packaging
RT42-01-226 | Dressing pad, absorbent, acetate fabric with antimicrobial hydrophobic coating (e.g. Dialkyl Carbamoyl Chloride) that physically binds and removes bacteria, size 70mm x 90mm, packaging individually wrapped, sterile. Each | Not applicable | 3 samples in an original packaging
RT42-01-227 | Dressing pad, absorbent, acetate fabric with antimicrobial hydrophobic coating that physically binds and removes bacteria, size 100mm x 100mm, packaging individually wrapped, sterile. Each | Not applicable | 3 samples in an original packaging
RT42-01-228 | Dressing pad, absorbent, acetate fabric with antimicrobial hydrophobic coating that physically binds and removes bacteria, size 100mm x 200mm, packaging individually wrapped, sterile. Each | Not applicable | 3 samples in an original packaging
RT42-01-229 | Dressing, swab, acetate fabric with antimicrobial hydrophobic coating that physically binds and removes bacteria, size 40mm x 60mm x 8 ply, packaging individually wrapped, sterile. Each | Not applicable | 3 samples in an original packaging
RT42-01-230 | Dressing, swab, acetate fabric with antimicrobial hydrophobic coating that physically binds and removes bacteria, size 70mm x90mm x 8 ply, packaging individually wrapped, sterile.. Each | Not applicable | 3 samples in an original packaging
RT42-01-231 | Dressing, absorbent hydropolymer core, acetate fabric with antimicrobial hydrophobic coating that physically binds and removes bacteria, size 70mm x 85mm, packaging individually wrapped, sterile. Each | Not applicable | 3 samples in an original packaging
RT42-01-232 | Dressing, absorbent hydropolymer core, acetate fabric with antimicrobial hydrophobic coating that physically binds and removes bacteria, size 140mm x 140mm, packaging individually wrapped, sterile. Each | Not applicable | 3 samples in an original packaging
RT42-01-233 | Dressing, absorbent hydropolymer core, acetate fabric with antimicrobial hydrophobic coating that physically binds and removes bacteria, size 190mm x 190mm, packaging individually wrapped, sterile. Each | Not applicable | 3 samples in an original packaging
RT42-01-234 | Dressing, absorbent hydropolymer core, acetate fabric with antimicrobial hydrophobic coating that physically binds and removes bacteria, size 240mm x 240mm, packaging individually wrapped, sterile. Each | Not applicable | 3 samples in an original packaging
RT42-01-235 | Dressing, absorbent hydropolymer core, acetate fabric with antimicrobial hydrophobic coating that physically binds and removes bacteria, size 140mm x 240mm, packaging individually wrapped, sterile. Each | Not applicable | 3 samples in an original packaging
RT42-01-236 | Dressing, transparent, wide area fixation. Dressing must be breathable, flexible and stretchable, chemically inert to rubber and plastics and highly adhesive. Must also be permeable to air and water vapour, prevents constriction, congestion and distubance to blood circulation and radiotransparent. Size: 50mm x 10m Each | Not applicable | 3 samples in an original packaging
RT42-01-237 | Dressing, transparent, wide are fixation. Hypo-allergenic and coated with a polyacrylate adhesive. Backing material is made of transversely stretchable non-woven polyester. Dressing must be breathable, flexible and stretchable, chemically inert to rubber and plastics andhighly adhesive. Must also be permeable to air and water vapor, prevents constriction, congestion and distubance to blood circulation and radiotransparent. Size: 100mm x 10m Each | Not applicable | 3 samples in an original packaging
RT42-01-238 | Dressing, non-adherent, contains 3% tribromophanate in special petrolatum blem ona fine mesh gauze which is non-adherent to wound. As an initial layer for wounds with light exudate where prtection from contamination and or deodorization is desired. Sterile. Must be individual peel packed. Size 10mm x 3m. Each | Not applicable | 3 samples in an original packaging
RT42-01-239 | Dressing post operative transparent waterproof border with absorbent pad with silver sulphate with safe tac technology, flexible and stretchable pad with y cuts for surgical incisions, disposable and sterile. Individually peel packed. Size 10cm x15cm Each | Not applicable | 3 samples in an original packaging
RT42-01-240 | Dressing post operative transparent waterproof border with absorbent pad with silver sulphate with safe tac technology, flexible and stretchable pad with y cuts for surgical incisions, disposable and sterile. Individually peel packed. Size 10cm x25cm Each | Not applicable | 3 samples in an original packaging
RT42-01-241 | Dressing post operative transparent waterproof border with absorbent pad with silver sulphate with safe tac technology, flexible and stretchable pad with y cuts for surgical incisions, disposable and sterile. Individually peel packed. Size 10cm x30cm Each | Not applicable | 3 samples in an original packaging
RT42-01-242 | Dressing post operative transparent waterproof border with absorbent pad with silver sulphate with safe tac technology, flexible and stretchable pad with y cuts for surgical incisions, disposable and sterile. Individually peel packed. Size 10cm x35cm Each | Not applicable | 3 samples in an original packaging
RT42-01-243 | Dressing, Wound, Ceramics. Non-adhesive pad .Non-woven sachet filled with micro-porous ceramic. Sterile, individually peel packed Size: 5cmx5cm Box of 10 dressings. | 100% | 3 samples in an original packaging
RT42-01-244 | Dressing, Wound, Ceramics. Non-adhesive pad Non-woven sachet filled with micro-porous ceramic. Sterile, individually peel packed Size: 5cmx10cm. Box of 10 dressings. | 100% | 3 samples in an original packaging
RT42-01-245 | Dressing, Wound, Ceramics. Non-adhesive pad Non-woven sachet filled with micro-porous ceramic. Sterile, individually peel packed Size: 5cmx15cm. Box of 10 dressings. | 100% | 3 samples in an original packaging
RT42-01-246 | Dressing, Wound, Ceramics. Non-adhesive pad Non-woven sachet filled with micro-porous ceramic. Sterile, individually peel packed Size: 5cmx20cm. Box of 10 dressings. | 100% | 3 samples in an original packaging
RT42-01-247 | Dressing, Wound, Ceramics. Non-adhesive pad .Non-woven sachet filled with micro-porous ceramic. Sterile, individually peel packed Size: 5cmx30cm. Box of 10 dressings. | 100% | 3 samples in an original packaging
RT42-01-248 | Dressing, Wound, Ceramics. Non-adhesive pad . Non-woven sachet filled with micro-porous ceramic. Sterile, individually peel packed Size: 10cmx10cm. Box of 10 dressings. | 100% | 3 samples in an original packaging
RT42-01-249 | Dressing, Wound, Ceramics. Non-adhesive pad . Non-woven sachet filled with micro-porous ceramic. Sterile, individually peel packed Size: 10cmx15cm. Box of 10 dressings. | 100% | 3 samples in an original packaging
RT42-01-250 | Dressing, Wound, Ceramics. Non-adhesive pad. Non-woven sachet filled with micro-porous ceramic. Sterile, individually peel packed Size: 10cmx75cm. Box of 2 dressings. | 100% | 3 samples in an original packaging
RT42-01-251 | Dressing, Wound, Ceramics. Non-adhesive pad . Non-woven sachet filled with micro-porous ceramic. Sterile, individually peel packed Size: 21cmx30cm. Box of 2 dressings. | 100% | 3 samples in an original packaging
RT42-01-252 | Dressing, Wound, Ceramics. Pad mounted on cloth adhesive. Non-woven sachet filled with micro-porous ceramic. Sterile, individually peel packed Dressing Size: 10cmx10cm Pad Size: 5cmx5cm. Box of 10 dressings. | 100% | 3 samples in an original packaging
RT42-01-253 | Dressing, Wound, Ceramics. Pad mounted on cloth adhesive. Non-woven sachet filled with micro-porous ceramic. Sterile, individually peel packed Dressing Size: 10cmx16cm Pad Size: 5cmx10cm. Box of 10 dressings. | 100% | 3 samples in an original packaging
RT42-01-254 | Dressing, Wound, Ceramics. Pad mounted on cloth adhesive. Non-woven sachet filled with micro-porous ceramic. Sterile, individually peel packed Dressing Size: 10cmx21cm Pad Size: 5cmx15cm. Box of 10 dressings. | 100% | 3 samples in an original packaging
RT42-01-255 | Dressing, Wound, Ceramics. Pad mounted on cloth adhesive . Non-woven sachet filled with micro-porous ceramic. Sterile, individually peel packed Dressing Size: 10cmx26cm Pad Size: 5cmx20cm. Box of 10 dressings. | 100% | 3 samples in an original packaging
RT42-01-256 | Dressing, Wound, Ceramics. Pad mounted on water resistant adhesive . Non-woven sachet filled with micro-porous ceramic. Sterile, individually peel packed Dressing Size: 10cmx10cm Pad Size: 5cmx5cm. Box of 10 dressings. | 100% | 3 samples in an original packaging
RT42-01-257 | Dressing, Wound, Ceramics. Pad mounted on water resistant adhesive . Non-woven sachet filled with micro-porous ceramic. Sterile, individually peel packed Dressing Size: 10cmx16cm Pad Size: 5cmx10cm. Box of 10 dressings. | 100% | 3 samples in an original packaging
RT42-01-258 | Dressing, Wound, Ceramics. Pad mounted on water resistant adhesive . Non-woven sachet filled with micro-porous ceramic. Sterile, individually peel packed Dressing Size: 10cmx21cm Pad Size: 5cmx15cm. Box of 10 dressings. | 100% | 3 samples in an original packaging
RT42-01-259 | Dressing, Wound, Ceramics. Pad mounted on water resistant adhesive . Non-woven sachet filled with micro-porous ceramic. Sterile, individually peel packed Dressing Size: 10cmx26cm Pad Size: 5cmx20cm. Box of 10 dressings. | 100% | 3 samples in an original packaging
RT42-01-260 | Dressing, Wound, Ceramics. Pad mounted on water resistant adhesive . Non-woven sachet filled with micro-porous ceramic. Sterile, individually peel packed Dressing Size: 10cmx35cm Pad Size: 5cmx30cm. Box of 10 dressings. | 100% | 3 samples in an original packaging
RT42-01-261 | Dressing, Finger, Ceramics. Non-woven sachet filled with micro-porous ceramic. Sachets to conform to the shape of finger and toe digits. Cloth adhisive to assist application. Sterile, individually peel packed . Box of 5 dressings. | 100% | 3 samples in an original packaging
RT42-01-262 | Dressing, Cavity wound, Ceramics. Non-adhesive pad .Non-woven sachet filled with micro-porous ceramic. Used as cavity dressing with high absorbing capacity. Sterile, individually peel packed Size: 10cmx10cm. Box of 5 dressings. | 100% | 3 samples in an original packaging
RT42-01-263 | Dressing, Cavity wound, Ceramics. Non-adhesive pad .Non-woven sachet filled with micro-porous ceramic. Used as cavity dressing with high absorbing capacity. Sterile, individually peel packed Size: 10cmx15cm. Box of 5 dressings. | 100% | 3 samples in an original packaging
RT42-01-264 | Dressing, Tracheostomy Wound, Ceramics. . Non-woven sachet filled with micro-porous ceramic shaped to facilitate tracheostomy tube. Sterile, individually peel packed Size: 7cmx7cm. Box of 5 dressings. | 100% | 3 samples in an original packaging
RT42-02-001 | Bandage, compression, flexible 2-layer , non-adhesive, self-adherent, felastic wrap "latex free". Inner layer 100mm x 2.7m unstretched. Compression layer 100mm x 4.5m fully stretched and 100mm x 3.5m unstretched. Consisting of porous, non-woven, polyester. Size: 100mm x 2.7m Individually wrapped Roll | Not applicable | 3 samples in an original packaging
RT42-02-002 | Bandage, compression, flexible 2-layer , non-adhesive, self-adherent, elastic wrap "latex free". Inner layer 75mm x 2.7m unstretched. Compression layer 75mm x 4.5m fully stretched and 75mm x 3.5m unstretched. Consisting of porous, non-woven, polyester. Size: 75mm x 2.7m Roll | Not applicable | 3 samples in an original packaging
RT42-02-003 | Bandage, porous, conforming, surgical. To consist with a woven cotton cloth backed with a porous, zinc oxide filled hypoallergenic adhesive, B.P. Bandage to have strong support with moderate compression Size: 25 mm x not less than 3m stretched Individually wrapped Roll | Not applicable | 3 samples in an original packaging
RT42-02-004 | Bandage, porous, conforming, surgical. To consist with a woven cotton cloth backed with a porous, zinc oxide filled hypoallergenic adhesive, B.P. Bandage to have strong support with moderate compression Size: 50mm x not less than 2.7m unstretched, 4.5m stretched Individually wrapped Roll | 100% | 3 samples in an original packaging
RT42-02-005 | Bandage, porous, conforming, surgical. To consist with a woven cotton cloth backed with a porous, zinc oxide filled hypoallergenic adhesive, B.P. Bandage to have strong support with moderate compression Size: 75mm x not less than 2.7m unstretched, 4.5m stretched Individually wrapped Roll | 100% | 3 samples in an original packaging
RT42-02-006 | Bandage, porous, conforming, surgical. To consist with a woven cotton cloth backed with a porous, zinc oxide filled hypoallergenic adhesive, B.P. Bandage to have strong support with moderate compression Size: 100mm x not less than 2.7m unstretched, 4.5m stretched Individually wrapped Roll | 100% | 3 samples in an original packaging
RT42-02-007 | Bandage, elastic, adhesive, porous, conforming, surgical. To consist of a woven cotton cloth backed with a porous, zinc oxide filled hypoallergenic adhesive, B.P. Bandage to have strong support, with moderate compression Size: 50mm x 4.5m stretched Laminated to an easy release silicone backing paper Individually wrapped Roll | 100% | 3 samples in an original packaging
RT42-02-008 | Bandage, elastic, adhesive, porous, conforming, surgical. To consist of a woven cotton cloth backed with a porous, zinc oxide filled hypoallergenic adhesive, B.P. Bandage to have strong support, with moderate compression Size: 75mm x 4.5m stretched Laminated to an easy release silicone backing paper Individually wrapped Roll | 100% | 3 samples in an original packaging
RT42-02-009 | Bandage, elastic, adhesive, porous, conforming, surgical. To consist of a woven cotton cloth backed with a porous, zinc oxide filled hypoallergenic adhesive, B.P. Bandage to have strong support, with moderate compression Size: 100mm x 4.5m stretched Laminated to an easy release silicone backing paper Individually wrapped Roll | 100% | 3 samples in an original packaging
RT42-02-010 | Bandage, retention, white, open wove, type 1 Size: 75mm x 5m Individually wrapped To comply with the latest issue of SANS 443 Compliance certificate to be submitted with bid Roll | Not applicable | 3 samples in an original packaging
RT42-02-011 | Bandage, retention, white, open wove, type 1 Size: 100mm x 5m Individually wrapped To comply with the latest issue of SANS 443 Compliance certificate to be submitted with bid Roll | Not applicable | 3 samples in an original packaging
RT42-02-012 | Bandage, tubular, elastic, mesh For light non-compression, retention of dressings. Size: 0 Unstretched length: 25M Unstretched width: 10mm To comply with the latest issue of CKS330 Compliance certificate to be submitted with bid Roll, individually boxed | Not applicable | 3 samples in an original packaging
RT42-02-013 | Bandage, tubular, elastic, mesh For light non-compression, retention of dressings. Size: 1 Unstretched length: 25m Unstretched width: 15mm To comply with the latest issue of CKS330 Compliance certificate to be submitted with bid Roll, individually boxed | Not applicable | 3 samples in an original packaging
RT42-02-014 | Bandage, tubular, elastic, mesh For light non-compression, retention of dressings. Size: 2 Unstretched length: 25m Unstretched width: 20mm To comply with the latest issue of CKS330 Compliance certificate to be submitted with bid Roll, individually boxed | Not applicable | 3 samples in an original packaging
RT42-02-015 | Bandage, tubular, elastic, mesh For light non-compression, retention of dressings. Size: 3 Unstretched length: 25m Unstretched width: 25mm To comply with the latest issue of CKS330 Compliance certificate to be submitted with bid Roll, individually boxed | Not applicable | 3 samples in an original packaging
RT42-02-016 | Bandage, tubular, elastic, mesh For light non-compression, retention of dressings. Size: 4 Unstretched length: 25m Unstretched width: 35mm To comply with the latest issue of CKS330 Compliance certificate to be submitted with bid Roll, individually boxed | Not applicable | 3 samples in an original packaging
RT42-02-017 | Bandage, tubular, elastic, mesh For light non-compression, retention of dressings. Size: 5 Unstretched length: 25m Unstretched width: 65mm To comply with the latest issue of CKS330 Compliance certificate to be submitted with bid Roll, individually boxed | Not applicable | 3 samples in an original packaging
RT42-02-018 | Bandage, tubular, elastic, mesh For light non-compression, retention of dressings. Size: 6 Unstretched length: 25m Unstretched width: 75mm To comply with the latest issue of CKS330 Compliance certificate to be submitted with bid Roll, individually boxed | Not applicable | 3 samples in an original packaging
RT42-02-019 | Bandage, tubular, elastic, mesh For light non-compression, retention of dressings. Size: 7 Unstretched length: 25m Unstretched width: 85mm To comply with the latest issue of CKS330 Compliance certificate to be submitted with bid Roll, individually boxed | Not applicable | 3 samples in an original packaging
RT42-02-020 | Bandage, tubular, orthopaedic, knitted, cotton-type, stockinette, unbleached Size: width: 50mm Size: length: 25m To comply with the latest issue of SANS 925 Compliance certificate to be submitted with bid Roll, individually boxed | Not applicable | 3 samples in an original packaging
RT42-02-021 | Bandage, tubular, orthopaedic, knitted, cotton-type, stockinette, unbleached Size:width: 80mm Size: length: 25m To comply with the latest issue of SANS 925 Compliance certificate to be submitted with bid Roll, individually boxed | Not applicable | 3 samples in an original packaging
RT42-02-022 | Bandage, tubular, orthopaedic, knitted, cotton-type, stockinette, unbleached Size:width: 100mm Size: length: 25m To comply with the latest issue of SANS 925 Compliance certificate to be submitted with bid Roll, individually boxed | Not applicable | 3 samples in an original packaging
RT42-02-023 | Bandage, tubular, orthopaedic, knitted, cotton-type, stockinette, unbleached Size:width: 150mm Size: ength: 25m To comply with the latest issue of SANS 925 Compliance certificate to be submitted with bid Roll, individually boxed | Not applicable | 3 samples in an original packaging
RT42-02-024 | Bandage, tubular, orthopaedic, knitted, cotton-type, stockinette, unbleached Size:width: 200mm Size: length: 25m To comply with the latest issue of SANS 925 Compliance certificate to be submitted with bid Roll, individually boxed | Not applicable | 3 samples in an original packaging
RT42-02-025 | Bandage, tubular, orthopaedic, knitted, cotton-type, stockinette, unbleached Size: width: 250mm Size: length: 25m To comply with the latest issue of SANS 925 Compliance certificate to be submitted with bid Roll, individually boxed | Not applicable | 3 samples in an original packaging
RT42-02-026 | Bandage Tubular Size: A: 4,5cm x 10m Elasticised for firm support and moderate compression Light weight, autoclavable. Well ventilated, crease free with elastic yarns protected Must conform to body contours. Must be washable and reusable Able to be applied by hand without the use of an applicator Must include a measuring device To comply with the latest issue of CKS641 Compliance certificate to be submitted with bid Roll, individually boxed | Not applicable | 3 samples in an original packaging
RT42-02-027 | Bandage Tubular Size: B: 6.25cm x 10m Elasticised for firm support and moderate compression Light weight, autoclavable. Well ventilated, crease free with elastic yarns protected Must conform to body contours. Must be washable and reusable Able to be applied by hand without the use of an applicator Must include a measuring device To comply with the latest issue of CKS641 Compliance certificate to be submitted with bid Roll, individually boxed | Not applicable | 3 samples in an original packaging
RT42-02-028 | Bandage Tubular Size: C: 6.75cm x 10m Elasticised for firm support and moderate compression Light weight, autoclavable. Well ventilated, crease free with elastic yarns protected Must conform to body contours. Must be washable and reusable Able to be applied by hand without the use of an applicator Must include a measuring device To comply with the latest issue of CKS641 Compliance certificate to be submitted with bid Roll, individually boxed | Not applicable | 3 samples in an original packaging
RT42-02-029 | Bandage Tubular Size: D: 7.5cm x 10m Elasticised for firm support and moderate compression Light weight, autoclavable. Well ventilated, crease free with elastic yarns protected Must conform to body contours. Must be washable and reusable Able to be applied by hand without the use of an applicator Must include a measuring device To comply with the latest issue of CKS641 Compliance certificate to be submitted with bid Roll, individually boxed | Not applicable | 3 samples in an original packaging
RT42-02-030 | Bandage Tubular Size: E: 8.75cm x 10m Elasticised for firm support and moderate compression Light weight, autoclavable. Well ventilated, crease free with elastic yarns protected Must conform to body contours. Must be washable and reusable Able to be applied by hand without the use of an applicator Must include a measuring device To comply with the latest issue of CKS641 Compliance certificate to be submitted with bid Roll, individually boxed | Not applicable | 3 samples in an original packaging
RT42-02-031 | Bandage Tubular Size: F: 10cm x 10m Elasticised for firm support and moderate compression Light weight, autoclavable. Well ventilated, crease free with elastic yarns protected Must conform to body contours. Must be washable and reusable Able to be applied by hand withoCMut the use of an applicator Must include a measuring device To comply with the latest issue of CKS641 Compliance certificate to be submitted with bid Roll, individually boxed | Not applicable | 3 samples in an original packaging
RT42-02-032 | Bandage Tubular Size: G: 12cm x 10m Elasticised for firm support and moderate compression Light weight, autoclavable. Well ventilated, crease free with elastic yarns protected Must conform to body contours. Must be washable and reusable Able to be applied by hand without the use of an applicator Must include a measuring device To comply with the latest issue of CKS641 Compliance certificate to be submitted with bid Roll, individually boxed | Not applicable | 3 samples in an original packaging
RT42-02-033 | Bandage Tubular Size: J: 17.5cm x 10m Elasticised for firm support and moderate compression Light weight, autoclavable. Well ventilated, crease free with elastic yarns protected Must conform to body contours. Must be washable and reusable Able to be applied by hand without the use of an applicator Must include a measuring device To comply with the latest issue of CKS641 Compliance certificate to be submitted with bid Roll, individually boxed | Not applicable | 3 samples in an original packaging
RT42-02-034 | Bandage Tubular Size: L: 32.5cm x 10m Elasticised for firm support and moderate compression Light weight, autoclavable. Well ventilated, crease free with elastic yarns protected Must conform to body contours. Must be washable and reusable Able to be applied by hand without the use of an applicator Must include a measuring device To comply with the latest issue of CKS641 Compliance certificate to be submitted with bid Roll, individually boxed | Not applicable | 3 samples in an original packaging
RT42-02-035 | Bandage Tubular Size: K: 21.5cm x 10m Elasticised for firm support and moderate compression Light weight, autoclavable. Well ventilated, crease free with elastic yarns protected Must conform to body contours. Must be washable and reusable Able to be applied by hand without the use of an applicator Must include a measuring device To comply with the latest issue of CKS641 Compliance certificate to be submitted with bid Roll, individually boxed | Not applicable | 3 samples in an original packaging
RT42-02-036 | Bandage, retention, crepe type, autoclavable Size: 50mm x not less than 4.5m stretched Individually wrapped To comply with the latest issue of SANS 1427, Table 2 Compliance certificate to be submitted with bid Roll | Not applicable | 3 samples in an original packaging
RT42-02-037 | Bandage, retention, crepe type, autoclavable Size: 75mm x not less than 4.5m stretched Individually wrapped To comply with the latest issue of SANS 1427, Table 2 Compliance certificate to be submitted with bid Roll | Not applicable | 3 samples in an original packaging
RT42-02-038 | Bandage, retention, crepe type, autoclavable Size: 100mm x not less than 4.5m stretched Individually wrapped To comply with the latest issue of SANS 1427, table 2 Compliance certificate to be submitted with bid Roll | Not applicable | 3 samples in an original packaging
RT42-02-039 | Bandage, retention, crepe type, autoclavable Size: 100mm x not less than 4.5m stretched Individually wrapped, sterile To comply with the latest issue of SANS 1427, table 2 Compliance certificate to be submitted with bid Roll | Not applicable | 3 samples in an original packaging
RT42-02-040 | Bandage, retention, crepe type, autoclavable Size: 150mm x not less than 4.5m stretched Individually wrapped To comply with the latest issue of SANS 1427, table 2 Compliance certificate to be submitted with bid Roll | Not applicable | 3 samples in an original packaging
RT42-02-041 | Bandage, retention, crepe type, autoclavable Size: 150mm x not less than 4.5m stretched Individually wrapped, sterile To comply with the latest issue of SANS 1427, table 2 Compliance certificate to be submitted with bid Roll | Not applicable | 3 samples in an original packaging
RT42-02-042 | Bandage, heavy support/ compression, cotton crepe type 1 Size: 75mm x not less than 4.5m stretched Individually wrapped To comply with the latest issue of SANS 945, part 1 Compliance certificate to be submitted with bid Roll | 100% | 3 samples in an original packaging
RT42-02-043 | Bandage, heavy support/ compression, cotton type, type 1 Size: 100mm x not less than 4.5m stretched Individually wrapped To comply with the latest issue of SANS 945, part 1 Compliance certificate to be submitted with bid Roll | 100% | 3 samples in an original packaging
RT42-02-044 | Bandage, light support, cotton crepe type 2, autoclavable Size: 50mm x not less than 4.5m stretched Individually wrapped To comply with the latest issue of SANS 945, Part 1 Compliance certificate to be submitted with bid Roll | Not applicable | 3 samples in an original packaging
RT42-02-045 | Bandage, light support, cotton crepe type 2, autoclavable Size: 75mm x not less than 4.5m stretched Individually wrapped To comply with the latest issue of SANS 945, Part 1 Compliance certificate to be submitted with bid Roll | Not applicable | 3 samples in an original packaging
RT42-02-046 | Bandage, light support, cotton crepe type 2, autoclavable Size: 100mm x not less than 4.5m stretched Individually wrapped To comply with the latest issue of SANS 945, Part 1 Compliance certificate to be submitted with bid Roll | Not applicable | 3 samples in an original packaging
RT42-02-047 | Bandage, light support, cotton crepe type 2, autoclavable Size: 150mm x not less than 4.5m stretched Individually wrapped To comply with the latest issue of SANS 945, Part 1 Compliance certificate to be submitted with bid Roll | Not applicable | 3 samples in an original packaging
RT42-02-048 | Bandage, support, triangular, calico, woven type, hemmed Size: 900mm x 900mm x 1 300mm Individually wrapped Each | 100% | 3 samples in an original packaging
RT42-02-049 | Bandage, support, triangular, non-woven, polypropylene Size: 900mm x 900mm x 1 300mm Minimum mass: 17 grams per square metre Individually wrapped To comply with the latest issue of SANS 79 Compliance certificate to be submitted with bid Each | Not applicable | 3 samples in an original packaging
RT42-02-050 | Bandage, cast, plaster of paris Leno weave base cloth, type 1 Size: 50mm x 3 m Individually wrapped and packed in a sealed container Comply with the latest issue of SANS 1308, Table 1 and 2 Compliance certificate to be submitted with bid Roll | Not applicable | 3 samples in an original packaging
RT42-02-051 | Bandage, cast, plaster of paris Leno weave base cloth, type 1 Size: 75mm x 3 m Individually wrapped and packed in a sealed container Comply with the latest issue of SANS 1308, Table 1 and 2 Compliance certificate to be submitted with bid Roll | Not applicable | 3 samples in an original packaging
RT42-02-052 | Bandage, cast, plaster of paris Leno weave base cloth, type 1 Size: 100mm x 3.5m Individually wrapped and packed in a sealed container To comply with the latest issue of SANS 1308, Table 1 and 2 Compliance certificate to be submitted with bid Roll | Not applicable | 3 samples in an original packaging
RT42-02-053 | Bandage, cast, plaster of paris Leno weave base cloth, type 1 Size: 150mm x 3.5m Individually wrapped and packed in a sealed container To comply with the latest issue of SANS 1308, Table 1 and 2 Compliance certificate to be submitted with bid Roll | Not applicable | 3 samples in an original packaging
RT42-02-054 | Bandage, cast, plaster of paris Leno weave base cloth, type 1 Size: 200mm x 3.5m Individually wrapped and packed in a sealed container To comply with the latest issue of SANS 1308, Table 1 and 2 Compliance certificate to be submitted with bid Roll | Not applicable | 3 samples in an original packaging
RT42-02-055 | Bandage, cast, plaster of paris impregnated slabs, plain woven base cloth Size: 75 mm x 750 mm Slabs wrapped in plastic, in a sealed container To comply with the latest issue of SANS 1308, Table 1 and 2 Compliance certificate to be submitted with bid Box of 50 strips | Not applicable | 3 samples in an original packaging
RT42-02-056 | Bandage, cast, plaster of paris impregnated slabs, plain woven base cloth Size: 100 mm x 750 mm Slabs wrapped in plastic, in a sealed container To comply with the latest issue of SANS 1308, Table 1 and 2 Compliance certificate to be submitted with bid Box of 50 strips | Not applicable | 3 samples in an original packaging
RT42-02-057 | Bandage, cast, plaster of paris impregnated slabs, plain woven base cloth Size: 150 mm x 750 mm Slabs wrapped in plastic, in a sealed container To comply with the latest issue of SANS 1308, Table 1 and 2 Compliance certificate to be submitted with bid Box of 50 strips | Not applicable | 3 samples in an original packaging
RT42-02-058 | Bandage casting waterproof, consisting of fibre-glass mix Light weight, fast setting, early load bearing, smooth finish Individually wrapped Size: 50mm x not less than 3.6m Roll | Not applicable | 3 samples in an original packaging
RT42-02-059 | Bandage casting waterproof, consisting of fibre-glass mix Light weight, fast setting, early load bearing, smooth finish Individually wrapped Size: 75mm x not less than 3.6m Roll | Not applicable | 3 samples in an original packaging
RT42-02-060 | Bandage casting waterproof, consisting of fibre-glass mix Light weight, fast setting, early load bearing, smooth finish Individually wrapped Size: 100mm x not less than 3.6m Roll | Not applicable | 3 samples in an original packaging
RT42-02-061 | Bandage casting waterproof, consisting of fibre-glass mix Light weight, fast setting, early load bearing, smooth finish Individually wrapped Size: 125mm x not less than 3.6m Roll | Not applicable | 3 samples in an original packaging
RT42-02-062 | Bandage, conforming knitted type Size: 50mm x not less than 4.5m stretch Table 3, KC30 Individually wrapped To comply with the latest issue of SANS 1427 Compliance certificate to be submitted with bid Roll | Not applicable | 3 samples in an original packaging
RT42-02-063 | Bandage, conforming knitted type Size: 75mm x not less than 4.5m stretch Table 3, KC30 Individually wrapped To comply with the latest issue of SANS 1427 Compliance certificate to be submitted with bid Roll | Not applicable | 3 samples in an original packaging
RT42-02-064 | Bandage, conforming knitted type Size: 100mm x not less than 4.5m stretch Table 3, KC30 Individually wrapped To comply with the latest issue of SANS 1427 Compliance certificate to be submitted with bid Roll | Not applicable | 3 samples in an original packaging
RT42-02-065 | Bandage, conforming knitted type Size: 150mm x not less than 4.5m stretch Table 3, KC30 Individually wrapped To comply with the latest issue of SANS 1427 Compliance certificate to be submitted with bid Roll | Not applicable | 3 samples in an original packaging
RT42-02-066 | Bandage, non-stretch, adhesive strapping, consist of rayon fabric, which is spread evenly with a adhesive mass containing zinc oxide. Cut to required legth. Must be uniformly spooled with adhesive face innermost onto cardboard cores. Hand-tearable, size: 40mm x 14m Each | 100% | 3 samples in an original packaging
RT42-02-067 | Bandage, non-stretch, adhesive strapping, consist of rayon fabric, which is spread evenly with a adhesive mass containing zinc oxide. Cut to required legth. Must be uniformly spooled with adhesive face innermost onto cardboard cores. Hand-tearable, size: 50mm x 14m Each | 100% | 3 samples in an original packaging
RT42-02-068 | Bandage, compression, flexible co-hesive elastic wrap, non-adhesive, selfadherent, "latex free" Consisting of porous, non-woven, polyester. Size: 25mm x 4,5m Individually wrapped Roll | Not applicable | 3 samples in an original packaging
RT42-02-069 | Bandage, compression, flexible co-hesive elastic wrap, non-adhesive, selfadherent, "latex free" Consisting of porous, non-woven, polyester. Size: 50mm x 4,5m Individually wrapped Roll | Not applicable | 3 samples in an original packaging
RT42-02-070 | Bandage, compression, flexible co-hesive elastic wrap, non-adhesive, selfadherent, "latex free" Consisting of porous, non-woven, polyester. Size: 75mm x 4,5m Individually wrapped Roll | Not applicable | 3 samples in an original packaging
RT42-02-071 | Bandage, compression, flexible co-hesive elastic wrap, non-adhesive, selfadherent, "latex free" Consisting of porous, non-woven, polyester. Size: 100mm x 4,5m Individually wrapped Roll | Not applicable | 3 samples in an original packaging
RT42-02-072 | Bandage, compression, flexible co-hesive elastic wrap, non-adhesive, selfadherent, "latex free" Consisting of porous, non-woven, polyester. Size: 150mm x 4,5m Individually wrapped Roll | Not applicable | 3 samples in an original packaging
RT42-02-073 | Bandage, compression, 3-layer, to treat complicated venous leg ulcers, mixed ulcers and arterial insufficiency ulcers, the system must consist of the following components: • 1 x short stretch, zinc paste bandage: 100mm x 5000mm stretched • 1 x viscose padding bandage: 150mm x 3000mm • 1 x cotton cohesive bandage: 100mm x 5000mm • Components must be individually packed in a suitable labelled box Pack size: Kit | Not applicable | 3 samples in an original packaging
RT42-03-001 | Cotton bud, double tipped on rolled paper stick, autoclavable Length: Not less than 80mm Container of 100 buds (Ear buds) | Not applicable | 2 containers of samples in an original packaging
RT42-03-002 | Cotton wool, absorbent, first grade, non-interleaved, 50g To comply with the latest issue of SANS 228 Compliance certificate to be submitted with bid Roll | 100% | 2 rolls in an original packaging packaging
RT42-03-003 | Cotton wool, absorbent, second grade, interleaved, 500g To comply with the latest issue of SANS 228 Compliance certificate to be submitted with bid Roll | 100% | 3 rolls in an original packaging packaging
RT42-03-004 | Cotton wool, absorbent balls, type 2, non-sterile, autoclavable, 500g package Size: Medium 1g To comply with the latest issue of SANS 149 Compliance certificate to be submitted with bid Packet | 100% | 2 packet of samples in an original packaging
RT42-03-005 | Cotton wool, absorbent balls, type 2, autoclavable, white, sterile, 5 balls per pouch, peel packed Size: Medium 1g To comply with the latest issue of SANS 149 Compliance certificate to be submitted with bid Pouch | 100% | 2 pouch of samples in an original packaging
RT42-03-006 | Cotton wool, absorbent balls, type 2, autoclavable, 500g package Size: Maternity 2g To comply with the latest issue of SANS 149 Compliance certificate to be submitted with bid Packet | 100% | 2 packet of samples in an original packaging
RT42-03-007 | Cotton wool, absorbent balls, type 2, autoclavable, white, sterile, 5 balls per pouch, peel packed Size: Maternity 2g To comply with the latest issue of SANS 149 Compliance certificate to be submitted with bid Pouch | 100% | 2 pouch of samples in an original packaging
RT42-03-008 | Tampon, vaginal, autoclavable Absorbent cotton wool with gauze made into a tube Cotton wool to comply with the latest SANS228 Compliance certificate to be submitted with bid Each | 100% | 2 samples in an original packaging
RT42-04-001 | Swabs, alcohol, 200 swabs Sterile, composed of viscose rayon non-woven fabric impregnated with minimum of 0,5ml Isopropyl Alcohol 70% in water Minimum dimensions: 24mm x 30mm Individually packed in aluminium foil pouches Box of 200 swabs | Not applicable | 2 box of samples in an original packaging
RT42-04-002 | Swab, abdominal, surgical, cotton, X-ray detectable, sterile, 8g, 4-ply Size: 170mm x 200mm Pack size: Pouch of 5 swabs Scope • Swabs for single use, all layers consist of white woven bandage cloth, to comply with the latest ;issue of SANS 443, type 3 Construction • Edges: All cut edges of the material to be folded in and stitched, leaving no frayed edges Cotton tape • Length of tape: 400mm • Width of tape: 12mm • The cotton tape must be embedded in a 50mm packet i.e. 100mm in length of tape folder double to 50mm stitched into the corner of the swab leaving a 300mm length of cotton tape clear • Breaking strength of tape – minimum 7.95kg X-Ray detectable material • X-Ray detectable material thread shall be woven in/impregnated Packaging • The swab shall be bound in bundles of five by means of either a paper band consisting of one piece secured by autoclavable glue (and not by a separate adhesive tape) or by means of an acceptable tying material which can withstand autoclaving. • The swabs shall be suitable wrapped so as to be adequately protected against soiling and contamination • Only swabs of the same size shall be packed together in a bulk container. • Packing shall comply with the relevant requirements of the latest issue of SANS 228- double swab • Swabs shall be packed in a red plastic bag  Size 280mm x 400mm, thickness at least 19micron, colour: red, Marking • The following information shall appear in legible and indelible marking on each package:  Manufacture’s name or trade mark or both  Type of swabs  Size of swabs  Quantity of swabs Bulk containers • The following information shall appear in legible and indelible marking on the outside of each bulk container. Base – Peel pouches • Clear PVC film, not less than 90u Lidding – Peel pouches / tray • Peel cover type (heat seal grid), 60g/m² medical Kraft paper lacquered with 11u adhesive Sterility indicator • Pack to contain tamperproof sterility indicator on the tray or pack, confirm sterility, bidder to indicate the type of indicator used Sterile Peel packed in bundles of 5 Pack size: Each | Not applicable | 2 packet of samples in an original packaging
RT42-04-003 | Swab, abdominal, surgical, cotton, X-ray detectable, sterile, 8g, 6-ply Size: 170mm x 200mm Pack size: Pouch of 5 swabs Scope • Swabs for single use, all layers consist of white woven bandage cloth, to comply with the latest ;issue of SANS 443, type 3 Construction • Edges: All cut edges of the material to be folded in and stitched, leaving no frayed edges Cotton tape • Length of tape: 400mm • Width of tape: 12mm • The cotton tape must be embedded in a 50mm packet i.e. 100mm in length of tape folder double to 50mm stitched into the corner of the swab leaving a 300mm length of cotton tape clear • Breaking strength of tape – minimum 7.95kg X-Ray detectable material • X-Ray detectable material thread shall be woven in/impregnated Packaging • The swab shall be bound in bundles of five by means of either a paper band consisting of one piece secured by autoclavable glue (and not by a separate adhesive tape) or by means of an acceptable tying material which can withstand autoclaving. • The swabs shall be suitable wrapped so as to be adequately protected against soiling and contamination • Only swabs of the same size shall be packed together in a bulk container. • Packing shall comply with the relevant requirements of the latest issue of SANS 228- double swab • Swabs shall be packed in a red plastic bag  Size 280mm x 400mm, thickness at least 19micron, colour: red, Marking • The following information shall appear in legible and indelible marking on each package:  Manufacture’s name or trade mark or both  Type of swabs  Size of swabs  Quantity of swabs Bulk containers • The following information shall appear in legible and indelible marking on the outside of each bulk container. Base – Peel pouches • Clear PVC film, not less than 90u Lidding – Peel pouches / tray • Peel cover type (heat seal grid), 60g/m² medical Kraft paper lacquered with 11u adhesive Sterility indicator • Pack to contain tamperproof sterility indicator on the tray or pack, confirm sterility, bidder to indicate the type of indicator used Sterile Peel packed in bundles of 5 Pack size: Each | 100% | 2 packet of samples in an original packaging
RT42-04-004 | Swab, abdominal, surgical, cotton, X-ray detectable, sterile, 8g, 8-ply Size: 170mm x 200mm Pack size: Pouch of 5 swabs Scope • Swabs for single use, all layers consist of white woven bandage cloth, to comply with the latest ;issue of SANS 443, type 3 Construction • Edges: All cut edges of the material to be folded in and stitched, leaving no frayed edges Cotton tape • Length of tape: 400mm • Width of tape: 12mm • The cotton tape must be embedded in a 50mm packet i.e. 100mm in length of tape folder double to 50mm stitched into the corner of the swab leaving a 300mm length of cotton tape clear • Breaking strength of tape – minimum 7.95kg X-Ray detectable material • X-Ray detectable material thread shall be woven in/impregnated Packaging • The swab shall be bound in bundles of five by means of either a paper band consisting of one piece secured by autoclavable glue (and not by a separate adhesive tape) or by means of an acceptable tying material which can withstand autoclaving. • The swabs shall be suitable wrapped so as to be adequately protected against soiling and contamination • Only swabs of the same size shall be packed together in a bulk container. • Packing shall comply with the relevant requirements of the latest issue of SANS 228- double swab • Swabs shall be packed in a red plastic bag  Size 280mm x 400mm, thickness at least 19micron, colour: red, Marking • The following information shall appear in legible and indelible marking on each package:  Manufacture’s name or trade mark or both  Type of swabs  Size of swabs  Quantity of swabs Bulk containers • The following information shall appear in legible and indelible marking on the outside of each bulk container. Base – Peel pouches • Clear PVC film, not less than 90u Lidding – Peel pouches / tray • Peel cover type (heat seal grid), 60g/m² medical Kraft paper lacquered with 11u adhesive Sterility indicator • Pack to contain tamperproof sterility indicator on the tray or pack, confirm sterility, bidder to indicate the type of indicator used Sterile Peel packed in bundles of 5 Pack size: Each | Not applicable | 2 packet of samples in an original packaging
RT42-04-005 | Swab, abdominal, surgical, cotton, X-ray detectable, sterile, 26g, 6-ply Size: 350mm x 350mm Pack size: Pouch of 5 swabs Scope • Swabs for single use, all layers consist of white woven bandage cloth, to comply with the latest issue of SANS 443, type 3 Construction • Edges: All cut edges of the material to be folded in and stitched, leaving no frayed edges Cotton tape • Length of tape: 400mm • Width of tape: 12mm • The cotton tape must be embedded in a 50mm packet i.e. 100mm in length of tape folder double to 50mm stitched into the corner of the swab leaving a 300mm length of cotton tape clear • Breaking strength of tape – minimum 7.95kg X-Ray detectable material • X-Ray detectable material thread shall be woven in/impregnated Packaging • The swab shall be bound in bundles of five by means of either a paper band consisting of one piece secured by autoclavable glue (and not by a separate adhesive tape) or by means of an acceptable tying material which can withstand autoclaving. • The swabs shall be suitable wrapped so as to be adequately protected against soiling and contamination • Only swabs of the same size shall be packed together in a bulk container. • Packing shall comply with the relevant requirements of the latest issue of SANS 228- double swab • Swabs shall be packed in a red plastic bag  Size 280mm x 400mm, thickness at least 19micron, colour: red, Marking • The following information shall appear in legible and indelible marking on each package:  Manufacture’s name or trade mark or both  Type of swabs  Size of swabs  Quantity of swabs Bulk containers • The following information shall appear in legible and indelible marking on the outside of each bulk container. Base – Peel pouches • Clear PVC film, not less than 90u Lidding – Peel pouches / tray • Peel cover type (heat seal grid), 60g/m² medical Kraft paper lacquered with 11u adhesive Sterility indicator • Pack to contain tamperproof sterility indicator on the tray or pack, confirm sterility, bidder to indicate the type of indicator used Sterile Peel packed in bundles of 5 Pack size: Each | Not applicable | 2 packet of samples in an original packaging
RT42-04-006 | Swab, abdominal, surgical, cotton, X-ray detectable, sterile, 36g, 6-ply Size: 370mm x 450mm Pack size: Pouch of 5 swabs Scope • Swabs for single use, all layers consist of white woven bandage cloth, to comply with the latest issue of SANS 443, type 3 Construction • Edges: All cut edges of the material to be folded in and stitched, leaving no frayed edges Cotton tape • Length of tape: 400mm • Width of tape: 12mm • The cotton tape must be embedded in a 50mm packet i.e. 100mm in length of tape folder double to 50mm stitched into the corner of the swab leaving a 300mm length of cotton tape clear • Breaking strength of tape – minimum 7.95kg X-Ray detectable material • X-Ray detectable material thread shall be woven in/impregnated Packaging • The swab shall be bound in bundles of five by means of either a paper band consisting of one piece secured by autoclavable glue (and not by a separate adhesive tape) or by means of an acceptable tying material which can withstand autoclaving. • The swabs shall be suitable wrapped so as to be adequately protected against soiling and contamination • Only swabs of the same size shall be packed together in a bulk container. • Packing shall comply with the relevant requirements of the latest issue of SANS 228- double swab • Swabs shall be packed in a red plastic bag  Size 280mm x 400mm, thickness at least 19micron, colour: red, Marking • The following information shall appear in legible and indelible marking on each package:  Manufacture’s name or trade mark or both  Type of swabs  Size of swabs  Quantity of swabs Bulk containers • The following information shall appear in legible and indelible marking on the outside of each bulk container. Base – Peel pouches • Clear PVC film, not less than 90u Lidding – Peel pouches / tray • Peel cover type (heat seal grid), 60g/m² medical Kraft paper lacquered with 11u adhesive Sterility indicator • Pack to contain tamperproof sterility indicator on the tray or pack, confirm sterility, bidder to indicate the type of indicator used Sterile Peel packed in bundles of 5 Pack size: Each | 100% | 2 packet of samples in an original packaging
RT42-04-007 | Swab, abdominal, surgical, cotton, X-ray detectable, Non sterile, 8g, 6-ply Size: 170mm x 200mm Pack size: Pouch of 5 swabs Scope • Swabs for single use, all layers consist of white woven bandage cloth, to comply with the latest issue of SANS 443, type 3 Construction • Edges: All cut edges of the material to be folded in and stitched, leaving no frayed edges Cotton tape • Length of tape: 400mm • Width of tape: 12mm • The cotton tape must be embedded in a 50mm packet i.e. 100mm in length of tape folder double to 50mm stitched into the corner of the swab leaving a 300mm length of cotton tape clear • Breaking strength of tape – minimum 7.95kg X-Ray detectable material • X-Ray detectable material thread shall be woven in/impregnated Packaging • The swab shall be bound in bundles of five by means of either a paper band consisting of one piece secured by autoclavable glue (and not by a separate adhesive tape) or by means of an acceptable tying material which can withstand autoclaving. • The swabs shall be suitable wrapped so as to be adequately protected against soiling and contamination • Only swabs of the same size shall be packed together in a bulk container. • Packing shall comply with the relevant requirements of the latest issue of SANS 228- double swab • Swabs shall be packed in a red plastic bag  Size 280mm x 400mm, thickness at least 19micron, colour: red, Marking • The following information shall appear in legible and indelible marking on each package:  Manufacture’s name or trade mark or both  Type of swabs  Size of swabs  Quantity of swabs Bulk containers • The following information shall appear in legible and indelible marking on the outside of each bulk container. Base – Peel pouches • Clear PVC film, not less than 90u Lidding – Peel pouches / tray • Peel cover type (heat seal grid), 60g/m² medical Kraft paper lacquered with 11u adhesive Sterility indicator • Pack to contain tamperproof sterility indicator on the tray or pack, confirm sterility, bidder to indicate the type of indicator used Sterile Peel packed in bundles of 5 Pack size: Each | 100% | 2 packet of samples in an original packaging
RT42-04-008 | Swab, abdominal, surgical, cotton, X-ray detectable, Non sterile, 36g, 6-ply Size: 370mm x 450mm Pack size: Pouch of 5 swabs Scope • Swabs for single use, all layers consist of white woven bandage cloth, to comply with the latest issue of SANS 443, type 3 Construction • Edges: All cut edges of the material to be folded in and stitched, leaving no frayed edges Cotton tape • Length of tape: 400mm • Width of tape: 12mm • The cotton tape must be embedded in a 50mm packet i.e. 100mm in length of tape folder double to 50mm stitched into the corner of the swab leaving a 300mm length of cotton tape clear • Breaking strength of tape – minimum 7.95kg X-Ray detectable material • X-Ray detectable material thread shall be woven in/impregnated Packaging • The swab shall be bound in bundles of five by means of either a paper band consisting of one piece secured by autoclavable glue (and not by a separate adhesive tape) or by means of an acceptable tying material which can withstand autoclaving. • The swabs shall be suitable wrapped so as to be adequately protected against soiling and contamination • Only swabs of the same size shall be packed together in a bulk container. • Packing shall comply with the relevant requirements of the latest issue of SANS 228- double swab • Swabs shall be packed in a red plastic bag  Size 280mm x 400mm, thickness at least 19micron, colour: red, Marking • The following information shall appear in legible and indelible marking on each package:  Manufacture’s name or trade mark or both  Type of swabs  Size of swabs  Quantity of swabs Bulk containers • The following information shall appear in legible and indelible marking on the outside of each bulk container. Base – Peel pouches • Clear PVC film, not less than 90u Lidding – Peel pouches / tray • Peel cover type (heat seal grid), 60g/m² medical Kraft paper lacquered with 11u adhesive Sterility indicator • Pack to contain tamperproof sterility indicator on the tray or pack, confirm sterility, bidder to indicate the type of indicator used Sterile Peel packed in bundles of 5 Pack size: Each | Not applicable | 2 packet of samples in an original packaging
RT42-04-009 | Swabs, cleansing, Type 3, 100 swabs (Balls) Diameter: 30mm Materials: Absorbent cotton wool to comply to SANS 228-1st Grade Absorbent gauze to comply to SANS 446 Compliance certificate to be submitted with bid Diameter: 30mm Construction: The cotton wool shall be packed in single layer of gauze covering material. The covering material shall be secured with cotton thread to prevent cotton wool from spilling out of the outer covering. The swabs shall be loosely stitched together in bundles of five by means of an easy to break cotton thread. Packet/Bag | 100% | 2 packet of samples in an original packaging
RT42-04-010 | Swabs, cleansing, Type 3 100 swabs (balls) Diameter: 65mm Materials: Absorbent cotton wool as per latest issue SANS 228-1st Grade Absorbent gauze as per latest issue of SANS 446 Compliance certificate to be submitted with bid Construction: The cotton wool shall be packed in single layer of gauze covering material. The covering material shall be secured with cotton thread to prevent cotton wool from spilling out of the outer covering. The swabs shall be loosely stitched together in bundles of five by means of an easy to break cotton thread. Packet/Bag | 100% | 2 packet of samples in an original packaging
RT42-04-011 | Swabs, surgical, dissecting, Type 1, X-Ray detectable, 100 swabs Size: 10mm Materials: Outer layer The outer layer shall be gauze as per the latest issue of SANS 446 Cotton wool The cotton wool shall be cotton wool Grade 2 as per the latest issue of SANS 228 Compliance certificate to be submitted with bid Construction: The cotton wool shall be packed in single layer of gauze X-ray detectable material of size 5mm x 5mm to be placed between the cotton wool and the outer layer of gauze so as to be visible from the outside. The gauze must be secured with cotton thread of acceptable quality. The swabs shall be loosely stitched together in bundles of five by means of an uneasy to break cotton thread. Packet/bag | 100% | 2 packet of samples in an original packaging
RT42-04-012 | Swabs, gauze, absorbent, type 3, bleached, 8 ply, 100 swabs Size: 100mm x 100mm Non sterile To comply with the latest issue of SANS 446 Compliance certificate to be submitted with bid Packet | 100% | 2 packet of samples in an original packaging
RT42-04-013 | Swabs, gauze, absorbent, type 3, bleached, 12 ply, 100 swabs Size: 75mm x 75mm Non sterile To comply with the latest issue of SANS 446 Compliance certificate to be submitted with bid Packet | 100% | 2 packet of samples in an original packaging
RT42-04-014 | Swabs, gauze, absorbent, type 3, bleached, 12 ply, 100 swabs Size: 100mm x 100mm Non-Sterile To comply with the latest issue of SANS 446 Compliance certificate to be submitted with bid Packet | 100% | 2 packet of samples in an original packaging
RT42-04-015 | Swabs, gauze, absorbent, type 3, bleached, sterile, 8 ply, 1x 5 swabs Size: 100mm x 100mm Sterile peel packed To comply with the latest issue of SANS 446 Compliance certificate to be submitted with bid Packet | 100% | 2 packet of samples in an original packaging
RT42-04-016 | Swabs, gauze, absorbent, type 3, bleached, sterile, 12 ply, 1 x 5 swabs Size: 75mm x 75mm Sterile peel packed To comply with the latest issue of SANS 446 Compliance certificate to be submitted with bid Packet | 100% | 2 packet of samples in an original packaging
RT42-04-017 | Swabs, gauze, absorbent, type 3, bleached, sterile, 12 ply, 1 x 5 swabs Size: 100mm x 100mm Sterile peel packed To comply with the latest issue of SANS 446 Compliance certificate to be submitted with bid Packet | 100% | 2 packet of samples in an original packaging
RT42-04-018 | Swabs, gauze, absorbent, type 3, bleached, sterile, X-Ray detectable, 8 ply, 5 swabs Size: 100mm x 100mm Sterile peel packed To comply with the latest issue of SANS 446 Compliance certificate to be submitted with bid Packet | 100% | 2 packet of samples in an original packaging
RT42-04-019 | Swabs, gauze, absorbent, type 1, bleached, sterile X-Ray detectable, 12 ply, 1 x 5 swabs Size: 100mm x 100mm Sterile peel packed To comply with the latest issue of SANS 446 Compliance certificate to be submitted with bid Packet | 100% | 2 packet of samples in an original packaging
RT42-04-020 | Swabs, gauze, absorbent, type 3, X-Ray detectable, 12 ply, 100 swabs Size: 75mm x 75mm Non Sterile, autoclavable To comply with the latest issue of SANS 446 Compliance certificate to be submitted with bid Packet | 100% | 2 packet of samples in an original packaging
RT42-04-021 | Swabs, gauze, absorbent, type 3, X-Ray detectable, 12 ply, 5 swabs Size: 75mm x 75mm Sterile, autoclavable To comply with the latest issue of SANS 446 Compliance certificate to be submitted with bid Packet | 100% | 2 packet of samples in an original packaging
RT42-04-022 | Swabs, gauze, absorbent, type 1, X-Ray detectable, 12 ply, 100 swabs Size: 100mm x 100mm Non sterile, autoclavable To comply with the latest issue of SANS 446 Compliance certificate to be submitted with bid Packet | 100% | 2 packet of samples in an original packaging
RT42-04-023 | Swabs, gauze, absorbent, type 1, X-Ray detectable, 12 ply, 5 swabs Size: 100mm x 100mm Sterile, autoclavable To comply with the latest issue of SANS 446 Compliance certificate to be submitted with bid Packet | 100% | 2 packet of samples in an original packaging
RT42-04-024 | Swabs, gauze, absorbent, type 1, dyed green, 12 ply, 100 swabs Size: 100mm x 100mm Non sterile, autoclavable To comply with the latest issue of SANS 446 Compliance certificate to be submitted with bid Packet | 100% | 2 packet of samples in an original packaging
RT42-04-025 | Cleansing Balls 65mm diameter impregnated with 0,5% Chlorhexidine Gluconate in 70% Alcohol. Skin Preparation, Pre-Operation. Sterile, double peel packed. 5 loose balls per package | Not applicable | 2 packet of samples in an original packaging
RT42-04-026 | Cleansing Balls 65mm diameter impregnated with 2% Chlorhexidine Gluconate in water. Skin Preparation, Pre-Operation. Sterile, double peel packed. 5 loose balls per package | Not applicable | 2 packet of samples in an original packaging
RT42-04-027 | Cleansing Balls 65mm diameter impregnated with 7.5% Povidone Iodine.Skin Preparation, Pre-Operation. Sterile, double peel packed. 5 loose balls per package | Not applicable | 2 packet of samples in an original packaging
RT42-04-028 | Cleansing Balls 65mm diameter impregnated with 10% Povidone Iodine.Skin Preparation, Pre-Operation. Sterile, double peel packed. 5 loose balls per package | Not applicable | 2 packet of samples in an original packaging
RT42-04-029 | Cleansing Balls 30mm diameter impregnated with 0,5% Chlorhexidine Gluconate in 70% Alcohol. Skin Preparation, Pre-Operation. Sterile, double peel packed. 5 loose balls per package | Not applicable | 2 packet of samples in an original packaging
RT42-04-030 | Cleansing Balls 30mm diameter impregnated with 2% Chlorhexidine Gluconate in water. Skin Preparation, Pre-Operation. Sterile, double peel packed. 5 loose balls per package | Not applicable | 2 packet of samples in an original packaging
RT42-04-031 | Cleansing Balls 30mm diameter impregnated with 7.5% Povidone Iodine.Skin Preparation, Pre-Operation. Sterile, double peel packed. 5 loose balls per package | Not applicable | 2 packet of samples in an original packaging
RT42-04-032 | Cleansing Balls 30mm diameter impregnated with 10% Povidone Iodine.Skin Preparation, Pre-Operation. Sterile, double peel packed. 5 loose balls per package | Not applicable | 2 packet of samples in an original packaging
RT42-05-001 | Gauze, ribbon, absorbent cotton Autoclavable, fast edge Size: 12mm x 10m To comply with the latest issue of SANS 446 Compliance certificate to be submitted with bid Roll | 100% | 2 rolls of samples in an original packaging
RT42-05-002 | Gauze, ribbon, absorbent cotton, roll Autoclavable, fast edge Size: 25mm x 10m To comply with the latest issue of SANS 446 Compliance certificate to be submitted with bid Roll | 100% | 2 rolls of samples in an original packaging
RT42-05-003 | Gauze, ribbon, absorbent cotton Autoclavable, fast edge Size: 50mm x 10m To comply with the latest issue of SANS 446 Compliance certificate to be submitted with bid Roll | 100% | 2 rolls of samples in an original packaging
RT42-05-004 | Gauze, ribbon, absorbent cotton Autoclavable, fast edge Size: 75mm x 10m To comply with the latest issue of SANS 446 Compliance certificate to be submitted with bid Roll | 100% | 2 rolls of samples in an original packaging
RT42-05-005 | Gauze, ribbon, absorbent cotton Autoclavable, fast edge Size: 100mm x 10m To comply with the latest issue of SANS 446 Compliance certificate to be submitted with bid Roll | 100% | 2 rolls of samples in an original packaging
RT42-05-006 | Gauze, absorbent cotton, type 1, roll Size: 0.9m x 5m Individually packed Non-sterile To comply with the latest issue of SANS 446 Compliance certificate to be submitted with bid Each | 100% | 3 samples in an original packaging
RT42-05-007 | Gauze, plugging, cotton, type 1, absorbent, X-Ray detectable, 100 swabs Size: 38mm x 700mm x 12 ply Non sterile To comply with the latest issue of SANS 446 Compliance certificate to be submitted with bid Packet | 100% | 2 packet of samples in an original packaging
RT42-05-008 | Gauze, plugging, cotton, type 1, absorbent, X-Ray detectable, 100 swabs Size: 75mm x 900mm x 12 ply Non sterile To comply with the latest issue of SANS 446 Compliance certificate to be submitted with bid Packet | 100% | 2 packet of samples in an original packaging
RT42-05-009 | Gauze, Oxidised cellulose, BP, USP Absorbable, haemostatic Sterile, individually peel packed Size: 50mm x 12mm Each | Not applicable | 3 samples in an original packaging
RT42-05-010 | Gauze, Oxidised cellulose, BP, USP Absorbable, haemostatic Sterile, individually peel packed Size: with allowance of 5mm , 70mm x 50mm Each | Not applicable | 3 samples in an original packaging
RT42-05-011 | Gauze, Oxidised cellulose, BP, USP Absorbable, haemostatic Sterile, individually peel packed Size: with allowance of 5mm , 350mm x 50mm Each | Not applicable | 3 samples in an original packaging
RT42-05-012 | Gauze, oxidised cellulose, BP, USP, absorbable, haemostatic. Size: with allowance of 5mm, 200mm x 100mm. Sterile. Each | Not applicable | 3 samples in an original packaging
RT42-06-001 | Pads, eye, oval, autoclavable, 144 pads Composed of a 6mm to 8mm layer of cotton wool Grade 1 To comply with the latest issue of SANS 228 Enclosed between two layers of absorbent gauze Type 3, to comply with the latest issue of SANS 446 Compliance certificate to be submitted with bid Packet | 100% | 2 packets of samples in an original packaging
RT42-06-002 | Pads, eye, oval, sterile Size: 67mm x 54mm at widest point Sterile, individually peel packed To comply with the latest issue of SANS 228 Enclosed between two layers of absorbent gauze Type 3, to comply with the latest issue of SANS 446 Compliance certificate to be submitted with bid Each | 100% | 3 samples in an original packaging
RT42-06-003 | Pad, surgical, absorbent, sterile Rayon filler with a non-woven fabric sleeve Size: 100mm x 100mm Sterile, individually peel packed To comply with the latest issue of SANS 8840 Compliance certificate to be submitted with bid Each | Not applicable | 3 samples in an original packaging
RT42-06-004 | Pad, surgical, absorbent, sterile Rayon filler with a non-woven fabric sleeve Size: 100mm x 200mm Sterile, individually peel packed To comply with the latest issue of SANS 8840 Compliance certificate to be submitted with bid Each | Not applicable | 3 samples in an original packaging
RT42-06-005 | Pad, surgical, autoclavable non-woven, type B For wound and swabbing Mass: 57g per 10 pads Size: 100mm x 200mm To comply with the latest issue of SANS 8840 Compliance certificate to be submitted with bid Each | Not applicable | 3 samples in an original packaging
RT42-06-006 | Pad, surgical, autoclavable non-woven, type B For wound and swabbing Mass: 109g per 10 pads Size: 200mm x 200mm To comply with the latest issue of SANS 8840 Compliance certificate to be submitted with bid Each | Not applicable | 3 samples in an original packaging
RT42-06-007 | Pad, wound debridement, the pad must be able to debride the wound bed and wound edges. The pad must be able to hydrate the perilesional area and the entire limb. The pad must contain a poloxamer based solution and an isotonic surfactant which is non-cytotoxic to wounds. Individually packed Sterile Size of pad: ±20cm x ±20cm Dispenser box with 10 pouches Each | Not applicable | 3 samples in an original packaging
RT42-07-001 | Drape, surgical, adhesive, U-slot aperture, 1 x Wrap Drape: Size: 1200mm x 1300mm Pear-shaped split with incise Poly Reinforcement Material, adhesive and handles as per Drape surgical incise adhesive above Each | 100% | 2 samples in an original packaging
RT42-07-002 | Drape, surgical, adhesive, body split, 1 x Wrap Drape: Size: 1910mm x 2950mm Split: 254mm x 635 mm Absorbent Reinforcement: 1016mm x 1118mm Pear-shaped split with incise Poly Reinforcement Material, adhesive and handles as per Drape surgical incise adhesive above Each | 100% | 2 samples in an original packaging
RT42-07-003 | Drape, surgical, chest/breast with fluid control pouch, 1 x Wrap Drape: Size: 1960mm x 3050mm Fenestration Pear Shaped Poly Reinforcement Tube holders Each | 100% | 2 samples in an original packaging
RT42-07-004 | Drape, surgical, adhesive patch with circle aperture centrally situated, Size: 400mm x 400mm Adhesive area: 100mm x 125mm Circle aperture: 60mm diameter Material, adhesive and handles as per Drape surgical incise adhesive above Each | 100% | 2 samples in an original packaging
RT42-07-005 | Drape, surgical, adhesive patch with oval aperture situated off-centre, Size: 800mm x 1300mm Adhesive area: 100mm x 125mm Oval aperture: 70mm x 50mm Material, adhesive and handles as per Drape surgical incise adhesive above Each | 100% | 2 samples in an original packaging
RT42-07-006 | Drape, surgical, adhesive patch with oval aperture, drainage pouch as part of drape, Size: 230mm x 450mm Adhesive area: 230mm x 250mm Oval aperture: 80mm x 100mm Material, adhesive and handles as per Drape surgical incise adhesive above Each | 100% | 2 samples in an original packaging
RT42-07-007 | Drape, disposable, with double instrument pouch compartments Overall size: 180cm (width) x 300cm (length) 2 x adhesive strips 5cm x 18cm Sterile Individually peel packed with description on outer pack Pack size: each | 100% | 2 samples in an original packaging
RT42-07-008 | Drape/dressing: Sterile. Self adhesive, transparent, latex free. Material to be an adhesive viscose, polyurethane or polyester or polyethylene film, occlusive to bacteria. The adhesive shall be a hypoallergenic acrylic copolymer. The film to be waterproof and have a siliconised release paper without a“window”frame, Individually packed in a peel pouch. To comply with the latest issue of SANS 53795 Compliance certificate to be submitted with bid Size:120mm x 250mm Adhesive area:100mm x 250mm Each | 100% | 2 samples in an original packaging
RT42-07-009 | Drape/dressing: Sterile. Self adhesive, transparent, latex free. Material to be an adhesive viscose, polyurethane or polyester or polyethylene film, occlusive to bacteria. The adhesive shall be a hypoallergenic acrylic copolymer. The film to be waterproof and have a siliconised release paper without a“window”frame, Individually packed in a peel pouch. To comply with the latest issue of SANS 53795 Compliance certificate to be submitted with bid Size: 300mm x 280mm Adhesive area:300mm x 200mm Each | 100% | 2 samples in an original packaging
RT42-07-010 | Drape/dressing: Sterile. Self adhesive, transparent, latex free. Material to be an adhesive viscose, polyurethane or polyester or polyethylene film, occlusive to bacteria. The adhesive shall be a hypoallergenic acrylic copolymer. The film to be waterproof and have a siliconised release paper without a“window”frame, Individually packed in a peel pouch. Size:400mm x 420mm Adhesive area:350mm x 420mm Each | 100% | 2 samples in an original packaging
RT42-07-011 | Drape/dressing: Sterile. Self adhesive, transparent, latex free. Material to be an adhesive viscose, polyurethane or polyester or polyethylene film, occlusive to bacteria. The adhesive shall be a hypoallergenic acrylic copolymer. The film to be waterproof and have a siliconised release paper without a“window”frame, Individually packed in a peel pouch. Size:450mm x 280mm Adhesive area:450mm x 200mm Each | 100% | 2 samples in an original packaging
RT42-07-012 | Drape/dressing: Sterile. Self adhesive, transparent, latex free. Material to be an adhesive viscose, polyurethane or polyester or polyethylene film, occlusive to bacteria. The adhesive shall be a hypoallergenic acrylic copolymer. The film to be waterproof and have a siliconised release paper without a“window”frame, Individually packed in a peel pouch. Size:450mm x 550mm Adhesive area:450mm x 480mm Each | 100% | 2 samples in an original packaging
RT42-07-013 | Drape/dressing: Sterile. Self adhesive, transparent, latex free. Material to be an adhesive viscose, polyurethane or polyester or polyethylene film, occlusive to bacteria. The adhesive shall be a hypoallergenic acrylic copolymer. The film to be waterproof and have a siliconised release paper without a“window”frame, Individually packed in a peel pouch. Size:560mm x 840mm Adhesive area:560mm x 760mm Each | 100% | 2 samples in an original packaging
RT42-07-014 | Drape, surgical incise, with iodine incorporated into the adhesive. Barrier material to be: Tear resistant, able to stretch and regain in any direction, transparent, able to conform to irregular surfaces, matt finish, antistatic, permeable to oxygen and water vapour, waterproof and antistatic Adhesive: Non irritant, non-allergenic, to adhere firmly during prolonged operations also along cut edges, covered with silicone treated paper Handles: Backing firm enough to guide adhesive area into proper positioning Sterile, individually peel packed with description on outer pack Size: 600mm x 350mm Adhesive area: 350mm x 350mm: Each | 100% | 2 samples in an original packaging
RT42-07-015 | Drape, surgical incise, with iodine incorporated into the adhesive. Barrier material to be: Tear resistant, able to stretch and regain in any direction, transparent, able to conform to irregular surfaces, matt finish, antistatic, permeable to oxygen and water vapour, waterproof and antistatic Adhesive: Non irritant, non-allergenic, to adhere firmly during prolonged operations also along cut edges, covered with silicone treated paper Handles: Backing firm enough to guide adhesive area into proper positioning Sterile, individually peel packed with description on outer pack Size: 900mm x 450mm Adhesive area: 350mm x 350mm Each | 100% | 2 samples in an original packaging
RT42-07-016 | Drape, surgical incise, with iodine incorporated into the adhesive. Barrier material to be: Tear resistant, able to stretch and regain in any direction, transparent, able to conform to irregular surfaces, matt finish, antistatic, permeable to oxygen and water vapour, waterproof and antistatic Adhesive: Non irritant, non-allergenic, to adhere firmly during prolonged operations also along cut edges, covered with silicone treated paper Handles: Backing firm enough to guide adhesive area into proper positioning Sterile, individually peel packed with description on outer pack Size: 400mm x 400mm Adhesive area: 340mm x 400mm Each | 100% | 2 samples in an original packaging
RT42-07-017 | Drape, surgical incise, with iodine incorporated into the adhesive. Barrier material to be: Tear resistant, able to stretch and regain in any direction, transparent, able to conform to irregular surfaces, matt finish, antistatic, permeable to oxygen and water vapour, waterproof and antistatic Adhesive: Non irritant, non-allergenic, to adhere firmly during prolonged operations also along cut edges, covered with silicone treated paper Handles: Backing firm enough to guide adhesive area into proper positioning Sterile, individually peel packed with description on outer pack Size: 700mm x 450mm Adhesive area: 600mm x 450mm Each | 100% | 2 samples in an original packaging
RT42-07-018 | Drape, surgical incise, with iodine incorporated into the adhesive. Barrier material to be: Tear resistant, able to stretch and regain in any direction, transparent, able to conform to irregular surfaces, matt finish, antistatic, permeable to oxygen and water vapour, waterproof and antistatic Adhesive: Non irritant, non-allergenic, to adhere firmly during prolonged operations also along cut edges, covered with silicone treated paper Handles: Backing firm enough to guide adhesive area into proper positioning Sterile, individually peel packed with description on outer pack Size: 900mm x 600mm Adhesive area: 600mm x 600mm Each | 100% | 2 samples in an original packaging
RT42-07-019 | Drape, surgical incise, with iodine incorporated into the adhesive. Barrier material to be: Tear resistant, able to stretch and regain in any direction, transparent, able to conform to irregular surfaces, matt finish, antistatic, permeable to oxygen and water vapour, waterproof and antistatic Adhesive: Non irritant, non-allergenic, to adhere firmly during prolonged operations also along cut edges, covered with silicone treated paper Handles: Backing firm enough to guide adhesive area into proper positioning Sterile, individually peel packed with description on outer pack Size: 900mm x 850mm Adhesive area: 600mm x 850mm Each | 100% | 2 samples in an original packaging
RT42-07-020 | Drape craniotomy, with povidone iodine incise film. Size:1830mm x 1830mm With adhesive strip, 50mm wide with fluid collection pouch and drainage tap. Each | Not applicable | 2 samples in an original packaging
RT42-07-021 | Drape, disposable, craniotomy, reinforced, absorbent material around fenestration filtering into a fluid collection pouch, integrated adhesive incise, wire attached to pouch to keep shape and allow for effective fluid control, filter inside pouch to catch bone chips and minimize blockage, drainage port attached to pouch, gutters of excess fluid control, full patient drape Fluid collection pouch: 600mm x 600mm; Reinforced area: 1600mm x 2000mm; Fenestration size: 250mm in diameter; Overall size: 230cm (width) x 320cm (length) Each | 100% | 2 samples in an original packaging
RT42-07-022 | Drape craniotomy, incise film. Size:550mm x 550mm overall size 300cm x 250cm With adhesive strip, 50mm wide with fluid collection pouch and drainage tap. Each | 100% | 2 samples in an original packaging
RT42-07-023 | Drape, cardiovascular bilateral split sheet with incise film, Size: 230mm x 320mm overall size 225cm x 300cm Each | 100% | 2 samples in an original packaging
RT42-07-024 | Drape, cardiovascular bilateral split sheet with incise film, Size: 368mm x 214mm overall size 200cm x 300cm Each | 100% | 2 samples in an original packaging
RT42-07-025 | Drape craniotomy, with povidone iodine incise film. Incise area:24cm(L) x 37cm(W) overall size 200cm (W) x 400cm (L) With adhesive strip, 50mm wide with fluid collection pouch and drainage tap. Each | Not applicable | 2 samples in an original packaging
RT42-07-026 | Drape surgical, under buttock with drainage pouch. Size: 850mm x 1050mm. For obstetric and gynecological procedures. Drainage pouch fitted with drainage tap and pliable to secure for fluid control. Materials must be waterproof. Sterile, individual peelpacked with description on outer pack. Each | 100% | 2 samples in an original packaging
RT42-07-027 | Drape disposable surgical, under buttocks pouch, with hand cuffs to position the drape in a sterile setting. 2 towels (800mm x1200mm) materila treated to be hydrophobic, soft and comfortable to drape over patients legs. 1 absorbent sheet (800mm x 1000mm) material highly absorbent and laminated to insure impervious nature. Sterile, peel packed with description on outer pack. Each | 100% | 2 samples in an original packaging
RT42-07-028 | Drape, disposable, C-section with clear incise pouch and drainage outlet. Size: 1960mm x 3100mm Adhesiv: 300mm x 300mm 1X dry wipe ( wipe paper toweling) 1x mayo table cover 75cmx140cm (blue plastic with laminated reinforcing) 1xtrolley cover 160cmx205cm with leminated reinforcing 60cmx80cm 1x baby towel All towel drapes are reinforced with laminated spunbond and be absorbable around the critical area Sterile, peelpacked with description on outer pack. Pack | 100% | 2 samples in an original packaging
RT42-07-029 | Drape, disposable, C-section with clear incise pouch and drainage outlet.Size: 1960mm x 3100mm Adhesiv: 300mm x 300mm Sterile, individual peelpacked with description on outer pack. Each | 100% | 2 samples in an original packaging
RT42-07-030 | Drape, knee arthroscopy bilateral, with a fluid collection pouch and drainable tap, perforated at the end so the surgeon can tie the drape around his back if he prefers, Overall size: 2290mm (width) x 3200mm (length) with fenestration 100mm x 100mm Barrier material to be: Tear resistant, able to stretch and regain in any direction transparent, able to conform to irregular surfaces, matt finish, antistatic, permeable to oxygen and water vapour, waterproof and antistatic. Critical area to be absorbable Sterile, individual peelpacked with description on outer pack. Each | 100% | 2 samples in an original packaging
RT42-07-031 | Drape, disposable, knee, with a fluid pouch, pouch has a wire strip attached to keep its shape so as to control fluid effectively, is reinforce area under the leg to ensure no risk for strikethrough, the drape is perforated at the end so the surgeon can tie the drape around his back if he prefers, Aperture size: 80mm Overall size: 2290mm (width) x 3200mm (length) Critical area to be absorbable Sterile, individual peelpacked with description on outer pack. Each | 100% | 2 samples in an original packaging
RT42-07-032 | Drape, disposable, shoulder, with a fluid pouch, 3 adhesive strips, to be use on both shoulders, the pouch has a wire strip attached to keep its shape so as to control fluid effectively, Velcro cable holders attached on either side of the pouch for effective cable management Adhesive strip size: 100mm x 500mm Overall size: 220cm (width) x 410cm (length) Critical area to be absorbable Each | 100% | 2 samples in an original packaging
RT42-07-033 | Drape, disposable, ophthalmic, double pouch with wire strips to keep them open and handle fluid effectively, additional wire strip place over nose area so as to lift material off patient’s nose thus allowing for easy breathing Adhesive size: 125mm x 125mm Overall size: 150cm (width) x 275cm (length) Sterile, Individually peel packed with description on outer pack Critical area to be absorbable Each | 100% | 2 samples in an original packaging
RT42-07-034 | Drape, hip sheet: Size: 2200mm x 3200mm with two dislocation pouches. Critical area to be absorbable Each | 100% | 2 samples in an original packaging
RT42-07-035 | Drape, disposable, spinal, fenestrated, with loose long instrument pouch, dry wipe, 2 x integrated cable holders, 3 layers around critical area – impervious and hydrophilic Instrument pouch size: 260mm x 480mm Adhesive incise: 295mm x 395mm Overall size: 210cm (width) x 300cm (length) Critical area to be absorbable Each | 100% | 2 samples in an original packaging
RT42-07-036 | Drape, disposable, surgical, X-Ray, Lap incise sheet, non-woven, water repellent, low linting, pore size allowing air but preventing bacteria passage through fabric, integrated cord holders, separated instrument pouch, 2 pouches to be indicated Overall size: 2100mm (width) x 3200mm (length) Incise adhesive film fenestration: size 270mm x 270mm Sterile Individually peel packed with description on outer pack Critical area to be absorbable Each | 100% | 2 samples in an original packaging
RT42-07-037 | Drape, disposable, surgical, X-Ray, C-arm, 4 adhesive patches, 2 plastic strips for securing drape Overall size: 1650mm (width) x 1060mm (length) Each | 100% | 2 samples in an original packaging
RT42-07-038 | Drape, extremity. Size: 2280mm (width) x 3000mm (length), Flexable aperture 60mm x 80mm Each | 100% | 2 samples in an original packaging
RT42-07-039 | Drape, extremity sheet with fluid collction pouch with tiebacks. Size: 2280mm (width) x 3040mm (length) Each | 100% | 2 samples in an original packaging
RT42-07-040 | Drape, disposable, sterile, sheet, extremity, water repellant, non-woven fabric, that allows air but prevent bacteria passage through fabric with hinge and cord holders Reinforce at key points,critical area to be absorbable Expandable fenestration Overall size: 1830mm (width) x 3020mm (length) Sterile Individually peel packed with description on outer pack Each | 100% | 2 samples in an original packaging
RT42-07-041 | Drape, disposable, sheet, spunlace, water repellent, non-woven fabric, low linting, pore size that allows air but prevent bacteria passage through the fabric, hypo-allergenic and latex free adhesive, positioned centrally on the edge of long side, sterile, description on outer pack, Adhesive size: 55mm x 1000mm Drape size: 1000mm (width) x 1700mm (length) Pack size: Packet of 5 sheets critical area to be absorbable Pack | Not applicable | 2 samples in an original packaging
RT42-07-042 | Drape, disposable, for minor procedures, with positioning tape and integrated cable holder, 3-layers around critical area – impervious and hydrophilic adhesive must be latex free, ideal for multiple procedures Aperture size: 80mm x 80mm Fenestration: 150mm x 150mm Overall size: 2000mm (width) x 3200mm (length critical area to be absorbable Each | 100% | 2 samples in an original packaging
RT42-07-043 | Drape, disposable, basic pack, consisting of the following: 1 x large drape towel: 160cm x 220cm with adhesive tape 1 x medium drape towel: 160cm x 180cm with adhesive tape 2 x small drape sheet: 90cm x 80cm with adhesive tape 1 x dry wipe (white paper toweling) 1 x Mayo table cover 75cm x 140cm (blue plastic with laminated reinforcing) 1 x trolley cover: 160cm x 205cm with laminated reinforcing 60cm x 80cm 1 x suture bag (clear plastic with adhesive tape) 1 x luxfab: 90cm x 90cm (blue crepe paper) 1 x adhesive strip with adhesive backing All towels drapes are reinforced / absorbable with laminated spunbond around the critical area Pack | 100% | 2 samples in an original packaging
RT42-07-044 | Drape, disposable, angiography, consisting of the following: 1 x angiography drape: 230cm x 320cm absorbent, adhesive incise, LDPE 1 x back table cover: 160cm x 210cm reinforced 1 x dry wipe (white paper towel) 2 x cathlab cover small round (clear LDPE plastic) 1 x cathlab cover: 105cm x 75cm (clear LDPE plastic) Full patient drape, blue absorbent material around femoral area with integrated gutter, adhesive incise for both legs Drape trimmed with clear plastic to cover control panel alongside bed disposable, for minor procedures, with positioning tape and integrated cable holder, 3-layers around critical area – impervious and hydrophilic adhesive must be latex free, ideal for multiple procedures Aperture size: 80mm x 80mm Fenestration: 150mm x 150mm Overall size: 2000mm (width) x 3200mm (length critical area to be absorbable Pack | 100% | 2 samples in an original packaging
RT42-07-045 | Drape, disposable, angiography, Brachial: Size 1120mm x 1450mm one central aperture, surrounded by adhesive strip. Barrier material to be: tear resistant, able to stretch and regain in any direction, transparent able to conform to irregular surfaces, matt finish, antistatic, impermeable to bacteria permeable to oxygen and water vapour, waterproof. Adhesive: low irritant, hypo-allergenic, to adhere firmly during prolonged operations, also on cut edges covered with silicone treated paper Handles: backing firm enough to guide adhesive area into proper positioning Sterile, individually peel packed with description on outer pack Critical area to be absorbable Each | 100% | 2 samples in an original packaging
RT42-07-046 | Drape, disposable, angiography, Femoral: Size 3250mm x 2030mm Two self adhesive apertures 60mm x 90mm. Barrier material to be: tear resistant, able to stretch and regain in any direction, transparent able to conform to irregular surfaces, matt finish, antistatic, impermeable to bacteria permeable to oxygen and water vapour, waterproof. Adhesive: low irritant, hypo-allergenic, to adhere firmly during prolonged operations, also on cut edges covered with silicone treated paper Handles: backing firm enough to guide adhesive area into proper positioning Sterile, individually peel packed with description on outer pack Critical area to be absorbable Each | 100% | 2 samples in an original packaging
RT42-07-047 | Drape, disposable, utility set, consisting of 4 towel drapes 42cm x 30cm with adhesive tape Barrier material must be: • Fluid repellant • Impermeable to bacteria • Permeable to oxygen and water vapor Sterile Individually peel packed Set | 100% | 2 samples in an original packaging
RT42-07-048 | Drape disposable surgical, under buttocks, with fluid control pouch 1 x Wrap Drape: Absorbent Reinforcement: 178 mm x 508mm and 254mm x 508mm Overall size: 940mm x 1040mm Sterile, peel packed with description on outer pack. Critical area to be absorbable Each | 100% | 2 samples in an original packaging
RT42-07-049 | Drape disposable surgical, under buttocks, with fluid control pouch 1 x Wrap Drape: Adhesive Overall size: 990mm x 1120mm Sterile, peel packed with description on outer pack. Critical area to be absorbable Each | 100% | 2 samples in an original packaging
RT42-07-050 | Drape, disposable.Laparoscopic Assisted Vaginal Hysterectomy with fluid control pouch, Tube holders Absorbent Reinforcements Instrument Pad: 254mm x 445mm overall size to be 200ccm x 220cm Sterile, peel packed with description on outer pack. Each | 100% | 2 samples in an original packaging
RT42-07-051 | Drape disposable surgical, under buttocks 1 x Wrap Drape: Under buttock drape with Absorbent Reinforcement Overall size: 890 x 1240mm Sterile, peel packed with description on outer pack. Each | 100% | 2 samples in an original packaging
RT42-07-052 | Drape, disposable, knee arthroscopy, with a fluid control pouch, tube holders 1 x Wrap Drape: Fenestration Absorbent Reinforcement size: 890mm x 1194mm Overall size: 2870mm (width) x 2920mm (length) Sterile, individual peelpacked with description on outer pack. Each | 100% | 2 samples in an original packaging
RT42-07-053 | Drape, disposable, surgical, Laparascopy/Cholecystectomy, 1 x Wrap Drape: Absorbent Reinforcement: 1041mm x 1422mm Instrument Pad: 254mm x 457mm Pouches Adhesive Overall size: 2690mm x 3100mm Sterile Individually peel packed with description on outer pack Each | 100% | 2 samples in an original packaging
RT42-07-054 | Drape, disposable, surgical, Laparascopy T-Drape with armboard covers, 1 x Wrap Absorbent Reinforcement Fenestration: 152mm x 203mm Fenestration: 76mm x 152mm Pouches Tube Holders Adhesive Overall size: 2460mm x 2590mm Sterile Individually peel packed with description on outer pack Each | 100% | 2 samples in an original packaging
RT42-07-055 | Drape, disposable, shoulder, with a fluid control pouch and tube holders 1 x Wrap Drape: Fenestration Pear Shaped Poly Reinforcement Overall size: 1570mm (width) x 2540 (length) Sterile, Individually peel packed with description on outer pack Each | 100% | 2 samples in an original packaging
RT42-07-056 | Drape, disposable, shoulder split, with incise and tube holders 1 x Wrap Drape: Pear Shaped Split with incise Poly Reinforcement Overall size: 1600mm (width) x 2540mm (length) Sterile, Individually peel packed with description on outer pack Each | 100% | 2 samples in an original packaging
RT42-07-057 | Drape, disposable, shoulder split with fluid pouch 1 x Wrap Drape: Split size: 152mm x 940 Fluid Control Pouch with Malleable Wire Overall size: 2290mm (width) x 2540mm (length) Sterile, Individually peel packed with description on outer pack Each | 100% | 2 samples in an original packaging
RT42-07-058 | Drape, disposable, beach chair shoulder with fluid control pouch 1 x Wrap Drape: Fenestration Pear Shaped Poly Reinforcement Velcro Tabs Fluid Control Pouch Overall size: 2590mm (width) x 4110mm (length) Sterile, Individually peel packed with description on outer pack EacH | 100% | 2 samples in an original packaging
RT42-07-059 | Drape, disposable, ophthalmic, incise with fluid control pouches, 1 x Wrap . Drape: Icise size: 102mm x 114mm Overall size: 1780mm (width) x 2540(length) Sterile, Individually peel packed with description on outer pack Each | 100% | 2 samples in an original packaging
RT42-07-060 | Drape, disposable, ophthalmic,with fluid collection saddlebag pouch and 1 x Wrap Drape: Tube holders. Fenestration size: 102mm x 127mm Absorbent Reinforcement size: 279mm x 457mm Overall size: 1220mm (width) x 1700(length) Sterile, Individually peel packed with description on outer pack Each | 100% | 2 samples in an original packaging
RT42-07-061 | Drape, disposable, hip 1 x Wrap Drape: Tube holders. Teardrop Fenestration Absorbent Reinforcement Overall size: 2950mm (width) x 3330(length) Sterile, Individually peel packed with description on outer pack Each | 100% | 2 samples in an original packaging
RT42-07-062 | Drape, disposable, hip with fluid control pouch 1 x Wrap Drape: Tube holders. Circular Fenestration Absorbent Reinforcement: 558mm x 1372mm Overall size: 2820mm (width) x 3380(length) Sterile, Individually peel packed with description on outer pack Each | 100% | 2 samples in an original packaging
RT42-07-063 | Drape, disposable, angiography: Size 1910mm x 3180mm 1 x Wrap 1 x Heavy Duty Absorbent Pad 914mm x 3175mm Circular Fenestration Each | 100% | 2 samples in an original packaging
RT42-07-064 | Drape, disposable camera cover size 1400mm (width) x 2500mm (length), drape camera/arm drape adhesive poly strap to hold in place adhesive strip at base clear poly film sterile. individually peel packed, description on outer pack. Sterile, Individually peel packed with description on outer pack Each | 100% | 2 samples in an original packaging
RT42-07-065 | Assessment services for the supply, installation and maintenance of hospital curtain rails CURTAIN RAIL Supply and installation of curtain rail: • Conduct thorough curtain rail assessment as required • Submit assessment report outlining areas for the need for installation of Hospital curtain rails or repair existing curtain rails or utilize existing curtain rails • Curtain rail must be made of Aluminium alloy, extruded with natural anodised finish • All components must be supplied and fitted • Provide maintenance plan for the duration of the contract • Size: 1m width curtain rail Price must be inclusive of supply, installation and maintanance of curtain rail for the duration of the contract. | 100% | 1 sample in an original packaging
RT42-07-066 | Assessment services for the supply, installation and maintenance of hospital and antimicrobial and sporicidal disposable curtains CURTAIN Supply and installation of Antimicrobial and Sporicidal Curtain: • Disposable curtain must have Antimicrobial and Sporicidal properties • Disposable curtains must not have silver as a properties • Made from polypropylene • Be fire retardant • Lightweight • Replaced every 2 years for the duration of the contract • Tested on installation against spores, bacteria and fungi • Traps and kills pathogens (bacteria, spores, fungi and mycobacteria) • Effective at ambient temperature • 100% antimicrobial coverage on polypropylene curtain • Provide in-service training • Size: 1m width X 2m length Price must be inclusive of supply, installation and maintenance of curtain for the duration of the contract. | 100% | 1 sample in an original packaging
RT42-07-067 | Assessment services for the supply, installation and maintenance of hospital and antimicrobial and sporicidal disposable curtains CURTAIN Supply and installation of Antimicrobial and Sporicidal Curtain: • Disposable curtain must have Antimicrobial and Sporicidal properties • Disposable curtains must not have silver as a properties • Made from polypropylene • Be fire retardant • Lightweight • Replaced every 2 years for the duration of the contract • Tested on installation against spores, bacteria and fungi • Traps and kills pathogens (bacteria, spores, fungi and mycobacteria) • Effective at ambient temperature • 100% antimicrobial coverage on polypropylene curtain • Provide in-service training • Size: 1m width X 2.5 m length Price must be inclusive of supply, installation and maintenance of curtain for the duration of the contract. | 100% | 1 sample in an original packaging
RT42-07-068 | Pillow cover, disposable, 50gsm spunbonded (hydrophobic), water repellent, non-woven fabric, low linting, pore size that allows air but prevent bacteria passage through the fabric, hypo-allergenic and latex free, individually packed non-sterile. Size: 80 x 55cm Mass to comply with SANS 79 Diimensions to comply with SANS 81 $ 83 Box of 50 pillow covers | Not applicable | 1 box in an original packaging
RT42-07-069 | Sheet, Flat, disposable, 25-40gsm spunbonded (hydrophobic), water repellent, non-woven fabric, low linting, pore size that allows air but prevent bacteria passage through the fabric, hypo-allergenic and latex free, with adhesive tabs, individually packed, Size: 200cm x 156cm Pack size: Each Mass to comply with SANS 79 Diimensions to comply with SANS 81 $ 83 Box of 50 Sheet,Flat | Not applicable | 1 box in an original packaging
RT42-07-070 | Sheet, Trolley disposable, 65gsm spunbonded (hydrophobic), water repellent, non-woven fabric, low linting, pore size that allows air but prevent bacteria passage through the fabric, hypo-allergenic and latex free, with adhesive tabs, individually packed, sterile Size: 100 x 240cm Mass to comply with SANS 79 Diimensions to comply with SANS 81 $ 83 Box of 50 Sheet, Trolley | Not applicable | 1 box in an original packaging
RT42-07-071 | Sheet, Fitted, disposable, 25-40gsm spunbonded (hydrophobic), water repellent, non-woven fabric, low linting, pore size that allows air but prevent bacteria passage through the fabric, hypo-allergenic and latex free, with adhesive tabs, individually packed, elastic sewn all around on fitted sheets allows covers to fit various widths / lengths, single bed Size: 210cm x90cm x 20cm Mass to comply with SANS 79 Diimensions to comply with SANS 81 $ 83 Box of 50 Sheet, Fitted | Not applicable | 1 box in an original packaging
RT42-07-072 | Duvet cover, Disposable,Big enough to cover normal single bed blanket Various widths / lengths, single bed Size: 170cm x 200cm Box of 50 Duvet cober | Not applicable | 1 box in an original packaging
RT42-08-001 | Tape, surgical, adhesive film, microporous paper consisting of non-woven synthetic fibres spread with a non-occlusive acrylic adhesive Surgical tape must: permit skin ventilation, be hypo-allergenic, be tolerated by skin for prolonged periods, not irritate the skin Size: 12 mm x 5 m Individually wrapped Roll | Not applicable | 3 samples in an original packaging
RT42-08-002 | Tape, surgical, adhesive film, microporous paper consisting of non-woven synthetic fibres spread with a non-occlusive acrylic adhesive Surgical tape must: permit skin ventilation, be hypo-allergenic, be tolerated by skin for prolonged periods, not irritate the skin Size: 25mm x 5 m Individually wrapped Roll | Not applicable | 3 samples in an original packaging
RT42-08-003 | Tape, surgical, adhesive film, microporous paper consisting of non-woven synthetic fibres spread with a non-occlusive acrylic adhesive Surgical tape must: permit skin ventilation, be hypo-allergenic, be tolerated by skin for prolonged periods, not irritate the skin Size: 50 mm x 5 m Individually wrapped Roll | Not applicable | 3 samples in an original packaging
RT42-08-004 | Tape, surgical, adhesive film, microporous paper consisting of non-woven synthetic fibres spread with a non-occlusive acrylic adhesive Surgical tape must: permit skin ventilation, be hypo-allergenic, be tolerated by skin for prolonged periods, not irritate the skin Size: 75 mm x 5 m Individually wrapped Roll | Not applicable | 3 samples in an original packaging
RT42-08-005 | Tape, surgical, adhesive film, synthetic must be hypo-allergenic, permeable and water repellent Size: 25 mm x 5 m Individually wrapped Roll | Not applicable | 3 samples in an original packaging
RT42-08-006 | Tape, surgical, adhesive film, synthetic must be hypo-allergenic, permeable and water repellent Size: 50 mm x 5 m Individually wrapped Roll | Not applicable | 3 samples in an original packaging
RT42-08-007 | Tape, surgical, adhesive film, synthetic must be hypo-allergenic, permeable and water repellent Size: 75 mm x 5 m Individually wrapped Roll | Not applicable | 3 samples in an original packaging
RT42-08-008 | Tape, surgical, adhesive film, synthetic must be hypo-allergenic, permeable and water repellent Size: 100 mm x 5 m Individually wrapped Roll | Not applicable | 3 samples in an original packaging
RT42-08-009 | Tape, zinc oxide, surgical, adhesive, B.P., woven cotton cloth, non-stretch Size: 25 mm x 5 m Individually wrapped Roll | Not applicable | 3 samples in an original packaging
RT42-08-010 | Tape, zinc oxide, surgical, adhesive, B.P., woven cotton cloth, non-stretch Size: 50 mm x 5 m Individually wrapped Roll | Not applicable | 3 samples in an original packaging
RT42-08-011 | Tape, zinc oxide, surgical, adhesive, B.P., woven cotton cloth, non-stretch Size: 75 mm x 5 m Individually wrapped Roll | Not applicable | 3 samples in an original packaging
RT42-08-012 | Tape neuro-proprioceptive, elastic, adhesive Size: 25mm x 5m • Latex Free. • Have a wavelike pattern of adhesive coating and be extensible. • Have a release paper lamination • Have a polyacrylate adhesive with a cotton substrate. • Variety of colours Pack size: Roll | Not applicable | 3 samples in an original packaging
RT42-08-013 | Tape neuro-proprioceptive, elastic, adhesive Size: 50mm x 5m • Latex Free. • Have a wavelike pattern of adhesive coating and be extensible. • Have a release paper lamination • Have a polyacrylate adhesive with a cotton substrate. • Variety of colours Pack size: Roll | Not applicable | 3 samples in an original packaging
RT42-08-014 | Tape neuro-proprioceptive, elastic, adhesive Size: 75mm x 5m • Latex Free. • Have a wavelike pattern of adhesive coating and be extensible. • Have a release paper lamination • Have a polyacrylate adhesive with a cotton substrate. • Variety of colours Pack size: Roll | Not applicable | 3 samples in an original packaging
RT42-08-015 | Tape, strapping • Highly adhesive, • Rigid support • Consisting of 100% rayon ,substrate that has been spread evenly with an adhesive mass containing zinc oxide. • Must have very high adhesive power, • Very high tensile strength, non-elastic with a water-repellent backing, • Providing secure fixation Size: 40cm x 14m Pack size: Roll | 100% | 3 samples in an original packaging
RT42-08-016 | Tape, retention, design to secure dressings and bandage to the body, use for wide area fixation, tape is constracted from non-woven polyester which is coated with polyacrylate adhesive and must have a protective laminated paper unsterile. Size: 25mm x 10m Each | Not applicable | 3 samples in an original packaging
RT42-08-017 | Tape, retention, design to secure dressings and bandage to the body, use for wide area fixation, tape is constracted from non-woven polyester which is coated with polyacrylate adhesive and must have a protective laminated paper unsterile. Size: 50mm x 10m Each | Not applicable | 3 samples in an original packaging
RT42-08-018 | Tape, retention, design to secure dressings and bandage to the body, use for wide area fixation, tape is constracted from non-woven polyester which is coated with polyacrylate adhesive and must have a protective laminated paper unsterile. Size: 100mm x 2m. Each | Not applicable | 3 samples in an original packaging
RT42-09-001 | Brain patties with radio opaque string, latex free, surgical, highly absorbent, X-Ray detectable Size: 6mm x 6mm Sterile peel packed Pack of 5 | Not applicable | 2 pack of samples in an original packaging
RT42-09-002 | Brain patties with radio opaque string, latex free, surgical, highly absorbent, X-Ray detectable Size: 2.5mm x2.5mm Sterile peel packed Pack of 5 | Not applicable | 2 pack of samples in an original packaging
RT42-09-003 | Brain patties with radio opaque string, latex free, surgical, highly absorbent, X-Ray detectable Size: 2mm x 2mm Sterile peel packed Pack of 5 | Not applicable | 2 pack of samples in an original packaging
RT42-09-004 | Brain patties with radio opaque string, latex free, surgical, highly absorbent, X-Ray detectable Size: 12mm x 12mm Sterile peel packed Pack of 5 | Not applicable | 2 pack of samples in an original packaging
RT42-09-005 | Brain patties with radio opaque string, latex free, surgical, highly absorbent, X-Ray detectable Size: 12mm x 50mm Sterile peel packed Pack of 5 | Not applicable | 2 pack of samples in an original packaging
RT42-09-006 | Brain patties with radio opaque string, latex free, surgical, highly absorbent, X-Ray detectable Size: 12mm x 75mm Sterile peel packed Pack of 5 | Not applicable | 2 pack of samples in an original packaging
RT42-09-007 | Brain patties with radio opaque string, latex free, surgical, highly absorbent, X-Ray detectable Size: 20mm x 20mm Sterile peel packed Pack of 5 | Not applicable | 2 pack of samples in an original packaging
RT42-10-001 | Sponge, surgical, polyvinyl alcohol type X-ray detectable Nasal tampon 45mm - double compressed Sterile peel packed Each | Not applicable | 3 samples in an original packaging
RT42-10-002 | Sponge, surgical, polyvinyl alcohol type X-ray detectable Nasal tampon 50mm - double compressed Sterile peel packed Each | Not applicable | 3 samples in an original packaging
RT42-10-003 | Sponge, surgical, polyvinyl alcohol type X-ray detectable Nasal tampon 75mm - double compressed Sterile peel packed Each | Not applicable | 3 samples in an original packaging
RT42-10-004 | Sponge, surgical, polyvinyl alcohol type X-ray detectable Nasal tampon 90mm - double compressed Sterile peel packed Each | Not applicable | 3 samples in an original packaging
RT42-10-005 | Spongette, triangular with pliable handle Consisting of highly asorbent, non- linting, fine grain compressed cellular material Sterile peel packed Packet of 5 spongettes | Not applicable | 3 samples in an original packaging
RT42-11-001 | Plaster, waterproof, self-adhesive, B.P. Rolled with adhesive inwards on metal or plastic spool, suitably protected by metal or plastic cover Size: 25 mm x 3 m Roll | 100% | 2 samples in an original packaging
RT42-11-002 | Plaster, waterproof, self-adhesive, B.P. Rolled with adhesive inwards on metal or plastic spool, suitably protected by metal or plastic cover Size: 50 mm x 3 m Roll | 100% | 2 samples in an original packaging
RT42-11-003 | Plaster, waterproof, self-adhesive, B.P. with non-stick wound padding (island) assorted Washable plasters Individually wrapped strips Packet of 20 plasters | Not applicable | 2 samples in an original packaging
RT42-12-001 | Skin closure strips, non-allergenic Strips shall have a high initial adherence for short term use up to 5 days Size: 3 mm x 75 mm Sterile peel packed Packaging: 5 Strips per envelope | 100% | 2 envelope samples in an original packaging
RT42-12-002 | Skin closure strips, non-allergenic Strips shall have a high initial adherence for short term use up to 5 days Size: 12 mm x 100 mm Sterile peel packed Packaging: 6 Strips per envelope | 100% | 2 envelope samples in an original packaging
RT42-13-001 | Support, intravenous, paediatric For paediatric use Size: 65 mm x 70mm Aluminium rod of 400 mm width encased in foam 3 mm thick covered with medical felt 2 Velcro strips on each end Each | Not applicable | 2 samples in an original packaging
RT42-13-002 | Support, intravenous, paediatric For paediatric use Size: 85 mm x 115mm Aluminium rod of 400 mm width encased in foam 3 mm thick covered with medical felt 2 Velcro strips on each end Each | Not applicable | 2 samples in an original packaging
RT42-13-003 | Support, intravenous, paediatric For paediatric use Size: 100 mm x 120 mm Aluminium rod of 400 mm width encased in foam 3 mm thick covered with medical felt 2 Velcro strips on each end Each | Not applicable | 2 samples in an original packaging
RT42-13-004 | Support, intravenous, paediatric For paediatric use Size: 150 mm x 50 mm Aluminium rod of 400 mm width encased in foam 3 mm thick covered with medical felt 2 Velcro strips on each end Each | Not applicable | 2 samples in an original packaging
RT42-13-005 | Support, intravenous,neonatal For neonatal use Size:50 mm x 57 mm Aluminium rod of 400 mm width encased in foam 3 mm thick covered with medical felt 2 Velcro strips on each end Each | Not applicable | 2 samples in an original packaging
RT42-14-001 | Padding, natural, orthopaedic, bonded type, autoclavable, type 2 Size: 50 mm x 3 m To comply with the latest issue of SANS 196 Compliance certificate to be submitted Roll | 100% | 2 samples in an original packaging
RT42-14-002 | Padding, natural, orthopaedic, bonded type, autoclavable, type 2 Size: 75 mm x 3 m To comply with the latest issue of SANS 196 Compliance certificate to be submitted Roll | 100% | 2 samples in an original packaging
RT42-14-003 | Padding, natural, orthopaedic, bonded type, autoclavable, type 2 Size: 100 mm x 3 m To comply with the latest issue of SANS 196 Compliance certificate to be submitted Roll | 100% | 2 samples in an original packaging
RT42-14-004 | Padding, natural, orthopaedic, bonded type, autoclavable, type 2 Size: 150 mm x 3 m To comply with the latest issue of SANS 196 Compliance certificate to be submitted Roll | 100% | 2 samples in an original packaging
RT42-14-005 | Padding, natural, orthopaedic, bonded type, autoclavable, type 2 Size: 200mm x 3m To comply with the latest issue of SANS 196 Compliance certificate to be submitted Roll | 100% | 2 samples in an original packaging
RT42-14-006 | Padding, orthopaedic, synthetic cast, type 4 or type 5 Size: 50mm x 3m To comply with the latest issue of SANS 196 Compliance certificate to be submitted Roll | 100% | 2 samples in an original packaging
RT42-14-007 | Padding, orthopaedic, synthetic cast, type 4 or type 5 Size: 75mm x 3m To comply with the latest issue of SANS 196 Compliance certificate to be submitted Roll | 100% | 2 samples in an original packaging
RT42-14-008 | Padding, orthopaedic, synthetic cast, type 4 or type 5 Size: 100mm x 3m To comply with the latest issue of SANS 196 Compliance certificate to be submitted Roll | 100% | 2 samples in an original packaging
RT42-14-009 | Padding, orthopaedic, synthetic cast, type 4 or type 5 Size: 150mm x 3m To comply with the latest issue of SANS 196 Compliance certificate to be submitted Roll | 100% | 2 samples in an original packaging
RT42-15-001 | Skin traction kit Size: Children Composed of extension plaster B.P. with a square plastic spreader in the centre on the adhesive side and covered with soft foam padding 10mm thick, the rest of the adhesive being covered by a flexible plastic protector. The spreader should have holes through which traction cord can be threaded. To be supplied with a cotton crepe bandage, type 1, to comply with the latest issue of SANS 945 Compliance certificate to be submitted Kit: Components with the following dimensions: Extension plaster: 1.8m long x 50 mm wide, Foam padding: 360mm long x ±50mm wide, Plastic spreader: 50mm square Traction cord: 1.65m long Cotton crepe bandage: not less than 4.5m x 50mm wide Packed in suitable labelled boxes Kit (Each) | 100% | 1 sample in an original packaging
RT42-15-002 | Skin traction kit Size: Adult Composed of extension plaster B.P with a square plastic spreader in the centre on adhesive side and covered with soft foam padding 10mm thick, the rest of the adhesive being covered by a flexible plastic protector. The spreader should have holes through which traction cord can be threaded. To be supplied with a cotton crepe bandage, type 1, to comply with the latest issue of SANS 945 Compliance certificate to be submitted Kit: Components with the following dimensions: Extension plaster: 1.8m long x 75 mm wide, Foam padding: 1.8 m long x 90 mm wide Plastic spreader: 75mm square Traction cord: 1.65m long Cotton crepe bandage: not less than 4.5m x 75mm wide Packed in suitable labelled boxes Each (Kit) | 100% | 1 sample in an original packaging
RT42-16-001 | Support, scrotal, suspensory Elastic waistband and legbands Size: Medium Individually packed Each | Not applicable | 2 sample in an original packaging
RT42-16-002 | Support, scrotal, suspensory Elastic waistband and legbands Size: Large Individually packed Each | Not applicable | 2 sample in an original packaging
RT42-17-001 | Soft silicone wound contract layer, atraumatic removal, adherent to dry skin. Size: 50mm x 70mm. Sterile, Individually peel packed, Each | Not applicable | 3 sample in an original packaging
RT42-17-002 | Soft silicone wound contract layer, atraumatic removal, adherent to dry skin. Size: 80mm x 100mm,Sterile, Individually peel packed, Each | Not applicable | 3 sample in an original packaging
RT42-17-003 | Soft silicone wound contract layer, atraumatic removal, adherent to dry skin. Size: 100mm x 180mm, Sterile, Individually peel packed, Each | Not applicable | 3 sample in an original packaging
RT42-17-004 | Soft silicone wound contract layer, atraumatic removal, adherent to dry skin. Size: 180mm x 200mm, Sterile, Individually peel packed, Each | Not applicable | 3 sample in an original packaging
RT42-17-005 | Soft silicone wound contract layer, atraumatic removal, adherent to dry skin. Size: 200mm x 310mm, Sterile, Individually peel packed, Each | Not applicable | 3 sample in an original packaging
RT42-17-006 | Soft silicone wound contract layer, atraumatic removal, adherent to dry skin. Size: 350mm x 600mm, Sterile, Individually peel packed, Each | Not applicable | 3 sample in an original packaging
RT42-18-001 | Silicone Non-Adhesive Foam Dressing. Size: 50mm x 70mm, Sterile, Individually peel packed, Each | Not applicable | 3 sample in an original packaging
RT42-18-002 | Silicone Non-Adhesive Foam Dressing. Size: 100mm x 100mm , Sterile, Individually peel packed, Each | Not applicable | 3 sample in an original packaging
RT42-18-003 | Silicone Non-Adhesive Foam Dressing. Size:100mm x 200mm, Sterile, Individually peel packed, Each | Not applicable | 3 sample in an original packaging
RT42-18-004 | Silicone Non-Adhesive Foam Dressing. Size: 150mm x 150mm, Sterile, Individually peel packed, Each | Not applicable | 3 sample in an original packaging
RT42-18-005 | Silicone Non-Adhesive Foam Dressing. Size: 200mm x 200mm, Sterile, Individually peel packed, Each | Not applicable | 3 sample in an original packaging
RT42-18-006 | Silicone Non-Adhesive Foam Dressing. Heel: 100mm x 130mm, Sterile, Individually peel packed, Each | Not applicable | 3 sample in an original packaging
RT42-18-007 | Silicone Adhesive Foam Dressing (Border), Size: 75mm x 75mm,Sterile, Individually peel packed, Each | Not applicable | 3 sample in an original packaging
RT42-18-008 | Silicone Adhesive Foam Dressing (Border), Size: 125mm x125mm, Sterile, Individually peel packed, Each | Not applicable | 3 sample in an original packaging
RT42-18-009 | Silicone Adhesive Foam Dressing (Border), Size: 225mm x 225mm, Sterile, Individually peel packed, Each | Not applicable | 3 sample in an original packaging
RT42-18-010 | Silicone Adhesive Foam Dressing (Border), Sacral 230mm x 230mm, Sterile, Individually peel packed, Each | Not applicable | 3 sample in an original packaging
RT42-18-011 | Silicone Adhesive Foam Dressing (Border), Size: 100mm x 300mm, Sterile, Individually peel packed, Each | Not applicable | 3 sample in an original packaging
RT42-18-012 | Silicone Adhesive Foam Dressing (Border), Size: 100mm x 200mm, Sterile, Individually peel packed, Each | Not applicable | 3 sample in an original packaging
RT42-19-001 | Casting Synthetic Splint,Size: 100mm x 400mm consist of multiple layers of knitted fabric, impregnated with moisture-curable resin, must have double-sided padding, exposure of splint to moisture of water initiates a chemical reaction, which causes the splint to become rigid, highly comfortable, water -repellent, lightweight , portable, setting time ± 4 min, functional strength ± 20 min, individually wrapped in a sealed foil pouch. Each | Not applicable | 2 sample in an original packaging
RT42-19-002 | Casting Synthetic Splint,Size: with allowance of 5mm,100mm x, 760mm consist of multiple layers of knitted fabric, impregnated with moisture-curable resin, must have double-sided padding, exposure of splint to moisture of water initiates a chemical reaction, which causes the splint to become rigid, highly comfortable, water -repellent, lightweight , portable, setting time ± 4 min, functional strength ± 20 min, individually wrapped in a sealed foil pouch. Each | Not applicable | 2 sample in an original packaging
RT42-19-003 | Casting Synthetic Splint,Size: with allowance of 5mm,125mm x, 760mm consist of multiple layers of knitted fabric, impregnated with moisture-curable resin, must have double-sided padding, exposure of splint to moisture of water initiates a chemical reaction, which causes the splint to become rigid, highly comfortable, water -repellent, lightweight , portable, setting time ± 4 min, functional strength ± 20 min, individually wrapped in a sealed foil pouch. Each | Not applicable | 2 sample in an original packaging
RT42-20-001 | Wipes, to remove adhesiveness from the skin. The wipe must also moisturizes and conditions the skin without greasiness. Gently cleans hydrocolloid, acrylic and rubber-based adhesive residues from the skin without irritation, discomfort or trauma. Applies easily and removes quickly with less need for reapplication. Non-irritating and non-sensitizing, even for ostomies and geriatric patients. Can be used anywhere there's a need for an adhesive dressing. Individually sterile wrapped Single use Size of wipe: 400mm x 400mm Box of 100 wipes | Not applicable | 2 box of samples in an original packaging
RT42-20-002 | Wipes, to protect the skin, is impregnated with a liquid film-forming dressing, with application to intact skin, a protective film is formed to help reduce friction during removal of tapes and dressings. Must be able to be used to prepare skin attachment sites for drainage tubes, external catheters, surrounding ostomy sites and adhesive dressings. Helps with tape and film. Removes easily when using skin cleanser or soap and water, so unsanitary residue is not left in or around the wound area. Allows skin to ""breathe"" so tapes and films adhere better. Must be tested to be use also on paediatric patients. Individually sterile wrapped Single use Size of wipe: 400mmx400mm Box of 100 wipes | Not applicable | 2 box of samples in an original packaging
RT42-20-003 | Adhesive removal liquid, to remove adhesives from the skin without irritation, discomfort or trauma. Liquid must be an alcohol based solvent. Applied solvent over tapes for easy and painless removal Bottle of 400 ml Each | Not applicable | 2 bottle of samples in an original packaging

## Item Specification (2) (hidden)
ANNEXURE A: ITEM SPECIFICATION
RFI-RT42-2021: SUPPLY AND DELIVERY OF BANDAGES AND DRESSINGS TO THE STATE
COMPANY NAME:
# | ITEM NUMBER | ITEM DESCRIPTION | COMMENTS ON ITEM SPECIFICATION | OTHER APPLICABLE QUALITY ASSURANCE COMPLIANCE
1 | RT42-01-001 | Dressing, burn, first aid non-adherent, each Sterile dressing for emergency or primary application to list 1st, 2nd, 3rd degree burns and scalds Size: 10cm x 10cm Properties of dressing: water based, absorb and dissipate heat, reduce burn progression, limit skin damage, inhibit fluid loss Properties of Gel: colourless to white, non-oily, contain 90 - 98% water, non-toxic, hypo-allergenic, low irritant, even to eyes, viscosity: 15 000 to 30 000 cPs pH: 5.5 to 7.0 Each
2 | RT42-01-002 | Dressing, burn, first aid, non-adherent, each Sterile dressing for emergency or primary application to list 1st, 2nd, 3rd degree burns and scalds Size: 20cm x 20cm Properties of dressing: water based, absorb and dissipate heat reduce burn progression, limit skin damage, inhibit fluid loss Properties of Gel: colourless to white, non-oily, contain 90-98% water, non-toxic, hypo-allergenic, low irritant, even to eyes, bacteriostatic viscosity: 15 000 - 30 000 cPs pH: 5.5 to 7.0 Label to state bacteriostatic Each
3 | RT42-01-003 | Dressing, burn, first aid, non-adherent, each Sterile dressing for emergency or primary application to list 1st, 2nd, 3rd degree burns and scalds Size: 60cm x 40cm Properties of dressing: water based, absorb and dissipate heat, reduce burn progression, limit skin damage, inhibit fluid loss Properties of Gel: colourless to white, non-oily, contain 90-98% water, non-toxic, hypo-allergenic, low irritant, even to eyes, viscosity: 15 000 - 30 000 cPs pH: 5.5 to 7.0 Each
4 | RT42-01-004 | Dressing, burn, first aid, multiple kit, Sterile dressing for pre-medical application to 1st, 2nd and 3rd degree burns and scalds Kit = 6 dressing: 2 x 100mm x 100mm 2 x 200mm x 200mm 2 x 600mm x 400mm Properties of dressing: water based, absorb and dissipate heat, reduce burn progression, limit skin damage, inhibit fluid loss Properties of Gel: colourless to white, non-oily, contain 96% water, non-toxic, hypo-allergenic, low irritant, even to eyes, viscosity: 15 000 - 30 000 cPs pH: 5.5 to 7.0 Kit
5 | RT42-01-005 | Dressing, burn, occlusive, hydrogel sheet, non-adherent, each Sterile dressing for application to 1st, 2nd degree burns and low to moderate draining wounds Size: 95mm x 95mm Properties of dressing: * water based absorb and dissipate heat * reduce burn progression * limit skin damage * inhibit fluid loss Properties of Gel: * transparent, colourless to white * non-oily * contain 90-98% water * non-toxic, hypy-allergenic * low irritant, even to eyes Each
6 | RT42-01-006 | Dressing, burn, occlusive, hydrogel sheet, non-adherent, each Sterile dressing for application to 1st, 2nd degree burns and low to moderate draining wounds Size: 150cm x 225cm Properties of dressing: * water based absorb and dissipate heat reduce burn progression * limit skin damage * inhibit fluid loss Properties of Gel: * Transparent, colourless to white * Non-oily * Contain 90-98% water * Non-toxic, hypo allergenic * low irritant even to eyes Each
7 | RT42-01-007 | Dressing, burn, first aid non-adherent, each Sterile dressing for emergency or primary application to list 1st, 2nd, 3rd degree burns and scalds Size: 5cm x15cm Properties of dressing: water based, absorb and dissipate heat, reduce burn progression, limit skin damage, inhibit fluid loss Properties of Gel: colourless to white, non-oily, contain 90-98% water, non-toxic, hypo-allergenic, low irritant, even to eyes v iscosity: 15 000 to 30 000 cPs pH: 5.5 to 7.0 Each
8 | RT42-01-008 | Dressing, burn, first aid non-adherent, each Sterile dressing for emergency or primary application to list 1st, 2nd, 3rd degree burns and scalds Size: 10cm x 40cm Properties of dressing: water based, absorb and dissipate heat, reduce burn progression, limit skin damage, inhibit fluid loss Properties of Gel: colourless to white, non-oily, contain 90-98% water, non-toxic, hypo-allergenic, low irritant, even to eyes, viscosity: 15 000 to 30 000 cPs pH: 5.5 to 7.0 Each
9 | RT42-01-009 | Dressing, burn, first aid non-adherent, each Sterile dressing for emergency or primary application to list 1st, 2nd, 3rd degree burns and scalds Size: 20cm x 45cm Properties of dressing: water based, absorb and dissipate heat, reduce burn progression, limit skin damage, inhibit fluid loss Properties of Gel: colourless to white, non-oily, contain 90-98% water, non-toxic, hypo-allergenic, low irritant, even to eyes, viscosity: 15 000 to 30 000 cPs pH: 5.5 to 7.0 Each
10 | RT42-01-010 | Dressing, burn, first aid non-adherent, each Sterile dressing for emergency or primary application to list 1st, 2nd, 3rd degree burns and scalds Size: 20cm x 55cm (hand dressing) Properties of dressing: water based, absorb and dissipate heat, reduce burn progression, limit skin damage, inhibit fluid loss Properties of Gel: colourless to white, non-oily, contain 90-98% water, non-toxic, hypo-allergenic, low irritant, even to eyes, viscosity: 15 000 to 30 000 cPs pH: 5.5 to 7.0 Each
11 | RT42-01-011 | Dressing, burn, first aid non-adherent, each Sterile dressing for emergency or primary application to list 1st, 2nd, 3rd degree burns and scalds Size: 30cm x 40cm (face mask) Properties of dressing: water based, absorb and dissipate heat, reduce burn progression, limit skin damage, inhibit fluid loss Properties of Gel: colourless to white, non-oily, contain 90-98% water, non-toxic, hypo-allergenic, low irritant, even to eyes, viscosity: 15 000 to 30 000 cPs pH: 5.5 to 7.0 Each
12 | RT42-01-012 | Dressing, burn, first aid non-adherent, each Sterile dressing, wrap, for emergency or primary application to list 1st, 2nd, 3rd degree burns and scalds Size: ±1m x 1m Properties of dressing: water based, absorb and dissipate heat, reduce burn progression, limit skin damage, inhibit fluid loss Properties of Gel: colourless to white, non-oily, contain 90-98% water, non-toxic, hypo-allergenic, low irritant, even to eyes, viscosity: 15 000 to 30 000 cPs pH: 5.5 to 7.0 Each
13 | RT42-01-013 | Dressing, burn, first aid non-adherent, each Sterile dressing, wrap, for emergency or primary application to list 1st, 2nd, 3rd degree burns and scalds Size: ±244cm x 183cm Properties of dressing: water based, absorb and dissipate heat, reduce burn progression, limit skin damage, inhibit fluid loss Properties of Gel: colourless to white, non-oily, contain 90-98% water, non-toxic, hypo-allergenic, low irritant, even to eyes, viscosity: 15 000 to 30 000 cPs pH: 5.5 to 7.0 Each
14 | RT42-01-014 | Dressing, burn, first aid non-adherent, each Sterile dressing for emergency or primary application to list 1st, 2nd, 3rd degree burns and scalds Size: 50cm x 1m Properties of dressing: water based, absorb and dissipate heat, reduce burn progression, limit skin damage, inhibit fluid loss Properties of Gel: colourless to white, non-oily, contain 96% water, non-toxic, hypo-allergenic, low irritant, even to eyes, viscosity: 15 000 to 30 000 cPs pH: 5.5 to 7.0 Each
15 | RT42-01-015 | Dressing, standard, first aid, sterile Consisting of an absorbent cotton wool pad to comply with the latest issue of SANS 228, enclosed in absorbent woven gauze to comply with the latest issue of SANS 446, attached to a conforming crepe or knitted type bandage to comply with the latest issue of SANS 1427 Compliance certificate to be submitted with bid Size: 50mm x 75mm pad 50mm x 1m bandage Sterile, individually peeled packed Kit
16 | RT42-01-016 | Dressing, standard, first aid, sterile Consisting of an absorbent cotton wool pad to comply with the latest issue of SANS 228, enclosed in absorbent woven gauze to comply with the latest issue of of SANS 446, attached to a conforming crepe or knitted type bandage to comply with the latest issue of SANS 1427 Compliance certificate to be submitted with bid Size: 75mm x 100mm pad; 75mm x 2m bandage Sterile, individually peeled packed Kit
17 | RT42-01-017 | Dressing, standard, first aid, sterile Consisting of an absorbent cotton wool pad to comply with the latest issue of SANS 228, enclosed in absorbent woven gauze o comply with the latest issue of of SANS 446, attached to a conforming crepe or knitted type bandage to comply with the latest issue of SANS 1427 Compliance certificate to be submitted with bid Size: 150mm x 200mm pad, 100mm x 4.5m bandage Sterile, individually peeled packed Kit
18 | RT42-01-018 | Dressing, standard, first aid, sterile, Consisting of an absorbent cotton wool pad to comply with the latest issue of SANS 228, enclosed in absorbent woven gauze to comply with the latest issue of of SANS 446, attached to a conforming crepe or knitted type bandage to comply with the latest issue of SANS 1427 Compliance certificate to be submitted with bid Size: 200mm x 300mm pad, 200mm x 4.5m bandage Sterile, individually peeled packed Kit
19 | RT42-01-019 | Dressing, burn, occlusive, hydrogel sheet, non-adherent, each Sterile dressing for application to 1st, 2nd degree burns and low to moderate draining wounds Size: 95mm x 95mm Properties of dressing: * water based absorb and dissipate heat * reduce burn progression * limit skin damage * inhibit fluid loss Properties of Gel: * transparent, colourless to white * non-oily * contain 90% water * non-toxic, hypy-allergenic * low irritant, even to eyes bacteriostatic Each
20 | RT42-01-020 | Dressing for burn wounds. Dressing must have open fingers which is an oversized glove. Dressing must have an antimicrobial silver contact wound barrier. Dressing should be able to release a continues flow of silver ions to wound. Dressing should be able to stayed for up to 7 days in place Latex Free individually wrapped Size: Extra Small Each
21 | RT42-01-021 | Dressing for burn wounds. Dressing must have open fingers which is an oversized glove. Dressing must have an antimicrobial silver contact wound barrier. Dressing should be able to release a continues flow of silver ions to wound. Dressing should be able to stayed for up to 7 days in place Latex Free individually wrapped Size: Small Each
22 | RT42-01-022 | Dressing for burn wounds. Dressing must have open fingers which is an oversized glove. Dressing must have an antimicrobial silver contact wound barrier. Dressing should be able to release a continues flow of silver ions to wound. Dressing should be able to stayed for up to 7 days in place Latex Free individually wrapped Size: Medium Each
23 | RT42-01-023 | Dressing for burn wounds. Dressing must have open fingers which is an oversized glove. Dressing must have an antimicrobial silver contact wound barrier. Dressing should be able to release a continues flow of silver ions to wound. Dressing should be able to stayed for up to 7 days in place Latex Free individually wrapped Size: Large Each
24 | RT42-01-024 | Dressing for burn wounds. Dressing must have open fingers which is an oversized glove. Dressing must have an antimicrobial silver contact wound barrier. Dressing should be able to release a continues flow of silver ions to wound. Dressing should be able to stayed for up to 7 days in place Latex Free individually wrapped Size: X-large Each
25 | RT42-02-001 | Dressing strip, fabric, elastic adhesive with a soft non-adherent, absorbent wound padding , B.P. Size: 40 mm x 1 m Each
26 | RT42-02-002 | Dressing strip, fabric, elastic adhesive with a soft non-adherent, absorbent wound padding , B.P. Size: 65 mm x 1 m Each
27 | RT42-02-003 | Dressing, absorbent gauze, type 1, bleached, 100 dressings Non sterile, autoclavable Size: 225mm x 225mm x 16ply To comply with the latest issue of SANS 446 Compliance certificate to be submitted with bid Packet
28 | RT42-02-004 | Dressing, paraffin gauze, B.P., 10 dressings in a tray Size: 100mm x 100mm Tray
29 | RT42-02-005 | Dressing, paraffin gauze, B.P., in a tray Size: 100mm x 7m One dressing per tray Tray
30 | RT42-02-006 | Dressing, paraffin gauze, B.P., 10 dressings Size: 100mm x 100mm Sterile individually peel packed Box
31 | RT42-02-007 | Dressing, paraffin gauze B.P. 10 dressings Size: 100mm x 400mm Sterile, individuallly peel packed Box
32 | RT42-02-008 | Dressing, wound, conformable, adhesive, hypo-allergenic, 20 dressings Sterile with a central highly absorbent non-adherent pad Size: 83mm x 120mm dressing 45mm x 75mm pad Sterile, individually peel packed Box
33 | RT42-02-009 | Dressing, wound, conformable, adhesive, hypo-allergenic, 20 dressings Sterile with a central highly absorbent non-adherent pad Size: 83mm x 180mm dressing 45mm x 120mm pad Sterile, individually peel packed Box
34 | RT42-02-010 | Dressing, wound, conformable, adhesive, hypo-allergenic, 50 dressings Sterile with a central highly absorbent non-adherent pad Size: 83mm x 60mm dressing 50mm x 35mm pad Sterile, individuallly peel packed Box
35 | RT42-02-011 | Dressing, wound, conformable, adhesive, hypo-allergenic, 20 dressings Sterile with a central highly absorbent non-adherent pad Size: 83mm x 240mm dressing 45mm x 170mm pad Sterile, individuallly peel packed Box
36 | RT42-02-012 | Dressing, wound, conformable, adhesive, hypo-allergenic, 20 dressings Sterile with a central highly absorbent non-adherent pad Size: 83mm x 415mm dressing 45mm x 345mm pad Sterile, individually peel packed Box
37 | RT42-02-013 | Dressing, tracheostomy, absorbent, non-adherent, waterproof Size: 90mm x 90mm Dressing to be: easy to apply and removable Sterile, individually peel packed Must be clearly marked "FOR TRACHEOSTOMY" Each
38 | RT42-02-014 | Dressing, tulle gras, 10 dressings Impregnated with chlorhexidine 0.5% m/m Size: 100mm x 100mm Sterile, individually peel packed SAHPRA product registration to be submitted Box
39 | RT42-02-015 | Dressing, tulle grass Impregnated with chlorhexidine 0.5% m/m Size: 100mm x 400mm Sterile, individually peel packed SAHPRA product registration to be submitted Each
40 | RT42-02-016 | Dressing, wound Non-adherent dressing polypropylene/polythylene wound contact layer Size: 100mm x 100mm x 2ply Sterile, individually peel packed Each
41 | RT42-02-017 | Dressing, film, transparent, barrier material, For intravenous catheter placement Barrier material to be: tear resistant, conform to irregular surfaces, antistatic, permiable to oxygen and water vapour and waterproof Adhesive: low-irritant hypo-allergenic pressure sensitive, to adhere firmly so as to keep catheter in position covered with protective silicone treated release backing Packing: Must include instructions for use Sterile, individually peel packed with description on outer pack Size: ±60mm x 70mm Each
42 | RT42-02-018 | Dressing, film, transparent, barrier material, For intravenous catheter placement Barrier material to be: tear resistant, conform to irregular surfaces, antistatic, permiable to oxygen and water vapour and waterproof Adhesive: low-irritant hypo-allergenic pressure sensitive, to adhere firmly so as to keep catheter in position covered with protective silicone treated release backing Packing: Must include instructions for use Sterile, individually peel packed with description on outer pack Size: ±60mm x 85mm Each
43 | RT42-02-019 | Dressing, film, transparent, barrier material, For intravenous catheter placement Barrier material to be: tear resistant, conform to irregular surfaces, antistatic, permiable to oxygen and water vapour and waterproof Adhesive: low-irritant hypo-allergenic pressure sensitive, to adhere firmly so as to keep catheter in position covered with protective silicone treated release backing Packing: Must include instructions for use Sterile, individually peel packed with description on outer pack Size: ±100mm x 120mm Each
44 | RT42-02-020 | Dressing, film, transparent, barrier material, For intravenous catheter placement Barrier material to be: tear resistant, conform to irregular surfaces, antistatic, permiable to oxygen and water vapour and waterproof Adhesive: low-irritant hypo-allergenic pressure sensitive, to adhere firmly so as to keep catheter in position covered with protective silicone treated release backing Packing: Must include instructions for use Sterile, individually peel packed with description on outer pack Size: ±100mm x 140mm Each
45 | RT42-02-021 | Dressing, film, transparent, barrier material, For intravenous catheter placement Barrier material to be: tear resistant, conform to irregular surfaces, antistatic, permiable to oxygen and water vapour and waterproof Adhesive: low-irritant hypo-allergenic pressure sensitive, to adhere firmly so as to keep catheter in position covered with protective silicone treated release backing Packing: Must include instructions for use Sterile, individually peel packed with description on outer pack Size: ±100mm x 200mm Each
46 | RT42-02-022 | Dressing, film, transparent, barrier material, For intravenous catheter placement Barrier material to be: tear resistant, conform to irregular surfaces, antistatic, permiable to oxygen and water vapour and waterproof Adhesive: low-irritant hypo-allergenic pressure sensitive, to adhere firmly so as to keep catheter in position covered with protective silicone treated release backing Packing: Must include instructions for use Sterile, individually peel packed with description on outer pack Size: ±100mm x 250mm Each
47 | RT42-02-023 | Dressing, film, transparent, barrier material, For intravenous catheter placement Barrier material to be: tear resistant, conform to irregular surfaces, antistatic, permiable to oxygen and water vapour and waterproof Adhesive: low-irritant hypo-allergenic pressure sensitive, to adhere firmly so as to keep catheter in position covered with protective silicone treated release backing Packing: Must include instructions for use Sterile, individually peel packed with description on outer pack Size: ±150mm x 200mm Each
48 | RT42-02-024 | Dressing, film, transparent, barrier material, For intravenous catheter placement Barrier material to be: tear resistant, conform to irregular surfaces, antistatic, permiable to oxygen and water vapour and waterproof Adhesive: low-irritant hypo-allergenic pressure sensitive, to adhere firmly so as to keep catheter in position covered with protective silicone treated release backing Packing: Must include instructions for use Sterile, individually peel packed with description on outer pack Size: ±200mm x 300mm Each
49 | RT42-02-025 | Dressing, film, transparent, Island Dressing, Central Island to be: highly absorbent, non-adherent pad Barrier material to be: transparent, matt finished and tear resistant conformable to irregular surfaces antistatic permiable to oxygen and water vapour waterproof Adhesive: low-irritant hypo-allergenic pressure sensitive, to adhere firmly so as to keep catheter in position covered with protective silicone treated release backing two tape handles to facilitate easy tear off after positioning Packing: must include instructions for use on outer pack Sterile, individually peel packed with description on outer pack Size Dressing: ±50mm x 70mm Pad Size: ±50mm x 35mm Each
50 | RT42-02-026 | Dressing, film, transparent, Island Dressing, Central Island to be: highly absorbent, non-adherent pad Barrier material to be: transparent, matt finished and tear resistant conformable to irregular surfaces antistatic permiable to oxygen and water vapour waterproof Adhesive: low-irritant hypo-allergenic pressure sensitive, to adhere firmly so as to keep catheter in position covered with protective silicone treated release backing two tape handles to facilitate easy tear off after positioning Packing: must include instructions for use on outer pack Sterile, individually peel packed with description on outer pack Size Dressing: ±83mm x 60mm Pad Size: ±50mm x 35mm Each
51 | RT42-02-027 | Dressing, film, transparent, Island Dressing, Central Island to be: highly absorbent, non-adherent pad Barrier material to be: transparent, matt finished and tear resistant conformable to irregular surfaces antistatic permiable to oxygen and water vapour waterproof Adhesive: low-irritant hypo-allergenic pressure sensitive, to adhere firmly so as to keep catheter in position covered with protective silicone treated release backing two tape handles to facilitate easy tear off after positioning Packing: must include instructions for use on outer pack Sterile, individually peel packed with description on outer pack Size Dressing: ±83mm x 240mm Pad Size: ±45mm x 170mm Each
52 | RT42-02-028 | Dressing, film, transparent, Island Dressing, Central Island to be: highly absorbent, non-adherent pad Barrier material to be: transparent, matt finished and tear resistant conformable to irregular surfaces antistatic permiable to oxygen and water vapour waterproof Adhesive: low-irritant hypo-allergenic pressure sensitive, to adhere firmly so as to keep catheter in position covered with protective silicone treated release backing two tape handles to facilitate easy tear off after positioning Packing: must include instructions for use on outer pack Sterile, individually peel packed with description on outer pack Size Dressing: ±100mm x 60mm Pad Size: ±50mm x 35mm Each
53 | RT42-02-029 | Dressing, film, transparent, Island Dressing, Central Island to be: highly absorbent, non-adherent pad Barrier material to be: transparent, matt finished and tear resistant conformable to irregular surfaces antistatic permiable to oxygen and water vapour waterproof Adhesive: low-irritant hypo-allergenic pressure sensitive, to adhere firmly so as to keep catheter in position covered with protective silicone treated release backing two tape handles to facilitate easy tear off after positioning Packing: must include instructions for use on outer pack Sterile, individually peel packed with description on outer pack Size Dressing: ±100mm x 120mm Pad Size: ±45mm x 75mm Each
54 | RT42-02-030 | Dressing, film, transparent, Island Dressing, Central Island to be: highly absorbent, non-adherent pad Barrier material to be: transparent, matt finished and tear resistant conformable to irregular surfaces antistatic permiable to oxygen and water vapour waterproof Adhesive: low-irritant hypo-allergenic pressure sensitive, to adhere firmly so as to keep catheter in position covered with protective silicone treated release backing two tape handles to facilitate easy tear off after positioning Packing: must include instructions for use on outer pack Sterile, individually peel packed with description on outer pack Size Dressing: ±100mm x 180mm Pad Size: ±45mm x 75mm Each
55 | RT42-02-031 | Dressing, film, transparent, Island Dressing, Central Island to be: highly absorbent, non-adherent pad Barrier material to be: transparent, matt finished and tear resistant conformable to irregular surfaces antistatic permiable to oxygen and water vapour waterproof Adhesive: low-irritant hypo-allergenic pressure sensitive, to adhere firmly so as to keep catheter in position covered with protective silicone treated release backing two tape handles to facilitate easy tear off after positioning Packing: must include instructions for use on outer pack Sterile, individually peel packed with description on outer pack Size Dressing: ±100mm x 200mm Pad Size: ±45mm x 75mm Each
56 | RT42-02-032 | Dressing, film, transparent, Island Dressing, Central Island to be: highly absorbent, non-adherent pad Barrier material to be: transparent, matt finished and tear resistant conformable to irregular surfaces antistatic permiable to oxygen and water vapour waterproof Adhesive: low-irritant hypo-allergenic pressure sensitive, to adhere firmly so as to keep catheter in position covered with protective silicone treated release backing two tape handles to facilitate easy tear off after positioning Packing: must include instructions for use on outer pack Sterile, individually peel packed with description on outer pack Size Dressing: ±100mm x 240mm Pad Size: ±45mm x 170mm Each
57 | RT42-02-033 | Dressing, film, transparent, Island Dressing, Central Island to be: highly absorbent, non-adherent pad Barrier material to be: transparent, matt finished and tear resistant conformable to irregular surfaces antistatic permiable to oxygen and water vapour waterproof Adhesive: low-irritant hypo-allergenic press covered with protective silicone treated release backing two tape handles to facilitate easy tear off after positioning Packing: must include instructions for use on outer pack Sterile, individually peel packed with description on outer pack Size Dressing: ±100mm x 350mm Pad Size: ±45mm x 75mm Each
58 | RT42-02-034 | Dressing, post-operative/wound, conformable, adhesive, hypoallergenic, Sterile with a central highly absorbent non-adherent non-stripe through pad must include instructions for use on outer pack Sterile, individually peel packed with description on outer pack Size Dressing: ±83mm x 60mm Pad Size: ±50mm x 35mm Each
59 | RT42-02-035 | Dressing, post-operative/wound, conformable, adhesive, hypoallergenic, Sterile with a central highly absorbent non-adherent non-stripe through pad must include instructions for use on outer pack Sterile, individually peel packed with description on outer pack Size Dressing: ±83mm x 120mm Pad Size: ±45mm x 75mm Each
60 | RT42-02-036 | Dressing, post-operative/wound, conformable, adhesive, hypoallergenic, Sterile with a central highly absorbent non-adherent non-stripe through pad must include instructions for use on outer pack Sterile, individually peel packed with description on outer pack Size Dressing: ±83mm x 180mm Pad Size: ±45mm x 120mm Each
61 | RT42-02-037 | Dressing, post-operative/wound, conformable, adhesive, hypoallergenic, Sterile with a central highly absorbent non-adherent non-stripe through pad Must include instructions for use on outer pack Sterile, individually peel packed with description on outer pack Size Dressing: ±83mm x 240mm Pad Size: ±45mm x 170mm Each
62 | RT42-02-038 | Dressing, post-operative/wound, conformable, adhesive, hypoallergenic, Sterile with a central highly absorbent non-adherent non-stripe through pad Must include instructions for use on outer pack Sterile, individually peel packed with description on outer pack Size Dressing: ±83mm x 415mm Pad Size: ±45mm x 345mm Each
63 | RT42-02-039 | Dressing, post-operative/wound, conformable, adhesive, hypoallergenic, Sterile with a central highly absorbent non-adherent non-stripe through pad Must include instructions for use on outer pack Sterile, individually peel packed with description on outer pack Size Dressing: ±100mm x 300mm Pad Size: ±45mm x 250mm Each
64 | RT42-02-040 | Dressing, Wound Polymer dressing, Highly absorbent Sterile non-adherent, moisture-vapour permeable Waterproof polyurethane. Dressing contains silver complex Must include instructions for use on outer pack Sterile, individually peel packed Size Dressing: ±50mm x 70mm Each
65 | RT42-02-041 | Dressing, Wound Polymer dressing, Highly absorbent Sterile non-adherent, moisture-vapour permeable Waterproof polyurethane. Dressing contains silver complex Must include instructions for use on outer pack Sterile, individually peel packed Size Dressing: ±100mm x 100mm Each
66 | RT42-02-042 | Dressing,Wound, Polymer dressing, Highly absorbent Sterile non-adherent, moisture-vapour permeable Waterproof polyurethane. Dressing contains silver complex Must include instructions for use on outer pack Sterile, individually peel packed Size Dressing: ±100mm x 120mm Each
67 | RT42-02-043 | Dressing, Wound Polymer dressing, Highly absorbent Sterile non-adherent, moisture-vapour permeable Waterproof polyurethane. Dressing contains silver complex Must include instructions for use on outer pack Sterile, individually peel packed Size Dressing: ±100mm x 200mm Each
68 | RT42-02-044 | Dressing, Wound Polymer dressing, Highly absorbent Sterile non-adherent, moisture-vapour permeable Waterproof polyurethane. Dressing contains silver complex Must include instructions for use on outer pack Sterile, individually peel packed Size Dressing: ±150mm x 150mm Each
69 | RT42-02-045 | Dressing, Wound, Polymer dressing, Highly absorbent Sterile non-adherent, moisture-vapour permeable Waterproof polyurethane. Dressing contains silver complex Must include instructions for use on outer pack Sterile, individually peel packed Size Dressing: ±150mm x 200mm Each
70 | RT42-02-046 | Dressing, Wound Polymer dressing, Highly absorbent Sterile non-adherent, moisture-vapour permeable Waterproof polyurethane. Dressing contains silver complex Must include instructions for use on outer pack Sterile, individually peel packed Size Dressing: ±200mm x 200mm Each
71 | RT42-02-047 | Dressing, Wound Polymer dressing, Highly absorbent Sterile non-adherent, moisture-vapour permeable Waterproof polyurethane. Dressing contains Ibuprofen, released at 0.5g/cm². Sterile, individually peel packed Size Dressing: ±100mm x 200mm Each
72 | RT42-02-048 | Dressing, Wound Polymer dressing, Highly absorbent Sterile non-adherent, moisture-vapour permeable Waterproof polyurethane. Dressing contains Ibuprofen, released at 0.5g/cm². Sterile, individually peel packed Size Dressing: ±150mm x 150mm Each
73 | RT42-02-049 | Dressing, Wound Polymer dressing, Highly absorbent Sterile non-adherent, moisture-vapour permeable Waterproof polyurethane. Dressing contains Ibuprofen, released at 0.5g/cm². Sterile, individually peel packed Size Dressing: ±200mm x 200mm Each
74 | RT42-02-050 | Dressing Wound, Foam, absorbent, non-adherent Highly absorbent centre layer to accommodate large volume exudate and odour, polypropylene/polyethylene wound contact layer of dressing to be water proof. Non-adherent wound contact layer. Sterile, individually peel packed Size Dressing: ±50mm x 70mm Each
75 | RT42-02-051 | Dressing Wound, Foam, absorbent, non-adherent Highly absorbent centre layer to accommodate large volume exudate and odour, polypropylene/polyethylene wound contact layer of dressing to be water proof. Non-adherent wound contact layer. Sterile, individually peel packed Size Dressing: ±100mm x 100mm Each
76 | RT42-02-052 | Dressing Wound, Foam, absorbent, non-adherent Highly absorbent centre layer to accommodate large volume exudate and odour, polypropylene/polyethylene wound contact layer of dressing to be water proof. Non-adherent wound contact layer. Sterile, individually peel packed Size Dressing: ±100mm x 200mm Each
77 | RT42-02-053 | Dressing Wound, Foam, absorbent, non-adherent Highly absorbent centre layer to accommodate large volume exudate and odour, polypropylene/polyethylene wound contact layer of dressing to be water proof. Non-adherent wound contact layer. Sterile, individually peel packed Size Dressing: ±150mm x 150mm Each
78 | RT42-02-054 | Dressing Wound, Foam, absorbent, non-adherent Highly absorbent centre layer to accommodate large volume exudate and odour, polypropylene/polyethylene wound contact layer of dressing to be water proof. Non-adherent wound contact layer. Sterile, individually peel packed Size Dressing: ±180mm x 180mm Each
79 | RT42-02-055 | Dressing Wound, Foam, absorbent, non-adherent Highly absorbent centre layer to accommodate large volume exudate and odour, polypropylene/polyethylene wound contact layer of dressing to be water proof. Non-adherent wound contact layer. Sterile, individually peel packed Size Dressing: ±200mm x 200mm Each
80 | RT42-02-056 | Dressing Wound, Foam, absorbent, non-adherent, Tracheostomy Highly absorbent centre layer to accommodate large volume exudate and odour, polypropylene/polyethylene wound contact layer of dressing to be water proof. Non-adherent wound contact layer. Sterile, individually peel packed Size Dressing: +-100mm x +-100mm Each
81 | RT42-02-057 | Dressing Wound, Foam, absorbent, non-adherent, Heel Highly absorbent centre layer to accommodate large volume exudate and odour, polypropylene/polyethylene wound contact layer of dressing to be water proof. Non-adherent wound contact layer. Sterile, individually peel packed Size Dressing: ±100mm x 100mm Each
82 | RT42-02-058 | Dressing Wound, Foam, absorbent, non-adherent Highly absorbent centre layer to accommodate large volume exudate and odour, polypropylene/polyethylene wound contact layer of dressing to be water proof. Non-adherent wound contact layer. Sterile, individually peel packed Size Dressing: ±100mm x 600mm Roll
83 | RT42-02-059 | Dressing Wound, Foam, absorbent, self-adherent Highly absorbent centre layer to accommodate large volume exudate and odour, polypropylene/polyethylene wound contact layer impregbated with silver, non-adherent and sterile. Outerlayer to be water proof. Non-adherent wound contact layer. Sterile, individually peel packed Size Dressing: ±75mm x 75mm Each
84 | RT42-02-060 | Dressing Wound, Foam, absorbent, self-adherent Highly absorbent centre layer to accommodate large volume exudate and odour, polypropylene/polyethylene wound contact layer impregbated with silver, non-adherent and sterile. Outerlayer to be water proof. Non-adherent wound contact layer. Sterile, individually peel packed Size Dressing: ±100mm x 100mm Each
85 | RT42-02-061 | Dressing Wound, Foam, absorbent, self-adherent Highly absorbent centre layer to accommodate large volume exudate and odour, polypropylene/polyethylene wound contact layer impregbated with silver, non-adherent and sterile. Outerlayer to be water proof. Non-adherent wound contact layer. Sterile, individually peel packed Size Dressing: ±125mm x 125mm Each
86 | RT42-02-062 | Dressing Wound, Foam, absorbent, self-adherent Highly absorbent centre layer to accommodate large volume exudate and odour, polypropylene/polyethylene wound contact layer impregbated with silver, non-adherent and sterile. Outerlayer to be water proof. Non-adherent wound contact layer. Sterile, individually peel packed Size Dressing: ±150mm x 150mm Each
87 | RT42-02-063 | Dressing Wound, Foam, absorbent, self-adherent Highly absorbent centre layer to accommodate large volume exudate and odour, polypropylene/polyethylene wound contact layer impregbated with silver, non-adherent and sterile. Outerlayer to be water proof. Non-adherent wound contact layer. Sterile, individually peel packed Size Dressing: ±180mm x 180mm Each
88 | RT42-02-064 | Dressing Wound, Foam, absorbent, self-adherent Highly absorbent centre layer to accommodate large volume exudate and odour, polypropylene/polyethylene wound contact layer impregbated with silver, non-adherent and sterile. Outerlayer to be water proof. Non-adherent wound contact layer. Sterile, individually peel packed Size Dressing: ±225mm x 225mm Each
89 | RT42-02-065 | Dressing Wound, Foam, absorbent, self-adherent , Heel Highly absorbent centre layer to accommodate large volume exudate and odour, polypropylene/polyethylene wound contact layer impregbated with silver, non-adherent and sterile. Outerlayer to be water proof. Non-adherent wound contact layer. Sterile, individually peel packed Size Dressing: 100mm x130mm Each
90 | RT42-02-066 | Dressing Wound, Foam, absorbent, self-adherent , Sacral Highly absorbent centre layer to accommodate large volume exudate and odour, polypropylene/polyethylene wound contact layer impregbated with silver, non-adherent and sterile. Outerlayer to be water proof. Non-adherent wound contact layer. Sterile, individually peel packed Size Dressing: 180mm x180mm Each
91 | RT42-02-067 | Dressing Wound, absorbent, self-adherent Highly absorbent centre layer to accommodate large volume exudate and odour, layer of dressing to be waterproof. Outer layer waterproof, tapered boders to prevent leakage. Non-adherent wound contact layer. Sterile, individually peel packed Size of Dressing: ±75mm x 75mm Each
92 | RT42-02-068 | Dressing Wound, absorbent, self-adherent Highly absorbent centre layer to accommodate large volume exudate and odour, layer of dressing to be waterproof. Outer layer waterproof, tapered boders to prevent leakage. Non-adherent wound contact layer.. Sterile, individually peel packed Size of Dressing: ±100mm x 100mm Each
93 | RT42-02-069 | Dressing Wound, absorbent, self-adherent Highly absorbent centre layer to accommodate large volume exudate and odour, layer of dressing to be waterproof. Outer layer waterproof, tapered boders to prevent leakage. Non-adherent wound contact layer. Sterile, individually peel packed Size of Dressing: ±125mm x 125mm Each
94 | RT42-02-070 | Dressing Wound, absorbent, self-adherent, Heel Highly absorbent centre layer to accommodate large volume exudate and odour, layer of dressing to be waterproof. Outer layer waterproof, tapered boders to prevent leakage. Non-adherent wound contact layer. Sterile, individually peel packed Size of Dressing: Each
95 | RT42-02-071 | Dressing Wound, absorbent, self-adherent, Sacral Highly absorbent centre layer to accommodate large volume exudate and odour, layer of dressing to be waterproof. Outer layer waterproof, tapered boders to prevent leakage. Non-adherent wound contact layer. Sterile, individually peel packed Each
96 | RT42-02-072 | Dressing, hydrocolloid Sodium Carboxymethyl cellulose hydrocolloid layer bonded to a polyurethane film, transparent for visual inspection, flexible dressing. Sterile, individually peel packed Size of Dressing: ±40mm x 60mm Each
97 | RT42-02-073 | Dressing, hydrocolloid Sodium Carboxymethyl cellulose hydrocolloid layer bonded to a polyurethane film, transparent for visual inspection, flexible dressing. Sterile, individually peel packed Size of Dressing: ±100mm x 100mm Each
98 | RT42-02-074 | Dressing, hydrocolloid Sodium Carboxymethyl cellulose hydrocolloid layer bonded to a polyurethane film, transparent for visual inspection, flexible dressing. Sterile, individually peel packed Size of Dressing: ±150mm x 150mm Each
99 | RT42-02-075 | Dressing, hydrocolloid Sodium Carboxymethyl cellulose hydrocolloid layer bonded to a polyurethane film, transparent for visual inspection, flexible dressing. Sterile, individually peel packed Size of Dressing: ±200mm x 200mm Each
100 | RT42-02-076 | Dressing, hydrocolloid, Sacral dressing Sodium Carboxymethyl cellulose hydrocolloid layer bonded to a polyurethane film, transparent for visual inspection, flexible dressing. Sterile, individually peel packed Size of Dressing: ±180mm x 150mm Each
101 | RT42-02-077 | Dressing, hydrocolloid, Sacral dressing Sodium Carboxymethyl cellulose hydrocolloid layer bonded to a polyurethane film, transparent for visual inspection, flexible dressing. Sterile, individually peel packed Size of Dressing: ±200mm x 230mm Each
102 | RT42-02-078 | Dressing, hydrocolloid, with border Sodium Carboxymethyl cellulose, pectin, gelatin Adhesive to be: hydrocolloid layer bonded to a polyurethane foam layer for protection, flexible, hypo-allergenic. Sterile, individually peel packed Size of Dressing: ±100mm x 100mm Each
103 | RT42-02-079 | Dressing, hydrocolloid, with border Sodium Carboxymethyl cellulose, pectin, gelatin Adhesive to be: hydrocolloid layer bonded to a polyurethane foam layer for protection, flexible, hypo-allergenic. Sterile, individually peel packed Size of Dressing: ±150mm x 150mm Each
104 | RT42-02-080 | Dressing, hydrocolloid, with border Sodium Carboxymethyl cellulose, pectin, gelatin Adhesive to be: hydrocolloid layer bonded to a polyurethane foam layer for protection, flexible, hypo-allergenic. Sterile, individually peel packed Size of Dressing: ±150mm x 180mm Each
105 | RT42-02-081 | Dressing Wound, absorbent, self-adherent Highly absorbent centre layer to accommodate large volume exudate and odour, layer of dressing to be waterproof. Outer layer waterproof, tapered boders to prevent leakage. Non-adherent wound contact layer. Sterile, individually peel packed Size of Dressing: ±150mm x 150mm Each
106 | RT42-02-082 | Dressing Wound, absorbent, self-adherent Highly absorbent centre layer to accommodate large volume exudate and odour, layer of dressing to be waterproof. Outer layer waterproof, tapered boders to prevent leakage. Non-adherent wound contact layer. Sterile, individually peel packed Size of Dressing:±180mm x 180mm Each
107 | RT42-02-083 | Dressing, hydrocolloid, with border Sodium Carboxymethyl cellulose, pectin, gelatin Adhesive to be: hydrocolloid layer bonded to a polyurethane foam layer for protection, flexible, hypo-allergenic. Sterile, individually peel packed Size of Dressing: ±200mm x 200mm Each
108 | RT42-02-084 | Dressing, hydrocolloid, with border Sodium Carboxymethyl cellulose, pectin, gelatin Adhesive to be: hydrocolloid layer bonded to a polyurethane foam layer for protection, flexible, hypo-allergenic. Sterile, individually peel packed Size of Dressing: ±200mm x 230mm Each
109 | RT42-02-085 | Dressing, hydrocolloid, Transparent, Sodium Carboxymethyl cellulose, pectin, gelatin, thin, hydrocolloid layer bonded to a polyurethane film, transparent for visual inspection, flexible dressing. Sterile, individually peel packed Size of Dressing: ±50mm x 70mm Each
110 | RT42-02-086 | Dressing, hydrocolloid, Transparent, Sodium Carboxymethyl cellulose, pectin, gelatin, thin, hydrocolloid layer bonded to a polyurethane film, transparent for visual inspection, flexible dressing. Sterile, individually peel packed Size of Dressing: ±50mm x 100mm Each
111 | RT42-02-087 | Dressing, hydrocolloid, Transparent, Sodium Carboxymethyl cellulose, pectin, gelatin, thin, hydrocolloid layer bonded to a polyurethane film, transparent for visual inspection, flexible dressing. Sterile, individually peel packed Size of Dressing: ±50mm x 150mm Each
112 | RT42-02-088 | Dressing, hydrocolloid, Transparent, Sodium Carboxymethyl cellulose, pectin, gelatin, thin, hydrocolloid layer bonded to a polyurethane film, transparent for visual inspection, flexible dressing. Sterile, individually peel packed Size of Dressing: ±50mm x 200mm Each
113 | RT42-02-089 | Dressing, hydrocolloid, Transparent, Sodium Carboxymethyl cellulose, pectin, gelatin, thin, hydrocolloid layer bonded to a polyurethane film, transparent for visual inspection, flexible dressing. Sterile, individually peel packed Size of Dressing: ±50mm x 250mm Each
114 | RT42-02-090 | Dressing, hydrocolloid, Transparent, Sodium Carboxymethyl cellulose, pectin, gelatin, thin, hydrocolloid layer bonded to a polyurethane film, transparent for visual inspection, flexible dressing. Sterile, individually peel packed Size of Dressing: ±90mm x 140mm Each
115 | RT42-02-091 | Dressing, hydrocolloid, Transparent, Sodium Carboxymethyl cellulose, pectin, gelatin, thin, hydrocolloid layer bonded to a polyurethane film, transparent for visual inspection, flexible dressing. Sterile, individually peel packed Size of Dressing: ±100mm x 100mm Each
116 | RT42-02-092 | Dressing, hydrocolloid, Transparent, Sodium Carboxymethyl cellulose, pectin, gelatin, thin, hydrocolloid layer bonded to a polyurethane film, transparent for visual inspection, flexible dressing. Sterile, individually peel packed Size of Dressing: ±150mm x 150mm Each
117 | RT42-02-093 | Dressing, hydrocolloid, Transparent, Sodium Carboxymethyl cellulose, pectin, gelatin, thin, hydrocolloid layer bonded to a polyurethane film, transparent for visual inspection, flexible dressing. Sterile, individually peel packed Size of Dressing: ±200mm x 200mm Each
118 | RT42-02-094 | Dressing, Alginate Absorbent, non-adherent, sterile, Sterile, individually peel packed Size of Dressing: ±27mm x 340mm Each
119 | RT42-02-095 | Dressing, Alginate Absorbent, non-adherent, sterile, Sterile, individually peel packed Size of Dressing: ±60mm x 60mm Each
120 | RT42-02-096 | Dressing, Alginate Absorbent, non-adherent, sterile, Sterile, individually peel packed Size of Dressing: ±100mm x 100mm Each
121 | RT42-02-097 | Dressing, Alginate Absorbent, non-adherent, sterile, Sterile, individually peel packed Size of Dressing: ±100mm x 200mm Each
122 | RT42-02-098 | Dressing, Alginate Absorbent, non-adherent, sterile, Sterile, individually peel packed Size of Dressing: ±150mm x 150mm Each
123 | RT42-02-099 | Dressing, Alginate Absorbent, non-adherent, sterile, Sterile, individually peel packed Size of Dressing: ±150mm x 250mm Each
124 | RT42-02-100 | Dressing, Silver Meshed Non-woven, porous, contains Silver Sulphate ±8mg/g, sterile. Sterile, individually peel packed Size of Dressing: ±100mm x 120mm Each
125 | RT42-02-101 | Dressing, Silver Meshed Non-woven, porous, contains Silver Sulphate ±8mg/g, sterile. Sterile, individually peel packed Size of Dressing: ±100mm x 200mm Each
126 | RT42-02-102 | Dressing, Silver Meshed Non-woven, porous, contains Silver Sulphate ±8mg/g, sterile. Sterile, individually peel packed Size of Dressing: ±150mm x 200mm Each
127 | RT42-02-103 | Dressing, Silver Meshed Non-woven, porous, contains Silver Sulphate ±8mg/g, sterile. Sterile, individually peel packed Size of Dressing: ±200mm x 200mm Each
128 | RT42-02-104 | Dressing, Silver Meshed Non-woven, porous, contains Silver Sulphate ±8mg/g, sterile. Sterile, individually peel packed Size of Dressing: ±400mm x 400mm Each
129 | RT42-02-105 | Dressing Wound, Povidone-Iodine impregnated, non-adherent Sterile, individually peel packed Size of Dressing: ±50mm x 50mm , less than 10%m/m Each
130 | RT42-02-106 | Dressing Wound, Povidone-Iodine impregnated, non-adherent Sterile, individually peel packed Size of Dressing: ±95mm x 950mm , less than 10%m/m Each
131 | RT42-02-107 | Dressing Wound, Honey, non-adherent. Sterile, individually peel packed Size of Dressing: ±100mm x 100mm Each
132 | RT42-02-108 | Dressing Wound, Honey, non-adherent.. Sterile, individually peel packed Size of Dressing: ±100mm x 200mm Each
133 | RT42-02-109 | Dressing Wound, Alginate, Impregnated with Honey, non-adherent.. Sterile, individually peel packed Size of Dressing: ±50mm x 50mm Each
134 | RT42-02-110 | Dressing Wound, Alginate, Impregnated with Honey, non-adherent.. Sterile, individually peel packed Size of Dressing: ±100mm x 100mm Each
135 | RT42-02-111 | Dressing Gauze, Impregnated with Honey, non-adherent.. Sterile, individually peel packed Size of Dressing: ±50mm x 50mm Each
136 | RT42-02-112 | Dressing, Gauze, Impregnated with Honey, non-adherent.. Sterile, individually peel packed Size of Dressing: ±100mm x 100mm Each
137 | RT42-02-113 | Dressing, gauze, non-woven, with haemostatic properties 3 ply trauma pad Size: ±30,5cm x ±30,5cm x 3-ply Individually packed Sterile Box of 10 dressings
138 | RT42-02-114 | Dressing Wound, , absorbent, with adhesive border Highly absorbent centre layer to accommodate large volume exudate and odour, polypropylene/polyethylene wound contact layer,non-adherent. Outerlayer of dressing to be waterproof. Sterile, individually peel packed Size Dressing: ±150mm x 150mm Each
139 | RT42-02-115 | Dressing Wound, , absorbent, with adhesive border Highly absorbent centre layer to accommodate large volume exudate and odour, polypropylene/polyethylene wound contact layer,non-adherent. Outerlayer of dressing to be waterproof. Sterile, individually peel packed Size Dressing: ±200mm x 200mm Each
140 | RT42-02-116 | Dressing Wound, , hydrocapilary, woven Absorbent, layered dressing to assist in the removal of exudate and dead tissue, be non-adherent to wound surface, sterile. . Should have description on outer pack Sterile, individually peel packed Size Dressing: ±50mm x 50mm Each
141 | RT42-02-117 | Dressing Wound, , hydrocapilary, woven Absorbent, layered dressing to assist in the removal of exudate and dead tissue, be non-adherent to wound surface, sterile. . Should have description on outer pack Sterile, individually peel packed Size Dressing: ±75mm x 75mm Each
142 | RT42-02-118 | Dressing Wound, , hydrocapilary, woven Absorbent, layered dressing to assist in the removal of exudate and dead tissue, be non-adherent to wound surface, sterile. . Should have description on outer pack Sterile, individually peel packed Size Dressing: ±100mm x 100mm Each
143 | RT42-02-119 | Dressing Wound, , hydrocapilary, woven Absorbent, layered dressing to assist in the removal of exudate and dead tissue, be non-adherent to wound surface, sterile. . Should have description on outer pack Sterile, individually peel packed Size Dressing: ±150mm x 150mm Each
144 | RT42-02-120 | Dressing Wound, , hydrocapilary, woven Absorbent, layered dressing to assist in the removal of exudate and dead tissue, be non-adherent to wound surface, sterile. Should have description on outer pack Sterile, individually peel packed Size Dressing: ±100mm x 1m Each
145 | RT42-02-121 | Dressing wound, , hydrocapilary, woven, stretch, absorbent, layered dressing to assist in the removal of exudate and dead tissue, be non-adherent to wound surface, sterile, should have description on outer pack, sterile, individually peel packed. Size100mm x 1m
146 | RT42-02-122 | Dressing Wound, , hydrocapilary, woven, Tracheostomy Absorbent, layered dressing to assist in the removal of exudate and dead tissue, be non-adherent to wound surface, sterile. Should have description on outer pack Sterile, individually peel packed Size Dressing :100mm x100mm Each
147 | RT42-02-123 | Dressing Wound, , hydrocapilary, woven Absorbent, layered dressing to assist in the removal of exudate and dead tissue, be non-adherent to wound surface, sterile. . Should have description on outer pack Sterile, individually peel packed Size Dressing: ±200mm x 200mm Each
148 | RT42-02-124 | Dressing Wound, , hydrocapilary, woven , Cavity Absorbent, layered dressing to assist in the removal of exudate and dead tissue, be non-adherent to wound surface, sterile. . Should have description on outer pack Sterile, individually peel packed Size Dressing: ±100mm x 100mm Each
149 | RT42-02-125 | Dressing wound, imprenaged with antimicrobial/antibacterial absorbent properties, must be have an open-pored sterile, freeze-dried matrix, must contains ionically bound silver, which provides antimicrobial protection against bacteria and infection. The absorbent collagen sheets manage wound moisture which must converts to soft gel sheet that stays in intimate contact with the wound bed, Three-dimensional structure helps support cellular growth Dressing must be composed of 55% collagen & and 45% oxidised regenerated cellulose Size: 123m3 Sterile, Individually wrapped Box of 10 dressings
150 | RT42-02-126 | Dressing wound, imprenaged with antimicrobial/antibacterial absorbent properties, must be have an open-pored sterile, freeze-dried matrix, must contains ionically bound silver, which provides antimicrobial protection against bacteria and infection. The absorbent collagen sheets manage wound moisture which must converts to soft gel sheet that stays in intimate contact with the wound bed, Three-dimensional structure helps support cellular growth Dressing must be composed of 55% collagen & and 45% oxidised regenerated cellulose Size: 28m3 Sterile, Individually wrapped Box of 10 dressings
151 | RT42-02-127 | Dressing, wound, construct of a gelling fibre, must have an absorbent wound covers, which combined, manage wound drainage. Dressing must be able to remove of dead skin, damage and infected tissue. Dressing must have an antimicrobial action with to clean. Dressing must be able to easy removal removes easily and reduce risk of skin damage. Dressing must be able to be used on treating of medium to highly fluid exudating wounds Individually wrapped Sterile Size: 10x10cm, Box of 10 dressings,
152 | RT42-02-128 | Dressing, wound, construct of a gelling fibre, must have an absorbent wound covers, which combined, manage wound drainage. Dressing must be able to remove of dead skin, damage and infected tissue. Dressing must have an antimicrobial action with to clean. Dressing must be able to easy removal removes easily and reduce risk of skin damage. Dressing must be able to be used on treating of medium to highly fluid exudating wounds Individually wrapped Sterile Size: 15x15cm, Box of 10 dressings,
153 | RT42-02-129 | Dressing, wound, construct of a gelling fibre, must have an absorbent wound covers, which combined, manage wound drainage. Dressing must be able to remove of dead skin, damage and infected tissue. Dressing must have an antimicrobial action with to clean. Dressing must be able to easy removal removes easily and reduce risk of skin damage. Dressing must be able to be used on treating of medium to highly fluid exudating wounds Individually wrapped Sterile Size: 20x30cm, Box of 10 dressings,
154 | RT42-02-130 | Dressing, wound, construct of a gelling fibre, must have an absorbent wound covers, which combined, manage wound drainage. Dressing must be able to remove of dead skin, damage and infected tissue. Dressing must have an antimicrobial action with to clean. Dressing must be able to easy removal removes easily and reduce risk of skin damage. Dressing must be able to be used on treating of medium to highly fluid exudating wounds Individually wrapped Sterile Size: 2x45cm, Box of 5 dressings,
155 | RT42-02-131 | Dressing, wound, non-adherent, fenestrated of highly absorbant hydrophyllic antimicrobial foam is impregnated with 0.3% polyhexamethylene biguanide 0.3%, Individually wrapped Sterile Size: ±900mm x ±750mm Box of 5 dressings
156 | RT42-02-132 | Dressing, wound, non-adherent, fenestrated of highly absorbant hydrophyllic antimicrobial foam is impregnated with 0.3% polyhexamethylene biguanide 0.3%, Individually wrapped Sterile Size: ±100mm x ±100mm Box of 5 dressings
157 | RT42-02-133 | Dressing, wound, non-adherent, fenestrated of highly absorbant hydrophyllic antimicrobial foam is impregnated with 0.3% polyhexamethylene biguanide 0.3%, Individually wrapped Sterile Size: ±100mm x ±200mm Box of 5 dressings
158 | RT42-02-134 | Dressing, wound, non-adherent, fenestrated of highly absorbant hydrophyllic antimicrobial foam is impregnated with 0.3% polyhexamethylene biguanide 0.3%, Individually wrapped Sterile Size: ±200mm x ±200mm Box of 5 dressings
159 | RT42-02-135 | Dressing, wound, non-adherent, fenestrated of highly absorbant hydrophyllic antimicrobial foam is impregnated with 0.3% polyhexamethylene biguanide 0.3%, Individually wrapped Sterile Size: ±50mm x ±50mm Box of 5 dressings
160 | RT42-02-136 | Dressing wound, negative pressure wound therapy system. Kit must include: *Dressing consist of a patented four layer dressing or sponge with a size: ±10cm x ±20cm *Pocket sized pump battery powered pump which must last for a maximum of 7 days Individually packed Single use Sterile Kit
161 | RT42-02-137 | Dressing wound, negative pressure wound therapy system. Kit must include: *Dressing consist of a patented four layer dressing or sponge with a size: ±10cm x ±30cm *Pocket sized pump battery powered pump which must last for a maximum of 7 days Individually packed Single use Sterile Kit
162 | RT42-02-138 | Dressing wound,negative pressure wound therapy system. Kit must include: *Dressing consist of a patented four layer dressing or sponge with a size: ±10cm x ±40cm *Pocket sized pump battery powered pump which must last for a maximum of 7 days Individually packed Single use Sterile Kit
163 | RT42-02-139 | Dressing wound, Single use negative pressure wound therapy system. Kit must include: *Dressing consist of a patented four layer dressing or sponge with a size: ±15cm x ±15cm *Pocket sized pump battery powered pump which must last for a maximum of 7 days Individually packed Single use Sterile Kit
164 | RT42-02-140 | Dressing wound, Single use negative pressure wound therapy system. Kit must include: *Dressing consist of a patented four layer dressing or sponge with a size: ±15cm x ±20cm *Pocket sized pump battery powered pump which must last for a maximum of 7 days Individually packed Single use Sterile Kit
165 | RT42-02-141 | Dressing, keliod and hypertropic scars, Size: ±100mm x180mm transparent medical grade silicone rubber sheet, soft adhesive ,washable and reusable. Reinforced with silicone membrane backing, conformable. Sterile, indivually peel- packed. Each
166 | RT42-02-142 | Dressing, keliod and hypertropic scars, Size: ±400mm x300mm transparent medical grade silicone rubber sheet, soft adhesive ,washable and reusable. Reinforced with silicone membrane backing, conformable. Sterile, indivually peel- packed. Each
167 | RT42-02-143 | Dressing, keliod and hypertropic scars, Size: ±500mm x750mm transparent medical grade silicone rubber sheet, soft adhesive ,washable and reusable. Reinforced with silicone membrane backing, conformable. Sterile, indivually peel- packed. Each
168 | RT42-02-144 | Dressing, incorporating ionic silver and calcium alginate with polyurethrane foam, Size: ±100mm x 100m , sterile, individually peel pack, Each
169 | RT42-02-145 | Dressing, incorporating ionic silver and calcium alginate with polyurethrane foam, Size: ±150mm x 150mm , sterile, individually peel pack, Each
170 | RT42-02-146 | Dressing, incorporating ionic silver and calcium alginate with polyurethrane foam, Size: ±200mm x 200mm , sterile, individually peel pack, Each
171 | RT42-02-147 | Dressing, incorporating ionic silver and calcium alginate with polyurethrane foam, Size: ±60mm x 60mm, sterile, individually peel pack, Each
172 | RT42-02-148 | Dressing, incorporating ionic silver and calcium alginate with polyurethrane foam, Size: ±80mm x 80mm , sterile, individually peel pack, Each
173 | RT42-02-149 | Self adhering foam designed to distribute body weight and prevent skin break down. Size:±12mm thick, (216 mm x 279mm).
174 | RT42-02-150 | Dressing, antimicrobial Size: ±85mm x 115mm . Semipermeable polyurethrane see through pad with outer adhesive border, contains Chlorhexidine Gluconate 2%.Sterile Individually Peel Packed, Each
175 | RT42-02-151 | Dressing, antimicrobial, Size: ±100mm x 120mm. Semipermeable polyurethrane see through pad with outer adhesive border, contains Chlorhexidine Gluconate 2%.Sterile Individually Peel Packed,
176 | RT42-02-152 | Dressing, antimicrobial , Size: ±100mm x 155mm . Semipermeable polyurethrane see through pad with outer adhesive border, contains Chlorhexidine Gluconate 2%.Sterile Individually Peel Packed..
177 | RT42-02-153 | Dressing Paediatric , IV Dressing, Transparent polyurethane film, water proof, moisture vapour permeable, unique frame delivery system, softcloth reinforcement for added secural, latex free, hypoallergenic, additional medipore cloth for added secural, pre-cut deep notch for catheter anchorage, transparent window for easy monitoring, single use dressing for 3-5 day wear time. . Size 5cm x 5.7cm.
178 | RT42-02-154 | Dressing , IV for peripheral Line, Transparent polyurethane film, water proof, moisture vapour permeable, unique frame delivery system, softcloth reinforcement for added secural, latex free, hypoallergenic, additional medipore cloth for added secural, pre-cut deep notch for catheter anchorage, transparent window for easy monitoring, single use dressing for 3-5 day wear time. . Size 70mm x 85mm.
179 | RT42-02-155 | Dressing , IV for CVC Line, Transparent polyurethane film, water proof, moisture vapour permeable, unique frame delivery system, softcloth reinforcement for added secural, latex free, hypoallergenic, additional medipore cloth for added secural, pre-cut deep notch for catheter anchorage, transparent window for easy monitoring, single use dressing for 3-5 day wear time. Size 85mm x 105mm.
180 | RT42-02-156 | Dressing , Universal IV . Transparent polyurethane film, water proof, moisture vapour permeable, unique frame delivery system, softcloth reinforcement for added secural, latex free, hypoallergenic, additional medipore cloth for added secural, pre-cut deep notch for catheter anchorage, transparent window for easy monitoring, single use dressing for 3-5 day wear time. Size 8.9cm x 11.5cm.
181 | RT42-02-157 | Dressing wound, non-adherent, Size: ±105mm x, 105mm, absorbent, contains 100% pure activated charcoal, sterile,individually peel packed. Each
182 | RT42-02-158 | Dressing Foam: Adhesive nanocrystalline silver antimicrobial barrier; release at a rate of 72-100 ppm over a period of 3 days Size: ±100mm x 100mm, Sterile, Individually peel packed. Each
183 | RT42-02-159 | Dressing Foam: Adhesive nanocrystalline silver antimicrobial barrier; release at a rate of 72-100 ppm over a period of 3 days Size: ±100mm x 200mm, Sterile, Individually peel packed. Each
184 | RT42-02-160 | Dressing Foam: Adhesive nanocrystalline antimicrobial barrier; Size: ±200mm x 400mm, Sustained silver release for 3 days at a rate of 70-100ppm; Sterile.
185 | RT42-02-161 | Dressing Foam: Adhesive nanocrystalline silver antimicrobial barrier; release at a rate of 72-100 ppm over a period of 3 days Size: ±400mm x 400mm, Sterile, Individually peel packed. Each
186 | RT42-02-162 | Dressing Foam: Adhesive nanocrystalline antimicrobial barrier; Size: ±100mm x 120mm, Sustained silver release for 7 days at a rate of 70-100ppm; Sterile.
187 | RT42-02-163 | Dressing, polyurethrane foam, white: Size: ±590mm x 490mm x 10mm, non-adherent, non adhesive, sterile, Individually peel packed, Each
188 | RT42-02-164 | Dressing, polyurethrane foam, white: Size: ±590mm x 490mm x 20mm, non-adherent, non adhesive, sterile, Individually peel packed, Each
189 | RT42-02-165 | Dressing, polyurethrane foam, white: Size: ±55mm x 6mm x 300mm, non-adherent, non adhesive, sterile, Individually peel packed, Each
190 | RT42-02-166 | Dressing, elastic expanded polyurethrane foam, Size: with allowance±550mm x 450mm x 20mm, : non-adherent, non adhesive,sterile, autoclavable.
191 | RT42-02-167 | Dressing, polyurethrane foam, white: Size: ±50mm x 3mm x 300mm, non-adherent, non adhesive, sterile, Individually peel packed, Each
192 | RT42-02-168 | Dressing, Acrylic:Size: ±76mm x 95mm, allows for moist wound healing promoting autolytic debridement , absorbent, clear, unique construction - dressing made up of two layers of transparent film sandwitched between a layer of acrylic polymer, self adherent, latex free, sterile,, moisture vapour permeable, hypoallergenic, individually packed, 5 dressings per box.
193 | RT42-02-169 | Dressing, Acrylic:Size: ±150mm x 150mm, allows for moist wound healing promoting autolytic debridement , absorbent, clear, unique construction - dressing made up of two layers of transparent film sandwitched between a layer of acrylic polymer, self adherent, latex free, sterile,, moisture vapour permeable, hypoallergenic, individually packed, 5 dressings per box.
194 | RT42-02-170 | Dressing, Acrylic:Size: ±170mm x 190mm, allows for moist wound healing promoting autolytic debridement , absorbent, clear, unique construction - dressing made up of two layers of transparent film sandwitched between a layer of acrylic polymer, self adherent, latex free, sterile,, moisture vapour permeable, hypoallergenic, individually packed, 5 dressings per box.
195 | RT42-02-171 | Dressing, Acrylic:Size: ±200mm x 200mm, allows for moist wound healing promoting autolytic debridement , absorbent, clear, unique construction - dressing made up of two layers of transparent film sandwitched between a layer of acrylic polymer, self adherent, latex free, sterile,, moisture vapour permeable, hypoallergenic, individually packed, 5 dressings per box.
196 | RT42-02-172 | Dressing, occlusive, adhesive, Size: ±80mm x 100mm, transparent polyurethrane-based synthetic, hydrophilic layer adheres without adhesive, sterile, Indivicually peel packed.
197 | RT42-02-173 | Dressing, occlusive, adhesive, Size: ±210mm x 310mm, transparent polyurethrane-based synthetic, hydrophilic layer adheres w/out adhesive, sterile, Indivicually peel packed.
198 | RT42-02-174 | Dressing, occlusive, adhesive, Size: with allowance of 5 mm, 50mm x 70mm, transparent polyurethrane-based synthetic, hydrophilic layer adheres without adhesive, sterile, Indivicually peel packed.
199 | RT42-02-175 | Dressing Pack: Surgical: Tray must be packed as follows 2x hand towels 1x red plastic bag 1x Pair of latex free gloves, size Medium, non powdered 2x Disposable soft drape 90x90cm Cleansing Swabs 30mm - 1 bundle of 5 balls 10 X Plain gauze 100mm x 100mm 8ply 2x Disposable forceps Sterile, peel packed , Comply with the latest SANS 446 Each
200 | RT42-02-176 | Dressing Pack: Maternity Tray must be packed as follows: 2x hand towels 1x red plastic bag 1x Pair of latex free gloves, size Medium, non powdered Cleansing Swabs 30mm - 1 bundle of 5 balls 10 X Plain gauze 100mm x 100mm 8ply 2 x Maternity sanitary towels loops Sterile, peel packed , Comply with the latest SANS 446 Each
201 | RT42-02-177 | Dressing, Acrylic:Size: ±110mm x 130mm, allows for moist wound healing promoting autolytic debridement , absorbent, clear, unique construction - dressing made up of two layers of transparent film sandwitched between a layer of acrylic polymer, self adherent, latex free, sterile, transparent film to be barrier to contamination. moisture vapour permeable, hypoallergenic, individually packed, 5 dressings per box.
202 | RT42-02-178 | Dressing finger and toe, small, 70% polyamide, Polyethersulfoae (PES); 21% elastane/lycra, Each.
203 | RT42-02-179 | Dressing finger and toe, medium, 70% polyamide, Polyethersulfoae (PES); 21% elastane/lycra, Each.
204 | RT42-02-180 | Dressing finger and toe, large, 70% polyamide, Polyethersulfoae (PES); 21% elastane/lycra, Each.
205 | RT42-02-181 | Dressing finger and toe, x-large, 70% polyamide, Polyethersulfoae (PES); 21% elastane/lycra, Each.
206 | RT42-02-182 | Dressing finger and toe, Size: 750mm x 1500mm, 70% polyamide,Polyethersulfoae ( PES); 21% elastane/lycra, Each, roll.
207 | RT42-03-001 | Dressing pad, absorbent, acetate fabric with antimicrobial hydrophobic coating (e.g. Dialkyl Carbamoyl Chloride) that physically binds and removes bacteria, size 70mm x 90mm, packaging individually wrapped, sterile. Each
208 | RT42-03-002 | Dressing pad, absorbent, acetate fabric with antimicrobial hydrophobic coating (e.g. Dialkyl Carbamoyl Chloride) that physically binds and removes bacteria, size 100mm x 100mm, packaging individually wrapped, sterile.
209 | RT42-03-003 | Dressing pad, absorbent, acetate fabric with antimicrobial hydrophobic coating (e.g. Dialkyl Carbamoyl Chloride) that physically binds and removes bacteria, size 100mm x 200mm, packaging individually wrapped, sterile.
210 | RT42-03-004 | Dressing, swab, acetate fabric with antimicrobial hydrophobic coating (e.g. Dialkyl Carbamoyl Chloride)that physically binds and removes bacteria, size 40mm x 60mm x 8 ply, packaging individually wrapped, sterile.
211 | RT42-03-005 | Dressing, swab, acetate fabric with antimicrobial hydrophobic coating (e.g. Dialkyl Carbamoyl Chloride) that physically binds and removes bacteria, size 70mm x90mm x 8 ply, packaging individually wrapped, sterile.
212 | RT42-03-006 | Dressing, ribbon gauze, cotton fabric with antimicrobial hydrophobic coating (e.g. Dialkyl Carbamoyl Chloride) that physically binds and removes bacteria, size 20mm x 500mm, packaging: roll, individually wrapped, sterile.
213 | RT42-03-007 | Dressing, ribbon gauze, cotton fabric with antimicrobial hydrophobic coating (e.g. dialkyl carbamoyl chloride)that physically binds and removes bacteria, size 50mm x 2000mm, packaging: roll, individually wrapped, sterile.
214 | RT42-03-008 | Dressing, absorbent hydropolymer core, acetate fabric with antimicrobial hydrophobic coating (e.g. Dialkyl Carbamoyl Chloride)that physically binds and removes bacteria, size 70mm x 85mm, packaging individually wrapped, sterile.
215 | RT42-03-009 | Dressing, absorbent hydropolymer core, acetate fabric with antimicrobial hydrophobic coating (e.g. Dialkyl Carbamoyl Chloride)that physically binds and removes bacteria, size 140mm x 140mm, packaging individually wrapped, sterile.
216 | RT42-03-010 | Dressing, absorbent hydropolymer core, acetate fabric with antimicrobial hydrophobic coating (e.g. Dialkyl Carbamoyl Chloride)that physically binds and removes bacteria, size 190mm x 190mm, packaging individually wrapped, sterile.
217 | RT42-03-011 | Dressing, absorbent hydropolymer core, acetate fabric with antimicrobial hydrophobic coating (e.g. Dialkyl Carbamoyl Chloride)that physically binds and removes bacteria, size 240mm x 240mm, packaging individually wrapped, sterile.
218 | RT42-03-012 | Dressing, absorbent hydropolymer core, acetate fabric with antimicrobial hydrophobic coating (e.g. Dialkyl Carbamoyl Chloride)that physically binds and removes bacteria, size 140mm x 240mm, packaging individually wrapped, sterile.
219 | RT42-03-013 | Dressing, self-adhesive, absorbent hydropolymer core, acetate fabric with antimicrobial hydrophobic coating (e.g. Dialkyl Carbamoyl Chloride) that physically binds and removes bacteria, size 70mm x 85mm, packaging individually wrapped, sterile.
220 | RT42-03-014 | Dressing, self-adhesive, absorbent hydropolymer core, acetate fabric with antimicrobial hydrophobic coating (e.g. Dialkyl Carbamoyl Chloride) that physically binds and removes bacteria, size 140mm x 140mm, packaging individually wrapped, sterile.
221 | RT42-03-015 | Dressing, self-adhesive, absorbent hydropolymer core, acetate fabric with antimicrobial hydrophobic coating (e.g. Dialkyl Carbamoyl Chloride) that physically binds and removes bacteria, size 190mm x 190mm, packaging individually wrapped, sterile.
222 | RT42-03-016 | Dressing, self-adhesive, absorbent hydropolymer core, acetate fabric with antimicrobial hydrophobic coating (e.g. Dialkyl Carbamoyl Chloride) that physically binds and removes bacteria, size 140mm x 240mm, packaging individually wrapped, sterile.
223 | RT42-03-017 | Dressing, self-adhesive, highly absorbent polyurethane foam core, acetate fabric with antimicrobial hydrophobic coating (e.g. Dialkyl Carbamoyl Chloride) that physically binds and removes bacteria, size 75mm x 75mm, packaging individually wrapped, sterile.
224 | RT42-03-018 | Dressing, self-adhesive, highly absorbent polyurethane foam core, acetate fabric with antimicrobial hydrophobic coating (e.g. Dialkyl Carbamoyl Chloride) that physically binds and removes bacteria, size 125mm x 125mm, packaging individually wrapped, sterile.
225 | RT42-03-019 | Dressing, self-adhesive, highly absorbent polyurethane foam core, acetate fabric with antimicrobial hydrophobic coating (e.g. Dialkyl Carbamoyl Chloride) that physically binds and removes bacteria, size 150mm x 150mm, packaging individually wrapped, sterile.
226 | RT42-03-020 | Dressing, self-adhesive, highly absorbent polyurethane foam core, acetate fabric with antimicrobial hydrophobic coating (e.g. Dialkyl Carbamoyl Chloride) that physically binds and removes bacteria, size 175mm x 175mm, packaging individually wrapped, sterile.
227 | RT42-03-021 | Dressing, self-adhesive, highly absorbent polyurethane foam core, acetate fabric with antimicrobial hydrophobic coating (e.g. Dialkyl Carbamoyl Chloride) that physically binds and removes bacteria, size 225 x 225cm, packaging individually wrapped, sterile.
228 | RT42-03-022 | Dressing wound, hydrokinetic, super absorbent, layered mechanically modified cellulose and gel forming polymerdressing to assist in the removal of exudate, bacteria and dead tissue, be non-adherent to wound surface, sterile, individually peel packed, size dressing: ± 7.5 x 7.5cm.
229 | RT42-03-023 | Dressing wound, hydrokinetic, super absorbent, layered mechanically modified cellulose and gel forming polymerdressing to assist in the removal of exudate, bacteria and dead tissue, be non-adherent to wound surface, sterile, individually peel packed, size dressing: ± 12 x 5cm.
230 | RT42-03-024 | Dressing wound, hydrokinetic, super absorbent, layered mechanically modified cellulose and gel forming polymerdressing to assist in the removal of exudate, bacteria and dead tissue, be non-adherent to wound surface, sterile, individually peel packed, size dressing: ± 10 x 10cm.
231 | RT42-03-025 | Dressing wound, hydrokinetic, super absorbent, layered mechanically modified cellulose and gel forming polymerdressing to assist in the removal of exudate, bacteria and dead tissue, be non-adherent to wound surface, sterile, individually peel packed, size dressing: ± 20 x 10cm.
232 | RT42-03-026 | Dressing wound, hydrokinetic, super absorbent, layered mechanically modified cellulose and gel forming polymerdressing to assist in the removal of exudate, bacteria and dead tissue, be non-adherent to wound surface, sterile, individually peel packed, size dressing: ± 15 x 15cm
233 | RT42-03-027 | Dressing wound, hydrokinetic, super absorbent, layered mechanically modified cellulose and gel forming polymerdressing to assist in the removal of exudate, bacteria and dead tissue, be non-adherent to wound surface, sterile, individually peel packed, size dressing: ± 20 x 20cm.
234 | RT42-03-028 | Dressing wound, hydrokinetic, super absorbent, layered mechanically modified cellulose and gel forming polymerdressing to assist in the removal of exudate, bacteria and dead tissue, be non-adherent to wound surface, sterile, individually peel packed, size dressing: ± 30 x 20cm
235 | RT42-03-029 | Dressing wound, hydrokinetic, super absorbent, layered mechanically modified cellulose and gel forming polymerdressing to assist in the removal of exudate, bacteria and dead tissue, be non-adherent to wound surface, sterile, individually peel packed, size dressing: ± round 8cm diameter.
236 | RT42-03-030 | Dressing wound, hydrokinetic, super absorbent, layered mechanically modified cellulose and gel forming polymerdressing to assist in the removal of exudate, bacteria and dead tissue, be non-adherent to wound surface, sterile, individually peel packed, size dressing: ± round 4cm diameter.
237 | RT42-03-031 | Dressing wound, self-adherent, hydrokinetic, super absorbent, layered mechanically modified cellulose and gel forming polymerdressing to assist in the removal of exudate, bacteria and dead tissue, be non-adherent to wound surface, sterile. individually peel packed. size dressing: ± 10 x 10cm.
238 | RT42-03-032 | Dressing wound, self-adherent, hydrokinetic, super absorbent, layered mechanically modified cellulose and gel forming polymerdressing to assist in the removal of exudate, bacteria and dead tissue, be non-adherent to wound surface, sterile. individually peel packed. size dressing: ± 15 x 15cm3.
239 | RT42-03-033 | Dressing wound, self-adherent, hydrokinetic, super absorbent, layered mechanically modified cellulose and gel forming polymerdressing to assist in the removal of exudate, bacteria and dead tissue, be non-adherent to wound surface, sterile. individually peel packed. size dressing: ± 15 x 25cm.
240 | RT42-03-034 | Dressing Surgical set: tray ,swabbing Tray must be packed as follows: 2x paper towels, 1x red plastic (280mm x400mm), 2x powder free latex gloves (medium), 1x protection sheet, 10x cotton wool balls 2g, 1x sanitary towel looped, 2x gallipots (75ml). Sterile, peel packed , Comply with the latest SANS 446 Each
241 | RT42-03-035 | Dressing tray sterile. Tray must be packed as follows, 2x paper towels, 1x red plastic (280mm x 400mm), 2x powder free latex gloves (medium), 1x protection sheet, 5 x gauze swabs (10mm x 100mm x 8 ply), 2 x gallipots (75ml), 1 x forcepts (plastic).Sterile, peel packed , Comply with the latest SANS 446 Each
242 | RT42-03-036 | Dressing, transparent film, post-operative/wound, adhesive, hypoallergenic, allows visibility of incision site, Sterile with a central see through absorbent pad, on-adherent pad will not disturb staples, sutures or surgical glue, Sterile, individually peel packed with description peel packed with description on outer pack Size Dressing: + 80mm x 100mm Each
243 | RT42-03-037 | Dressing, transparent film, post-operative/wound, adhesive, hypoallergenic, allows visibility of incision site, Sterile with a central see through absorbent pad, on-adherent pad will not disturb staples, sutures or surgical glue, Sterile, individually peel packed with description peel packed with description on outer pack Size Dressing: + 150mm x 100mm Each
244 | RT42-03-038 | Dressing, transparent film, post-operative/wound, adhesive, hypoallergenic, allows visibility of incision site, Sterile with a central see through absorbent pad, on-adherent pad will not disturb staples, sutures or surgical glue, Sterile, individually peel packed with description peel packed with description on outer pack Size Dressing: + 200mm x 100mm Each
245 | RT42-03-039 | Dressing, transparent film, post-operative/wound, adhesive, hypoallergenic, allows visibility of incision site, Sterile with a central see through absorbent pad, on-adherent pad will not disturb staples, sutures or surgical glue, Sterile, individually peel packed with description peel packed with description on outer pack Size Dressing: + 250mm x 100mm Each
246 | RT42-03-040 | Dressing, transparent film, post-operative/wound, adhesive, hypoallergenic, allows visibility of incision site, Sterile with a central see through absorbent pad, on-adherent pad will not disturb staples, sutures or surgical glue, Sterile, individually peel packed with description peel packed with description on outer pack Size Dressing: + 350mm x 100mm Each
247 | RT42-04-001 | Bandage, compression, flexible 2-layer co-hesive elastic wrap, non-adhesive, self-adherent, "latex free". Inner layer 100mm x 2.7m unstretched. Compression layer 100mm x 4.5m fully stretched and 100mm x 3.5m unstretched. Consisting of porous, non-woven, polyester. Size: 100mm x 2.7m Individually wrapped Roll
248 | RT42-04-002 | Bandage, compression, flexible 2-layer co-hesive elastic wrap, non-adhesive, self-adherent, "latex free". Inner layer 75mm x 2.7m unstretched. Compression layer 75mm x 4.5m fully stretched and 75mm x 3.5m unstretched. Consisting of porous, non-woven, polyester. Size: 75mm x 2.7m Roll
249 | RT42-04-003 | Bandage, porous, conforming, surgical. To consist with a woven cotton cloth backed with a porous, zinc oxide filled hypoallergenic adhesive, B.P. Bandage to have strong support with moderate compression Size: 25 mm x not less than 3m stretched Individually wrapped Roll
250 | RT42-04-004 | Bandage, porous, conforming, surgical. To consist with a woven cotton cloth backed with a porous, zinc oxide filled hypoallergenic adhesive, B.P. Bandage to have strong support with moderate compression Size: 50mm x not less than 2.7m unstretched, 4.5m stretched Individually wrapped Roll
251 | RT42-04-005 | Bandage, porous, conforming, surgical. To consist with a woven cotton cloth backed with a porous, zinc oxide filled hypoallergenic adhesive, B.P. Bandage to have strong support with moderate compression Size: 75mm x not less than 2.7m unstretched, 4.5m stretched Individually wrapped Roll
252 | RT42-04-006 | Bandage, porous, conforming, surgical. To consist with a woven cotton cloth backed with a porous, zinc oxide filled hypoallergenic adhesive, B.P. Bandage to have strong support with moderate compression Size: 100mm x not less than 2.7m unstretched, 4.5m stretched Individually wrapped Roll
253 | RT42-04-007 | Bandage, elastic, adhesive, porous, conforming, surgical. To consist of a woven cotton cloth backed with a porous, zinc oxide filled hypoallergenic adhesive, B.P. Bandage to have strong support, with moderate compression Size: 50mm x 4.5m stretched Laminated to an easy release silicone backing paper Individually wrapped Roll
254 | RT42-04-008 | Bandage, elastic, adhesive, porous, conforming, surgical. To consist of a woven cotton cloth backed with a porous, zinc oxide filled hypoallergenic adhesive, B.P. Bandage to have strong support, with moderate compression Size: 75mm x 4.5m stretched Laminated to an easy release silicone backing paper Individually wrapped Roll
255 | RT42-04-009 | Bandage, elastic, adhesive, porous, conforming, surgical. To consist of a woven cotton cloth backed with a porous, zinc oxide filled hypoallergenic adhesive, B.P. Bandage to have strong support, with moderate compression Size: 100mm x 4.5m stretched Laminated to an easy release silicone backing paper Individually wrapped Roll
256 | RT42-04-010 | Bandage, retention, white, open wove, type 1 Size: 75mm x 5m Individually wrapped To comply with the latest issue of SANS 443 Compliance certificate to be submitted with bid Roll
257 | RT42-04-011 | Bandage, retention, white, open wove, type 1 Size: 100mm x 5m Individually wrapped To comply with the latest issue of SANS 443 Compliance certificate to be submitted with bid Roll
258 | RT42-04-012 | Bandage, tubular, elastic, mesh For light non-compression, retention of dressings. Size: 6 Unstretched length: 25m Unstretched width: +-80mm To comply with the latest issue of CKS330 Compliance certificate to be submitted with bid Roll, individually boxed
259 | RT42-04-013 | Bandage, tubular, orthopaedic, knitted, cotton-type, stockinette, unbleached Size: width: 50mm Size: length: 25m To comply with the latest issue of SANS 925 Compliance certificate to be submitted with bid Roll, individually boxed
260 | RT42-04-014 | Bandage, tubular, orthopaedic, knitted, cotton-type, stockinette, unbleached Size:width: 80mm Size: length: 25m To comply with the latest issue of SANS 925 Compliance certificate to be submitted with bid Roll, individually boxed
261 | RT42-04-015 | Bandage, tubular, orthopaedic, knitted, cotton-type, stockinette, unbleached Size:width: 100mm Size: length: 25m To comply with the latest issue of SANS 925 Compliance certificate to be submitted with bid Roll, individually boxed
262 | RT42-04-016 | Bandage, tubular, orthopaedic, knitted, cotton-type, stockinette, unbleached Size:width: 150mm Size: ength: 25m To comply with the latest issue of SANS 925 Compliance certificate to be submitted with bid Roll, individually boxed
263 | RT42-04-017 | Bandage, tubular, orthopaedic, knitted, cotton-type, stockinette, unbleached Size:width: 200mm Size: length: 25m To comply with the latest issue of SANS 925 Compliance certificate to be submitted with bid Roll, individually boxed
264 | RT42-04-018 | Bandage, tubular, orthopaedic, knitted, cotton-type, stockinette, unbleached Size: width: 250mm Size: length: 25m To comply with the latest issue of SANS 925 Compliance certificate to be submitted with bid Roll, individually boxed
265 | RT42-04-019 | Bandage Tubular Size: C: 6.75cm x 10m Elasticised for firm support and moderate compression Light weight, autoclavable. Well ventilated, crease free with elastic yarns protected Must conform to body contours. Must be washable and reusable Able to be applied by hand without the use of an applicator Must include a measuring device To comply with the latest issue of CKS641 Compliance certificate to be submitted with bid Roll, individually boxed
266 | RT42-04-020 | Bandage, retention, crepe type, autoclavable Size: 50mm x not less than 4.5m stretched Individually wrapped To comply with the latest issue of SANS 1427, Table 2 Compliance certificate to be submitted with bid Roll
267 | RT42-04-021 | Bandage, retention, crepe type, autoclavable Size: 75mm x not less than 4.5m stretched Individually wrapped To comply with the latest issue of SANS 1427, Table 2 Compliance certificate to be submitted with bid Roll
268 | RT42-04-022 | Bandage, retention, crepe type, autoclavable Size: 100mm x not less than 4.5m stretched Individually wrapped To comply with the latest issue of SANS 1427, table 2 Compliance certificate to be submitted with bid Roll
269 | RT42-04-023 | Bandage, retention, crepe type, autoclavable Size: 150mm x not less than 4.5m stretched Individually wrapped To comply with the latest issue of SANS 1427, table 2 Compliance certificate to be submitted with bid Roll
270 | RT42-04-024 | Bandage, heavy support/ compression, cotton crepe type 1 Size: 75mm x not less than 4.5m stretched Individually wrapped To comply with the latest issue of SANS 945, part 1 Compliance certificate to be submitted with bid Roll
271 | RT42-04-025 | Bandage, heavy support/ compression, cotton type, type 1 Size: 100mm x not less than 4.5m stretched Individually wrapped To comply with the latest issue of SANS 945, part 1 Compliance certificate to be submitted with bid Roll
272 | RT42-04-026 | Bandage, light support, cotton crepe type 2, autoclavable Size: 50mm x not less than 4.5m stretched Individually wrapped To comply with the latest issue of SANS 945, Part 1 Compliance certificate to be submitted with bid Roll
273 | RT42-04-027 | Bandage, light support, cotton crepe type 2, autoclavable Size: 75mm x not less than 4.5m stretched Individually wrapped To comply with the latest issue of SANS 945, Part 1 Compliance certificate to be submitted with bid Roll
274 | RT42-04-028 | Bandage, light support, cotton crepe type 2, autoclavable Size: 100mm x not less than 4.5m stretched Individually wrapped To comply with the latest issue of SANS 945, Part 1 Compliance certificate to be submitted with bid Roll
275 | RT42-04-029 | Bandage, light support, cotton crepe type 2, autoclavable Size: 150mm x not less than 4.5m stretched Individually wrapped To comply with the latest issue of SANS 945, Part 1 Compliance certificate to be submitted with bid Roll
276 | RT42-04-030 | Bandage, support, triangular, calico, woven type, hemmed Size: 900mm x 900mm x 1 300mm Individually wrapped Each
277 | RT42-04-031 | Bandage, support, triangular, non-woven, polypropylene Size: 900mm x 900mm x 1 300mm Minimum mass: 17 grams per square metre Individually wrapped To comply with the latest issue of SANS 79 Compliance certificate to be submitted with bid Each
278 | RT42-04-032 | Bandage, cast, plaster of paris Leno weave base cloth, type 1 Size: 50mm x 3 m Individually wrapped and packed in a sealed container Comply with the latest issue of SANS 1308, Table 1 and 2 Compliance certificate to be submitted with bid Roll
279 | RT42-04-033 | Bandage, cast, plaster of paris Leno weave base cloth, type 1 Size: 75mm x 3 m Individually wrapped and packed in a sealed container Comply with the latest issue of SANS 1308, Table 1 and 2 Compliance certificate to be submitted with bid Roll
280 | RT42-04-034 | Bandage, cast, plaster of paris Leno weave base cloth, type 1 Size: 100mm x 3.5m Individually wrapped and packed in a sealed container To comply with the latest issue of SANS 1308, Table 1 and 2 Compliance certificate to be submitted with bid Roll
281 | RT42-04-035 | Bandage, cast, plaster of paris Leno weave base cloth, type 1 Size: 150mm x 3.5m Individually wrapped and packed in a sealed container To comply with the latest issue of SANS 1308, Table 1 and 2 Compliance certificate to be submitted with bid Roll
282 | RT42-04-036 | Bandage, cast, plaster of paris Leno weave base cloth, type 1 Size: 200mm x 3.5m Individually wrapped and packed in a sealed container To comply with the latest issue of SANS 1308, Table 1 and 2 Compliance certificate to be submitted with bid Roll
283 | RT42-04-037 | Bandage, cast, plaster of paris impregnated slabs, plain woven base cloth Size: 75 mm x 750 mm Slabs wrapped in plastic, in a sealed container To comply with the latest issue of SANS 1308, Table 1 and 2 Compliance certificate to be submitted with bid Box of 50 strips
284 | RT42-04-038 | Bandage, cast, plaster of paris impregnated slabs, plain woven base cloth Size: 100 mm x 750 mm Slabs wrapped in plastic, in a sealed container To comply with the latest issue of SANS 1308, Table 1 and 2 Compliance certificate to be submitted with bid Box of 50 strips
285 | RT42-04-039 | Bandage, cast, plaster of paris impregnated slabs, plain woven base cloth Size: 150 mm x 750 mm Slabs wrapped in plastic, in a sealed container To comply with the latest issue of SANS 1308, Table 1 and 2 Compliance certificate to be submitted with bid Box of 50 strips
286 | RT42-04-040 | Bandage casting waterproof, consisting of fibre-glass mix Light weight, fast setting, early load bearing, smooth finish Individually wrapped Size: 50mm x not less than 3.6m Roll
287 | RT42-04-041 | Bandage casting waterproof, consisting of fibre-glass mix Light weight, fast setting, early load bearing, smooth finish Individually wrapped Size: 75mm x not less than 3.6m Roll
288 | RT42-04-042 | Bandage casting waterproof, consisting of fibre-glass mix Light weight, fast setting, early load bearing, smooth finish Individually wrapped Size: 100mm x not less than 3.6m Roll
289 | RT42-04-043 | Bandage casting waterproof, consisting of fibre-glass mix Light weight, fast setting, early load bearing, smooth finish Individually wrapped Size: 100mm x not less than 3.6m Roll
290 | RT42-04-044 | Bandage casting waterproof, consisting of fibre-glass mix Light weight, fast setting, early load bearing, smooth finish Individually wrapped Size: 125mm x not less than 3.6m Roll
291 | RT42-04-045 | Bandage, conforming knitted type Size: 75mm x not less than 4.5m stretch Table 3, KC30 Individually wrapped To comply with the latest issue of SANS 1427 Compliance certificate to be submitted with bid Roll
292 | RT42-05-001 | Cotton bud, double tipped on rolled paper stick, autoclavable Length: Not less than 80mm Container of 100 buds
293 | RT42-05-002 | Cotton wool, absorbent, first grade, non-interleaved, 50g To comply with the latest issue of SANS 228 Compliance certificate to be submitted with bid Roll
294 | RT42-05-003 | Cotton wool, absorbent, second grade, interleaved, 500g To comply with the latest issue of SANS 228 Compliance certificate to be submitted with bid Roll
295 | RT42-05-004 | Cotton wool, absorbent balls, type 2, non-sterile, autoclavable, 500g package Size: Medium 1g To comply with the latest issue of SANS 149 Compliance certificate to be submitted with bid Packet
296 | RT42-05-005 | Cotton wool, absorbent balls, type 2, autoclavable, white, sterile, 5 balls per pouch, peel packed Size: Medium 1g To comply with the latest issue of SANS 149 Compliance certificate to be submitted with bid Pouch
297 | RT42-05-006 | Cotton wool, absorbent balls, type 2, autoclavable, 500g package Size: Maternity 2g To comply with the latest issue of SANS 149 Compliance certificate to be submitted with bid Packet
298 | RT42-05-007 | Cotton wool, absorbent balls, type 2, autoclavable, white, sterile, 5 balls per pouch, peel packed Size: Maternity 2g To comply with the latest issue of SANS 149 Compliance certificate to be submitted with bid Pouch
299 | RT42-05-008 | Tampon, vaginal, autoclavable Absorbent cotton wool with gauze made into a tube Cotton wool to comply with the latest issue of SANS 228 Absorbent cotton gauze to comply with the latest issue of SANS 446 Compliance certificate to be submitted with bid Each
300 | RT42-06-001 | Swabs, alcohol, 200 swabs Sterile, composed of viscose rayon non-woven fabric impregnated with minimum of 0,5ml Isopropyl Alcohol 70% in water Minimum dimensions: 24mm x 30mm Individually packed in aluminium foil pouches SAHPRA product registration to be submitted Box of 200 swabs
301 | RT42-06-002 | Swab, abdominal, surgical, cotton, X-ray detectable, sterile, 8g, 4-ply Size: 170mm x 200mm Pack size: Pouch of 5 swabs Scope • Swabs for single use, all layers consist of white woven bandage cloth, to comply with the latest ;issue of SANS 443, type 3 Construction • Edges: All cut edges of the material to be folded in and stitched, leaving no frayed edges Cotton tape • Length of tape: 400mm • Width of tape: 12mm • The cotton tape must be embedded in a 50mm packet i.e. 100mm in length of tape folder double to 50mm stitched into the corner of the swab leaving a 300mm length of cotton tape clear • Breaking strength of tape – minimum 7.95kg X-Ray detectable material • X-Ray detectable material thread shall be woven in/impregnated Packaging • The swab shall be bound in bundles of five by means of either a paper band consisting of one piece secured by autoclavable glue (and not by a separate adhesive tape) or by means of an acceptable tying material which can withstand autoclaving. • The swabs shall be suitable wrapped so as to be adequately protected against soiling and contamination • Only swabs of the same size shall be packed together in a bulk container. • Packing shall comply with the relevant requirements of the latest issue of CKS 228- double swab • Swabs shall be packed in a red plastic bag  Size 280mm x 400mm, thickness at least 19micron, colour: red, CKS 559 Marking • The following information shall appear in legible and indelible marking on each package:  Manufacture’s name or trade mark or both  Type of swabs  Size of swabs  Quantity of swabs Bulk containers • The following information shall appear in legible and indelible marking on the outside of each bulk container. Base – Peel pouches • Clear PVC film, not less than 90u Lidding – Peel pouches / tray • Peel cover type (heat seal grid), 60g/m² medical Kraft paper lacquered with 11u adhesive Sterility indicator • Pack to contain tamperproof sterility indicator on the tray or pack, confirm sterility, bidder to indicate the type of indicator used Sterile Peel packed in bundles of 5 Pack size: Each
302 | RT42-06-003 | Swab, abdominal, surgical, cotton, X-ray detectable, sterile, 8g, 6-ply Size: 170mm x 200mm Pack size: Pouch of 5 swabs Scope • Swabs for single use, all layers consist of white woven bandage cloth, to comply with the latest ;issue of SANS 443, type 3 Construction • Edges: All cut edges of the material to be folded in and stitched, leaving no frayed edges Cotton tape • Length of tape: 400mm • Width of tape: 12mm • The cotton tape must be embedded in a 50mm packet i.e. 100mm in length of tape folder double to 50mm stitched into the corner of the swab leaving a 300mm length of cotton tape clear • Breaking strength of tape – minimum 7.95kg X-Ray detectable material • X-Ray detectable material thread shall be woven in/impregnated Packaging • The swab shall be bound in bundles of five by means of either a paper band consisting of one piece secured by autoclavable glue (and not by a separate adhesive tape) or by means of an acceptable tying material which can withstand autoclaving. • The swabs shall be suitable wrapped so as to be adequately protected against soiling and contamination • Only swabs of the same size shall be packed together in a bulk container. • Packing shall comply with the relevant requirements of the latest issue of CKS 228- double swab • Swabs shall be packed in a red plastic bag  Size 280mm x 400mm, thickness at least 19micron, colour: red, CKS 559 Marking • The following information shall appear in legible and indelible marking on each package:  Manufacture’s name or trade mark or both  Type of swabs  Size of swabs  Quantity of swabs Bulk containers • The following information shall appear in legible and indelible marking on the outside of each bulk container. Base – Peel pouches • Clear PVC film, not less than 90u Lidding – Peel pouches / tray • Peel cover type (heat seal grid), 60g/m² medical Kraft paper lacquered with 11u adhesive Sterility indicator • Pack to contain tamperproof sterility indicator on the tray or pack, confirm sterility, bidder to indicate the type of indicator used Sterile Peel packed in bundles of 5 Pack size: Each
303 | RT42-06-004 | Swab, abdominal, surgical, cotton, X-ray detectable, sterile, 8g, 8-ply Size: 170mm x 200mm Pack size: Pouch of 5 swabs Scope • Swabs for single use, all layers consist of white woven bandage cloth, to comply with the latest ;issue of SANS 443, type 3 Construction • Edges: All cut edges of the material to be folded in and stitched, leaving no frayed edges Cotton tape • Length of tape: 400mm • Width of tape: 12mm • The cotton tape must be embedded in a 50mm packet i.e. 100mm in length of tape folder double to 50mm stitched into the corner of the swab leaving a 300mm length of cotton tape clear • Breaking strength of tape – minimum 7.95kg X-Ray detectable material • X-Ray detectable material thread shall be woven in/impregnated Packaging • The swab shall be bound in bundles of five by means of either a paper band consisting of one piece secured by autoclavable glue (and not by a separate adhesive tape) or by means of an acceptable tying material which can withstand autoclaving. • The swabs shall be suitable wrapped so as to be adequately protected against soiling and contamination • Only swabs of the same size shall be packed together in a bulk container. • Packing shall comply with the relevant requirements of the latest issue of CKS 228- double swab • Swabs shall be packed in a red plastic bag  Size 280mm x 400mm, thickness at least 19micron, colour: red, CKS 559 Marking • The following information shall appear in legible and indelible marking on each package:  Manufacture’s name or trade mark or both  Type of swabs  Size of swabs  Quantity of swabs Bulk containers • The following information shall appear in legible and indelible marking on the outside of each bulk container. Base – Peel pouches • Clear PVC film, not less than 90u Lidding – Peel pouches / tray • Peel cover type (heat seal grid), 60g/m² medical Kraft paper lacquered with 11u adhesive Sterility indicator • Pack to contain tamperproof sterility indicator on the tray or pack, confirm sterility, bidder to indicate the type of indicator used Sterile Peel packed in bundles of 5 Pack size: Each
304 | RT42-06-005 | Swab, abdominal, surgical, cotton, X-ray detectable, sterile, 26g, 6-ply Size: 350mm x 350mm Pack size: Pouch of 5 swabs Scope • Swabs for single use, all layers consist of white woven bandage cloth, to comply with the latest issue of SANS 443, type 3 Construction • Edges: All cut edges of the material to be folded in and stitched, leaving no frayed edges Cotton tape • Length of tape: 400mm • Width of tape: 12mm • The cotton tape must be embedded in a 50mm packet i.e. 100mm in length of tape folder double to 50mm stitched into the corner of the swab leaving a 300mm length of cotton tape clear • Breaking strength of tape – minimum 7.95kg X-Ray detectable material • X-Ray detectable material thread shall be woven in/impregnated Packaging • The swab shall be bound in bundles of five by means of either a paper band consisting of one piece secured by autoclavable glue (and not by a separate adhesive tape) or by means of an acceptable tying material which can withstand autoclaving. • The swabs shall be suitable wrapped so as to be adequately protected against soiling and contamination • Only swabs of the same size shall be packed together in a bulk container. • Packing shall comply with the relevant requirements of the latest issue of CKS 228- double swab • Swabs shall be packed in a red plastic bag  Size 280mm x 400mm, thickness at least 19micron, colour: red, CKS 559 Marking • The following information shall appear in legible and indelible marking on each package:  Manufacture’s name or trade mark or both  Type of swabs  Size of swabs  Quantity of swabs Bulk containers • The following information shall appear in legible and indelible marking on the outside of each bulk container. Base – Peel pouches • Clear PVC film, not less than 90u Lidding – Peel pouches / tray • Peel cover type (heat seal grid), 60g/m² medical Kraft paper lacquered with 11u adhesive Sterility indicator • Pack to contain tamperproof sterility indicator on the tray or pack, confirm sterility, bidder to indicate the type of indicator used Sterile Peel packed in bundles of 5 Pack size: Each
305 | RT42-06-006 | Swab, abdominal, surgical, cotton, X-ray detectable, sterile, 36g, 6-ply Size: 370mm x 450mm Pack size: Pouch of 5 swabs Scope • Swabs for single use, all layers consist of white woven bandage cloth, to comply with the latest issue of SANS 443, type 3 Construction • Edges: All cut edges of the material to be folded in and stitched, leaving no frayed edges Cotton tape • Length of tape: 400mm • Width of tape: 12mm • The cotton tape must be embedded in a 50mm packet i.e. 100mm in length of tape folder double to 50mm stitched into the corner of the swab leaving a 300mm length of cotton tape clear • Breaking strength of tape – minimum 7.95kg X-Ray detectable material • X-Ray detectable material thread shall be woven in/impregnated Packaging • The swab shall be bound in bundles of five by means of either a paper band consisting of one piece secured by autoclavable glue (and not by a separate adhesive tape) or by means of an acceptable tying material which can withstand autoclaving. • The swabs shall be suitable wrapped so as to be adequately protected against soiling and contamination • Only swabs of the same size shall be packed together in a bulk container. • Packing shall comply with the relevant requirements of the latest issue of CKS 228- double swab • Swabs shall be packed in a red plastic bag  Size 280mm x 400mm, thickness at least 19micron, colour: red, CKS 559 Marking • The following information shall appear in legible and indelible marking on each package:  Manufacture’s name or trade mark or both  Type of swabs  Size of swabs  Quantity of swabs Bulk containers • The following information shall appear in legible and indelible marking on the outside of each bulk container. Base – Peel pouches • Clear PVC film, not less than 90u Lidding – Peel pouches / tray • Peel cover type (heat seal grid), 60g/m² medical Kraft paper lacquered with 11u adhesive Sterility indicator • Pack to contain tamperproof sterility indicator on the tray or pack, confirm sterility, bidder to indicate the type of indicator used Sterile Peel packed in bundles of 5 Pack size: Each
306 | RT42-06-007 | Swabs, cleansing, Type 3, 100 swabs Diameter: 30mm Materials: Absorbent cotton wool to comply to SANS 228-1st Grade Absorbent gauze to comply to SANS 446 Compliance certificate to be submitted with bid Diameter: 30mm Construction: The cotton wool shall be packed in single layer of gauze covering material. The covering material shall be secured with cotton thread to prevent cotton wool from spilling out of the outer covering. The swabs shall be loosely stitched together in bundles of five by means of an easy to break cotton thread. Packet/Bag
307 | RT42-06-008 | Swabs, cleansing, Type 3 100 swabs Diameter: 65mm Materials: Absorbent cotton wool as per latest issue SANS 228-1st Grade Absorbent gauze as per latest issue of SANS 446 Compliance certificate to be submitted with bid Construction: The cotton wool shall be packed in single layer of gauze covering material. The covering material shall be secured with cotton thread to prevent cotton wool from spilling out of the outer covering. The swabs shall be loosely stitched together in bundles of five by means of an easy to break cotton thread. Packet/Bag
308 | RT42-06-009 | Swabs, surgical, dissecting, Type 1, X-Ray detectable, 100 swabs Size: 10mm Materials: Outer layer The outer layer shall be gauze as per the latest issue of SANS 446 Cotton wool The cotton wool shall be cotton wool Grade 2 as per the latest issue of SANS 228 Compliance certificate to be submitted with bid Construction: The cotton wool shall be packed in single layer of gauze X-ray detectable material of size 5mm x 5mm to be placed between the cotton wool and the outer layer of gauze so as to be visible from the outside. The gauze must be secured with cotton thread of acceptable quality. The swabs shall be loosely stitched together in bundles of five by means of an uneasy to break cotton thread. Packet/bag
309 | RT42-06-010 | Swabs, gauze, absorbent, type 3, bleached, 8 ply, 100 swabs Size: 100mm x 100mm Non sterile To comply with the latest issue of SANS 446 Compliance certificate to be submitted with bid Packet
310 | RT42-06-011 | Swabs, gauze, absorbent, type 3, bleached, 12 ply, 100 swabs Size: 75mm x 75mm Non sterile To comply with the latest issue of SANS 446 Compliance certificate to be submitted with bid Packet
311 | RT42-06-012 | Swabs, gauze, absorbent, type 3, bleached, 12 ply, 100 swabs Size: 100mm x 100mm Non-Sterile To comply with the latest issue of SANS 446 Compliance certificate to be submitted with bid Packet
312 | RT42-06-013 | Swabs, gauze, absorbent, type 3, bleached, sterile, 8 ply, 1x 5 swabs Size: 100mm x 100mm Sterile peel packed To comply with the latest issue of SANS 446 Compliance certificate to be submitted with bid Packet
313 | RT42-06-014 | Swabs, gauze, absorbent, type 3, bleached, sterile, 12 ply, 1 x 5 swabs Size: 75mm x 75mm Sterile peel packed To comply with the latest issue of SANS 446 Compliance certificate to be submitted with bid Packet
314 | RT42-06-015 | Swabs, gauze, absorbent, type 3, bleached, sterile, 12 ply, 1 x 5 swabs Size: 100mm x 100mm Sterile peel packed To comply with the latest issue of SANS 446 Compliance certificate to be submitted with bid Packet
315 | RT42-06-016 | Swabs, gauze, absorbent, type 3, bleached, sterile, X-Ray detectable, 8 ply, 5 swabs Size: 100mm x 100mm Sterile peel packed To comply with the latest issue of SANS 446 Compliance certificate to be submitted with bid Packet
316 | RT42-06-017 | Swabs, gauze, absorbent, type 1, bleached, sterile X-Ray detectable, 12 ply, 1 x 5 swabs Size: 100mm x 100mm Sterile peel packed To comply with the latest issue of SANS 446 Compliance certificate to be submitted with bid Packet
317 | RT42-06-018 | Swabs, gauze, absorbent, type 3, X-Ray detectable, 12 ply, 100 swabs Size: 75mm x 75mm Non Sterile, autoclavable To comply with the latest issue of SANS 446 Compliance certificate to be submitted with bid Packet
318 | RT42-06-019 | Swabs, gauze, absorbent, type 3, X-Ray detectable, 12 ply, 100 swabs Size: 75mm x 75mm Sterile, autoclavable To comply with the latest issue of SANS 446 Compliance certificate to be submitted with bid Packet
319 | RT42-06-020 | Swabs, gauze, absorbent, type 1, X-Ray detectable, 12 ply, 100 swabs Size: 100mm x 100mm Non sterile, autoclavable To comply with the latest issue of SANS 446 Compliance certificate to be submitted with bid Packet
320 | RT42-06-021 | Swabs, gauze, absorbent, type 1, X-Ray detectable, 12 ply, 100 swabs Size: 100mm x 100mm Sterile, autoclavable To comply with the latest issue of SANS 446 Compliance certificate to be submitted with bid Packet
321 | RT42-06-022 | Swabs, gauze, absorbent, type 1, dyed green, 12 ply, 100 swabs Size: 100mm x 100mm Non sterile, autoclavable To comply with the latest issue of SANS 446 Compliance certificate to be submitted with bid Packet
322 | RT42-07-001 | Gauze, ribbon, absorbent cotton Autoclavable, fast edge Size: 12mm x 10m To comply with the latest issue of SANS 446 Compliance certificate to be submitted with bid Roll
323 | RT42-07-002 | Gauze, ribbon, absorbent cotton, roll Autoclavable, fast edge Size: 25mm x 10m To comply with the latest issue of SANS 446 Compliance certificate to be submitted with bid Roll
324 | RT42-07-003 | Gauze, ribbon, absorbent cotton Autoclavable, fast edge Size: 50mm x 10m To comply with the latest issue of SANS 446 Compliance certificate to be submitted with bid Roll
325 | RT42-07-004 | Gauze, ribbon, absorbent cotton Autoclavable, fast edge Size: 75mm x 10m To comply with the latest issue of SANS 446 Compliance certificate to be submitted with bid Roll
326 | RT42-07-005 | Gauze, ribbon, absorbent cotton Autoclavable, fast edge Size: 100mm x 10m To comply with the latest issue of SANS 446 Compliance certificate to be submitted with bid Roll
327 | RT42-07-006 | Gauze, absorbent cotton, type 1, roll Size: 0.9m x 5m Individually packed Non-sterile To comply with the latest issue of SANS 446 Compliance certificate to be submitted with bid Each
328 | RT42-07-007 | Gauze, plugging, cotton, type 1, absorbent, X-Ray detectable, 100 swabs Size: 38mm x 700mm x 12 ply Non sterile To comply with the latest issue of SANS 446 Compliance certificate to be submitted with bid Packet
329 | RT42-07-008 | Gauze, plugging, cotton, type 1, absorbent, X-Ray detectable, 100 swabs Size: 75mm x 900mm x 12 ply Non sterile To comply with the latest issue of SANS 446 Compliance certificate to be submitted with bid Packet
330 | RT42-07-009 | Gauze, Oxidised cellulose, BP, USP Absorbable, haemostatic Sterile, individually peel packed Size: 50mm x 12mm Each
331 | RT42-07-010 | Gauze, Oxidised cellulose, BP, USP Absorbable, haemostatic Sterile, individually peel packed Size: with allowance of 5mm , 70mm x 50mm Each
332 | RT42-07-011 | Gauze, Oxidised cellulose, BP, USP Absorbable, haemostatic Sterile, individually peel packed Size: with allowance of 5mm , 350mm x 50mm Each
333 | RT42-07-012 | Gauze, oxidised cellulose, BP, USP, absorbable, haemostatic. Size: with allowance of 5mm, 200mm x 100mm. Sterile.
334 | RT42-08-001 | Pads, eye, oval, autoclavable, 144 pads Composed of a 6mm to 8mm layer of cotton wool Grade 1 To comply with the latest issue of SANS 228 Enclosed between two layers of absorbent gauze Type 3, to comply with the latest issue of SANS 446 Compliance certificate to be submitted with bid Packet
335 | RT42-08-002 | Pads, eye, oval, sterile Size: 67mm x 54mm at widest point Sterile, individually peel packed To comply with the latest issue of 8840 Compliance certificate to be submitted with bid Each
336 | RT42-08-003 | Pad, surgical, absorbent, sterile Rayon filler with a non-woven fabric sleeve Size: 100mm x 100mm Sterile, individually peel packed To comply with the latest issue of 8840 Compliance certificate to be submitted with bid Each
337 | RT42-08-004 | Pad, surgical, absorbent, sterile Rayon filler with a non-woven fabric sleeve Size: 100mm x 200mm Sterile, individually peel packed To comply with the latest issue of 8840 Compliance certificate to be submitted with bid Each
338 | RT42-08-005 | Pad, surgical, autoclavable non-woven, type B For wound and swabbing Mass: 57g per 10 pads Size: 100mm x 200mm To comply with the latest issue of 8840 Compliance certificate to be submitted with bid Each
339 | RT42-08-006 | Pad, surgical, autoclavable non-woven, type B For wound and swabbing Mass: 109g per 10 pads Size: 200mm x 200mm To comply with the latest issue of 8840 Compliance certificate to be submitted with bid Each
340 | RT42-08-007 | Pad, wound debridement, the pad must be able to debride the wound bed and wound edges. The pad must be able to hydrate the perilesional area and the entire limb. The pad must contain a poloxamer based solution and an isotonic surfactant which is non-cytotoxic to wounds. Individually packed Sterile Size of pad: ±20cm x ±20cm Dispenser box with 10 pouches
341 | RT42-09-001 | Drape, surgical, adhesive, U-slot aperture, 1 x Wrap Drape: Size: 1200mm x 1300mm Pear-shaped split with incise Poly Reinforcement Material, adhesive and handles as per Drape surgical incise adhesive above Each
342 | RT42-09-002 | Drape, surgical, adhesive, body split, 1 x Wrap Drape: Size: 1910mm x 2950mm Split: 254mm x 635 mm Absorbent Reinforcement: 1016mm x 1118mm Pear-shaped split with incise Poly Reinforcement Material, adhesive and handles as per Drape surgical incise adhesive above Each
343 | RT42-09-003 | Drape, surgical, chest/breast with fluid control pouch, 1 x Wrap Drape: Size: 1960mm x 3050mm Fenestration Pear Shaped Poly Reinforcement Tube holders Each
344 | RT42-09-004 | Drape, surgical, adhesive patch with circle aperture centrally situated, Size: 400mm x 400mm Adhesive area: 100mm x 125mm Circle aperture: 60mm diameter Material, adhesive and handles as per Drape surgical incise adhesive above Each
345 | RT42-09-005 | Drape, surgical, adhesive patch with oval aperture situated off-centre, Size: 800mm x 1300mm Adhesive area: 100mm x 125mm Oval aperture: 70mm x 50mm Material, adhesive and handles as per Drape surgical incise adhesive above Each
346 | RT42-09-006 | Drape, surgical, adhesive patch with oval aperture, drainage pouch as part of drape, Size: 230mm x 450mm Adhesive area: 230mm x 250mm Oval aperture: 80mm x 100mm Material, adhesive and handles as per Drape surgical incise adhesive above Each
347 | RT42-09-007 | Drape, disposable, with double instrument pouch compartments Overall size: 180cm (width) x 300cm (length) 2 x adhesive strips 5cm x 18cm Sterile Individually peel packed with description on outer pack Pack size: each
348 | RT42-09-008 | Drape/dressing: Sterile. Self adhesive, transparent, latex free. Material to be an adhesive viscose, polyurethane or polyester or polyethylene film, occlusive to bacteria. The adhesive shall be a hypoallergenic acrylic copolymer. The film to be waterproof and have a siliconised release paper without a“window”frame, Individually packed in a peel pouch. To comply with the latest issue of CKS676 Compliance certificate to be submitted with bid Size:120mm x 250mm Adhesive area:100mm x 250mm Each
349 | RT42-09-009 | Drape/dressing: Sterile. Self adhesive, transparent, latex free. Material to be an adhesive viscose, polyurethane or polyester or polyethylene film, occlusive to bacteria. The adhesive shall be a hypoallergenic acrylic copolymer. The film to be waterproof and have a siliconised release paper without a“window”frame, Individually packed in a peel pouch. To comply with the latest issue of CKS676 Compliance certificate to be submitted with bid Size: 300mm x 280mm Adhesive area:300mm x 200mm Each
350 | RT42-09-010 | Drape/dressing: Sterile. Self adhesive, transparent, latex free. Material to be an adhesive viscose, polyurethane or polyester or polyethylene film, occlusive to bacteria. The adhesive shall be a hypoallergenic acrylic copolymer. The film to be waterproof and have a siliconised release paper without a“window”frame, Individually packed in a peel pouch. To comply with the latest issue of CKS676 Compliance certificate to be submitted with bid Size:400mm x 420mm Adhesive area:350mm x 420mm Each
351 | RT42-09-011 | Drape/dressing: Sterile. Self adhesive, transparent, latex free. Material to be an adhesive viscose, polyurethane or polyester or polyethylene film, occlusive to bacteria. The adhesive shall be a hypoallergenic acrylic copolymer. The film to be waterproof and have a siliconised release paper without a“window”frame, Individually packed in a peel pouch. To comply with the latest issue of CKS676 Compliance certificate to be submitted with bid Size:450mm x 280mm Adhesive area:450mm x 200mm Each
352 | RT42-09-012 | Drape/dressing: Sterile. Self adhesive, transparent, latex free. Material to be an adhesive viscose, polyurethane or polyester or polyethylene film, occlusive to bacteria. The adhesive shall be a hypoallergenic acrylic copolymer. The film to be waterproof and have a siliconised release paper without a“window”frame, Individually packed in a peel pouch. To comply with the latest issue of CKS676 Compliance certificate to be submitted with bid Size:450mm x 550mm Adhesive area:450mm x 480mm Each
353 | RT42-09-013 | Drape/dressing: Sterile. Self adhesive, transparent, latex free. Material to be an adhesive viscose, polyurethane or polyester or polyethylene film, occlusive to bacteria. The adhesive shall be a hypoallergenic acrylic copolymer. The film to be waterproof and have a siliconised release paper without a“window”frame, Individually packed in a peel pouch. To comply with the latest issue of CKS676 Compliance certificate to be submitted with bid Size:560mm x 840mm Adhesive area:560mm x 760mm Each
354 | RT42-09-014 | Drape, surgical incise, with iodine incorporated into the adhesive. Barrier material to be: Tear resistant, able to stretch and regain in any direction, transparent, able to conform to irregular surfaces, matt finish, antistatic, permeable to oxygen and water vapour, waterproof and antistatic Adhesive: Non irritant, non-allergenic, to adhere firmly during prolonged operations also along cut edges, covered with silicone treated paper Handles: Backing firm enough to guide adhesive area into proper positioning Sterile, individually peel packed with description on outer pack Size: 600mm x 350mm Adhesive area: 350mm x 350mm: Each
355 | RT42-09-015 | Drape, surgical incise, with iodine incorporated into the adhesive. Barrier material to be: Tear resistant, able to stretch and regain in any direction, transparent, able to conform to irregular surfaces, matt finish, antistatic, permeable to oxygen and water vapour, waterproof and antistatic Adhesive: Non irritant, non-allergenic, to adhere firmly during prolonged operations also along cut edges, covered with silicone treated paper Handles: Backing firm enough to guide adhesive area into proper positioning Sterile, individually peel packed with description on outer pack Size: 900mm x 450mm Adhesive area: 350mm x 350mm Each
356 | RT42-09-016 | Drape, surgical incise, with iodine incorporated into the adhesive. Barrier material to be: Tear resistant, able to stretch and regain in any direction, transparent, able to conform to irregular surfaces, matt finish, antistatic, permeable to oxygen and water vapour, waterproof and antistatic Adhesive: Non irritant, non-allergenic, to adhere firmly during prolonged operations also along cut edges, covered with silicone treated paper Handles: Backing firm enough to guide adhesive area into proper positioning Sterile, individually peel packed with description on outer pack Size: 900mm x 600mm Adhesive area: 600mm x 600mm Each
357 | RT42-09-017 | Drape, surgical incise, with iodine incorporated into the adhesive. Barrier material to be: Tear resistant, able to stretch and regain in any direction, transparent, able to conform to irregular surfaces, matt finish, antistatic, permeable to oxygen and water vapour, waterproof and antistatic Adhesive: Non irritant, non-allergenic, to adhere firmly during prolonged operations also along cut edges, covered with silicone treated paper Handles: Backing firm enough to guide adhesive area into proper positioning Sterile, individually peel packed with description on outer pack Size: 900mm x 850mm Adhesive area: 600mm x 850mm Each
358 | RT42-09-018 | Drape craniotomy, with povidone iodine incise film. Size:1830mm x 1830mm With adhesive strip, 50mm wide with fluid collection pouch and drainage tap. Each
359 | RT42-09-019 | Drape, disposable, craniotomy, reinforced, absorbent material around fenestration filtering into a fluid collection pouch, integrated adhesive incise, wire attached to pouch to keep shape and allow for effective fluid control, filter inside pouch to catch bone chips and minimize blockage, drainage port attached to pouch, gutters of excess fluid control, full patient drape Fluid collection pouch: 600mm x 600mm; Reinforced area: 1600mm x 2000mm; Fenestration size: 250mm in diameter; Overall size: 230cm (width) x 320cm (length) Each
360 | RT42-09-020 | Drape craniotomy, incise film. Size:550mm x 550mm With adhesive strip, 50mm wide with fluid collection pouch and drainage tap. Each
361 | RT42-09-021 | Drape, cardiovascular bilateral split sheet with incise film, Size: 230mm x 320mm Each
362 | RT42-09-022 | Drape, cardiovascular bilateral split sheet with incise film, Size: 368mm x 214mm Each
363 | RT42-09-023 | Drape craniotomy, with povidone iodine incise film. Size:1000mm x 850mm With adhesive strip, 50mm wide with fluid collection pouch and drainage tap. Each
364 | RT42-09-024 | Drape surgical, under buttock with drainage pouch. Size: 850mm x 1050mm. For obstetric and gynecological procedures. Drainage pouch fitted with drainage tap and pliable to secure for fluid control. Materials must be waterproof. Sterile, individual peelpacked with description on outer pack. Each
365 | RT42-09-025 | Drape disposable surgical, under buttocks pouch, with hand cuffs to position the drape in a sterile setting. 2 towels (800mm x1200mm) materila treated to be hydrophobic, soft and comfortable to drape over patients legs. 1 absorbent sheet (800mm x 1000mm) material highly absorbent and laminated to insure impervious nature. Sterile, peel packed with description on outer pack. Each
366 | RT42-09-026 | Drape, disposable, C-section with clear incise pouch and drainage outlet. Size: 1960mm x 3100mm Adhesiv: 300mm x 300mm 1X dry wipe ( wipe paper toweling) 1x mayo table cover 75cmx140cm (blue plastic with laminated reinforcing) 1xtrolley cover 160cmx205cm with leminated reinforcing 60cmx80cm 1x baby towel All towel drapes are reinforced with laminated spunbond and be absorbable around the critical area Sterile, peelpacked with description on outer pack. Pack
367 | RT42-09-027 | Drape, knee arthroscopy bilateral, with a fluid collection pouch and drainable tap, perforated at the end so the surgeon can tie the drape around his back if he prefers, Overall size: 2290mm (width) x 3200mm (length) with fenestration 100mm x 100mm Barrier material to be: Tear resistant, able to stretch and regain in any direction transparent, able to conform to irregular surfaces, matt finish, antistatic, permeable to oxygen and water vapour, waterproof and antistatic. Critical area to be absorbable Sterile, individual peelpacked with description on outer pack. Each
368 | RT42-09-028 | Drape, disposable, knee, with a fluid pouch, pouch has a wire strip attached to keep its shape so as to control fluid effectively, is reinforce area under the leg to ensure no risk for strikethrough, the drape is perforated at the end so the surgeon can tie the drape around his back if he prefers, Aperture size: 80mm Overall size: 2290mm (width) x 3200mm (length) Critical area to be absorbable Sterile, individual peelpacked with description on outer pack. Each
369 | RT42-09-029 | Drape, disposable, shoulder, with a fluid pouch, 3 adhesive strips, to be use on both shoulders, the pouch has a wire strip attached to keep its shape so as to control fluid effectively, Velcro cable holders attached on either side of the pouch for effective cable management Adhesive strip size: 100mm x 500mm Overall size: 220cm (width) x 410cm (length) Critical area to be absorbable Each
370 | RT42-09-030 | Drape, disposable, ophthalmic, double pouch with wire strips to keep them open and handle fluid effectively, additional wire strip place over nose area so as to lift material off patient’s nose thus allowing for easy breathing Adhesive size: 125mm x 125mm Overall size: 150cm (width) x 275cm (length) Sterile, Individually peel packed with description on outer pack Critical area to be absorbable Each
371 | RT42-09-031 | Drape, hip sheet: Size: 2200mm x 3200mm with two dislocation pouches. Critical area to be absorbable Each
372 | RT42-09-032 | Drape, disposable, spinal, fenestrated, with loose long instrument pouch, dry wipe, 2 x integrated cable holders, 3 layers around critical area – impervious and hydrophilic Instrument pouch size: 260mm x 480mm Adhesive incise: 295mm x 395mm Overall size: 210cm (width) x 300cm (length) Critical area to be absorbable Each
373 | RT42-09-033 | Drape, disposable, surgical, X-Ray, Lap incise sheet, non-woven, water repellent, low linting, pore size allowing air but preventing bacteria passage through fabric, integrated cord holders, separated instrument pouch, 2 pouches to be indicated Overall size: 2100mm (width) x 3200mm (length) Incise adhesive film fenestration: size 270mm x 270mm Sterile Individually peel packed with description on outer pack Critical area to be absorbable Each
374 | RT42-09-034 | Drape, disposable, surgical, X-Ray, C-arm, 4 adhesive patches, 2 plastic strips for securing drape Overall size: 1650mm (width) x 1060mm (length) Each
375 | RT42-09-035 | Drape, extremity. Size: 2280mm (width) x 3000mm (length), Flexable aperture 60mm x 80mm Each
376 | RT42-09-036 | Drape, extremity sheet with fluid collction pouch with tiebacks. Size: 2280mm (width) x 3040mm (length) Each
377 | RT42-09-037 | Drape, disposable, sterile, sheet, extremity, water repellant, non-woven fabric, that allows air but prevent bacteria passage through fabric with hinge and cord holders Reinforce at key points,critical area to be absorbable Expandable fenestration Overall size: 1830mm (width) x 3020mm (length) Sterile Individually peel packed with description on outer pack Each
378 | RT42-09-038 | Drape, disposable, sheet, spunlace, water repellent, non-woven fabric, low linting, pore size that allows air but prevent bacteria passage through the fabric, hypo-allergenic and latex free adhesive, positioned centrally on the edge of long side, sterile, description on outer pack, Adhesive size: 55mm x 1000mm Drape size: 1000mm (width) x 1700mm (length) Pack size: Packet of 5 sheets critical area to be absorbable Pack
379 | RT42-09-039 | Drape, disposable, for minor procedures, with positioning tape and integrated cable holder, 3-layers around critical area – impervious and hydrophilic adhesive must be latex free, ideal for multiple procedures Aperture size: 80mm x 80mm Fenestration: 150mm x 150mm Overall size: 2000mm (width) x 3200mm (length critical area to be absorbable Each
380 | RT42-09-040 | Drape, disposable, basic pack, consisting of the following: 1 x large drape towel: 160cm x 220cm with adhesive tape 1 x medium drape towel: 160cm x 180cm with adhesive tape 2 x small drape sheet: 90cm x 80cm with adhesive tape 1 x dry wipe (white paper toweling) 1 x Mayo table cover 75cm x 140cm (blue plastic with laminated reinforcing) 1 x trolley cover: 160cm x 205cm with laminated reinforcing 60cm x 80cm 1 x suture bag (clear plastic with adhesive tape) 1 x luxfab: 90cm x 90cm (blue crepe paper) 1 x adhesive strip with adhesive backing All towels drapes are reinforced / absorbable with laminated spunbond around the critical area Pack
381 | RT42-09-041 | Drape, disposable, angiography, consisting of the following: 1 x angiography drape: 230cm x 320cm absorbent, adhesive incise, LDPE 1 x back table cover: 160cm x 210cm reinforced 1 x dry wipe (white paper towel) 2 x cathlab cover small round (clear LDPE plastic) 1 x cathlab cover: 105cm x 75cm (clear LDPE plastic) Full patient drape, blue absorbent material around femoral area with integrated gutter, adhesive incise for both legs Drape trimmed with clear plastic to cover control panel alongside bed disposable, for minor procedures, with positioning tape and integrated cable holder, 3-layers around critical area – impervious and hydrophilic adhesive must be latex free, ideal for multiple procedures Aperture size: 80mm x 80mm Fenestration: 150mm x 150mm Overall size: 2000mm (width) x 3200mm (length critical area to be absorbable Pack
382 | RT42-09-042 | Drape, disposable, angiography, Brachial: Size 1120mm x 1450mm one central aperture, surrounded by adhesive strip. Barrier material to be: tear resistant, able to stretch and regain in any direction, transparent able to conform to irregular surfaces, matt finish, antistatic, impermeable to bacteria permeable to oxygen and water vapour, waterproof. Adhesive: low irritant, hypo-allergenic, to adhere firmly during prolonged operations, also on cut edges covered with silicone treated paper Handles: backing firm enough to guide adhesive area into proper positioning Sterile, individually peel packed with description on outer pack Critical area to be absorbable Each
383 | RT42-09-043 | Drape, disposable, angiography, Femoral: Size 3250mm x 2030mm Two self adhesive apertures 60mm x 90mm. Barrier material to be: tear resistant, able to stretch and regain in any direction, transparent able to conform to irregular surfaces, matt finish, antistatic, impermeable to bacteria permeable to oxygen and water vapour, waterproof. Adhesive: low irritant, hypo-allergenic, to adhere firmly during prolonged operations, also on cut edges covered with silicone treated paper Handles: backing firm enough to guide adhesive area into proper positioning Sterile, individually peel packed with description on outer pack Critical area to be absorbable Each
384 | RT42-09-044 | Drape, disposable, utility set, consisting of 4 towel drapes 42cm x 30cm with adhesive tape Barrier material must be: • Fluid repellant • Impermeable to bacteria • Permeable to oxygen and water vapor Sterile Individually peel packed Set
385 | RT42-09-045 | Drape disposable surgical, under buttocks, with fluid control pouch 1 x Wrap Drape: Absorbent Reinforcement: 178 mm x 508mm and 254mm x 508mm Overall size: 940mm x 1040mm Sterile, peel packed with description on outer pack. Critical area to be absorbable Each
386 | RT42-09-046 | Drape disposable surgical, under buttocks, with fluid control pouch 1 x Wrap Drape: Adhesive Overall size: 990mm x 1120mm Sterile, peel packed with description on outer pack. Critical area to be absorbable Each
387 | RT42-09-047 | Drape, disposable.Laparoscopic Assisted Vaginal Hysterectomy with fluid control pouch, Tube holders Absorbent Reinforcements Instrument Pad: 254mm x 445mm Sterile, peel packed with description on outer pack.
388 | RT42-09-048 | Drape disposable surgical, under buttocks 1 x Wrap Drape: Under buttock drape with Absorbent Reinforcement Overall size: 890 x 1240mm Sterile, peel packed with description on outer pack. Each
389 | RT42-09-049 | Drape, disposable, knee arthroscopy, with a fluid control pouch, tube holders 1 x Wrap Drape: Fenestration Absorbent Reinforcement size: 890mm x 1194mm Overall size: 2870mm (width) x 2920mm (length) Sterile, individual peelpacked with description on outer pack. Each
390 | RT42-09-050 | Drape, disposable, surgical, Laparascopy/Cholecystectomy, 1 x Wrap Drape: Absorbent Reinforcement: 1041mm x 1422mm Instrument Pad: 254mm x 457mm Pouches Adhesive Overall size: 2690mm x 3100mm Sterile Individually peel packed with description on outer pack Each
391 | RT42-09-051 | Drape, disposable, surgical, Laparascopy T-Drape with armboard covers, 1 x Wrap Absorbent Reinforcement Fenestration: 152mm x 203mm Fenestration: 76mm x 152mm Pouches Tube Holders Adhesive Overall size: 2460mm x 2590mm Sterile Individually peel packed with description on outer pack Each
392 | RT42-09-052 | Drape, disposable, shoulder, with a fluid control pouch and tube holders 1 x Wrap Drape: Fenestration Pear Shaped Poly Reinforcement Overall size: 1570mm (width) x 2540 (length) Sterile, Individually peel packed with description on outer pack Each
393 | RT42-09-053 | Drape, disposable, shoulder split, with incise and tube holders 1 x Wrap Drape: Pear Shaped Split with incise Poly Reinforcement Overall size: 1600mm (width) x 2540mm (length) Sterile, Individually peel packed with description on outer pack Each
394 | RT42-09-054 | Drape, disposable, shoulder split with fluid pouch 1 x Wrap Drape: Split size: 152mm x 940 Fluid Control Pouch with Malleable Wire Overall size: 2290mm (width) x 2540mm (length) Sterile, Individually peel packed with description on outer pack Each
395 | RT42-09-055 | Drape, disposable, beach chair shoulder with fluid control pouch 1 x Wrap Drape: Fenestration Pear Shaped Poly Reinforcement Velcro Tabs Fluid Control Pouch Overall size: 2590mm (width) x 4110mm (length) Sterile, Individually peel packed with description on outer pack EacH
396 | RT42-09-056 | Drape, disposable, ophthalmic, incise with fluid control pouches, 1 x Wrap . Drape: Icise size: 102mm x 114mm Overall size: 1780mm (width) x 2540(length) Sterile, Individually peel packed with description on outer pack Each
397 | RT42-09-057 | Drape, disposable, ophthalmic,with fluid collection saddlebag pouch and 1 x Wrap Drape: Tube holders. Fenestration size: 102mm x 127mm Absorbent Reinforcement size: 279mm x 457mm Overall size: 1220mm (width) x 1700(length) Sterile, Individually peel packed with description on outer pack Each
398 | RT42-09-058 | Drape, disposable, hip 1 x Wrap Drape: Tube holders. Teardrop Fenestration Absorbent Reinforcement Overall size: 2950mm (width) x 3330(length) Sterile, Individually peel packed with description on outer pack Each
399 | RT42-09-059 | Drape, disposable, hip with fluid control pouch 1 x Wrap Drape: Tube holders. Circular Fenestration Absorbent Reinforcement: 558mm x 1372mm Overall size: 2820mm (width) x 3380(length) Sterile, Individually peel packed with description on outer pack Each
400 | RT42-09-060 | Drape, disposable, angiography: Size 1910mm x 3180mm 1 x Wrap 1 x Heavy Duty Absorbent Pad 914mm x 3175mm Circular Fenestration Each
401 | RT42-09-061 | Drape, disposable camera cover size 1400mm (width) x 2500mm (length), drape camera/arm drape adhesive poly strap to hold in place adhesive strip at base clear poly film sterile. individually peel packed, description on outer pack. Sterile, Individually peel packed with description on outer pack Each
402 | RT42-10-001 | Tape, surgical, adhesive film, microporous paper consisting of non-woven synthetic fibres spread with a non-occlusive acrylic adhesive Surgical tape must: permit skin ventilation, be hypo-allergenic, be tolerated by skin for prolonged periods, not irritate the skin Size: 12 mm x 5 m Individually wrapped Roll
403 | RT42-10-002 | Tape, surgical, adhesive film, microporous paper consisting of non-woven synthetic fibres spread with a non-occlusive acrylic adhesive Surgical tape must: permit skin ventilation, be hypo-allergenic, be tolerated by skin for prolonged periods, not irritate the skin Size: 25mm x 5 m Individually wrapped Roll
404 | RT42-10-003 | Tape, surgical, adhesive film, microporous paper consisting of non-woven synthetic fibres spread with a non-occlusive acrylic adhesive Surgical tape must: permit skin ventilation, be hypo-allergenic, be tolerated by skin for prolonged periods, not irritate the skin Size: 50 mm x 5 m Individually wrapped Roll
405 | RT42-10-004 | Tape, surgical, adhesive film, microporous paper consisting of non-woven synthetic fibres spread with a non-occlusive acrylic adhesive Surgical tape must: permit skin ventilation, be hypo-allergenic, be tolerated by skin for prolonged periods, not irritate the skin Size: 75 mm x 5 m Individually wrapped Roll
406 | RT42-10-005 | Tape, surgical, adhesive film, synthetic must be hypo-allergenic, permeable and water repellent Size: 25 mm x 5 m Individually wrapped Roll
407 | RT42-10-006 | Tape, surgical, adhesive film, synthetic must be hypo-allergenic, permeable and water repellent Size: 50 mm x 5 m Individually wrapped Roll
408 | RT42-10-007 | Tape, surgical, adhesive film, synthetic must be hypo-allergenic, permeable and water repellent Size: 75 mm x 5 m Individually wrapped Roll
409 | RT42-10-008 | Tape, surgical, adhesive film, synthetic must be hypo-allergenic, permeable and water repellent Size: 100 mm x 5 m Individually wrapped Roll
410 | RT42-10-009 | Tape, zinc oxide, surgical, adhesive, B.P., woven cotton cloth, non-stretch Size: 25 mm x 5 m Individually wrapped Roll
411 | RT42-10-010 | Tape, zinc oxide, surgical, adhesive, B.P., woven cotton cloth, non-stretch Size: 50 mm x 5 m Individually wrapped Roll
412 | RT42-10-011 | Tape, zinc oxide, surgical, adhesive, B.P., woven cotton cloth, non-stretch Size: 75 mm x 5 m Individually wrapped Roll
413 | RT42-11-001 | Patties, surgical, highly absorbent for neuro surgical use X-Ray detectable Size: 6mm x 6mm Sterile peel packed Each
414 | RT42-11-002 | Brain patties with radio opaque string, latex free, surgical, highly absorbent, X-Ray detectable Size: 2.5mm x2.5mm Pack of 5 Sterile peel packed Each
415 | RT42-11-003 | Brain patties with radio opaque string, latex free, surgical, highly absorbent, X-Ray detectable Size: 2mm x 2mm Pack of 5 Sterile peel packed Each
416 | RT42-11-004 | Brain patties with radio opaque string, latex free, surgical, highly absorbent, X-Ray detectable Size: 6mm x 6mm Pack of 5 Sterile peel packed Each
417 | RT42-11-005 | Brain patties with radio opaque string, latex free, surgical, highly absorbent, X-Ray detectable Size: 12mm x 12mm Pack of 5 Sterile peel packed Each
418 | RT42-11-006 | Brain patties with radio opaque string, latex free, surgical, highly absorbent, X-Ray detectable Size: 12mm x 50mm Pack of 5 Sterile peel packed Each
419 | RT42-11-007 | Brain patties with radio opaque string, latex free, surgical, highly absorbent, X-Ray detectable Size: 13mm x 13mm Pack of 5 Sterile peel packed Each
420 | RT42-11-008 | Brain patties with radio opaque string, latex free, surgical, highly absorbent, X-Ray detectable Size: 12mm x 75mm Pack of 5 Sterile peel packed Each
421 | RT42-11-009 | Brain patties with radio opaque string, latex free, surgical, highly absorbent, X-Ray detectable Size: 20mm x 20mm Pack of 5 Sterile peel packed Each
422 | RT42-12-001 | Sponge, surgical, polyvinyl alcohol type X-ray detectable Nasal tampon 45mm - double compressed Sterile peel packed Each
423 | RT42-12-002 | Sponge, surgical, polyvinyl alcohol type X-ray detectable Nasal tampon 50mm - double compressed Sterile peel packed Each
424 | RT42-12-003 | Sponge, surgical, polyvinyl alcohol type X-ray detectable Nasal tampon 75mm - double compressed Sterile peel packed Each
425 | RT42-12-004 | Sponge, surgical, polyvinyl alcohol type X-ray detectable Nasal tampon 90mm - double compressed Sterile peel packed Each
426 | RT42-12-005 | Spongette, triangular with pliable handle Consisting of highly asorbent, non- linting, fine grain compressed cellular material Sterile peel packed Packet of 10 spongettes
427 | RT42-13-001 | Plaster, waterproof, self-adhesive, B.P. Rolled with adhesive inwards on metal or plastic spool, suitably protected by metal or plastic cover Size: 25 mm x 3 m Roll
428 | RT42-13-002 | Plaster, waterproof, self-adhesive, B.P. Rolled with adhesive inwards on metal or plastic spool, suitably protected by metal or plastic cover Size: 50 mm x 3 m Roll
429 | RT42-13-003 | Plaster, waterproof, self-adhesive, B.P. with non-stick wound padding (island) assorted Washable plasters Individually wrapped strips Packet of 20 plasters
430 | RT42-14-001 | Skin closure strips, non-allergenic Strips shall have a high initial adherence for short term use up to 5 days Size: 3 mm x 75 mm Sterile peel packed Packaging: 5 Strips per envelope
431 | RT42-14-002 | Skin closure strips, non-allergenic Strips shall have a high initial adherence for short term use up to 5 days Size: 12 mm x 100 mm Sterile peel packed Packaging: 6 Strips per envelope
432 | RT42-15-001 | Support, intravenous, paediatric For paediatric use Size: 65 mm x 70mm Aluminium rod of 400 mm width encased in foam 3 mm thick covered with medical felt 2 Velcro strips on each end Each
433 | RT42-15-002 | Support, intravenous, paediatric For paediatric use Size: 85 mm x 115mm Aluminium rod of 400 mm width encased in foam 3 mm thick covered with medical felt 2 Velcro strips on each end Each
434 | RT42-15-003 | Support, intravenous, paediatric For paediatric use Size: 100 mm x 120 mm Aluminium rod of 400 mm width encased in foam 3 mm thick covered with medical felt 2 Velcro strips on each end Each
435 | RT42-15-004 | Support, intravenous, paediatric For paediatric use Size: 150 mm x 50 mm Aluminium rod of 400 mm width encased in foam 3 mm thick covered with medical felt 2 Velcro strips on each end Each
436 | RT42-15-005 | Support, intravenous,neonatal For neonatal use Size:50 mm x 57 mm Aluminium rod of 400 mm width encased in foam 3 mm thick covered with medical felt 2 Velcro strips on each end Each
437 | RT42-16-001 | Padding, natural, orthopaedic, bonded type, autoclavable, type 2 Size: 50 mm x 3 m To comply with the latest issue of SANS 196 Compliance certificate to be submitted Roll
438 | RT42-16-002 | Padding, natural, orthopaedic, bonded type, autoclavable, type 2 Size: 75 mm x 3 m To comply with the latest issue of SANS 196 Compliance certificate to be submitted Roll
439 | RT42-16-003 | Padding, natural, orthopaedic, bonded type, autoclavable, type 2 Size: 100 mm x 3 m To comply with the latest issue of SANS 196 Compliance certificate to be submitted Roll
440 | RT42-16-004 | Padding, natural, orthopaedic, bonded type, autoclavable, type 2 Size: 150 mm x 3 m To comply with the latest issue of SANS 196 Compliance certificate to be submitted Roll
441 | RT42-16-005 | Padding, natural, orthopaedic, bonded type, autoclavable, type 2 Size: 200mm x 3m To comply with the latest issue of SANS 196 Compliance certificate to be submitted Roll
442 | RT42-16-006 | Padding, orthopaedic, synthetic cast, type 4 or type 5 Size: 50mm x 3m To comply with the latest issue of SANS 196 Compliance certificate to be submitted Roll
443 | RT42-16-007 | Padding, orthopaedic, synthetic cast, type 4 or type 5 Size: 75mm x 3m To comply with the latest issue of SANS 196 Compliance certificate to be submitted Roll
444 | RT42-16-008 | Padding, orthopaedic, synthetic cast, type 4 or type 5 Size: 100mm x 3m To comply with the latest issue of SANS 196 Compliance certificate to be submitted Roll
445 | RT42-16-009 | Padding, orthopaedic, synthetic cast, type 4 or type 5 Size: 150mm x 3m To comply with the latest issue of SANS 196 Compliance certificate to be submitted Roll
446 | RT42-17-001 | Skin traction kit Size: Children Composed of extension plaster B.P. with a square plastic spreader in the centre on the adhesive side and covered with soft foam padding 10mm thick, the rest of the adhesive being covered by a flexible plastic protector. The spreader should have holes through which traction cord can be threaded. To be supplied with a cotton crepe bandage, type 1, to comply with the latest issue of SANS 945 Compliance certificate to be submitted Kit: Components with the following dimensions: Extension plaster: 1.8m long x 50 mm wide, Foam padding: 360mm long x ±50mm wide, Plastic spreader: 50mm square Traction cord: 1.65m long Cotton crepe bandage: not less than 4.5m x 50mm wide Packed in suitable labelled boxes Kit (Each)
447 | RT42-17-002 | Skin traction kit Size: Adult Composed of extension plaster B.P with a square plastic spreader in the centre on adhesive side and covered with soft foam padding 10mm thick, the rest of the adhesive being covered by a flexible plastic protector. The spreader should have holes through which traction cord can be threaded. To be supplied with a cotton crepe bandage, type 1, to comply with the latest issue of SANS 945 Compliance certificate to be submitted Kit: Components with the following dimensions: Extension plaster: 1.8m long x 75 mm wide, Foam padding: 1.8 m long x 90 mm wide Plastic spreader: 75mm square Traction cord: 1.65m long Cotton crepe bandage: not less than 4.5m x 75mm wide Packed in suitable labelled boxes Each (Kit)
448 | RT42-18-001 | Support, scrotal, suspensory Elastic waistband and legbands Size: Medium Individually packed Each
449 | RT42-18-002 | Support, scrotal, suspensory Elastic waistband and legbands Size: Large Individually packed Each
450 | RT42-19-001 | Soft silicone wound contract layer, atraumatic, adherent to dry skin. Size: ±50mm x 70mm. Sterile, Individually peel packed, Each
451 | RT42-19-002 | Soft silicone wound contract layer, atraumatic, adherent to dry skin. Size: ±80mm x 100mm,Sterile, Individually peel packed, Each
452 | RT42-19-003 | Soft silicone wound contract layer, atraumatic, adherent to dry skin. Size: ±100mm x 180mm, Sterile, Individually peel packed, Each
453 | RT42-19-004 | Soft silicone wound contract layer, atraumatic, adherent to dry skin. Size: ±180mm x 200mm, Sterile, Individually peel packed, Each
454 | RT42-19-005 | Soft silicone wound contract layer, atraumatic, adherent to dry skin. Size: ±200mm x 310mm, Sterile, Individually peel packed, Each
455 | RT42-19-006 | Soft silicone wound contract layer, atraumatic, adherent to dry skin. Size: ±350mm x 600mm, Sterile, Individually peel packed, Each
456 | RT42-20-001 | Silicone Non-Adhesive Foam Dressing. Size: ±50mm x 70mm, Sterile, Individually peel packed, Each
457 | RT42-20-002 | Silicone Non-Adhesive Foam Dressing. Size: ±100mm x 100mm , Sterile, Individually peel packed, Each
458 | RT42-20-003 | Silicone Non-Adhesive Foam Dressing. Size:±100mm x 200mm, Sterile, Individually peel packed, Each
459 | RT42-20-004 | Silicone Non-Adhesive Foam Dressing. Size: ±150mm x 150mm, Sterile, Individually peel packed, Each
460 | RT42-20-005 | Silicone Non-Adhesive Foam Dressing. Size: ±200mm x 200mm, Sterile, Individually peel packed, Each
461 | RT42-20-006 | Silicone Non-Adhesive Foam Dressing. Heel: 100mm x 130mm, Sterile, Individually peel packed, Each
462 | RT42-20-007 | Silicone Adhesive Foam Dressing (Border), Size: ±75mm x 75mm,Sterile, Individually peel packed, Each
463 | RT42-20-008 | Silicone Adhesive Foam Dressing (Border), Size: ±125mm x125mm, Sterile, Individually peel packed, Each
464 | RT42-20-009 | Silicone Adhesive Foam Dressing (Border), Size: ±225mm x 225mm, Sterile, Individually peel packed, Each
465 | RT42-20-010 | Silicone Adhesive Foam Dressing (Border), Sacral 230mm x 230mm, Sterile, Individually peel packed, Each
466 | RT42-20-011 | Silicone Adhesive Foam Dressing (Border), Size: 100mm x 300mm, Sterile, Individually peel packed, Each
467 | RT42-20-012 | Silicone Adhesive Foam Dressing (Border), Size: 100mm x 200mm, Sterile, Individually peel packed, Each
468 | RT42-21-001 | Casting Synthetic Splint,Size: ±100mm x 400mm consist of multiple layers of knitted fabric, impregnated with moisture-curable resin, must have double-sided padding, exposure of splint to moisture of water initiates a chemical reaction, which causes the splint to become rigid, highly comfortable, water -repellent, lightweight , portable, setting time ± 4 min, functional strength ± 20 min, individually wrapped in a sealed foil pouch. Each
469 | RT42-21-002 | Casting Synthetic Splint,Size: with allowance of 5mm,100mm x, 760mm consist of multiple layers of knitted fabric, impregnated with moisture-curable resin, must have double-sided padding, exposure of splint to moisture of water initiates a chemical reaction, which causes the splint to become rigid, highly comfortable, water -repellent, lightweight , portable, setting time ± 4 min, functional strength ± 20 min, individually wrapped in a sealed foil pouch. Each
470 | RT42-21-003 | Casting Synthetic Splint,Size: with allowance of 5mm,125mm x, 760mm consist of multiple layers of knitted fabric, impregnated with moisture-curable resin, must have double-sided padding, exposure of splint to moisture of water initiates a chemical reaction, which causes the splint to become rigid, highly comfortable, water -repellent, lightweight , portable, setting time ± 4 min, functional strength ± 20 min, individually wrapped in a sealed foil pouch. Each
471 | RT42-22-001 | Wipe, to remove adhesiveness from the skin. The wipe must also moisturizes and conditions the skin without greasiness. Gently cleans hydrocolloid, acrylic and rubber-based adhesive residues from the skin without irritation, discomfort or trauma. Applies easily and removes quickly with less need for reapplication. Non-irritating and non-sensitizing, even for ostomies and geriatric patients. Can be used anywhere there's a need for an adhesive dressing. Individually sterile wrapped Single use Size of wipe: 400mm x 400mm Box of 100 wipes
472 | RT42-22-002 | Wipes, to protect the skin, is impregnated with a liquid film-forming dressing, with application to intact skin, a protective film is formed to help reduce friction during removal of tapes and dressings. Must be able to be used to prepare skin attachment sites for drainage tubes, external catheters, surrounding ostomy sites and adhesive dressings. Helps with tape and film. Removes easily when using skin cleanser or soap and water, so unsanitary residue is not left in or around the wound area. Allows skin to ""breathe"" so tapes and films adhere better. Must be tested to be use also on paediatric patients. Individually sterile wrapped Single use Size of wipe: 400mmx400mm Box of 100 wipes
473 | RT42-22-003 | Adhesive removal liquid, to remove adhesives from the skin without irritation, discomfort or trauma. Liquid must be an alcohol based solvent. Applied solvent over tapes for easy and painless removal Bottle of 350 ml Each