## ECT Tray
ECT Tray
Description
Description
Device used in the Post Anaesthesia Care Unit (PACU) for inducing seizures in patients by applying a therapeutic electric current to the brain through scalp electrodes
Specifications
The system shall be equipped automatic self-diagnostic program upon start-up, which shall detect and clearly indicate any defects or malfunction
The system shall be well constructed with durable materials to withstand typical abuse and cleaning
The system shall have connections that are secure enough to resist accidental disconnection
The unit shall consist of a console with controls for the stimulus parameters, displays and printers with non-thermal paper and a minimum of two electrodes
The unit shall be integrated with the following monitoring features:
Vital signs
Electrocardiograph(ECG) monitoring
Electroencephalograph(EEG) monitoring
Electromyograph(EMG) monitoring
The electrodes (handheld or assemblies) shall have no conductive areas except on the surfaces intended to contact the patient for the desired therapy
The unit shall provide a digital display of dose delivered and shall permit all parameters to be adjusted
The unit shall be equipped with audible and visible stimulus indicators
The unit shall have incorporated impedance tests to measure the quality of electrode to patient contact
The stimulus parameters of the system shall include:
Voltage: 50 - 400V
Current: 500 - 800MA
Frequency: 20 - 120Hz
Ultra Pulse Width: 0.5 - 1.5m.sec
Duration: 0.1 - 5.9 sec
Energy: 0.6 - 205J (220 ohm patient impedance)
The system shall be equipped with the following features:
Postictal Suppression Index
4-Channel Monitor/Printer
Optimized Stimulus Programs
Power Spectral Analysis
Upgrade by Changing Microchip
Maximum Sustained EEG Power
Continuous Digital Heart Rate Display & Printout
Baseline and Peak Heart Rate Printout
Seizure Duration from Heart Rate
EEG Coherence Analysis
Seizure Energy Index
Computer-Determined EEG Seizure Duration
Computer-Determined Motor Seizure Duration
True EMG Monitor
EEG Ictal Line Seizure Indicator
Light-Emitting Elapsed Time Display
Audible EEG monitor
Instant Impedance Test
Extended Seizure Alert
The system shall be easy to clean, disinfect, and/or sterilize, as appropriate
The unit shall operate directly from a 230V ± 10%, 50 ± 2 Hertz, single phase AC supply as well as its own internal, rechargeable battery
Switches and controls should be protected against penetration of fluids
STANDARD ACCESSORIES
Adherent stimulus electrodes
Audible EEG audio cassette tape
Non-Thermal recording paper
Monitoring and treatment cables
Monitoring lead wires for ECG, clip style, set of 9
Monitoring lead wires for EMG, clip style, set of 2
Oral protector
EEG/ECG/EMG self-adhesive disposable recording electrodes

## Treatment Electrodes
Treatment Electrodes for ECT
Pretac gel for use with electrodes
Recording paper
Vent A disposable mouth guard
Mouth guard oral protector
Monitoring cable for system 4
Monitoring electrode leads for system 4
Monitoring cable for sysytem 4 (60 inch)
Treatment cable for thymapads

## Coat and Towel Hook
Ligature resistant Coat and Towel hook
Description
Used to hang towel and coats for easy storage
Specification
Suicide resistant design.
Fixed in place with two Stainless Steel Torx Head Countersunk security screws.
Ligature resistant coat hook that is able to carry the weight of approximately 6 kilograms before releasing the load.
The surface should be easy to clean and maintain while not affecting colour and finish.
Durable, long lasting with low maintenance and replacement costs.
No sharp edges, points or protruding parts that could potentially cause harm.
Made from a low shore rating rubber with a high silicone content and mounted in a Stainless Steel Housing unit.
Dimensions
65mm in diameter
Stainless Steel - 8mm thick

## Observation Mirror
Full domed convex observation mirror
Description
An equipment used to monitor patients movements
Specification
Tamperproof and durable full dome observation mirror, mounted on the roof of rooms for safety and security.
Virtually indestructible and unbreakable, mirror quality polycarbonate face.
Pure reflective coating.
Powder coated steel frame for enhanced long life and appearance.
Mounted with Stainless Steel Tamperproof Torx Head screws.
Distortion free reflective coating for excellent 360° viewing quality.
Recessed, tamper proof fixings increases security and reduces ligature risk.
The frame must have Counter-bore holes designed to take security head fasteners.
Foam insert to fill cavity and enhance strength.
Upon installation, the mirror shall be sealed to the roof with a tamper and pick resistant silicone based sealant.
Dimensions
Mirror Size 600mm diameter

## Half Observation Mirror
Half domed convex observation mirror
Description
An equipment used to monitor patients movements
Specification
Tamperproof and durable half dome observation mirror, mounted in the roof and wall of rooms for safety and security.
Virtually indestructible and unbreakable, mirror quality polycarbonate face.
Pure reflective coating.
Powder coated steel frame for enhanced long life and appearance.
Mounted with stainless steel tamperproof Torx Head screws.
Distortion free reflective coating for excellent 180° viewing quality.
Recessed, tamper proof fixings increases security and reduces ligature risk.
The Frame must have counter-bore holes designed to take security head fasteners.
Upon installation, the mirror shall be sealed to the wall with a tamper and pick resistant silicone based sealant.
Foam insert to fill cavity and enhance strength.
Dimensions
Mirror Size 600x300mm

## Quarter Mirror
Quarter domed convex observation mirror
Description
An equipment used to monitor patients movements
Specification
Tamperproof and durable quarter dome observation mirror, mounted in the corner of rooms for safety and security.
Virtually indestructible and unbreakable, mirror quality polycarbonate face.
Pure reflective coating.
Powder coated steel frame for enhanced long life and appearance.
Mounted with Stainless Steel Tamperproof Torx Head screws.
Distortion free reflective coating for excellent viewing quality.
Recessed, tamper proof fixings increases security and reduces ligature risk.
The Frame must have Counter-bore holes designed to take security head fasteners.
Upon installation, the mirror shall be sealed to the wall with a tamper and pick resistant silicone based sealant.
Foam Insert to fill cavity and enhance strength.
Dimensions
Mirror Size 450mm x 450mm