## Category A
Item Number | Category | Item Description | Bidder Name | Name of Manufacturer | Couuntry of Origin | Brand Name | Model/Model Number | Year of Release | Expected Life Span of the Equipment | No. of Offers | Unit of Measure | Equipment/Accessory price inclusive of VAT, 1st and 2nd Year Maintenance Costs | 3rd Year Maintenance Costs | 4th Year Maintenance Costs | 5th Year Maintenance Costs | 6th Year Maintenance Costs | 7th Year Maintenance Costs | Calibration costs annum | Total Costs | sort field | Total Costs | LOWEST ACCEPTABLE BID | POINTS FOR PRICE | HDI | Total Points | Ranking | Comment
RT55-01-001 | Audiology | DIAGNOSTIC AUDIOMETER (WITH EXTENDED HIGH FREQUENCY TESTING CAPABILITIES) Description:PC-based audiometer used to test and diagnose hearing thresholds. Technical Specifications:PC-controlled, clinical two-channel audiometer that must test in increments of no more than 5db. Capable of performing masked and unmasked pure tone and warble tone air and bone conduction and speech . Masking sounds must include narrowband, white and speech noise. Frequency range for air conduction must be within a range of 125Hz to minimum of 16 000Hz with intensity range from -10dB to a minimum of 100dB -10dB to 80 or 100dB SPL. Intensity range for air conduction (125-250Hz): -10dB to 100dB decibel range for bone conduction: 250- minimum of 4000Hz. Must have the option to add following special tests: SISI. ABLB. stenger bekesy weber.tone-in-noise speech-in-noise. Must come with earphones. (including high frequency earphones) paired insert phones bone conductor response button talk-forward built-in or gooseneck microphone or operator's headset with microphone pair of sound-field speakers. The audiometer offered must allow for user selectable attenuation over the entire range. Must provide auditory as well as visual patient response monitoring from its frequency and intensity controls in order to further enhance user control of its operation. Pre-recorded speech wordlists in English, with the option of South African languages in male and female voice, must be built-in, failing which the audiometer must have a built-in ability to play recorded speech tests .Contralateral white noise, speech noise and narrow band noise for air / bone and free field audiometry masking must be provided on the unit offered. Must be supplied with field speakers Specifications of field speakers are as follows:Space saving, high efficiency, flat response, low distortion speakers. Speakers should have a frequency response range of 250Hz to 6000Hz with extendable range up to 8000Hz. Meets ANSI or equivalent standards.Impedance of 8 Ω (Ohm) and efficiency of 100dB HTL Speech Level Must be compatible with booths and diagnostic audiometers. Must include mounting brackets for installation. Accessories: earphones paired insert phones bone conductor response button. talk-forward built-in or gooseneck microphone or operator's headset with microphone field speakers in terms of the specifications above. Optional accessories: SISI . C239ABLB,stenger bekesy weber,tone-in-noise speech-in-noise. | Navitrons | Electronique du Mazet | France | Echodia | AudioSmart (Article code 303244) with DD450 (HF) (Free Field not availabe) with Laptop | 2011 | 10 years | 2 (Offer 2) | 1 | 126000 | 15120 | 16632 | 18295.2 | 20124.72 | 22137.192 |  | 196225.92 | RT55-01-001 | 196225.92 | 196225.92 | 90 | 0 | 90 | 1 | Recommended
RT55-01-001 | Audiology | DIAGNOSTIC AUDIOMETER (WITH EXTENDED HIGH FREQUENCY TESTING CAPABILITIES) Description:PC-based audiometer used to test and diagnose hearing thresholds. Technical Specifications:PC-controlled, clinical two-channel audiometer that must test in increments of no more than 5db. Capable of performing masked and unmasked pure tone and warble tone air and bone conduction and speech . Masking sounds must include narrowband, white and speech noise. Frequency range for air conduction must be within a range of 125Hz to minimum of 16 000Hz with intensity range from -10dB to a minimum of 100dB -10dB to 80 or 100dB SPL. Intensity range for air conduction (125-250Hz): -10dB to 100dB decibel range for bone conduction: 250- minimum of 4000Hz. Must have the option to add following special tests: SISI. ABLB. stenger bekesy weber.tone-in-noise speech-in-noise. Must come with earphones. (including high frequency earphones) paired insert phones bone conductor response button talk-forward built-in or gooseneck microphone or operator's headset with microphone pair of sound-field speakers. The audiometer offered must allow for user selectable attenuation over the entire range. Must provide auditory as well as visual patient response monitoring from its frequency and intensity controls in order to further enhance user control of its operation. Pre-recorded speech wordlists in English, with the option of South African languages in male and female voice, must be built-in, failing which the audiometer must have a built-in ability to play recorded speech tests .Contralateral white noise, speech noise and narrow band noise for air / bone and free field audiometry masking must be provided on the unit offered. Must be supplied with field speakers Specifications of field speakers are as follows:Space saving, high efficiency, flat response, low distortion speakers. Speakers should have a frequency response range of 250Hz to 6000Hz with extendable range up to 8000Hz. Meets ANSI or equivalent standards.Impedance of 8 Ω (Ohm) and efficiency of 100dB HTL Speech Level Must be compatible with booths and diagnostic audiometers. Must include mounting brackets for installation. Accessories: earphones paired insert phones bone conductor response button. talk-forward built-in or gooseneck microphone or operator's headset with microphone field speakers in terms of the specifications above. Optional accessories: SISI . C239ABLB,stenger bekesy weber,tone-in-noise speech-in-noise. | Navitrons | Electronique du Mazet | France | Echodia | Elios (Article code: 300644) - Audiometry with HF including DD450 (HF)(Free Field not availabe) with Laptop | 2011 | 10 years | 2 (Offer 1) | 1 | 170000 | 20400 | 22440 | 24684 | 27152.4 | 29867.64 |  | 264729.4 | RT55-01-001 | 264729.4 | 196225.92 | 58.5805 | 0 | 58.5805 | 2 | Recommended
RT55-01-002 | Audiology | DIAGNOSTIC OAE WITH STANDARD FREQUENCY .Description:Device used for diagnostic assessment of oto-acoustic emissions within the cochlear region. Technical Specification:Advanced clinical oto-acoustic emission system, able to record and archive completed tests Must be able to perform TEOAEs and DPOAEs. DPOAE frequencies: 1.5- 6 kHz, Intensities: 40-70 dB SPL SNR: 3-10 dB, Averaging Time: 0.5,1,2,4 sec. TEOAE frequencies: 0.7-4 kHz Intensity: 83 dB pe SPL SNR: 3-10 dB, Averaging Time: 8, 16, 32, 64 sec.Must include the following: bilateral probes.probe cable clips.nubs and probe pouch.Ability to connect to a computer. Must be able to indicate noise flow measurements.Laptop requirements are a minimum of 5GB free disk space, minimum 1GB of RAM, Latest Intel version (for the option of connecting to a laptop or PC for printing purposes) Accessories bilateral probes probe cable clips nubs and probe pouch ability to connect to a computer . | Navitrons | Electronique du Mazet | France | Echodia | BABYSCREEN - OAE (TEOAE and DPOAE) with Laptop | 2016 | 10 years | 2 (Offer 2) | 1 | 113500 | 13620 | 14982 | 16480.2 | 18128.22 | 19941.042 |  | 176766.42 | RT55-01-002 | 176766.42 | 176766.42 | 90 | 0 | 90 | 1 | Recommended
RT55-01-002 | Audiology | DIAGNOSTIC OAE WITH STANDARD FREQUENCY .Description:Device used for diagnostic assessment of oto-acoustic emissions within the cochlear region. Technical Specification:Advanced clinical oto-acoustic emission system, able to record and archive completed tests Must be able to perform TEOAEs and DPOAEs. DPOAE frequencies: 1.5- 6 kHz, Intensities: 40-70 dB SPL SNR: 3-10 dB, Averaging Time: 0.5,1,2,4 sec. TEOAE frequencies: 0.7-4 kHz Intensity: 83 dB pe SPL SNR: 3-10 dB, Averaging Time: 8, 16, 32, 64 sec.Must include the following: bilateral probes.probe cable clips.nubs and probe pouch.Ability to connect to a computer. Must be able to indicate noise flow measurements.Laptop requirements are a minimum of 5GB free disk space, minimum 1GB of RAM, Latest Intel version (for the option of connecting to a laptop or PC for printing purposes) Accessories bilateral probes probe cable clips nubs and probe pouch ability to connect to a computer . | Navitrons | Electronique du Mazet | France | Echodia | Elios (Article code: 300644) -OAE (TEOAE and DPOAE) with Laptop | 2011 | 10 years | 2 (Offer 1) | 1 | 145000 | 17400 | 19140 | 21054 | 23159.4 | 25475.34 |  | 225808.4 | RT55-01-002 | 225808.4 | 176766.42 | 65.0304 | 0 | 65.0304 | 2 | Recommended
RT55-01-003 | Audiology | ELECTROPHYSIOLOGY TESTING EQUIPMENT WITH ABR (AIR AND BONE CONDUCTION), ASSR, DPOAE Description:Device used for assessing the peripheral auditory system from the cochlear. Complete standalone unit which must be pre-loaded with diagnostic ABR, ASSR, middle latency response (MLR), late latency responses (LLR) and OAE software with the ability to add additional protocols of eABR, eCochG, VEMPs. Technical Specification: Complete evoked potential system able to perform diagnostic ABR (click, tone burst, chirp), auditory steady state response (ASSR), diagnostic TEOAES and DPOAEs, middle latency response (MLR), late latency responses (LLR), eABR, eCochG, VEMPs or more protocols ,Must include laptop, software, and free field aided testing capacity.Must come with a compatible laptop and software as required ,Must include a minimum of: 2 sets of cup electrodes 2 sets of disposable electrodes ,pack of electrode pads OAE nubs 2 packets of large insert earphones (yellow) 2 packets of small insert earphones (pink) 2 packets of infant insert earphones (red) 2 packets Goldfoil electrodes. Circumaural headset bone conductor probe cables. Trigger cables (eABR),carry case,2x Nuprep gel,2 tubs of EEG conduction paste. Accessories: TDH Headphones Bone vibrator,2 sets of cup electrodes 2 sets of disposable electrodes pack of electrode pads ,OAE nubs,2 packets of large insert earphones (yellow) 2 packets of small insert earphones (pink) 2 packets of infant insert earphones (red).2 packets Goldfoil electrodes,Circumaural headset bone conductor probe cables. Trigger cables (eABR),carry case Consumables 2x Nuprep gel. 2 tubs of EEG conduction paste Optional accessories Freefield speakers to be able to aided ABR’s .Available trigger cable to do eABR | Navitrons | Electronique du Mazet | France | Echodia | Elios (Article code: 300644) - Diagnostic ABR + ASSR + Screening AABR kit + OAE (TEOAE+DPOAE) + EcochG+oVEMS+cVEMPS With Laptop | 2011 | 10 years | 1 | 1 | 408500 | 49020 | 53922 | 59314.2 | 65245.62 | 71770.182 |  | 636058.82 | RT55-01-003 | 636058.82 | 636058.82 | 90 | 0 | 90 | 1 | Recommended
RT55-01-004 | Audiology | FRENZEL LENSES .Description : Hand-held, light source used to disable patient's ability to visually fixate on an object while allowing the examiner to adequately visualise the eye, with appropriate power source. Technical Specifications :Must include lenses and bulbs which allow for a large viewing area. Must be supplied with appropriate power source Accessories:lenses,Bulbs. | Navitrons | Instrumentation Difra | Belgium | Difra | Frenzel Lenses (003062-30) | 1993 | 10 Years | 3 (Offer 1) | 1 | 35000 | 4200 | 4620 | 5082 | 5590.2 | 6149.22 |  | 54550.2 | RT55-01-004 | 54550.2 | 54550.2 | 90 | 0 | 90 | 1 | Recommended
RT55-01-004 | Audiology | FRENZEL LENSES .Description : Hand-held, light source used to disable patient's ability to visually fixate on an object while allowing the examiner to adequately visualise the eye, with appropriate power source. Technical Specifications :Must include lenses and bulbs which allow for a large viewing area. Must be supplied with appropriate power source Accessories:lenses,Bulbs. | Lebone Medical Supplies (Pty) Ltd t/a Amtronix Diagnostics | Inventis S.R.L. | Italy | Inventis | VideoScope model – digital Wired video-Frenzel and Laptop | 2022 | 10 Years | 2 | 1 Unit | 79384.5 | 12650 | 13915 | 14674 | 15433 | 16445 | 0 | 136082.5 | RT55-01-004 | 136082.5 | 54550.2 | -44.5166 | 5.103 | -39.4136 | 2 | Recommended
RT55-01-005 | Audiology | Hearing Aid Earmould Drill Electrical Small Handheld with Bits .Description. Small electric hand drill complete with drill bits suitable for drilling earmoulds, pack of 5 different sizes of approximately 0.5, 1, 1.5, 2 and 2.5mm drill bits, buffer ball, grindstone and sand paper.Technical Specifications:With variable electronic speed control for high torque, even in the lower speed range. Various Bits and cutters for use in drill.Supplied in a carry case .Appropriate power supply. Accessories :Set of drillbits and accessories for a small electric earmould drill consisting of 5 different sizes of approximately 0.5, 1, 1.5, 2 and 2.5mm drill bits, buffer ball, grindstone and sand paper, for ear mould. To be compatible with awarded drill. | Navitrons | PROXONN | Germany | PROXONN | MICROMOT 60/E model building and engraving set - NO 28 515 |  | 5 Years | 1 | 1 | 6000 | [6 empty] | 6059 | RT55-01-005 | 6059 | 6059 | 90 | 0 | 90 | 1 | Recommended
RT55-01-008 | Audiology | Screening Audiometer :Description:Portable audiometer for hearing screening. Technical Specification: Compact, lightweight and portable screening audiometer Must be supplied with protective carry case.Must be supplied with a response button and 1 set of spare audiocups. 1 channel audiometer with air conduction, with a testable frequency range of 125Hz - 8000Hz. Must have decibel intensity range of -10dBHL to at least 100dBHL. Must be supplied with TDH39 or equivalent noise cancelling headphones. Must have A/C power supply and battery. Accessories:a response button.1 set of spare audiocups.Must be compliant to ANSI S3-6 (Type 4) or equivalent (certificate required).. | Navitrons | Amplivox | United Kingdom | Amplivox | Model 116-B with Audiocups | 2002 | 10 years | 3 (Offer 2) | 1 | 43000 | 5160 | 5676 | 6243.6 | 6867.96 | 7554.756 |  | 67008.56 | RT55-01-008 | 67008.56 | 67008.56 | 90 | 0 | 90 | 1 | Recommended
RT55-01-008 | Audiology | Screening Audiometer :Description:Portable audiometer for hearing screening. Technical Specification: Compact, lightweight and portable screening audiometer Must be supplied with protective carry case.Must be supplied with a response button and 1 set of spare audiocups. 1 channel audiometer with air conduction, with a testable frequency range of 125Hz - 8000Hz. Must have decibel intensity range of -10dBHL to at least 100dBHL. Must be supplied with TDH39 or equivalent noise cancelling headphones. Must have A/C power supply and battery. Accessories:a response button.1 set of spare audiocups.Must be compliant to ANSI S3-6 (Type 4) or equivalent (certificate required).. | Navitrons | Amplivox | United Kingdom | Amplivox | Model 170 with Audiocups (Without Battery) | 2005 | 10 years | 3 (Offer 3) | 1 | 53500 | 6420 | 7062 | 7768.2 | 8545.02 | 9399.522 |  | 83357.22 | RT55-01-008 | 83357.22 | 67008.56 | 68.0419 | 0 | 68.0419 | 2 | Recommended
RT55-01-008 | Audiology | Screening Audiometer :Description:Portable audiometer for hearing screening. Technical Specification: Compact, lightweight and portable screening audiometer Must be supplied with protective carry case.Must be supplied with a response button and 1 set of spare audiocups. 1 channel audiometer with air conduction, with a testable frequency range of 125Hz - 8000Hz. Must have decibel intensity range of -10dBHL to at least 100dBHL. Must be supplied with TDH39 or equivalent noise cancelling headphones. Must have A/C power supply and battery. Accessories:a response button.1 set of spare audiocups.Must be compliant to ANSI S3-6 (Type 4) or equivalent (certificate required).. | Navitrons | Electronique du Mazet | France | Echodia | AudioSmart (Article code 303246) with DD65 | 2011 | 10 Years | 3 (Offer 1) | 1 | 63000 | 7560 | 8316 | 9147.6 | 10062.36 | 11068.596 |  | 98145.96 | RT55-01-008 | 98145.96 | 67008.56 | 48.179 | 0 | 48.179 | 3 | Recommended
RT55-01-008 | Audiology | Screening Audiometer :Description:Portable audiometer for hearing screening. Technical Specification: Compact, lightweight and portable screening audiometer Must be supplied with protective carry case.Must be supplied with a response button and 1 set of spare audiocups. 1 channel audiometer with air conduction, with a testable frequency range of 125Hz - 8000Hz. Must have decibel intensity range of -10dBHL to at least 100dBHL. Must be supplied with TDH39 or equivalent noise cancelling headphones. Must have A/C power supply and battery. Accessories:a response button.1 set of spare audiocups.Must be compliant to ANSI S3-6 (Type 4) or equivalent (certificate required).. | Lebone Medical Supplies (Pty) Ltd t/a Amtronix Diagnostics | Maico Diagnostics GmbH | Germany | Maico | Maico MA25 | 2018 | 10 Years | 1 | 1 Unit | 46301.3 | 12650 | 13915 | 14674 | 15433 | 16445 | 4000 | 107000.3 | RT55-01-008 | 107000.3 | 67008.56 | 36.2866 | 5.103 | 41.3896 | 4 | Recommended
RT55-01-008 | Audiology | Screening Audiometer :Description:Portable audiometer for hearing screening. Technical Specification: Compact, lightweight and portable screening audiometer Must be supplied with protective carry case.Must be supplied with a response button and 1 set of spare audiocups. 1 channel audiometer with air conduction, with a testable frequency range of 125Hz - 8000Hz. Must have decibel intensity range of -10dBHL to at least 100dBHL. Must be supplied with TDH39 or equivalent noise cancelling headphones. Must have A/C power supply and battery. Accessories:a response button.1 set of spare audiocups.Must be compliant to ANSI S3-6 (Type 4) or equivalent (certificate required).. | Lebone Medical Supplies (Pty) Ltd t/a Amtronix Diagnostics | Inventis S.R.L. | Italy | Inventis | Inventis Triangle | 2017 | 10 Years | 2 | 1 Unit | 55443.593 | 12650 | 13915 | 14674 | 15433 | 16445 | 4000 | 116143.593 | RT55-01-008 | 116143.593 | 67008.56 | 24.0061 | 5.103 | 29.1091 | 5 | Recommended
RT55-01-010 | Audiology | Screening Audiometer and Tympanometer(226HZ AND 1000HZ). Description:Audiometry equipment used for screening of hearing, tympanometry and acoustic reflexes Technical Specification:Compact, lightweight and portable screening tympanometer and audiometer combination unit .Must be supplied with a carry case.Must have A/C power supply and battery.Tympanometry measurement capability of 226Hz and 1000Hz. Must have printing compatibility for audiometer and tympanometer.Must have a memory for at least 12 tests, with probe assembly Must be supplied with tympanometry probes/ear tips.Must be supplied with a box of tympanometry ear tips of at least six different sizes, with at least three tips of each size.Must include ipsilateral reflexes with frequency range from 500Hz to 4000Hz. Screening audiometer module must include a testable frequency range of 125Hz - 8000Hz and decibel intensity range of -10dBHL to at least 100dBHL.Must be supplied with TDH39 or equivalent headphones Must be supplied with 1 set of spare audiocups.Must be supplied with patient response button. Accessories:set of headphones set of insert earphones set of audio cups patient response button. | Lebone Medical Supplies (Pty) Ltd t/a Amtronix Diagnostics | Path Medical GmbH | Germany | Path Sentiero | Path Senteiro Desktop with Inserts and Headphones | 2012 - Continious firmware upgrades | 10 Years | 1 | 1 Unit | 123234 | 23000 | 25300 | 26680 | 28060 | 29900 | 5000 | 231303 | RT55-01-010 | 231303 | 231303 | 90 | 5.103 | 95.103 | 1 | Recommended
RT55-01-011 | Audiology | Screening Audiometer(With Insert Earphones.Description:Portable audiometer for hearing screening Technical Specification: Compact, lightweight and portable screening audiometer Must be supplied with protective carry case.Is a 1 channel audiometer with air conduction, and testable frequency range of 125Hz - 8000Hz Must have decibel intensity range of -10dBHL to at least 100dBHL.Must be supplied with TDH39 or equivalent noise cancelling headphones and insert earphones (EAR 3A or equivalent insert earphones to be included). Must have A/C power supply and battery. Accessories:a response button,1 set of spare audiocups to be included. | Lebone Medical Supplies (Pty) Ltd t/a Amtronix Diagnostics | Inventis S.R.L. | Italy | Inventis | Inventis Triangle | 2017 | 10 Years | 2 | 1 Unit | 56593.593 | 12650 | 13915 | 14674 | 15433 | 16445 | 4000 | 117295.593 | RT55-01-011 | 117295.593 | 117295.593 | 90 | 5.103 | 95.103 | 1 | Recommended
RT55-01-012 | Audiology | Technical Specifications of screening tympanometer (226Hz tone): Tympanometry measurement capability of 226Hz.Must test tympanometry across range of +200 -to -400daPa with accuracy of ±10daPa, Ipsi-lateral testing range from 500,1000,2000 and 4000Hz.Must include probe/ear tips .Must be supplied with protective carry case.Must have A/C power supply. Must have printing capabilities and supplied with two rolls of paper.Must have a memory for at least 12 tests, with probe assembly.Must be supplied with a box of ear tips of at least six different sizes, with at least three tips of each size.Accessories:probe/ear tips of at least six different sizes, with at least three tips of each size,protective carry case,two rolls of paper,memory for at least 12 tests,with probe assembly.Box of ear tips .Optional accessory :Printer. | Lebone Medical Supplies (Pty) Ltd t/a Amtronix Diagnostics | Inventis S.R.L. | Italy | Inventis | Inventis Timpani | 2012 - Continious firmware upgrades | 10 Years | 1 | 1 Unit | 67827 | 27025 | 29727.5 | 31349 | 32970.5 | 35132.5 | 4000 | 192930 | RT55-01-012 | 192930 | 192930 | 90 | 5.103 | 95.103 | 1 | Recommended
RT55-01-012 | Audiology | Technical Specifications of screening tympanometer (226Hz tone): Tympanometry measurement capability of 226Hz.Must test tympanometry across range of +200 -to -400daPa with accuracy of ±10daPa, Ipsi-lateral testing range from 500,1000,2000 and 4000Hz.Must include probe/ear tips .Must be supplied with protective carry case.Must have A/C power supply. Must have printing capabilities and supplied with two rolls of paper.Must have a memory for at least 12 tests, with probe assembly.Must be supplied with a box of ear tips of at least six different sizes, with at least three tips of each size.Accessories:probe/ear tips of at least six different sizes, with at least three tips of each size,protective carry case,two rolls of paper,memory for at least 12 tests,with probe assembly.Box of ear tips .Optional accessory :Printer. | Lebone Medical Supplies (Pty) Ltd t/a Amtronix Diagnostics | Maico Diagnostics GmbH | Germany | Maico | Maico easyTymp | 2012 - Continious firmware upgrades | 10 Years | 2 | 1 Unit | 73140 | 27025 | 29727.5 | 31349 | 32970.5 | 35132.5 | 4000 | 198244 | RT55-01-012 | 198244 | 192930 | 87.5211 | 5.103 | 92.6241 | 2 | Recommended
RT55-01-013 | Audiology | VISUAL RESPONSE AUDIOMETRY UNIT (VIDEO).Description:Equipment used together with audiometer for paediatric hearing evaluations consisting of toy cabinets or video re-inforcers and sound-field speakers. Technical Specifications :A Video VRA reinforcer.Two colour LED flatscreen approximately 15" - 20" monitors that display motion action cartoon images for the use during paediatric hearing testing Switch must be able to be installed in the observation room to control animation (i.e. on and off), and switch on the screen. No time elapse should occur between the initiation of VRA system and illumination and activation of toy reinforcer/ screen.Must include x2 mounting bracket for installation of the screens Each screen should be able to be lit up independently.2 colour monitors, compatible with diagnostic audiometer and intermediary compute/laptop system with HDMI connecting cables to be included if system requires.Reinforcement activated by push of button by either wireless remote or through computer system.Silent switch that can be connected to the audiometer for easy changes between right and left monitors.Variety of animated characters in reinforcement playlist. Must include x2 mounting bracket for installation to mount flush against wall or on arm lever Must be compatible with audiometer installed and sound field speakers if not included. Must include AC power supplies Must include switch for easy changes between right and left monitors. | Lebone Medical Supplies (Pty) Ltd t/a Amtronix Diagnostics | Lebone Medical Supplies (Pty) Ltd t/a Amtronix Diagnostics | South Africa | Amtronix | Amtronix Digital VRA | 2022 | 10 Years | 1 | 1 Unit | 36225 | 12075 | 13282.5 | 14007 | 14731.5 | 15697.5 | 0 | 90354 | RT55-01-013 | 90354 | 90354 | 90 | 5.103 | 95.103 | 1 | Recommended
RT55-02-001 | Occupational Therapy | 90-DEGREE WIRE BENDER Description: Handheld tool to bend wire of various thickness Technical Specification:90° wire bender suited for bending copper wire to 90 degrees, used for dynamic outriggers Ergonomic, soft grip handle for comfort and reduced stress on hand. Safety latch to keep handles and jaws closed. | Duromed cc | North Coast Medical | USA | NCM | NC12569 | N/A | 2 YEARS | OFFER 1 | EACH | 1897.5 | 0 | 0 | 0 | 0 | 0 |  | 1901.5 | RT55-02-001 | 1901.5 | 1901.5 | 90 | 6.9 | 96.9 | 1 | Recommended
RT55-02-001 | Occupational Therapy | 90-DEGREE WIRE BENDER Description: Handheld tool to bend wire of various thickness Technical Specification:90° wire bender suited for bending copper wire to 90 degrees, used for dynamic outriggers Ergonomic, soft grip handle for comfort and reduced stress on hand. Safety latch to keep handles and jaws closed. | Sanbonani Holdings (Pty) Ltd | North Coast Medical Inc. | USA | North Coast Medical Inc. | NC12569 | 2004 | 6 years | 1 | Each | 1894 | [6 empty] | 1977 | RT55-02-001 | 1977 | 1901.5 | 86.4265 | 10 | 96.4265 | 2 | Recommended
RT55-02-001 | Occupational Therapy | 90-DEGREE WIRE BENDER Description: Handheld tool to bend wire of various thickness Technical Specification:90° wire bender suited for bending copper wire to 90 degrees, used for dynamic outriggers Ergonomic, soft grip handle for comfort and reduced stress on hand. Safety latch to keep handles and jaws closed. | HITECH THERAPY | PERFORMANCE HEALTH | UK | PERFORMANCE HEALTH | 7031640 | 2014 | 7 YEARS | 1 | N/A | 2990 | 0 | 0 | 0 | 0 | 0 |  | 3105 | RT55-02-001 | 3105 | 1901.5 | 33.0371 | 10 | 43.0371 | 3 | Recommended
RT55-02-002 | Occupational Therapy | COMBINATION PLIERS, LONG NOSE Description: Long nose combination pliers are the same as traditional, but instead of a flat nose, the pliers have a long, tapered nose, which allows you to get into tighter spaces. Technical Specification: Robust and temperature resistant steel.Approximately 180mm-200mm total length Ergonomic handles covered in soft durable rubber/plastic | KUSASALETU INVESTMENT GROUP | BUILT IT | RSA | BUILT IT | 10011349 | 2024 | LIFETIME | 1 | 1 | 180 | 0 | 0 | 0 | 0 | 0 |  | 201 | RT55-02-002 | 201 | 201 | 90 | 10 | 100 | 1 | Recommended
RT55-02-008 | Occupational Therapy | SCOOTER BOARD Description: Therapeutic equipment used in the treatment of balance and gross motor development for paediatrics Technical Specification:Wooden or plastic durable base padded with soft foam and covered with vinyl or similar waterproof easy to clean material mounted on 4 durable castors. Approximate size 40x60x10cm Suitable to hold the weight of a child. | Sanbonani Holdings (Pty) Ltd | Geppetto's World of Developments cc t/a Playon | South Africa | Playon | GW-SB4 | 2003 | 8 years | 1 | Each | 1180 | [6 empty] | 1264 | RT55-02-008 | 1264 | 1264 | 90 | 10 | 100 | 1 | Recommended
RT55-02-011 | Occupational Therapy | SLIDING WEDGE WITHOUT VELCRO (Set of 3) Description: Therapeutic tool used for functional physical rehabilitation Technical Specification: Wedge to provide positioning in supine or prone therapy activities. High density wedges: in various sizes, multiple colours. Easy to clean, quality leather or flame-retardant vinyl. Hard firm wedges padded and covered with smooth vinyl, PVC or similar 3 dimensional blocks; Without velcro attachment in the following sizes: a) 15cm x 60cmx65cm b) 20cmx60cmx65cm c) 25cmx60cmx65cm | HITECH THERAPY | GEPPETTO | SOUTH AFRICA | SYNERGY | HTT6015 | 2003 | 7 YEARS | 1 | CM | 6555 | 0 | 0 | 0 | 0 | 0 |  | 6671 | RT55-02-011 | 6671 | 6671 | 90 | 10 | 100 | 1 | Recommended
RT55-02-012 | Occupational Therapy | Therapy Foam Rollers (set of 5 X 60 cm length) Description:A set of therapeutic tools used for functional physical rehabilitation Technical Specification:High density foam rollers: in various sizes (long & round), multiple colours. Cleanable quality leather or flame-retardant vinyl.Rollers are made of Polystyrene, covered with chip foam and cover. Therapy Foam Rollers: Approximate dimensions: a) 60cm length and 50cm diameter. Therapy Foam Rollers: Approximate dimensions b) 60cm length and 40cm diameter.Therapy Foam Rollers: Approximate dimensions c) 60cm length and 30cm diameter.Therapy Foam Rollers :Approximate dimensions d) 60cm length and 20cm diameter.Therapy Foam Rollers :Approximate dimensions e) 60cm length and 15cm diameter | Sanbonani Holdings (Pty) Ltd | Geppetto's World of Developments cc t/a Playon | South Africa | Playon | GP-ROS605 | 2003 | 8 years | 1 | Set of 5 | 4925 | [6 empty] | 5010 | RT55-02-012 | 5010 | 5010 | 90 | 0 | 90 | 1 | Recommended
RT55-02-012 | Occupational Therapy | Therapy Foam Rollers (set of 5 X 60 cm length) Description:A set of therapeutic tools used for functional physical rehabilitation Technical Specification:High density foam rollers: in various sizes (long & round), multiple colours. Cleanable quality leather or flame-retardant vinyl.Rollers are made of Polystyrene, covered with chip foam and cover. Therapy Foam Rollers: Approximate dimensions: a) 60cm length and 50cm diameter. Therapy Foam Rollers: Approximate dimensions b) 60cm length and 40cm diameter.Therapy Foam Rollers: Approximate dimensions c) 60cm length and 30cm diameter.Therapy Foam Rollers :Approximate dimensions d) 60cm length and 20cm diameter.Therapy Foam Rollers :Approximate dimensions e) 60cm length and 15cm diameter | HITECH THERAPY | GEPPETTO | SOUTH AFRICA | SYNERGY | HTT4015 | 2003 | 7 YEARS | 1 | CM | 6555 | 0 | 0 | 0 | 0 | 0 |  | 6672 | RT55-02-012 | 6672 | 5010 | 60.1437 | 0 | 60.1437 | 2 | Recommended
RT55-02-015 | Occupational Therapy | BALANCE BEAM SET Description: Therapeutic equipment used in the treatment of balance and gross motor development for paediatrics Technical specification:2 bases with removable half round or square beam.Covered in waterproof, durable, easy to clean, vinyl or similar material. Hard non-yielding inner with high density foam padding.Velcro to secure beam to bases .Each Base: 60 x 45x 45cm, with a cut out to fit the balance beam Balance Beam: 180cm (L) x 30cm x 15cm | Sanbonani Holdings (Pty) Ltd | Geppetto's World of Developments cc t/a Playon | South Africa | Playon | GP-BARP18 | 2012 | 8 years | 1 | Set | 6465 | [6 empty] | 6551 | RT55-02-015 | 6551 | 6551 | 90 | 0 | 90 | 1 | Recommended
RT55-02-015 | Occupational Therapy | BALANCE BEAM SET Description: Therapeutic equipment used in the treatment of balance and gross motor development for paediatrics Technical specification:2 bases with removable half round or square beam.Covered in waterproof, durable, easy to clean, vinyl or similar material. Hard non-yielding inner with high density foam padding.Velcro to secure beam to bases .Each Base: 60 x 45x 45cm, with a cut out to fit the balance beam Balance Beam: 180cm (L) x 30cm x 15cm | HITECH THERAPY | GEPPETTO | SOUTH AFRICA | PLAY ON | HTT7015 | 2003 | 7 YEARS | 1 | CM | 11500 | 0 | 0 | 0 | 0 | 0 |  | 11618 | RT55-02-015 | 11618 | 6551 | 20.3877 | 0 | 20.3877 | 2 | Recommended
RT55-02-016 | Occupational Therapy | BALL PIT Description: Equipment, typically filled with hollow balls, used in the treatment of paediatrics to promote gross motor development, sensory input or as a positioning device. Technical Specification:Ball pond 150x150x45cm.Square ball pit, Firm inner with padding to prevent injury .Durable, Tarpaulin, vinyl or similar easy to clean waterproof material. Accessories:100 Hollow plastic balls: manufactured of durable, lightweight plastic, supplied in a variety of colours, balls approximately 10cm diameter. | Sanbonani Holdings (Pty) Ltd | Geppetto's World of Developments cc t/a Playon | South Africa | Playon | GW-BPH+B | 2018 | 8 years | 1 | Set | 7100 | [6 empty] | 7187 | RT55-02-016 | 7187 | 7187 | 90 | 0 | 90 | 1 | Recommended
RT55-02-016 | Occupational Therapy | BALL PIT Description: Equipment, typically filled with hollow balls, used in the treatment of paediatrics to promote gross motor development, sensory input or as a positioning device. Technical Specification:Ball pond 150x150x45cm.Square ball pit, Firm inner with padding to prevent injury .Durable, Tarpaulin, vinyl or similar easy to clean waterproof material. Accessories:100 Hollow plastic balls: manufactured of durable, lightweight plastic, supplied in a variety of colours, balls approximately 10cm diameter. | HITECH THERAPY | GEPPETTO | SOUTH AFRICA | PLAY ON | 7035 | 2003 | 7 YEARS | 1 | CM | 8625 | 0 | 0 | 0 | 0 | 0 |  | 8744 | RT55-02-016 | 8744 | 7187 | 70.5023 | 0 | 70.5023 | 2 | Recommended
RT55-02-017 | Occupational Therapy | COMBINATION PLIERS, FLAT Description:Pliers are a hand tool used to hold objects firmly, bend and physically compress a wide range of materials. Combination pliers’ dual functionality allows for both cutting and gripping. Technical Specification:These pliers have one scissor-type jaw opening, where half of the length is flat and used for gripping, and the other half is a cutting edge used for cutting wire. Robust and temperature resistant steel.Approximately 180mm-200mm total length. Ergonomic handles covered in soft durable rubber/plastic | KUSASALETU INVESTMENT GROUP | BUILT IT | RSA | BUILT IT | 1101000000 | 2024 | LIFETIME | 1 | 1 | 210 | 0 | 0 | 0 | 0 | 0 |  | 232 | RT55-02-017 | 232 | 232 | 90 | 10 | 100 | 1 | Recommended
RT55-02-019 | Occupational Therapy | EYELET PUNCH Description: A tool used for the securing of eyelets in place. Technical Specifications:Hand-held punch to be used to securing eyelets. Suitable for use with 5mm eyelets.Ergonomic, soft grip handle. Accessories: 100-125 piece (pack) steel eyelets.All above listed accessories must be supplied standard with the unit.Accessories must be available to purchase separately. | Sanbonani Holdings (Pty) Ltd | North Coast Medical Inc. | USA | North Coast Medical Inc. | NC12714 | 2004 | 6 years | 1 | Each | 892 | [6 empty] | 980 | RT55-02-019 | 980 | 980 | 90 | 0 | 90 | 1 | Recommended
RT55-02-020 | Occupational Therapy | FINGER CIRCUMFERENCE GAUGE READING Description: A set of Instrument used to test grip strength and other components of hand function Technical Specification:Used to measure the circumference of a finger Finger circumference gauge webbing loop wraps around the finger and the circumference is read off the ruler Must be measured in inches and centimetres. | Duromed cc | North Coast Medical | USA | NCM EXACTA | NC70158 | N/A | 2 YEARS | OFFER 1 | EACH | 379.5 | 0 | 0 | 0 | 0 | 0 |  | 384.5 | RT55-02-020 | 384.5 | 384.5 | 90 | 6.9 | 96.9 | 1 | Recommended
RT55-02-020 | Occupational Therapy | FINGER CIRCUMFERENCE GAUGE READING Description: A set of Instrument used to test grip strength and other components of hand function Technical Specification:Used to measure the circumference of a finger Finger circumference gauge webbing loop wraps around the finger and the circumference is read off the ruler Must be measured in inches and centimetres. | Mediotronics Physical Medicine Pty Ltd | Saehan Corporation | Korean | Saehan | SH5023-1 | 2016 | 5 Years | 1 | 1 / each | 365 | 0 | 0 | 0 | 0 | 0 |  | 399 | RT55-02-020 | 399 | 384.5 | 86.606 | 0 | 86.606 | 2 | Recommended
RT55-02-020 | Occupational Therapy | FINGER CIRCUMFERENCE GAUGE READING Description: A set of Instrument used to test grip strength and other components of hand function Technical Specification:Used to measure the circumference of a finger Finger circumference gauge webbing loop wraps around the finger and the circumference is read off the ruler Must be measured in inches and centimetres. | HITECH THERAPY | FABRICATION ENTERPRISES | USA | BASELINE | 121221 | 2016 | 7 YEARS | 1 | INCH + CM | 644 | 0 | 0 | 0 | 0 | 0 |  | 764 | RT55-02-020 | 764 | 384.5 | 1.1704 | 0 | 1.1704 | 3 | Recommended
RT55-02-021 | Occupational Therapy | FINGER GONIOMETER (PLASTIC) Description: Used to measure range of motion of the fingers. Technical Specification:Made of durable hard plastic Must have 180 degrees scale marked in 1-degree increments that read in either direction. Must be able to measure hyperextension of the finger (0 - 120° with hyperextension 0 - 30°) Linear in inches and centimeters. | HITECH THERAPY | FABRICATION ENTERPRISES | USA | BASELINE | 121014 | 2016 | 7 YEARS | 1 | INCH + CM | 644 | 0 | 0 | 0 | 0 | 0 |  | 765 | RT55-02-021 | 765 | 765 | 90 | 0 | 90 | 1 | Recommended
RT55-02-022 | Occupational Therapy | FINGER GONIOMETER (STAINLESS STEEL) Description:An instrument used to measure joint range of motion Technical Specification:Used to measure range of motion of the fingers. Made of stainless steel. Must have 180 degrees scale marked in 1-degree increments that read in either direction. Increment markings etched into metal (not only painted on) Must be able to measure hyperextension of the finger (0 - 120° with hyperextension 0 - 30°) Linear in inches and centimeters. | Mediotronics Physical Medicine Pty Ltd | Saehan Corporation | Korean | Saehan | SH5100 | 2016 | 5 Years | 1 | 1 / each | 461 | 0 | 0 | 0 | 0 | 0 |  | 496 | RT55-02-022 | 496 | 496 | 90 | 0 | 90 | 1 | Recommended
RT55-02-022 | Occupational Therapy | FINGER GONIOMETER (STAINLESS STEEL) Description:An instrument used to measure joint range of motion Technical Specification:Used to measure range of motion of the fingers. Made of stainless steel. Must have 180 degrees scale marked in 1-degree increments that read in either direction. Increment markings etched into metal (not only painted on) Must be able to measure hyperextension of the finger (0 - 120° with hyperextension 0 - 30°) Linear in inches and centimeters. | Sanbonani Holdings (Pty) Ltd | North Coast Medical Inc. | USA | Saehan | NC70098 | 2001 | 10 years | 2 | Each | 625 | [6 empty] | 714 | RT55-02-022 | 714 | 496 | 50.4435 | 0 | 50.4435 | 2 | Recommended
RT55-02-022 | Occupational Therapy | FINGER GONIOMETER (STAINLESS STEEL) Description:An instrument used to measure joint range of motion Technical Specification:Used to measure range of motion of the fingers. Made of stainless steel. Must have 180 degrees scale marked in 1-degree increments that read in either direction. Increment markings etched into metal (not only painted on) Must be able to measure hyperextension of the finger (0 - 120° with hyperextension 0 - 30°) Linear in inches and centimeters. | Duromed cc | North Coast Medical | USA | NCM | NC70108 | N/A | 2 YEARS | OFFER 1 | EACH | 828 | 0 | 0 | 0 | 0 | 0 |  | 834 | RT55-02-022 | 834 | 496 | 28.6694 | 6.9 | 35.5694 | 3 | Recommended
RT55-02-022 | Occupational Therapy | FINGER GONIOMETER (STAINLESS STEEL) Description:An instrument used to measure joint range of motion Technical Specification:Used to measure range of motion of the fingers. Made of stainless steel. Must have 180 degrees scale marked in 1-degree increments that read in either direction. Increment markings etched into metal (not only painted on) Must be able to measure hyperextension of the finger (0 - 120° with hyperextension 0 - 30°) Linear in inches and centimeters. | HITECH THERAPY | FABRICATION ENTERPRISES | USA | BASELINE | 121016 | 2016 | 7 YEARS | 1 | INCH + CM | 851 | 0 | 0 | 0 | 0 | 0 |  | 973 | RT55-02-022 | 973 | 496 | 3.4476 | 0 | 3.4476 | 4 | Recommended
RT55-02-023 | Occupational Therapy | GONIOMETER (30CM) Description:An instrument used to measure joint range of motion Technical Specification: Used to measure range of motion of large joints of the human body Made of durable hard plastic Must have 0-360° scale.Must have 1-2° degree increments Linear in inches and centimetres | Mediotronics Physical Medicine Pty Ltd | Saehan Corporation | Korean | Saehan | SH5203 | 2016 | 5 Years | 1 | 1 / each | 225 | 0 | 0 | 0 | 0 | 0 |  | 261 | RT55-02-023 | 261 | 261 | 90 | 0 | 90 | 1 | Recommended
RT55-02-023 | Occupational Therapy | GONIOMETER (30CM) Description:An instrument used to measure joint range of motion Technical Specification: Used to measure range of motion of large joints of the human body Made of durable hard plastic Must have 0-360° scale.Must have 1-2° degree increments Linear in inches and centimetres | Duromed cc | North Coast Medical | USA | NCM EXACTA | NC70102 | N/A | 2 YEARS | OFFER 1 | EACH | 391 | 0 | 0 | 0 | 0 | 0 |  | 398 | RT55-02-023 | 398 | 261 | 42.7586 | 6.9 | 49.6586 | 2 | Recommended
RT55-02-024 | Occupational Therapy | GONIOMETER 180° (15-20CM) Description:An instrument used to measure joint range of motion Technical Specification:Used to measure range of motion of medium joints of the body. Made of durable hard plastic Must have 0-180° scale.Must have 1-2° degree increments (Rulong) Must have linear in inches and centimeters | Mediotronics Physical Medicine Pty Ltd | Saehan Corporation | Korean | Saehan | SH5200 | 2016 | 5 Years | 1 | 1 / each | 185 | 0 | 0 | 0 | 0 | 0 |  | 222 | RT55-02-024 | 222 | 222 | 90 | 0 | 90 | 1 | Recommended
RT55-02-025 | Occupational Therapy | GONIOMETER 360° (15-20CM) Description:An instrument used to measure joint range of motion. Technical Specification:Used to measure range of motion of medium joints of the body. Made of durable hard plastic Must have 0-360° scale.Must have 1-2° degree increments Linear in inches and centimetres | Mediotronics Physical Medicine Pty Ltd | Saehan Corporation | Korean | Saehan | SH5202 | 2016 | 5 Years | 1 | 1 / each | 155 | 0 | 0 | 0 | 0 | 0 |  | 193 | RT55-02-025 | 193 | 193 | 90 | 0 | 90 | 1 | Recommended
RT55-02-026 | Occupational Therapy | GONIOMETER EXTENDABLE Description:A versatile instrument used to measure large and small joint range of motion Technical Specification:Made of durable hard plastic.Easy to read Includes 4x magnifying lens to read the dual 180-degree high resolution scale, in 1° increments Includes 2 stainless steel arms with a total collapsed length of approximately 20 cm and total extended length of approximately 70 cm. | HITECH THERAPY | FABRICATION ENTERPRISES | USA | LAFAYETTE INSTRUMENTS | 121035 | 2016 | 7 YEARS | 1 | DEGREES | 2300 | 0 | 0 | 0 | 0 | 0 |  | 2423 | RT55-02-026 | 2423 | 2423 | 90 | 0 | 90 | 1 | Recommended
RT55-02-026 | Occupational Therapy | GONIOMETER EXTENDABLE Description:A versatile instrument used to measure large and small joint range of motion Technical Specification:Made of durable hard plastic.Easy to read Includes 4x magnifying lens to read the dual 180-degree high resolution scale, in 1° increments Includes 2 stainless steel arms with a total collapsed length of approximately 20 cm and total extended length of approximately 70 cm. | Mediotronics Physical Medicine Pty Ltd | MVS in Motion | Belgium | Baseline | MVS12-1035 | 2016 | 5 years | 1 | 1 / each | 3000 | 0 | 0 | 0 | 0 | 0 |  | 3039 | RT55-02-026 | 3039 | 2423 | 67.1193 | 0 | 67.1193 | 2 | Recommended
RT55-02-026 | Occupational Therapy | GONIOMETER EXTENDABLE Description:A versatile instrument used to measure large and small joint range of motion Technical Specification:Made of durable hard plastic.Easy to read Includes 4x magnifying lens to read the dual 180-degree high resolution scale, in 1° increments Includes 2 stainless steel arms with a total collapsed length of approximately 20 cm and total extended length of approximately 70 cm. | Duromed cc | North Coast Medical | USA | NCM | NC70110 | N/A | 2 YEARS | OFFER 1 | EACH | 3220 | 0 | 0 | 0 | 0 | 0 |  | 3228 | RT55-02-026 | 3228 | 2423 | 60.0991 | 6.9 | 66.9991 | 3 | Recommended
RT55-02-027 | Occupational Therapy | GONIOMETER SET.Description: An instrument used to measure joint range of motion. Set should include at least 4 commonly used goniometers in a carrying case with compartments that holds goniometers firmly in place:Technical Specification:A complete set must include the below:a.Finger goniometer,Used to measure range of motion of the fingers. Made of durable hard plastic, Must have 180 degrees scale marked in 1-degree increments that read in either direction. Must be able to measure hyperextension of the finger (0 - 120° with hyperextension 0 - 30°). Linear in inches and centimeters.Goniometer 15-20cm,Used to measure range of motion of medium joints of the body. Made of durable hard plastic.Must have 0-180° scale. Must have 1-2° degree increments (Rulong) Must have linear in inches and centimeters.Goniometer 15-20cm,Used to measure range of motion of medium joints of the body. Made of durable hard plastic. Must have 0-360° scale. Must have 1-2° degree increments Linear in inches and centimetres. Goniometer 30 cm,Used to measure range of motion of large joints of the human body. Made of durable hard plastic. Must have 0-360° scale,Must have 1-2° degree increments Linear in inches and centimetres. Accessory:Must include carry case: durable fabric or plastic case with compartments that holds goniometers firmly in place with fasteners or slots. | Mediotronics Physical Medicine Pty Ltd | Saehan Corporation | Korean | Saehan | SH5206 | 2016 | 5 Years | 1 | 1 / each | 1522.5 | 0 | 0 | 0 | 0 | 0 |  | 1562.5 | RT55-02-027 | 1562.5 | 1562.5 | 90 | 0 | 90 | 1 | Recommended
RT55-02-029 | Occupational Therapy | HALF WHEEL Description: Therapeutic equipment used to promote gross motor development, sensory input or as a positioning device. Technical Specification :Half wheel shaped climbing and positioning apparatus. Hard foam/Compact non-yielding foam base padded with high density foam and covered with vinyl or similar waterproof easy to clean material. Velcro secured to the flat surface to allow for 2 half wheels to be combined to create a tunnel.Measurements approximately 120 x 45 x 60cm | Sanbonani Holdings (Pty) Ltd | Geppetto's World of Developments cc t/a Playon | South Africa | Playon | GP-BAHW120 | 2018 | 8 years | 1 | Each | 3100 | [6 empty] | 3190 | RT55-02-029 | 3190 | 3190 | 90 | 0 | 90 | 1 | Recommended
RT55-02-030 | Occupational Therapy | HAMMER, CLAW, STEEL Description:Basic hand-held, claw hammer, metal head to drive in or remove nails from wood, shape metal, or to crush rock. Technical Specification:Steel head with claw.Ergonomic, soft grip handle | KUSASALETU INVESTMENT GROUP | BUILT IT | RSA | BUILT IT | 110100025 | 2024 | LIFETIME | 1 | 1 | 284 | 0 | 0 | 0 | 0 | 0 |  | 307 | RT55-02-030 | 307 | 307 | 90 | 10 | 100 | 1 | Recommended
RT55-02-031 | Occupational Therapy | HEAT GUN Description:A device used in the construction of thermoplastic splints and warm up splinting material for adjustments. Technical Specification:Electrical heat gun with adjustable variable temperature (50 - 600 deg C) Air flow minimum 2000 watts.Flat base or built in safety stand. Compact and light weight with insulated ergonomically designed handle Additional air gun nozzles for directed heat flow.Built-in self-cooling system. Automatic switch off during overheating Must be supplied with various sizes and shapes of nozzles On/off switch, which does not need to be held in during use. Accessories:Spare reduction nozzles with a minimum of 3 with variety of sizes Stand, if not flatbase.Must include Hard carry case. Standard length electrical cord SA compatible with 3 pin plug, 220 -250 volts, minimum 1-year warranty All above listed accessories must be supplied standard with the unit. Accessories highlighted in bold must be available to purchase separate price list. Any other additional accessories and spare parts not listed above but applicable for the functionality of the unit must be supplied on a separate price list | Sanbonani Holdings (Pty) Ltd | Steinel | Gernany | Steinel | ST351502 | 2016 | 10 years | 2 | Kit | 6987.95 | 2400 | 2592 | 2747.52 | 2912.3712 | 3145.3609 |  | 17730.8412 | RT55-02-031 | 17730.8412 | 17730.8412 | 90 | 0 | 90 | 1 | Recommended
RT55-02-032 | Occupational Therapy | KNIFE, RETRACTABLE BLADE (Utility knife).Description:A tool used for scoring/cutting thermoplastic material, masonite, tubing and similar materials. Technical Specification:Durable, heavy duty, sharp stainless-steel utility knife with quick change steel heavy duty blades that is retractable and locks into place to ensure safety. Must be supplied with a pack of 5 blades.C246The knife handle must be curved and user friendly. Offers variable cut depth, with a thumb pad and ergonomically cutting angle and works with most standard utility blades. Accessories:Pack of 5 blades. Blades for retractable blade Knife.Pack of 5 compatible blades as replacement blades for retractable blade knife | Duromed cc | North Coast Medical | USA | STANLEY | NC12712 | N/A | 2 YEARS | OFFER 1 | EACH | 805 | 0 | 0 | 0 | 0 | 0 |  | 814 | RT55-02-032 | 814 | 814 | 90 | 6.9 | 96.9 | 1 | Recommended
RT55-02-033 | Occupational Therapy | KNIFE, RETRACTABLE BLADE (Utility knife).Description:A tool used for scoring/cutting thermoplastic material, masonite, tubing and similar materials. Technical Specification:Durable, heavy duty, sharp stainless-steel utility knife with quick change steel heavy duty blades that is retractable and locks into place to ensure safety. Must be supplied with a pack of 5 blades.The knife handle must be curved and user friendly.Offers variable cut depth, with a thumb pad and ergonomically cutting angle and works with most standard utility blades. Accessories:Pack of 5 blades. Blades for retractable blade Knife.Pack of 5 compatible blades as replacement blades for retractable blade knife. | Duromed cc | North Coast Medical | USA | STANLEY | NC12713 (Blades Only) | N/A | 2 YEARS | OFFER 1 | PACK OF 5 | 149.5 | 0 | 0 | 0 | 0 | 0 |  | 159.5 | RT55-02-033 | 159.5 | 159.5 | 90 | 6.9 | 96.9 | 1 | Recommended
RT55-02-036 | Occupational Therapy | MOBILE POSTURE MIRROR (STEEL).Description:A therapy tool used for visual feedback during mirror therapy and functional rehabilitation Technical Specification:At least 4 four replaceable castor wheels, with swivel/ tilt mirror (0-360°) on a hinged frame with angle/mirror tilt lock mechanism. Mirror to be made from minimum of 15mm toughened, shatter proof glass and have suitable backing. Stand is made from durable powder coated steel allowing it to be cleaned and sterilized. Rectangular mirror: Approximate dimensions: (W) 40-50cm (H) 1,5-2m.Free standing on lockable, castor wheels of size approximate 75-100mm Made of durable material ,easy to clean and disinfect. Accessories:Set of 4 compatible castors. Pair of Adjustable Tilt Knobs. | Mediotronics Physical Medicine Pty Ltd | Enraf-Nonius | Netherlands/Germany | Enraf-Nonius | 3446-311SA | 2010 | 10 Years | 1 | 1 / Each | 3880 | 0 | 0 | 0 | 0 | 0 |  | 3921 | RT55-02-036 | 3921 | 3921 | 90 | 0 | 90 | 1 | Recommended
RT55-02-036 | Occupational Therapy | MOBILE POSTURE MIRROR (STEEL).Description:A therapy tool used for visual feedback during mirror therapy and functional rehabilitation Technical Specification:At least 4 four replaceable castor wheels, with swivel/ tilt mirror (0-360°) on a hinged frame with angle/mirror tilt lock mechanism. Mirror to be made from minimum of 15mm toughened, shatter proof glass and have suitable backing. Stand is made from durable powder coated steel allowing it to be cleaned and sterilized. Rectangular mirror: Approximate dimensions: (W) 40-50cm (H) 1,5-2m.Free standing on lockable, castor wheels of size approximate 75-100mm Made of durable material ,easy to clean and disinfect. Accessories:Set of 4 compatible castors. Pair of Adjustable Tilt Knobs. | New Horizon Metals cc | New Horizon Metals cc | South Africa | NHM | MPM001 | 2022/2023 | 7 years | 1 | Each | 4738 | 0 | 0 | 0 | 0 | 0 |  | 4809 | RT55-02-036 | 4809 | 3921 | 69.6174 | 0 | 69.6174 | 2 | Recommended
RT55-02-038 | Occupational Therapy | MUNCH MAT/FLOP MAT Description: A foam filled mat placed on the floor to promote safety as a soft surface during the use of suspended equipment. Technical Specification:Outer: Tarpaulin or similar waterproof easy to clean material Inner: Soft malleable material, chip foam pieces or similar.Suitable for use as safety equipment below swings Approximate size: 200x150x15cm. | Sanbonani Holdings (Pty) Ltd | Geppetto's World of Developments cc t/a Playon | South Africa | Playon | GP-MA21515 | 2016 | 8 years | 1 | Each | 3625 | [6 empty] | 3717 | RT55-02-038 | 3717 | 3717 | 90 | 0 | 90 | 1 | Recommended
RT55-02-045 | Occupational Therapy | RAMP Description: Therapeutic equipment used in the treatment of balance and gross motor development for paediatrics Technical Specification:Approximately 200cm (L) x 35 cm (H) x 60 cm (W) Wood or hard durable plastic ramp with platform at the top. Raised edges to keep the scooter board within the boundaries of the ramp and platform. Durable wood with no exposed nails or screws. Suitable attachments to attach a rope or similar apparatus at the top of the ramp. Minimum weight capacity of approximately 100kgs | Sanbonani Holdings (Pty) Ltd | Geppetto's World of Developments cc t/a Playon | South Africa | Playon | GW-SB4 | 2003 | 8 years | 1 | Each | 6750 | [6 empty] | 6843 | RT55-02-045 | 6843 | 6843 | 90 | 0 | 90 | 1 | Recommended
RT55-02-046 | Occupational Therapy | REVOLVING PUNCH Description: Handheld tool used to punch holes in various materials such as leather, thermoplastic & vinyl Technical Specifications:High quality heavy duty durable revolving punch made from stainless steel,Side spring that automatically adjust the tubes in position for punching. Total length approximately 22cm, ergonomic, soft grip handle. Must have 6 tube sizes ( 2.4mm, 2.8mm, 3.2mm, 4mm, 4,4mm, 4,8mm). Accessories:Revolving punch tubes.Description: Replacement tubes must be compatible with the above revolving hole punch. Technical Specifications:Dimensions: 2.4mm, 2.8mm, 3.2mm, 4.0mm, 4.4mm, 4.8mm. All accessories and consumables must be supplied standard with the unit. A separate price of each accessory and consumables listed above must also be supplied with this offer. | Duromed cc | North Coast Medical | USA | NCM | NC12345 | N/A | 2 YEARS | OFFER 1 | EACH | 3933 | 0 | 0 | 0 | 0 | 0 |  | 3944 | RT55-02-046 | 3944 | 3944 | 90 | 6.9 | 96.9 | 1 | Recommended
RT55-02-046 | Occupational Therapy | REVOLVING PUNCH Description: Handheld tool used to punch holes in various materials such as leather, thermoplastic & vinyl Technical Specifications:High quality heavy duty durable revolving punch made from stainless steel,Side spring that automatically adjust the tubes in position for punching. Total length approximately 22cm, ergonomic, soft grip handle. Must have 6 tube sizes ( 2.4mm, 2.8mm, 3.2mm, 4mm, 4,4mm, 4,8mm). Accessories:Revolving punch tubes.Description: Replacement tubes must be compatible with the above revolving hole punch. Technical Specifications:Dimensions: 2.4mm, 2.8mm, 3.2mm, 4.0mm, 4.4mm, 4.8mm. All accessories and consumables must be supplied standard with the unit. A separate price of each accessory and consumables listed above must also be supplied with this offer. | Sanbonani Holdings (Pty) Ltd | North Coast Medical Inc. | USA | North Coast Medical Inc. | NC12345 | 2008 | 10 years | 1 | Each | 3953.84 | [6 empty] | 4047.84 | RT55-02-046 | 4047.84 | 3944 | 87.6304 | 0 | 87.6304 | 2 | Recommended
RT55-02-050 | Occupational Therapy | SENSORY EVALUATOR HAND KIT Description:A set of instruments used to sensation and other components of hand function Technical Specification:The evaluation kit must include 5 basic monofilaments for objective testing of the cutaneious sensory perception threshold. Each monofilament must represent a unique amount of force and be marked as such. It must be easy to fold nylon filaments for protection, easy storage and portability .Includes a 5 piece hand set. 1x 2.83 (0.07g) - green, 1 x 3.61(0.4g)- blue, 1 x 4.31(2.0g)-purple, 1x4,56(4.0g)-red, 1x6.65(300g)-red Accessories:Must be accompanied with a soft-shell protective carrying case 1x user instructions.All accessories must be supplied standard with the unit | HITECH THERAPY | FABRICATION ENTERPRISES | USA | BASELINE | 121740 | 2018 | 7 YEARS | 1 | GRAM | 1840 | 0 | 0 | 0 | 0 | 0 |  | 1964 | RT55-02-050 | 1964 | 1964 | 90 | 0 | 90 | 1 | Recommended
RT55-02-050 | Occupational Therapy | SENSORY EVALUATOR HAND KIT Description:A set of instruments used to sensation and other components of hand function Technical Specification:The evaluation kit must include 5 basic monofilaments for objective testing of the cutaneious sensory perception threshold. Each monofilament must represent a unique amount of force and be marked as such. It must be easy to fold nylon filaments for protection, easy storage and portability .Includes a 5 piece hand set. 1x 2.83 (0.07g) - green, 1 x 3.61(0.4g)- blue, 1 x 4.31(2.0g)-purple, 1x4,56(4.0g)-red, 1x6.65(300g)-red Accessories:Must be accompanied with a soft-shell protective carrying case 1x user instructions.All accessories must be supplied standard with the unit | Duromed cc | North Coast Medical | USA | TOUCH TEST | NC12772 | N/A | 2 YEARS | OFFER 1 | EACH | 5359 | 0 | 0 | 0 | 0 | 0 |  | 5371 | RT55-02-050 | 5371 | 1964 | -66.1253 | 6.9 | -59.2253 | 2 | Recommended
RT55-02-050 | Occupational Therapy | SENSORY EVALUATOR HAND KIT Description:A set of instruments used to sensation and other components of hand function Technical Specification:The evaluation kit must include 5 basic monofilaments for objective testing of the cutaneious sensory perception threshold. Each monofilament must represent a unique amount of force and be marked as such. It must be easy to fold nylon filaments for protection, easy storage and portability .Includes a 5 piece hand set. 1x 2.83 (0.07g) - green, 1 x 3.61(0.4g)- blue, 1 x 4.31(2.0g)-purple, 1x4,56(4.0g)-red, 1x6.65(300g)-red Accessories:Must be accompanied with a soft-shell protective carrying case 1x user instructions.All accessories must be supplied standard with the unit | Mediotronics Physical Medicine Pty Ltd | MVS in Motion | China | Baseline | MVS12-0119 | 2015 | 5 Years | 1 | 1 / each including 7 components | 21980 | 0 | 0 | 0 | 0 | 0 |  | 22022 | RT55-02-050 | 22022 | 1964 | -829.1548 | 0 | -829.1548 | 3 | Recommended
RT55-02-053 | Occupational Therapy | SHIFTING SPANNER (0-30) Description:Also known as an adjustable spanner, its jaws can be expanded or narrowed down to tighten or loose the nut bolts. They are used for those nuts where no other spanner fits well.Technical Specification:These are generally made of drop forge steel or carbon steel. Should allow for 30mm adjustability.Standard: To DIN and ISO | KUSASALETU INVESTMENT GROUP | BUILT IT | RSA | BUILT IT | 10011310 | 2024 | LIFETIME | 1 | 1 | 145 | 0 | 0 | 0 | 0 | 0 |  | 169 | RT55-02-053 | 169 | 169 | 90 | 10 | 100 | 1 | Recommended
RT55-02-054 | Occupational Therapy | SLIDING WEDGE WITH VELCRO (Set of 3) Description: Therapeutic tool used for functional physical rehabilitation Technical Specification: Wedge with Velcro straps to provide positioning in supine or prone therapy activities. High density wedges: in various sizes, multiple colours. Easy to clean, quality leather or flame-retardant vinyl. Hard firm wedges padded and covered with smooth vinyl, PVC or similar , 3 dimensional blocks: Velcro attachment: Must provide a minimum of 2 Velcro (8-10cm broad ) straps to hold children in place for alignment. a) 15cm x 60cmx65cm b) 20cmx60cmx65cm c) 25cmx60cmx65cm | Sanbonani Holdings (Pty) Ltd | Geppetto's World of Developments cc t/a Playon | South Africa | Playon | GP-WES65V | 2016 | 8 years | 1 | Set of 3 | 3600 | [6 empty] | 3695 | RT55-02-054 | 3695 | 3695 | 90 | 10 | 100 | 1 | Recommended
RT55-02-056 | Occupational Therapy | SPANDEX TUNNEL Description: Therapeutic equipment to facilitate gross motor development Technical Specification:200 (length) x 40cm (diameter) Spandex tunnel manufactured from quality spandex material secured to a solid loop on one end. | Sanbonani Holdings (Pty) Ltd | Geppetto's World of Developments cc t/a Playon | South Africa | Playon | GP-TUSPA2 | 2018 | 6 years | 1 | Each | 1190 | [6 empty] | 1286 | RT55-02-056 | 1286 | 1286 | 90 | 10 | 100 | 1 | Recommended
RT55-02-057 | Occupational Therapy | SPLINTING SCISSOR, CURVED, SMALL Description: Tool used for cutting thermoplastic splinting material Technical Specification:Splinting scissors( small) with curved blades to cut easily from right or left. Heavy duty chrome plated. Able to easily cut through warmed splinting material. Dimensions: Overall length approximate 14cm, length of blade: approximate 5-8cm | Duromed cc | North Coast Medical | USA | NCM | NC12727 | N/A | 2 YEARS | OFFER 1 | EACH | 230 | 0 | 0 | 0 | 0 | 0 |  | 243 | RT55-02-057 | 243 | 243 | 90 | 6.9 | 96.9 | 1 | Recommended
RT55-02-057 | Occupational Therapy | SPLINTING SCISSOR, CURVED, SMALL Description: Tool used for cutting thermoplastic splinting material Technical Specification:Splinting scissors( small) with curved blades to cut easily from right or left. Heavy duty chrome plated. Able to easily cut through warmed splinting material. Dimensions: Overall length approximate 14cm, length of blade: approximate 5-8cm | Sanbonani Holdings (Pty) Ltd | North Coast Medical Inc. | USA | North Coast Medical Inc. | NC12727 | 2010 | 8 years | 1 | Each | 360 | [6 empty] | 457 | RT55-02-057 | 457 | 243 | 10.7407 | 10 | 20.7407 | 2 | Recommended
RT55-02-057 | Occupational Therapy | SPLINTING SCISSOR, CURVED, SMALL Description: Tool used for cutting thermoplastic splinting material Technical Specification:Splinting scissors( small) with curved blades to cut easily from right or left. Heavy duty chrome plated. Able to easily cut through warmed splinting material. Dimensions: Overall length approximate 14cm, length of blade: approximate 5-8cm | HITECH THERAPY | PERFORMANCE HEALTH | UK | SKLAR | 091046531 | 2019 | 3 YEARS | 1 | N/A | 2900 | 0 | 0 | 0 | 0 | 0 |  | 3025 | RT55-02-057 | 3025 | 243 | -940.3704 | 10 | -930.3704 | 3 | Recommended
RT55-02-058 | Occupational Therapy | SPLINTING SCISSOR, SPRINGLOADED (LEFT-HANDED ) Description:Tool used for cutting thermoplastic splinting material. Technical Specification:The tool must be designed for left-handed individuals.Heavy duty, high quality ultra-light weight scissors (spring loaded action) with stainless steel blades and handles (handles insulated with durable heavy-duty hard plastic). Capable of cutting 3.2 mm splinting material evenly and cleanly when heated and 1.6 mm splinting material when cold, leaves a smooth edge.20cm long, enlarged bottom handle to fit 4 fingers and locking mechanism when not in use.Ergonomic friendly handle.Left-handed handle high quality ultra-light weight scissors (spring action) with stainless steel blades and handles (handles insulated with durable heavy duty hard plastic). | Duromed cc | North Coast Medical | USA | FISKARS SOFTOUCH | NC12565 | N/A | 2 YEARS | OFFER 1 | EACH | 782 | 0 | 0 | 0 | 0 | 0 |  | 796 | RT55-02-058 | 796 | 796 | 90 | 6.9 | 96.9 | 1 | Recommended
RT55-02-058 | Occupational Therapy | SPLINTING SCISSOR, SPRINGLOADED (LEFT-HANDED ) Description:Tool used for cutting thermoplastic splinting material. Technical Specification:The tool must be designed for left-handed individuals.Heavy duty, high quality ultra-light weight scissors (spring loaded action) with stainless steel blades and handles (handles insulated with durable heavy-duty hard plastic). Capable of cutting 3.2 mm splinting material evenly and cleanly when heated and 1.6 mm splinting material when cold, leaves a smooth edge.20cm long, enlarged bottom handle to fit 4 fingers and locking mechanism when not in use.Ergonomic friendly handle.Left-handed handle high quality ultra-light weight scissors (spring action) with stainless steel blades and handles (handles insulated with durable heavy duty hard plastic). | Sanbonani Holdings (Pty) Ltd | North Coast Medical Inc. | USA | North Coast Medical Inc. | NC12565-1 | 2010 | 8 years | 2 | Each | 965 | [6 empty] | 1063 | RT55-02-058 | 1063 | 796 | 59.8116 | 10 | 69.8116 | 2 | Recommended
RT55-02-058 | Occupational Therapy | SPLINTING SCISSOR, SPRINGLOADED (LEFT-HANDED ) Description:Tool used for cutting thermoplastic splinting material. Technical Specification:The tool must be designed for left-handed individuals.Heavy duty, high quality ultra-light weight scissors (spring loaded action) with stainless steel blades and handles (handles insulated with durable heavy-duty hard plastic). Capable of cutting 3.2 mm splinting material evenly and cleanly when heated and 1.6 mm splinting material when cold, leaves a smooth edge.20cm long, enlarged bottom handle to fit 4 fingers and locking mechanism when not in use.Ergonomic friendly handle.Left-handed handle high quality ultra-light weight scissors (spring action) with stainless steel blades and handles (handles insulated with durable heavy duty hard plastic). | HITECH THERAPY | PERFORMANCE HEALTH | UK | FISKARS | 091022896 | 2020 | 3 YEARS | 1 | N/A | 2900 | 0 | 0 | 0 | 0 | 0 |  | 3026 | RT55-02-058 | 3026 | 796 | -162.1357 | 10 | -152.1357 | 3 | Recommended
RT55-02-059 | Occupational Therapy | SPLINTING SCISSOR, SPRINGLOADED (RIGHT-HANDED) Description:Tool used for cutting thermoplastic splinting material. Technical Specification:The tool must be designed for right-handed individuals Heavy duty, high quality ultra-light weight scissors (spring loaded action) with stainless steel blades and handles (handles insulated with durable heavy-duty hard plastic).Capable of cutting 3.2 mm splinting material evenly and cleanly when heated and 1.6 mm splinting material when cold, leaves a smooth edge.20cm long, enlarged bottom handle to fit 4 fingers and locking mechanism when not in use. Ergonomic friendly handle. Right-handed handle high quality ultra-light weight scissors (spring action) with stainless steel blades and handles (handles insulated with durable heavy-duty hard plastic) | Duromed cc | North Coast Medical | USA | FISKARS SOFTOUCH | NC12565 | N/A | 2 YEARS | OFFER 1 | EACH | 782 | 0 | 0 | 0 | 0 | 0 |  | 797 | RT55-02-059 | 797 | 797 | 90 | 6.9 | 96.9 | 1 | Recommended
RT55-02-059 | Occupational Therapy | SPLINTING SCISSOR, SPRINGLOADED (RIGHT-HANDED) Description:Tool used for cutting thermoplastic splinting material. Technical Specification:The tool must be designed for right-handed individuals Heavy duty, high quality ultra-light weight scissors (spring loaded action) with stainless steel blades and handles (handles insulated with durable heavy-duty hard plastic).Capable of cutting 3.2 mm splinting material evenly and cleanly when heated and 1.6 mm splinting material when cold, leaves a smooth edge.20cm long, enlarged bottom handle to fit 4 fingers and locking mechanism when not in use. Ergonomic friendly handle. Right-handed handle high quality ultra-light weight scissors (spring action) with stainless steel blades and handles (handles insulated with durable heavy-duty hard plastic) | Sanbonani Holdings (Pty) Ltd | North Coast Medical Inc. | USA | North Coast Medical Inc. | NC12565-1 | 2010 | 8 years | 2 | Each | 965 | [6 empty] | 1064 | RT55-02-059 | 1064 | 797 | 59.8494 | 10 | 69.8494 | 2 | Recommended
RT55-02-059 | Occupational Therapy | SPLINTING SCISSOR, SPRINGLOADED (RIGHT-HANDED) Description:Tool used for cutting thermoplastic splinting material. Technical Specification:The tool must be designed for right-handed individuals Heavy duty, high quality ultra-light weight scissors (spring loaded action) with stainless steel blades and handles (handles insulated with durable heavy-duty hard plastic).Capable of cutting 3.2 mm splinting material evenly and cleanly when heated and 1.6 mm splinting material when cold, leaves a smooth edge.20cm long, enlarged bottom handle to fit 4 fingers and locking mechanism when not in use. Ergonomic friendly handle. Right-handed handle high quality ultra-light weight scissors (spring action) with stainless steel blades and handles (handles insulated with durable heavy-duty hard plastic) | HITECH THERAPY | PERFORMANCE HEALTH | UK | FISKARS | 091022896 | 2020 | 3 YEARS | 1 | N/A | 2900 | 0 | 0 | 0 | 0 | 0 |  | 3027 | RT55-02-059 | 3027 | 797 | -161.8193 | 10 | -151.8193 | 3 | Recommended
RT55-02-060 | Occupational Therapy | SPLINTING SCISSORS, CURVED, LARGE Description: Tool used for cutting thermoplastic splinting material Technical Specification: Splinting scissors (large) with curved blades to cut easily from right or left. Approximately 10 degrees to cut easily from right or left. Heavy duty chrome plated. Able to easily cut through warmed splinting material. Dimensions: Overall approx. 20cm, length of blade: approximately 8 - 10cm | Duromed cc | North Coast Medical | USA | NCM | NC12719 | N/A | 2 YEARS | OFFER 1 | EACH | 1656 | 0 | 0 | 0 | 0 | 0 |  | 1672 | RT55-02-060 | 1672 | 1672 | 90 | 6.9 | 96.9 | 1 | Recommended
RT55-02-060 | Occupational Therapy | SPLINTING SCISSORS, CURVED, LARGE Description: Tool used for cutting thermoplastic splinting material Technical Specification: Splinting scissors (large) with curved blades to cut easily from right or left. Approximately 10 degrees to cut easily from right or left. Heavy duty chrome plated. Able to easily cut through warmed splinting material. Dimensions: Overall approx. 20cm, length of blade: approximately 8 - 10cm | Sanbonani Holdings (Pty) Ltd | North Coast Medical Inc. | USA | North Coast Medical Inc. | NC12719 | 2010 | 8 years | 2 | Each | 1774 | [6 empty] | 1874 | RT55-02-060 | 1874 | 1672 | 79.1268 | 8.66 | 87.7868 | 2 | Recommended
RT55-02-061 | Occupational Therapy | SPLINTING STATION Description: Mobile and versatile splinting station for the storage and transporting of splinting materials and tools. Technical Specification:Cabinet suitable for the storage of splinting material and consumables with top on which to place and use a splinting pan. Grab-bar handle for mobility on left side or right side for ease of transporting.(3) vertical cubby sections for splinting material behind hinged door with lock on right. Minimum of 1 draw for storage of consumables with 2 horizontal shelves below the draw on the left. Manufactured from water resistant material.Manufactured from melamine or similar material with durable dowel and glue construction. 100mm twin-wheel locking casters. Approximate measurement: 107-130cm x 61-64cm x 79cm-110cm | Sanbonani Holdings (Pty) Ltd | North Coast Medical Inc. | USA | North Coast Medical Inc. | NCUEWS-1 | 2008 | 12 years | 1 | Each | 18500 | [6 empty] | 18601 | RT55-02-061 | 18601 | 18601 | 90 | 8.66 | 98.66 | 1 | Recommended
RT55-02-061 | Occupational Therapy | SPLINTING STATION Description: Mobile and versatile splinting station for the storage and transporting of splinting materials and tools. Technical Specification:Cabinet suitable for the storage of splinting material and consumables with top on which to place and use a splinting pan. Grab-bar handle for mobility on left side or right side for ease of transporting.(3) vertical cubby sections for splinting material behind hinged door with lock on right. Minimum of 1 draw for storage of consumables with 2 horizontal shelves below the draw on the left. Manufactured from water resistant material.Manufactured from melamine or similar material with durable dowel and glue construction. 100mm twin-wheel locking casters. Approximate measurement: 107-130cm x 61-64cm x 79cm-110cm | HITECH THERAPY | GJ BARNARD | SOUTH AFRICA | SYNERGY | 0145 | 2019 | 7 YEARS | 1 | N/A | 19550 | 0 | 0 | 0 | 0 | 0 |  | 19678 | RT55-02-061 | 19678 | 18601 | 84.789 | 8.66 | 93.449 | 2 | Recommended
RT55-02-062 | Occupational Therapy | SPLINTING TROLLEY, STAINLESS STEEL Description: Equipment to allow for mobile splinting Technical Specification :Durable, stainless steel trolley, with 2 shelves and 1cm shallow lip edge, 4 lockable castors.Approximate dimensions: width 550mm width by 670mm length | New Horizon Metals cc | New Horizon Metals cc | South Africa | NHM | STSS001 | 2022/2023 | 7 years | 1 | Each | 5922.5 | 0 | 0 | 0 | 0 | 0 |  | 5994.5 | RT55-02-062 | 5994.5 | 5994.5 | 90 | 0 | 90 | 1 | Recommended
RT55-02-063 | Occupational Therapy | SUSPENSION FRAME, FREE STANDING & PORTABLE Description: Must accommodate most therapeutic swings or equipment Technical Specification:Must include incorporates a 360-degree vertical rotation system, allowing smooth, unrestricted range of motion for both vertical and rotational stimulation simultaneously. Made of durable, sturdy metal.Minimum weight capacity of 100 kilograms, clearly indicated in frame. English instruction manual to be included. Accessories: DAISY CHAINS: Accessory required for the suspension and height adjustment of swings. Durable nylon material. Multiple loops to allow for height adjustment of swings.Approximate length: 150cm.Approximate width 20mm. CARABINER CLIPS: Accessory required for the suspension of swings. A specialized type of shackle, a metal loop with a spring-loaded gate used to quickly and reversibly connect components, most notably in safety-critical systems. Made of durable metal, easy to use clips. Minimum weight capacity of 150kgs. Available in single units. | HITECH THERAPY | MEDEN -INMED | POLAND | ELEVEO | 01-ELV-0101-003 | 2019 | 7 YEARS | 1 | KG | 250000 | 2500 | 2500 | 2500 | 2500 | 2500 |  | 260129 | RT55-02-063 | 260129 | 260129 | 90 | 8.66 | 98.66 | #REF! | Recommended
RT55-02-064 | Occupational Therapy | TAPE MEASURE Description:A set of Instrument used to test grip strength and other components of hand function. Technical Specification:Small, compact measuring tape used to measure the circumference of the digits and hand . Measures in inches and centimetres. Must be flexible to wrap around the digit/ hand Must be retractable.Measures up to a minimum of 1 metre. | Duromed cc | North Coast Medical | USA | NCM EXACTA | NC12708 | N/A | 2 YEARS | OFFER 1 | EACH | 109.25 | 0 | 0 | 0 | 0 | 0 |  | 126.25 | RT55-02-064 | 126.25 | 126.25 | 90 | 6.9 | 96.9 | 1 | Recommended
RT55-02-064 | Occupational Therapy | TAPE MEASURE Description:A set of Instrument used to test grip strength and other components of hand function. Technical Specification:Small, compact measuring tape used to measure the circumference of the digits and hand . Measures in inches and centimetres. Must be flexible to wrap around the digit/ hand Must be retractable.Measures up to a minimum of 1 metre. | HITECH THERAPY | FABRICATION ENTERPRISES | USA | BASELINE | 121210 | 2003 | 7 YEARS | 1 | CM & INCHES | 115 | 0 | 0 | 0 | 0 | 0 |  | 245 | RT55-02-064 | 245 | 126.25 | 5.3465 | 8.66 | 14.0065 | 2 | Recommended
RT55-02-064 | Occupational Therapy | TAPE MEASURE Description:A set of Instrument used to test grip strength and other components of hand function. Technical Specification:Small, compact measuring tape used to measure the circumference of the digits and hand . Measures in inches and centimetres. Must be flexible to wrap around the digit/ hand Must be retractable.Measures up to a minimum of 1 metre. | Sanbonani Holdings (Pty) Ltd | North Coast Medical Inc. | USA | Exacta™ | NC12708 | 2012 | 6 years | 3 | Each | 90 | [6 empty] | 192 | RT55-02-064 | 192 | 192 | 90 | 8.66 | 98.66 | 1 | Recommended
RT55-02-065 | Occupational Therapy | THERAPEUTIC SWINGS (Set of 8) Description: Therapeutic tools used for functional paediatric physical and or sensory rehabilitation Technical Specification:All swings can either be ceiling mounted (installed) or provided with a suspension frame compatible for the swings below using the designated amount of daisy chains and carabiner clips/hooks which should be supplied with each swing unit. Should be used with a soft-landing space on the floor by means of a suitable mat (munch mat). Therapeutic Swings Sizes: a) Bolster swing: A cylindrical swing suspended from 2 points .150cm (Length) x 30cm (Diameter),Solid bolster/cylinder padded with high density foam, positioned horizontally with suspension ropes on either side. Upholstered in waterproof durable vinyl or similar material.Minimum weight capacity of 100 kg. b) Platform swing: A flat platform swing suspended from 2 points.60 x 90 x 25cm solid lightly padded platform, Vinyl covered.Suspension ropes attached at each corner, 2 ropes at each long end of the platform joined by a solid bar suitable for a child to stabilize him/herself on.2 points for suspension.Minimum weight capacity of 100 kg. c) Flexion swing: A swing suspended from one point.Round 60-70cm solid base with attached central upright solid cylinder.Padded with high density foam. Upholstered with waterproof vinyl or similar material.Height adjustable attachment.Minimum weight capacity of 80 kg. d) Flying saucer swing: A flat platform swing with removable padded safety barrier suspended from. 2 points.90cm diameter, removable side barrier 20cm in height, hard firm platform with light high-density foam padding, with removable side barrier, side barrier made of hard firm inner with soft high-density foam padding. Covered with smooth cleanable vinyl.Minimum weight capacity of 100 kg. e) Barrel Swing: A barrel, open on both ends, to be used in hammock swing. Available in 2 sizes: Barrel 70 - 70x65 diam cm Barrel 90 - 90x65 diam cm. f) Hammock, spandex: A swing suspended from 4 points.3 layers of quality spandex.4-point attachment.Approximately 220cm x 150cm.Minimum weight capacity of 150 kg. g) Hammock, denim: A swing suspended from 2 points .Manufactured from durable denim material.Suspended from two separate points to create an open hammock or one central point for a cocoon like effect. Must be supplied with two daisy chains and 2 carabiner hooks for suspension. Approximate measurements:Available in 2 sizes: Small - 200x150 cm and Large - 300x150 cm.Minimum weight capacity of 150 kg. h) Trapeze: Equipment to promote gross motor development Wooden or durable hard plastic 60cm horizontal trapeze bar.Secured on each side by a rope to a single attachment point for attachment to suspension unit.Splinter free surface. | Sanbonani Holdings (Pty) Ltd | Geppetto's World of Developments cc t/a Playon | South Africa | Playon | GP-SWS8 | 2018 | 6 years | 1 | Set | 18465 | [6 empty] | 18568 | RT55-02-065 | 18568 | 18568 | 90 | 8.66 | 98.66 | 1 | Recommended
RT55-02-065 | Occupational Therapy | THERAPEUTIC SWINGS (Set of 8) Description: Therapeutic tools used for functional paediatric physical and or sensory rehabilitation Technical Specification:All swings can either be ceiling mounted (installed) or provided with a suspension frame compatible for the swings below using the designated amount of daisy chains and carabiner clips/hooks which should be supplied with each swing unit. Should be used with a soft-landing space on the floor by means of a suitable mat (munch mat). Therapeutic Swings Sizes: a) Bolster swing: A cylindrical swing suspended from 2 points .150cm (Length) x 30cm (Diameter),Solid bolster/cylinder padded with high density foam, positioned horizontally with suspension ropes on either side. Upholstered in waterproof durable vinyl or similar material.Minimum weight capacity of 100 kg. b) Platform swing: A flat platform swing suspended from 2 points.60 x 90 x 25cm solid lightly padded platform, Vinyl covered.Suspension ropes attached at each corner, 2 ropes at each long end of the platform joined by a solid bar suitable for a child to stabilize him/herself on.2 points for suspension.Minimum weight capacity of 100 kg. c) Flexion swing: A swing suspended from one point.Round 60-70cm solid base with attached central upright solid cylinder.Padded with high density foam. Upholstered with waterproof vinyl or similar material.Height adjustable attachment.Minimum weight capacity of 80 kg. d) Flying saucer swing: A flat platform swing with removable padded safety barrier suspended from. 2 points.90cm diameter, removable side barrier 20cm in height, hard firm platform with light high-density foam padding, with removable side barrier, side barrier made of hard firm inner with soft high-density foam padding. Covered with smooth cleanable vinyl.Minimum weight capacity of 100 kg. e) Barrel Swing: A barrel, open on both ends, to be used in hammock swing. Available in 2 sizes: Barrel 70 - 70x65 diam cm Barrel 90 - 90x65 diam cm. f) Hammock, spandex: A swing suspended from 4 points.3 layers of quality spandex.4-point attachment.Approximately 220cm x 150cm.Minimum weight capacity of 150 kg. g) Hammock, denim: A swing suspended from 2 points .Manufactured from durable denim material.Suspended from two separate points to create an open hammock or one central point for a cocoon like effect. Must be supplied with two daisy chains and 2 carabiner hooks for suspension. Approximate measurements:Available in 2 sizes: Small - 200x150 cm and Large - 300x150 cm.Minimum weight capacity of 150 kg. h) Trapeze: Equipment to promote gross motor development Wooden or durable hard plastic 60cm horizontal trapeze bar.Secured on each side by a rope to a single attachment point for attachment to suspension unit.Splinter free surface. | HITECH THERAPY | GEPPETTO | SOUTH AFRICA | PLAY ON | RT55SW | 2003 | 7 YEARS | 1 | CM | 32775 | 0 | 0 | 0 | 0 | 0 |  | 32906 | RT55-02-065 | 32906 | 18568 | 20.503 | 8.66 | 29.163 | 2 | Recommended
RT55-02-066 | Occupational Therapy | Therapy Blocks, set of 6 Description: A set of therapeutic tools used for functional physical rehabilitation Technical Specification:High density foam blocks: in various sizes (long & round), multiple colours. Easy to clean, quality leather or flame-retardant vinyl.Hard firm blocks padded and covered with smooth vinyl, PVC or similar ; 3 dimensional blocks; Approximate dimensions: a. 30cm x 45cm x 45cm b. 40cm x 40cm x 40cm c. 50cm x 50cm x10cm d. 55cm x 55cm x 15cm e. 60cm x 60cm x 20cm f. 65cm x 65cm x 25cm | HITECH THERAPY | GEPPETTO | SOUTH AFRICA | PLAY ON | RT55BSET | 2003 | 2 YEARS | 1 | CM | 8131 | 0 | 0 | 0 | 0 | 0 |  | 8263 | RT55-02-066 | 8263 | 8263 | 90 | 8.66 | 98.66 | 1 | Recommended
RT55-02-067 | Occupational Therapy | THERAPY FOAM ROLLERS (set of 5 X 90 cm length) Description:A set of therapeutic tools used for functional physical rehabilitation Technical Specification:High density foam rollers: in various sizes (long & round), multiple colours. Cleanable quality leather or flame-retardant vinyl. Rollers are made of Polystyrene, covered with chip foam and cover. Approximate dimensions: a.90cm length and 50cm diameter b· 90cm length and 40cm diameter. c·90cm length and 30cm diameter d·90cm length and 20cm diameter. d. 90cm length and 15cm diameter . | Sanbonani Holdings (Pty) Ltd | Geppetto's World of Developments cc t/a Playon | South Africa | Playon | GP-RO5905 | 2003 | 8 years | 1 | Set | 6500 | [6 empty] | 6604 | RT55-02-067 | 6604 | 6604 | 90 | 8.66 | 98.66 | 1 | Recommended
RT55-02-067 | Occupational Therapy | THERAPY FOAM ROLLERS (set of 5 X 90 cm length) Description:A set of therapeutic tools used for functional physical rehabilitation Technical Specification:High density foam rollers: in various sizes (long & round), multiple colours. Cleanable quality leather or flame-retardant vinyl. Rollers are made of Polystyrene, covered with chip foam and cover. Approximate dimensions: a.90cm length and 50cm diameter b· 90cm length and 40cm diameter. c·90cm length and 30cm diameter d·90cm length and 20cm diameter. d. 90cm length and 15cm diameter . | HITECH THERAPY | GEPPETTO | SOUTH AFRICA | PLAY ON | HTT3015 | 2003 | 2 YEARS | 1 | CM | 8510 | 0 | 0 | 0 | 0 | 0 |  | 8643 | RT55-02-067 | 8643 | 6604 | 62.2123 | 8.66 | 70.8723 | 2 | Recommended
RT55-02-068 | Occupational Therapy | THERAPY WEDGES (set of 3).Item Description:Therapeutic tool used for functional physical rehabilitation Technical Specification:Therapy wedges: in various sizes (triangular), multiple colours. Wedges are made of Polystyrene, covered with foam and cover.Upholstered surface made with imitation leather upholstery or vinyl, contains no phthalates and is flame retardant. Easy to clean, durable and light weight. At least one angle must be 90 degree Fluid resistant (medical grade water resistant). To include the following approximate sizes: a) Small (45cmx45cm, height 8-10cm). b) Medium (45cm x45cm, height 15cm). c) Large (45x45cm, height 20cm) . | HITECH THERAPY | GEPPETTO | SOUTH AFRICA | PLAY ON | WEDRT55 | 2003 | 2 YEARS | 1 | CM | 3289 | 0 | 0 | 0 | 0 | 0 |  | 3423 | RT55-02-068 | 3423 | 3423 | 90 | 8.66 | 98.66 | 1 | Recommended
RT55-02-073 | Occupational Therapy | TWO POINT DISCRIMINATOR Description:Instrument used to test grip strength and other components of hand function Technical Specification:Consists of two sturdy, rotating, plastic disks that are joined together. Rounded tips are spaced at standard testing intervals from 1 to 15mm apart A 20 and 25mm spacing are also given. One disk setting tests from 1 to 8mm. To change to the 9 to 15mm disk setting, simply rotate the top disk until it clicks in place. Accessories:Instructions included.All accessories must be supplied standard with the unit. | Duromed cc | North Coast Medical | USA | TOUCH TEST | NC12776 | N/A | 2 YEARS | OFFER 1 | EACH | 954.5 | 0 | 0 | 0 | 0 | 0 |  | 972.5 | RT55-02-073 | 972.5 | 972.5 | 90 | 6.9 | 96.9 | 1 | Recommended
RT55-02-073 | Occupational Therapy | TWO POINT DISCRIMINATOR Description:Instrument used to test grip strength and other components of hand function Technical Specification:Consists of two sturdy, rotating, plastic disks that are joined together. Rounded tips are spaced at standard testing intervals from 1 to 15mm apart A 20 and 25mm spacing are also given. One disk setting tests from 1 to 8mm. To change to the 9 to 15mm disk setting, simply rotate the top disk until it clicks in place. Accessories:Instructions included.All accessories must be supplied standard with the unit. | Sanbonani Holdings (Pty) Ltd | North Coast Medical Inc. | USA | Touch Test® | NC12776 | 2012 | 8 years | 2 | Each | 955 | [6 empty] | 1060 | RT55-02-073 | 1060 | 972.5 | 81.9023 | 8.66 | 90.5623 | 2 | Recommended
RT55-02-073 | Occupational Therapy | TWO POINT DISCRIMINATOR Description:Instrument used to test grip strength and other components of hand function Technical Specification:Consists of two sturdy, rotating, plastic disks that are joined together. Rounded tips are spaced at standard testing intervals from 1 to 15mm apart A 20 and 25mm spacing are also given. One disk setting tests from 1 to 8mm. To change to the 9 to 15mm disk setting, simply rotate the top disk until it clicks in place. Accessories:Instructions included.All accessories must be supplied standard with the unit. | HITECH THERAPY | FABRICATION ENTERPRISES | USA | BASELINE | 121496 | 2016 | 7 YEARS | 1 | N/A | 1050 | 0 | 0 | 0 | 0 | 0 |  | 1185 | RT55-02-073 | 1185 | 972.5 | 70.3342 | 8.66 | 78.9942 | 3 | Recommended
RT55-02-073 | Occupational Therapy | TWO POINT DISCRIMINATOR Description:Instrument used to test grip strength and other components of hand function Technical Specification:Consists of two sturdy, rotating, plastic disks that are joined together. Rounded tips are spaced at standard testing intervals from 1 to 15mm apart A 20 and 25mm spacing are also given. One disk setting tests from 1 to 8mm. To change to the 9 to 15mm disk setting, simply rotate the top disk until it clicks in place. Accessories:Instructions included.All accessories must be supplied standard with the unit. | Mediotronics Physical Medicine Pty Ltd | MVS in Motion | Belgium | Baseline | MVS12-1496 | 2015 | 5 Years | 1 | 1 / each | 1192.75 | 0 | 0 | 0 | 0 | 0 |  | 1235.75 | RT55-02-073 | 1235.75 | 972.5 | 65.6375 | 0 | 65.6375 | 4 | Recommended
RT55-02-074 | Occupational Therapy | UPPER EXTREMITY WORKSTATION Description:Multi-functional therapeutic unit allowing simultaneous treatment of upper limb function of multiple clients .Technical Specification:Upper extremity, four-sided mobile workstation on lockable castors. Workstation Features:Accommodate two wheelchairs comfortably, Height and angle adjustable side shelves .Slide-out rear shelf suitable for a seated client.Lockable double doors for security.Two pairs of height adjustable interior shelves.Sturdy construction with durable Formica work surfaces Smooth-rolling, locking castors for easy mobility.Approximate dimensions: Height 99cm, Length 96cm and Depth 61cm. Accessories: 10 therapeutic items: a). Manipulation board :Various types of locks, switches and latches mounted to a durable wooden board that has option for wall mounting and have desk top holder. b). Two-tiered horizontal bolt board: High-quality wood with various size bolts, from coarse to fine, with matching nuts and washers. Bolted on a 2-tiered durable, wooden base. c.) Shoulder abduction ladder :Wooden ‘laddered’ tool that can be wall mounted or lay flat on a table surface. With 33 steps, each measuring approximately 6cm W x 4cm H, users can climb the ladder with their fingers to improve flexibility and mobility. d). Bilateral sander: Approximately 25cm (L) x 22cm (W) 10cm (H).Includes 5 pairs of interchangeable handles that vary in diameter Constructed of durable wood that sits on smooth plastic gliders which will protect the working area. Perfect tool for improving upper extremity coordination and strength. Difficulty can be adjusted by using the interchangeable handles or by adding weight to the box. e). Stacking cones:Washable, non-toxic, and impact resistant acrylic cones with a smooth surface measure approximately 10cm high with approximately 6cm base diameter.Cones come in six translucent colors (blue, green, red, amber, smoke and clear) to be used for color-coordination exercises.Cones fit on a grooved durable, wooden board base. f). Graded pegboard:Features five rows of graded-height pegs, each size painted a different color. Helpful for advanced training of size and depth perception.Board, measuring approximately 25cm x 30cm, has 30 pegs; the tallest size is approximately 7cm.Durable wooden board base with cut out holes to fit pegs. g). Extend Your Reach Exerciser: Includes a standard wooden yard stick, a wood box base, and 24 clothespins.The resistance of the clothespins can be adjusted by adding or removing rubber bands. h) Finger extension remedial game: Two-sided board, allowing for two game options: checkers and a solitaire jumping game. The hook and loop fastener playing surfaces give desired resistance and engage the finger in position for active extension.Heavy pressed board with nonskid feet.Includes 32 square plastic checkers with hook and loop playing surfaces and checkers with a soft vinyl loop.Checkers are 2.5cm squares, Entire board is approximately 23 cm x23 cm. i)Graded Range of Motion Arch:Arc adjusts to approximately 38cm, 50cm, 64cm & 76 cm heights.Durable wooden base measures approximately 92cm, X 2cm X 15cm.Plastic tube is approximately 2cm in diameter and includes 12 rings j). E-Z exercise board:Made of a wooden board, approximately 22 cm x 33 cm Hook & loop fastener base, provides a stable platform for 7 tools that offer various resistance. | HITECH THERAPY | GJ BARNARD | SOUTH AFRICA | SYNERGY | 0273 | 2019 | 7 YEARS | 1 | N/A | 39675 | 0 | 0 | 0 | 0 | 0 |  | 39811 | RT55-02-074 | 39811 | 39811 | 90 | 8.66 | 98.66 | 1 | Recommended
RT55-02-074 | Occupational Therapy | UPPER EXTREMITY WORKSTATION Description:Multi-functional therapeutic unit allowing simultaneous treatment of upper limb function of multiple clients .Technical Specification:Upper extremity, four-sided mobile workstation on lockable castors. Workstation Features:Accommodate two wheelchairs comfortably, Height and angle adjustable side shelves .Slide-out rear shelf suitable for a seated client.Lockable double doors for security.Two pairs of height adjustable interior shelves.Sturdy construction with durable Formica work surfaces Smooth-rolling, locking castors for easy mobility.Approximate dimensions: Height 99cm, Length 96cm and Depth 61cm. Accessories: 10 therapeutic items: a). Manipulation board :Various types of locks, switches and latches mounted to a durable wooden board that has option for wall mounting and have desk top holder. b). Two-tiered horizontal bolt board: High-quality wood with various size bolts, from coarse to fine, with matching nuts and washers. Bolted on a 2-tiered durable, wooden base. c.) Shoulder abduction ladder :Wooden ‘laddered’ tool that can be wall mounted or lay flat on a table surface. With 33 steps, each measuring approximately 6cm W x 4cm H, users can climb the ladder with their fingers to improve flexibility and mobility. d). Bilateral sander: Approximately 25cm (L) x 22cm (W) 10cm (H).Includes 5 pairs of interchangeable handles that vary in diameter Constructed of durable wood that sits on smooth plastic gliders which will protect the working area. Perfect tool for improving upper extremity coordination and strength. Difficulty can be adjusted by using the interchangeable handles or by adding weight to the box. e). Stacking cones:Washable, non-toxic, and impact resistant acrylic cones with a smooth surface measure approximately 10cm high with approximately 6cm base diameter.Cones come in six translucent colors (blue, green, red, amber, smoke and clear) to be used for color-coordination exercises.Cones fit on a grooved durable, wooden board base. f). Graded pegboard:Features five rows of graded-height pegs, each size painted a different color. Helpful for advanced training of size and depth perception.Board, measuring approximately 25cm x 30cm, has 30 pegs; the tallest size is approximately 7cm.Durable wooden board base with cut out holes to fit pegs. g). Extend Your Reach Exerciser: Includes a standard wooden yard stick, a wood box base, and 24 clothespins.The resistance of the clothespins can be adjusted by adding or removing rubber bands. h) Finger extension remedial game: Two-sided board, allowing for two game options: checkers and a solitaire jumping game. The hook and loop fastener playing surfaces give desired resistance and engage the finger in position for active extension.Heavy pressed board with nonskid feet.Includes 32 square plastic checkers with hook and loop playing surfaces and checkers with a soft vinyl loop.Checkers are 2.5cm squares, Entire board is approximately 23 cm x23 cm. i)Graded Range of Motion Arch:Arc adjusts to approximately 38cm, 50cm, 64cm & 76 cm heights.Durable wooden base measures approximately 92cm, X 2cm X 15cm.Plastic tube is approximately 2cm in diameter and includes 12 rings j). E-Z exercise board:Made of a wooden board, approximately 22 cm x 33 cm Hook & loop fastener base, provides a stable platform for 7 tools that offer various resistance. | Sanbonani Holdings (Pty) Ltd | North Coast Medical Inc. | USA | North Coast Medical Inc. | NCUEWS10 | 2002 | 12 years | 1 | Set | 43930 | [6 empty] | 44036 | RT55-02-074 | 44036 | 39811 | 80.4486 | 8.66 | 89.1086 | 2 | Recommended
RT55-02-075 | Occupational Therapy | VISUAL IMAGERY SMART DESKTOP MIRROR Description:A therapy tool used for visual feedback during mirror therapy and functional rehabilitation, also referred to as a smart rehabilitation mirror. Technical Specification:Must have an Adjustable mirror between 90 and 65 degrees Easy set up for one handed use in clinic or home.Mirror cover must be sturdy, nonslip and covered in black nylon, or PVC which is easy to clean Visible shatterproof mirror area .Easily folds flat for safe and compact storage Range of mirror (20-25cm x30-35cm) | Sanbonani Holdings (Pty) Ltd | North Coast Medical Inc. | USA | Achieva | NC95110 | 2020 | 10 years | 2 | Each | 2980.8 | [6 empty] | 3087.8 | RT55-02-075 | 3087.8 | 3087.8 | 90 | 8.66 | 98.66 | 1 | Recommended
RT55-02-076 | Occupational Therapy | WEIGHTS, ANKLE/WRIST.Description: Therapeutic tools available in a pair/set to strap on ankles or wrists to facilitate stabilisation or muscle strengthening. Technical Specification:Easy to don/doff with Velcro and /or D-rings or similar fastenings which allow for adjustability to different limb sizes yet fit securely and snugly. Covered in soft, durable material that is comfortable and sweat or moisture resistant.Available in various weights (pair): a. 0.25kg b. 0.5kg c. 1 kg d. 2 kg | Sanbonani Holdings (Pty) Ltd | North Coast Medical Inc. | USA | Fizique | NC02004 | 2002 | 12 years | 1 | Set | 1557.5 | [6 empty] | 1665.5 | RT55-02-076 | 1665.5 | 1665.5 | 90 | 8.66 | 98.66 | 1 | Recommended
RT55-02-077 | Occupational Therapy | WEIGHTED BLANKET Description: A therapeutic weighted blanket that assists with concentration, organisation, body awareness and sensory processing and provides calming proprioceptive input. Technical Specification:Should be made out of a strong material such as denim. Size: Various, approximately 100-150x80-90 cm Available in two different weights: +-4 Kg and +-7 Kg | Sanbonani Holdings (Pty) Ltd | Geppetto's World of Developments cc t/a Playon | South Africa | Playon | GP-WBS2 | 2018 | 8 years | 1 | Set | 1590 | [6 empty] | 1699 | RT55-02-077 | 1699 | 1699 | 90 | 8.66 | 98.66 | 1 | Recommended
RT55-02-078 | Occupational Therapy | WEIGHTED LAP PAD Description: Weighted lap pad used for self-regulation. Technical Specifications:Should be made out of durable material such as denim.Size: approximately 60 cm x 30 cm.Weight: approximately 1.5 kg | Sanbonani Holdings (Pty) Ltd | Geppetto's World of Developments cc t/a Playon | South Africa | Playon | GP-WLP | 2018 | 8 years | 1 | Each | 750 | [6 empty] | 860 | RT55-02-078 | 860 | 860 | 90 | 8.66 | 98.66 | 1 | Recommended
RT55-03-006 | Orthotist and Prosthetist | 6 Piece Screwdriver Set with \Striking Cap: Hexagon blade design for strength with an additional Bolster Ergonomic three component Power Grip handle for an ergonomic comfortable grip and reduced fatigue over extended periods.Positive fit joint of handle and blade for maximum force transmission.Constructed from High Alloy Chrome Vanadium Molybdenum Steel.Screwdriver Type.Engineers,Number of Pieces6.Tip Size:5, 5.5, 6.5, 8mm, PH2, Phillips: PH1, Slotted: 4.5 | KUSASALETU INVESTMENT GROUP | BUILT IT | RSA | BUILT IT | 110100238 | 2024 | LIFETIME | 1 | 1 | 250 | 0 | 0 | 0 | 0 | 0 |  | 275 | RT55-03-006 | 275 | 275 | 90 | 10 | 100 | 1 | Recommended
RT55-04-002 | Physiotherapy | Active and Passive Motion (ASPM) Pedal Exerciser for Upper and lower body Description :A rehabilitation unit used for exercising upper and lower limbs Technical Specification :Portable unit that weighs between 13 - 15kg Dimensions (LxWxH): 50x46x43cm 100 - 240V, 50/60Hz power connection Electric motor drive: 240 Watt Control & display unit with a colour screen .Remote control shall be small and handy Must be a single unit with detachable components suitable for upper and lower limbs The unit shall have the following features: Spasm detection and release. Feedback on speed, duration, muscle tone, distance covered, right to left symmetry, resistance and performance.Active or passive motor assistance .Pre-programmed exercise programs and control programs customizable programs.Rotary operating pane with screen .Safety foot shells or stabilising straps with a comfortable non-slip surface . Spring mounted calf straps/shells Non-slip stand .Knee bending manual adjustment .Hydraulic or manual height adjustment Leg guides to secure guidance of leg movement .Forearm shells or safety straps with wrist or hand cuffs/shells with a quick release function . Emergency stop function .Extra high safety edging on foot shells to prevent slipping Height adjustable handlebars with non slip material cover .Control and display unit with a colour screen Foot rests and V Handles .The unit shall have the option of adding a rechargeable battery in line with end user requirements Accessories: Foot shells and straps Forearm shells and straps The unit shall operate from 220V and the power cords shall have a 3-point plug approved by SABS The unit shall be FDS or CE approved (Attach certification) All accessories and consumables must be supplied standard with the unit A separate price of each accessory and consumables listed above must also be supplied with this offer Any other additional accessories, consumables and spare parts not listed above but applicable for the functionality of the unit must be supplied on a separate price list | HITECH THERAPY | MEDICA MEDIZINTECHNIK trading as THERA TRAINER | GERMANY | THERA TRAINER | MOBI 540 | 2015 | 7 YEARS | 1 | RPM | 78200 | 1500 | 1500 | 1500 | 1500 | 1500 |  | 84337 | RT55-04-002 | 84337 | 84337 | 90 | 8.66 | 98.66 | 1 | Recommended
RT55-04-002 | Physiotherapy | Active and Passive Motion (ASPM) Pedal Exerciser for Upper and lower body Description :A rehabilitation unit used for exercising upper and lower limbs Technical Specification :Portable unit that weighs between 13 - 15kg Dimensions (LxWxH): 50x46x43cm 100 - 240V, 50/60Hz power connection Electric motor drive: 240 Watt Control & display unit with a colour screen .Remote control shall be small and handy Must be a single unit with detachable components suitable for upper and lower limbs The unit shall have the following features: Spasm detection and release. Feedback on speed, duration, muscle tone, distance covered, right to left symmetry, resistance and performance.Active or passive motor assistance .Pre-programmed exercise programs and control programs customizable programs.Rotary operating pane with screen .Safety foot shells or stabilising straps with a comfortable non-slip surface . Spring mounted calf straps/shells Non-slip stand .Knee bending manual adjustment .Hydraulic or manual height adjustment Leg guides to secure guidance of leg movement .Forearm shells or safety straps with wrist or hand cuffs/shells with a quick release function . Emergency stop function .Extra high safety edging on foot shells to prevent slipping Height adjustable handlebars with non slip material cover .Control and display unit with a colour screen Foot rests and V Handles .The unit shall have the option of adding a rechargeable battery in line with end user requirements Accessories: Foot shells and straps Forearm shells and straps The unit shall operate from 220V and the power cords shall have a 3-point plug approved by SABS The unit shall be FDS or CE approved (Attach certification) All accessories and consumables must be supplied standard with the unit A separate price of each accessory and consumables listed above must also be supplied with this offer Any other additional accessories, consumables and spare parts not listed above but applicable for the functionality of the unit must be supplied on a separate price list | Sanbonani Holdings (Pty) Ltd | Kinetec SAS | France | Kinetec Kinevia Duo™ | 4621009002 | 2014 | 12 years | 1 | Each | 174459.6 | 4800 | 5184 | 5598.72 | 6046.6176 | 6530.347 |  | 196199.9376 | RT55-04-002 | 196199.9376 | 84337 | -29.3742 | 8.66 | -20.7142 | 2 | Recommended
RT55-04-003 | Physiotherapy | Atlas Walker (chrome plated steel frame) Description:A mobility Rehabilitation device used to assist patients with gait Technical Specification :The unit shall be a height adjustable atlas walking frame The unit shall have a chrome plated steel frame The unit shall be easy to clean The unit shall have a padded rest pad that is shaped to fit the human body Padded rest pad for the upper limbs to rest on, ergonomically shaped, to disperse weight between upper body and legs Padded rest pad to be covered in durable vinyl, which is easily disinfected and flame retardant The unit shall have 4 heavy duty swivel castors, 2 in front and 2 in the rear for easy manoeuvrability Rear castors will have brakes Quality requirements: Product shall be free of burrs, sharp edges, oil, grease and/or any foreign matter prior to packaging The unit shall have 2 adjustable handgrips Specifications: Maximum user weight atleast: 160kg Height adjustment from approximately 98cm – 145cm Overall width approximately 63cm Overall depth approximately 75cm Dimension: height adjustable 93-142cm; overall width 63cm. Overall depth: 76cm Weight of unit 8-10kg Bidder shall supply document/certificate from supplier along with samples for testing, to certify type of aluminium alloy and strength used. Accessories Castors The unit shall be FDS or CE approved (Attach certification) All accessories and consumables must be supplied standard with the unit A separate price of each accessory and consumables listed above must also be supplied with this offer Any other additional accessories, consumables and spare parts not listed above but applicable for the functionality of the unit must be supplied on a separate price list | New Horizon Metals cc | New Horizon Metals cc | South Africa | NHM | AW001 | 2022/2023 | 7 years | 1 | Each | 4145.75 | 0 | 0 | 0 | 0 | 0 |  | 4219.75 | RT55-04-003 | 4219.75 | 4219.75 | 90 | 0 | 90 | 1 | Recommended
RT55-04-004 | Physiotherapy | Balance Foam Pad Description: A device used for standing balance training Technical Specification: The unit shall be manufactured from a dense closed cell foam that is antibacterial and water impermeable. Dimensions (LxWxD) shall be at least 50cm x 40cm x 5cm The unit shall be light weight ≤800grams The unit shall have a non-slip surface | Mediotronics Physical Medicine Pty Ltd | MVS in Motion | Belgium | Sirex | Balance Pad : CB-0240 | 2010 | 5 Years | 1 | Each | 684.25 | 0 | 0 | 0 | 0 | 0 |  | 730.25 | RT55-04-004 | 730.25 | 730.25 | 90 | 0 | 90 | 1 | Recommended
RT55-04-004 | Physiotherapy | Balance Foam Pad Description: A device used for standing balance training Technical Specification: The unit shall be manufactured from a dense closed cell foam that is antibacterial and water impermeable. Dimensions (LxWxD) shall be at least 50cm x 40cm x 5cm The unit shall be light weight ≤800grams The unit shall have a non-slip surface | Duromed cc | North Coast Medical | USA | NCM | NC64677 | N/A | 2 YEARS | OFFER 1 | EACH | 2012.5 | 0 | 0 | 0 | 0 | 0 |  | 2031.5 | RT55-04-004 | 2031.5 | 730.25 | -70.3732 | 6.9 | -63.4732 | 2 | Recommended
RT55-04-005 | Physiotherapy | Corner Steps/Stairs with Height Adjustable Handrails Description:Rehabilitation unit used for gait training on stairs Technical Specification :Stairs 90° manufactured with solid wood treated for pests, sealed and varnished The unit shall have a non-slip surface (preferably carpeted) Width of steps will be a minimum of 60 cm and a maximum of 90cm The unit shall be sturdy with and approximate weight of ≥50kg Platform dimensions should be at least 75cm x 90cm The unit shall have powder coated metal height adjustable grab rails on either side Metallic hand rails will be adjustable from 40cm - 90cm in height There will be one set of 4 steps with a height of ±15cm on the one side There will be one set of 3 steps with a height of ±20cm on the other side Minimum loading capacity of 250 kg | New Horizon Metals cc | New Horizon Metals cc | South Africa | NHM | CS001 | 2022/2023 | 7 years | 1 | Each | 15990.75 | 0 | 0 | 0 | 0 | 0 |  | 16065.75 | RT55-04-005 | 16065.75 | 16065.75 | 90 | 0 | 90 | 1 | Recommended
RT55-04-006 | Physiotherapy | Free Standing Foldable Parallel Bar Description:A foldable free standing unit used for gait training (walking) of patients Technical Specification :Parallel bars designed to be free standing and easily folded to the floor Handrails measure at least 3m in length, and safely fold and latched against the wall vertically when not in use Folded dimension approximately 25cm deep when folded Folding mechanism to be ergonomic and durable Dimensions: ± 300cm length ±40-70cm width ±50cm - 110cm height adjustments to be made via telescoping tubing with durable locking pins. Height should offer 3-4cm adjustment increments within the specified range ,adjustable width of approximately 30-70cm Handrails shall be durable, smooth non chafing, powder or chrome coated steel .Handrails shall be approximately 4cm in diameter.Heavy gauge steel base plates, uprights and fittings .Base plates to not cause obstruction to ambulation. | New Horizon Metals cc | New Horizon Metals cc | South Africa | NHM | FPB001 | 2022/2023 | 7 years | 1 | Each | 17250 | 0 | 0 | 0 | 0 | 0 |  | 17326 | RT55-04-006 | 17326 | 17326 | 90 | 0 | 90 | 1 | Recommended
RT55-04-007 | Physiotherapy | Hydrocollator 12 packs, fixed unit Description:A unit used to warm moist heat packs Technical Specification :This specification establishes the requirements and supply of a mobile PHYSIOTHERAPY – HYDROCOLLATOR UNIT. The unit offered shall be specifically designed to provide consistent heat for HOTPACKS that will be used to treat patients with conditions requiring moist heat therapy treatment The unit shall be suitable for a high usage clinic.The hydrocollator heating unit shall have the capacity to accommodate at least 12 heat packs .The unit shall have a durable stainless steel water tank that houses heat packs The unit shall have a low water sensor that shuts off the unit when the water level gets too low The unit shall have full insulation to prevent heat loss The unit offered shall be fitted with a reliable temperature control device to regulate the temperature at a temperature range of between 70º C to 75º C. The accuracy of the temperature regulation must be within ± 10% or better The heat up time to 70 º C must not exceed twofour hours on the unit offered The cool down time to 70 º C must not exceed two hours on the unit offered The unit offered must be fitted with a safety temperature cut out device that will shut off power to the heater when it exceeds a temperature in a range of between 82º C to 85º C The unit offered must incorporate a reliable drain mechanism to drain the water The unit shall be easy to clean .The unit shall not require any plumbing The unit shall be easy to maintain. Maintenance specifications to be supplied by bidder The bidder shall specify the dimensions of the unit .Each unit delivered shall include 8 standard sized (25cm x 30cm) hot packs, 2 cervical hot packs and 2 large hot packs .Each unit delivered shall include heat resistant tongs Accessories: castors with locks , heat pack teri covers: standard & cervical (individually priced) all sizes heat packs The unit shall operate from 220V and the power cords shall have a 3-point plug approved by SABS The unit offered must be supplied with internal removable stainless steel racks that will accommodate all sizes and shapes of hot packs . All accessories and consumables must be supplied standard with the unit A separate price of each accessory and consumables listed above must also be supplied with this offer Any other additional accessories, consumables and spare parts not listed above but applicable for the functionality of the unit must be supplied on a separate price list | HITECH THERAPY | FABRICATION ENTERPRISES | SOUTH KOREA | RELIEF PACK | 111975 | 2017 | 7 YEARS | 1 | DEGREES CELSIUS | 54500 | 750 | 750 | 750 | 750 | 750 |  | 57638 | RT55-04-007 | 57638 | 57638 | 90 | 0 | 90 | 1 | Recommended
RT55-04-008 | Physiotherapy | Hydrocollator 12 Packs, mobile unit DescriptionA mobile unit used to warm moist heat packs Technical Specification :This specification establishes the requirements and supply of a mobile PHYSIOTHERAPY – HYDROCOLLATOR UNIT. The unit offered shall be specifically designed to provide consistent heat for HOTPACKS that will be used to treat patients with conditions requiring moist heat therapy treatment .The unit shall be suitable for a high usage clinic The hydrocollator heating unit shall have the capacity to accommodate at least 12 heat packs The unit shall have a durable stainless steel water tank that houses heat packs The unit shall have a low water sensor that shuts off the unit when the water level gets too low The unit shall have full insulation to prevent heat loss .The unit offered shall be fitted with a reliable temperature control device to regulate the temperature at a temperature range of between 70º C to 75º C. The accuracy of the temperature regulation must be within ± 10% or better .The heat up time to 70 º C must not exceed four two hours on the unit offered The cool down time to 70 º C must not exceed two hours on the unit offered .The unit offered must be fitted with a safety temperature cut out device that will shut off power to the heater when it exceeds a temperature in a range of between 82º C to 85º C .The unit offered must incorporate a reliable drain mechanism to drain the water .The unit shall be mobile and equipped with at least four swivel castors .At least two of the castors shall have locks The unit shall be easy to clean .The unit shall not require any plumbing The unit shall be easy to maintain. Maintenance specifications to be supplied by bidder The bidder shall specify the dimensions of the unit .Each unit delivered shall include 8 standard sized (25cm x 30cm) hot packs, 2 cervical hot packs and 2 large hot packs .Each unit delivered shall include heat resistant tongs Accessories: castors with locks ,heat pack teri covers: standard & cervical (individually priced) all sizes heat packs The unit shall operate from 220V and the power cords shall have a 3-point plug approved by SABS The unit offered must be supplied with internal removable stainless steel racks that will accommodate all sizes and shapes of hot packs . All accessories and consumables must be supplied standard with the unit A separate price of each accessory and consumables listed above must also be supplied with this offer Any other additional accessories, consumables and spare parts not listed above but applicable for the functionality of the unit must be supplied on a separate price list | HITECH THERAPY | FABRICATION ENTERPRISES | SOUTH KOREA | RELIEF PACK | 111975 | 2017 | 7 YEARS | 1 | DEGREES CELSIUS | 54500 | 750 | 750 | 750 | 750 | 750 |  | 57639 | RT55-04-008 | 57639 | 57639 | 90 | 0 | 90 | 1 | Recommended
RT55-04-009 | Physiotherapy | Hydrocollator 4 packs, fixed unit Description:A unit used to warm moist heat packs Technical Specification :This specification establishes the requirements and supply of a mobile PHYSIOTHERAPY – HYDROCOLLATOR UNIT. The unit offered shall be specifically designed to provide consistent heat for HOTPACKS that will be used to treat patients with conditions requiring moist heat therapy treatment The unit shall be suitable for a high usage clinic .The hydrocollator heating unit shall have the capacity to accommodate at least 4 heat packs The unit shall have a durable stainless steel water tank that houses heat packs The unit shall have a low water sensor that shuts off the unit when the water level gets too low The unit shall have full insulation to prevent heat loss .The unit offered shall be fitted with a reliable temperature control device to regulate the temperature at a temperature range of between 70º C to 75º C. The accuracy of the temperature regulation must be within ± 10% or better .The heat up time to 70 º C must not exceed four two hours on the unit offered .The cool down time to 70 º C must not exceed two hours on the unit offered .The unit offered must be fitted with a safety temperature cut out device that will shut off power to the heater when it exceeds a temperature in a range of between 82º C to 85º C .The unit offered must incorporate a reliable drain mechanism to drain the water .The unit shall be easy to clean .The unit shall not require any plumbing The unit shall be easy to maintain. Maintenance specifications to be supplied by bidder The bidder shall specify the dimensions of the unit .Each unit delivered shall include 3 standard sized (25cm x 30cm) hot packs and 1 cervical hot pack .Each unit delivered shall include heat resistant tongs Accessories: castors with locks ,heat pack teri covers: standard & cervical (individually priced) all sizes heat packs The unit shall operate from 220V and the power cords shall have a 3-point plug approved by SABS The unit offered must be supplied with internal removable stainless steel racks that will accommodate all sizes and shapes of hot packs . All accessories and consumables must be supplied standard with the unit A separate price of each accessory and consumables listed above must also be supplied with this offer Any other additional accessories, consumables and spare parts not listed above but applicable for the functionality of the unit must be supplied on a separate price list | HITECH THERAPY | FABRICATION ENTERPRISES | SOUTH KOREA | RELIEF PACK | 111971 | 2017 | 7 YEARS | 1 | DEGREES CELSIUS | 15500 | 750 | 750 | 750 | 750 | 750 |  | 18640 | RT55-04-009 | 18640 | 18640 | 90 | 0 | 90 | 1 | Recommended
RT55-04-010 | Physiotherapy | Hydrocollator 4 Packs, mobile unit Description :A mobile unit used to warm moist heat packs Technical Specification :This specification establishes the requirements and supply of a mobile Physiotherapy – Hydrocollator Unit. The unit offered shall be specifically designed to provide consistent heat for HOTPACKS that will be used to treat patients with conditions requiring moist heat therapy treatment .The unit shall be suitable for a high usage clinic The hydrocollator heating unit shall have the capacity to accommodate at least 4 heat packs The unit shall have a durable stainless steel water tank that houses heat packs .The unit shall have a low water sensor that shuts off the unit when the water level gets too low .The unit shall have full insulation to prevent heat loss The unit offered shall be fitted with a reliable temperature control device to regulate the temperature at a temperature range of between 70º C to 75º C. including an auto on off function.The accuracy of the temperature regulation must be within ± 10% or better .The quick heat up time to 70 º C must not exceed twofour hourss on the unit offered The cool down time to 70 º C must not exceed two hours on the unit offered The unit offered must be fitted with a safety temperature cut out device that will shut off power to the heater when it exceeds a temperature in a range of between 82º C to 85º C .The unit offered must incorporate a reliable drain mechanism to drain the water .The unit shall be mobile and equipped with at least four swivel castors At least two of the castors shall have locks.The unit shall be easy to clean The unit shall not require any plumbing .The unit shall be easy to maintain. Maintenance specifications to be supplied by bidder .The bidder shall specify the dimensions of the unit Each unit delivered shall include 3 standard sized (25cm x 30cm) hot packs and 1 cervical hot pack Each unit delivered shall include heat resistant tongs Accessories: castors with locks heat pack teri covers: standard & cervical (individually priced) all sizes heat packs The unit shall operate from 220V and the power cords shall have a 3-point plug approved by SABS The unit offered must be supplied with internal removable stainless steel racks that will accommodate all sizes and shapes of hot packs .All accessories and consumables must be supplied standard with the unit A separate price of each accessory and consumables listed above must also be supplied with this offer Any other additional accessories, consumables and spare parts not listed above but applicable for the functionality of the unit must be supplied on a separate price list . | HITECH THERAPY | FABRICATION ENTERPRISES | SOUTH KOREA | RELIEF PACK | 111971 | 2017 | 7 YEARS | 1 | DEGREES CELSIUS | 15500 | 750 | 750 | 750 | 750 | 750 |  | 18641 | RT55-04-010 | 18641 | 18641 | 90 | 0 | 90 | 1 | Recommended
RT55-04-011 | Physiotherapy | Hydrocollator 8 packs, fixed unit Description:A unit used to warm moist heat packs Technical Specification:This specification establishes the requirements and supply of a mobile PHYSIOTHERAPY – HYDROCOLLATOR UNIT. The unit offered shall be specifically designed to provide consistent heat for HOTPACKS that will be used to treat patients with conditions requiring moist heat therapy treatment The unit shall be suitable for a high usage clinic .The hydrocollator heating unit shall have the capacity to accommodate at least 8 heat packs .The unit shall have a durable stainless steel water tank that houses heat packs The unit shall have a low water sensor that shuts off the unit when the water level gets too low The unit shall have full insulation to prevent heat loss The unit offered shall be fitted with a reliable temperature control device to regulate the temperature at a temperature range of between 70º C to 75º C. The accuracy of the temperature regulation must be within ± 10% or better The heat up time to 70 º C must not exceed two four hours on the unit offered The cool down time to 70 º C must not exceed two hours on the unit offered The unit offered must be fitted with a safety temperature cut out device that will shut off power to the heater when it exceeds a temperature in a range of between 82º C to 85º C.The unit offered must incorporate a reliable drain mechanism to drain the water .The unit shall be easy to clean .The unit shall not require any plumbing The unit shall be easy to maintain. Maintenance specifications to be supplied by bidder The bidder shall specify the dimensions of the unit. Each unit delivered shall include 6 standard sized (25cm x 30cm) hot packs and 2 cervical hot pack .Each unit delivered shall include heat resistant tongs Accessories: castors with locks, heat pack teri covers: standard & cervical (individually priced) all sizes heat packs The unit shall operate from 220V and the power cords shall have a 3-point plug approved by SABS The unit offered must be supplied with internal removable stainless steel racks that will accommodate all sizes and shapes of hot packs . All accessories and consumables must be supplied standard with the unit A separate price of each accessory and consumables listed above must also be supplied with this offer Any other additional accessories, consumables and spare parts not listed above but applicable for the functionality of the unit must be supplied on a separate price list | HITECH THERAPY | FABRICATION ENTERPRISES | SOUTH KOREA | RELIEF PACK | 111974 | 2017 | 7 YEARS | 1 | DEGREES CELSIUS | 46000 | 750 | 750 | 750 | 750 | 750 |  | 49142 | RT55-04-011 | 49142 | 49142 | 90 | 0 | 90 | 1 | Recommended
RT55-04-012 | Physiotherapy | Hydrocollator 8 Packs, mobile unit Description:A mobile unit used to warm moist heat packs Technical Specification:This specification establishes the requirements and supply of a mobile Physiotherapy – Hydrocollator Unit. The unit offered shall be specifically designed to provide consistent heat for HOTPACKS that will be used to treat patients with conditions requiring moist heat therapy treatment The unit shall be suitable for a high usage clinic .The hydrocollator heating unit shall have the capacity to accommodate at least 8 heat packs .The unit shall have a durable stainless steel water tank that houses heat packs The unit shall have a low water sensor that shuts off the unit when the water level gets too low The unit shall have full insulation to prevent heat loss The unit offered shall be fitted with a reliable temperature control device to regulate the temperature at a temperature range of between 70º C to 75º C. The accuracy of the temperature regulation must be within ± 10% or better The heat up time to 70 º C must not exceed four hourstwo hours on the unit offered The cool down time to 70 º C must not exceed two hours on the unit offered The unit offered must be fitted with a safety temperature cut out device that will shut off power to the heater when it exceeds a temperature in a range of between 82º C to 85º C .The unit offered must incorporate a reliable drain mechanism to drain the water .The unit shall be mobile and equipped with at least four swivel castors At least two of the castors shall have locks .The unit shall be easy to clean .The unit shall not require any plumbing .The unit shall be easy to maintain. Maintenance specifications to be supplied by bidder .The bidder shall specify the dimensions of the unit .Each unit delivered shall include 6 standard sized (25cm x 30cm) hot packs and 2 cervical hot pack .Each unit delivered shall include heat resistant tongs Accessories: castors with locks heat pack teri covers: standard & cervical (individually priced) all sizes heat packs The unit shall operate from 220V and the power cords shall have a 3-point plug approved by SABS The unit offered must be supplied with internal removable stainless steel racks that will accommodate all sizes and shapes of hot packs .All accessories and consumables must be supplied standard with the unit A separate price of each accessory and consumables listed above must also be supplied with this offer Any other additional accessories, consumables and spare parts not listed above but applicable for the functionality of the unit must be supplied on a separate price list | HITECH THERAPY | FABRICATION ENTERPRISES | SOUTH KOREA | RELIEF PACK | 111974 | 2017 | 7 YEARS | 1 | DEGREES CELSIUS | 46000 | 750 | 750 | 750 | 750 | 750 |  | 49143 | RT55-04-012 | 49143 | 49143 | 90 | 0 | 90 | 1 | Recommended
RT55-04-014 | Physiotherapy | PARAFIN WAX BATH Description : Warm paraffin wax bath used to treat Musculo skeletal pain and inflammation Specifications: Quantity: 1 x Paraffin wax bath 5 x refill wax 1kg each Description • PARAFFIN WAX BATH - 220V • At least 3KG warming unit capacity • Adjustable temperature settings. • With clearly visible thermostat and temperature gauge • Power indicator light • Safety auto shutdown/overheat function • Lid to cover wax bath. • Electrical cord and plug. Hospital grade. • Accommodates hands, elbow and feet. • Fully insulated. • Cool to touch exterior • Compatible Paraffin wax available to purchase separately in 1kg packages | Sanbonani Holdings (Pty) Ltd | North Coast Medical Inc. | USA | Therabath® | NC15494-220V | 2012 | 10 years | 1 | Kit | 10650 | 2800 | 3024 | 3265.92 | 3527.1936 | 3879.913 |  | 23379.1136 | RT55-04-014 | 23379.1136 | 23379.1136 | 90 | 0 | 90 | 1 | Recommended
RT55-04-015 | Physiotherapy | Parallel Bars 3m, Wall Mounted, Foldable (powder/chrome) Description :A wall mounted, vertically foldable unit used for gait training (walking) of patients Technical Specification:Parallel bars designed to be wall mounted and easily folded to the floor and back to the wall, vertically Handrails measure at least 3m in length, and safely fold and latched against the wall vertically when not in use Folded dimension approximately 25cm deep when folded .Folding mechanism to be ergonomic and durable Dimensions: ± 300cm length ±40-70cm width ±50cm - 110cm height adjustments to be made via telescoping tubing with durable locking pins. Height should offer 3-4cm adjustment increments within the specified range adjustable width of approximately 30-70cm .Handrails shall be durable, powder or chrome coated steel Handrails shall be 3.8cm in diameter .Heavy gauge steel base plates, uprights and fittings The unit shall have a weight-bearing capacity of at least 250kg, with sturdy, durable base plates All accessories and consumables must be supplied standard with the unit A separate price of each accessory and consumables listed above must also be supplied with this offer Any other additional accessories, consumables and spare parts not listed above but applicable for the functionality of the unit must be supplied on a separate price list . | New Horizon Metals cc | New Horizon Metals cc | South Africa | NHM | PBW002 | 2022/2023 | 7 years | 1 | Each | 22099.91 | 0 | 0 | 0 | 0 | 0 |  | 22176.91 | RT55-04-015 | 22176.91 | 22176.91 | 90 | 0 | 90 | 1 | Recommended
RT55-04-018 | Physiotherapy | Parallel Bars 4m, Floor Mounted (stainless steel) Description :A fixed unit used for gait training (walking) of patients Technical Specification :Parallel bars designed to be floor mounted The unit shall have 4 floor mounted uprights with mounting plates Mounting plates should not cause an obstruction to ambulation The unit shall have heavy gauge stainless steel base plates, uprights and fittings Adjustable with durable mechanisms and locking pins, with ergonomic control knobs Dimensions: Length 400cm width ±40-70cm (min-max) height ±50cm - 110cm(min-max) The unit shall have a weight bearing capacity of at least 250kg Width adjusts with ergonomic control knobs on each upright adjustments to be made via telescoping tubing with durable locking pins. Height should offer 3-4cm adjustment increments within the specified range . Handrails shall be approx. 4cm in diameter Handrails shall be one-piece, stainless steel All spares including adjustment knobs, pins and screws should be listed and priced Any spare parts not listed above but applicable for the functionality of the unit must be supplied on a separate price list | New Horizon Metals cc | New Horizon Metals cc | South Africa | NHM | PBFSS002 | 2022/2023 | 7 years | 1 | Each | 19951.11 | 0 | 0 | 0 | 0 | 0 |  | 20029.11 | RT55-04-018 | 20029.11 | 20029.11 | 90 | 0 | 90 | 1 | Recommended
RT55-04-019 | Physiotherapy | Parallel Bars m, Floor Mounted (powder/chrome) Description :A fixed unit used for gait training (walking) of patients Technical Specification: Parallel bars designed to be floor mounted .The unit shall have 4 floor mounted uprights with mounting plates .The unit shall have heavy gauge steel base plates, uprights and fittings Adjustable height and width, with ergonomic, durable mechanisms and locking pins, with ergonomic control knobs Dimensions: L x W x H: 3m x 38cm – 71cm x 70cm – 108cm, approximately The unit shall have a weight bearing capacity of approximately 250kg Width shall be adjustable from ±15"(38cm) - ±28"(71cm) Width adjusts with ergonomic control knobs on each upright Each upright telescope will extend in 3-4cm increments with durable locking pins Handrails shall be approx. 3.8cm in diameter .Handrails shall be one-piece, made of powder-coated / chrome plated steel . All accessories and consumables must be supplied standard with the unit . A separate price of each accessory and consumables listed above must also be supplied with this offer. Any other additional accessories, consumables and spare parts not listed above but applicable for the functionality of the unit must be supplied on a separate price list. | HITECH THERAPY | PEBJ | SOUTH AFRICA | SYNERGY | 0010 | 2015 | 7 YEARS | 1 | METER | 17595 | 0 | 0 | 0 | 0 | 0 |  | 17739 | RT55-04-019 | 17739 | 17739 | 90 | 0 | 90 | #REF! | Recommended
RT55-04-020 | Physiotherapy | Physiotherapy Exercise Software computer based Description :Software used to design personalised self exercise home programmes Technical Specification :User-friendly software exercise programme, compatible with public sector available hardware Should be supplied as a DVD and , installable on a computer The software shall be compatible with Windows7 / Window8/ Windows10 compatible software (otherwise it must be clearly stated in offer and be communicated with end user) The software shall be installable to computer, not internet based or dependable on internet access Availability of a comprehensive range of multi-professional content of exercises and education encompassing different specialised areas regularly treated by health care workers, including but not limited to (should be available to purchased as a full package or separately as individual modules : general exercises back care balance re-education cardiovascular exercises flexibility exercises contracture prevention incontinence and pelvic floor exercises amputations respiratory care walking with walking aids therapeutic balls exercises resistance band exercises stick exercises rehabilitation post CVA motor delay for paeds CP related therapy burns care and exercises geriatrics posture and ergonomics The software shall have a lifelong license .It shall be possible to install the software on multiple (at least 5) computers Available in English and at least 1 other official language of South Africa Updates regularly available .Logo and contact details of health care provider can be added Sharing of single, lifelong license for use on multiple computers (at least 5) .Designed programmes must be printable and can be emailed to clients Illustrations of exercises shall be clear and accurate .Instructions for exercises shall be clear and accurate with possibility of amending / personalisingpersonalizing Product training should be included .The unit shall be FDS or CE approved (Attach certification) All accessories and consumables must be supplied standard with the unit .A separate price of each accessory and consumables listed above must also be supplied with this offer .Any other additional accessories, consumables and spare parts not listed above but applicable for the functionality of the unit must be supplied on a separate price list. | HITECH THERAPY | PHYSIOTOOLS | FINLAND | PHYSIOTOOLS | PHYSIOTOOLS - RT55A66 | 2014 | 7 YEARS | 1 | N/A | 44850 | 0 | 0 | 0 | 0 | 0 |  | 44995 | RT55-04-020 | 44995 | 44995 | 90 | 0 | 90 | 1 | Recommended
RT55-04-021 | Physiotherapy | Saddle Stool, Mobile Description :A seating device used for optimal positioning of the therapist Technical Specification : The unit shall be a mobile saddle stool The unit shall be height adjustable (approximately 53cm - 73cm) The unit shall have 5 castors The seat shall be saddle shaped The seat shall be covered with vinyl or leather Accessories :Replaceable castors All accessories and consumables must be supplied standard with the unit A separate price of each accessory and consumables listed above must also be supplied with this offer Any other additional accessories, consumables and spare parts not listed above but applicable for the functionality of the unit must be supplied on a separate price list | New Horizon Metals cc | New Horizon Metals cc | South Africa | NHM | SS001 | 2022/2023 | 7 years | 1 | Each | 2305.75 | 0 | 0 | 0 | 0 | 0 |  | 2385.75 | RT55-04-021 | 2385.75 | 2385.75 | 90 | 0 | 90 | 1 | Recommended
RT55-04-021 | Physiotherapy | Saddle Stool, Mobile Description :A seating device used for optimal positioning of the therapist Technical Specification : The unit shall be a mobile saddle stool The unit shall be height adjustable (approximately 53cm - 73cm) The unit shall have 5 castors The seat shall be saddle shaped The seat shall be covered with vinyl or leather Accessories :Replaceable castors All accessories and consumables must be supplied standard with the unit A separate price of each accessory and consumables listed above must also be supplied with this offer Any other additional accessories, consumables and spare parts not listed above but applicable for the functionality of the unit must be supplied on a separate price list | Arjo Huntliegh South Africa | Seers Medical | UK | Economy saddle stool | RT55-04-21 | 2015 | 5 | 1 | EA | 4804.53 | 501.92 | 501.92 | 501.92 | 501.92 | 501.92 |  | 6814.21 | RT55-04-021 | 6814.21 | 2385.75 | -77.0592 | 0 | -77.0592 | 2 | Recommended
RT55-04-022 | Physiotherapy | Stand-Up Manual Patient Lift Description:A stand-up unit used for patient transfer Technical Specification:Non-powered (non-electric) Safe minimum loading capacity of 180kg Weigh no more than 30kg Length x height: ±900mm x ± 1050mm Turning radius: ±1150mm Minimum door requirement 630mm Pedal operating force <300N Pedal operated adjustment of leg width Internal width legs open ±730mm.Parallel Bars 3m, Floor Mounted (powder/chrome). Low friction castors .Compliance with legislature and applicable product standards Shall have the following features: Must be easily cleaned using soap water or alcohol.ManoeuvringManeuvering handles. Cross bar/handle bar.Sling attachment lugs/clips for back support strap Pivoting seat halves/seat pads .Knee pad ,2 Front castors 2 Back castors with brakes Footboard/base (reinforced/heavy duty plastic) Leg pivot foot levers/pedals Expected frame life of at least 10 years Expected operational life of slings/backrest of at least 2 years Accessories: sling/back rests in a variety of sizes/adjustable size sling/back rests to be manufactured of a durable material sling/back rests shall be compatible with the purchased stand-up manual patient lift secure buckle for fastening around patient secure clips for fastening to frame ,sling attachment lugs/clips for back support strap knee pad castors . All accessories and consumables must be supplied standard with the unit . A separate price of each accessory and consumables listed above must also be supplied with this offer. Any other additional accessories, consumables and spare parts not listed above but applicable for the functionality of the unit must be supplied on a separate price list | Arjo Huntliegh South Africa | Arjo | Canada & Dominican Republic | Sara stedy | RT55-04-22 | 1995 | 15 | 1 | EA | 19358.11 | 1831.72 | 1831.72 | 1831.72 | 1831.72 | 1831.72 |  | 26687.99 | RT55-04-022 | 26687.99 | 26687.99 | 90 | 0 | 90 | 1 | Recommended
RT55-04-024 | Physiotherapy | Therapy Mats (4 sizes) Description:A mat used for rehabilitation .Technical Specification : The unit shall be available in the following approximate sizes: A: 10mm (thickness) x 55cm (width) x 140cm (length) B: 10mm x 55cm x 180cm C: 15mm x 100cm x 185cm D: 15mm x 100cm x 200cm Can be purchased as a set or individually. Different sizes mats can be procured separately as single units The unit shall be manufactured from a non-slip dense closed cell foam that is antibacterial and water impermeable Ribbed pattern for increased traction on floor Available in various colour options The unit shall be hard-wearing with a long life span The unit shall be comfortable to use - soft, supple, yielding and supportive The unit shall have a slip-proof surface Rounded or square corners . The unit shall conform to CE- high quality materials, strict quality controls | HITECH THERAPY | AIREX AG | SWITZERLAND | AIREX | 0486 0495 0409 200PL2W | 2012 | 5 YEARS | 1 | N/A | 10465 | 0 | 0 | 0 | 0 | 0 |  | 10611 | RT55-04-024 | 10611 | 10611 | 90 | 0 | 90 | 1 | Recommended
RT55-04-025 | Physiotherapy | WOODEN BOBATH PLINTH (NON ADJUSTABLE) Description:Wooden non adjustable wide (bobath) neurology plinth Technical specifications :Solid wood frame with non slip legs Plinth dimensions L 180cm W 150cm The unit shall have a minimum load capacity of atleast 250kg Wood to be treated appropriately for pests and safe for patient use. The unit shall be upholstered in durable hygienic, non-slip and washable (e.g. vinyl, or imitation leather) Upholstery high density foam shall be at least 4-6cm thick. The upholstery and frame shall be free of phthalates and shall be flame retardant | Sanbonani Holdings (Pty) Ltd | Geppetto's World of Developments cc t/a Playon | South Africa | Playon | GP-FHWBP200 | 2022 | 6 years | 1 | Set | 8250 | [6 empty] | 8363 | RT55-04-025 | 8363 | 8363 | 90 | 0 | 90 | 1 | Recommended
RT55-04-025 | Physiotherapy | WOODEN BOBATH PLINTH (NON ADJUSTABLE) Description:Wooden non adjustable wide (bobath) neurology plinth Technical specifications :Solid wood frame with non slip legs Plinth dimensions L 180cm W 150cm The unit shall have a minimum load capacity of atleast 250kg Wood to be treated appropriately for pests and safe for patient use. The unit shall be upholstered in durable hygienic, non-slip and washable (e.g. vinyl, or imitation leather) Upholstery high density foam shall be at least 4-6cm thick. The upholstery and frame shall be free of phthalates and shall be flame retardant | New Horizon Metals cc | New Horizon Metals cc | South Africa | NHM | WBP001 | 2022/2023 | 7 years | 1 | Each | 9522 | 0 | 0 | 0 | 0 | 0 |  | 9603 | RT55-04-025 | 9603 | 8363 | 76.6555 | 0 | 76.6555 | 2 | Recommended
RT55-05-008 | Podiatry | Goniometer, plastic.Description:"6 piece goniometer set- 14cm Finger Goniometer: Measures range-of-motion of finger joints (metacarpophalangeal and interphalangeal) and other small joints including toes. Head has two opposing 180° scales marked in 5° increments. All linear scales in inches and centimetres on the arm. 18cm Robinson Pocket Goniometer:Pocket-Sized goniometer. 180° scale in 5° increments. 4 inch linear scale and 10cm linear scale on the arm. 19 cm Goniometer – 180°:Two opposing scales marked in 1° increments. Tension between arms is controlled by the thumb knob. 36cm Goniometer – 180°: Two opposing scales marked in 1° increments. Extra-long arms for measuring all joints. Non-locking friction arm. 36cm Goniometer – 360°: One 360° and two 180° scales that read in opposite directions. All marked in 1° increments. Locking Arm. Knurled knob adjusts arm tension. 20cm X-Ray Goniometer: Document Angles. Two Opposing Scales marked in 2.5° increments. Black aluminium with easy-to-read white radiopaque marking.Hard Carry Case." Technical specification:Goniometer plastic, angle measure, 360 degrees | HITECH THERAPY | FABRICATION ENTERPRISES | USA | BASELINE | 121043 | 2021 | 7 YEARS | 1 | INCH + CM | 5635 | 0 | 0 | 0 | 0 | 0 |  | 5782 | RT55-05-008 | 5782 | 5782 | 90 | 0 | 90 | 1 | Recommended
RT55-05-047 | Podiatry | Patella Hammer. Description: Medical istrument used to test the deep tendon reflexes, which help indicate the health of overall nervous system A.utoclavable. Technical Specification:Patella hammer, stainless steel, high grade, rust free, buck, neurological, rubber percussion head with needle probe and brush, 19-23cm long, can be sterilised .Patella hammer, round head, plastic handle with probe end, 36cm long .Reflex hammer, with a 18-20cm long octagonal chrome plated handle with probe and synthetic rubber head. | HITECH THERAPY | PERFORMANCE HEALTH | UK | PERFORMANCE HEALTH | HT9015 | 2017 | 7 YEARS | 1 | N/A | 1725 | 0 | 0 | 0 | 0 | 0 |  | 1873 | RT55-05-047 | 1873 | 1873 | 90 | 0 | 90 | 1 | Recommended
RT55-05-048 | Podiatry | Podiatrist Chair.Description: Clinician or operators chair, used by practitioner while consulting the patient that supports the spine. Technical Specification:Podiatrist chair, adjustable tilt and height so that it can be adapted to a person’s individual working position.Has a plastic, 5-star castor base for support and balance. Fire retardant upholstery to meet current Health & Safety regulations. | New Horizon Metals cc | New Horizon Metals cc | South Africa | NHM | PC001 | 2022/2023 | 7 years | 1 | Each | 4427.5 | 885.5 | 956.34 | 1032.85 | 1115.47 | 1204.71 |  | 8499.66 | RT55-05-048 | 8499.66 | 8499.66 | 90 | 0 | 90 | 1 | Recommended
RT55-05-050 | Podiatry | Precision Medical Fiber (Monofilament) Description: Is a device used for testing of peripheral neuropathy in diabetic patients Technical Specification:nylon, retractable, 10g | Duromed cc | North Coast Medical | USA | TOUCH TEST | NC12775-14 | N/A | 2 YEARS | OFFER 1 | EACH | 770.5 | 0 | 0 | 0 | 0 | 0 |  | 790.5 | RT55-05-050 | 790.5 | 790.5 | 90 | 6.9 | 96.9 | 1 | Recommended
RT55-06-010 | Speech | Cause and Effect Cards. Description: A game used for adults with neurological impairments for communication therapy. Technical Specification: Preference will be given to images that are relevant to the South African context.Cause and effect cards consisting of 45-50 cards Images on cards must be colour photographs.Must reflect a variety of cultures, races, daily activities and everyday objects. Cards must be high quality laminated or gloss with rounded corners. Each set of cards must come in a sturdy box with a manual. Card dimensions must be at least 15x10cm | Navitrons | Routledge | United Kingdom | SpeechMark: Cause and Effect | ISBN: 9780863883552 | 1997 | Not Applicable | 1 | 1 | 2250 | [6 empty] | 2313 | RT55-06-010 | 2313 | 2313 | 90 | 0 | 90 | 1 | Recommended
RT55-06-019 | Speech | Human Brain Model3D anatomical model used for patient counselling and education. Technical Specification: Educational model of the human brain. Three dimensional model which can be dismantled to see vessels and nerves and re-assembled. Clearly details the different structures/landmarks of the brain such as frontal, parietal, occipital and temporal lobes, cerebellum, corpus callosum, pons, medulla oblongata etc. Depicts both internal and external anatomy of the brain. Anatomically accurateLife size (adult)Made from durable materialStand included. The model must be supplied with clear labeling of anatomical structures and a literature of supporting structure, general location and functions. | HITECH THERAPY | FABRICATION ENTERPRISES | USA | 3B SCIENTIFIC | 124562 | 2011 | 7 YEARS | 1 | N/A | 10580 | 0 | 0 | 0 | 0 | 0 |  | 10729 | RT55-06-019 | 10729 | 10729 | 90 | 0 | 90 | 1 | Recommended
RT55-06-020 | Speech | Human Ear Model3D anatomical model used for patient counselling and education.Technical Specification: Educational model of the human ear Four dimensional model which can be dismantled and re-assembled.Clearly details the outer, middle and inner ear, which can be dismantled into four pieces.Removable eardrum, malleus, incus and stapes.Anatomically accurateLarge size. Made from durable material.Stand included.The model must be supplied with clear labeling of anatomical structures and a literature of supporting structure, general location and functions. | HITECH THERAPY | FABRICATION ENTERPRISES | USA | 3B SCIENTIFIC | 124565 | 2016 | 7 YEARS | 1 | N/A | 10580 | 0 | 0 | 0 | 0 | 0 |  | 10730 | RT55-06-020 | 10730 | 10730 | 90 | 0 | 90 | 1 | Recommended
RT55-06-022 | Speech | Human Skull Model 3D anatomical model used for patient counselling and education . Technical Specification: Educational model of the human skull. Three dimensional model which can be dismantled and re-assembled. Clearly details bone structures, joints, sutures, fissures etc. Separates into at least 3 parts (e.g. calvaria, base of skull, mandible) and details both internal and external anatomy of the skull. Anatomically accurate.Life size (adult). Made from durable material. Stand included. The model must be supplied with clear labeling of anatomical structures and a literature of supporting structure, general location and functions. | HITECH THERAPY | FABRICATION ENTERPRISES | USA | 3B SCIENTIFIC | 124547 | 2016 | 7 YEARS | 1 | N/A | 6325 | 0 | 0 | 0 | 0 | 0 |  | 6476 | RT55-06-022 | 6476 | 6476 | 90 | 0 | 90 | 1 | Recommended
RT55-06-023 | Speech | LCD writing pad. Electronic LCD writing/drawing pad. LCD Minimum 8.5 inch LCD screen with stylus attached. Battery operated; must come with appropriately sized batteries | Navitrons | Parrot Products (Pty) LTd. | South Africa | Parrot: 10" LCD Writing Tablet Slate | SKU: BD2810 |  | Not Applicable | 1 | 1 | 750 | [6 empty] | 814 | RT55-06-023 | 814 | 814 | 90 | 0 | 90 | 1 | Recommended
RT55-06-023 | Speech | LCD writing pad. Electronic LCD writing/drawing pad. LCD Minimum 8.5 inch LCD screen with stylus attached. Battery operated; must come with appropriately sized batteries | Morele Group | Union Toys | China | Union Toys | ND7PHKZ | 2024 | 2-5 Years | 1 | CM | 1200 | 0 | 0 | 0 | 0 | 0 |  | 1252 | RT55-06-023 | 1252 | 814 | 41.5725 | 0 | 41.5725 | 2 | Recommended
RT55-06-025 | Speech | Paediatric Basic Sequence Cards: (3 steps). Description: A game used for paediatrics with neurological impairments for communication therapy . Technical Specifications: Basic sequences cards: 3 step consisting of 45-50 cards Images on cards must be colour photographs .Must reflect a variety of cultures, races, daily activities and everyday objects.Cards must be high quality laminated or gloss with rounded corners. Each set of cards must come in a sturdy box with a manual Card dimensions must be at least 15x10cm | Navitrons | Routledge | United Kingdom | SpeechMark: Basic Sequence | ISBN: 9780863883514 | 2001 | Not Applicable | 1 | 1 | 2250 | [6 empty] | 2315 | RT55-06-025 | 2315 | 2315 | 90 | 0 | 90 | 1 | Recommended
RT55-06-026 | Speech | Paediatric Basic Verbs Cards : Description: A collection of pictures and board games used for paediatrics with neurological impairments for communication therapy . Technical Specifications: Basic verbs cards consisting of 45-50 cards . Images on cards must be colour photographs. Must reflect a variety of cultures, races, daily activities and everyday objects .Cards must be high quality laminated or gloss with rounded corners.Each set of cards must come in a sturdy box with a manual Card dimensions must be at least 15x10cm | Navitrons | Routledge | United Kingdom | SpeechMark: Basic Verbs | ISBN: 9781909301948 | 2016 | Not Applicable | 1 | 1 | 2250 | [6 empty] | 2316 | RT55-06-026 | 2316 | 2316 | 90 | 0 | 90 | 1 | Recommended
RT55-06-029 | Speech | Paediatric Everyday Objects Cards . Description: A game used for paediatrics with neurological impairments for communication therapy. Technical Specifications :Everyday objects cards consisting of 45-50 cards . Images on cards must be colour photographs Must reflect a variety of cultures, races, daily activities and everyday objects Cards must be high quality laminated or gloss with rounded corners Each set of cards must come in a sturdy box with a manual Card dimensions must be at least 15x10cm | Navitrons | Routledge | United Kingdom | SpeechMark: Everyday Objects | ISBN: 9781909301955 | 2016 | Not Applicable | 1 | 1 | 2250 | [6 empty] | 2317 | RT55-06-029 | 2317 | 2317 | 90 | 0 | 90 | 1 | Recommended
RT55-06-030 | Speech | Paediatric Familiar Verbs Cards. Description: A game used for paediatrics with neurological impairments for communication therapy. Technical Specifications: Familiar verbs cards consisting of 45-50 cards Images on cards must be colour photographs Must reflect a variety of cultures, races, daily activities and everyday objects Cards must be high quality laminated or gloss with rounded corners Each set of cards must come in a sturdy box with a manual Card dimensions must be at least 15x10cm | Navitrons | Routledge | United Kingdom | SpeechMark: Familiar Verbs | ISBN: 9780863885198 | 1999 | Not Applicable | 1 | 1 | 2250 | [6 empty] | 2318 | RT55-06-030 | 2318 | 2318 | 90 | 0 | 90 | 1 | Recommended
RT55-06-031 | Speech | Paediatric Feeling Cards . Description: A game used for paediatrics with neurological impairments for communication therapy. Technical Specifications :Feelings cards consisting of 45-50 cards . Images on cards must be colour photographs.Must reflect a variety of cultures, races, daily activities and everyday objects. Cards must be high quality laminated or gloss with rounded corners. Each set of cards must come in a sturdy box with a manual Card dimensions must be at least 15x10cm | Navitrons | Routledge | United Kingdom | SpeechMark: Emotions | ISBN: 9781138102958 | 2017 | Not Applicable | 1 | 1 | 2250 | [6 empty] | 2319 | RT55-06-031 | 2319 | 2319 | 90 | 0 | 90 | 1 | Recommended
RT55-06-032 | Speech | Paediatric Food Cards. Description: A game used for paediatrics with neurological impairments for communication therapy. Technical Specifications: Food cards consisting of 95-100 cards. Images on cards must be colour photographs. Must reflect a variety of cultures, races, daily activities and everyday objects. Cards must be high quality laminated or gloss with rounded corners.Each set of cards must come in a sturdy box with a manual Card dimensions must be at least 15x10cm | Navitrons | Routledge | United Kingdom | SpeechMark: Food | ISBN: 9781138570412 | 2018 | Not Applicable | 1 | 1 | 2850 | [6 empty] | 2920 | RT55-06-032 | 2920 | 2920 | 90 | 0 | 90 | 1 | Recommended
4485428.2984
8774630.6768
13260058.9752

## Category B
Item Number | Category | Item Description | Bidder Name | Name of Manufacturer | Couuntry of Origin | Brand Name | Model/Model Number | Year of Release | Expected Life Span of the Equipment | No. of Offers | Unit of Measure | Equipment/Accessory price inclusive of VAT, 1st and 2nd Year Maintenance Costs | 3rd Year Maintenance Costs | 4th Year Maintenance Costs | 5th Year Maintenance Costs | 6th Year Maintenance Costs | 7th Year Maintenance Costs | Calibration costs annum | Total Costs | Sort field | Total Costs | LOWEST ACCEPTABLE BID | POINTS FOR PRICE | HDI | Total Points | Ranking | Comment
RT55-01-016 | Audiology | Soundproof booth with extraction fan and soundproof observation/clinician area (large) Description:Soundproof booth with ventilation and extractor fan (ventilation unit for TB protection with minimal noise output).Technical Specification:Soundproof booth with ventilation and extractor fan (with minimal noise output).Should contain observation safety glass window and work platform, large enough to attachment in the soundproof observation/ clinician area. Placement of access door to be determined at site visit .Observation soundproof area should have external dimensions of width: 2200mm, height 2400mm, length: 2200mm or smaller as per customer request before manufacture internal dimensions of width: 2000mm, height: 2000mm, length: 2000mm or smaller as per customer requests before manufacture. The booth should have sound absorption in accordance with ISO 9001, SANS or equivalent.Inside ambient noise levels compliant with the ANSI S.1-1999 or equivalent.Minimum booth external dimensions of length: 2200mm, width: 2200mm, height 2400mm. Minimal booth internal dimensions of length: 2000mm, width: 2000mm, height: 2000mm.At least 12 input and output jacks (for air conduction left and right, bone conduction, inserts left and right, speakers left and right, extended high frequency headphones left and right, VRA left and right, talk forward, and talk back).Should contain observation safety glass window which is appropriately acoustically sealed and work platform attachment which is appropriately sized to accommodate audiometer, accessories and a writing surface. The booth should have sound absorption in accordance with ISO 9001 or equivalent.The proposed audiology booth is to comply with SANS 10182: 2006 as amended– the measurements and assessment of acoustic environments for audiometric test.The booth must be mounted on anti-vibration mounts to avoid any structural transfer of sound from the main structure. Should have adequate silent ventilation system that does not generate sound in the booth or conduct sound through the envelope of the booth resulting in an ambient sound pressure level above the prescribed limit per octave band.Outward opening door that can be opened from inside as well as outside of the booth, is not lockable from the inside, is at least 900 mm wide, and must appropriately acoustically seal so as not to compromise the required ambient sound pressure level inside the booth. Comes standard with wheelchair access ramp that does not impede door closure. Input jacks for connecting equipment on the outside of the booth . Output jacks for connecting transducers inside the booth.Adequate lighting supplied within the booth and clinician test area when appropriate. Electrical fittings and wiring to have standard certified safety measures, including trip switches.All penetrations through the envelope for electrical services must be sealed and must not compromise the ambient sound pressure level inside the booth.On/off switches for lighting and ventilation to be installed in easy to reach location for audiologist. Two standard 3 pin electrical outlet to be installed at suitable location in relation to audio logical equipment. | Lebone Medical Supplies (Pty) Ltd t/a Amtronix Diagnostics | Amtronix Diagnostics | South Africa | Amtronix Duagnostics AX | AX4 Dual room booth | 2020 | 10 Years | 1 | 1 Unit | 373750 | 9775 | 10752.5 | 11339 | 11925.5 | 12707.5 | 0 | 430249.5 | RT55-01-016 | 430249.5 | 430249.5 | 90 | 5.103 | 95.103 | 1 | Recommended
RT55-01-017 | Audiology | VIDEO HEAD IMPULSE TEST (VHIT). Description:Synapsys Video Head Impulse Test (VHIT) allows to assess the vestibular-ocular-reflex (VOR) by measuring, recording, displaying, and analyzing eye and head movements. Technical Specifications :Must provide at least 2 product configurations to test for example: 6 canals versus lateral canals. Must include required lenses and camera regarding the testing system. Must come with relevant accessories, and consumables. Relevant spec Laptop/ PC and software to made available. Must include relevant power supply and connectivity options for both PC and Printer. | Lebone Medical Supplies (Pty) Ltd t/a Amtronix Diagnostics | Inventis S.R.L. | Italy | Inventis | Synapsys Video Head Impulse Test (VHIT) with Laptop | 2022 - Continious firmware upgrades | 10 Years | 1 | 1 Unit | 289225 | 38525 | 42377.5 | 44689 | 47000.5 | 50082.5 | 4000 | 515899.5 | RT55-01-017 | 515899.5 | 515899.5 | 90 | 5.103 | 95.103 | 1 | Recommended
RT55-01-017 | Audiology | VIDEO HEAD IMPULSE TEST (VHIT). Description:Synapsys Video Head Impulse Test (VHIT) allows to assess the vestibular-ocular-reflex (VOR) by measuring, recording, displaying, and analyzing eye and head movements. Technical Specifications :Must provide at least 2 product configurations to test for example: 6 canals versus lateral canals. Must include required lenses and camera regarding the testing system. Must come with relevant accessories, and consumables. Relevant spec Laptop/ PC and software to made available. Must include relevant power supply and connectivity options for both PC and Printer. | Lebone Medical Supplies (Pty) Ltd t/a Amtronix Diagnostics | Natus Sensory Ltd | USA | Natus | ICS Impulse Video Head Impulse Test (VHIT) with Laptop | 2014 - Continious firmware upgrades | 10 Years | 2 | 1 Unit | 323724.08 | 38525 | 42377.5 | 44689 | 47000.5 | 50082.5 | 5000 | 551398.58 | RT55-01-017 | 551398.58 | 515899.5 | 83.8071 | 5.103 | 88.9101 | 2 | Recommended
RT55-01-018 | Audiology | DIAGNOSTIC VIDEONYSTAGMOGRAPHY (VNG):Diagnostic equipment videonystagmography (VNG) Description:Diagnostic assessment equipment, videonystagmography, equipment for the assessment of balance. Technical Specification: To include air calorics and digital light screen or equivalent. The test battery must include saccades, optokinetics, gaze, static positional, dix hallpike, calorics and other specialised tests. Relevant goggles to be supplied, must have ability to record, review, and transfer of information to allow for sharing of information. Compatible PC/Laptop with software to be supplied. Accessories to be included . | Lebone Medical Supplies (Pty) Ltd t/a Amtronix Diagnostics | Inventis S.R.L. | Italy | Inventis | Synapsys Videonystagmograph (VNG) with Laptop, Monitor and Air Caloric | 2018 - Continious firmware upgrades | 10 Years | 1 | 1 Unit | 398130 | 38525 | 42377.5 | 44689 | 47000.5 | 50082.5 | 5000 | 625804.5 | RT55-01-018 | 625804.5 | 625804.5 | 90 | 5.103 | 95.103 | 1 | Recommended
RT55-01-018 | Audiology | DIAGNOSTIC VIDEONYSTAGMOGRAPHY (VNG):Diagnostic equipment videonystagmography (VNG) Description:Diagnostic assessment equipment, videonystagmography, equipment for the assessment of balance. Technical Specification: To include air calorics and digital light screen or equivalent. The test battery must include saccades, optokinetics, gaze, static positional, dix hallpike, calorics and other specialised tests. Relevant goggles to be supplied, must have ability to record, review, and transfer of information to allow for sharing of information. Compatible PC/Laptop with software to be supplied. Accessories to be included . | Navitrons | Instrumentation Difra | Belgium | Difra | DI 140603: Easy-Headstar Monocular VNG System - Complete System | 204 | 10 years | 4 (Offer 4) | 1 | 479000 | 57480 | 63228 | 69550.8 | 76505.88 | 84156.468 | 0 | 829921.148 | RT55-01-018 | 829921.148 | 625804.5 | 60.645 | 6.9 | 67.545 | 2 | Recommended
RT55-01-018 | Audiology | DIAGNOSTIC VIDEONYSTAGMOGRAPHY (VNG):Diagnostic equipment videonystagmography (VNG) Description:Diagnostic assessment equipment, videonystagmography, equipment for the assessment of balance. Technical Specification: To include air calorics and digital light screen or equivalent. The test battery must include saccades, optokinetics, gaze, static positional, dix hallpike, calorics and other specialised tests. Relevant goggles to be supplied, must have ability to record, review, and transfer of information to allow for sharing of information. Compatible PC/Laptop with software to be supplied. Accessories to be included . | Navitrons | Instrumentation Difra | Belgium | Difra | DI 140601: Easy-VNG II Binocular VNG System - Complete System | 2017 | 10 years | 4 (Offer 2) | 1 | 521500 | 62580 | 68838 | 75721.8 | 83293.98 | 91623.378 | 0 | 903557.158 | RT55-01-018 | 903557.158 | 625804.5 | 50.055 | 6.9 | 56.955 | 3 | Recommended
RT55-01-018 | Audiology | DIAGNOSTIC VIDEONYSTAGMOGRAPHY (VNG):Diagnostic equipment videonystagmography (VNG) Description:Diagnostic assessment equipment, videonystagmography, equipment for the assessment of balance. Technical Specification: To include air calorics and digital light screen or equivalent. The test battery must include saccades, optokinetics, gaze, static positional, dix hallpike, calorics and other specialised tests. Relevant goggles to be supplied, must have ability to record, review, and transfer of information to allow for sharing of information. Compatible PC/Laptop with software to be supplied. Accessories to be included . | Navitrons | Instrumentation Difra | Belgium | Difra | DI 140504: Headstar V3 Monocular VNG System - Complete System | 2024 | 10 years | 4 (Offer 3) | 1 | 682000 | 81840 | 90024 | 99026.4 | 108929.04 | 119821.944 | 0 | 1181641.384 | RT55-01-018 | 1181641.384 | 625804.5 | 10.0624 | 6.9 | 16.9624 | 4 | Recommended
RT55-01-018 | Audiology | DIAGNOSTIC VIDEONYSTAGMOGRAPHY (VNG):Diagnostic equipment videonystagmography (VNG) Description:Diagnostic assessment equipment, videonystagmography, equipment for the assessment of balance. Technical Specification: To include air calorics and digital light screen or equivalent. The test battery must include saccades, optokinetics, gaze, static positional, dix hallpike, calorics and other specialised tests. Relevant goggles to be supplied, must have ability to record, review, and transfer of information to allow for sharing of information. Compatible PC/Laptop with software to be supplied. Accessories to be included . | Navitrons | Instrumentation Difra | Belgium | Difra | DI 140501: NysStar II Binocular VNG System - Complete System | 2016 | 10 years | 4 (Offer 1) | 1 | 695500 | 83460 | 91806 | 100986.6 | 111085.26 | 122193.786 | 0 | 1205031.646 | RT55-01-018 | 1205031.646 | 625804.5 | 6.6985 | 6.9 | 13.5985 | 5 | Recommended
RT55-01-020 | Audiology | Portable Diagnostic Audiometer with extended high frequencies (tele-audiometry) Description:Portable audiometer used for testing and diagnosing hearing thresholds including suprathresholds level testing that allows for tele-audiometry capabilities. Technical Specification:Two-channel audiometer that must test a range of frequencies from 125-16 000Hz, with intensities of -10 to 110dB, in increments of no less than 5db. Must be capable of performing masked pure-tone and warble-tone air and bone conduction and speech testing. Must come with a with:PC/laptop-integrated or controlled system protective carry case.earphones paired insert earphones bone conductor response button talk-forward built-in or gooseneck microphone. Earphones must have individual audiometers with active noise monitoring technology to allow diagnostic testing without a soundproof booth. Earphones must have individual active noise monitoring technology to allow diagnostic testing without a soundproof booth . Must have a built-in pre-recorded speech material in atleast 1 or more South African accents to be connected via an external port or built-in option for recorded tests Must allow for tele-audiometry. Must allow for automated and manual testing.Must allow for printing capabilities.The ability to store and manage large amounts of patient data Accessories:protective carry case earphones, paired insert earphones,bone conductor ,response button,built-in or gooseneck microphone. Optional accessory.Printer. | Lebone Medical Supplies (Pty) Ltd t/a Amtronix Diagnostics | Inventis S.R.L. | Italy | Inventis | Inventis Cello with Laptop | 2010 - Continious firmware upgrades | 10 Years | 1 | 1 Unit | 152914.695 | 21275 | 23402.5 | 24679 | 25955.5 | 27657.5 | 5000 | 280884.195 | RT55-01-020 | 280884.195 | 280884.195 | 90 | 5.103 | 95.103 | 1 | Recommended
RT55-01-022 | Audiology | Screening Tympanometer (226HZ and 1000HZ): Description:Portable screener for tympanometry and ipsilateral acoustic reflex testing. Technical Specification: Tympanometry measurement capability of 226Hz and 1000Hz.Must test tympanometry across range of +200 -to -400daPa with accuracy of ±10daPa Ipsi-lateral testing range from 500,1000,2000 and 4000Hz. Must include probe/ear tips. Must be supplied with protective carry case .Must have A/C power supply.Must have printing compatibility Must have a memory for at least 12 tests, with probe assembly.Must be supplied with a box of ear tips of at least six different sizes, with at least three tips of each size. Accessories: probe/ear tips of at least six different sizes, with at least three tips of each size protective carry case.two rolls of paper.memory for at least 12 tests probe assembly box of ear tips. | Navitrons | Amplivox | United Kingdom | Amplivox | Otowave 202-H with printer | 2015 | 10 years | 1 | 1 | 91000 | 10920 | 12012 | 13213.2 | 14534.52 | 15987.972 | 4950 | 162617.692 | RT55-01-022 | 162617.692 | 162617.692 | 90 | 0 | 90 | 1 | Recommended
RT55-01-023 | Audiology | Soundproof booth with extraction fan (medium) Description: Soundproof booth with ventilation and extractor fan (ventilation unit for TB protection with minimal noise output). Technical Specification: Soundproof booth with ventilation and extractor fan (with minimal noise output). Should contain observation safety glass window and work platform attachment.The booth should have sound absorption in accordance with ISO 9001, SANS or equivalent. Inside ambient noise levels compliant with the ANSI S.1-1999 or equivalent.Fire resistance in compliance with EN13823 and SANS 10177.Minimum external dimensions of length: 1998mm, width: 1998mm, height 2277mm.Minimal internal dimensions of length: 1783mm, width: 1783mm, height: 1960mm. At least 12 input and output jacks (for air conduction left and right, bone conduction, inserts left and right, speakers left and right, VRA left and right, talk forward, and talk back).Should contain observation safety glass window which is appropriately acoustically sealed and work platform attachment which is appropriately sized to accommodate audiometer, accessories and a writing surface.The booth should have sound absorption in accordance with ISO 9001 or equivalent. The proposed audiology booth is to comply with SANS 10182: 2006 as amended– the measurements and assessment of acoustic environments for audiometric test.The booth must be mounted on anti-vibration mounts to avoid any structural transfer of sound from the main structure.Inside ambient noise levels compliant with the ANSI S.1-1999 or equivalent.Fire resistance in compliance with EN13823 and SANS 10177. Should have adequate silent ventilation system that does not generate sound in the booth or conduct sound through the envelope of the booth resulting in an ambient sound pressure level above the prescribed limit per octave band.Outward opening door that can be opened from inside as well as outside of the booth, is not lockable from the inside, is at least 900 mm wide, and must appropriately acoustically seal so as not to compromise the required ambient sound pressure level inside the booth. Comes standard with wheelchair access ramp that does not impede door closure.Input jacks for connecting equipment on the outside of the booth .Output jacks for connecting transducers inside the booth. Adequate lighting supplied within the booth and clinician test area when appropriate.Electrical fittings and wiring to have standard certified safety measures, including trip switches. All penetrations through the envelope for electrical services must be sealed and must not compromise the ambient sound pressure level inside the booth. On/off switches for lighting and ventilation to be installed in easy to reach location for audiologist.Two standard 3 pin electrical outlet to be installed at suitable location in relation to audio logical equipment. | Lebone Medical Supplies (Pty) Ltd t/a Amtronix Diagnostics | Amtronix Diagnostics | South Africa | Amtronix Duagnostics AX | AX2 Single room booth | 2020 | 10 Years | 1 | 1 Unit | 229425 | 9775 | 10752.5 | 11339 | 11925.5 | 12707.5 | 0 | 285924.5 | RT55-01-023 | 285924.5 | 285924.5 | 90 | 5.103 | 95.103 | 1 | Recommended
RT55-02-002 | Occupational Therapy | COMBINATION PLIERS, LONG NOSE Description: Long nose combination pliers are the same as traditional, but instead of a flat nose, the pliers have a long, tapered nose, which allows you to get into tighter spaces. Technical Specification: Robust and temperature resistant steel.Approximately 180mm-200mm total length Ergonomic handles covered in soft durable rubber/plastic | KUSASALETU INVESTMENT GROUP | BUILT IT | RSA | BUILT IT | 10011349 | 2024 | LIFETIME | 1 | 1 | 180 | 0 | 0 | 0 | 0 | 0 |  | 180 | RT55-02-002 | 180 | 180 | 90 | 10 | 100 | 1 | Recommended
RT55-02-017 | Occupational Therapy | COMBINATION PLIERS, FLAT Description:Pliers are a hand tool used to hold objects firmly, bend and physically compress a wide range of materials. Combination pliers’ dual functionality allows for both cutting and gripping. Technical Specification:These pliers have one scissor-type jaw opening, where half of the length is flat and used for gripping, and the other half is a cutting edge used for cutting wire. Robust and temperature resistant steel.Approximately 180mm-200mm total length. Ergonomic handles covered in soft durable rubber/plastic | KUSASALETU INVESTMENT GROUP | BUILT IT | RSA | BUILT IT | 1101000000 | 2024 | LIFETIME | 1 | 1 | 210 | 0 | 0 | 0 | 0 | 0 |  | 210 | RT55-02-017 | 210 | 210 | 90 | 10 | 100 | 1 | Recommended
RT55-02-030 | Occupational Therapy | HAMMER, CLAW, STEEL Description:Basic hand-held, claw hammer, metal head to drive in or remove nails from wood, shape metal, or to crush rock. Technical Specification:Steel head with claw.Ergonomic, soft grip handle | KUSASALETU INVESTMENT GROUP | BUILT IT | RSA | BUILT IT | 110100025 | 2024 | LIFETIME | 1 | 1 | 284 | 0 | 0 | 0 | 0 | 0 |  | 284 | RT55-02-030 | 284 | 284 | 90 | 10 | 100 | 1 | Recommended
RT55-02-053 | Occupational Therapy | SHIFTING SPANNER (0-30) Description:Also known as an adjustable spanner, its jaws can be expanded or narrowed down to tighten or loose the nut bolts. They are used for those nuts where no other spanner fits well.Technical Specification:These are generally made of drop forge steel or carbon steel. Should allow for 30mm adjustability.Standard: To DIN and ISO | KUSASALETU INVESTMENT GROUP | BUILT IT | RSA | BUILT IT | 10011310 | 2024 | LIFETIME | 1 | 1 | 145 | 0 | 0 | 0 | 0 | 0 |  | 145 | RT55-02-053 | 145 | 145 | 90 | 10 | 100 | 1 | Recommended
RT55-02-083 | Occupational Therapy | Hand Evaluation Kit (electronic) Description: A set of electronically linked instruments used to test grip strength and other components of hand function. Technical Specification:The E-link hand kit comprises a dynamometer and pinch gauge, precision devices that are designed for accuracy, ease and speed of data collection as well as being sensitive to record very small measurements.The Dynamometer accurately registers force applied form 0.1-90 kg, pinchmeter 0.1-22kg in 0.1 increments. Sensitive to detect very small readings. Tests include standard peak and sustained force assessment for Grip and Pinch, RET for Grip E-link Dynamometer is an industry standard design and the software enables peak force test result to be compared to normative data, must be compatible with Microsoft Office Progress reports are easily generated and utilised as definitive outcome measures. Data is readily exported for audit and research purposes.Used to promote therapeutic intervention post evaluation and for monitoring and evaluating therapeutic outcomes. | Sanbonani Holdings (Pty) Ltd | Kinetec SAS | France | GripAble Pro | M921001202 | 2022 | 10 years | 1 | Kit | 46730 | 3800 | 4028 | 4269.68 | 4525.8608 | 4797.4124 | 2800 | 70950.9532 | RT55-02-083 | 70950.9532 | 70950.9532 | 90 | 0 | 90 | 1 | Recommended
RT55-02-084 | Occupational Therapy | HAND EVALUATION KIT (Hydraulic) Description:A set of instruments used to test grip strength and other components of hand function Technical Specification:Kit includes: hydraulic hand dynamometer, pinch gauge, finger goniometer, Sensory evaluator hand kit, Two point Discriminator, Tape measure, finger circumference gauge (cm reading). Hydraulic hand dynamometer: Maximum strength indicator remains after each reading until reset. Registers in pounds and kilograms. Approximate Dimensions: 28cm x 10cm x 20cm. Reliable grip force evaluations with adjustable handles which can be placed in 5 grip positions from 35mm to 87mm in 13mm increments to accommodate the hand size. To include calibration certificate and norm ranges. Hydraulic pinch Gauge: It offers accurate and repeatable pinch strength readings (tip, key and palmer). Maximum strength indicator remains after each reading until reset.Registers in pounds and kilograms. Approximate Dimensions: 13cm x 5cm x 6cm. To include calibration certificate and norm ranges. Finger goniometer:6’ stainless steel finger goniometer. Has 180 degrees scale marked in 1 degree increments that read in either direction. Linear in inches and centimeters. Must be able to measure hyperextension of the finger. (0 - 120° with hyperextension 0 - 30°) Sensory evaluator hand kit.5 piece sensory evaluator kit. Each evaluator is individually calibrated within a 5% standard deviation of a predetermined level of force.The sensory evaluator handles are marked with the corresponding color-coded pressure level. The following colour coded pressure levels must be included: 2.83, 3.61, 4.31, 4.56, 6.65.A small carry case muse be included for storage Two-point Discriminator.Consists of two sturdy, rotating, plastic disks that are joined together Rounded tips are spaced at standard testing intervals from 1 to 15mm apart A 20 and 25mm spacing are also given.One disk setting tests from 1 to 8mm. To change to the 9 to 15mm disk setting, simply rotate the top disk until it clicks in place. Instructions included .Tape Measure.Small, compact measuring tape used to measure the circumference of the digits and hand Measures in inches and centimetres. Must be flexible to wrap around the digit/ hand.Must be retractable.Measures up to a minimum of 1 metre Finger circumference gauge reading.Used to measure the circumference of a finger.Finger circumference gauge webbing loop wraps around the finger and the circumference is read off the ruler.Must be measured in inches and centimetres Accessories: Must include hard carry, moldable case for storage of the 7-piece kit | Mediotronics Physical Medicine Pty Ltd | MVS in Motion | Belgium | Baseline | 8 Piece Hand Evaluation Kit | 2015 | 5 Years | 1 | 1/each | 19950 | 0 | 0 | 0 | 0 | 0 | 0 | 19950 | RT55-02-084 | 19950 | 19950 | 90 | 0 | 90 | 1 | Recommended
RT55-02-084 | Occupational Therapy | HAND EVALUATION KIT (Hydraulic) Description:A set of instruments used to test grip strength and other components of hand function Technical Specification:Kit includes: hydraulic hand dynamometer, pinch gauge, finger goniometer, Sensory evaluator hand kit, Two point Discriminator, Tape measure, finger circumference gauge (cm reading). Hydraulic hand dynamometer: Maximum strength indicator remains after each reading until reset. Registers in pounds and kilograms. Approximate Dimensions: 28cm x 10cm x 20cm. Reliable grip force evaluations with adjustable handles which can be placed in 5 grip positions from 35mm to 87mm in 13mm increments to accommodate the hand size. To include calibration certificate and norm ranges. Hydraulic pinch Gauge: It offers accurate and repeatable pinch strength readings (tip, key and palmer). Maximum strength indicator remains after each reading until reset.Registers in pounds and kilograms. Approximate Dimensions: 13cm x 5cm x 6cm. To include calibration certificate and norm ranges. Finger goniometer:6’ stainless steel finger goniometer. Has 180 degrees scale marked in 1 degree increments that read in either direction. Linear in inches and centimeters. Must be able to measure hyperextension of the finger. (0 - 120° with hyperextension 0 - 30°) Sensory evaluator hand kit.5 piece sensory evaluator kit. Each evaluator is individually calibrated within a 5% standard deviation of a predetermined level of force.The sensory evaluator handles are marked with the corresponding color-coded pressure level. The following colour coded pressure levels must be included: 2.83, 3.61, 4.31, 4.56, 6.65.A small carry case muse be included for storage Two-point Discriminator.Consists of two sturdy, rotating, plastic disks that are joined together Rounded tips are spaced at standard testing intervals from 1 to 15mm apart A 20 and 25mm spacing are also given.One disk setting tests from 1 to 8mm. To change to the 9 to 15mm disk setting, simply rotate the top disk until it clicks in place. Instructions included .Tape Measure.Small, compact measuring tape used to measure the circumference of the digits and hand Measures in inches and centimetres. Must be flexible to wrap around the digit/ hand.Must be retractable.Measures up to a minimum of 1 metre Finger circumference gauge reading.Used to measure the circumference of a finger.Finger circumference gauge webbing loop wraps around the finger and the circumference is read off the ruler.Must be measured in inches and centimetres Accessories: Must include hard carry, moldable case for storage of the 7-piece kit | Duromed cc | North Coast Medical | USA | NCM EXACTA | NC70142-HK +NC12772 +NC12776 +NC12708 +NC70158 | 2000 | 10 YEARS | OFFER 1 | EACH | 26243 | 2000 | 2500 | 2500 | 2500 | 2500 | 2500 | 40743 | RT55-02-084 | 40743 | 19950 | -3.803 | 6.9 | 3.097 | 2 | Recommended
RT55-02-085 | Occupational Therapy | HYDRAULIC HAND DYNAMOMETER Description:An instrument used for assessing grip strength Technical Specification:Offers accurate and repeatable grip strength readings. The unit has an adjustable 5-position handle. Maximum strength indicator remains after each reading until reset. Registers in pounds and kilograms. Approximate Dimensions: 28cm x 10cm x 20cm Features include: Dual scale readout displays isometric grip force from 0 to 90 kg Handle which is easily adjustable to five grip positions from 35mm to 87mm in 13mm increments Accessories:Must include hard plastic moldable carrier case for storage and norm ranges booklet | Mediotronics Physical Medicine Pty Ltd | Saehan Corporation | Korea | Saehan | SH5001 : Hydraulic Hand Dynamometer | 2015 | 5 Years | 1 | 1/each | 6957.5 | 0 | 0 | 0 | 0 | 0 |  | 6957.5 | RT55-02-085 | 6957.5 | 6957.5 | 90 | 0 | 90 | 1 | Recommended
RT55-02-085 | Occupational Therapy | HYDRAULIC HAND DYNAMOMETER Description:An instrument used for assessing grip strength Technical Specification:Offers accurate and repeatable grip strength readings. The unit has an adjustable 5-position handle. Maximum strength indicator remains after each reading until reset. Registers in pounds and kilograms. Approximate Dimensions: 28cm x 10cm x 20cm Features include: Dual scale readout displays isometric grip force from 0 to 90 kg Handle which is easily adjustable to five grip positions from 35mm to 87mm in 13mm increments Accessories:Must include hard plastic moldable carrier case for storage and norm ranges booklet | Duromed cc | North Coast Medical | USA | NCM EXACTA | NC70142 | 2000 | 10 YEARS | OFFER 1 | EACH | 8832 | 2070 | 2530 | 2530 | 2530 | 2530 | 2530 | 23552 | RT55-02-085 | 23552 | 6957.5 | -124.6612 | 6.9 | -117.7612 | 2 | Recommended
RT55-02-086 | Occupational Therapy | HYDRAULIC PINCH GAUGE Description: Instrument used to test grip strength and other components of hand function Technical Specification:It offers accurate and repeatable pinch strength readings (tip, key and palmer). Maximum strength indicator remains after each reading until reset. Registers in pounds and kilograms. Approximate Dimensions: 13cm x 5cm x 6cm Accessories:Must include hard carry case and norm ranges. | Mediotronics Physical Medicine Pty Ltd | Saehan Corporation | Korea | Saehan | SH5005 : Hydraulic Pinch Gauge | 2015 | 5 Years | 1 | 1/each | 5025 | 0 | 0 | 0 | 0 | 0 |  | 5025 | RT55-02-086 | 5025 | 5025 | 90 | 0 | 90 | 1 | Recommended
RT55-02-086 | Occupational Therapy | HYDRAULIC PINCH GAUGE Description: Instrument used to test grip strength and other components of hand function Technical Specification:It offers accurate and repeatable pinch strength readings (tip, key and palmer). Maximum strength indicator remains after each reading until reset. Registers in pounds and kilograms. Approximate Dimensions: 13cm x 5cm x 6cm Accessories:Must include hard carry case and norm ranges. | Duromed cc | North Coast Medical | USA | NCM EXACTA | NC70141 | 2000 | 10 YEARS | OFFER 1 | EACH | 6434.25 | 1725 | 2070 | 2070 | 2070 | 2070 | 2070 | 18509.25 | RT55-02-086 | 18509.25 | 5025 | -151.509 | 6.9 | -144.609 | 2 | Recommended
RT55-02-092 | Occupational Therapy | SPLINT PAN (11-20 LITRES) (With hinged Lid) Description:A device used for molding thermoplastic splinting materials by means of hot water during therapy sessions, generally but not limited to hand therapy. Technical Specification: Concealed heating element with a perforated tray with non-stick liner to prevent splinting material from adhering to pan base. Should not have water circulated or causing a flow current Have built in syphon drainage.Hinged lid with insulated plastic handles for safe opening (if lid is stainless steel). Unit must be made of stainless steel.To indicate maximum boiling temperature Adjustable thermostat controls with thermometer on control panel & indicator light. Must be durable and rust proof.Electrical cord with SA compatible 3 pin plug, 220 -250 Volt, Electrical cord with SABS approved standard.CE marking standard or equivalent . Approximate capacity of 11-20 litres. Minimum internal dimensions 500mmx 300mm x 180mm. Accessories:Spatula:For commissioning purposes, suppliers must supply 10cm x10cm thermoplastic materials. Accessories: Perforated tray :For commissioning purposes, suppliers must supply 10cm x10cm thermoplastic materials . Accessories:Non-stick liner:For commissioning purposes, suppliers must supply 10cm x10cm thermoplastic materials . | Sanbonani Holdings (Pty) Ltd | Kinetec SAS | France | Kinetec SAS | M200000056 | 2021 | 10 years | 1 | Each | 59491.52 | 3200 | 3456 | 3732.48 | 4031.0784 | 4353.5647 |  | 78264.6431 | RT55-02-092 | 78264.6431 | 78264.6431 | 90 | 0 | 90 | 1 | Recommended
RT55-02-093 | Occupational Therapy | SPLINT PAN (11-20 LITRES) (With removable Lid) Description:A device used for molding thermoplastic splinting materials by means of hot water during therapy sessions, generally but not limited to hand therapy. Technical Specification: Concealed heating element with a perforated tray with non-stick liner to prevent splinting material from adhering to pan base. Should not have water circulated or causing a flow current Have built in syphon drainage. Completely removable lid with insulated plastic handles for safe removal, if lid is stainless steel. Unit must be made of stainless steel.To indicate maximum boiling temperature Adjustable thermostat controls with thermometer on control panel & indicator light. Must be durable and rust proof. Electrical cord with SA compatible 3 pin plug, 220 -250 Volt, Electrical cord with SABS approved standard.CE marking standard or equivalent . Approximate capacity of 11-20 litres.Minimum internal dimensions 500mmx 300mm x 180mm. Accessories:Spatula:For commissioning purposes, suppliers must supply 10cm x10cm thermoplastic materials. Accessories:Non-stick liner:For commissioning purposes, suppliers must supply 10cm x10cm thermoplastic materials. Accessories:Perforated Tray:For commissioning purposes, suppliers must supply 10cm x10cm thermoplastic materials. | Sanbonani Holdings (Pty) Ltd | North Coast Medical Inc. | USA | Splint-Form™ | NC14000 | 2014 | 10 years | 1 | Each | 35707.5 | 3200 | 3456 | 3732.48 | 4031.0784 | 4353.5647 |  | 54480.6231 | RT55-02-093 | 54480.6231 | 54480.6231 | 90 | 0 | 90 | 1 | Recommended
RT55-02-094 | Occupational Therapy | SPLINT PAN (21-35 LITRES) (with Hinged Lid) Description:A device used for molding thermoplastic splinting materials by means of hot water during therapy sessions, generally but not limited to hand therapy. Technical Specification: Concealed heating element with a perforated tray with non-stick liner to prevent splinting material from adhering to pan base. Should not have water circulated or causing a flow current Have built in syphon drainage.Hinged lid with insulated plastic handles for safe opening, if lid is stainless steel. Unit must be made of stainless steel.To indicate maximum boiling temperature.Adjustable thermostat controls with thermometer on control panel & indicator light. Must be durable and rust proof.Electrical cord with SA compatible 3 pin plug, 220 -250 Volt, Electrical cord with SABS approved standard.CE marking standard or equivalent Approximately size of 20-35 litres.Minimum internal dimensions 700mm x 500m x 200mm. Accessories:Spatula .For commissioning purposes, suppliers must supply 10cm x10cm thermoplastic materials. Accessories:Perforated Tray.For commissioning purposes, suppliers must supply 10cm x10cm thermoplastic materials. Accessories:Non Stick liner.For commissioning purposes, suppliers must supply 10cm x10cm thermoplastic materials | Sanbonani Holdings (Pty) Ltd | Kinetec SAS | France | Kinetec SAS | M200000033 | 2014 | 10 years | 1 | Each | 51197.4 | 3200 | 3456 | 3732.48 | 4031.0784 | 4353.5647 |  | 69970.5231 | RT55-02-094 | 69970.5231 | 69970.5231 | 90 | 0 | 90 | 1 | Recommended
RT55-02-094 | Occupational Therapy | SPLINT PAN (21-35 LITRES) (with Hinged Lid) Description:A device used for molding thermoplastic splinting materials by means of hot water during therapy sessions, generally but not limited to hand therapy. Technical Specification: Concealed heating element with a perforated tray with non-stick liner to prevent splinting material from adhering to pan base. Should not have water circulated or causing a flow current Have built in syphon drainage.Hinged lid with insulated plastic handles for safe opening, if lid is stainless steel. Unit must be made of stainless steel.To indicate maximum boiling temperature.Adjustable thermostat controls with thermometer on control panel & indicator light. Must be durable and rust proof.Electrical cord with SA compatible 3 pin plug, 220 -250 Volt, Electrical cord with SABS approved standard.CE marking standard or equivalent Approximately size of 20-35 litres.Minimum internal dimensions 700mm x 500m x 200mm. Accessories:Spatula .For commissioning purposes, suppliers must supply 10cm x10cm thermoplastic materials. Accessories:Perforated Tray.For commissioning purposes, suppliers must supply 10cm x10cm thermoplastic materials. Accessories:Non Stick liner.For commissioning purposes, suppliers must supply 10cm x10cm thermoplastic materials | Duromed cc | KINETEC | UK | Manosplint | M200000062 | 2016 | 10 YEARS | OFFER 1 | EACH | 78775 | 2875 | 3450 | 3450 | 3450 | 3450 | 0 | 95450 | RT55-02-094 | 95450 | 69970.5231 | 57.2269 | 6.9 | 64.1269 | 2 | Recommended
RT55-02-095 | Occupational Therapy | SPLINT PAN (21-35 LITRES) (with removable Lid) Description:A device used for molding thermoplastic splinting materials by means of hot water during therapy sessions, generally but not limited to hand therapy. Technical Specification:Concealed heating element with a perforated tray with non-stick liner to prevent splinting material from adhering to pan base.Should not have water circulated or causing a flow current Have built in syphon drainage Completely removable lid with insulated plastic handles for safe removal, if lid is stainless steel. Unit must be made of stainless steel.To indicate maximum boiling temperature Adjustable thermostat controls with thermometer on control panel & indicator light. Must be durable and rust proof.Electrical cord with SA compatible 3 pin plug, 220 -250 Volt, Electrical cord with SABS approved standard. CE marking standard or equivalent Approximately size of 20-35 litres.Minimum internal dimensions 700mm x 500m x 200mm. Accessories:Spatula :For commissioning purposes, suppliers must supply 10cm x10cm thermoplastic materials. Accessories:Perforated tray:For commissioning purposes, suppliers must supply 10cm x10cm thermoplastic materials. Accessories:Non-stick liner:For commissioning purposes, suppliers must supply 10cm x10cm thermoplastic materials . | Sanbonani Holdings (Pty) Ltd | Kinetec SAS | France | Kinetec SAS | M200000034 | 2014 | 10 years | 1 | Each | 45398.38 | 3200 | 3456 | 3732.48 | 4031.0784 | 4353.5647 |  | 64171.5031 | RT55-02-095 | 64171.5031 | 64171.5031 | 90 | 10 | 100 | 1 | Recommended
RT55-02-096 | Occupational Therapy | SPLINT PAN (5-10 LITRES) (splinting pans with hinged lid) Description:A device used for molding thermoplastic splinting materials by means of hot water during therapy sessions, generally but not limited to hand therapy. Technical Specification: Concealed heating element with a perforated tray with non-stick liner to prevent splinting material from adhering to pan base.Should not have water circulated or causing a flow current Have built in syphon drainage Hinged lid with insulated plastic handles for safe opening, if lid is stainless steel. Unit must be made of stainless steel.To indicate maximum boiling temperature.Adjustable thermostat controls with thermometer on control panel & indicator light. Must be durable and rust proof.Electrical cord with SA compatible 3 pin plug, 220 -250 Volt, Electrical cord with SABS approved standard.CE marking standard or equivalent . Approximate capacity of 5-10 litres.Minimum internal dimensions 330mm x 200mmx 100mm. Accessories:Spatula:For commissioning purposes, suppliers must supply 10cm x10cm thermoplastic materials . Accessories: Perforated tray : For commissioning purposes, suppliers must supply 10cm x10cm thermoplastic materials . Accessories: Non Stick liner : For commissioning purposes, suppliers must supply 10cm x10cm thermoplastic materials . | Sanbonani Holdings (Pty) Ltd | Kinetec SAS | France | Kinetec SAS | M200000032 | 2014 | 10 years | 1 | Each | 29503.17 | 3200 | 3456 | 3732.48 | 4031.0784 | 4353.5647 |  | 48276.2931 | RT55-02-096 | 48276.2931 | 48276.2931 | 90 | 10 | 100 | 1 | Recommended
RT55-02-097 | Occupational Therapy | SPLINT PAN (5-10 LITRES) (splinting pans with removable lid) Description:A device used for molding thermoplastic splinting materials by means of hot water during therapy sessions, generally but not limited to hand therapy. Technical Specification:Concealed heating element with a perforated tray with non-stick liner to prevent splinting material from adhering to pan base.Should not have water circulated or causing a flow current Have built in syphon drainage Completely removable lid with insulated plastic handles for safe removal, if lid is stainless steel. Unit must be made of stainless steel.To indicate maximum boiling temperature.Adjustable thermostat controls with thermometer on control panel & indicator light. Must be durable and rust proof.Electrical cord with SA compatible 3 pin plug, 220 -250 Volt, Electrical cord with SABS approved standard.CE marking standard or equivalent .Approximate capacity of 5-10 litres Minimum internal dimensions 330mm x 200mmx 100mm. Accessories:Spatula:For commissioning purposes, suppliers must supply 10cm x10cm thermoplastic materials Accessories:Perforated tray :For commissioning purposes, suppliers must supply 10cm x10cm thermoplastic materials. Accessories: Non Stick liner:For commissioning purposes, suppliers must supply 10cm x10cm thermoplastic materials | Sanbonani Holdings (Pty) Ltd | Kinetec SAS | France | Kinetec SAS | M200000030 | 2014 | 10 years | 1 | Each | 24212.47 | 3200 | 3456 | 3732.48 | 4031.0784 | 4353.5647 |  | 42985.5931 | RT55-02-097 | 42985.5931 | 42985.5931 | 90 | 10 | 100 | #REF! | Recommended
RT55-02-098 | Occupational Therapy | SUSPENSION FRAME, HEIGHT ADJUSTABLE SYSTEM Description: This system is designed for therapeutic use and can securely suspend various therapeutic swings or rehabilitation tools. Ceiling mounted frame, height adjustable that various therapeutic swings or rehabilitation tools can be suspended from. Technical specification:Made of sturdy alloy materials, Minimum weight capacity of 250 kilograms, weight limit clearly indicated on frame Components to include a Swing Rotational Device (for smooth rotation), Safety Connectors, a Therapy Rope (strong nylon), and a Height Adjuster . Ergonomic handgrip on the Height Adjuster ensures therapist comfort. Ceiling bolt to allow approximately 3.5 meter from the center suspension point to walls and other surfaces. Recommended Height: About 2.5 metres from the floor. Accessories: Swing Rotational Device,Safety Connectors, Therapy Rope (strong nylon), Height Adjuster ,Ceiling bolt | New Horizon Metals cc | New Horizon Metals cc | South Africa | NHM | SF001 | 2022/2023 | 7 years | 1 | Each | 27600 | 0 | 0 | 0 | 0 | 0 | 0 | 27600 | RT55-02-098 | 27600 | 27600 | 90 | 0 | 90 | 1 | Recommended
RT55-04-028 | Physiotherapy | Dual Channel TENS (Trans Electrical Neuromuscular Stimulation) and NMS (Neuromuscular Stimulation) Unit Description :A device used for pain relieve and muscle rehabilitation Technical Specification: The unit shall comply with at least the following specifications: Waveform: Asymmetrical, rectangular, biphasic with zero DC current Max Voltage: 180Volts +10/-30V ,Pulse Amplitude: Adjustable from 0-90mA Pulse Width: Adjustable from 50 - 450µS(microseconds) Pulse Rate: Adjustable from 2 - 200Hz .Lock mode function The unit shall have the following features: At least 2 totally independent channels Possible to run to different programs at the same time ,A minimum of 10 present TENS programs A minimum of 2 present EMS programs ,A minimum of 2 customisable programs A large digital display ,Clear backlit LCD display ,Rechargeable Battery operated for ease of portability Must be supplied with a battery charger ,protective carrycase . Accessories: Re-usable electrodes in different sizes should be available to order individually Disposable electrodes in different sizes should be available to order individually Self-adhesive electrodes of different sizes should be available to order independently Dual conductor lead wire for TENS/NMS Optional accessories .Rechargeable batteries included (Lithium ion preferable) Battery charger included. | HITECH THERAPY | VERITY MEDICAL | IRELAND | NEUROTRAC | MULTI-TENS | 2011 | 7 YEARS | 1 | MA | 4025 | 650 | 650 | 650 | 650 | 650 | 0 | 7275 | RT55-04-028 | 7275 | 7275 | 90 | 10 | 100 | 1 | Recommended
RT55-04-029 | Physiotherapy | COMBINATION ELECTROTERAPY UNIT (without Ultrasound) Technical Specification:The unit must offer two (2) channel alternating current , and must be suitable for the commonly used forms of therapy based on INTERFERENTIAL THERAPY, TENS, RUSSIAN STIMULATION, RECTANGULAR AND TRIANGULAR PULSED CURRENTS, IONTOPHORESIS, Diadynamic currents and medium frequency currents. A two channel vacuum unit must be also available and compatible for use with the Interferential Therapy Application. The unit must offer the latest technology and it must be extremely user-friendly.The unit must be fully operable off 220 Volt + 10%, 50Hz single phase Alternating Current (ac) supply. The unit must be fused in BOTH the LIVE and NEUTRAL. Supply the unit with a 3 (three) pin plug.N.B. Units that employ an external power supply will not be considered.The two channels of the unit must provide independent treatment control and the patient current must be continuously displayed on a front panel numerical digital display.The functional design of the unit must be such that the controls are logically arranged on the front CONTROL PANEL, enabling the user to switch easily and quickly from one current to another.Tthe unit must provide a display for indication of all instrument settings selected. Both displays must provide clear viewing under all lighting conditions.All the front panel controls must be clearly labeled and should symbols be used, the user must find these symbols easy to understand.The unit must provide memory storage, thus offering the user the ability to store frequently used settings. The ability to store a minimum of ten (10) complete therapy settings must be possible, with the possibility of calling these programs up at any given time.The following features must be provided on the unit:All types of currents must be provided as follows: 2 (two) Channels, Constant current/Constant voltage with one (1) channel switchable to either, while the other is always Constant Current.Muscle training programs.Timer that must provide user settings from 0 to 60 minutes. INTERFERENTIAL THERAPY must be provided typically and desirably as follows: 4 Pole, 4 Pole with Vector Scan, 2 Pole alternating current, Muscle training programs.Amplitude per channel of 0 – 110Ma.Carrier Frequency of 2.5 and 4kHz.Amplitude Modulating .Frequency adjustable in the range of 0 to 200Hz.Frequency Modulation (spectrum) control over a range of at least 0 to 100Hz.User Selectable Frequency modulation programs.The unit must ensure that electrotherapy to body parts, where the positioning of rubber electrodes is not possible, are easy to treat with the application of vacuum electrodes. The unit must provide continuous and intermittent adjustable suction strength. The unit must also provide quick application of vacuum electrodes on positions where fixation straps cannot be used. A cable connection for fast application of rubber electrodes must be provided.PULSED CURRENTS (TENS): Asymmetrical/Symmetrical biphasic pulse, 2 (two) channel muscle-training programs must be provided typically and desirably as follows:Amplitude per channel of 0 – 140mA. Amplitude Modulating Frequency adjustable in the range of 0 to 200Hz. Frequency Modulation (Spectrum) control over the range of at least 0 to 100Hz.User Selectable Frequency Modulation programs.Intermittent Alternating Current must be available for 2 (two) channel muscle-training programs (Russian Stimulation) typically and desirably as follows: Amplitude per channel of 0–110mA. Burst Frequency must be selectable in the range 0 to 100Hz .User Selectable Burst/Interval ratio. | HITECH THERAPY | Shenzhen Dongdixin Technology Co., Ltd. | CHINA | NU-TEK | MT2201 | 2018 | 7 YEARS | 1 | MA | 45000 | 1200 | 1200 | 1200 | 1200 | 1200 | 2500 | 53500 | RT55-04-029 | 53500 | 53500 | 90 | 10 | 100 | 1 | Recommended
RT55-04-029 | Physiotherapy | COMBINATION ELECTROTERAPY UNIT (without Ultrasound) Technical Specification:The unit must offer two (2) channel alternating current , and must be suitable for the commonly used forms of therapy based on INTERFERENTIAL THERAPY, TENS, RUSSIAN STIMULATION, RECTANGULAR AND TRIANGULAR PULSED CURRENTS, IONTOPHORESIS, Diadynamic currents and medium frequency currents. A two channel vacuum unit must be also available and compatible for use with the Interferential Therapy Application. The unit must offer the latest technology and it must be extremely user-friendly.The unit must be fully operable off 220 Volt + 10%, 50Hz single phase Alternating Current (ac) supply. The unit must be fused in BOTH the LIVE and NEUTRAL. Supply the unit with a 3 (three) pin plug.N.B. Units that employ an external power supply will not be considered.The two channels of the unit must provide independent treatment control and the patient current must be continuously displayed on a front panel numerical digital display.The functional design of the unit must be such that the controls are logically arranged on the front CONTROL PANEL, enabling the user to switch easily and quickly from one current to another.Tthe unit must provide a display for indication of all instrument settings selected. Both displays must provide clear viewing under all lighting conditions.All the front panel controls must be clearly labeled and should symbols be used, the user must find these symbols easy to understand.The unit must provide memory storage, thus offering the user the ability to store frequently used settings. The ability to store a minimum of ten (10) complete therapy settings must be possible, with the possibility of calling these programs up at any given time.The following features must be provided on the unit:All types of currents must be provided as follows: 2 (two) Channels, Constant current/Constant voltage with one (1) channel switchable to either, while the other is always Constant Current.Muscle training programs.Timer that must provide user settings from 0 to 60 minutes. INTERFERENTIAL THERAPY must be provided typically and desirably as follows: 4 Pole, 4 Pole with Vector Scan, 2 Pole alternating current, Muscle training programs.Amplitude per channel of 0 – 110Ma.Carrier Frequency of 2.5 and 4kHz.Amplitude Modulating .Frequency adjustable in the range of 0 to 200Hz.Frequency Modulation (spectrum) control over a range of at least 0 to 100Hz.User Selectable Frequency modulation programs.The unit must ensure that electrotherapy to body parts, where the positioning of rubber electrodes is not possible, are easy to treat with the application of vacuum electrodes. The unit must provide continuous and intermittent adjustable suction strength. The unit must also provide quick application of vacuum electrodes on positions where fixation straps cannot be used. A cable connection for fast application of rubber electrodes must be provided.PULSED CURRENTS (TENS): Asymmetrical/Symmetrical biphasic pulse, 2 (two) channel muscle-training programs must be provided typically and desirably as follows:Amplitude per channel of 0 – 140mA. Amplitude Modulating Frequency adjustable in the range of 0 to 200Hz. Frequency Modulation (Spectrum) control over the range of at least 0 to 100Hz.User Selectable Frequency Modulation programs.Intermittent Alternating Current must be available for 2 (two) channel muscle-training programs (Russian Stimulation) typically and desirably as follows: Amplitude per channel of 0–110mA. Burst Frequency must be selectable in the range 0 to 100Hz .User Selectable Burst/Interval ratio. | Sanbonani Holdings (Pty) Ltd | BTL Industries Limited / BTL Medical SA (Pty) Ltd | United Kingdom | BTL-4620 Premium | P2620.701 | 2007 | 10 years | 2 | Set | 107019 | 3800 | 4104 | 4435 | 4790 | 5170 | 3800 | 133118 | RT55-04-029 | 133118 | 53500 | -43.9368 | 8.66 | -35.2768 | 2 | Recommended
RT55-04-030 | Physiotherapy | Combination Electrotherapy and Ultrasound Unit .Description :A multimodal electrotherapy unit used for pain relief, muscle stimulation and to promote healing Technical Specification :The unit shall offer ultrasound and electrotherapy combination therapy . The combination electrotherapy unit shall include 3 modalities: electrotherapy and, ultrasound modalities. Electrotherapy modality shall provide for at least the following current types: Interferential, TENS, Russian stimulation, diadynamic and galvanic current. The unit shall have at least two independent electrotherapy channels .Intensity: 0-3/cm2 (ultrasound); 0-50mA (IF);0-100MA (LF). Ultrasound output: 1&3 MHz frequencies to be available .Ultrasound head: 2cm2 and 5cm22 . Pulse Width: 30-400 microseconds .Pulse rate-carrier frequency: 2KHz-10KHz for IFC; 2.5KHz for Russian; 0Hz-250Hz for TENS and EMS .Pulse rate-beat frequency: 1Hz-200Hz for IFC; 20Hz-100Hz for Russian; 0Hz-250Hz for TENS and EMS. .The unit shall be equipped with body part navigation feature. The unit should include preprogrammed as well as customizable treatment options. The unit shall be equipped with a patient database . The unit shall have the option of 2 or 4 pole interference .The unit shall have a built in vacuum unit . The unit shall be battery operated for ease of portability .All accessories necessary to perform all the functionalities of the unit shall be included in the offer . Accessories: Ultrasound head 5cm ,Ultrasound head 1cm ,Vacuum electrodes ,Hands free sono applicator .Rubber electrodes ,Self-Adhesive electrodes,Russian stimulator,5cm2 ultrasound head ,1cm2 ultrasound head Adapter 100-240V; 47-64Hz .Power cable .User manual .Self adhesive electrodes in various sizes (50x100mm; 50x50mm) Carbon electrodes in various sizes (70x110mm; 70x100mm).Envelope sponges to fit carbon electrodes. Skin leads wires .Straps 75x1200mm and 75x600mm .Combo lead wires .Optional extras.Vacuum unit compatible with the supplied unit. C60 Optional Accessories: 2 tier trolley . The unit shall operate from 220V and the power cords shall have a 3-point plug approved by SABS | HITECH THERAPY | Shenzhen Dongdixin Technology Co., Ltd. | CHINA | NU-TEK | CT2200 | 2018 | 7 YEARS | 1 | WATT | 69000 | 1200 | 1200 | 1200 | 1200 | 1200 | 2500 | 77500 | RT55-04-030 | 77500 | 77500 | 90 | 8.66 | 98.66 | 1 | Recommended
RT55-04-030 | Physiotherapy | Combination Electrotherapy and Ultrasound Unit .Description :A multimodal electrotherapy unit used for pain relief, muscle stimulation and to promote healing Technical Specification :The unit shall offer ultrasound and electrotherapy combination therapy . The combination electrotherapy unit shall include 3 modalities: electrotherapy and, ultrasound modalities. Electrotherapy modality shall provide for at least the following current types: Interferential, TENS, Russian stimulation, diadynamic and galvanic current. The unit shall have at least two independent electrotherapy channels .Intensity: 0-3/cm2 (ultrasound); 0-50mA (IF);0-100MA (LF). Ultrasound output: 1&3 MHz frequencies to be available .Ultrasound head: 2cm2 and 5cm22 . Pulse Width: 30-400 microseconds .Pulse rate-carrier frequency: 2KHz-10KHz for IFC; 2.5KHz for Russian; 0Hz-250Hz for TENS and EMS .Pulse rate-beat frequency: 1Hz-200Hz for IFC; 20Hz-100Hz for Russian; 0Hz-250Hz for TENS and EMS. .The unit shall be equipped with body part navigation feature. The unit should include preprogrammed as well as customizable treatment options. The unit shall be equipped with a patient database . The unit shall have the option of 2 or 4 pole interference .The unit shall have a built in vacuum unit . The unit shall be battery operated for ease of portability .All accessories necessary to perform all the functionalities of the unit shall be included in the offer . Accessories: Ultrasound head 5cm ,Ultrasound head 1cm ,Vacuum electrodes ,Hands free sono applicator .Rubber electrodes ,Self-Adhesive electrodes,Russian stimulator,5cm2 ultrasound head ,1cm2 ultrasound head Adapter 100-240V; 47-64Hz .Power cable .User manual .Self adhesive electrodes in various sizes (50x100mm; 50x50mm) Carbon electrodes in various sizes (70x110mm; 70x100mm).Envelope sponges to fit carbon electrodes. Skin leads wires .Straps 75x1200mm and 75x600mm .Combo lead wires .Optional extras.Vacuum unit compatible with the supplied unit. C60 Optional Accessories: 2 tier trolley . The unit shall operate from 220V and the power cords shall have a 3-point plug approved by SABS | Sanbonani Holdings (Pty) Ltd | BTL Industries Limited / BTL Medical SA (Pty) Ltd | United Kingdom | BTL-4820S Premium | P2820.771 | 2007 | 10 years | 2 | Set | 102397 | 3800 | 4104 | 4435 | 4790 | 5170 | 3800 | 128496 | RT55-04-030 | 128496 | 77500 | 30.7788 | 8.66 | 39.4388 | 2 | Recommended
RT55-04-031 | Physiotherapy | Electronic Height Adjustable Plinth 2 Section Description :An electrically operated 2 section unit used to position patients during physical rehabilitation procedure . Technical Specification :The unit shall have an electric high-low height adjustment with high position ≥90cm and low position ≤45cm .Height adjustment via a remote control and foot control unit .The unit shall have a minimum load capacity of 200kg .Overall length shall be a minimum of 180cm .Overall width shall be a minimum of 60cm (Optional tapering of head and foot piece to a minimum of 50cm) .The frame shall be high grade and durable .The frame shall have a ground clearance of ±15cm to allow for use of patient transfers units such as hoists .The unit shall have heavy duty castors .easily accessible foot operated brakes for easy movement and positioning of the unit . The castors shall be fully retractable .The feet shall be heavy duty and robust with heavy duty non scratch nonslip ferrules.The unit shall be upholstered in durable hygienic, non-slip and washable (e.g. vinyl, or imitation leather) Upholstery high density foam shall be at least 4-6cm thick. The upholstery shall be free of phthalates and shall be flame retardant . The unit shall comprise of two sections that are adjustable at one point .All sections to be independently adjustable either manually with hydraulic controls, electronically or a combination of manual and electronic controls .If hydraulically adjusted, All adjustment handles shall be durable and made of stainless steel or equivalent quality material . Electronic controls shall be via a bed mounted control box.The unit shall accommodate lying and sitting position. The head section shall have a comfortable breathing slot with a plug .The head section shall have backrest and angle shall allow tilt from at least -30° to +90° . Accessories: replacement castors available as spares should be available as a set or individually.The unit shall operate from 220V and the power cords shall have a 3-point plug approved by SABS standard manufacturers length (extra length built in power cord should be available on request) (not an extension lead). | HITECH THERAPY | HUANGHSHAN JINFU MEDICAL EQUIPMENT | CHINA | COINFYCARE | EL02R | 2019 | 7 YEARS | 1 | N/A | 28175 | 950 | 950 | 950 | 950 | 950 | 0 | 32925 | RT55-04-031 | 32925 | 32925 | 90 | 8.66 | 98.66 | 1 | Recommended
RT55-04-031 | Physiotherapy | Electronic Height Adjustable Plinth 2 Section Description :An electrically operated 2 section unit used to position patients during physical rehabilitation procedure . Technical Specification :The unit shall have an electric high-low height adjustment with high position ≥90cm and low position ≤45cm .Height adjustment via a remote control and foot control unit .The unit shall have a minimum load capacity of 200kg .Overall length shall be a minimum of 180cm .Overall width shall be a minimum of 60cm (Optional tapering of head and foot piece to a minimum of 50cm) .The frame shall be high grade and durable .The frame shall have a ground clearance of ±15cm to allow for use of patient transfers units such as hoists .The unit shall have heavy duty castors .easily accessible foot operated brakes for easy movement and positioning of the unit . The castors shall be fully retractable .The feet shall be heavy duty and robust with heavy duty non scratch nonslip ferrules.The unit shall be upholstered in durable hygienic, non-slip and washable (e.g. vinyl, or imitation leather) Upholstery high density foam shall be at least 4-6cm thick. The upholstery shall be free of phthalates and shall be flame retardant . The unit shall comprise of two sections that are adjustable at one point .All sections to be independently adjustable either manually with hydraulic controls, electronically or a combination of manual and electronic controls .If hydraulically adjusted, All adjustment handles shall be durable and made of stainless steel or equivalent quality material . Electronic controls shall be via a bed mounted control box.The unit shall accommodate lying and sitting position. The head section shall have a comfortable breathing slot with a plug .The head section shall have backrest and angle shall allow tilt from at least -30° to +90° . Accessories: replacement castors available as spares should be available as a set or individually.The unit shall operate from 220V and the power cords shall have a 3-point plug approved by SABS standard manufacturers length (extra length built in power cord should be available on request) (not an extension lead). | New Horizon Metals cc | New Horizon Metals cc | South Africa | NHM | 2SEP001 | 2022/2023 | 7 years | 1 | Each | 20870.2 | 4174.04 | 4507.96 | 4868.6 | 5258.09 | 5678.74 | 0 | 45357.63 | RT55-04-031 | 45357.63 | 32925 | 56.0156 | 0 | 56.0156 | 2 | Recommended
RT55-04-031 | Physiotherapy | Electronic Height Adjustable Plinth 2 Section Description :An electrically operated 2 section unit used to position patients during physical rehabilitation procedure . Technical Specification :The unit shall have an electric high-low height adjustment with high position ≥90cm and low position ≤45cm .Height adjustment via a remote control and foot control unit .The unit shall have a minimum load capacity of 200kg .Overall length shall be a minimum of 180cm .Overall width shall be a minimum of 60cm (Optional tapering of head and foot piece to a minimum of 50cm) .The frame shall be high grade and durable .The frame shall have a ground clearance of ±15cm to allow for use of patient transfers units such as hoists .The unit shall have heavy duty castors .easily accessible foot operated brakes for easy movement and positioning of the unit . The castors shall be fully retractable .The feet shall be heavy duty and robust with heavy duty non scratch nonslip ferrules.The unit shall be upholstered in durable hygienic, non-slip and washable (e.g. vinyl, or imitation leather) Upholstery high density foam shall be at least 4-6cm thick. The upholstery shall be free of phthalates and shall be flame retardant . The unit shall comprise of two sections that are adjustable at one point .All sections to be independently adjustable either manually with hydraulic controls, electronically or a combination of manual and electronic controls .If hydraulically adjusted, All adjustment handles shall be durable and made of stainless steel or equivalent quality material . Electronic controls shall be via a bed mounted control box.The unit shall accommodate lying and sitting position. The head section shall have a comfortable breathing slot with a plug .The head section shall have backrest and angle shall allow tilt from at least -30° to +90° . Accessories: replacement castors available as spares should be available as a set or individually.The unit shall operate from 220V and the power cords shall have a 3-point plug approved by SABS standard manufacturers length (extra length built in power cord should be available on request) (not an extension lead). | Arjo Huntliegh South Africa | Seers Medical | UK | Therapy 2 section | RT55-04-031 | 2008 | 15 | 1 | EA | 33797.21 | 2977.47 | 2977.47 | 2977.47 | 2977.47 | 2977.47 |  | 48684.56 | RT55-04-031 | 48684.56 | 32925 | 46.9215 | 8.66 | 55.5815 | 3 | Recommended
RT55-04-032 | Physiotherapy | Electronic Height Adjustable Plinth 3 Section Description :An electrical operated 3 section unit used to position patients during the physical rehabilitation procedure . Technical Specification :The unit shall have an electric high-low height adjustment with high position ≥90cm and low position ≤45cm .Height adjustment via a remote control and foot control unit .The unit shall have a minimum load capacity of 200kg .Overall length shall be a minimum of 180cm .Overall width shall be a minimum of 60cm (Optional tapering of head and foot piece to a minimum of 50cm) . The frame shall be high grade and durable .The frame shall have a ground clearance of ±15cm to allow for use of patient transfers units such as hoists . The unit shall have heavy duty castors ,easily accessible foot operated brakes for easy movement and positioning of the unit .The castors shall be fully retractable . The feet shall be heavy duty and robust with heavy duty non scratch nonslip ferrules.The unit shall be upholstered in durable hygienic, non-slip and washable (e.g. vinyl, or imitation leather) Upholstery high density foam shall be at least 4-6cm thick. The upholstery shall be free of phthalates and shall be flame retardant . The unit shall comprise of three sections that are adjustable at two points .A ll sections to be independently adjustable either manually with hydraulic controls, electronically or a combination of manual and electronic controls .If hydraulically adjusted, All adjustment handles shall be durable and made of stainless steel or equivalent quality material . Electronic controls shall be via a bed mounted control box.The unit shall accommodate lying and sitting position. The head section shall have a comfortable breathing slot with a plug .The head section shall have backrest and angle shall allow tilt from at least -30° to +90° .Accessories :replacement castors available as spares should be available as a set or individually.The unit shall operate from 220V and the power cords shall have a 3-point plug approved by SABS standard manufacturers length (extra length built in power cord should be available on request) (not an extension lead) | New Horizon Metals cc | New Horizon Metals cc | South Africa | NHM | 3SEP002 | 2022/2023 | 7 years | 1 | Each | 21321 | 4264.2 | 4605.34 | 4973.76 | 5371.66 | 5801.4 | 0 | 46337.36 | RT55-04-032 | 46337.36 | 46337.36 | 90 | 0 | 90 | #REF! | Recommended
RT55-04-033 | Physiotherapy | Electronic Height Adjustable Plinth 5 Section Description :An electrically operated 5 section unit used to position patients during the physical rehabilitation procedure Technical Specification:The unit shall have an electric high-low height adjustment with high position ≥90cm and low position ≤45cm .Height adjustment via a remote control and foot control unit .The unit shall have a minimum load capacity of 200kg .Overall length shall be a minimum of 180cm .Overall width shall be a minimum of 60cm (Optional tapering of head and foot piece to a minimum of 50cm) The frame shall be high grade and durable .The frame shall have a ground clearance of ±15cm to allow for use of patient transfers units such as hoists . The unit shall have heavy duty castors ,easily accessible foot operated brakes for easy movement and positioning of the unit .The castors shall be fully retractable. The feet shall be heavy duty and robust with heavy duty non scratch nonslip ferrules.The unit shall be upholstered in durable hygienic, non-slip and washable (e.g. vinyl, or imitation leather) .Upholstery high density foam shall be at least 4-6cm thick. The upholstery shall be free of phthalates and shall be flame retardant. The unit shall comprise of five sections that are adjustable at four points (head and lumbar and two arm rests) . All sections to be independently adjustable either manually with hydraulic controls, electronically or a combination of manual and electronic controls. If hydraulically adjusted, All adjustment handles shall be durable and made of stainless steel or equivalent quality material .Electronic controls shall be via a bed mounted control box.The unit shall accommodate lying and sitting position.The head section shall have a comfortable breathing slot with a plug .The head section shall have backrest and angle shall allow tilt from at least -30° to +90° . Accessories :replacement castors available as spares should be available as a set or individually. The unit shall operate from 220V and the power cords shall have a 3-point plug approved by SABS standard manufacturers length (extra length built in power cord should be available on request) (not an extension lead). | New Horizon Metals cc | New Horizon Metals cc | South Africa | NHM | 5SEP003 | 2022/2023 | 7 years | 1 | Each | 21913.25 | 4382.65 | 4733.26 | 5111.92 | 5520.88 | 5962.55 | 0 | 47624.51 | RT55-04-033 | 47624.51 | 47624.51 | 90 | 0 | 90 | 1 | Recommended
RT55-04-034 | Physiotherapy | LASER Description :Electrotherapy modality utilizing laser to treat Musculo skeletal conditions Specifications:Easy to transport Handpiece power : MLS power of 1.1Watt – 25 Watt Peak power Lithium Chargeable battery operated Backlit LCD color graphic touch screen display & 2 cm beam area 4 treatment modalities : specific pathologies , pain management, anti-edema, bio stimulation Simultaneous management of two independent channels Able to customize each treatment parameters Synchronized emission of continuous pulsed waves and frequented pulsed waves Modulation frequency from 1 to 2000Hz with 1Hz steps Power levels : 25%,50%,75% and 100%.Fixed duty cycle 50% and variable. Treatment durations : 1min – 60 min.Laser emission warning light, Interlock& Acoustic signal.36cm x 28cm x 10 cm 3kg weight.Qualified technician to attend to problems onsite within 48-72hours | Mediotronics Physical Medicine Pty Ltd | Astar | Poland | Astar | ASTAR-PGLA : PhysioGo Lite Laser with Battery |  | 10 Years | 1 | 1 Each | 68050.5 | 2875 | 2875 | 2875 | 2875 | 2875 |  | 82425.5 | RT55-04-034 | 82425.5 | 82425.5 | 90 | 0 | 90 | 1 | Recommended
RT55-04-034 | Physiotherapy | LASER Description :Electrotherapy modality utilizing laser to treat Musculo skeletal conditions Specifications:Easy to transport Handpiece power : MLS power of 1.1Watt – 25 Watt Peak power Lithium Chargeable battery operated Backlit LCD color graphic touch screen display & 2 cm beam area 4 treatment modalities : specific pathologies , pain management, anti-edema, bio stimulation Simultaneous management of two independent channels Able to customize each treatment parameters Synchronized emission of continuous pulsed waves and frequented pulsed waves Modulation frequency from 1 to 2000Hz with 1Hz steps Power levels : 25%,50%,75% and 100%.Fixed duty cycle 50% and variable. Treatment durations : 1min – 60 min.Laser emission warning light, Interlock& Acoustic signal.36cm x 28cm x 10 cm 3kg weight.Qualified technician to attend to problems onsite within 48-72hours | Mediotronics Physical Medicine Pty Ltd | Enraf-Nonius | Netherlands/Germany | Enraf-Nonius | 1633-901 : Endolaser 120 1,2 Watts Output |  | 10 Years | 2 | 1 Each | 120750 | 2875 | 2875 | 2875 | 2875 | 2875 |  | 135125 | RT55-04-034 | 135125 | 82425.5 | 32.4577 | 0 | 32.4577 | 2 | Recommended
RT55-04-034 | Physiotherapy | LASER Description :Electrotherapy modality utilizing laser to treat Musculo skeletal conditions Specifications:Easy to transport Handpiece power : MLS power of 1.1Watt – 25 Watt Peak power Lithium Chargeable battery operated Backlit LCD color graphic touch screen display & 2 cm beam area 4 treatment modalities : specific pathologies , pain management, anti-edema, bio stimulation Simultaneous management of two independent channels Able to customize each treatment parameters Synchronized emission of continuous pulsed waves and frequented pulsed waves Modulation frequency from 1 to 2000Hz with 1Hz steps Power levels : 25%,50%,75% and 100%.Fixed duty cycle 50% and variable. Treatment durations : 1min – 60 min.Laser emission warning light, Interlock& Acoustic signal.36cm x 28cm x 10 cm 3kg weight.Qualified technician to attend to problems onsite within 48-72hours | HITECH THERAPY | ASA SRL | ITALY | MPHI | MPHI | 2015 | 7 YEARS | 1 | WATT | 179975 | 950 | 950 | 950 | 950 | 950 | 2500 | 187225 | RT55-04-034 | 187225 | 82425.5 | -24.4301 | 8.66 | -15.7701 | 3 | Recommended
RT55-04-035 | Physiotherapy | Neurology (wide Bobath) 2 Section Plinth-Electrical Height Adjustable Description :A wide electrical operated 2 section unit used to position patients during active physical rehabilitation procedure . Technical Specification :The unit shall have an electric high-low height adjustment with high position ≥90cm and low position ≤45cm .Height adjustment via a remote control and foot control unit .The unit shall have a minimum load capacity of 250kg .Overall length shall be a minimum of 180cm .Overall width shall be a minimum of 120cm .The frame shall be high grade and durable metal.The frame shall have a ground clearance of ±15cm to allow for use of patient transfers units such as hoists .The unit shall have heavy duty castors ,easily accessible foot operated brakes for easy movement and positioning of the unit .The castors shall be fully retractable .The feet shall be heavy duty and robust with heavy duty non scratch nonslip ferrules.The unit shall be upholstered in durable hygienic, non-slip and washable (e.g. vinyl, or imitation leather) .Upholstery high density foam shall be at least 4-6cm thick. The upholstery shall be free of phthalates and shall be flame retardant.The unit shall comprise of five sections that are adjustable at four points (head and lumbar and two arm rests) .All sections to be independently adjustable electronically or a combination of manual and electronic controls .Electronic controls shall be via a bed mounted control box.The unit shall accommodate lying and sitting position.The head section shall have a comfortable breathing slot with a plug . The head section shall have backrest and angle shall allow tilt from at least -30° to +90° Accessories :replacement castors available as spares should be available as a set or individually The unit shall operate from 220V and the power cords shall have a 3-point plug approved by SABS standard manufacturers length (extra length built in power cord should be available on request) (not an extension lead) | New Horizon Metals cc | New Horizon Metals cc | South Africa | NHM | BBPEW001 | 2022/2023 | 7 years | 1 | Each | 26059 | 5211.8 | 5628.74 | 6079.04 | 6565.37 | 7090.6 | 0 | 56634.55 | RT55-04-035 | 56634.55 | 56634.55 | 90 | 0 | 90 | #REF! | Recommended
RT55-04-035 | Physiotherapy | Neurology (wide Bobath) 2 Section Plinth-Electrical Height Adjustable Description :A wide electrical operated 2 section unit used to position patients during active physical rehabilitation procedure . Technical Specification :The unit shall have an electric high-low height adjustment with high position ≥90cm and low position ≤45cm .Height adjustment via a remote control and foot control unit .The unit shall have a minimum load capacity of 250kg .Overall length shall be a minimum of 180cm .Overall width shall be a minimum of 120cm .The frame shall be high grade and durable metal.The frame shall have a ground clearance of ±15cm to allow for use of patient transfers units such as hoists .The unit shall have heavy duty castors ,easily accessible foot operated brakes for easy movement and positioning of the unit .The castors shall be fully retractable .The feet shall be heavy duty and robust with heavy duty non scratch nonslip ferrules.The unit shall be upholstered in durable hygienic, non-slip and washable (e.g. vinyl, or imitation leather) .Upholstery high density foam shall be at least 4-6cm thick. The upholstery shall be free of phthalates and shall be flame retardant.The unit shall comprise of five sections that are adjustable at four points (head and lumbar and two arm rests) .All sections to be independently adjustable electronically or a combination of manual and electronic controls .Electronic controls shall be via a bed mounted control box.The unit shall accommodate lying and sitting position.The head section shall have a comfortable breathing slot with a plug . The head section shall have backrest and angle shall allow tilt from at least -30° to +90° Accessories :replacement castors available as spares should be available as a set or individually The unit shall operate from 220V and the power cords shall have a 3-point plug approved by SABS standard manufacturers length (extra length built in power cord should be available on request) (not an extension lead) | Arjo Huntliegh South Africa | Seers Medical | UK | Therapy Bobath Electric | RT55-04-035 | 2012 | 15 | 1 | EA | 41070.36 | 4652.3 | 4652.3 | 4652.3 | 4652.3 | 4652.3 |  | 64331.86 | RT55-04-035 | 64331.86 | 56634.55 | 77.7679 | 0 | 77.7679 | #REF! | Recommended
RT55-04-036 | Physiotherapy | Neurology (wide Bobath) Plinth-Hydraulic Description :A wide hydraulic operated unit used to position patients during active physical rehabilitation procedure . Technical Specification :The unit shall have a hydraulic lifting system that is manually operated The unit shall have an electric high-low height adjustment with high position ≥90cm and low position ≤45cm Height adjustment foot control unit .The unit shall have a minimum load capacity of atleast 250kg Overall length shall be a minimum of 180cm.Overall width shall be a minimum of 120cm The frame shall be high grade and durable metal.The frame shall have a ground clearance of ±15cm to allow for use of patient transfers units such as hoists . The unit shall have heavy duty castors .easily accessible foot operated brakes for easy movement and positioning of the unit .The castors shall be fully retractable . The feet shall be heavy duty and robust with heavy duty non scratch nonslip ferrules.The unit shall be upholstered in durable hygienic, non-slip and washable (e.g. vinyl, or imitation leather) .Upholstery high density foam shall be at least 4-6cm thick. The upholstery shall be free of phthalates and shall be flame retardant . The unit shall comprise of five sections that are adjustable at four points (head and lumbar and two arm rests) . All sections to be independently adjustable hydraulically.The unit shall accommodate lying and sitting position. The head section shall have a comfortable breathing slot with a plug .The head section shall have backrest and angle shall allow tilt from at least -30° to +90 Accessories: replacement castors available as spares should be available as a set or individually The unit shall operate from 220V and the power cords shall have a 3-point plug approved by SABS standard manufacturers length (extra length built in power cord should be available on request) (not an extension lead) | Arjo Huntliegh South Africa | Seers Medical | UK | Therapy Bobath Hydraulic | RT55-04-036 | 2012 | 15 | 1 | EA | 41076.28 | 3108.66 | 3108.66 | 3108.66 | 3108.66 | 3108.66 |  | 56619.58 | RT55-04-036 | 56619.58 | 56619.58 | 90 | 0 | 90 | 1 | Recommended
RT55-04-036 | Physiotherapy | Neurology (wide Bobath) Plinth-Hydraulic Description :A wide hydraulic operated unit used to position patients during active physical rehabilitation procedure . Technical Specification :The unit shall have a hydraulic lifting system that is manually operated The unit shall have an electric high-low height adjustment with high position ≥90cm and low position ≤45cm Height adjustment foot control unit .The unit shall have a minimum load capacity of atleast 250kg Overall length shall be a minimum of 180cm.Overall width shall be a minimum of 120cm The frame shall be high grade and durable metal.The frame shall have a ground clearance of ±15cm to allow for use of patient transfers units such as hoists . The unit shall have heavy duty castors .easily accessible foot operated brakes for easy movement and positioning of the unit .The castors shall be fully retractable . The feet shall be heavy duty and robust with heavy duty non scratch nonslip ferrules.The unit shall be upholstered in durable hygienic, non-slip and washable (e.g. vinyl, or imitation leather) .Upholstery high density foam shall be at least 4-6cm thick. The upholstery shall be free of phthalates and shall be flame retardant . The unit shall comprise of five sections that are adjustable at four points (head and lumbar and two arm rests) . All sections to be independently adjustable hydraulically.The unit shall accommodate lying and sitting position. The head section shall have a comfortable breathing slot with a plug .The head section shall have backrest and angle shall allow tilt from at least -30° to +90 Accessories: replacement castors available as spares should be available as a set or individually The unit shall operate from 220V and the power cords shall have a 3-point plug approved by SABS standard manufacturers length (extra length built in power cord should be available on request) (not an extension lead) | New Horizon Metals cc | New Horizon Metals cc | South Africa | NHM | BBPHW002 | 2022/2023 | 7 years | 1 | Each | 27243.5 | 5448.7 | 5884.6 | 6355.36 | 6863.79 | 7412.9 | 0 | 59208.85 | RT55-04-036 | 59208.85 | 56619.58 | 85.8842 | 0 | 85.8842 | 2 | Recommended
RT55-04-037 | Physiotherapy | Oil free compressor Description:Oil free compressor which works together with shockwave Specifications:Low noise level (30dB),100% oil free,5L receiver with internal coating 8 bar output,12.5 L/min air fill rate.Clean air 5um air filter with automatic drain,Low maintenance Constant pressure.Plug and play.Designed to run at 100% duty cycle,portable. | HITECH THERAPY | EMS ELECTRO | SWITZERLAND | EMS | DW034 | 2012 | 7 YEARS | 1 | BAR | 43000 | 950 | 950 | 950 | 950 | 950 | 2500 | 50250 | RT55-04-037 | 50250 | 50250 | 90 | 0 | 90 | 1 | Recommended
RT55-04-038 | Physiotherapy | Shockwave : Longest Radial Shockwave Unit Description :Electrotherapy modality used to treat musculoskeletal conditions Specifications:Bars/mJ : 1.0 – 4.0 bar (adjustable by 0.1 bar) Max Frequency: 1 – 17 Hz (adjustable by 1 Hz) Pulse: 1-3000 (step by 100 pulses) Ergonomic handpiece with impact sleeve to reduce strain on the practitioner Shock counter.Body Map indications menu with preset treatment protocols 11 pre-set protocols + add up to 5,000 of your own Fully customizable treatment parameters with simple adjustments What is included : PowerShocker Handpiece x 1 15mm Radial transmitter x 1 20mm Radial transmitter x 1 Projectile x 2 Metal bullet x 2 Vibration Handpiece x 1 25mm Vibration transmitter x 1 Cleaning Brush x 1 Reversion box x 1 Built in Compressor | HITECH THERAPY | Guangzhou Longest Medical Technology Co., Ltd | CHINA | LONGEST | POWERSHOCKER LGT2500S | 2013 | 7-10 YEARS | 1 | BAR | 175000 | 15000 | 15000 | 15000 | 15000 | 15000 | 2500 | 252500 | RT55-04-038 | 252500 | 252500 | 90 | 0 | 90 | 1 | Recommended
RT55-04-038 | Physiotherapy | Shockwave : Longest Radial Shockwave Unit Description :Electrotherapy modality used to treat musculoskeletal conditions Specifications:Bars/mJ : 1.0 – 4.0 bar (adjustable by 0.1 bar) Max Frequency: 1 – 17 Hz (adjustable by 1 Hz) Pulse: 1-3000 (step by 100 pulses) Ergonomic handpiece with impact sleeve to reduce strain on the practitioner Shock counter.Body Map indications menu with preset treatment protocols 11 pre-set protocols + add up to 5,000 of your own Fully customizable treatment parameters with simple adjustments What is included : PowerShocker Handpiece x 1 15mm Radial transmitter x 1 20mm Radial transmitter x 1 Projectile x 2 Metal bullet x 2 Vibration Handpiece x 1 25mm Vibration transmitter x 1 Cleaning Brush x 1 Reversion box x 1 Built in Compressor | Sanbonani Holdings (Pty) Ltd | BTL Industries Limited / BTL Medical SA (Pty) Ltd | United Kingdom | BTL-6000 RSWT | 844-POEAEA-M | 2007 | 10 years | 1 | Set | 283486.5 | 3800 | 4104 | 4435 | 4790 | 5170 | 4500 | 310285.5 | RT55-04-038 | 310285.5 | 252500 | 69.4032 | 0 | 69.4032 | 2 | Recommended
RT55-04-039 | Physiotherapy | SHOCKWAVE THERAPY Description: Electrotherapy modality used to treat musculoskeletal conditions Specifications :FDA Approved:Shockwave.Evo Blue Handpiece for constant energy output Radial ESWT Pneumatic technology.External compressor Frequency range:4Hz/8Hz/12Hz/20Hz External compressor Pressure range:1.5bar-4bar Impulse range 500 shots -2500 shots Energy output:15mm headpiece(up to 0.18mJ/mm2; 15mm focused(up to 0.28mJ/mm2 | HITECH THERAPY | E.M.S. ELECTROMEDICAL SYSTEMS | SWITZERLAND | DOLORCLAST | SMART 20 | 2015 | 7 YEARS | 1 | BAR | 333500 | 15000 | 15000 | 15000 | 15000 | 15000 | 2500 | 411000 | RT55-04-039 | 411000 | 411000 | 90 | 0 | 90 | 1 | Recommended
RT55-04-039 | Physiotherapy | SHOCKWAVE THERAPY Description: Electrotherapy modality used to treat musculoskeletal conditions Specifications :FDA Approved:Shockwave.Evo Blue Handpiece for constant energy output Radial ESWT Pneumatic technology.External compressor Frequency range:4Hz/8Hz/12Hz/20Hz External compressor Pressure range:1.5bar-4bar Impulse range 500 shots -2500 shots Energy output:15mm headpiece(up to 0.18mJ/mm2; 15mm focused(up to 0.28mJ/mm2 | Sanbonani Holdings (Pty) Ltd | BTL Industries Limited / BTL Medical SA (Pty) Ltd | United Kingdom | BTL-6000 RSWT Elite | 844-POELMA-M | 2007 | 10 years | 1 | Set | 368874 | 12000 | 12960 | 13998 | 15120 | 16330 | 9000 | 448282 | RT55-04-039 | 448282 | 411000 | 81.8361 | 0 | 81.8361 | 2 | Recommended
RT55-04-041 | Physiotherapy | ULTRASOUND The unit offered must incorporate the latest technology with microprocessor control and monitoring of the treatment functions. The unit must be fully portable.The functional design of the unit must be such that all controls must be clearly and logically arranged on the front panel-making switchover from one treatment mode to another quick and simple.All controls must be clearly labeled and where symbols are used for the controls, the user must find the symbols easy to understand.It is important that the controls panel and controls are flush in order to avoid liquid spillages from gaining entry into the internal P.C.B.’s etc.The unit offered must incorporate an automatic self-test immediately after the unit is switched on in order to confirm that the unit is in good functional working condition.The unit offered must allow for selective ultrasound therapy treatment of both superficial and deeper lying tissue structures. Ultrasound therapy on the unit offered must be provided with ultrasound transducers at a minimum of two (2) frequencies typically as follows: 1) 1MHz or 1,1MHz. 2) 3MHz or 3,1MHz. The unit must provide continuous and pulsed ultrasound modes and a treatment timer, which is built internally into the unit, must be included with the unit offered.The unit must be provided with a Front Panel Alphanumeric Display, which must provide clear viewing under all lighting conditions, for the display of either the intensity in W/sq. cm. or the total power output in watts, which ever is selected or preferred by the user.The Front Panel Digital Display must also display the user selected treatment time in a minimum range of 0 to 30 minutes.The density of the ultrasound radiation must be user selectable within the limits of 0 to 3 W/sq. cm. for pulsed mode and 0 to 2 W/sq. cm. for continuous mode.The effective radiating area of the treatment head must be well defined.The unit must be fully operable off 220 Volt + 10%, 50hz. AC single phase supply. The unit must be fused in both the Live and Neutral. The power cords shall have a 3-point plug approved by SABS. N.B. Units offered which uses an external power supply will not be considered.Treatment transducers must be supplied in a minimum of two different sizes i.e. ONE LARGE and ONE SMALL radiating area for each of the specified frequency and the cost of these must be included in the final bid price.N.B. The bidder must also state the correct description, catalogue number and price inclusive of VAT of these transducers on the schedule of optional accessories.It must also be possible for two treatment heads to be connected simultaneously e.g. a large 1MHz and a small 3MHz in order that practically every application is within easy reach by simply selecting the required frequency on the front panel control.The Treatment transducer must be clearly labeled indicating the frequency.The Parasitic sidewall radiation of the ultrasound heads must be less than 10mW per sq. cm.The unit must provide an indication of the quality of contact between the treatment head and the treatment area.In any event where the treatment transducer head surface contact is insufficient and poor, the treatment power to the transducer head must be automatically reduced to zero and the treatment timer must be automatically stopped and the treatment power and timer operation must only resume once adequate surface contact of the transducer head is restored.The unit must provide an indication of the quality of contact between the treatment head and the treatment area.In any event where the treatment transducer head surface contact is insufficient and poor, the treatment power to the transducer head must be automatically reduced to zero and the treatment timer must be automatically stopped and the treatment power and timer operation must only resume once adequate surface contact of the transducer head is restored.The unit offered must conform to I.E.C. 601-2-5 particular requirements for the Safety of Ultrasound Therapy Equipment. | HITECH THERAPY | Shenzhen Dongdixin Technology Co., Ltd. | CHINA | NU-TEK | UT1021 | 2024 | 7 YEARS | 1 | WATT | 29950 | 1200 | 1200 | 1200 | 1200 | 1200 | 2500 | 38450 | RT55-04-041 | 38450 | 38450 | 90 | 0 | 90 | 1 | Recommended
RT55-04-042 | Physiotherapy | Tilt Table, electronically adjustable 2 section bed with activity table Description :Electronically adjustable tilting table for passively standing patients at varying degrees Technical Specifications: The unit shall have an electric high-low height adjustment with high position ≥90cm and low position ≤45cm .Height and tilt adjustment via a remote control .The unit shall have a minimum load capacity of 200kg .Overall length shall be a minimum of 180cm .Overall width shall be a minimum of 60cm .The frame shall be high grade and durable .The frame shall have a ground clearance of ±15cm to allow for use of patient transfers units such as hoists . The unit shall have heavy duty castor,easily accessible foot operated brakes for easy movement and positioning of the unit.The castors shall be fully retractable . The feet shall be heavy duty and robust with heavy duty non scratch nonslip ferrules.The unit shall be upholstered in durable hygienic, non-slip and washable (e.g. vinyl, or imitation leather) .Upholstery high density foam shall be at least 4-6cm thick. The upholstery shall be free of phthalates and shall be flame retardant . The unit shall comprise of two sections that are adjustable at one point .All sections to be independently adjustable either manually with hydraulic controls, electronically or a combination of manual and electronic controls .If hydraulically adjusted, All adjustment handles shall be durable and made of stainless steel or equivalent quality material . Electronic controls shall be via a bed mounted control box.The unit shall accommodate lying and sitting position. The head section shall have a comfortable breathing slot with a plug.The head section shall have backrest and angle shall allow tilt from at least -30° to +90° .Minimum of three(3) safety straps to support, chest, pelvis and knees.Straps should be soft/padded webbing ,material adjustable via Velcro straps and height and width adjustable to customize for each user. Adjustable foot box by means of quick release spring loaded locking pins to accommodate for ankle contractures.Height and angle adjustable activity table for additional Upper body support and for patients to engage on activities.Additional height adjustable arm supports for stronger patients who do not need a table. Accessories :replacement castors available as spares should be available as a set or individually.The unit shall operate from 220V and the power cords shall have a 3-point plug approved by SABS standard manufacturers length (extra length built in power cord should be available on request) (not an extension lead).The unit shall be FDS or CE approved (Attach certification) .The unit shall be supplied with operators and service & maintenance manuals .All accessories and consumables must be supplied standard with the unit .A separate price of each accessory and consumables listed above must also be supplied with this offer.Any other additional accessories, consumables and spare parts not listed above but applicable for the functionality of the unit must be supplied on a separate price list . | HITECH THERAPY | HUANGHSHAN JINFU MEDICAL EQUIPMENT | CHINA | COINFYCARE | EL12 | 2019 | 7 YEARS | 1 | N/A | 47150 | 950 | 950 | 950 | 950 | 950 | 0 | 51900 | RT55-04-042 | 51900 | 51900 | 90 | 10 | 100 | 1 | Recommended
RT55-04-042 | Physiotherapy | Tilt Table, electronically adjustable 2 section bed with activity table Description :Electronically adjustable tilting table for passively standing patients at varying degrees Technical Specifications: The unit shall have an electric high-low height adjustment with high position ≥90cm and low position ≤45cm .Height and tilt adjustment via a remote control .The unit shall have a minimum load capacity of 200kg .Overall length shall be a minimum of 180cm .Overall width shall be a minimum of 60cm .The frame shall be high grade and durable .The frame shall have a ground clearance of ±15cm to allow for use of patient transfers units such as hoists . The unit shall have heavy duty castor,easily accessible foot operated brakes for easy movement and positioning of the unit.The castors shall be fully retractable . The feet shall be heavy duty and robust with heavy duty non scratch nonslip ferrules.The unit shall be upholstered in durable hygienic, non-slip and washable (e.g. vinyl, or imitation leather) .Upholstery high density foam shall be at least 4-6cm thick. The upholstery shall be free of phthalates and shall be flame retardant . The unit shall comprise of two sections that are adjustable at one point .All sections to be independently adjustable either manually with hydraulic controls, electronically or a combination of manual and electronic controls .If hydraulically adjusted, All adjustment handles shall be durable and made of stainless steel or equivalent quality material . Electronic controls shall be via a bed mounted control box.The unit shall accommodate lying and sitting position. The head section shall have a comfortable breathing slot with a plug.The head section shall have backrest and angle shall allow tilt from at least -30° to +90° .Minimum of three(3) safety straps to support, chest, pelvis and knees.Straps should be soft/padded webbing ,material adjustable via Velcro straps and height and width adjustable to customize for each user. Adjustable foot box by means of quick release spring loaded locking pins to accommodate for ankle contractures.Height and angle adjustable activity table for additional Upper body support and for patients to engage on activities.Additional height adjustable arm supports for stronger patients who do not need a table. Accessories :replacement castors available as spares should be available as a set or individually.The unit shall operate from 220V and the power cords shall have a 3-point plug approved by SABS standard manufacturers length (extra length built in power cord should be available on request) (not an extension lead).The unit shall be FDS or CE approved (Attach certification) .The unit shall be supplied with operators and service & maintenance manuals .All accessories and consumables must be supplied standard with the unit .A separate price of each accessory and consumables listed above must also be supplied with this offer.Any other additional accessories, consumables and spare parts not listed above but applicable for the functionality of the unit must be supplied on a separate price list . | Arjo Huntliegh South Africa | Seers Medical | UK | 2 Section tilt table | RT55-04-042 | 2012 | 15 | 1 | EA | 64285.63 | 4988.33 | 4988.33 | 4988.33 | 4988.33 | 4988.33 |  | 89227.28 | RT55-04-042 | 89227.28 | 51900 | 25.2706 | 10 | 35.2706 | #REF! | Recommended
RT55-05-118 | Podiatry | Cabinet .Description:Podiatry cabinets are used to store clinical equipment and podiatry tools that a podiatrist routinely uses for efficiency of work within a clinical setting.Technical Specification: Cabinet, mobile workstation laminated top, on castors, sliding shelf with power socket for drill. 2 large cupboards and additional 3 drawers, light socket.One large retractable compartment and one retractable motor compartment (empty).One large compartment, complete and ready for use. Dimensions: Height – 73cm / Width – 73cm / Depth – 44cm Full comprehensive minimum 2 year warranty. Cabinet, podiatry, workstation with UV compartment, 3 drawers, white wood decoration, 73cm width, 73cm height, 44cm depth, one large retractable compartment and one retractable motor compartment (empty), one large UV compartment, complete and ready for use, compact footcare cabinet with smooth gliding, braked castor roller feet for effortless change of location. Full comprehensive minimum 2 year warranty. | New Horizon Metals cc | New Horizon Metals cc | South Africa | NHM | PC001 | 2022/2023 | 7 years | 1 | Each | 26450 | 0 | 0 | 0 | 0 | 0 | 0 | 26450 | RT55-05-118 | 26450 | 26450 | 90 | 0 | 90 | 1 | Recommended