## Dual head Gamma camera
Dual Head Gamma Camera
Description
ITEM | Technical Specifications
1 | IT IS ALSO CRITICAL THAT THE PROSPECTIVE BIDDERS OF THIS SPECIALIZED RADIOLOGY EQUIPMENT MUST GUARANTEE THAT THE PRODUCT OFFERED HAS INTERNAL AND EXTERNAL NON – CORROSIVE MATERIAL/COMPONENTS, WHEN ARE EXPOSED TO EXTREME ENVIRONMENTAL FACTORS, SUCH AS THE SEA ETC
1.1 | The gamma camera shall be capable of:
Static, dynamic, whole body and gated planar acquisition and processing
SPECT (including gated SPECT) acquisition, reconstruction, processing and display.
2 | The gamma cameras will be required for, but not limited to, the following studies (i.e. 3phase, dynamic, static and whole body imaging for the various regions):
Head and Neck Studies
e) Thyroid Scans;
f) Parathyroid Imaging;
n) Cerebral perfusion and receptor imaging;
Chest and Cardiac Imaging
c) Myocardial perfusion imaging and thrombus imaging;
d) Gated cardiac blood pool imaging;
g) Lung ventilation and perfusion imaging including quantitative studies;
t) Scintimammography;
s) Sentinel Node imaging;
Abdominal Studies
i) Gastro-oesophageal reflux
j) Gastric emptying studies
k) Gastro-intestinal bleeding
b) Renograms and renal cortex functional imaging;
h) Milk scans;
p) Liver and spleen imaging
q) Hepatobiliary studies;
u) Adrenal Imaging;
Other Studies
a) Skeletal;
l) Gallium imaging for infection and tumours;
m) Infection / inflammation imaging with White Cell / leukoscan/ other available agents for this indication;
o) Tumour imaging in adults and children;
r) Lymphoscintigraphy
v) Pre- / post radionuclide therapy imaging
w) SPECT studies of any of the above.
3 | SPECT Camera
Must be upgradeable to hybrid SPECT/CT technology and this be priced as an option in the bid.
The CT component when installed should be capable of producing CT images that will be used primarily for attenuation correction and anatomic localization.
The supported CT technology options available should be clearly described, bidder to specify.
4 | Must be able to transfer images in DICOM and/or interfile format.
The systems should be able to connect to the planned PACS system (Department of Diagnostic Radiology) and be able to transfer images within the institution and other external institutions
5 | It shall also be possible to interface to the laser printer from the processing stations for hardcopies and to CD/DVD for soft copies
7 | The offered dual head SPECT gamma camera system shall consist of at least the following components:
a) Two large, rectangular field of view NaI detectors with a crystal thickness of 9.5mm
b) One set of low-energy high-resolution collimators
One set of low energy general purpose collimators
c) One set of medium-energy collimators
d) One set of high-energy general purpose collimators for imaging
e) One set of fan beam collimators
f) Pin- hole collimator
g) Storage and trolley for collimator change
h) Data acquisition computer and complete suite of acquisition software
i) Image Processing Computers and complete suite of standard processing software
j) Imaging table
k) ECG for cardiac imaging
l) Mammography pallet for scintimammography
m) Paediatric pallet
n) Integrated DVD/R RW
8 | The Gamma Camera shall have 2 rectangular NaI (TI) detectors with a useful field of view (UFOV) of at least 52 x 37 cm or greater.
9 | The number of photomultiplier tubes per detector shall be at least 56 or greater
10 | The intrinsic performance measured according to National Electrical Manufacturers Association (NEMA) specifications in normal operational mode for 140keV photons in the UFOV shall equal or better the following figures:
(a) Spatial Resolution of less than 4mm FWHM.
(b) Spatial linearity of 0.6 mm absolute and 0.2 mm differential.
(c) Energy resolution <10%.
(d) Uniformity of less then 4% integral and less then 3% differential.
11 | The extrinsic performance (measured according to NEMA specifications) for 140 keV with Tc-99m photons in the UFOV shall be quoted.
12 | The system spatial resolution with scatter (LEHR collimator distance 10 cm) for FWHM for CFOV should be < 9.0 mm and without scatter < 8.0mm.
13 | The system sensitivity per detector with Tc-99m LEHR collimator is > 160 cts/min/µCi.
14 | The reconstructed tomographic spatial resolution shall be, Tangential <9 mm; Radial <12 mm; Central <12mm.
15 | The energy range should be from 40 keV to 580 keV
16 | The variation of PHA response to a point source at 140 keV during a 360° rotation of the detector shall be 1% or better.
17 | The variation in count rate from a point source at 140 keV during a 360° rotation of the detector shall be 5% or better.
18 | The number of PHA windows which can be used to acquire and display data in a single image or image set must be specified with any limitation on dynamic, static, tomographic or gated tomographic acquisitions.
19 | Specify the number of discrete image sets which can be acquired from simultaneously applied PHA windows for static, dynamic, whole body, tomographic or gated tomographic acquisitions (“dual isotope acquisitions”).
Collimators:
20 | The following collimators must be included
a) A pair low energy (140 keV) high resolution parallel hole collimators.
One set of low energy general purpose collimators
b) A pair of medium energy (300 keV) general purpose parallel hole collimators
c) A pair of high energy (360 keV) general purpose parallel hole collimators
d) A pair of Fan beam collimators
e) Pin- hole collimator
21 | Collimators must be electronically coded in the acquisition system.
22 | The effective hole angulation of parallel hole collimators shall be preferably less than 0.15°.
23 | Storage rack and cart shall be supplied. The effective hole angulation of parallel hole collimators shall be preferably less than 0.15°.
Gantry and Detector Movements:
24 | All controls for movement of the detector head must be accessible for any head position.
25 | It must be possible for movements of the camera head and imaging table by a remote control device.
26 | The detector must remain in a fixed position unless activated by the operator.
27 | This must also apply in the event of a power failure.
28 | A patient must be able to be removed from the system in the event of a power failure easily
29 | The head must be equipped with appropriate sensors to detect contact of the head with the patient or an obstruction. This must prevent any further motion of the head.
30 | It must be possible to restart acquisition without loss of data.
31 | Independent detector angulations giving 90° to 180° for SPECT imaging shall be available. Tilting option shall also be available. Specify any other relative detector angles and any other limitations on system operation.
32 | The gantry and detector movements (e.g. rotational, translational) which are manual and those which are motor driven and remotely controlled shall be described. The speeds, range and directions of each of the motor driver movements shall be specified.
33 | Detector alignment with patient shall be automatically maintained while the detector / patient are varied.
34 | The system must have an auto contouring device for automatically sensing the patient-detector distance and adjusting the detector to maintain this distance constant as part of the basic system. The Bidder to state the technology employed.
35 | The acquisition modes for which the auto contouring modes can be used for:
a) SPECT;
b) Whole body.
c) Other specify
36 | Auto contouring device can be used with all the collimators
37 | It should be possible to deactivate auto contouring device
38 | State the distance between collimator and patient that the auto contouring device can achieve;
39 | The maximum and minimum radius of tomography shall be specified.
40 | Any additional general safety features must be separately described.
41 | For tomographic acquisition
a) Clockwise and anticlockwise motions shall be available.
b) Step and shoot plus continouos acquisitions shall be available.
c) Specify if other tomography angles are offered.
d) Specify the matrix sizes with which the system can acquire and store SPECT data and gated SPECT data.
e) An option to upgrade to SPECT /CT imaging SHOULD BE AVAILABLE and specified with regards to compatibility and technology
42 | Whole body Acquisition:
a) The range of scanning speeds should be between 5-60 cm/min.
b) Whole body scans are acquired continuous scanning motion
c) Maximum scan length shall be a minimum of 200 cm.
d) The matrix sizes available shall be a minimum 512 x 512
43 | Imaging Table:
a) The imaging table shall be suitable for static, dynamic, SPECT and whole body acquisitions.
b) The table shall be able to handle a patient load of at least 220 kg or more.
c) The imaging table should be able to be completely detached from the camera and wheeled to the side of the room.
d) The material used in the construction of the table; e.g. carbon fibre, aluminium. Specify the attenuation (%) of 140 keV photons.
44 | Physiological Triggering:
a) The system must be provided with a trigger input and appropriate software to enable acquired data to be gated for ECG gating cardiac studies.
b) Indicate if the ECG gating is controlled by an independent ECG monitor or an ECG monitor built into the gamma camera. Indicate whether it is possible to store the ECG signal as part of the patient file and if the ECG signal can be viewed afterwards
45 | Data Acquisition Hardware:
a) Specify the details of the computer supplied.
b) Indicate the hardware platform that the data acquisition workstation runs on (e.g. PC, MAC, SUN).
c) The latest computer hardware must be supplied.
d) Specify the CPU speed. The system must have a minimum speed of 2.8GHz.
e) Specify the amount of memory installed.
f) Specify the capacity of the hard disc drive.
g) Specify the operating system as well as the version that will be installed.
h) Specify any other storage devices available.
i) A flat panel LCD model with a minimum size of 19 inch must be supplied.
j) The monitor must have a minimum resolution of 1920 X 1200.
46 | Acquisition:
a) The system should be able to display images in persistence mode to enable ease of patient positioning before initiating acquisition and also be available for each detector head.
b) Rotate the acquiring image in multiples of 90° (i.e 90, 180, 270 and 360 degrees).
c) The range of acquisition software for the camera shall include at least the following: static, dynamic, multiple gated acquisition (planar) and gated SPECT, first pass, whole body, tomography.
d) The software shall allow the creation of predefined user protocols for routine acquisitions.
e) The following should be able to be specified in the pre-defined protocols: Study name, view name, energy window, matrix size, matrix depth, acquisition termination condition, image zoom, collimator, gantry and imaging table position.
f) It should be possible for images to be acquired from each head independently and stored separately.
g) It should be possible to store any study halted prematurely and to pause and restart acquisition without loss of data.
h) It should be possible for acquisitions to be terminable either by preset count, preset time, accepted heart beats, manually or by count density within a user-defined ROI.
i) Various matrix sizes and pixel depths should be available for the acquisition types that the system is capable of performing.
j) Various magnification factors should be available for zoomed acquisitions and indicate limitations thereof.
k) The system shall have a range of pre-defined radionuclides such as 99mTc, 67Ga, 57Co, 111In, 123I, 131I, 201TI and 81mKr and it shall be possible for the user to add radionuclides not included in the list.
l) The system shall be able to correct for isotope decay for tomographic as well as dynamic acquisition.
m) All acquisitions must be possible using conventional energy window settings that shall preferably be calibrated in keV.
n) For dynamic acquisitions indicate the minimum and maximum time per frame as well as the increments.
o) Indicate the number of different framing rates that would be possible to use during a single dynamic study.
p) Automatic body contouring shall be possible for whole body and SPECT images. Provide details of the method used for automatic body contouring.
q) Anterior and posterior images shall be acquired simultaneously for whole body acquisitions.
r) Gated planar and SPECT acquisition is required. A minimum of 16 frames is necessary. Specify the minimum and maximum number of frames that may be acquired during a cardiac cycle.
s) The user shall be able to define a window of acceptable beat lengths. Data from all cardiac cycles that fall outside this window shall be rejected.
t) The orbital movement for tomography shall include circular orbit, elliptical orbit and body contouring.
u) Gantry and bed positions shall be stored within the acquisition data file so that it can be used for patient re-positioning.
v) It should be possible for all collimators in tendered list to be used for SPECT.
w) Step and shoot acquisition shall be available and step angle increments and minimum and maximum shoot times and number of images indicated
47 | Quality Control Software:
The system shall have software for monitoring
a) the system as required by NEMA.
b) Software for the calibration of energy, linearity, uniformity and centre of rotation correction shall be available.
c) State if axial alignment correction is possible
48 | Image Display and ROI Management:
a) The system shall have several standard supplied colour tables, including gray scales, to display images.
b) The software should allow for the creation of user-defined colour tables.
c) It shall be possible to display several images simultaneously and to adjust display levels independently in each window.
d) It shall be possible to convert images from one matrix size to another for static, dynamic, tomographic and gated tomographic studies.
e) It shall be possible to create a new dynamic study from frames of an existing dynamic study.
f) Summation of contiguous frames in a dynamic series shall be executable.
g) It must be possible to produce condensed images from a dynamic series.
h) Indicate if variable width profile generation and analysis thereof is possible.
i) Spatial and temporal smoothing of images shall be possible.
j) Image algebra on static and dynamic images shall be available.
k) Image rotation and mirroring shall be available as well as linear translation for static, dynamic and tomography acquisitions so that correction for patient motion is possible.
l) It shall be possible to view SPECT data in a 3-D volume/cine display.
m) The system shall have software for the creation of ROI’s on images or summed images using manual methods, preset geometrical shapes, isocount contouring and image thresholding.
n) Indicate if it is possible to store ROI’s in the database so that they can be retrieved and manipulated at a later stage.
o) Indicate if ROI’s can be superimposed on a cine display of images and any limitations thereof.
p) Indicate if two separate cine-gated cardiac displays (e.g. stress and rest) can be synchronized so that corresponding frames are displayed simultaneously.
q) ROI arithmetic shall be performable and data production on total counts, size and average counts per pixel shall be possible. ROI manipulation and mirroring shall also be possible
r) Generation of activity-time curves from ROI’s shall be possible. The storage and display of such curves shall also be possible.
s) Image quantitation ability shall be available
t) It shall be possible to perform curve algebra, to smooth, integrate and differentiate curves.
u) Indicate if curve fitting is possible and if so, indicate the types.
49 | Reconstruction Software:
a) Reconstruction of tomography images shall be possible using filtered back projection as well as iterative techniques. Iterative reconstruction software shall be possible
b) Indicate the maximum tolerable centre of rotation offset that can be corrected by software.
c) Indicate if scatter corrections are available and what this entails.
d) Specify the method(s) of attenuation correction that are available.
e) SPECT reconstruction and storage of axial, coronal, sagital and oblique planes shall be possible.
f) SPECT reconstruction of circular, elliptical and body contour orbits shall be possible.
g) Indicate the filters that are available for reconstruction purposes.
h) Indicate if alignment of sequential tomographic images is possible using either landmarks or manually.
I) CT reconstruction and storage of axial, coronal, sagittal and oblique planes shall be possible.
50 | Clinical Processing Software:
Indicate if the system has the capability to perform the following processing functions for dynamic renal imaging:
a) Calculation of relative function;
b) Calculation of retention function;
c) Calculation of parenchymal transit time;
d) Calculation of output efficiency;
e) Calculation of whole kidney transit time;
f) Determination of furosemide response;
g) Determination of the washout halftime;
h) Generation of Patlak-Rutland plots and
i) Determination of the perfusion index.
51 | Indicate if the system has software for calculation of ventilation/perfusion ratio in lung imaging and calculation of washout rates from ventilation studies
52 | Indicate if the system is capable of
a) Esophageal parametric images and the calculation of the transit time.
b) Gastric emptying rate and halftime can be calculated including use of geometric mean images.
c) Parathyroid subtraction imaging software is available.
d) Brain Processing Software.
53 | For cardiac
a) Create Bull’s-eye images for myocardial perfusion
b) 4DM/QGS cardiac processing software.
c) Software for the 3-axis display of myocardial perfusion images.
d) MUGA can be determined both semi & fully automatically and if the Fourier phase & amplitude images can be displayed.
e) Beat histogram can be displayed.
f) MUGA calculations can be performed for LVEF, PFR, PER and regional EF.
g) Calculate First pass RVEF.
h) Perform First pass shunt quantitation.
54 | It shall be indicated if software updates will be supplied and installed and what the relevant costs will be.
55 | File Handling and Data Processing:
a) The system must have standard directory and file handling utilities.
b) The system must have archiving and indexing software and allow searches using several different criteria such as patient name, ID number, study type etc.
c) Archiving can be initiated e.g. automatically or manually.
d) The system shall have the software to convert images to TIFF, JPEG, GIF or AVI formats that are compatible with MS Windows.
e) The system shall have the capability and necessary software for CD or DVD writing.
f) Installation of third party software such as antivirus software shall be possible.
g) The system shall be capable of connecting to other systems using standard Ethernet protocols.
h) The system shall be able to reliably import and export data from and to other processors currently in use with Interfile or DICOM facilities.
i) The system shall have the necessary TCP/IP software to make communication by a standard network system possible.
j) Indicate if automatic transfer of completed acquisitions is possible.
k) Indicate if it is possible for the user to transfer studies from the system to a remote station (push) and/or remotely transfer studies from the system (pull).
l) The Bidder to provide details about the image archiving and storage facilities, as well as the number of 512 x 512 resolution images that can be stored, on the internal hard drive(s). The hard drive(s) shall be large-capacity industry standard, and easily upgradeable.
m) It shall also be possible to store/archive images on commercially available CD-ROM or DVD disks in a generally recognizable and readable DICOM format.
56 | DICOM COMPATIBLE LASER CAMERA.
DICOM Compatible networked laser camera must be provided. (455 DPI or better, for printing on 18 x 24cm film size is required).
51 | APPLICABLE DOCUMENTS
The offered product shall comply with the National Year 2000 Conformity Requirements as defined in the Republic of South Africa. PD2000-1: 1998 Year 2000 definition. Please clearly state if the offered product complies
52 | POWER SUPPLY
The bidder must install adequate electrical power supply for the optimal functionality of the equipment.
The mains cable of the unit being quoted for must be the hospital grade type and it must be of an adequate length (minimum length of three (3) meters).
The power input must be 220-240V, 50Hz AC. Bidders must ensure that the product quoted for is fitted appropriately with a 15 Amp SABS approved mains plug that is held together by two screws.
A 3 phase in / out power supply must be offered. State details.
The equipment tendered for must not overload and trip the hospital power in the area where it is installed. Resettable overcurrent breaker/s must be fitted for protection. State details.
All available online Uninterrupted Power Supply (UPS system) and line voltage regulators must be included for the system with a support time of 50 to 60 minutes. Attach brochure.
The lifetime of UPS must be at least 15 years.
An Uninterrupted Power Supply (UPS system) conforming to IS-302 must be offered for the computer system.
The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade type and it must be a minimum length of three (3) meters.
The mains cable of the unit tendered for must be SABS colour coded. State details.
The equipment quoted for must be protected against electromagnetic interference. State details of measures that will be taken to protect the unit against electromagnetic interference.
Maintenance free batteries must be offered.
58 | OPTIONS
The Bidder is to offer the following as optional extras (state price and details) for each separately. The offered options shall be fully functional options that require no additional hardware or software to function:
OPTION A
Compatible CT component technology to upgrade the currect SPECT to hybrid SPECT/CT
The different available options of CT technology packages that can be incorporated on the system to upgrade to a hybrid SPECT/CT camera should be clearly stated and described for comprehensive consideration
OPTION B
Image Registration Software
It must be able to automatically register SPECT and CT images
It must be able to register images which are obtained from other modalities such as MRI, CT and PET and other gamma cameras.
OPTION C
If there are any additional available software or hardware options not specified in this bid, the bidder is requested to motivate the benefits of each option.
OPTION D
If a separate persistence scope is not supplied with the main offer, quote the price for a persistence scope.
OPTION E
Must be supplied with fully automated collimator exchange system.

## Single head Gamma Camera
Single Head Gamma Camera
Description
ITEM | Technical Specifications
1 | IT IS ALSO CRITICAL THAT THE PROSPECTIVE BIDDERS OF THIS SPECIALIZED RADIOLOGY EQUIPMENT MUST GUARANTEE THAT THE PRODUCT OFFERED HAS INTERNAL AND EXTERNAL NON – CORROSIVE MATERIAL/COMPONENTS, WHEN ARE EXPOSED TO EXTREME ENVIRONMENTAL FACTORS, SUCH AS THE SEA ETC
1.1 | The gamma camera shall be capable of:
Static, dynamic, whole body and gated planar acquisition and processing
SPECT (including gated SPECT) acquisition, reconstruction, processing and display.
2 | The gamma cameras will be required for, but not limited to, the following studies (i.e. 3phase, dynamic, static and whole body imaging for the various regions):
Head and Neck Studies
e) Thyroid Scans;
f) Parathyroid Imaging;
n) Cerebral perfusion and receptor imaging;
Chest and Cardiac Imaging
c) Myocardial perfusion imaging and thrombus imaging;
d) Gated cardiac blood pool imaging;
g) Lung ventilation and perfusion imaging including quantitative studies;
t) Scintimammography;
s) Sentinel Node imaging;
Abdominal Studies
i) Gastro-oesophageal reflux
j) Gastric emptying studies
k) Gastro-intestinal bleeding
b) Renograms and renal cortex functional imaging;
h) Milk scans;
p) Liver and spleen imaging
q) Hepatobiliary studies;
u) Adrenal Imaging;
Other Studies
a) Skeletal;
l) Gallium imaging for infection and tumours;
m) Infection / inflammation imaging with White Cell / leukoscan/ other available agents for this indication;
o) Tumour imaging in adults and children;
r) Lymphoscintigraphy
v) Pre- / post radionuclide therapy imaging
w) SPECT studies of any of the above.
3 | SPECT Camera
Must be upgradeable to hybrid SPECT/CT technology and this be priced as an option in the bid.
The CT component when installed should be capable of producing CT images that will be used primarily for attenuation correction and anatomic localization.
The supported CT technology options available should be clearly described, bidder to specify.
4 | Must be able to transfer images in DICOM and/or interfile format.
The systems should be able to connect to the planned PACS system (Department of Diagnostic Radiology) and be able to transfer images within the institution and other external institutions
5 | It shall also be possible to interface to the laser printer from the processing stations for hardcopies and to CD/DVD for soft copies
7 | The offered dual head SPECT gamma camera system shall consist of at least the following components:
a) Large, rectangular field of view NaI detectors with a crystal thickness of 9.5mm
b) Low-energy high-resolution collimator
Low energy general purpose collimator
c) Medium-energy collimator
d) High-energy general purpose collimator for imaging
e) Fan beam collimator
f) Pin- hole collimator
g) Storage and trolley for collimator change
h) Data acquisition computer and complete suite of acquisition software
i) Image Processing Computers and complete suite of standard processing software
j) Imaging table
k) ECG for cardiac imaging
l) Mammography pallet for scintimammography
m) Paediatric pallet
n) Integrated DVD/R RW
8 | The Gamma Camera shall have rectangular NaI (TI) detector with a useful field of view (UFOV) of at least 52 x 37 cm or greater.
9 | The number of photomultiplier tubes per detector shall be at least 56 or greater
10 | The intrinsic performance measured according to National Electrical Manufacturers Association (NEMA) specifications in normal operational mode for 140keV photons in the UFOV shall equal or better the following figures:
(a) Spatial Resolution of less than 4mm FWHM.
(b) Spatial linearity of 0.6 mm absolute and 0.2 mm differential.
(c) Energy resolution <10%.
(d) Uniformity of less then 4% integral and less then 3% differential.
11 | The extrinsic performance (measured according to NEMA specifications) for 140 keV with Tc-99m photons in the UFOV shall be quoted.
12 | The system spatial resolution with scatter (LEHR collimator distance 10 cm) for FWHM for CFOV should be < 9.0 mm and without scatter < 8.0mm.
13 | The system sensitivity per detector with Tc-99m LEHR collimator is > 160 cts/min/µCi.
14 | The reconstructed tomographic spatial resolution shall be, Tangential <9 mm; Radial <12 mm; Central <12mm.
15 | The energy range should be from 40 keV to 580 keV
16 | The variation of PHA response to a point source at 140 keV during a 360° rotation of the detector shall be 1% or better.
17 | The variation in count rate from a point source at 140 keV during a 360° rotation of the detector shall be 5% or better.
18 | The number of PHA windows which can be used to acquire and display data in a single image or image set must be specified with any limitation on dynamic, static, tomographic or gated tomographic acquisitions.
19 | Specify the number of discrete image sets which can be acquired from simultaneously applied PHA windows for static, dynamic, whole body, tomographic or gated tomographic acquisitions (“dual isotope acquisitions”).
Collimators:
20 | The following collimators must be included
a) Low energy (140 keV) high resolution parallel hole collimator.
Low energy general purpose collimator
b) Medium energy (300 keV) general purpose parallel hole collimator
c) High energy (360 keV) general purpose parallel hole collimator
d) Fan beam collimator
e) Pin- hole collimator
21 | Collimators must be electronically coded in the acquisition system.
22 | The effective hole angulation of parallel hole collimators shall be preferably less than 0.15°.
23 | Storage rack and cart shall be supplied. The effective hole angulation of parallel hole collimators shall be preferably less than 0.15°.
Gantry and Detector Movements:
24 | All controls for movement of the detector head must be accessible for any head position.
25 | It must be possible for movements of the camera head and imaging table by a remote control device.
26 | The detector must remain in a fixed position unless activated by the operator.
27 | This must also apply in the event of a power failure.
28 | A patient must be able to be removed from the system in the event of a power failure easily
29 | The head must be equipped with appropriate sensors to detect contact of the head with the patient or an obstruction. This must prevent any further motion of the head.
30 | It must be possible to restart acquisition without loss of data.
31 | Independent detector angulations giving 90° to 180° for SPECT imaging shall be available. Tilting option shall also be available. Specify any other relative detector angles and any other limitations on system operation.
32 | The gantry and detector movements (e.g. rotational, translational) which are manual and those which are motor driven and remotely controlled shall be described. The speeds, range and directions of each of the motor driver movements shall be specified.
33 | Detector alignment with patient shall be automatically maintained while the detector / patient are varied.
34 | The system must have an auto contouring device for automatically sensing the patient-detector distance and adjusting the detector to maintain this distance constant as part of the basic system. The Bidder to state the technology employed.
35 | The acquisition modes for which the auto contouring modes can be used for:
a) SPECT;
b) Whole body.
c) Other specify
36 | Auto contouring device can be used with all the collimators
37 | It should be possible to deactivate auto contouring device
38 | State the distance between collimator and patient that the auto contouring device can achieve;
39 | The maximum and minimum radius of tomography shall be specified.
40 | Any additional general safety features must be separately described.
41 | For tomographic acquisition
a) Clockwise and anticlockwise motions shall be available.
b) Step and shoot plus continouos acquisitions shall be available.
c) Specify if other tomography angles are offered.
d) Specify the matrix sizes with which the system can acquire and store SPECT data and gated SPECT data.
e) An option to upgrade to SPECT /CT imaging SHOULD BE AVAILABLE and specified with regards to compatibility and technology
42 | Whole body Acquisition:
a) The range of scanning speeds should be between 5-60 cm/min.
b) Whole body scans are acquired continuous scanning motion
c) Maximum scan length shall be a minimum of 200 cm.
d) The matrix sizes available shall be a minimum 512 x 512
43 | Imaging Table:
a) The imaging table shall be suitable for static, dynamic, SPECT and whole body acquisitions.
b) The table shall be able to handle a patient load of at least 220 kg or more.
c) The imaging table should be able to be completely detached from the camera and wheeled to the side of the room.
d) The material used in the construction of the table; e.g. carbon fibre, aluminium. Specify the attenuation (%) of 140 keV photons.
44 | Physiological Triggering:
a) The system must be provided with a trigger input and appropriate software to enable acquired data to be gated for ECG gating cardiac studies.
b) Indicate if the ECG gating is controlled by an independent ECG monitor or an ECG monitor built into the gamma camera. Indicate whether it is possible to store the ECG signal as part of the patient file and if the ECG signal can be viewed afterwards
45 | Data Acquisition Hardware:
a) Specify the details of the computer supplied.
b) Indicate the hardware platform that the data acquisition workstation runs on (e.g. PC, MAC, SUN).
c) The latest computer hardware must be supplied.
d) Specify the CPU speed. The system must have a minimum speed of 2.8GHz.
e) Specify the amount of memory installed.
f) Specify the capacity of the hard disc drive.
g) Specify the operating system as well as the version that will be installed.
h) Specify any other storage devices available.
i) A flat panel LCD model with a minimum size of 19 inch must be supplied.
j) The monitor must have a minimum resolution of 1920 X 1200.
46 | Acquisition:
a) The system should be able to display images in persistence mode to enable ease of patient positioning before initiating acquisition and also be available for each detector head.
b) Rotate the acquiring image in multiples of 90° (i.e 90, 180, 270 and 360 degrees).
c) The range of acquisition software for the camera shall include at least the following: static, dynamic, multiple gated acquisition (planar) and gated SPECT, first pass, whole body, tomography.
d) The software shall allow the creation of predefined user protocols for routine acquisitions.
e) The following should be able to be specified in the pre-defined protocols: Study name, view name, energy window, matrix size, matrix depth, acquisition termination condition, image zoom, collimator, gantry and imaging table position.
f) It should be possible for images to be acquired from each head independently and stored separately.
g) It should be possible to store any study halted prematurely and to pause and restart acquisition without loss of data.
h) It should be possible for acquisitions to be terminable either by preset count, preset time, accepted heart beats, manually or by count density within a user-defined ROI.
i) Various matrix sizes and pixel depths should be available for the acquisition types that the system is capable of performing.
j) Various magnification factors should be available for zoomed acquisitions and indicate limitations thereof.
k) The system shall have a range of pre-defined radionuclides such as 99mTc, 67Ga, 57Co, 111In, 123I, 131I, 201TI and 81mKr and it shall be possible for the user to add radionuclides not included in the list.
l) The system shall be able to correct for isotope decay for tomographic as well as dynamic acquisition.
m) All acquisitions must be possible using conventional energy window settings that shall preferably be calibrated in keV.
n) For dynamic acquisitions indicate the minimum and maximum time per frame as well as the increments.
o) Indicate the number of different framing rates that would be possible to use during a single dynamic study.
p) Automatic body contouring shall be possible for whole body and SPECT images. Provide details of the method used for automatic body contouring.
q) Anterior and posterior images shall be acquired simultaneously for whole body acquisitions.
r) Gated planar and SPECT acquisition is required. A minimum of 16 frames is necessary. Specify the minimum and maximum number of frames that may be acquired during a cardiac cycle.
s) The user shall be able to define a window of acceptable beat lengths. Data from all cardiac cycles that fall outside this window shall be rejected.
t) The orbital movement for tomography shall include circular orbit, elliptical orbit and body contouring.
u) Gantry and bed positions shall be stored within the acquisition data file so that it can be used for patient re-positioning.
v) It should be possible for all collimators in tendered list to be used for SPECT.
w) Step and shoot acquisition shall be available and step angle increments and minimum and maximum shoot times and number of images indicated
47 | Quality Control Software:
The system shall have software for monitoring
a) the system as required by NEMA.
b) Software for the calibration of energy, linearity, uniformity and centre of rotation correction shall be available.
c) State if axial alignment correction is possible
48 | Image Display and ROI Management:
a) The system shall have several standard supplied colour tables, including gray scales, to display images.
b) The software should allow for the creation of user-defined colour tables.
c) It shall be possible to display several images simultaneously and to adjust display levels independently in each window.
d) It shall be possible to convert images from one matrix size to another for static, dynamic, tomographic and gated tomographic studies.
e) It shall be possible to create a new dynamic study from frames of an existing dynamic study.
f) Summation of contiguous frames in a dynamic series shall be executable.
g) It must be possible to produce condensed images from a dynamic series.
h) Indicate if variable width profile generation and analysis thereof is possible.
i) Spatial and temporal smoothing of images shall be possible.
j) Image algebra on static and dynamic images shall be available.
k) Image rotation and mirroring shall be available as well as linear translation for static, dynamic and tomography acquisitions so that correction for patient motion is possible.
l) It shall be possible to view SPECT data in a 3-D volume/cine display.
m) The system shall have software for the creation of ROI’s on images or summed images using manual methods, preset geometrical shapes, isocount contouring and image thresholding.
n) Indicate if it is possible to store ROI’s in the database so that they can be retrieved and manipulated at a later stage.
o) Indicate if ROI’s can be superimposed on a cine display of images and any limitations thereof.
p) Indicate if two separate cine-gated cardiac displays (e.g. stress and rest) can be synchronized so that corresponding frames are displayed simultaneously.
q) ROI arithmetic shall be performable and data production on total counts, size and average counts per pixel shall be possible. ROI manipulation and mirroring shall also be possible
r) Generation of activity-time curves from ROI’s shall be possible. The storage and display of such curves shall also be possible.
s) Image quantitation ability shall be available
t) It shall be possible to perform curve algebra, to smooth, integrate and differentiate curves.
u) Indicate if curve fitting is possible and if so, indicate the types.
49 | Reconstruction Software:
a) Reconstruction of tomography images shall be possible using filtered back projection as well as iterative techniques. Iterative reconstruction software shall be possible
b) Indicate the maximum tolerable centre of rotation offset that can be corrected by software.
c) Indicate if scatter corrections are available and what this entails.
d) Specify the method(s) of attenuation correction that are available.
e) SPECT reconstruction and storage of axial, coronal, sagital and oblique planes shall be possible.
f) SPECT reconstruction of circular, elliptical and body contour orbits shall be possible.
g) Indicate the filters that are available for reconstruction purposes.
h) Indicate if alignment of sequential tomographic images is possible using either landmarks or manually.
I) CT reconstruction and storage of axial, coronal, sagittal and oblique planes shall be possible.
50 | Clinical Processing Software:
Indicate if the system has the capability to perform the following processing functions for dynamic renal imaging:
a) Calculation of relative function;
b) Calculation of retention function;
c) Calculation of parenchymal transit time;
d) Calculation of output efficiency;
e) Calculation of whole kidney transit time;
f) Determination of furosemide response;
g) Determination of the washout halftime;
h) Generation of Patlak-Rutland plots and
i) Determination of the perfusion index.
51 | Indicate if the system has software for calculation of ventilation/perfusion ratio in lung imaging and calculation of washout rates from ventilation studies
52 | Indicate if the system is capable of
a) Esophageal parametric images and the calculation of the transit time.
b) Gastric emptying rate and halftime can be calculated including use of geometric mean images.
c) Parathyroid subtraction imaging software is available.
d) Brain Processing Software.
53 | For cardiac
a) Create Bull’s-eye images for myocardial perfusion
b) 4DM/QGS cardiac processing software.
c) Software for the 3-axis display of myocardial perfusion images.
d) MUGA can be determined both semi & fully automatically and if the Fourier phase & amplitude images can be displayed.
e) Beat histogram can be displayed.
f) MUGA calculations can be performed for LVEF, PFR, PER and regional EF.
g) Calculate First pass RVEF.
h) Perform First pass shunt quantitation.
54 | It shall be indicated if software updates will be supplied and installed and what the relevant costs will be.
55 | File Handling and Data Processing:
a) The system must have standard directory and file handling utilities.
b) The system must have archiving and indexing software and allow searches using several different criteria such as patient name, ID number, study type etc.
c) Archiving can be initiated e.g. automatically or manually.
d) The system shall have the software to convert images to TIFF, JPEG, GIF or AVI formats that are compatible with MS Windows.
e) The system shall have the capability and necessary software for CD or DVD writing.
f) Installation of third party software such as antivirus software shall be possible.
g) The system shall be capable of connecting to other systems using standard Ethernet protocols.
h) The system shall be able to reliably import and export data from and to other processors currently in use with Interfile or DICOM facilities.
i) The system shall have the necessary TCP/IP software to make communication by a standard network system possible.
j) Indicate if automatic transfer of completed acquisitions is possible.
k) Indicate if it is possible for the user to transfer studies from the system to a remote station (push) and/or remotely transfer studies from the system (pull).
l) The Bidder to provide details about the image archiving and storage facilities, as well as the number of 512 x 512 resolution images that can be stored, on the internal hard drive(s). The hard drive(s) shall be large-capacity industry standard, and easily upgradeable.
m) It shall also be possible to store/archive images on commercially available CD-ROM or DVD disks in a generally recognizable and readable DICOM format.
56 | DICOM COMPATIBLE LASER CAMERA.
DICOM Compatible networked laser camera must be provided. (455 DPI or better, for printing on 18 x 24cm film size is required).
51 | APPLICABLE DOCUMENTS
The offered product shall comply with the National Year 2000 Conformity Requirements as defined in the Republic of South Africa. PD2000-1: 1998 Year 2000 definition. Please clearly state if the offered product complies
52 | POWER SUPPLY
The bidder must install adequate electrical power supply for the optimal functionality of the equipment.
The mains cable of the unit being quoted for must be the hospital grade type and it must be of an adequate length (minimum length of three (3) meters).
The power input must be 220-240V, 50Hz AC. Bidders must ensure that the product quoted for is fitted appropriately with a 15 Amp SABS approved mains plug that is held together by two screws.
A 3 phase in / out power supply must be offered. State details.
The equipment tendered for must not overload and trip the hospital power in the area where it is installed. Resettable overcurrent breaker/s must be fitted for protection. State details.
All available online Uninterrupted Power Supply (UPS system) and line voltage regulators must be included for the system with a support time of 50 to 60 minutes. Attach brochure.
The lifetime of UPS must be at least 15 years.
An Uninterrupted Power Supply (UPS system) conforming to IS-302 must be offered for the computer system.
The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade type and it must be a minimum length of three (3) meters.
The mains cable of the unit tendered for must be SABS colour coded. State details.
The equipment quoted for must be protected against electromagnetic interference. State details of measures that will be taken to protect the unit against electromagnetic interference.
Maintenance free batteries must be offered.
58 | OPTIONS
The Bidder is to offer the following as optional extras (state price and details) for each separately. The offered options shall be fully functional options that require no additional hardware or software to function:
OPTION A
Compatible CT component technology to upgrade the currect SPECT to hybrid SPECT/CT
The different available options of CT technology packages that can be incorporated on the system to upgrade to a hybrid SPECT/CT camera should be clearly stated and described for comprehensive consideration
OPTION B
Image Registration Software
It must be able to automatically register SPECT and CT images
It must be able to register images which are obtained from other modalities such as MRI, CT and PET and other gamma cameras.
OPTION C
If there are any additional available software or hardware options not specified in this bid, the bidder is requested to motivate the benefits of each option.
OPTION D
If a separate persistence scope is not supplied with the main offer, quote the price for a persistence scope.
OPTION E
Must be supplied with fully automated collimator exchange system.

## Gamma SPECT (16 slice or more)
Dual Detector Stationary Gamma Camera (SPECT CT) (CT = 16 slice)
Item | Technical Specifications | Weight
1 | General Requirements
1.1 | The system is designed for general purpose whole body and planar Nuclear Medicine Studies | 5
1.2 | The system offers SPECT CT (CT = 16 slice) and GSPECT imaging | #
1.3 | The System offers ECG gating functionality | 5
2 | Detector and Detector assembly
At least one of the two detectors shall allow for caudal/cephalic tilt.
2.1 | The minimum UFOV dimensions (min 50x38cm) | 3
2.1.1 | Please state the crystal thickness ie 3/8" | 3
2.2 | Specify the number of PMTs (min 59) and their diameter in cm | #
2.3 | The maximum Count rate is at least 220000 cps | 5
2.4 | The energy range is 40 to 580 keV | 5
2.5 | Detector parameters (NEMA Standard)
Intrinsic spatial resolution (FWHM) with 20% energy window for Tc-99m
2.5.1 | FWHM, UFOV (<4mm) | 5
2.5.2 | FWTM, UFOV (<7mm) | 5
Intrinsic energy resolution for Tc-99m
2.5.3 | FWHM, 140keV (<10%) | 5
Intrinsic linearity
2.5.4 | Differential, UFOV (<0.2mm) | 5
2.5.5 | Absolute, UFOV (<0.6mm) | 5
System spatial resolution with LEHR collimator
2.5.6 | FWHM 10cm, without scatter (<8mm) | 5
Intrinsic flood field uniformity with 20% energy window for Tc-99m
2.5.7 | Differential, UFOV (<3%) | 5
2.5.8 | Integral, UFOV (<4%) | 5
Automatic corrections
2.5.9 | Energy correction | 5
2.5.10 | Linearity correction | 5
2.5.11 | Uniformity correction | 5
2.5.12 | Center of rotation correction | 5
3 | Gantry
3.1 | The Gantry contains dual detectors. Provide brief details regarding the method of gantry construction | #
3.2 | The gantry should include a persistence monitor mounted on gantry. | #
3.3 | The gantry shall support variable angle configurability of the detectors including 76°, 90° and 180°. | #
3.4 | The automatic contouring must be available | #
3.7 | Please state the distance range between both detectors | 3
3.8 | Should be able to image patients while standing or seated on wheelchairs | 3
3.9 | Specify any limitation on any for the motions in 3.2 that depend on the type of collimator | #
3.10 | The movement controls (hand controls) must be on both sides of the gantry. | #
The emergency button to stop the movements must be on both sides of the gantry
3.11 | Manual controls are available to move the detector head away from the patient in the event of a power failure. | 5
3.12 | Gantry weight (without collimators) [kg] | 3
3.13 | Typical heat dissipation [Watt/hr] | 3
3.14 | Gantry opening (diameter) must be atleast 70cm | 3
3.15 | Upright height (min & max) for both detectors | 3
3.16 | Table
The table must be motorised
It must have movement controls on both sides
The whole body scan length shall be at least 200 cm
Minimum patient table height shall be ≤ 55cm for easy loading/ unloading of patients
3.16.1 | The minimum patient weight of 220kg | 3
3.16.2 | Please state the Pallet size [length x width x thickness) [cm] | 3
3.16.3 | Attenuation correction percentage of the table palette for energy 140keV [7%] | 3
4 | Collimators
4.1 | The following Collimators are supplied with the system: | #
4.1.1 | - LEGP | 3
4.1.2 | - LEHR | 3
4.1.3 | - Pinhole | 3
4.1.4 | - HEGP | 3
4.1.5 | - MEGP | 3
Cart (or carts) for all collimators
4.2 | Describe the process of collimator exchange | 3
5 | Operator Console
5.1 | The operator Console provide the following functions: | #
5.1.1 | - Zoom Control | 2
5.1.2 | - Scaler | 2
5.1.3 | - Manual and Automatic energy selection | 2
5.1.4 | - Image rotation control | 2
Emergency stop botton
6 | Acquisition Workstation
6.1 | An independent acquisition workstation should be offerred separately from the processing workstation | #
6.2 | Customizable automatic acquisition setup allowing best workflow efficiency | 3
6.3 | A recordable DVD or CD Drive is required | 3
6.4 | Ability to acquire minimum of 2 datasets (with different settings such as matrix size, zoom, time or count settings and gating) simultaneously is preferred | 3
6.5 | Ability to launch collimator exchange automatically from acquisition menu | 3
6.6 | Number of energy windows available | 3
6.7 | User written acquisition protocols | 3
6.8 | Persistence scope monitor | 3
6.9 | System quality control tools | 3
6.10. | Static acquisitions | 3
| - Matrices (dimensions) [pixel x pixel] | 3
| - MUGA acquisitions | 3
6.11 | Dynamic acquisitions | 5
| - Maximum number of intervals (min 3) | 5
| - Matrices (dimensions) [pixel x pixel] | 5
| - Maximum rate [matrices/sec] | 5
6.12 | Whole body scintigraphy | 3
| - Maximum whole body scan length (min 190cm) | 3
| - Automatic body contouring | 3
6.13 | SPECT acquisitions | 5
| - Scanning diameter with LEHR collimators (min to max) [cm - cm] | 5
| - Detectors geometry (available angles) [min 180º, 90º, 0º] | 5
| - Detectors rotation range around camera axis (in degrees) | 5
| - Maximum rotation speed of detectors (rpm) | 5
| - ECG gated tomography | 5
| - Automatic body contouring with all types of collimators | 5
6.14 | SPECT CT
Must have the following:
image processing and analysis
image archiving-including auto archiving
image networking-including auto transfer
Auto view of newly reconstructed images
DVD writer
32 GB or better RAM for the main system
applications for dose reduction of paediatric patients must be available
Reconstruction must be multi modality (SPECT and CT)
7 | Processing Workstation
The unit must come with atleast 2 processing workstation
7.1 | An independent acquisition workstation should be offerred separated from the processing workstation and should include | #
| - Color monitor (min 19'') | 5
| - RAM memory [min 2GB] | 5
| - Hard Disk capacity [min 500GB] | 5
| - Data archiving media | 5
7.2 | Software for automatic SPECT data reconstruction iterative method | 5
7.3 | Software packages should include all routine Nuclear Medicine procedures for: | 5
| - Quantitative analysis for cardiac SPECT studies (including cardiac gated SPECT) with normal patient databases [QBS, QPS, QGS cardiac protocols] | 5
| - Planar gated data analysis | 5
| - First pass analysis | 5
| - Cardiac shunt analysis | 5
| - Quantitative analysis of lung perfusion and ventilation | 5
| - Gallbladder ejection fraction analysis | 5
| - Gastric emptying evaluation | 5
| - Comprehensive renal analysis (deconvolution, transit times, GFR, ERPF, renogram) | 5
| - Liver functin analysis | 5
| - Semi-quantitative brain perfusion analysis | 5
| - Thyroid uptake and parathyroid analysis | 5
| - Total body knitting | 5
- SPECT viewers for brain, liver, bone
- GI bleeding
7.4 | Software packages should include: | 5
| - Manual image registration tools | 5
| - Export of images in graphic formats (JPEG, BitMap, AVI) | 5
| - Microsoft Office software | 5
7.5 | Please describe any reconstruction algorithms that improve resolution in SPECT and in Planar imaging and quote as an option. Note improved resolution to below 5mm in SPECT and Below 5.5mm in Planar imaging is preferred | 5
7.6 | SPECT CT
Must have the following:
image processing and analysis
image archiving-including auto archiving
image networking-including auto transfer
Auto view of newly reconstructed images
DVD writer
32 GB or better RAM for the main system
applications for dose reduction of paediatric patients must be available
Reconstruction must be multi modality (SPECT and CT)
7.6 | Please state if customised image processing protocols can be created | 5
Full DICOM 3.0 compatibility (send/receive, print, archiving, query/retrieve, work list) with RIS/HIS and PACS connectivity
Allow fusion with PET, MR and/or CT data from external PACS access
8 | Presentation of Results
8.1 | Color laser printer for A4 format on paper | #
8.2 | Data export/import and DICOM 3.0 format (append Dicom Conformance Statement) | #
9 | Accessories
9.1 | Head and Arm holders for Whole body imaging to be included, cardiac Spect and Brain Spect | #
9.2 | Cardiac gating device | #
9.3 | QC accessories with their mounting bracket where necessasy (as per user requirements) | #
9.3.1 | 4 bar phantom (1 in the department) | #
9.3.2 | Cobalt 57 source ( 10mCi) (1 in department) | #
9.3.3 | Jaszczak SPECT
9.3.4 | Phantom CTDI phantom
9.3.5 | CT ionization chamber with appropriate eletrometer
9.3.6 | Catphan phantom
10 | Transmission Imaging
10.1 | Spatial resolution [lp/cm] | 5
10.2 | Low contrast resolution (mm) | 5
10.3 | Transaxial field of view (cm) | 5
10.4 | Axial field of view (coverage) [cm] | 5
10.5 | Number of slices with 1 tube rotation | 5
10.6 | Slice thickness (range) [mm] | 5
10.7 | Separation between SPECT and CT planes | 5
10.8 | Plan CT of NM P-scope | 5
10.9 | Breathing technique (tidal or breathhold) | 5
11 | CT scanner
11.1 | Be integrated, spiral, fully diagnostic, multi-slice, with 16 slices;
11.2 | Provide CT data collection in support of both SPECT/CT fusion for multi-modality imaging and validated CT- AC SPECT;
11.3 | 3D CT dose modulation;
11.4 | The possibility of sub 1 mm slice thicknesses in Axial Mode;
11.5 | The minimum CT slice thickness during spiral acquisition mode shall be < 1 mm.;
11.6 | Patient port diameter/aperture of at least 70 cm;
11.7 | FOV of at least 50 cm;
11.8 | Maximum voltage of at least 130kV;
11.9 | Tube current in range 20-300 mA
11.10. | Be capable of automatically matching the CT slice thickness to the SPECT slice thickness for accurate image fusion and attenuation correction
12 | X-RAY TUBE
12.1 | Anode heat capacity must be a minimum of 3 MHU. ( State anode heat capacity, higher values preferred )
12.2 | State tube replacement pricing as well as the exchange rate.
12.3 | The tube life must be guaranteed for a minimum of 3 years.
12.4 | State the maximum anode heat dissipation rate, assuming a cold start. Higher rate preferred.
13 | DETECTOR SYSTEM:
13.1 | State the detector sampling rate
13.2 | All service phantoms to be included.
13.3 | Detector geometric efficiency must be at least 80% for all modes of operation and have 99% absorption efficiency.
B - General Issues
14 | Safety Standards of the System
14.1 | The Unit should be CE marked | #
14.2 | The Unit should be manufacture according to FDA 1020.30 regulations | #
14.3 | The Unit should be built in Compliance to IEC 601-01 | #
15 | IT Integration
15.1 | The system must have the following DICOM compatibility: | #
15.1.1 | DICOM print class | #
15.1.2 | DICOM storage class | #
15.1.3 | DICOM query/retrieve | #
15.1.4 | DICOM work list | #
15.2 | The system must comply to the following IHE profiles (Integrating the Healthcare Enterprise)
15.2.1 | Scheduled Workflow (SWF) | 5
15.2.2 | Patient Information Reconciliation (PIR) | 5
15.2.3 | Simple image and Numeric Reports (SINR) | 5
15.2.4 | A safeguard method that protects the digital modality from potential malicious software attacks within the hospital network must be included. Describe method | 5
16 | POWER SUPPLY
16.1 | The bidder must install adequate electrical power supply for the optimal functionality of the equipment. | #
16.2 | The mains cable of the unit being quoted for must be the hospital grade type and it must be of an adequate length (minimum length of three (3) meters). | #
16.3 | The power input must be 220-240V, 50Hz AC. Bidders must ensure that the product quoted for is fitted appropriately with a 15 Amp SABS approved mains plug that is held together by two screws. | 5
16.4 | A 3 phase in / out power supply must be offered. State details.
16.5 | The equipment tendered for must not overload and trip the hospital power in the area where it is installed. Resettable overcurrent breaker/s must be fitted for protection. State details.
16.6 | All available online Uninterrupted Power Supply (UPS system) and line voltage regulators must be included for the system with a support time of 50 to 60 minutes. Attach brochure.
16.7 | The lifetime of UPS must be at least 15 years.
16.8 | An Uninterrupted Power Supply (UPS system) conforming to IS-302 must be offered for the computer system.
16.9 | The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade type and it must be a minimum length of three (3) meters.
16.10. | The mains cable of the unit tendered for must be SABS colour coded. State details.
16.11 | The equipment quoted for must be protected against electromagnetic interference. State details of measures that will be taken to protect the unit against electromagnetic interference.
16.12 | Maintenance free batteries must be offered.

## PET CT 16 slices stand alo
PET CT (16 slices) Stand Alone
Item | Technical Specifications
1 | Gantry and System
1.1 | Height between 160 cm to 210cm
1.2 | Width between 190 cm to 240cm
1.3 | Depth between 100 cm to 150 cm
1.4 | Patient port atleast 70 cm
1.5 | Total system weight must be from 3 kg to 5 kg
2 | Patient Handling System
2.1 | Minimum patient load of 200 kg
2.2 | Max co-scan range of atleast 180 cm and above
2.3 | Length between 350 cm to 390 cm
2.4 | Width between 30 cm to 60 cm
3 | PET Acquisition
3.1 | PET Service Tools
3.2 | 3D Acquisition
3.3 | Online histograming (static beds)
3.4 | Automated multi-bed acquisitions
3.5 | Acquisition with offline histogramming supports dynamic and gated images
4 | PET Reconstruction
4.1 | Asynchronous processing/reconstruction
4.1 | 3D image-based scatter correction
4.1 | Metal artifact reduction for cardiac imaging
4.1 | 3D iterative reconstruction
4.1 | ToF
4.1 | FORE + 2D Backprojection
4.1 | Rb-82 Prompts Gamma Correction
4.1 | Reconstruction of gated datasets
4.1 | Reconstruction of dynamic datasets
5 | PET Detector Assembly
5.1 | Detector material must be LSO
5.2 | Detector element dimension of a minimum 4 x 4 x 20 mm
5.3 | A minimum of 160 detector elements per block
5.4 | 4 Photomultiplier tubes (PMTs) per block
5.5 | Detector ring diameter between 820 to 860 mm
5.6 | 48 Blocks per detector ring
5.7 | Transaxial FoV of atleast 70 cm
5.8 | 2 mm Plane spacing
5.9 | Atleast 620 and above detectors elements per ring
5.10. | Atleast 35 and above detector element rings
5.11 | 144 total detector blocks
5.12 | Total number of detector elements minimum of 24 336 (4-Ring TrueV
5.13 | Axial FoV minimum of 160 (4-Ring TrueV)
5.14 | Atleast 75 and above image planes
6 | PET NEMA 2012 Performance
6.1 | Scatter fraction (peak NECR) a minimum 35%
6.2 | Scatter faction (low activity conc.) a minimum 30%
6.3 | Count rate accuracy (Δr mean bias ≤22 kBq/cc) a minimum ±4%
6.4 | Count rate accuracy (Δr mean bias @ peak NEC) a minimum ±4%
7 | PET Data Acquisition/Processing
7.1 | Coincidence window of atleast 3.8 nsec
7.2 | System energy resolution of ≤12% FWHM
7.3 | System time resolution (ToF) of 500 or better psec typical
7.4 | Acquisition modes (static multi-bed, dynamic FlowMotion of up to 20 and above pases and variable speed per bed region)
8 | PET Reconstruction Times
8.1 | Reconstruction time per bed (200 x 200)
8.2 | ToF 30 sec
8.3 | UltraHD•PET 35 sec
9 | Spatial Resolution–Axial
9.1 | Standard Processing (256x256)
9.2 | FWHM @ 1 cm (mm) less than 5.5 mm
9.3 | FWHM @ 10 cm (mm) minimum 6 mm
9.4 | HI-REZProcessing (400x400)
9.5 | FWHM @ 1 cm (mm) less than 5 mm
9.6 | FWHM @ 10 cm (mm) minimum 6 mm
10 | Spatial Resolution–Transverse
10.1 | Standard Processing (256x256)
10.2 | FWHM @ 1 cm (mm) less than 5.5 mm
10.3 | FWHM @ 10 cm (mm) minimum 6 mm
10.4 | HI-REZProcessing (400x400)
10.5 | FWHM @ 1 cm (mm) less than 5 mm
10.6 | FWHM @ 10 cm (mm) minimum 5 mm
10.7 | Sensitivity (cps/kBq) less than 5
10.8 | Peak NEC rate (kcps) 107 @ ≤30 kBq/cc
10.9 | Peak trues rate (kcps) 380 @ ≤46 kBq/cc
11 | Software
11.1 | Time of Flight (ToF)
11.2 | CT extended FoV up to 80 cm
11.3 | PET offline reconstruction
11.4 | Easy and intuitive scan parameter setting (FAST Scan Assistant)
11.5 | Direct scan parameter adjustment (FAST Adjust)
11.6 | Patient Dose management (CT)-(CARE kV; CARE Profile;CARE Dashboard;CARE Child; CARE Filter; CARE Topo; CARE Dose4D™; CARE Bolus CT)
11.7 | Workstream 4D (3D-Recon)- 4D workflow with direct generation of axial, sagittal, coronal, or double-oblique images from standard scanning protocols
11.8 | Visualization, volumetric analysis, and registration/fusion of PET·CT studies (syngo TrueD Basic)
11.9 | Volume Rendering Technique (syngo VRT)
11.10. | Volume Calculation (syngo Volume Calculation)
11.11 | Surface Shaded Display-syngo 3D SSD (Surface Shaded Display)
11.12 | Archiving and Network (syngo Archiving and Network)
11.13 | Dynamic Evaluation (syngo Dynamic Evaluation)
11.14 | DICOM viewer–Included on each CD; automatically started on the viewer’s PC (DICOM viewer–Included on each CD; automatically started on the viewer’s PC)
12 | CT General Specifications
12.1 | Aperture 70 cm and above
12.2 | Scan field (HD FoV 78 cm (extended FoV) / Diagnostic FoV min 50 cm)
12.3 | Rotation times 1 sec or better
12.4 | 16 slices or more per rotation
12.5 | Temporal resolution-down to 83 ms
12.6 | Multislice Detector
12.7 | Minimum of 24 detector rows
12.8 | Minimum of 20 detector electronic channels (DAS)
12.9 | Minimum of 14 500 detector elements
12.10. | 1 400 Total channels/slice
12.11 | Number of projections up to 2320/360°
12.12 | Generator maximum power minimum 50kW-80 kW
12.13 | X-ray tube
Tube current 20-666 mA
Tube voltages 70, 80, 100, 120, 140 kV
Tube focal spot 0.7 x 0.7 mm/7°, 0.9 x 1.1 mm/7°
13 | CT Reconstruction System
13.1 | Reconstruction speed minimum of 16-20 images/s
14 | CT Topogram
14.1 | Length of topogram field minimum of 128-2,200 mm
14.2 | Topogram scan times of 1.5-20 sec
15 | CT Sequence Acquisition
15.1 | Reconstructed slice widths in sequence acquisition (0.6, 0.75, 1, 1.2, 1.5, 2, 2.4, 3, 4, 4.8, 5, 6, 7, 7.2, 8, 9, 10,12, 14.4, 15 mm)
15.2 | Scan times (360° full scan) 1 sec or better
15.3 | Scan times (260° partial scan) 0.24, 0.36, 0.72 sec
15.4 | Atleast 90 uninterrupted scans per range
15.5 | Atleast 30 ranges per protocol
15.6 | Scan cycle time 0.75-60 ses (±10%)
16 | CT Spiral Acquisition
16.1 | Reconstructed slice widths in spiral acquisition (0.6, 0.75, 1, 1.5, 2, 3, 4, 5, 6, 7, 8, 10 mm)
16.2 | Scan times (360° full scan) 5 sec or better
16.3 | Slice increment 0.1-10 mm
16.4 | Pitch factor 0.35 -1.5, 0.15, 0.09
16.5 | Spiral scan time max. atleast 100 sec or better
16.6 | Minimum of 30 ranges per protocol
17 | CT Reconstruction
17.1 | Real-time display supported during spiral acquisition
17.2 | Reconstruction field 5-50 cm, 5-78 cm (HD FoV)
17.3 | Reconstruction matrix 512 x 512
17.4 | HU scale 1024 to 3071
17.5 | Extended HU scale 1024 to 3071
17.6 | Reconstruction time min 16-20 images/s
18 | CT Image Quality
18.1 | Low-contrast resolution (CATPHAN 20 cm) / Technique–10 mm, 120 kV
18.2 | Sequence
18.2.1 | Object size 5 mm
18.2.2 | Contrast difference 3 HU and above
18.2.3 | Dose at surface 10.7 mGy@180 mAs
18.3 | Spiral
18.3.1 | Object size 5 mm
18.3.2 | Contrast difference 3 HU and above
18.3.3 | Dose at surface 13.1 mGy@180 mAs
18.4 | High-contrast Resolution
18.4.1 | Technique (160 mA,/120 kV,/0.5 s,/2.4 mm/(12x1.2mm)/UHR/Kernel H70h)
18.4.2 | Plane 0% MTF (±10%) 17.4 lp/cm
18.4.3 | Plane 2% MTF (±10%) 16.4 lp/cm
18.5 | Scanner Room Environment
18.5.1 | Temperature range 20-26° C (68-79°F)
18.5.2 | Relative air humidity without condensation (recommended level 30-70%)
18.5.3 | HVAC requirement (heat dissipation during operation) 2.0 kW (6,667 BTU/hr)
18.5.4 | Air pressure 750-1,060 mbar
18.5.5 | Variation in floor levelness not to exceed 12.7 mm (0.5 in) over 5.2 m (17 ft)
18.5.6 | Minimum room size 4.6 m x 7.3 m (15 ft x 24 ft)–(requires angling gantry and PHS)
18.5.7 | Recommended room size ≥5.0 m (16.5 ft) x ≥7.9 m (26 ft)
19 | Utility Room Environment
19.1 | Temperature range 18-30° C (64-86° F)
19.2 | Relative air humidity without condensation (recommended level 30-70%)
19.3 | HVAC requirement removal of 5.0 kW (17,100 BTU/hr)
19.4 | Gantry heat exchanger–max. 15 kW to water cooling environment
19.5 | Minimum room size 3.2 m (10.5 ft) x 2.1 m (7 ft)
20 | Operator’s Room Environment
20.1 | Temperature range 20-26° C (68-79° F)
20.2 | Relative air humidity without condensation (recommended level 30-70%)
20.3 | HVAC requirement removal of 3.1 kW (10,600 BTU/hr)
20.4 | Recommended room size ≥3.2 m (10.5 ft) x ≥4 m (13 ft)
21 | System Power Requirements
21.1 | Nominal voltage: 3/N~ 380-480 V (±10%) in 20 V steps
21.2 | Nominal line frequency: 50; 60 Hz (+2 Hz)
21.3 | Line impedance: 80-140 mOhm (dependent on voltage)
21.4 | Maximal power consumption: 150 kVA
22 | Software Configuration–Optional
22.1 | HD•PET
22.2 | syngo Image Fusion
22.3 | syngo Security Package
22.4 | HD•Chest
22.5 | Cardiac PET
22.6 | Cardiac CT
22.7 | Cardiac PET/CT
22.8 | Dynamic PET
22.9 | SMART Auto Cardiac Registration
22.10. | SMART Neuro AC
22.11 | NEMA Test Suite
22.12 | CT SAFIRE
22.13 | CT single-source dual-energy scan mode
23 | Patient Handling System: Continuous Bed Motion
FlowMotion
Multiparametric PET
24 | Auto QC: QualityGuard
25 | Respiratory Motion Correction
25.1 | Oncofreeze
25.2 | PET Respiratory Gating
25.3 | CT Respiratory
25.4 | Gating/Triggering
25.5 | PET + CT Respiratory Gating
25.6 | Cardiac Gating: Cardiofreeze
26 | Upgradable to latest SiPM System: Forklift Upgrade

## PET CT 64 slice
PET CT (64 slices)
Item | Technical Specifications
1 | PHYSICAL FEATURES
1.1 | A full ring PET system plus integrated CT scanner offering 64 acquired CT slices per rotation including a patient bed and single combined PET/ CT acquisition console.
2 | Gantry
2.1 | PET Gantry
2.1.1 | For minimized count rate losses and high detection efficiency the detector crystal should consist of a lutetium based scintillator.
2.1.2 | The proposed PET scanner must be capable of acquiring at least 81 transverse cross-sectional slices simultaneously without undergoing any axial motion. (TrueV, 109 slices)
2.1.3 | The PET detector elements must be isotropic and have axial and transaxial dimensions less than or equal to 4 mm
2.1.4 | The geometric axial field of view (FOV) as measured from the outer edges of the crystals in the axial direction must be >16.2 cm. (TrueV, 22.1)
2.1.5 | The Gantry must have a 78 cm or greater aperture throughout the entire length of the tunnel for patient comfort, ease of positioning and RTP support.
2.1.6 | Detector design must employ a modular design for ease of service and field upgradability to a different size field of view.
2.1.7 | The (center-to-center) separation between slices acquired simultaneously without any axial motion must not exceed 2.0 mm.
2.1.8 | The scanner must have a Coincidence Timing Window no greater than 4.1 nsec.
2.1.9 | The scanner must have an Energy Window with lower level discriminator greater than or equal to 435 keV.
2.1.10 | The number of pixels (detector elements) per PMT must be less than 45 pixels/PMT.
2.1.11 | The system should contain 24,336 or more individual crystal pixels (32,448 with TrueV).
2.1.12 | The scanner must not use gantry detector wobble or axial shift of the patient to obtain the specified transaxial and axial specifications.
2.2 | CT Gantry
2.2.1 | The CT should support spiral scanning with a maximum spiral acquisition of 100 seconds or greater.
2.2.2 | The transverse field-of-view for CT used during PET attenuation correction should extend all the way to the edges of the gantry and offer uniformity to within 50 HU when imaging greater than 50 cm FOV.
2.2.3 | CT tube must cool at a rate of 5 MHU/min or greater
2.2.4 | The CT should provide at least 64 acquired slices
2.2.5 | The CT should have a 360 degree rotation speed of at least 0.30 seconds.
2.2.6 | The CT must have spiral (pitch) that is freely variable without a change in image quality.
2.2.7 | CT image quality must provide <0.4 mm resolution scanning for all clinical scans, and < 0.25 mm resolution in hi resolution mode.
2.2.8 | The CT system must be capable of 30 LP/cm
2.2.9 | The CT subsystem will have a dynamic collimator dose shield reduction technique to reduce pre and post scan dose to the patient
2.2.10 | The CT must have on the fly, real time dose modulation without additional radiogram procedures.
2.2.11 | The CT must be capable to scan at 70 kV for reduced dose to pediatrics
2.2.12 | The CT generator power must be at least 100 kW.
2.2.13 | The gantry thickness should not exceed 136 cm.
2.2.14 | The PET/CT gantry must require less than 7000 BTU/hr for HVAC room cooling.
2.2.15 | The gantry should have a closed design without vents/fans to reduce ambient noise and temperature restrictions for the scan room.
3 | Patient Bed
3.1 | The horizontal motion of the patient bed must be electrically motorized and computer-controlled with an independent operator control option as well. Operator controls accessible from both sides of the patient must be provided for both horizontal and vertical control.
3.2 | The patient bed must support the use of continuous bed motion during acquisitions to eliminate PET bed stops.
3.3 | A digital readout of the horizontal and vertical position of the bed must exist and must be located near the operator controls for the bed to provide ease in positioning.
3.4 | The patient bed must provide a constant deflection throughout the scanner axial field-of-view. The bed should support a maximum patient weight of 230 kg or greater at full scan coverage.
3.5 | The lowest vertical position should be 65 cm or less from the floor to the top of the side edge of the patient pallet.
3.6 | The horizontal travel of the bed should enable a co-scan axial field-of-view of 190 cm or greater.
3.7 | A method of preventing accidental vertical motion by an extraneous activation of the operator's controls must be provided to prevent the hazard of patient injury.
3.8 | A method of stopping any computer controlled or accidental motion must be provided and must be located at all operator positions of the system.
3.9 | A mechanism for manually removing the patient bed from the gantry bypassing the electric motor control must be supplied for emergency situations and described fully in the proposal.
3.10. | The patient bed should support the use of a flat RTP pallet (width > 50 cm)
3.11 | The bed should be usable for Radiotherapy threatment planning
4 | SOFTWARE
4.1 | Data Acquisition:
4.1.1 | The software must exist to acquire data in 3D (e.g., as a volume data set) modes. Reconstruction by the FORE and OSEM iterative algorithms must be included.
4.1.2 | The acquisition program should support pre-programmed scan protocols with acquisition and reconstruction parameters and patient information with simple, dynamic editing of parameters. These parameters would include all information necessary to acquire data on the PET scanner (e.g., scan duration, patient information, bed motion), as well as information necessary for reconstruction, which would be carried through to the reconstruction step.
4.1.3 | Continuous bed motion during a whole body PET scan is required.
4.1.4 | Whole body with multi-bed and/or continuous bed motion acquisition protocols should be performed such that operator interaction is not required to complete the scan once it has been initiated.
4.1.5 | As a part of static acquisition, a total count rate display should be updated at least every 5 sec.
4.1.6 | The system must include a single, integrated console, which controls acquisition for both the CT and PET.
4.1.7 | The software must support DICOM (PET and CT) export, import, and well as filming and archival.
4.1.8 | Online data storage capacity should be a minimum of 1 TB
4.1.9 | Metal artifact reduction must be included for the CT and the new images with reduced artifact, must be used for attenuation correction
4.1.10 | A raw data iterative CT reconstruction technique must be included with the CT system
4.1.11 | The system must have the ability to cover 9 cm of perfusion data during Perfusion CT studies
4.1.12 | The system must support dual energy CT.
5 | Reconstruction and Reformatting
5.1 | Static image reconstruction should commence automatically upon completion of a bed position or acquiring enough information during CBM to begin parallel with the acquisition, using the reconstruction parameters defined at acquisition.
5.2 | If the system cannot keep pace with the acquisition, there should be a method for selecting a specific scan to be reconstructed immediately.
5.3 | In that case, other scans should be reconstructed immediately after the requested reconstruction has finished.
5.4 | The methods for randoms estimation and subtraction must be by measuring randoms rather than calculations from measured singles.
5.5 | The system must include algorithms for utilizing the CT data to generate attenuation correction for the PET data.
5.6 | A PET frame (200x200) should be reconstructed in < 30 sec/bed with all data corrections, 3D AW-OSEM
5.7 | 512*512 CT images should be reconstructed at a rate of at least 40 images per second
5.8 | The user should have the capability to specify the transverse FOV for reconstruction.
5.9 | The reconstruction program should support reconstruction in image sizes of at least 360x360 for PET and 512x512 for CT.
5.10. | If the pixel and reconstruction matrix sizes are selected such that the reconstructed area is smaller than the transaxial FOV, it should be possible to select where in the transaxial FOV the reconstruction is centered.
5.11 | A scatter correction technique that is spatially variant and adjusts for patient geometry must be included. Scatter correction must be provided based on scan of the actual patient whose scan is being corrected and processed automatically in an iterative loop
5.12 | A fully automated, one-button protocol for whole body oncology studies should be provided, having the following features:
5.12.1 | permits multi-bed or CBM scans to proceed in an efficient way without operator intervention;
5.12.2 | processing software which proceeds automatically without operator intervention (processes include segmented attenuation correction, iterative reconstruction, scatter correction, normalization, smoothing, filtering, etc.)
5.12.3 | display of cine whole body projection views as well as coronal, transverse, and sagittal volume slices including multiple reconstruction jobs, auto archiving and networking.
5.13 | The system must be capable of performing a Full 8 gate cardiac perfusion/MBF acquisition with complete reconstruction of all the data necessary for the rest and stress portion of the study in less than 23 minutes from start to finish, barring no pharmaceutical stress issues.
5.14 | Measured Point Spread Functions of the system must be utilized in the iterative reconstruction algorithm.
5.15 | Respiratory gating must be included which can automatically find and automatically reconstruct a motion free image from the gated PET data.
5.16 | Cardiac gating must be included which can correct for respiratory motion using a deviceless method
5.17 | Gating inputs must be located on the patient bed to reduce hanging wires for safety.
5.18 | Gating device must be interfaced with PET/CT system eliminating the need to access the gating laptop for the UI in order to acquire data.
5.19 | A whole body scan must be fully capable of integrating respiratory gated acquisition/data into both the acquisition and display
5.20. | PET and CT respiratory gated data can be automatically reconstructed in a phased matched manner.
6 | Display and Analysis
6.1 | It is required that the acquisition, reconstruction and image display and analysis software to operate in Windows 10 or greater. A description of the software environment and flexibility should be given in the proposal.
6.2 | The system must be able to display CTDIw and DLP
6.3 | Region of interest (ROI) capability for circular, rectangular, and irregular ROI's must be provided.
6.4 | A patient database with such entries as patient name, patient weight, injected activity and time, identification number, protocol (disease), date of scan, data filename is required and should be discussed in the proposal.
6.5 | Utility programs for access to database information should also be discussed.
6.6 | The Patient Database should support DICOM worklist functionality.
6.7 | A visualization application that allows the comparison of 3 fused datasets (e.g. pre- and post- therapy PET/CTs) on the same screen with all gating and quantification capabilities.
7 | Quality Control
7.1 | PET QC can be set up and run automatically overnight (optional).
7.2 | A procedure to calibrate the scanner to a well counter must be provided with the system and described in the proposal.
7.3 | The QC phantoms must be part of the unit (optional).
7.4 | Daily quality control procedures are well defined and automated, and all necessary sources are included with the system
7.5 | An apparatus for reproducible positioning is required and should be described in the proposal.
8 | PET Performance Specifications
8.1 | Transverse Resolution
8.2 | Transaxial resolution (in-plane resolution) specification for 3D mode must be equal to or better than 4.5 mm FWHM at 1 cm from the central axis of the gantry.
8.3 | Axial Resolution
8.4 | Axial resolution specification for 3D mode must be equal to or better than 4.7 mm FWHM at 1 cm from the central axis of the gantry.
8.5 | Sensitivity
8.6 | System sensitivity specification must exceed 10.2 cps/kBq at 435 keV LLD.
8.7 | Count Rate
8.8 | The count rate peak NECR performance must achieve at least 180 kcps
8.9 | System must offer a 3D PSF PET reconstruction technique
8.10. | Time of flight capability must be available with better than 550 ps timing resolution at all clinical count rates
9 | POWER SUPPLY
9.1 | The bidder must install adequate electrical power supply for the optimal functionality of the equipment.
9.2 | The mains cable of the unit being quoted for must be the hospital grade type and it must be of an adequate length (minimum length of three (3) meters).
9.3 | The power input must be 220-240V, 50Hz AC. Bidders must ensure that the product quoted for is fitted appropriately with a 15 Amp SABS approved mains plug that is held together by two screws.
9.4 | A 3 phase in / out power supply must be offered. State details.
9.5 | The equipment tendered for must not overload and trip the hospital power in the area where it is installed. Resettable overcurrent breaker/s must be fitted for protection. State details.
9.6 | All available online Uninterrupted Power Supply (UPS system) and line voltage regulators must be included for the system with a support time of 50 to 60 minutes. Attach brochure.
9.7 | The lifetime of UPS must be at least 15 years.
9.8 | An Uninterrupted Power Supply (UPS system) conforming to IS-302 must be offered for the computer system.
9.9 | The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade type and it must be a minimum length of three (3) meters.
9.10. | The mains cable of the unit tendered for must be SABS colour coded. State details.
9.11 | The equipment quoted for must be protected against electromagnetic interference. State details of measures that will be taken to protect the unit against electromagnetic interference.
9.12 | Maintenance free batteries must be offered.
10 | Miscellaneous
10.1 | Software for archival and retrieval of scan and image data must be provided and described in the proposal.
10.2 | The system shall support the use of USB drives for fast easy storage

## Automatic injector (PETinfusion
Automatic Injector (PET Infusion System)
ITEM | Technical Specifications
1 | Automatic injector (PET infusion system)
1.1 | Weight approximately 330 kg
1.2 | Dimensions approximately≤ 64 cm wide, ≤ 81 cm deep, ≤ 90 cm high.
1.3 | Maximum bulk radioactivity 30GBq (811 mCi)
1.4 | maximun radiopharmaceutical volume 30mls
1.5 | Dose range1-30mci
1.6 | dose accuracy ≤ 5% of the measured dose
1.7 | system radiation profile ≤ 0.014mSv/hr at 30.5cm from the surface of the cart at maximum activity.
1.8 | flow rate ≤ 1ml/sec
1.9 | frequency 50/60 Hz
1.10. | power shall be ≤ 300VA
1.11 | VOLTAGE 110-240VAC
1.12 | maximum radiopharmaceutical concentration100mCi/ml
1.13 | Energy range: 25keV - 3 MeV
1.14 | POWER SUPPLY
1.14.1 | The bidder must install adequate electrical power supply for the optimal functionality of the equipment.
1.14.2 | The mains cable of the unit being quoted for must be the hospital grade type and it must be of an adequate length (minimum length of three (3) meters).
1.14.3 | The power input must be 220-240V, 50Hz AC. Bidders must ensure that the product quoted for is fitted appropriately with a 15 Amp SABS approved mains plug that is held together by two screws.
1.14.4 | A 3 phase in / out power supply must be offered. State details.
1.14.5 | The equipment tendered for must not overload and trip the hospital power in the area where it is installed. Resettable overcurrent breaker/s must be fitted for protection. State details.
1.14.6 | All available online Uninterrupted Power Supply (UPS system) and line voltage regulators must be included for the system with a support time of 50 to 60 minutes. Attach brochure.
1.14.7 | The lifetime of UPS must be at least 15 years.
1.14.8 | An Uninterrupted Power Supply (UPS system) conforming to IS-302 must be offered for the computer system.
1.14.9 | The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade type and it must be a minimum length of three (3) meters.
1.14.10 | The mains cable of the unit tendered for must be SABS colour coded. State details.
1.14.11 | The equipment quoted for must be protected against electromagnetic interference. State details of measures that will be taken to protect the unit against electromagnetic interference.
1.14.12 | Maintenance free batteries must be offered.

## Gamma Analyser
Gamma Analyser
Item | Technical Specifications
1.1 | DETECTOR
1.1.1 | The detector should consist of NaI crystal
1.1.1 | Single Photomultiplier tube
1.1.2 | The detector should provide optimum geometry for counting samples for GFR and Red cell mass and plasma volume
1.1.3 | Must be able to count I-131, I-123, Fe 59, Tc-99m, Cr-51 as used routinely
1.2 | DETECTOR SHIELDING
1.2.1 | Optimum Pb shield design with a 4-pi shielding geometry
1.2.2 | Pb shutter for shielding of cosmic radiation
1.2.3 | Copper shielding around chamber to eliminate x-rays from Pb shield
1.2.4 | There must be minimal interference from samples on the conveyor
1.3 | PERFORMANCE SPECIFICATIONS
1.3.1 | Should be able to detect gamma emmitters up to 2000keV
1.3.2 | State counting efficiency for I-129, Cs-137, Cr-51
1.3.3 | State resolution with Cs-137 and I-129
1.3.4 | State count rates in CPM
1.4 | SAMPLE VIALS AND RACKS
1.4.1 | Must be able to count from plastic or glass tubes used for drawing of blood samples
1.4.2 | Different studies coded with ID code plates and therefore be identified and processed accordingly
1.4.3 | State number of vial racks provided
1.4.4 | State number of ID codes available on your system
1.5 | INSTRUMENT OPERATION AND SOFTWARE
1.5.1 | Must be MS window based
1.5.2 | Must include software to count for GFR and RCMPV speciments
1.5.3 | Instruction manuals must be supplied for programmed software for counting of speciments and calculations thereafter
1.5.4 | MUST include printer – to display counts and final results
1.5.5 | Results must be printed out for GFR etc. preferably on the A5 format
1.5.6 | All files to be dicom compatible