## X-Ray B-Floating 80kW
Digital X-RAY UNIT (80kW generator)
Description
ITEM
1 | The said Unit must be fully automatic, digital, radiographic examination and evaluation workstation based on two fixed/wireless detectors technology for high image dynamics with excellent signal/noise ratio is required.
1.1 | The bid price quoted must include the supply, installation, demonstration and commissioning of the equipment as well as the building alterations that may be required for the successful installation of the unit/s.
1.2 | State make and model of main unit offered.
1.3 | The bidders are requested to structure their quote so that it reflects the price of the unit separately from any building alterations that are required. The price schedule at the end of this specification must be completed.
1.4 | The unit must be suitable for general and trauma radiography.
2 | GENERATOR:
2.1 | The unit should be a high frequency type generator.
2.2 | The output shall be 80kW. State maximum milliamps at different kilovolts.
2.3 | The unit must incorporate an electronic timer with minimum exposure time of 1mS and electronic switching
2.4 | The unit must include automatic overload protection
2.5 | The unit must have automatic mains compensation.
2.6 | A milliamp meter or indicator must be provided.
2.7 | The kV, mA and time must be independently selectable and digitally displayed.
2.8 | Bidders to state the range of the exposure timer. Acceptable range 0.001 sec – 6 sec (1 ms – 6sec).
2.9 | Bidders to state kV and mA ranges for radiography. The kV range to be at least 40 to 150 kV and mA range at least 10 to 630mA. The kV selection must be in steps of 1 and must be a single control.
2.10. | The unit offered must be programmable. Bidders to supply details of programming facility and the number of programs.
2.11 | The offered unit must include automatic exposure control.
2.12 | Tube overload protection is required
3 | X-Ray TUBE:
3.1 | The rotating anode tube must match the output of the generator.
3.2.1 | The maximum tube kV must be at least 150 kV.
3.2.2 | The unit must incorporate an electronic timer with minimum exposure time of 1mS and electronic switching
3.2.3 | The maximum tube mA must be at least 500 mA
3.2.4 | The speed of anode rotation must be at least 8000rpm
3.2.5 | Anode heat storage capacity must be at least 300kHU
3.2.6 | The tube housing heat storage capacity must be more than 400kHU
3.2.7 | The size of small focus must be 0,6 mm
3.2.8 | The maximum tube mA must be at least 500 mA
3.2.9 | The speed of anode rotation must be at least 8000rpm
3.2.10 | The size of large focus must be 1,2 mm
3.2.11 | Inherent filtration of at least 0.8mmAl
3.2.12 | Total filtration must exceed 2.5mmAl.
3.2.13 | Anode angle must be large enough to provide a uniform beam over a field of 43cm x 43cm at a FFD of 100cm.
3.3 | HIGH TENSION CABLE: One pair of high tension cable of suitable length must be supplied. The cable length must not restrict free movement of the tube for all radiographic procedures.
3.4 | The power supply must be a three phase voltage 380-400 V at 50Hz
4 | X-RAY TUBE WITH MULTILEAF COLLIMATOR
4.1 | Nominal voltage range (40 - 150kV)
4.2 | State the nominal power
4.3 | Focal spot parameter according to International Electrotechnical Commission (IEC) 336
4.4 | Small focus 0.6mm
4.5 | Large focus 1.2mm
4.6 | Anode angle 12o
4.8 | State the speed of rotating anode
4.9 | Heat storage capacity of the anode should not be less than 350kHU
4.10 | Heat storage capacity of the X-ray tube should not be less than 400kHU
4.11 | Radiation absorption (according to IEC 601-1-3)
4.12 | There must be a measuring tape on the X-Ray tube
4.13 | Light localizer on/off switching with timer to terminate the light after 30 seconds
4.14 | The table must be able to perform tomography. Please state the angles offered and the speeds related to each angle.
5 | CEILING MOUNTED TUBE SUPPORT
5.1 | The tube column must be ceiling mounted
5.2 | The X-ray tube assembly rotation must be 360° about the vertical and horizontal axis
5.3 | Tube movements:
5.3.1 | Longitudinal: at least 150 cm from the iso centre
5.3.2 | Transverse: at least 100 cm from the iso centre
5.3.3 | Vertical: Tube centre should be able to move down to the lowest of 30cm above the floor.
5.4 | Multiplanner movements controlled by electric lock
5.5 | All movements to be counter balanced
6 | CEILING SUSPENSION:
6.1 | The ceiling suspension must provide coverage of an area of approximately 3 X 4 meters.
6.2 | Stops must be provided for centering to table and vertical Bucky.
6.3 | Bidders to indicate focal film distance range for vertical and table Bucky.
6.4 | A safety device for possible failure of the counter-balance system must be incorporated.
6.5 | The unit must be able to perform horizontal beam radiography on both side of the table.
6.6 | Bidders to state:
6.6.1 | Tube rotation and movement around vertical axis
6.6.2 | Tube rotation around horizontal axis and
6.6.3 | Range of vertical movement.
7 | BUCKY TABLE:
7.1 | The bidder to state the make and model of the Bucky table offered.
7.2 | The Bucky table must have a flat, four way floating top.
7.3 | The table must be electronically height controlled. Foot switches must form part of the table.
7.4 | State the variable height of the table offered.
7.5 | The dimensions of the table top must have atleast length 210cm and 80 to 90 cm width
7.6 | State the lateral and longitudinal movement of the table top.
7.7 | The table offered must be able to carry patients with a weight of at least 220kg centred and off-centred.
7.8 | State the longitudinal movement of Bucky tray. Bidders must detail whether the distance relates to the outermost end of the Bucky tray, or from centre point to centre point. A data sheet must be included to clarify this point.
7.9 | The table-top movement must be controlled with electromagnetic brakes.
7.10. | The table-top film distance must be approximately 6 centimeters.
7.11 | The Bucky assembly must be fitted with the oscillating secondary radiation grids suitable for direct digital flat panel technology.
7.12 | A compression band with fitting mechanism compatible to the table top must be supplied.
7.13 | A full set of positioning sponges with washable water resistance covers must be supplied with the table.
7.14 | Emergency stop mechanism must be attached to the table
8 | VERTICAL BUCKY
8.1 | State the make and model of vertical Bucky offered.
8.2 | The vertical height should be adequate for a lateral cervical spine projection of a standing patient and therefore maximum height from centre of the Bucky to floor should be at least 180 cm.
8.3 | The Bidder to state range of vertical movement from floor to centre point of the Bucky.
8.4 | Vertical movement of the vertical Bucky must be via an electronic locking/unlocking switch. Analogue locking/unlocking levers or knobs will not be accepted
8.5 | The Bucky assembly must be fitted with the oscillating secondary radiation grids suitable for direct digital flat panel technology and must accommodate a chest radiography at 180cm.
8.6 | Radiography of a patient who is in a seated position on the stretcher or table top must be possible.
8.7 | The Bucky must be mounted centrally on the column with the choice of loading on either left or right.
8.8 | An overhang arm support must be included that is compatible with erect Bucky.
8.9 | An overhang frame for out of Bucky radiography must be included that is compatible with erect Bucky.
9 | DIGITAL FLAT PANEL DETECTOR:
9.1 | The detector technology must be of Cesium Iodide type
9.2 | The exposure formats must be 43cm x 43cm
9.3 | The full vertical rotation must be possible without removing detector for 43 x 43 format
9.4 | The full horizontal rotation must be possible without removing detector for 43 x 43 format
9.5 | The detector exposure matrix: pixels must be minimum 3.6 lp/mm
9.6 | The detector pixel pitch must be minimum 139um
9.7 | The reading speed at maximum format from the exposure to preview on the monitor must not be greater than 8 seconds
9.8 | The digital Flat panels must operate wirelessly
9.9 | The wireless flat panels must be supplied with x2 spare battery packs each that can be charged outside of the flat digital flat panel.
9.1 | The analog to digital conversion must not be less than 12 bits
9.11 | The wireless ditector must come with grids, detector cover and holder
9.12 | The flat panel battery:
9.12.1 | The battery should be able to acquire and transfer at least 90 images or more when fully charged. Details are stated in the supporting documents provided .
9.12.2 | The battery should be able to handle 5 hours of active use and 15 hours standby, when fully charged. Details are stated in the supporting documents provided .
9.12.3 | Charging time for each battery from fully depleted to fully charged should be less than 2 hours. Details are stated in the supporting documents provided.
9.12.4 | The batteries must be interchangeable between the 24 x30 cm and 43 x 43 cm panels.
9.12.5 | The battery charger must be light weight and should weigh less than 2 kg. Details of the weight of the battery charger are stated in the supporting documents provided .
9.12.6 | The battery must be able to operate at ambient temperature range of minumum of 5°C to maximum of 35 °C.
9.12.7 | The battery must be able to operate at ambient humidity of minimum humidity of 20 % to maximum of 85% RH.
9.12.8 | The battery must be able tolerate storage at a temperature range of minumum of -30 to to maximum of 60 degrees celsius.
9.12.9 | The battery must be able tolerate storage at a minimum humidity of 10 to maximum of 85% RH.
9.13 | The Grid
9.13.1 | An encasement with a built in grid that fits the 43cm x 43cm detector must be supplied.
9.13.2 | The grid must have at least:
9.13.3 | Grid ratio 8:1
9.13.4 | 100cm to 180cm focal distance
9.14 | Load capacity
9.14.1 | The uniform load that the detector must be able to withstand should be at least 150kg. Details are stated in the supporting documents provided .
9.14.2 | The local load (on an area of 40 mm diameter) that the detector must be able to withstand should be at least 100kg. Details are stated in the supporting documents provided .
11 | X-RAY EXAMINATION TABLE
11.1 | Table movements must be motorised (all directions)
11.2 | Minimum height (above the Floor) 40-60cm
11.3 | Must have a Floating table top
11.4 | Detector carriage should have a motorized scattered radiation grid.
11.5 | Film Focus Distance = 115cm
11.6 | Table top dimensions of at least 220L X 80W cm.
11.7 | Table top longitudinal movement of at least (50cm) in both direction with minimum patient weight of 220kg.
11.8 | Transversal movement of at least (20cm) in one direction with minimum patient weight of 220kg.
11.9 | Control of the table functions via footswitch attached to the table
11.10 | Electromagnetic brakes to be standard
11.11 | Table top to be of suitable low radiation absorption material
11.12 | Load capacity must be at least 220 kg
11.13 | Must have automatic exposure device
12 | IMAGE WORKSTATION
12.1 | Storage capacity must be minimum 4 tera bytes
12.2 | The system must have minimum of 4 GB memory
12.3 | Minimum 19 inch, maximum 23 inch colour LCD/LED display
12.4 | Must be a touch screen
12.5 | Resolution Minimum 1920x 1080 pixels
13 | IMAGING SOFTWARE
MUST INCLUDE:
13.1 | General Features
13.1.1 | Windows based graphic user interface
13.1.2 | Multi-image display (1x1 / 2x2 / 3x3 / 4x4 )
13.1.3 | Multi-image selection
13.1.4 | Unlimited procedure set ups
13.1.5 | ROI changing and creation feature
13.1.6 | English language support
13.1.7 | Unlimited PACS code ( CPT code )
13.1.8 | Default anatomic program more than 700
13.1.9 | Support DICOM Worklist SCU, DICOM Storage SCU and transfer function
13.1.10 | Support DICOM Multi-transfer function
13.1.11 | All hardware & software relating to DICOM compatibility, including licences, must be supplied.
13.2 | Post processing features
13.2.1 | Noise reduction parameter
13.2.2 | Global contrast parameter
13.2.3 | Detail contrast parameter
13.3 | Image Maintenance (All functions must be supported by the pop-up menu)
13.3.1 | ROI
13.3.2 | Mark
13.3.3 | Horizontal Flip
13.3.4 | Vertical Flip
13.3.5 | Rotate CW
13.3.6 | Rotate CCW
13.3.7 | Inverse (Black or White)
13.3.8 | Text Annotation
13.3.9 | Ruler : Distance tool
13.3.10 | Angle : Angle measurement tool
13.3.11 | Zoom : Image zoom in/out
13.3.12 | Magnify : Image magnify glass window
13.3.13 | Pan : Image panning
13.3.14 | Fit Image : Auto fitting to window size
13.3.15 | Image Cut : Image crop/cut function
13.3.16 | Image Copy : Copy of image in the region of interest(ROI)
13.3.17 | Image Recovery : Recover the original image
13.4 | CD Burning
13.4.1 | DICOMDIR based CDR data generation
13.4.2 | Support CD/DVD Recording
13.4.3 | Include internal DICOM Viewer
13.4.4 | Support multi-study data
13.5 | DICOM Features : DICOM PRINT
13.5.1 | DICOM 3.0 compatible
13.5.2 | Support Print Preview
13.5.3 | Support Film Layout : 1:1 / 1:2 / 2:1 / 2:2
13.5.4 | Support multi-page printing
13.5.5 | Support Real size printing
13.5.6 | Support image swap in layout
13.6 | DICOM Feature : DICOM STORAGE
13.6.1 | DICOM 3.0 compatible
13.6.2 | Support DX/CR modality
13.6.3 | Support the modification of Transfer Syntax
13.6.3.1 | Transfer status
13.6.4 | DICOM 3.0 compatible
13.6.5 | Support DX/CR modality
13.7 | DICOM Feature : WORKLIST
13.7.1 | Support DICOM Modality Worklist Standard
13.7.2 | Support DICOM Query/Retrieve
13.7.3 | Support Search Filter ( ID / Name / Access Number )
13.7.4 | Support Import Filter
13.7.5 | Support DICOM Modality Worklist Standard
13.8 | Overlay Display on image
13.8.1 | Projection description
13.8.2 | Hospital or Institution Name, Date and Time of examination
13.8.3 | Radiographer Name
13.8.4 | Patient Name, Sex and Age
13.8.5 | Machine Name
13.8.6 | kV / mA / Time / mAs
13.8.7 | Feed-back mAs / Feed-back Time for AEC
13.8.8 | EI(Exposure Index)
14 | THE FOLLOWING MUST BE PROVIDED WITH THE SYSTEM INTERFACES AND CONNECTIVITY
14.1 | A "Conformance Statement" for all DICOM interfaces must be included. (the document must be attached)
14.2 | Interface DICOM: "Basic Print"
14.3 | Interface DICOM: "Modality Work-list"
14.4 | Interface DICOM: "Send as Service Class User (SCU)"
14.5 | Interface DICOM: "Receive as Service Class User (SCU)"
14.6 | Interface DICOM: "Retrieve as Service Class User (SCU)"
14.7 | Interface for checking the system status or for loading new
14.8 | Software to all components should be DICOM Interfaced.
14.9 | The modality shall include the DICOM Service Class license Image Storage: SCU and SCP
14.10. | The modality shall include the DICOM Service Class license Query/Retrieve SCU
14.11 | The modality shall include the DICOM Service Class license Print SCU
14.12 | The modality shall include the DICOM Service Class license Storage Commitment SCP
14.13 | The vendor shall include annexures of Proof of DICOM Conformance Statement
14.14 | The succesful bidder must ensure that the configuration between the processed images on the Digital Imaging Workstation and that transferred to the receivers of these images, such as PACS , Laser Printer , CD/DVD are properly calibrated/ reconfigured during installation, service and maintenance,repairs, and future software updates/upgrades.
15 | GENERAL
15.1 | Remote diagnostics feature for fault finding / troubleshooting must be available
15.2 | All computer systems supplied must have UPS protection with at least 20 minutes backup
15.3 | The Unit must link to existing laser printer
15.4 | The Unit must link to existing CD Robot
15.5 | The Unit offered must be able to connect a PACS/ RIS and HIS
15.6 | The Unit must be able to connect to any vendor
17 | RADIATION PROTECTION
17.1 | The system must have the ability to display Exposure Indicator values (EI/ S-values), immediately after exposure.
17.2 | X-ray Warning Notice for display at the X-ray Room entrance (not branded).
17.3 | An advanced Retake / Reject / Repeat analysis software package that includes analysis as well as calculations in percentages, must be provided for applicable systems.
17.4 | The applicable systems must have an integrated Dose Area Product (DAP) meter that includes skin dose equivalent measurement functions.
17.5 | A radiation report must be automatically generated and transferable via DICOM.
17.6 | Should be FDA or CE approved product
17.7 | Electrical safety conforms to standards for electrical safety IEC-60601
17.8 | The unit must comply to IEC-61267
17.9 | The equipment/system must be approved and licensed by the Directorate Radiation Control of DoH: South Africa
17.10. | A copy of a valid license issued in terms of the Hazardous Substance Act, Act No 15 of 1973 must be submitted with the tender. Failure to submit such a valid license may result in a tender not being considered.
17.11 | Acceptance testing must be performed as part of the installation. (Failure to do will lead to failure to pay
17.12 | Lead Glass window to be installed as per room requirement
17.13 | Safety aspects of Radiation dosage leakage should be spelt out
17.14 | DB boards and electrical filters to be installed
17.15 | QA TOOLS
17.16 | Display Exposure index values
17.17 | Radiation protection Lead equivalence of the radiation protection screen must be 0.5m Pb and must be convex, covering the radiographer from various angles against scatter radiation.
17.18 | Patient Conventional face shield to be included.
17.19 | Patient Tomosynthesis face shield to be included, removable and retractable.
17.20. | Emergency stop controls must be supplied on both sides of the gantry and on the Acquisition Workstation
17.21 | Interlocks to stop system movement when compression is applied must be provided.
17.22 | Interlock for premature release of the X-ray switch must be available
17.23 | Interlock preventing an exposure whenever the filter or light field is not positioned correctly must be possible.
17.24 | Phantom for calibration should be provided.
18 | POWER SUPPLY
18.1 | The bidder must install adequate electrical power supply for the optimal functionality of the equipment.
18.2 | The mains cable of the unit being quoted for must be the hospital grade type and it must be of an adequate length (minimum length of three (3) meters).
18.3 | The power input must be 220-240V, 50Hz AC. Bidders must ensure that the product quoted for is fitted appropriately with a 15 Amp SABS approved mains plug that is held together by two screws.
18.4 | A 3 phase in / out power supply must be offered. State details.
18.5 | The equipment tendered for must not overload and trip the hospital power in the area where it is installed. Resettable overcurrent breaker/s must be fitted for protection. State details.
18.6 | All available online Uninterrupted Power Supply (UPS system) and line voltage regulators must be included for the system with a support time of 50 to 60 minutes. Attach brochure.
18.7 | The lifetime of UPS must be at least 15 years.
18.8 | An Uninterrupted Power Supply (UPS system) conforming to IS-302 must be offered for the computer system.
18.9 | The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade type and it must be a minimum length of three (3) meters.
18.10. | The mains cable of the unit tendered for must be SABS colour coded. State details.
18.11 | The equipment quoted for must be protected against electromagnetic interference. State details of measures that will be taken to protect the unit against electromagnetic interference.
18.12 | Maintenance free batteries must be offered.
19 | AIR CONDITIONER
19.1 | Bidder must include any air conditioning that is considered necessary for optimal functioning of the unit.
19.2 | Any such air conditioning must be covered by the guarantee bided for on the unit and servicing of the air conditioning units for the guarantee period and must be included in the bid price.
19.3 | The air conditioner must also be included in the maintenance plan.
19.4 | Successful bidder to replace existing air conditioner with an air conditioner of 24,000 BTU. (Where applicable).
19.5 | The air conditioner must be but no less than 24,000 BTU. State details on the air conditioner to be supplied according to the room size and machine requirements.
19.6 | The Successful bidder must install an outlet-plumbing pipe for water overflow from air conditioner to be re-directed to a drainage system.
19.7 | Air-conditioning must be included in Maintenance plan

## X-Ray B-Floating 60kW
Digital X-RAY UNIT (60kW generator)
Description
ITEM
1 | The said Unit must be fully automatic, digital, radiographic examination and evaluation workstation based on two fixed/wireless detectors technology for high image dynamics with excellent signal/noise ratio is required.
1.1 | The bid price quoted must include the supply, installation, demonstration and commissioning of the equipment as well as the building alterations that may be required for the successful installation of the unit/s.
1.2 | State make and model of main unit offered.
1.3 | The bidders are requested to structure their quote so that it reflects the price of the unit separately from any building alterations that are required. The price schedule at the end of this specification must be completed.
1.4 | The unit must be suitable for general and trauma radiography.
2 | GENERATOR:
2.1 | The unit should be a high frequency type generator.
2.2 | The output shall be 60kW. State maximum milliamps at different kilovolts.
2.3 | The unit must incorporate an electronic timer with minimum exposure time of 1mS and electronic switching
2.4 | The unit must include automatic overload protection
2.5 | The unit must have automatic mains compensation.
2.6 | A milliamp meter or indicator must be provided.
2.7 | The kV, mA and time must be independently selectable and digitally displayed.
2.8 | Bidders to state the range of the exposure timer. Acceptable range 0.001 sec – 6 sec (1 ms – 6sec).
2.9 | Bidders to state kV and mA ranges for radiography. The kV range to be at least 40 to 150 kV and mA range at least 10 to 630mA. The kV selection must be in steps of 1 and must be a single control.
2.10. | The unit offered must be programmable. Bidders to supply details of programming facility and the number of programs.
2.11 | The offered unit must include automatic exposure control.
2.12 | Tube overload protection is required
3 | X-Ray TUBE:
3.1 | The rotating anode tube must match the output of the generator.
3.2.1 | The maximum tube kV must be at least 150 kV.
3.2.2 | The unit must incorporate an electronic timer with minimum exposure time of 1mS and electronic switching
3.2.3 | The maximum tube mA must be at least 500 mA
3.2.4 | The speed of anode rotation must be at least 8000rpm
3.2.5 | Anode heat storage capacity must be at least 300kHU
3.2.6 | The tube housing heat storage capacity must be more than 400kHU
3.2.7 | The size of small focus must be 0,6 mm
3.2.8 | The maximum tube mA must be at least 500 mA
3.2.9 | The speed of anode rotation must be at least 8000rpm
3.2.10 | The size of large focus must be 1,2 mm
3.2.11 | Inherent filtration of at least 0.8mmAl
3.2.12 | Total filtration must exceed 2.5mmAl.
3.2.13 | Anode angle must be large enough to provide a uniform beam over a field of 43cm x 43cm at a FFD of 100cm.
3.3 | HIGH TENSION CABLE: One pair of high tension cable of suitable length must be supplied. The cable length must not restrict free movement of the tube for all radiographic procedures.
3.4 | The power supply must be a three phase voltage 380-400 V at 50Hz
4 | X-RAY TUBE WITH MULTILEAF COLLIMATOR
4.1 | Nominal voltage range (40 - 150kV)
4.2 | State the nominal power
4.3 | Focal spot parameter according to International Electrotechnical Commission (IEC) 336
4.4 | Small focus 0.6mm
4.5 | Large focus 1.2mm
4.6 | Anode angle 12o
4.8 | State the speed of rotating anode
4.9 | Heat storage capacity of the anode should not be less than 350kHU
4.10 | Heat storage capacity of the X-ray tube should not be less than 400kHU
4.11 | Radiation absorption (according to IEC 601-1-3)
4.12 | There must be a measuring tape on the X-Ray tube
4.13 | Light localizer on/off switching with timer to terminate the light after 30 seconds
4.14 | The table must be able to perform tomography. Please state the angles offered and the speeds related to each angle.
5 | CEILING MOUNTED TUBE SUPPORT
5.1 | The tube column must be ceiling mounted
5.2 | The X-ray tube assembly rotation must be 360° about the vertical and horizontal axis
5.3 | Tube movements:
5.3.1 | Longitudinal: at least 150 cm from the iso centre
5.3.2 | Transverse: at least 100 cm from the iso centre
5.3.3 | Vertical: Tube centre should be able to move down to the lowest of 30cm above the floor.
5.4 | Multiplanner movements controlled by electric lock
5.5 | All movements to be counter balanced
6 | CEILING SUSPENSION:
6.1 | The ceiling suspension must provide coverage of an area of approximately 3 X 4 meters.
6.2 | Stops must be provided for centering to table and vertical Bucky.
6.3 | Bidders to indicate focal film distance range for vertical and table Bucky.
6.4 | A safety device for possible failure of the counter-balance system must be incorporated.
6.5 | The unit must be able to perform horizontal beam radiography on both side of the table.
6.6 | Bidders to state:
6.6.1 | Tube rotation and movement around vertical axis
6.6.2 | Tube rotation around horizontal axis and
6.6.3 | Range of vertical movement.
7 | BUCKY TABLE:
7.1 | The bidder to state the make and model of the Bucky table offered.
7.2 | The Bucky table must have a flat, four way floating top.
7.3 | The table must be electronically height controlled. Foot switches must form part of the table.
7.4 | State the variable height of the table offered.
7.5 | The dimensions of the table top must have atleast length 210cm and 80 to 90 cm width
7.6 | State the lateral and longitudinal movement of the table top.
7.7 | The table offered must be able to carry patients with a weight of at least 220kg centred and off-centred.
7.8 | State the longitudinal movement of Bucky tray. Bidders must detail whether the distance relates to the outermost end of the Bucky tray, or from centre point to centre point. A data sheet must be included to clarify this point.
7.9 | The table-top movement must be controlled with electromagnetic brakes.
7.10. | The table-top film distance must be approximately 6 centimeters.
7.11 | The Bucky assembly must be fitted with the oscillating secondary radiation grids suitable for direct digital flat panel technology.
7.12 | A compression band with fitting mechanism compatible to the table top must be supplied.
7.13 | A full set of positioning sponges with washable water resistance covers must be supplied with the table.
7.14 | Emergency stop mechanism must be attached to the table
8 | VERTICAL BUCKY
8.1 | State the make and model of vertical Bucky offered.
8.2 | The vertical height should be adequate for a lateral cervical spine projection of a standing patient and therefore maximum height from centre of the Bucky to floor should be at least 180 cm.
8.3 | The Bidder to state range of vertical movement from floor to centre point of the Bucky.
8.4 | Vertical movement of the vertical Bucky must be via an electronic locking/unlocking switch. Analogue locking/unlocking levers or knobs will not be accepted
8.5 | The Bucky assembly must be fitted with the oscillating secondary radiation grids suitable for direct digital flat panel technology and must accommodate a chest radiography at 180cm.
8.6 | Radiography of a patient who is in a seated position on the stretcher or table top must be possible.
8.7 | The Bucky must be mounted centrally on the column with the choice of loading on either left or right.
8.8 | An overhang arm support must be included that is compatible with erect Bucky.
8.9 | An overhang frame for out of Bucky radiography must be included that is compatible with erect Bucky.
9 | DIGITAL FLAT PANEL DETECTOR:
9.1 | The detector technology must be of Cesium Iodide type
9.2 | The exposure formats must be 43cm x 43cm
9.3 | The full vertical rotation must be possible without removing detector for 43 x 43 format
9.4 | The full horizontal rotation must be possible without removing detector for 43 x 43 format
9.5 | The detector exposure matrix: pixels must be minimum 3.6 lp/mm
9.6 | The detector pixel pitch must be minimum 139um
9.7 | The reading speed at maximum format from the exposure to preview on the monitor must not be greater than 8 seconds
9.8 | The digital Flat panels must operate wirelessly
9.9 | The wireless flat panels must be supplied with x2 spare battery packs each that can be charged outside of the flat digital flat panel.
9.1 | The analog to digital conversion must not be less than 12 bits
9.11 | The wireless ditector must come with grids, detector cover and holder
9.12 | The flat panel battery:
9.12.1 | The battery should be able to acquire and transfer at least 90 images or more when fully charged. Details are stated in the supporting documents provided .
9.12.2 | The battery should be able to handle 5 hours of active use and 15 hours standby, when fully charged. Details are stated in the supporting documents provided .
9.12.3 | Charging time for each battery from fully depleted to fully charged should be less than 2 hours. Details are stated in the supporting documents provided.
9.12.4 | The batteries must be interchangeable between the 24 x30 cm and 43 x 43 cm panels.
9.12.5 | The battery charger must be light weight and should weigh less than 2 kg. Details of the weight of the battery charger are stated in the supporting documents provided .
9.12.6 | The battery must be able to operate at ambient temperature range of minumum of 5°C to maximum of 35 °C.
9.12.7 | The battery must be able to operate at ambient humidity of minimum humidity of 20 % to maximum of 85% RH.
9.12.8 | The battery must be able tolerate storage at a temperature range of minumum of -30 to to maximum of 60 degrees celsius.
9.12.9 | The battery must be able tolerate storage at a minimum humidity of 10 to maximum of 85% RH.
9.13 | The Grid
9.13.1 | An encasement with a built in grid that fits the 43cm x 43cm detector must be supplied.
9.13.2 | The grid must have at least:
9.13.3 | Grid ratio 8:1
9.13.4 | 100cm to 180cm focal distance
9.14 | Load capacity
9.14.1 | The uniform load that the detector must be able to withstand should be at least 150kg. Details are stated in the supporting documents provided .
9.14.2 | The local load (on an area of 40 mm diameter) that the detector must be able to withstand should be at least 100kg. Details are stated in the supporting documents provided .
11 | X-RAY EXAMINATION TABLE
11.1 | Table movements must be motorised (all directions)
11.2 | Minimum height (above the Floor) 40-60cm
11.3 | Must have a Floating table top
11.4 | Detector carriage should have a motorized scattered radiation grid.
11.5 | Film Focus Distance = 115cm
11.6 | Table top dimensions of at least 220L X 80W cm.
11.7 | Table top longitudinal movement of at least (50cm) in both direction with minimum patient weight of 220kg.
11.8 | Transversal movement of at least (20cm) in one direction with minimum patient weight of 220kg.
11.9 | Control of the table functions via footswitch attached to the table
11.10 | Electromagnetic brakes to be standard
11.11 | Table top to be of suitable low radiation absorption material
11.12 | Load capacity must be at least 220 kg
11.13 | Must have automatic exposure device
12 | IMAGE WORKSTATION
12.1 | Storage capacity must be minimum 4 tera bytes
12.2 | The system must have minimum of 4 GB memory
12.3 | Minimum 19 inch, maximum 23 inch colour LCD/LED display
12.4 | Must be a touch screen
12.5 | Resolution Minimum 1920x 1080 pixels
13 | IMAGING SOFTWARE
MUST INCLUDE:
13.1 | General Features
13.1.1 | Windows based graphic user interface
13.1.2 | Multi-image display (1x1 / 2x2 / 3x3 / 4x4 )
13.1.3 | Multi-image selection
13.1.4 | Unlimited procedure set ups
13.1.5 | ROI changing and creation feature
13.1.6 | English language support
13.1.7 | Unlimited PACS code ( CPT code )
13.1.8 | Default anatomic program more than 700
13.1.9 | Support DICOM Worklist SCU, DICOM Storage SCU and transfer function
13.1.10 | Support DICOM Multi-transfer function
13.1.11 | All hardware & software relating to DICOM compatibility, including licences, must be supplied.
13.2 | Post processing features
13.2.1 | Noise reduction parameter
13.2.2 | Global contrast parameter
13.2.3 | Detail contrast parameter
13.3 | Image Maintenance (All functions must be supported by the pop-up menu)
13.3.1 | ROI
13.3.2 | Mark
13.3.3 | Horizontal Flip
13.3.4 | Vertical Flip
13.3.5 | Rotate CW
13.3.6 | Rotate CCW
13.3.7 | Inverse (Black or White)
13.3.8 | Text Annotation
13.3.9 | Ruler : Distance tool
13.3.10 | Angle : Angle measurement tool
13.3.11 | Zoom : Image zoom in/out
13.3.12 | Magnify : Image magnify glass window
13.3.13 | Pan : Image panning
13.3.14 | Fit Image : Auto fitting to window size
13.3.15 | Image Cut : Image crop/cut function
13.3.16 | Image Copy : Copy of image in the region of interest(ROI)
13.3.17 | Image Recovery : Recover the original image
13.4 | CD Burning
13.4.1 | DICOMDIR based CDR data generation
13.4.2 | Support CD/DVD Recording
13.4.3 | Include internal DICOM Viewer
13.4.4 | Support multi-study data
13.5 | DICOM Features : DICOM PRINT
13.5.1 | DICOM 3.0 compatible
13.5.2 | Support Print Preview
13.5.3 | Support Film Layout : 1:1 / 1:2 / 2:1 / 2:2
13.5.4 | Support multi-page printing
13.5.5 | Support Real size printing
13.5.6 | Support image swap in layout
13.6 | DICOM Feature : DICOM STORAGE
13.6.1 | DICOM 3.0 compatible
13.6.2 | Support DX/CR modality
13.6.3 | Support the modification of Transfer Syntax
13.6.3.1 | Transfer status
13.6.4 | DICOM 3.0 compatible
13.6.5 | Support DX/CR modality
13.7 | DICOM Feature : WORKLIST
13.7.1 | Support DICOM Modality Worklist Standard
13.7.2 | Support DICOM Query/Retrieve
13.7.3 | Support Search Filter ( ID / Name / Access Number )
13.7.4 | Support Import Filter
13.7.5 | Support DICOM Modality Worklist Standard
13.8 | Overlay Display on image
13.8.1 | Projection description
13.8.2 | Hospital or Institution Name, Date and Time of examination
13.8.3 | Radiographer Name
13.8.4 | Patient Name, Sex and Age
13.8.5 | Machine Name
13.8.6 | kV / mA / Time / mAs
13.8.7 | Feed-back mAs / Feed-back Time for AEC
13.8.8 | EI(Exposure Index)
14 | THE FOLLOWING MUST BE PROVIDED WITH THE SYSTEM INTERFACES AND CONNECTIVITY
14.1 | A "Conformance Statement" for all DICOM interfaces must be included. (the document must be attached)
14.2 | Interface DICOM: "Basic Print"
14.3 | Interface DICOM: "Modality Work-list"
14.4 | Interface DICOM: "Send as Service Class User (SCU)"
14.5 | Interface DICOM: "Receive as Service Class User (SCU)"
14.6 | Interface DICOM: "Retrieve as Service Class User (SCU)"
14.7 | Interface for checking the system status or for loading new
14.8 | Software to all components should be DICOM Interfaced.
14.9 | The modality shall include the DICOM Service Class license Image Storage: SCU and SCP
14.10. | The modality shall include the DICOM Service Class license Query/Retrieve SCU
14.11 | The modality shall include the DICOM Service Class license Print SCU
14.12 | The modality shall include the DICOM Service Class license Storage Commitment SCP
14.13 | The vendor shall include annexures of Proof of DICOM Conformance Statement
14.14 | The succesful bidder must ensure that the configuration between the processed images on the Digital Imaging Workstation and that transferred to the receivers of these images, such as PACS , Laser Printer , CD/DVD are properly calibrated/ reconfigured during installation, service and maintenance,repairs, and future software updates/upgrades.
15 | GENERAL
15.1 | Remote diagnostics feature for fault finding / troubleshooting must be available
15.2 | All computer systems supplied must have UPS protection with at least 20 minutes backup
15.3 | The Unit must link to existing laser printer
15.4 | The Unit must link to existing CD Robot
15.5 | The Unit offered must be able to connect a PACS/ RIS and HIS
15.6 | The Unit must be able to connect to any vendor
17 | RADIATION PROTECTION
17.1 | The system must have the ability to display Exposure Indicator values (EI/ S-values), immediately after exposure.
17.2 | X-ray Warning Notice for display at the X-ray Room entrance (not branded).
17.3 | An advanced Retake / Reject / Repeat analysis software package that includes analysis as well as calculations in percentages, must be provided for applicable systems.
17.4 | The applicable systems must have an integrated Dose Area Product (DAP) meter that includes skin dose equivalent measurement functions.
17.5 | A radiation report must be automatically generated and transferable via DICOM.
17.6 | Should be FDA or CE approved product
17.7 | Electrical safety conforms to standards for electrical safety IEC-60601
17.8 | The unit must comply to IEC-61267
17.9 | The equipment/system must be approved and licensed by the Directorate Radiation Control of DoH: South Africa
17.10. | A copy of a valid license issued in terms of the Hazardous Substance Act, Act No 15 of 1973 must be submitted with the tender. Failure to submit such a valid license may result in a tender not being considered.
17.11 | Acceptance testing must be performed as part of the installation. (Failure to do will lead to failure to pay
17.12 | Lead Glass window to be installed as per room requirement
17.13 | Safety aspects of Radiation dosage leakage should be spelt out
17.14 | DB boards and electrical filters to be installed
17.15 | QA TOOLS
17.16 | Display Exposure index values
17.17 | Radiation protection Lead equivalence of the radiation protection screen must be 0.5m Pb and must be convex, covering the radiographer from various angles against scatter radiation.
17.18 | Patient Conventional face shield to be included.
17.19 | Patient Tomosynthesis face shield to be included, removable and retractable.
17.20. | Emergency stop controls must be supplied on both sides of the gantry and on the Acquisition Workstation
17.21 | Interlocks to stop system movement when compression is applied must be provided.
17.22 | Interlock for premature release of the X-ray switch must be available
17.23 | Interlock preventing an exposure whenever the filter or light field is not positioned correctly must be possible.
17.24 | Phantom for calibration should be provided.
18 | POWER SUPPLY
18.1 | The bidder must install adequate electrical power supply for the optimal functionality of the equipment.
18.2 | The mains cable of the unit being quoted for must be the hospital grade type and it must be of an adequate length (minimum length of three (3) meters).
18.3 | The power input must be 220-240V, 50Hz AC. Bidders must ensure that the product quoted for is fitted appropriately with a 15 Amp SABS approved mains plug that is held together by two screws.
18.4 | A 3 phase in / out power supply must be offered. State details.
18.5 | The equipment tendered for must not overload and trip the hospital power in the area where it is installed. Resettable overcurrent breaker/s must be fitted for protection. State details.
18.6 | All available online Uninterrupted Power Supply (UPS system) and line voltage regulators must be included for the system with a support time of 50 to 60 minutes. Attach brochure.
18.7 | The lifetime of UPS must be at least 15 years.
18.8 | An Uninterrupted Power Supply (UPS system) conforming to IS-302 must be offered for the computer system.
18.9 | The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade type and it must be a minimum length of three (3) meters.
18.10. | The mains cable of the unit tendered for must be SABS colour coded. State details.
18.11 | The equipment quoted for must be protected against electromagnetic interference. State details of measures that will be taken to protect the unit against electromagnetic interference.
18.12 | Maintenance free batteries must be offered.
19 | AIR CONDITIONER
19.1 | Bidder must include any air conditioning that is considered necessary for optimal functioning of the unit.
19.2 | Any such air conditioning must be covered by the guarantee bided for on the unit and servicing of the air conditioning units for the guarantee period and must be included in the bid price.
19.3 | The air conditioner must also be included in the maintenance plan.
19.4 | Successful bidder to replace existing air conditioner with an air conditioner of 24,000 BTU. (Where applicable).
19.5 | The air conditioner must be but no less than 24,000 BTU. State details on the air conditioner to be supplied according to the room size and machine requirements.
19.6 | The Successful bidder must install an outlet-plumbing pipe for water overflow from air conditioner to be re-directed to a drainage system.
19.7 | Air-conditioning must be included in Maintenance plan

## X-Ray B-Tube stand 80kW
ITEM | X-RAY BUCKY UNIT WITH A FREE - STANDING TUBE STAND and VERTICAL BUCKY STAND.
1
1.1 | The said Unit must be fully automatic, digital, radiographic examination and evaluation workstation based on two detector technology for high image dynamics with excellent signal/noise ratio is required.
1.2 | The bid price quoted must include the supply, installation, demonstration and commissioning of the equipment as well as the building alterations that may be required for the successful installation of the unit/s.
1.3 | State the make and model of the unit offered.
1.4 | The bidders are requested to structure their quote so that it reflects the price of the unit separately from any building alterations that are required. The price schedule at the end of this specification must be completed.
1.5 | The unit must be suitable for general and trauma radiography.
TECHNICAL SPECIFICATION
2 | GENERATOR:
2.1 | The unit should be a high frequency type generator.
2.2 | The output must be 80kW. State maximum milliamps at different kilovolts
2.3 | The unit must incorporate an electronic timer with a minimum exposure time of 0,001 sec or less and electronic switching.
2.4 | The unit must include automatic overload protection.
2.5 | The unit must include automatic mains compensation.
2.6 | A milliamp meter or indicator must be provided
2.7 | The kV, mA and time must be independently selectable and digitally displayed.
2.8 | Bidders to state the range of the exposure timer. Acceptable range 0.001 sec – 6 sec (1 ms – 6sec).
2.9 | Bidders to state kV and mA ranges for radiography. The kV range to be at least 40 to 125 kV and mA range at least 10 to 600mA. KV selection must be in steps of 1 and must be a single control.
2.10. | The unit must preferably be programmable. Supply details of the programming facility and the number of programmes.
2.11 | The offered unit must include automatic exposure control.
3 | ROTATING ANODE X‑RAY TUBE
3.1 | The rotating anode tube must match the output of the generator.
3.2 | Bidders to supply the following details:
3.2.1 | The maximum tube kV must be at least 125 kV
3.2.2 | The maximum tube mA must be at least 600 mA
3.2.3 | The speed of anode rotation must be at least 8000rpm.
3.2.4 | Anode heat storage capacity must be at least 300kHU
3.2.5 | The tube housing heat storage capacity must be more than 400kHU
3.2.6 | The size of small focus must be 0,6 mm
3.2.7 | The size of large focus must be 1,2 mm
3.2.8 | Inherent filtration of at least 0.8mmAl
3.2.9 | Total filtration must exceed 2.5mmAl.
3.2.10 | Anode angle must be large enough to provide a uniform beam over a field of 43cm x 43cm at a FFD of 100cm.
3.3 | One pair of high tension cables of suitable length must be supplied. The cable length must not restrict free movement of the tube for all radiographic procedures.
3.4 | The table must be able to perform tomography. Please state the angles offered and the speeds related to each angle.
4 | COLLIMATOR.
4.1 | The unit must have a multileaf collimator.
4.2 | Rotating flange between the collimator and tube must be supplied to provide diagonal collimation.
4.3 | One sinus cone must be supplied with the unit.
4.4 | A suitable handle bar with central release of all brakes as well as individual release must be provided.
4.6 | An electronic timer for the Halogen light source must be provided. The collimator to provide illumination in excess of 100 lux.
5 | TUBE STAND.
5.1 | A floor mounted tube stand must be supplied that will provide a focal film distance of at least 150 cms to the Bucky table.
The tube stand must not be integrated into the table.
The longitudinal movement of the tube stand must extend beyond the table at both ends.
5.2 | The unit must be fitted with electro‑magnetic brakes.
5.3 | Bidders to indicate the focal film distance range. Markings indicating the focal film distance on the tube stand must be provided.
5.4 | A safety device for possible failure of the counter‑balance system must be incorporated.
5.5 | The unit must be suitable to function with an x‑ray Bucky table, the vertical Bucky and provide a facility for horizontal beam radiography.
5.6 | Bidders to state:
i) Tube rotation around vertical axis
ii) Tube rotation around horizontal axis and
iii) Range of vertical movement.
iv) Range of longitudinal movement.
5.7 | Radiography of a patient who is in a seated position on the stretcher or table top must be possible.
5.8 | The unit must be suitable for radiography of patients on trolleys and therefore space must be available at both ends of the table – e.g. it must be possible to do a supine chest x-ray or a tibia and fibula x-ray on a hospital trolley. State the longitudinal movement of the tube column.
5.8 | Rotation of the tube to the lateral position for horizontal shoot through radiographs must be possible by means of a central electronic locking/unlocking device at the collimator. Units that have the locking/unlocking device on the tube stand close to the tube rail will not be accepted.
6 | BUCKY TABLE
6.1 | The Bucky table must have a flat, four way floating top.
6.2 | The table must be electronically height controlled. Foot switches must form part of the table.
6.3 | The bidder to state dimensions of table top offered (length and width).
6.4 | State the variable height of the table offered.
6.5 | State the lateral and longitudinal movement of the table top.
6.6 | The table offered must be able to carry patients with a weight of at least 220kg.
6.7 | State longitudinal movement of Bucky tray. Bidders must detail whether the distance relates to the outermost end of the Bucky tray or from centre point to centre point. A data sheet must be included to clarify this point.
6.8 | The table‑top movement must be controlled with electro‑magnetic brakes.
6.9 | The table‑top ‑ film distance must be less than 6 cms.
6.10. | The Bucky assembly must be fitted with the oscillating secondary radiation grids suitable for direct digital flat panel technology.
6.11 | A compression band with fitting mechanism compatible to the top must be supplied.
6.12 | A full set of positioning sponges with washable water resistance covers must be supplied with the table.
7 | VERTICAL BUCKY
7.1 | State the make and model of the vertical Bucky offered.
7.2 | The Bucky assembly must be fitted with a oscillating secondary radiation grid of 10: 1, and be suitable for chest in Bucky radiography.
7.3 | Radiography of a patient who is in a seated position must be possible.
7.4 | The Bucky assembly must be fitted with the oscillating secondary radiation grids suitable for direct digital flat panel technology and must accommodate a chest radiography at 180cm.
7.5 | The Bucky must be mounted centrally on the column with the choice of loading on either left or right
7.6 | Vertical movement of the vertical Bucky must be via an electronic locking/unlocking switch. Analogue locking/unlocking levers or knobs will not be accepted
7.7 | An overhang frame for out of Bucky radiography must be included that is compatible with erect Bucky.
7.8 | An overhang arm support must be included that is compatible with erect Bucky.
8 | RADIATION PROTECTION
8.1 | The system must have the ability to display Exposure Indicator values (EI/ S-values), immediately after exposure.
8.2 | X-ray Warning Notice for display at the X-ray Room entrance (not branded).
8.3 | An advanced Retake / Reject / Repeat analysis software package that includes analysis as well as calculations in percentages, must be provided for applicable systems.
8.4 | The applicable systems must have an integrated Dose Area Product (DAP) meter that includes skin dose equivalent measurement functions.
8.5 | A radiation report must be automatically generated and transferable via DICOM.
8.6 | Should be FDA or CE approved product
8.7 | Electrical safety conforms to standards for electrical safety IEC-60601
8.8 | The unit must comply to IEC-61267
8.9 | The equipment/system must be approved and licensed by the Directorate Radiation Control of DoH: South Africa
8.10. | A copy of a valid license issued in terms of the Hazardous Substance Act, Act No 15 of 1973 must be submitted with the tender. Failure to submit such a valid license may result in a tender not being considered.
8.11 | Acceptance testing must be performed as part of the installation. (Failure to do will lead to failure to pay
8.12 | Lead Glass window to be installed as per room requirement
8.13 | Safety aspects of Radiation dosage leakage should be spelt out
8.14 | DB boards and electrical filters to be installed
8.15 | QA TOOLS
8.16 | Display Exposure index values
8.17 | Radiation protection Lead equivalence of the radiation protection screen must be 0.5m Pb and must be convex, covering the radiographer from various angles against scatter radiation.
8.18 | Patient Conventional face shield to be included.
8.19 | Patient Tomosynthesis face shield to be included, removable and retractable.
8.20. | Emergency stop controls must be supplied on both sides of the gantry and on the Acquisition Workstation
8.21 | Interlocks to stop system movement when compression is applied must be provided.
8.22 | Interlock for premature release of the X-ray switch must be available
8.23 | Interlock preventing an exposure whenever the filter or light field is not positioned correctly must be possible.
8.24 | Phantom for calibration should be provided.
9 | POWER SUPPLY
9.1 | The bidder must install adequate electrical power supply for the optimal functionality of the equipment.
9.2 | The mains cable of the unit being quoted for must be the hospital grade type and it must be of an adequate length (minimum length of three (3) meters).
9.3 | The power input must be 220-240V, 50Hz AC. Bidders must ensure that the product quoted for is fitted appropriately with a 15 Amp SABS approved mains plug that is held together by two screws.
9.4 | A 3 phase in / out power supply must be offered. State details.
9.5 | The equipment tendered for must not overload and trip the hospital power in the area where it is installed. Resettable overcurrent breaker/s must be fitted for protection. State details.
9.6 | All available online Uninterrupted Power Supply (UPS system) and line voltage regulators must be included for the system with a support time of 50 to 60 minutes. Attach brochure.
9.7 | The lifetime of UPS must be at least 15 years.
9.8 | An Uninterrupted Power Supply (UPS system) conforming to IS-302 must be offered for the computer system.
9.9 | The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade type and it must be a minimum length of three (3) meters.
9.10. | The mains cable of the unit tendered for must be SABS colour coded. State details.
9.11 | The equipment quoted for must be protected against electromagnetic interference. State details of measures that will be taken to protect the unit against electromagnetic interference.
9.12 | Maintenance free batteries must be offered.
10 | AIR CONDITIONER
10.1 | Bidder must include any air conditioning that is considered necessary for optimal functioning of the unit.
10.2 | Any such air conditioning must be covered by the guarantee bided for on the unit and servicing of the air conditioning units for the guarantee period and must be included in the bid price.
10.3 | The air conditioner must also be included in the maintenance plan.
10.4 | Successful bidder to replace existing air conditioner with an air conditioner of 24,000 BTU. (Where applicable).
10.5 | The air conditioner must be but no less than 24,000 BTU. State details on the air conditioner to be supplied according to the room size and machine requirements.
10.6 | The Successful bidder must install an outlet-plumbing pipe for water overflow from air conditioner to be re-directed to a drainage system.
10.7 | Air-conditioning must be included in Maintenance plan

## X-Ray B-Tube stand 60kW
ITEM | X-RAY BUCKY UNIT WITH A FREE - STANDING TUBE STAND and VERTICAL BUCKY STAND.
1
1.1 | The said Unit must be fully automatic, digital, radiographic examination and evaluation workstation based on two detector technology for high image dynamics with excellent signal/noise ratio is required.
1.2 | The bid price quoted must include the supply, installation, demonstration and commissioning of the equipment as well as the building alterations that may be required for the successful installation of the unit/s.
1.3 | State the make and model of the unit offered.
1.4 | The bidders are requested to structure their quote so that it reflects the price of the unit separately from any building alterations that are required. The price schedule at the end of this specification must be completed.
1.5 | The unit must be suitable for general and trauma radiography.
TECHNICAL SPECIFICATION
2 | GENERATOR:
2.1 | The unit should be a high frequency type generator.
2.2 | The output must be 60kW. State maximum milliamps at different kilovolts
2.3 | The unit must incorporate an electronic timer with a minimum exposure time of 0,001 sec or less and electronic switching.
2.4 | The unit must include automatic overload protection.
2.5 | The unit must include automatic mains compensation.
2.6 | A milliamp meter or indicator must be provided
2.7 | The kV, mA and time must be independently selectable and digitally displayed.
2.8 | Bidders to state the range of the exposure timer. Acceptable range 0.001 sec – 6 sec (1 ms – 6sec).
2.9 | Bidders to state kV and mA ranges for radiography. The kV range to be at least 40 to 125 kV and mA range at least 10 to 600mA. KV selection must be in steps of 1 and must be a single control.
2.10. | The unit must preferably be programmable. Supply details of the programming facility and the number of programmes.
2.11 | The offered unit must include automatic exposure control.
3 | ROTATING ANODE X‑RAY TUBE
3.1 | The rotating anode tube must match the output of the generator.
3.2 | Bidders to supply the following details:
3.2.1 | The maximum tube kV must be at least 125 kV
3.2.2 | The maximum tube mA must be at least 600 mA
3.2.3 | The speed of anode rotation must be at least 8000rpm.
3.2.4 | Anode heat storage capacity must be at least 300kHU
3.2.5 | The tube housing heat storage capacity must be more than 400kHU
3.2.6 | The size of small focus must be 0,6 mm
3.2.7 | The size of large focus must be 1,2 mm
3.2.8 | Inherent filtration of at least 0.8mmAl
3.2.9 | Total filtration must exceed 2.5mmAl.
3.2.10 | Anode angle must be large enough to provide a uniform beam over a field of 43cm x 43cm at a FFD of 100cm.
3.3 | One pair of high tension cables of suitable length must be supplied. The cable length must not restrict free movement of the tube for all radiographic procedures.
3.4 | The table must be able to perform tomography. Please state the angles offered and the speeds related to each angle.
4 | COLLIMATOR.
4.1 | The unit must have a multileaf collimator.
4.2 | Rotating flange between the collimator and tube must be supplied to provide diagonal collimation.
4.3 | One sinus cone must be supplied with the unit.
4.4 | A suitable handle bar with central release of all brakes as well as individual release must be provided.
4.6 | An electronic timer for the Halogen light source must be provided. The collimator to provide illumination in excess of 100 lux.
5 | TUBE STAND.
5.1 | A floor mounted tube stand must be supplied that will provide a focal film distance of at least 150 cms to the Bucky table.
The tube stand must not be integrated into the table.
The longitudinal movement of the tube stand must extend beyond the table at both ends.
5.2 | The unit must be fitted with electro‑magnetic brakes.
5.3 | Bidders to indicate the focal film distance range. Markings indicating the focal film distance on the tube stand must be provided.
5.4 | A safety device for possible failure of the counter‑balance system must be incorporated.
5.5 | The unit must be suitable to function with an x‑ray Bucky table, the vertical Bucky and provide a facility for horizontal beam radiography.
5.6 | Bidders to state:
i) Tube rotation around vertical axis
ii) Tube rotation around horizontal axis and
iii) Range of vertical movement.
iv) Range of longitudinal movement.
5.7 | Radiography of a patient who is in a seated position on the stretcher or table top must be possible.
5.8 | The unit must be suitable for radiography of patients on trolleys and therefore space must be available at both ends of the table – e.g. it must be possible to do a supine chest x-ray or a tibia and fibula x-ray on a hospital trolley. State the longitudinal movement of the tube column.
5.8 | Rotation of the tube to the lateral position for horizontal shoot through radiographs must be possible by means of a central electronic locking/unlocking device at the collimator. Units that have the locking/unlocking device on the tube stand close to the tube rail will not be accepted.
6 | BUCKY TABLE
6.1 | The Bucky table must have a flat, four way floating top.
6.2 | The table must be electronically height controlled. Foot switches must form part of the table.
6.3 | The bidder to state dimensions of table top offered (length and width).
6.4 | State the variable height of the table offered.
6.5 | State the lateral and longitudinal movement of the table top.
6.6 | The table offered must be able to carry patients with a weight of at least 220kg.
6.7 | State longitudinal movement of Bucky tray. Bidders must detail whether the distance relates to the outermost end of the Bucky tray or from centre point to centre point. A data sheet must be included to clarify this point.
6.8 | The table‑top movement must be controlled with electro‑magnetic brakes.
6.9 | The table‑top ‑ film distance must be less than 6 cms.
6.10. | The Bucky assembly must be fitted with the oscillating secondary radiation grids suitable for direct digital flat panel technology.
6.11 | A compression band with fitting mechanism compatible to the top must be supplied.
6.12 | A full set of positioning sponges with washable water resistance covers must be supplied with the table.
7 | VERTICAL BUCKY
7.1 | State the make and model of the vertical Bucky offered.
7.2 | The Bucky assembly must be fitted with a oscillating secondary radiation grid of 10: 1, and be suitable for chest in Bucky radiography.
7.3 | Radiography of a patient who is in a seated position must be possible.
7.4 | The Bucky assembly must be fitted with the oscillating secondary radiation grids suitable for direct digital flat panel technology and must accommodate a chest radiography at 180cm.
7.5 | The Bucky must be mounted centrally on the column with the choice of loading on either left or right
7.6 | Vertical movement of the vertical Bucky must be via an electronic locking/unlocking switch. Analogue locking/unlocking levers or knobs will not be accepted
7.7 | An overhang frame for out of Bucky radiography must be included that is compatible with erect Bucky.
7.8 | An overhang arm support must be included that is compatible with erect Bucky.
8 | RADIATION PROTECTION
8.1 | The system must have the ability to display Exposure Indicator values (EI/ S-values), immediately after exposure.
8.2 | X-ray Warning Notice for display at the X-ray Room entrance (not branded).
8.3 | An advanced Retake / Reject / Repeat analysis software package that includes analysis as well as calculations in percentages, must be provided for applicable systems.
8.4 | The applicable systems must have an integrated Dose Area Product (DAP) meter that includes skin dose equivalent measurement functions.
8.5 | A radiation report must be automatically generated and transferable via DICOM.
8.6 | Should be FDA or CE approved product
8.7 | Electrical safety conforms to standards for electrical safety IEC-60601
8.8 | The unit must comply to IEC-61267
8.9 | The equipment/system must be approved and licensed by the Directorate Radiation Control of DoH: South Africa
8.10. | A copy of a valid license issued in terms of the Hazardous Substance Act, Act No 15 of 1973 must be submitted with the tender. Failure to submit such a valid license may result in a tender not being considered.
8.11 | Acceptance testing must be performed as part of the installation. (Failure to do will lead to failure to pay
8.12 | Lead Glass window to be installed as per room requirement
8.13 | Safety aspects of Radiation dosage leakage should be spelt out
8.14 | DB boards and electrical filters to be installed
8.15 | QA TOOLS
8.16 | Display Exposure index values
8.17 | Radiation protection Lead equivalence of the radiation protection screen must be 0.5m Pb and must be convex, covering the radiographer from various angles against scatter radiation.
8.18 | Patient Conventional face shield to be included.
8.19 | Patient Tomosynthesis face shield to be included, removable and retractable.
8.20. | Emergency stop controls must be supplied on both sides of the gantry and on the Acquisition Workstation
8.21 | Interlocks to stop system movement when compression is applied must be provided.
8.22 | Interlock for premature release of the X-ray switch must be available
8.23 | Interlock preventing an exposure whenever the filter or light field is not positioned correctly must be possible.
8.24 | Phantom for calibration should be provided.
9 | POWER SUPPLY
9.1 | The bidder must install adequate electrical power supply for the optimal functionality of the equipment.
9.2 | The mains cable of the unit being quoted for must be the hospital grade type and it must be of an adequate length (minimum length of three (3) meters).
9.3 | The power input must be 220-240V, 50Hz AC. Bidders must ensure that the product quoted for is fitted appropriately with a 15 Amp SABS approved mains plug that is held together by two screws.
9.4 | A 3 phase in / out power supply must be offered. State details.
9.5 | The equipment tendered for must not overload and trip the hospital power in the area where it is installed. Resettable overcurrent breaker/s must be fitted for protection. State details.
9.6 | All available online Uninterrupted Power Supply (UPS system) and line voltage regulators must be included for the system with a support time of 50 to 60 minutes. Attach brochure.
9.7 | The lifetime of UPS must be at least 15 years.
9.8 | An Uninterrupted Power Supply (UPS system) conforming to IS-302 must be offered for the computer system.
9.9 | The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade type and it must be a minimum length of three (3) meters.
9.10. | The mains cable of the unit tendered for must be SABS colour coded. State details.
9.11 | The equipment quoted for must be protected against electromagnetic interference. State details of measures that will be taken to protect the unit against electromagnetic interference.
9.12 | Maintenance free batteries must be offered.
10 | AIR CONDITIONER
10.1 | Bidder must include any air conditioning that is considered necessary for optimal functioning of the unit.
10.2 | Any such air conditioning must be covered by the guarantee bided for on the unit and servicing of the air conditioning units for the guarantee period and must be included in the bid price.
10.3 | The air conditioner must also be included in the maintenance plan.
10.4 | Successful bidder to replace existing air conditioner with an air conditioner of 24,000 BTU. (Where applicable).
10.5 | The air conditioner must be but no less than 24,000 BTU. State details on the air conditioner to be supplied according to the room size and machine requirements.
10.6 | The Successful bidder must install an outlet-plumbing pipe for water overflow from air conditioner to be re-directed to a drainage system.
10.7 | Air-conditioning must be included in Maintenance plan

## Digital Fixed X-Ray Chest 60kW
DIGITAL FIXED CHEST X-RAY UNIT, FLOOR MOUNTED X-RAY TUBE AND ERECT BUCKY STAND WITH FIXED DETECTOR
ITEM | TECHNICAL SPECIFICATIONS
1.1 | The system must be capable of performing all general radiographic examinations of adults and children
1.2 | The system must be configured with a floor mounted tube stand on a rail and vertical stand detector combination
1.3 | The system must incorporate one flat panel digital detector that must fit into the the erect buckey :one active 43cm x 43 cm field size range .
1.4 | The system must be capable to capture a "low-energy"image and a "high-energy" image during a single examination. (Dual energy substraction chest imaging)
1.6 | System include LCD 19 inch sized, flat screen colour monitor for operators console /acquisition review station
1.7 | UPS provided to be mounted on a bracket/cabinet, not to be placed directly on the floor.
1.8 | The system must have removable grids.
1.9 | A wall mounted hanger/cabinet for the storage of removable grids and detector must be provided .
2 | X-RAY GENERATOR
2.1 | Generator should be of latest technology with high frequency and high speed starter and a power output shall be 60kW. Details are stated in supporting documents provided.
2.2 | A monitoring and display for heat status must be provided.
2.3 | AEC Functionality with 3-field ionization chamber
2.4 | APR Functionality (Anatomical programming) . Details are stated in supporting documents provided
2.5 | Maximum mA Range of 800mA and above at 100kV. Maximum mA output at 100 kV according to IEC60601-2-7 is stated.
2.6 | Selectable kV range of 40-125 kVp in 1 kVp increments
2.7 | The system must have selectable mAs. Details are stated.
2.8 | Shortest exposure time with AEC. Details are stated.
2.9 | Automatic mains compensation capability must be included.
3 | X-RAY TUBE AND FLOOR MOUNTED TUBE COLLUMN STAND
3.1 | X-ray tube should be a rotating anode high speed, compatible with digital flat panel detector. X-ray tube must be affixed to the floor mounted tube stand. Details including latest type of mode, are stated in supporting documents provided.
3.2 | Dual focal spot (small and large focus sizes) must be provided. Details of nominal focal spot sizes and focal spot combinations ,according to IEC 60336 prescription, are stated in supporting documents provided.
3.3 | Real time of tube overloading should be displayed
3.4 | Anode heat storage capacity must be 300 kHU and above. Details are stated in supporting documents provided.
3.5 | Tube overload protection is required
3.6 | Total Filtration at 70kV is 2.5mm (IEC 60601-1-3 Standard)
3.7 | Manual multileaf collimator with rotating flange collimator
3.8 | ± 45° swivelling/rotation of collimator
3.9 | Light localizer on/off switching
3.10. | Collimator must have electronic timer to terminate the light.
3.11 | X-ray Tube rotation range between ± 120 ° up to ± 135 ° . (± symbol means in both directions)
3.12 | Variable SID of the system
3.13 | Floor mounted free standing tube with rotational column design and telescopic tube arm movement, including longitudinal floor rails with 250cm travel and above.
3.14 | The longitudinal floor rail must have permanent stoppers /brackets at each end of the rail set, with shock absorbers to ensure the unit does not derail.
3.15 | Column rotation around its own axis at ± 90 ° (both directions)with a locking mechanism.
3.16 | Transverse movement
3.17 | Electromagnetic brakes with fully counter balanced mechanism
3.18 | Interlocking mechanism for 115cm /100cm and 180 cm SID . These distances must be electronically displayed on the x-ray tube /collimator housing.
4 | VERTICAL STAND WITH DETECTOR UNIT
4.1 | Vertical stand fitted with a digital flat panel - detector and three AEC measuring chambers
4.2 | Built in tray/support for 43 cm x 43 cm digital detector and removable grid
4.3 | The digital flat panel detector must be fixed . Option one
4.4 | The digital flat panel detector must be wireless and removable . Option two
4.5 | Vertical height of stand must be 2 metres and above
4.6 | Range of vertical movements of buckey: ( 1) 35cm-45 cm (minimum range height, and (2) 180cm- 200cm (maximum range height).Details are stated
4.7 | Patient hand grips for arms above head and on side of detector holder
5 | ONE X DIGITAL FLAT PANEL DETECTOR
5.1 | One x Digital FP detector of between 35cm x 43cm to 43cm x 43 cm field size.
5.2 | Amorphous silicon with cesium iodide scintillator
5.3 | Detectors have acquisition depth of 14 bits and above grayscale. Details are stated in supporting documents provided.
5.4 | Pixel pitch up to 160 micrometer (µm).
5.5 | Detector to have matrix in pixels of 2000 x 2000 and above
5.6 | For Wireless and Removable Flat Panel Detector (Option 2), include the following :
5.6.1 | Wi-Fi and LAN capability of Wireless detector . A tethered cable for wireless FP-detector option in case of Wi-Fi disruption or detector discharge .
5.6.2 | A vendor neutral wireless detector must be offered. Details are stated.
5.6.3 | The wireless detector must be able to trigger for acquisition without connectivity to the X-Ray generator to allow the wireless detector to be used in different areas.
5.6.4 | Images must be stored on the system even during battery failure.
5.6.5 | Battery pack.
5.6.6 | Battery charger.
5.6 | DQE must be between a range of 60%-70% @ 0 line pairs per millimeter / 50%- 60 % @ 0.05 line pairs per millimeter / 40 %-50% @ 1 line pairs per millimeter (lp/mm). DQE and MTF details are stated in supporting documents.
5.7 | Description of the scatter radiation grids used in conjunction with the detectors, are detailed in supporting documents provided.
5.8 | The data transmission/ reading speed from exposure to preview up to 10 seconds .
5.9 | The data transmission/ reading speed from exposure to full image up to 15 seconds .
5.10. | The flat panel battery:
5.10.1 | The battery should be able to acquire and transfer at least 90 images or more when fully charged. Details are stated in the supporting documents provided .
5.10.2 | The battery should be able to handle 5 hours of active use and 15 hours standby, when fully charged. Details are stated in the supporting documents provided .
5.10.3 | Charging time for each battery from fully depleted to fully charged should be less than 2 hours. Details are stated in the supporting documents provided.
5.10.4 | The batteries must be interchangeable between the 24 x30 cm and 43 x 43 cm panels.
5.10.5 | The battery charger must be light weight and should weigh less than 2 kg. Details of the weight of the battery charger are stated in the supporting documents provided .
5.10.6 | The battery must be able to operate at ambient temperature range of minumum of 5°C to maximum of 35 °C.
5.10.7 | The battery must be able to operate at ambient humidity of minimum humidity of 20 % to maximum of 85% RH.
5.10.8 | The battery must be able tolerate storage at a temperature range of minumum of -30 to to maximum of 60 degrees celsius.
5.10.9 | The battery must be able tolerate storage at a minimum humidity of 10 to maximum of 85% RH.
5.11 | The Grid
5.11.1 | An encasement with a built in grid that fits the 43cm x 43cm detector must be supplied.
5.11.2 | The grid must have at least:
5.11.3 | Grid ratio 8:1
5.11.4 | 100cm to 180cm focal distance
5.12 | Load capacity
5.12.1 | The uniform load that the detector must be able to withstand should be at least 150kg. Details are stated in the supporting documents provided .
5.12.2 | The local load (on an area of 40 mm diameter) that the detector must be able to withstand should be at least 100kg. Details are stated in the supporting documents provided .
6 | OPERATOR CONSOLE
6.1 | High resolution digital radiography system with DICOM network connection for image processing and image display on a preview monitor
6.2 | A diagnostic radiology LCD 19 inch imaging workstation based on latest flat panel detector technology for high image dynamics with excellent signal/noise ratio is required
6.3 | Application and examination database sotfware included.Details are stated
6.4 | Image processing software are included
6.5 | 64 bit 3.2GHz i5 Core Dual Microprocessor or similar
6.6 | RAM: 16-32GB
6.7 | 64 bit Operating system included (Windows or Similar)
6.8 | Image storage capacity up to 10 000 full format images on internal hard drive. Details of image storage capacity are stated in GB, as well as in maximum number of images in the supporting documents provided.
6.9 | No Autodeletion of images stored on internal hard drive. System to alert when near maximum storage capacity is reached.
6.10. | System must be DICOM compatible and capable of HIS/HMS/ RIS and PACS integration
6.11 | DVD /CD Writer inclusive of software for functionality of the DVD/CD Writer..
6.12 | Virus free software
F.13 | Ergonomic keyboard with mouse and function buttons
F.14 | Memory stick support/USB PORT for QC
F.15 | Touch screen monitor
7 | USER INTERFACE
7.1 | Patient data entering
7.2 | Menu
7.3 | Integrated generator operation
7.4 | Menu control
7.5 | Organ program selection
7.6 | Window position/width
7.7 | Horizontal/vertical image mirroring
7.8 | Image manipulation: rotation, cropping, contrast,brightness,sharpness, inversion
7.9 | Left/right, ap/pa working
7.10. | Configurable text annotation
7.11 | Filter selection for edge enhancement
7.12 | Filter selection for noise suppression
7.13 | Image pan and zoom capability
7.14 | measurement tools
8 | APPLICATION AND EVALUATION PROGRAMS
8.1 | Organ-related application program
8.2 | Post processing functionality
8.3 | Windowing (centre/width)
8.4 | Paging forward and back
8.5 | Multi-image display on the monitor (mosaic)
8.6 | Retro-collimation facilities (shutter)
8.7 | Image documentation in the background
8.8 | Multi-tasking technique
8.9 | Edge enhancement
8.1 | Noise suppression
8.11 | pan, zoom and image manipulation
8.12 | Individual energy substraction programs provided
9 | SYSTEM INTERPHASE
9.1 | Detector interface
9.2 | PACS ready
9.3 | Interface DICOM: "Basic Print"
9.4 | Interface DICOM: "Modality Worklist"
9.5 | Interface DICOM: "Storage"
9.6 | Interface DICOM: "Send as SCU"
9.7 | Interface DICOM "Image export/ transfer/query/retrieval to PACS/HIS/RIS
9.8 | Conformance Statement for all DICOM interfaces . Conformance Statements are included.
9.9 | Firewall protection
9.10. | The succesful bidder must ensure that the configuration between the processed images on the Digital Imaging Workstation and that transferred to the receivers of these images, such as PACS , Laser Printer , CD/DVD are properly calibrated/ reconfigured during installation, service and maintenance,repairs, and future software updates/upgrades .
12 | OPTIONAL ITEMS
12.1 | 1 X TB HARD DRIVE(external) ; DICOM 3 CONFIGURED TO STORE, TRANSFER AND PRINT X-RAY IMAGES.
12.2 | Motorized collimation control
12.3 | Tilting buckey
12.4 | Synchronized tube and detector movements. auto detector tracking and autotube positioning
12.5 | The system must have the ability to perform digital long length stitching images(scoliosis/long bones)
12.6 | Auto image stitching program included in the software package.
12.7 | TOR Rad or similar Image Resolution QC Test Tool
12.8 | A trolley with lockable castors, radiolucent top and an integrated detector holder that must move across the length of the trolley.
13 | RADIATION PROTECTION
13.1 | The system must have the ability to display Exposure Indicator values (EI/ S-values), immediately after exposure.
13.2 | X-ray Warning Notice for display at the X-ray Room entrance (not branded).
13.3 | An advanced Retake / Reject / Repeat analysis software package that includes analysis as well as calculations in percentages, must be provided for applicable systems.
13.4 | The applicable systems must have an integrated Dose Area Product (DAP) meter that includes skin dose equivalent measurement functions.
13.5 | A radiation report must be automatically generated and transferable via DICOM.
13.6 | Should be FDA or CE approved product
13.7 | Electrical safety conforms to standards for electrical safety IEC-60601
13.8 | The unit must comply to IEC-61267
13.9 | The equipment/system must be approved and licensed by the Directorate Radiation Control of DoH: South Africa
13.10. | A copy of a valid license issued in terms of the Hazardous Substance Act, Act No 15 of 1973 must be submitted with the tender. Failure to submit such a valid license may result in a tender not being considered.
13.11 | Acceptance testing must be performed as part of the installation. (Failure to do will lead to failure to pay
13.12 | Lead Glass window to be installed as per room requirement
13.13 | Safety aspects of Radiation dosage leakage should be spelt out
13.14 | DB boards and electrical filters to be installed
13.15 | QA TOOLS
13.16 | Display Exposure index values
13.17 | Radiation protection Lead equivalence of the radiation protection screen must be 0.5m Pb and must be convex, covering the radiographer from various angles against scatter radiation.
13.18 | Patient Conventional face shield to be included.
13.19 | Patient Tomosynthesis face shield to be included, removable and retractable.
13.20. | Emergency stop controls must be supplied on both sides of the gantry and on the Acquisition Workstation
13.21 | Interlocks to stop system movement when compression is applied must be provided.
13.22 | Interlock for premature release of the X-ray switch must be available
13.23 | Interlock preventing an exposure whenever the filter or light field is not positioned correctly must be possible.
13.24 | Phantom for calibration should be provided.
14 | POWER SUPPLY
14.1 | The bidder must install adequate electrical power supply for the optimal functionality of the equipment.
14.2 | The mains cable of the unit being quoted for must be the hospital grade type and it must be of an adequate length (minimum length of three (3) meters).
14.3 | The power input must be 220-240V, 50Hz AC. Bidders must ensure that the product quoted for is fitted appropriately with a 15 Amp SABS approved mains plug that is held together by two screws.
14.4 | A 3 phase in / out power supply must be offered. State details.
14.5 | The equipment tendered for must not overload and trip the hospital power in the area where it is installed. Resettable overcurrent breaker/s must be fitted for protection. State details.
14.6 | All available online Uninterrupted Power Supply (UPS system) and line voltage regulators must be included for the system with a support time of 50 to 60 minutes. Attach brochure.
14.7 | The lifetime of UPS must be at least 15 years.
14.8 | An Uninterrupted Power Supply (UPS system) conforming to IS-302 must be offered for the computer system.
14.9 | The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade type and it must be a minimum length of three (3) meters.
14.10. | The mains cable of the unit tendered for must be SABS colour coded. State details.
14.11 | The equipment quoted for must be protected against electromagnetic interference. State details of measures that will be taken to protect the unit against electromagnetic interference.
14.12 | Maintenance free batteries must be offered.
15 | AIR CONDITIONER
15.1 | Bidder must include any air conditioning that is considered necessary for optimal functioning of the unit.
15.2 | Any such air conditioning must be covered by the guarantee bided for on the unit and servicing of the air conditioning units for the guarantee period and must be included in the bid price.
15.3 | The air conditioner must also be included in the maintenance plan.
15.4 | Successful bidder to replace existing air conditioner with an air conditioner of 24,000 BTU. (Where applicable).
15.5 | The air conditioner must be but no less than 24,000 BTU. State details on the air conditioner to be supplied according to the room size and machine requirements.
15.6 | The Successful bidder must install an outlet-plumbing pipe for water overflow from air conditioner to be re-directed to a drainage system.
15.7 | Air-conditioning must be included in Maintenance plan

## Digital Fixed X-Ray Chest 80kW
DIGITAL FIXED CHEST X-RAY UNIT, FLOOR MOUNTED X-RAY TUBE AND ERECT BUCKY STAND WITH FIXED DETECTOR
ITEM | TECHNICAL SPECIFICATIONS
1.1 | The system must be capable of performing all general radiographic examinations of adults and children
1.2 | The system must be configured with a floor mounted tube stand on a rail and vertical stand detector combination
1.3 | The system must incorporate one flat panel digital detector that must fit into the the erect buckey :one active 43cm x 43 cm field size range .
1.4 | The system must be capable to capture a "low-energy"image and a "high-energy" image during a single examination. (Dual energy substraction chest imaging)
1.6 | System include LCD 19 inch sized, flat screen colour monitor for operators console /acquisition review station
1.7 | UPS provided to be mounted on a bracket/cabinet, not to be placed directly on the floor.
1.8 | The system must have removable grids.
1.9 | A wall mounted hanger/cabinet for the storage of removable grids and detector must be provided .
2 | X-RAY GENERATOR
2.1 | Generator should be of latest technology with high frequency and high speed starter and a power output shall be 80kW. Details are stated in supporting documents provided.
2.2 | A monitoring and display for heat status must be provided.
2.3 | AEC Functionality with 3-field ionization chamber
2.4 | APR Functionality (Anatomical programming) . Details are stated in supporting documents provided
2.5 | Maximum mA Range of 800mA and above at 100kV. Maximum mA output at 100 kV according to IEC60601-2-7 is stated.
2.6 | Selectable kV range of 40-125 kVp in 1 kVp increments
2.7 | The system must have selectable mAs. Details are stated.
2.8 | Shortest exposure time with AEC. Details are stated.
2.9 | Automatic mains compensation capability must be included.
3 | X-RAY TUBE AND FLOOR MOUNTED TUBE COLLUMN STAND
3.1 | X-ray tube should be a rotating anode high speed, compatible with digital flat panel detector. X-ray tube must be affixed to the floor mounted tube stand. Details including latest type of mode, are stated in supporting documents provided.
3.2 | Dual focal spot (small and large focus sizes) must be provided. Details of nominal focal spot sizes and focal spot combinations ,according to IEC 60336 prescription, are stated in supporting documents provided.
3.3 | Real time of tube overloading should be displayed
3.4 | Anode heat storage capacity must be 300 kHU and above. Details are stated in supporting documents provided.
3.5 | Tube overload protection is required
3.6 | Total Filtration at 70kV is 2.5mm (IEC 60601-1-3 Standard)
3.7 | Manual multileaf collimator with rotating flange collimator
3.8 | ± 45° swivelling/rotation of collimator
3.9 | Light localizer on/off switching
3.10. | Collimator must have electronic timer to terminate the light.
3.11 | X-ray Tube rotation range between ± 120 ° up to ± 135 ° . (± symbol means in both directions)
3.12 | Variable SID of the system
3.13 | Floor mounted free standing tube with rotational column design and telescopic tube arm movement, including longitudinal floor rails with 250cm travel and above.
3.14 | The longitudinal floor rail must have permanent stoppers /brackets at each end of the rail set, with shock absorbers to ensure the unit does not derail.
3.15 | Column rotation around its own axis at ± 90 ° (both directions)with a locking mechanism.
3.16 | Transverse movement
3.17 | Electromagnetic brakes with fully counter balanced mechanism
3.18 | Interlocking mechanism for 115cm /100cm and 180 cm SID . These distances must be electronically displayed on the x-ray tube /collimator housing.
4 | VERTICAL STAND WITH DETECTOR UNIT
4.1 | Vertical stand fitted with a digital flat panel - detector and three AEC measuring chambers
4.2 | Built in tray/support for 43 cm x 43 cm digital detector and removable grid
4.3 | The digital flat panel detector must be fixed . Option one
4.4 | The digital flat panel detector must be wireless and removable . Option two
4.5 | Vertical height of stand must be 2 metres and above
4.6 | Range of vertical movements of buckey: ( 1) 35cm-45 cm (minimum range height, and (2) 180cm- 200cm (maximum range height).Details are stated
4.7 | Patient hand grips for arms above head and on side of detector holder
5 | ONE X DIGITAL FLAT PANEL DETECTOR
5.1 | One x Digital FP detector of between 35cm x 43cm to 43cm x 43 cm field size.
5.2 | Amorphous silicon with cesium iodide scintillator
5.3 | Detectors have acquisition depth of 14 bits and above grayscale. Details are stated in supporting documents provided.
5.4 | Pixel pitch up to 160 micrometer (µm).
5.5 | Detector to have matrix in pixels of 2000 x 2000 and above
5.6 | For Wireless and Removable Flat Panel Detector (Option 2), include the following :
5.6.1 | Wi-Fi and LAN capability of Wireless detector . A tethered cable for wireless FP-detector option in case of Wi-Fi disruption or detector discharge .
5.6.2 | A vendor neutral wireless detector must be offered. Details are stated.
5.6.3 | The wireless detector must be able to trigger for acquisition without connectivity to the X-Ray generator to allow the wireless detector to be used in different areas.
5.6.4 | Images must be stored on the system even during battery failure.
5.6.5 | Battery pack.
5.6.6 | Battery charger.
5.6 | DQE must be between a range of 60%-70% @ 0 line pairs per millimeter / 50%- 60 % @ 0.05 line pairs per millimeter / 40 %-50% @ 1 line pairs per millimeter (lp/mm). DQE and MTF details are stated in supporting documents.
5.7 | Description of the scatter radiation grids used in conjunction with the detectors, are detailed in supporting documents provided.
5.8 | The data transmission/ reading speed from exposure to preview up to 10 seconds .
5.9 | The data transmission/ reading speed from exposure to full image up to 15 seconds .
5.10. | The flat panel battery:
5.10.1 | The battery should be able to acquire and transfer at least 90 images or more when fully charged. Details are stated in the supporting documents provided .
5.10.2 | The battery should be able to handle 5 hours of active use and 15 hours standby, when fully charged. Details are stated in the supporting documents provided .
5.10.3 | Charging time for each battery from fully depleted to fully charged should be less than 2 hours. Details are stated in the supporting documents provided.
5.10.4 | The batteries must be interchangeable between the 24 x30 cm and 43 x 43 cm panels.
5.10.5 | The battery charger must be light weight and should weigh less than 2 kg. Details of the weight of the battery charger are stated in the supporting documents provided .
5.10.6 | The battery must be able to operate at ambient temperature range of minumum of 5°C to maximum of 35 °C.
5.10.7 | The battery must be able to operate at ambient humidity of minimum humidity of 20 % to maximum of 85% RH.
5.10.8 | The battery must be able tolerate storage at a temperature range of minumum of -30 to to maximum of 60 degrees celsius.
5.10.9 | The battery must be able tolerate storage at a minimum humidity of 10 to maximum of 85% RH.
5.11 | The Grid
5.11.1 | An encasement with a built in grid that fits the 43cm x 43cm detector must be supplied.
5.11.2 | The grid must have at least:
5.11.3 | Grid ratio 8:1
5.11.4 | 100cm to 180cm focal distance
5.12 | Load capacity
5.12.1 | The uniform load that the detector must be able to withstand should be at least 150kg. Details are stated in the supporting documents provided .
5.12.2 | The local load (on an area of 40 mm diameter) that the detector must be able to withstand should be at least 100kg. Details are stated in the supporting documents provided .
6 | OPERATOR CONSOLE
6.1 | High resolution digital radiography system with DICOM network connection for image processing and image display on a preview monitor
6.2 | A diagnostic radiology LCD 19 inch imaging workstation based on latest flat panel detector technology for high image dynamics with excellent signal/noise ratio is required
6.3 | Application and examination database sotfware included.Details are stated
6.4 | Image processing software are included
6.5 | 64 bit 3.2GHz i5 Core Dual Microprocessor or similar
6.6 | RAM: 16-32GB
6.7 | 64 bit Operating system included (Windows or Similar)
6.8 | Image storage capacity up to 10 000 full format images on internal hard drive. Details of image storage capacity are stated in GB, as well as in maximum number of images in the supporting documents provided.
6.9 | No Autodeletion of images stored on internal hard drive. System to alert when near maximum storage capacity is reached.
6.10. | System must be DICOM compatible and capable of HIS/HMS/ RIS and PACS integration
6.11 | DVD /CD Writer inclusive of software for functionality of the DVD/CD Writer..
6.12 | Virus free software
F.13 | Ergonomic keyboard with mouse and function buttons
F.14 | Memory stick support/USB PORT for QC
F.15 | Touch screen monitor
7 | USER INTERFACE
7.1 | Patient data entering
7.2 | Menu
7.3 | Integrated generator operation
7.4 | Menu control
7.5 | Organ program selection
7.6 | Window position/width
7.7 | Horizontal/vertical image mirroring
7.8 | Image manipulation: rotation, cropping, contrast,brightness,sharpness, inversion
7.9 | Left/right, ap/pa working
7.10. | Configurable text annotation
7.11 | Filter selection for edge enhancement
7.12 | Filter selection for noise suppression
7.13 | Image pan and zoom capability
7.14 | measurement tools
8 | APPLICATION AND EVALUATION PROGRAMS
8.1 | Organ-related application program
8.2 | Post processing functionality
8.3 | Windowing (centre/width)
8.4 | Paging forward and back
8.5 | Multi-image display on the monitor (mosaic)
8.6 | Retro-collimation facilities (shutter)
8.7 | Image documentation in the background
8.8 | Multi-tasking technique
8.9 | Edge enhancement
8.1 | Noise suppression
8.11 | pan, zoom and image manipulation
8.12 | Individual energy substraction programs provided
9 | SYSTEM INTERPHASE
9.1 | Detector interface
9.2 | PACS ready
9.3 | Interface DICOM: "Basic Print"
9.4 | Interface DICOM: "Modality Worklist"
9.5 | Interface DICOM: "Storage"
9.6 | Interface DICOM: "Send as SCU"
9.7 | Interface DICOM "Image export/ transfer/query/retrieval to PACS/HIS/RIS
9.8 | Conformance Statement for all DICOM interfaces . Conformance Statements are included.
9.9 | Firewall protection
9.10. | The succesful bidder must ensure that the configuration between the processed images on the Digital Imaging Workstation and that transferred to the receivers of these images, such as PACS , Laser Printer , CD/DVD are properly calibrated/ reconfigured during installation, service and maintenance,repairs, and future software updates/upgrades .
12 | OPTIONAL ITEMS
12.1 | 1 X TB HARD DRIVE(external) ; DICOM 3 CONFIGURED TO STORE, TRANSFER AND PRINT X-RAY IMAGES.
12.2 | Motorized collimation control
12.3 | Tilting buckey
12.4 | Synchronized tube and detector movements. auto detector tracking and autotube positioning
12.5 | The system must have the ability to perform digital long length stitching images(scoliosis/long bones)
12.6 | Auto image stitching program included in the software package.
12.7 | TOR Rad or similar Image Resolution QC Test Tool
12.8 | A trolley with lockable castors, radiolucent top and an integrated detector holder that must move across the length of the trolley.
13 | RADIATION PROTECTION
13.1 | The system must have the ability to display Exposure Indicator values (EI/ S-values), immediately after exposure.
13.2 | X-ray Warning Notice for display at the X-ray Room entrance (not branded).
13.3 | An advanced Retake / Reject / Repeat analysis software package that includes analysis as well as calculations in percentages, must be provided for applicable systems.
13.4 | The applicable systems must have an integrated Dose Area Product (DAP) meter that includes skin dose equivalent measurement functions.
13.5 | A radiation report must be automatically generated and transferable via DICOM.
13.6 | Should be FDA or CE approved product
13.7 | Electrical safety conforms to standards for electrical safety IEC-60601
13.8 | The unit must comply to IEC-61267
13.9 | The equipment/system must be approved and licensed by the Directorate Radiation Control of DoH: South Africa
13.10. | A copy of a valid license issued in terms of the Hazardous Substance Act, Act No 15 of 1973 must be submitted with the tender. Failure to submit such a valid license may result in a tender not being considered.
13.11 | Acceptance testing must be performed as part of the installation. (Failure to do will lead to failure to pay
13.12 | Lead Glass window to be installed as per room requirement
13.13 | Safety aspects of Radiation dosage leakage should be spelt out
13.14 | DB boards and electrical filters to be installed
13.15 | QA TOOLS
13.16 | Display Exposure index values
13.17 | Radiation protection Lead equivalence of the radiation protection screen must be 0.5m Pb and must be convex, covering the radiographer from various angles against scatter radiation.
13.18 | Patient Conventional face shield to be included.
13.19 | Patient Tomosynthesis face shield to be included, removable and retractable.
13.20. | Emergency stop controls must be supplied on both sides of the gantry and on the Acquisition Workstation
13.21 | Interlocks to stop system movement when compression is applied must be provided.
13.22 | Interlock for premature release of the X-ray switch must be available
13.23 | Interlock preventing an exposure whenever the filter or light field is not positioned correctly must be possible.
13.24 | Phantom for calibration should be provided.
14 | POWER SUPPLY
14.1 | The bidder must install adequate electrical power supply for the optimal functionality of the equipment.
14.2 | The mains cable of the unit being quoted for must be the hospital grade type and it must be of an adequate length (minimum length of three (3) meters).
14.3 | The power input must be 220-240V, 50Hz AC. Bidders must ensure that the product quoted for is fitted appropriately with a 15 Amp SABS approved mains plug that is held together by two screws.
14.4 | A 3 phase in / out power supply must be offered. State details.
14.5 | The equipment tendered for must not overload and trip the hospital power in the area where it is installed. Resettable overcurrent breaker/s must be fitted for protection. State details.
14.6 | All available online Uninterrupted Power Supply (UPS system) and line voltage regulators must be included for the system with a support time of 50 to 60 minutes. Attach brochure.
14.7 | The lifetime of UPS must be at least 15 years.
14.8 | An Uninterrupted Power Supply (UPS system) conforming to IS-302 must be offered for the computer system.
14.9 | The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade type and it must be a minimum length of three (3) meters.
14.10. | The mains cable of the unit tendered for must be SABS colour coded. State details.
14.11 | The equipment quoted for must be protected against electromagnetic interference. State details of measures that will be taken to protect the unit against electromagnetic interference.
14.12 | Maintenance free batteries must be offered.
15 | AIR CONDITIONER
15.1 | Bidder must include any air conditioning that is considered necessary for optimal functioning of the unit.
15.2 | Any such air conditioning must be covered by the guarantee bided for on the unit and servicing of the air conditioning units for the guarantee period and must be included in the bid price.
15.3 | The air conditioner must also be included in the maintenance plan.
15.4 | Successful bidder to replace existing air conditioner with an air conditioner of 24,000 BTU. (Where applicable).
15.5 | The air conditioner must be but no less than 24,000 BTU. State details on the air conditioner to be supplied according to the room size and machine requirements.
15.6 | The Successful bidder must install an outlet-plumbing pipe for water overflow from air conditioner to be re-directed to a drainage system.
15.7 | Air-conditioning must be included in Maintenance plan

## Digital Ceiling X-RayChest 80kW
DIGITAL FIXED CHEST X-RAY UNIT, CEILING SUSPENDED X-RAY TUBE AND ERECT BUCKY STAND WITH FIXED DETECTOR
ITEM | TECHNICAL SPECIFICATIONS
1.1 | The system must be capable of performing all general radiographic examinations of adults and children
1.2 | The system must be configured with the x-ray tube affixed to a ceiling suspended tube assembly and a vertical buckey stand combination.
1.3 | The system must incorporate one flat panel digital detector that must fit into the erect buckey : 43cm x 43 cm field size range .
1.4 | The system must be capable to capture a "low- energy"image and a "high-energy" image during a single examination. (Dual energy substraction chest imaging)
1.6 | System include LCD 19 inch sized, flat screen colour monitor for operators console /acquisition review station
1.7 | UPS provided to be mounted on a bracket/cabinet, not to be placed directly on the floor.
1.8 | The system must have removable grids.
1.9 | A wall mounted hanger/cabinet for the storage of removable grids and detector must be provided .
2 | X-RAY GENERATOR
2.1 | Generator should be of latest technology with high frequency and high speed starter and a power output of 80 kW and above range. Details are stated in supporting documents provided.
2.2 | A monitoring and display for heat status must be provided.
2.3 | AEC Functionality with 3-field ionization chamber
2.4 | APR Functionality (Anatomical programming) . Details are stated in supporting documents provided
2.5 | Maximum mA Range of 800mA and above at 100kV. Maximum mA output at 100 kV according to IEC60601-2-7 is stated.
2.6 | Selectable kV range of 40-125 kVp in 1 kVp increments
2.7 | The system must have selectable mAs. Details are stated.
2.8 | Shortest exposure time with AEC. Details are stated.
2.9 | Automatic mains compensation capability must be included.
3 | X-RAY TUBE HOUSING AFFIXED TO CEILING SUSPENSION: CEILING SUSPENDED TUBE ASSEMBLY
3.1 | X-ray tube should be a rotating anode high speed, compatible with flat panel detector. X-ray tube must be affixed to the ceiling suspended tube support .Details including latest type of mode, are stated in supporting documents provided.
3.2 | Dual focal spot (small and large focus sizes) must be provided. Details of nominal focal spot sizes and focal spot combinations ,according to IEC 60336 prescription, are stated in supporting documents provided.
3.3 | Real time of tube overloading should be displayed
3.4 | Anode heat storage capacity must be 300 kHU and above. Details are stated in supporting documents provided.
3.5 | Tube overload protection is required
3.6 | Total Filtration at 70kV is 2.5mm (IEC 60601-1-3 Standard)
3.7 | Manual multileaf collimator with rotating flange collimator
3.8 | ± 45° swivelling/rotation of collimator
3.9 | Light localizer on/off switching
3.10. | Collimator must have electronic timer to terminate the light.
3.11 | X-ray Tube rotation range between ± 120 ° up to ± 135 ° . (± symbol means in both directions)
3.12 | Variable SID of the system
3.13 | A ceiling suspended x-ray tube support with a telescopic column must be offered.
3.14 | The ceiling syspended x-ray tube system must have a transverse travel range. (3000mm to 4000mm). Details of range are stated.
3.15 | A ceiling suspended x-ray tube support with a telescopic column must be offered.
3.16 | A centering device in longitudinal and transversal directions to be provided.
3.17 | The vertical travel of the column must not be less than 50cm from the floor. State range of offered system.
3.18 | Electromagnetic brakes with fully counter balanced mechanism
3.19 | Interlocking mechanism for 100 cm/115cm and 180 cm SID . These distances must be electronically displayed on the x-ray tube /collimator housing.
3.20. | A solid control arm with integrated handles as well as an electronic digital angle display panel and individual brake release buttons including a single button releasing all locks to be offered.
3.21. | The digital display panel must also indicate the SID (source image distance) of the unit from the bucky table.
4 | VERTICAL STAND WITH DETECTOR UNIT
4.1 | Vertical stand fitted with a digital flat panel - detector and three AEC measuring chambers
4.2 | Built in tray/support for digital detector and removable grid
4.3 | The digital flat panel detector must be fixed.
4.5 | Vertical height of stand must be 2 metres and above
4.6 | Range of vertical movements of buckey: ( 1) 35cm-45 cm (minimum range height, and (2) 180cm- 200cm (maximum range height).Details are stated
4.7 | Patient hand grips for arms above head and on side of detector holder
5 | DIGITAL FLAT PANEL DETECTOR SYSTEM
5.1 | One x Digital FP detector of 43cm x 43 cm field size.
5.2 | Amorphous silicon with cesium iodide scintillator
5.3 | Detectors have acquisition depth of 14 bits and above grayscale. Details are stated in supporting documents provided.
5.4 | Pixel pitch up to 160 micrometer (µm).
5.5 | Detector to have matrix in pixels of 2000 x 2000 and above
5.7 | DQE must be between a range of 60%-70% @ 0 line pairs per millimeter / 50%- 60 % @ 0.05 line pairs per millimeter / 40 %-50% @ 1 line pairs per millimeter (lp/mm). DQE and MTF details are stated in supporting documents.
5.8 | Description of the scattered radiation grid are detailed in supporting documents provided.
5.9 | The data transmission/ reading speed from exposure to preview is 10 seconds and lower .
5.10. | The data transmission/ reading speed from exposure to full image is 15 seconds and lower .
5.11 | The flat panel battery:
5.11.1 | The battery should be able to acquire and transfer at least 90 images or more when fully charged. Details are stated in the supporting documents provided .
5.11.2 | The battery should be able to handle 5 hours of active use and 15 hours standby, when fully charged. Details are stated in the supporting documents provided .
5.11.3 | Charging time for each battery from fully depleted to fully charged should be less than 2 hours. Details are stated in the supporting documents provided.
5.11.4 | The batteries must be interchangeable between the 24 x30 cm and 43 x 43 cm panels.
5.11.5 | The battery charger must be light weight and should weigh less than 2 kg. Details of the weight of the battery charger are stated in the supporting documents provided .
5.11.6 | The battery must be able to operate at ambient temperature range of minumum of 5°C to maximum of 35 °C.
5.11.7 | The battery must be able to operate at ambient humidity of minimum humidity of 20 % to maximum of 85% RH.
5.11.8 | The battery must be able tolerate storage at a temperature range of minumum of -30 to to maximum of 60 degrees celsius.
5.11.9 | The battery must be able tolerate storage at a minimum humidity of 10 to maximum of 85% RH.
5.12 | The Grid
5.12.1 | An encasement with a built in grid that fits the 43cm x 43cm detector must be supplied.
5.12.2 | The grid must have at least:
5.12.3 | Grid ratio 8:1
5.12.4 | 100cm to 180cm focal distance
5.13 | Load capacity
5.13.1 | The uniform load that the detector must be able to withstand should be at least 150kg. Details are stated in the supporting documents provided .
5.13.2 | The local load (on an area of 40 mm diameter) that the detector must be able to withstand should be at least 100kg. Details are stated in the supporting documents provided .
6 | OPERATOR CONSOLE
6.1 | High resolution digital radiography system with DICOM network connection for image processing and image display on a preview monitor
6.2 | A diagnostic radiology LCD 19 inch imaging workstation based on latest flat panel detector technology for high image dynamics with excellent signal/noise ratio is required
6.3 | Application and examination database sotfware included.Details are stated
6.4 | Image processing software are included
6.5 | 64 bit 3.2GHz i5 Core Dual Microprocessor or similar
6.6 | RAM: 16-32 GB
6.7 | 64 bit Operating system included (Windows or Similar)
6.8 | Image storage capacity up to 10 000 full format images on internal hard drive. Details of image storage capacity are stated in GB, as well as in maximum number of images in the supporting documents provided.
6.9 | No Autodeletion of images stored on internal hard drive. System to alert when near maximum storage capacity is reached.
6.10. | System must be DICOM compatible and capable of HIS/HMS/ RIS and PACS integration
6.11 | DVD /CD Writer inclusive of software for functionality of the DVD/CD Writer..
6.12 | Virus free software
6.13 | Ergonomic keyboard with mouse and function buttons
6.14 | Memory stick support/USB PORT for QC
6.15 | Touch screen monitor
7 | USER INTERFACE
7.1 | Patient data entering
7.2 | Menu
7.3 | Integrated generator operation
7.4 | Menu control
7.5 | Organ program selection
7.6 | Window position/width
7.7 | Horizontal/vertical image mirroring
7.8 | Image manipulation: rotation, cropping, contrast,brightness,sharpness, inversion
7.9 | Left/right, ap/pa working
7.1 | Configurable text annotation
7.11 | Filter selection for edge enhancement
7.12 | Filter selection for noise suppression
7.13 | Image pan and zoom capability
7.14 | measurement tools
8 | CONTROL CONSOLE
8.1 | Patient data entering
8.2 | Manual
8.3 | Bar code reader
8.4 | Monitor size of 19 inch(diagonal) Touch screen
8.5 | Colour
8.6 | Keyboard with mouse and function buttons
8.7 | Spill proof keyboard design
8.8 | Mouse
8.9 | Menu
8.10. | Integrated generator operation
8.11 | Menu control
8.12 | Organ program selection
8.13 | Window position/width
8.14 | Horizontal/vertical image mirroring
8.15 | Image rotation
8.16 | Left/right, ap/pa working
8.17 | Configurable text annotation
8.18 | Filter selection for edge enhancement
8.19 | Filter selection for noise suppression
8.20. | Image zoom capability
9 | APPLICATION AND EVALUATION PROGRAMS
9.1 | Organ-related application program
9.2 | Post processing functionality
9.3 | Windowing (centre/width)
9.4 | Paging forward and back
9.5 | Multi-image display on the monitor (mosaic)
9.6 | Retro-collimation facilities (shutter)
9.7 | Image documentation in the background
9.8 | Multi-tasking technique
9.9 | Edge enhancement
9.1 | Noise suppression
9.11 | Zoom
9.12 | Individual energy substraction programs provided
10 | SYSTEM INTERPHASE
10.1 | Detector interphase
10.2 | Interface DICOM: "Basic Print"
10.3 | Interface DICOM: "Modality Worklist"
10.4 | Interface DICOM: "Storage"
10.5 | Interface DICOM: "Send as SCU"
10.6 | Interface DICOM "Image export/transfer to PACS/HIS/RIS
10.7 | Is there a "Confirmation Statement" for all DICOM interfaces
10.8 | Firewall protection
10.9 | The succesful bidder must ensure that the configuration between the processed images on the Digital Imaging Workstation and that transferred to the receivers of these images, such as PACS , Laser Printer , CD/DVD are properly calibrated/ reconfigured during installation, service and maintenance,repairs, and future software updates/upgrades .
11 | OPTIONAL ITEMS
11.1 | 1 X TB HARD DRIVE(external) ; DICOM CONFIGURED TO STORE, TRANSFER AND PRINT X-RAY IMAGES.
11.2 | Motorized collimation control as an option
11.3 | Tilting buckey as an option
11.4 | Synchronized tube and detector movements. Auto detector tracking and autotube positioning
11.5 | The system must have the ability to perform digital long length stitching images(scoliosis/long bones) -OPTIONAL
11.6 | Auto image stitching program included in the software package.
11.7 | TOR Rad or similar Image Resolution QC Test Tool
11.8 | A trolley with lockable castors, radiolucent top and an integrated detector holder that must move across the length of the trolley.
12 | RADIATION PROTECTION
12.1 | The system must have the ability to display Exposure Indicator values (EI/ S-values), immediately after exposure.
12.2 | X-ray Warning Notice for display at the X-ray Room entrance (not branded).
12.3 | An advanced Retake / Reject / Repeat analysis software package that includes analysis as well as calculations in percentages, must be provided for applicable systems.
12.4 | The applicable systems must have an integrated Dose Area Product (DAP) meter that includes skin dose equivalent measurement functions.
12.5 | A radiation report must be automatically generated and transferable via DICOM.
12.6 | Should be FDA or CE approved product
12.7 | Electrical safety conforms to standards for electrical safety IEC-60601
12.8 | The unit must comply to IEC-61267
12.9 | The equipment/system must be approved and licensed by the Directorate Radiation Control of DoH: South Africa
12.10. | A copy of a valid license issued in terms of the Hazardous Substance Act, Act No 15 of 1973 must be submitted with the tender. Failure to submit such a valid license may result in a tender not being considered.
12.11 | Acceptance testing must be performed as part of the installation. (Failure to do will lead to failure to pay
12.12 | Lead Glass window to be installed as per room requirement
12.13 | Safety aspects of Radiation dosage leakage should be spelt out
12.14 | DB boards and electrical filters to be installed
12.15 | QA TOOLS
12.16 | Display Exposure index values
12.17 | Radiation protection Lead equivalence of the radiation protection screen must be 0.5m Pb and must be convex, covering the radiographer from various angles against scatter radiation.
12.18 | Patient Conventional face shield to be included.
12.19 | Patient Tomosynthesis face shield to be included, removable and retractable.
12.20. | Emergency stop controls must be supplied on both sides of the gantry and on the Acquisition Workstation
12.21 | Interlocks to stop system movement when compression is applied must be provided.
12.22 | Interlock for premature release of the X-ray switch must be available
12.23 | Interlock preventing an exposure whenever the filter or light field is not positioned correctly must be possible.
12.24 | Phantom for calibration should be provided.
13 | POWER SUPPLY
13.1 | The bidder must install adequate electrical power supply for the optimal functionality of the equipment.
13.2 | The mains cable of the unit being quoted for must be the hospital grade type and it must be of an adequate length (minimum length of three (3) meters).
13.3 | The power input must be 220-240V, 50Hz AC. Bidders must ensure that the product quoted for is fitted appropriately with a 15 Amp SABS approved mains plug that is held together by two screws.
13.4 | A 3 phase in / out power supply must be offered. State details.
13.5 | The equipment tendered for must not overload and trip the hospital power in the area where it is installed. Resettable overcurrent breaker/s must be fitted for protection. State details.
13.6 | All available online Uninterrupted Power Supply (UPS system) and line voltage regulators must be included for the system with a support time of 50 to 60 minutes. Attach brochure.
13.7 | The lifetime of UPS must be at least 15 years.
13.8 | An Uninterrupted Power Supply (UPS system) conforming to IS-302 must be offered for the computer system.
13.9 | The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade type and it must be a minimum length of three (3) meters.
13.10. | The mains cable of the unit tendered for must be SABS colour coded. State details.
13.11 | The equipment quoted for must be protected against electromagnetic interference. State details of measures that will be taken to protect the unit against electromagnetic interference.
13.12 | Maintenance free batteries must be offered.
14 | AIR CONDITIONER
14.1 | Bidder must include any air conditioning that is considered necessary for optimal functioning of the unit.
14.2 | Any such air conditioning must be covered by the guarantee bided for on the unit and servicing of the air conditioning units for the guarantee period and must be included in the bid price.
14.3 | The air conditioner must also be included in the maintenance plan.
14.4 | Successful bidder to replace existing air conditioner with an air conditioner of 24,000 BTU. (Where applicable).
14.5 | The air conditioner must be but no less than 24,000 BTU. State details on the air conditioner to be supplied according to the room size and machine requirements.
14.6 | The Successful bidder must install an outlet-plumbing pipe for water overflow from air conditioner to be re-directed to a drainage system.
14.7 | Air-conditioning must be included in Maintenance plan

## Digital Ceiling X-RayChest 60kW
DIGITAL FIXED CHEST X-RAY UNIT, CEILING SUSPENDED X-RAY TUBE AND ERECT BUCKY STAND WITH FIXED DETECTOR
ITEM | TECHNICAL SPECIFICATIONS
1.1 | The system must be capable of performing all general radiographic examinations of adults and children
1.2 | The system must be configured with the x-ray tube affixed to a ceiling suspended tube assembly and a vertical buckey stand combination.
1.3 | The system must incorporate one flat panel digital detector that must fit into the erect buckey : 43cm x 43 cm field size range .
1.4 | The system must be capable to capture a "low- energy"image and a "high-energy" image during a single examination. (Dual energy substraction chest imaging)
1.6 | System include LCD 19 inch sized, flat screen colour monitor for operators console /acquisition review station
1.7 | UPS provided to be mounted on a bracket/cabinet, not to be placed directly on the floor.
1.8 | The system must have removable grids.
1.9 | A wall mounted hanger/cabinet for the storage of removable grids and detector must be provided .
2 | X-RAY GENERATOR
2.1 | Generator should be of latest technology with high frequency and high speed starter and a power output of 60 kW. Details are stated in supporting documents provided.
2.2 | A monitoring and display for heat status must be provided.
2.3 | AEC Functionality with 3-field ionization chamber
2.4 | APR Functionality (Anatomical programming) . Details are stated in supporting documents provided
2.5 | Maximum mA Range of 800mA and above at 100kV. Maximum mA output at 100 kV according to IEC60601-2-7 is stated.
2.6 | Selectable kV range of 40-125 kVp in 1 kVp increments
2.7 | The system must have selectable mAs. Details are stated.
2.8 | Shortest exposure time with AEC. Details are stated.
2.9 | Automatic mains compensation capability must be included.
3 | X-RAY TUBE HOUSING AFFIXED TO CEILING SUSPENSION: CEILING SUSPENDED TUBE ASSEMBLY
3.1 | X-ray tube should be a rotating anode high speed, compatible with flat panel detector. X-ray tube must be affixed to the ceiling suspended tube support .Details including latest type of mode, are stated in supporting documents provided.
3.2 | Dual focal spot (small and large focus sizes) must be provided. Details of nominal focal spot sizes and focal spot combinations ,according to IEC 60336 prescription, are stated in supporting documents provided.
3.3 | Real time of tube overloading should be displayed
3.4 | Anode heat storage capacity must be 300 kHU and above. Details are stated in supporting documents provided.
3.5 | Tube overload protection is required
3.6 | Total Filtration at 70kV is 2.5mm (IEC 60601-1-3 Standard)
3.7 | Manual multileaf collimator with rotating flange collimator
3.8 | ± 45° swivelling/rotation of collimator
3.9 | Light localizer on/off switching
3.10. | Collimator must have electronic timer to terminate the light.
3.11 | X-ray Tube rotation range between ± 120 ° up to ± 135 ° . (± symbol means in both directions)
3.12 | Variable SID of the system
3.13 | A ceiling suspended x-ray tube support with a telescopic column must be offered.
3.14 | The ceiling syspended x-ray tube system must have a transverse travel range. (3000mm to 4000mm). Details of range are stated.
3.15 | A ceiling suspended x-ray tube support with a telescopic column must be offered.
3.16 | A centering device in longitudinal and transversal directions to be provided.
3.17 | The vertical travel of the column must not be less than 50cm from the floor. State range of offered system.
3.18 | Electromagnetic brakes with fully counter balanced mechanism
3.19 | Interlocking mechanism for 100 cm/115cm and 180 cm SID . These distances must be electronically displayed on the x-ray tube /collimator housing.
3.20. | A solid control arm with integrated handles as well as an electronic digital angle display panel and individual brake release buttons including a single button releasing all locks to be offered.
3.21. | The digital display panel must also indicate the SID (source image distance) of the unit from the bucky table.
4 | VERTICAL STAND WITH DETECTOR UNIT
4.1 | Vertical stand fitted with a digital flat panel - detector and three AEC measuring chambers
4.2 | Built in tray/support for digital detector and removable grid
4.3 | The digital flat panel detector must be fixed.
4.5 | Vertical height of stand must be 2 metres and above
4.6 | Range of vertical movements of buckey: ( 1) 35cm-45 cm (minimum range height, and (2) 180cm- 200cm (maximum range height).Details are stated
4.7 | Patient hand grips for arms above head and on side of detector holder
5 | DIGITAL FLAT PANEL DETECTOR SYSTEM
5.1 | One x Digital FP detector of 43cm x 43 cm field size.
5.2 | Amorphous silicon with cesium iodide scintillator
5.3 | Detectors have acquisition depth of 14 bits and above grayscale. Details are stated in supporting documents provided.
5.4 | Pixel pitch up to 160 micrometer (µm).
5.5 | Detector to have matrix in pixels of 2000 x 2000 and above
5.7 | DQE must be between a range of 60%-70% @ 0 line pairs per millimeter / 50%- 60 % @ 0.05 line pairs per millimeter / 40 %-50% @ 1 line pairs per millimeter (lp/mm). DQE and MTF details are stated in supporting documents.
5.8 | Description of the scattered radiation grid are detailed in supporting documents provided.
5.9 | The data transmission/ reading speed from exposure to preview is 10 seconds and lower .
5.10. | The data transmission/ reading speed from exposure to full image is 15 seconds and lower .
5.11 | The flat panel battery:
5.11.1 | The battery should be able to acquire and transfer at least 90 images or more when fully charged. Details are stated in the supporting documents provided .
5.11.2 | The battery should be able to handle 5 hours of active use and 15 hours standby, when fully charged. Details are stated in the supporting documents provided .
5.11.3 | Charging time for each battery from fully depleted to fully charged should be less than 2 hours. Details are stated in the supporting documents provided.
5.11.4 | The batteries must be interchangeable between the 24 x30 cm and 43 x 43 cm panels.
5.11.5 | The battery charger must be light weight and should weigh less than 2 kg. Details of the weight of the battery charger are stated in the supporting documents provided .
5.11.6 | The battery must be able to operate at ambient temperature range of minumum of 5°C to maximum of 35 °C.
5.11.7 | The battery must be able to operate at ambient humidity of minimum humidity of 20 % to maximum of 85% RH.
5.11.8 | The battery must be able tolerate storage at a temperature range of minumum of -30 to to maximum of 60 degrees celsius.
5.11.9 | The battery must be able tolerate storage at a minimum humidity of 10 to maximum of 85% RH.
5.12 | The Grid
5.12.1 | An encasement with a built in grid that fits the 43cm x 43cm detector must be supplied.
5.12.2 | The grid must have at least:
5.12.3 | Grid ratio 8:1
5.12.4 | 100cm to 180cm focal distance
5.13 | Load capacity
5.13.1 | The uniform load that the detector must be able to withstand should be at least 150kg. Details are stated in the supporting documents provided .
5.13.2 | The local load (on an area of 40 mm diameter) that the detector must be able to withstand should be at least 100kg. Details are stated in the supporting documents provided .
6 | OPERATOR CONSOLE
6.1 | High resolution digital radiography system with DICOM network connection for image processing and image display on a preview monitor
6.2 | A diagnostic radiology LCD 19 inch imaging workstation based on latest flat panel detector technology for high image dynamics with excellent signal/noise ratio is required
6.3 | Application and examination database sotfware included.Details are stated
6.4 | Image processing software are included
6.5 | 64 bit 3.2GHz i5 Core Dual Microprocessor or similar
6.6 | RAM: 16-32 GB
6.7 | 64 bit Operating system included (Windows or Similar)
6.8 | Image storage capacity up to 10 000 full format images on internal hard drive. Details of image storage capacity are stated in GB, as well as in maximum number of images in the supporting documents provided.
6.9 | No Autodeletion of images stored on internal hard drive. System to alert when near maximum storage capacity is reached.
6.10. | System must be DICOM compatible and capable of HIS/HMS/ RIS and PACS integration
6.11 | DVD /CD Writer inclusive of software for functionality of the DVD/CD Writer..
6.12 | Virus free software
6.13 | Ergonomic keyboard with mouse and function buttons
6.14 | Memory stick support/USB PORT for QC
6.15 | Touch screen monitor
7 | USER INTERFACE
7.1 | Patient data entering
7.2 | Menu
7.3 | Integrated generator operation
7.4 | Menu control
7.5 | Organ program selection
7.6 | Window position/width
7.7 | Horizontal/vertical image mirroring
7.8 | Image manipulation: rotation, cropping, contrast,brightness,sharpness, inversion
7.9 | Left/right, ap/pa working
7.1 | Configurable text annotation
7.11 | Filter selection for edge enhancement
7.12 | Filter selection for noise suppression
7.13 | Image pan and zoom capability
7.14 | measurement tools
8 | CONTROL CONSOLE
8.1 | Patient data entering
8.2 | Manual
8.3 | Bar code reader
8.4 | Monitor size of 19 inch(diagonal) Touch screen
8.5 | Colour
8.6 | Keyboard with mouse and function buttons
8.7 | Spill proof keyboard design
8.8 | Mouse
8.9 | Menu
8.10. | Integrated generator operation
8.11 | Menu control
8.12 | Organ program selection
8.13 | Window position/width
8.14 | Horizontal/vertical image mirroring
8.15 | Image rotation
8.16 | Left/right, ap/pa working
8.17 | Configurable text annotation
8.18 | Filter selection for edge enhancement
8.19 | Filter selection for noise suppression
8.20. | Image zoom capability
9 | APPLICATION AND EVALUATION PROGRAMS
9.1 | Organ-related application program
9.2 | Post processing functionality
9.3 | Windowing (centre/width)
9.4 | Paging forward and back
9.5 | Multi-image display on the monitor (mosaic)
9.6 | Retro-collimation facilities (shutter)
9.7 | Image documentation in the background
9.8 | Multi-tasking technique
9.9 | Edge enhancement
9.1 | Noise suppression
9.11 | Zoom
9.12 | Individual energy substraction programs provided
10 | SYSTEM INTERPHASE
10.1 | Detector interphase
10.2 | Interface DICOM: "Basic Print"
10.3 | Interface DICOM: "Modality Worklist"
10.4 | Interface DICOM: "Storage"
10.5 | Interface DICOM: "Send as SCU"
10.6 | Interface DICOM "Image export/transfer to PACS/HIS/RIS
10.7 | Is there a "Confirmation Statement" for all DICOM interfaces
10.8 | Firewall protection
10.9 | The succesful bidder must ensure that the configuration between the processed images on the Digital Imaging Workstation and that transferred to the receivers of these images, such as PACS , Laser Printer , CD/DVD are properly calibrated/ reconfigured during installation, service and maintenance,repairs, and future software updates/upgrades .
11 | OPTIONAL ITEMS
11.1 | 1 X TB HARD DRIVE(external) ; DICOM CONFIGURED TO STORE, TRANSFER AND PRINT X-RAY IMAGES.
11.2 | Motorized collimation control as an option
11.3 | Tilting buckey as an option
11.4 | Synchronized tube and detector movements. Auto detector tracking and autotube positioning
11.5 | The system must have the ability to perform digital long length stitching images(scoliosis/long bones) -OPTIONAL
11.6 | Auto image stitching program included in the software package.
11.7 | TOR Rad or similar Image Resolution QC Test Tool
11.8 | A trolley with lockable castors, radiolucent top and an integrated detector holder that must move across the length of the trolley.
12 | RADIATION PROTECTION
12.1 | The system must have the ability to display Exposure Indicator values (EI/ S-values), immediately after exposure.
12.2 | X-ray Warning Notice for display at the X-ray Room entrance (not branded).
12.3 | An advanced Retake / Reject / Repeat analysis software package that includes analysis as well as calculations in percentages, must be provided for applicable systems.
12.4 | The applicable systems must have an integrated Dose Area Product (DAP) meter that includes skin dose equivalent measurement functions.
12.5 | A radiation report must be automatically generated and transferable via DICOM.
12.6 | Should be FDA or CE approved product
12.7 | Electrical safety conforms to standards for electrical safety IEC-60601
12.8 | The unit must comply to IEC-61267
12.9 | The equipment/system must be approved and licensed by the Directorate Radiation Control of DoH: South Africa
12.10. | A copy of a valid license issued in terms of the Hazardous Substance Act, Act No 15 of 1973 must be submitted with the tender. Failure to submit such a valid license may result in a tender not being considered.
12.11 | Acceptance testing must be performed as part of the installation. (Failure to do will lead to failure to pay
12.12 | Lead Glass window to be installed as per room requirement
12.13 | Safety aspects of Radiation dosage leakage should be spelt out
12.14 | DB boards and electrical filters to be installed
12.15 | QA TOOLS
12.16 | Display Exposure index values
12.17 | Radiation protection Lead equivalence of the radiation protection screen must be 0.5m Pb and must be convex, covering the radiographer from various angles against scatter radiation.
12.18 | Patient Conventional face shield to be included.
12.19 | Patient Tomosynthesis face shield to be included, removable and retractable.
12.20. | Emergency stop controls must be supplied on both sides of the gantry and on the Acquisition Workstation
12.21 | Interlocks to stop system movement when compression is applied must be provided.
12.22 | Interlock for premature release of the X-ray switch must be available
12.23 | Interlock preventing an exposure whenever the filter or light field is not positioned correctly must be possible.
12.24 | Phantom for calibration should be provided.
13 | POWER SUPPLY
13.1 | The bidder must install adequate electrical power supply for the optimal functionality of the equipment.
13.2 | The mains cable of the unit being quoted for must be the hospital grade type and it must be of an adequate length (minimum length of three (3) meters).
13.3 | The power input must be 220-240V, 50Hz AC. Bidders must ensure that the product quoted for is fitted appropriately with a 15 Amp SABS approved mains plug that is held together by two screws.
13.4 | A 3 phase in / out power supply must be offered. State details.
13.5 | The equipment tendered for must not overload and trip the hospital power in the area where it is installed. Resettable overcurrent breaker/s must be fitted for protection. State details.
13.6 | All available online Uninterrupted Power Supply (UPS system) and line voltage regulators must be included for the system with a support time of 50 to 60 minutes. Attach brochure.
13.7 | The lifetime of UPS must be at least 15 years.
13.8 | An Uninterrupted Power Supply (UPS system) conforming to IS-302 must be offered for the computer system.
13.9 | The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade type and it must be a minimum length of three (3) meters.
13.10. | The mains cable of the unit tendered for must be SABS colour coded. State details.
13.11 | The equipment quoted for must be protected against electromagnetic interference. State details of measures that will be taken to protect the unit against electromagnetic interference.
13.12 | Maintenance free batteries must be offered.
14 | AIR CONDITIONER
14.1 | Bidder must include any air conditioning that is considered necessary for optimal functioning of the unit.
14.2 | Any such air conditioning must be covered by the guarantee bided for on the unit and servicing of the air conditioning units for the guarantee period and must be included in the bid price.
14.3 | The air conditioner must also be included in the maintenance plan.
14.4 | Successful bidder to replace existing air conditioner with an air conditioner of 24,000 BTU. (Where applicable).
14.5 | The air conditioner must be but no less than 24,000 BTU. State details on the air conditioner to be supplied according to the room size and machine requirements.
14.6 | The Successful bidder must install an outlet-plumbing pipe for water overflow from air conditioner to be re-directed to a drainage system.
14.7 | Air-conditioning must be included in Maintenance plan

## CR Reader Single Slot
Single Feeder CR Reader
ITEM
TECHNICAL SPECIFICATION
1 | The system offered must have a compact footprint. State the dimensions of the system offered.
2 | The system must be able to accept patient and imaging data.
3 | The unit must be a 19 inch touch screen system.
4 | The unit weight must not exceed 30kg
5 | The unit must come with a auto-stitching software
6 | The system must have a configurable auto delete function that can alert when near maximum storage capacity is reached
7 | The unit must be equipped with a barcode reader or radio chipped, which will enable the radiographer to link patient data to imaging data on a specific cassette.
8 | If an exam digitization has been pre-ordered through Radiology Information System (RIS), the plate reader must provide the ability to acquire appropriate patient and exam identification from the RIS via a Dicom Modality work-list.
9 | The plate reader must also provide an option for the operator to enter exam and patient identification information manually, for unscheduled exams such as emergency cases.
10 | The unit must accept the following radiographic cassette/IP plates sizes: 35cm X 43cm; 24 cm X 30cm and 18cm X 24 cm. State the throughput of the unit for each different size of cassette/IP plate. State what other cassette/IP plate sizes are available.
11 | State the cost of the various size cassettes/IP plates on the Optional Schedule of Accessories.
12 | The minimum throughput must be 90 (35cm x 43cm) cassettes/imaging plates or higher per hour.
13 | The system must be capable of high resolution scanning from the imaging plates. A resolution of at least 10 pixels per mm is required for all image plate sizes. State resolution of IP sizes offered.
14 | The system must have the ability to do image control and system control, as well as patient data manipulation.
15 | Imaging plates must be automatically read after insertion into the plate reader.
16 | The CR Reader unit must be connected to a dedicated CR console, providing the operator with all necessary controls in respect of window/level and image transfer and formatting to a future Picture Archiving & Communications System (PACS) or Dry Laser Imager via the DICOM 3 standard for reporting, distribution, printing and archiving.
17 | This console must be mounted above or alongside the CR Reader.
18 | State all the functionalities of the console offered.
19 | Reject analysis and Quality Control software must be included.
20 | The console described in Clause 12 above must have the ability to auto-route image and patient information to either the radiologist’s or the radiographer’s workstation as well as automatic transfer of images to a Picture Archiving & Communications System (PACS).
21 | In the event of the Hospital network failure, it must be possible to directly print images from the CR system to a Dry Laser Imager/ Printer/CD writer as described later. The Network cost between the CR System and Dry Laser Imager/Printer/CD writer must be included as part of the main bid price.
22 | The network cable must be Cat 5/6 UTP Cable or higher.
23 | A minimum of 4 TB image storage capacity on-line on the CR Reader System.
24 | It must take less than 1 minute to retrieve an image from system storage for image review.
25 | A DVD/CD writer must be included for DICOM image distribution. The cost of this must be included in the bid.
26 | Inline/Online UPS to be included for computer system (that can run for at least 60 minutes)
27 | Where applicable:
27.1 | The CR Reader must be linked to the existing Radiology Information System (RIS)/Picture Archiving and Communication System (PACS).
27.2 | All Dicom connectivity costs to Radiology Information System (RIS)/Picture Archiving and Communication System (PACS) MUST be included in the offer.
28 | RANGE OF CR CASSETTES (INCLUDING IMAGING PLATES)
28.1 | The following range and quantity of cassettes with imaging plates must be included in the offer. State price per each cassette/IP plate as well as a total price for the following requirement. (Moderate variation in cassette sizes will be considered)
28.1.1 | Sizes and Quantity
(i) 24cm x 30 cm = 2
(ii) 18cm x 24 cm = 2
(iii) 35cm x 43 cm = 2
28.2 | The cassettes must be bar-coded or radio chipped to facilitate the operation of the system.
28.3 | The expected plate life must be 30 000 cycles or higher. State the expected plate life in terms of cycles.
28.4 | State what warranty is given on the number of cycles per IP plate and how this is digitally recorded. Supply details.
28.5 | State if the IP plates alone can be replaced, or if the complete cassette must be replaced.
28.6 | State the greyscale resolution of the plates.
28.7 | The standard resolution must be 4 pixels /mm or better for general purpose radiography and 9 pixels /mm or better for high resolution radiography.
28.8 | State the standard resolution per millimeter for general radiography and for high resolution radiography.
29 | POWER SUPPLY
29.1 | The bidder must install adequate electrical power supply for the optimal functionality of the equipment.
29.2 | The mains cable of the unit being quoted for must be the hospital grade type and it must be of an adequate length (minimum length of three (3) meters).
29.3 | The power input must be 220-240V, 50Hz AC. Bidders must ensure that the product quoted for is fitted appropriately with a 15 Amp SABS approved mains plug that is held together by two screws.
29.4 | A 3 phase in / out power supply must be offered. State details.
29.5 | The equipment tendered for must not overload and trip the hospital power in the area where it is installed. Resettable overcurrent breaker/s must be fitted for protection. State details.
29.6 | All available online Uninterrupted Power Supply (UPS system) and line voltage regulators must be included for the system with a support time of 50 to 60 minutes. Attach brochure.
29.7 | The lifetime of UPS must be at least 15 years.
29.8 | An Uninterrupted Power Supply (UPS system) conforming to IS-302 must be offered for the computer system.
29.9 | The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade type and it must be a minimum length of three (3) meters.
29.10. | The mains cable of the unit tendered for must be SABS colour coded. State details.
29.11 | The equipment quoted for must be protected against electromagnetic interference. State details of measures that will be taken to protect the unit against electromagnetic interference.
29.12 | Maintenance free batteries must be offered.
30 | AIR CONDITIONER
30.1 | Bidder must include any air conditioning that is considered necessary for optimal functioning of the unit.
30.2 | Any such air conditioning must be covered by the guarantee bided for on the unit and servicing of the air conditioning units for the guarantee period and must be included in the bid price.
30.3 | The air conditioner must also be included in the maintenance plan.
30.4 | Successful bidder to replace existing air conditioner with an air conditioner of 24,000 BTU. (Where applicable).
30.5 | The air conditioner must be but no less than 24,000 BTU. State details on the air conditioner to be supplied according to the room size and machine requirements.
30.6 | The Successful bidder must install an outlet-plumbing pipe for water overflow from air conditioner to be re-directed to a drainage system.
30.7 | Air-conditioning must be included in Maintenance plan

## Dry Laser Imager
DRY LASER IMAGER
Item | TECHNICAL SPECIFICATIONS
1 | A good quality dry laser imager must be included in this offer. State the dimensions of the system offered.
2 | The unit must use either laser or thermal technology. No units using chemical technology will be considered.
3 | The printer must automatically perform a calibration test at start-up or on the request of the operator.
4 | The printer must include a self diagnostic capability which shall indicate at least the following error and system status conditions; film low, film empty, memory full, film feed error, printing, and alarm.
5 | The unit must allow daylight loading of film. State the loading capacity of the printer supply magazine.
6 | Multi-format printing on all film sizes must be possible. State the printing formats available on the offered printer.
7 | The printer must print the Department name, patient identification, date (dd/mm/yy) and time onto the film.
8 | Requests for printing must not compromise any additional workstation operation or performance.
Must include an online densitometer for constant image quality assessment including self calibration and optical density consistency as per SAHPRA QC test requirements
9 | The printer must be DICOM 3.0 compatible and enabled.
10 | It must be possible to print the following film sizes: 35cm X 43cm, 24cm X 30cm or 18cm X 24cm. The printer should be capable of supporting 2 film sizes on - line permanently.
11 | The throughput must be at least 120 (35cm x 43cm) films per hour. State the throughput of the dry laser imager offered. (ie number of 35cm x 43cm films per hour).
12 | It must be possible to send a print command from both the CR reader system and all optional workstations.
13 | POWER SUPPLY
13.1 | The bidder must install adequate electrical power supply for the optimal functionality of the equipment.
13.2 | The mains cable of the unit being quoted for must be the hospital grade type and it must be of an adequate length (minimum length of three (3) meters).
13.3 | The power input must be 220-240V, 50Hz AC. Bidders must ensure that the product quoted for is fitted appropriately with a 15 Amp SABS approved mains plug that is held together by two screws.
13.4 | A 3 phase in / out power supply must be offered. State details.
13.5 | The equipment tendered for must not overload and trip the hospital power in the area where it is installed. Resettable overcurrent breaker/s must be fitted for protection. State details.
13.6 | All available online Uninterrupted Power Supply (UPS system) and line voltage regulators must be included for the system with a support time of 50 to 60 minutes. Attach brochure.
13.7 | The lifetime of UPS must be at least 15 years.
13.8 | An Uninterrupted Power Supply (UPS system) conforming to IS-302 must be offered for the computer system.
13.9 | The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade type and it must be a minimum length of three (3) meters.
13.10. | The mains cable of the unit tendered for must be SABS colour coded. State details.
13.11 | The equipment quoted for must be protected against electromagnetic interference. State details of measures that will be taken to protect the unit against electromagnetic interference.
13.12 | Maintenance free batteries must be offered.
14 | AIR CONDITIONER
14.1 | Bidder must include any air conditioning that is considered necessary for optimal functioning of the unit.
14.2 | Any such air conditioning must be covered by the guarantee bided for on the unit and servicing of the air conditioning units for the guarantee period and must be included in the bid price.
14.3 | The air conditioner must also be included in the maintenance plan.
14.4 | Successful bidder to replace existing air conditioner with an air conditioner of 24,000 BTU. (Where applicable).
14.5 | The air conditioner must be but no less than 24,000 BTU. State details on the air conditioner to be supplied according to the room size and machine requirements.
14.6 | The Successful bidder must install an outlet-plumbing pipe for water overflow from air conditioner to be re-directed to a drainage system.
14.7 | Air-conditioning must be included in Maintenance plan

## CR Reader Singl with Mamo
Single Feeder CR Reader with Mammography
ITEM | TECHNICAL SPECIFICATION
1 | The system offered must have a compact footprint. State the dimensions of the system offered.
2 | The system must be able to accept patient and imaging data.
3 | The unit must be a 19 inch touch screen system.
4 | The unit weight must not exceed 30kg
5 | The unit must come with a auto-stitching software
6 | The system must have a configurable auto delete function that can alert when near maximum storage capacity is reached
7 | The unit must be equipped with a barcode reader or radio chipped, which will enable the radiographer to link patient data to imaging data on a specific cassette.
8 | If an exam digitization has been pre-ordered through Radiology Information System (RIS), the plate reader must provide the ability to acquire appropriate patient and exam identification from the RIS via a Dicom Modality work-list.
9 | The plate reader must also provide an option for the operator to enter exam and patient identification information manually, for unscheduled exams such as emergency cases.
10 | The unit must accept the following radiographic cassette/IP plates sizes: 35cm X 43cm; 24 cm X 30cm and 18cm X 24 cm. State the throughput of the unit for each different size of cassette/IP plate. State what other cassette/IP plate sizes are available.
11 | State the cost of the various size cassettes/IP plates on the Optional Schedule of Accessories.
12 | The minimum throughput must be 90 (35cm x 43cm) cassettes/imaging plates or higher per hour.
13 | The system must be capable of high resolution scanning from the imaging plates. Resolutions of at least 10 pixels per mm is required for all image plate sizes. State resolution of IP sizes offered.
14 | The system must have the ability to do image control and system control, as well as patient data manipulation.
15 | Imaging plates must be automatically read after insertion into the plate reader.
16 | The CR Reader unit must be connected to a dedicated CR console, providing the operator with all necessary controls in respect of window/level and image transfer and formatting to a future Picture Archiving & Communications System (PACS) or Dry Laser Imager via the DICOM 3 standard for reporting, distribution, printing and archiving. This console must be mounted above or alongside the CR Reader.
17 | State all the functionalities of the console offered.
18 | Reject analysis and Quality Control software must be included.
19 | The console described in Clause 1.12 above must have the ability to auto-route image and patient information to either the radiologist’s or the radiographer’s workstation as well as automatic transfer of images to a Picture Archiving & Communications System (PACS).
20 | In the event of the Hospital network failure, it must be possible to directly print images from the CR system to a Dry Laser Imager/ Printer as described later. The Network cost between the CR System and Dry Laser Imager/Printer must be included as part of the main bid price. The network cable must be Cat 5/6 UTP Cable or higher.
21 | The bidder must state the image storage capacity on-line on the CR Reader System.
22 | State the time-frame required to retrieve an image from system storage for image review.
23 | A DVD writer must be included for image distribution. The cost of this must be included in the bid.
24 | RANGE OF CR CASSETTES (INCLUDING IMAGING PLATES)
24.1 | The following range and quantity of cassettes with imaging plates must be included in the offer. State price per each cassette/IP plate as well as a total price for the following requirement. (Moderate variation in cassette sizes will be considered)
24.1.1 | Sizes and Quantity
(i) 24cm x 30 cm = 2
(ii) 18cm x 24 cm = 2
(iii) 35cm x 43 cm = 2
24.1.2 | Mammography Cassettes
Size Quantity
(i) 24 x 30cm 2
(ii) 18 x 24cm 2
24.2 | The cassettes must be bar-coded or radio chipped to facilitate the operation of the system.
24.3 | The expected plate life must be 30 000 cycles or higher. State the expected plate life in terms of cycles.
24.4 | State what warranty is given on the number of cycles per IP plate and how this is digitally recorded. Supply details.
24.5 | State if the IP plates alone can be replaced, or if the complete cassette must be replaced.
24.6 | State the greyscale resolution of the plates.
24.7 | The standard resolution must be 4 pixels /mm or better for general purpose radiography and 9 pixels /mm or better for high resolution radiography and 10 pixels /mm or better for mammogarphy.
24.8 | The bidder to state the standard resolution Lp/mm for general radiography, for high resolution radiography and ultrahigh resolution for mammography.
25 | POWER SUPPLY
25.1 | The bidder must install adequate electrical power supply for the optimal functionality of the equipment.
25.2 | The mains cable of the unit being quoted for must be the hospital grade type and it must be of an adequate length (minimum length of three (3) meters).
25.3 | The power input must be 220-240V, 50Hz AC. Bidders must ensure that the product quoted for is fitted appropriately with a 15 Amp SABS approved mains plug that is held together by two screws.
25.4 | A 3 phase in / out power supply must be offered. State details.
25.5 | The equipment tendered for must not overload and trip the hospital power in the area where it is installed. Resettable overcurrent breaker/s must be fitted for protection. State details.
25.6 | All available online Uninterrupted Power Supply (UPS system) and line voltage regulators must be included for the system with a support time of 50 to 60 minutes. Attach brochure.
25.7 | The lifetime of UPS must be at least 15 years.
25.8 | An Uninterrupted Power Supply (UPS system) conforming to IS-302 must be offered for the computer system.
25.9 | The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade type and it must be a minimum length of three (3) meters.
25.10. | The mains cable of the unit tendered for must be SABS colour coded. State details.
25.11 | The equipment quoted for must be protected against electromagnetic interference. State details of measures that will be taken to protect the unit against electromagnetic interference.
25.12 | Maintenance free batteries must be offered.
26 | AIR CONDITIONER
26.1 | Bidder must include any air conditioning that is considered necessary for optimal functioning of the unit.
26.2 | Any such air conditioning must be covered by the guarantee bided for on the unit and servicing of the air conditioning units for the guarantee period and must be included in the bid price.
26.3 | The air conditioner must also be included in the maintenance plan.
26.4 | Successful bidder to replace existing air conditioner with an air conditioner of 24,000 BTU. (Where applicable).
26.5 | The air conditioner must be but no less than 24,000 BTU. State details on the air conditioner to be supplied according to the room size and machine requirements.
26.6 | The Successful bidder must install an outlet-plumbing pipe for water overflow from air conditioner to be re-directed to a drainage system.
26.7 | Air-conditioning must be included in Maintenance plan
26 | ACCESSORIES
Supplier to please specify on the Schedule of Optional Accessories

## CR Reader MultiSlot
Multislot Feeder CR Reader
ITEM | TECHNICAL SPECIFICATION
1 | The system must be able to accept patient and imaging data.
2 | The unit must be a 19 inch touch screen system.
3 | The unit must come with a auto-stitching software
4 | The system must have a configurable auto delete function that can alert when near maximum storage capacity is reached
5 | The unit must be equipped with a barcode reader or radio chipped, which will enable the radiographer to link patient data to imaging data on a specific cassette.
6 | If an exam has been pre-ordered through the RIS (Radiology Information System), the plate reader must provide the ability to acquire appropriate patient and exam identification from the RIS via a Dicom Modality work-list.
7 | The plate reader must also provide an option for the operator to enter exam and patient identification information manually, for unscheduled exams such as emergency cases.
8 | The unit must accept the following radiographic cassette/IP plates sizes : 35 X 43; 24 X 30 and 18 X 24; State the throughput of the unit for each different size of cassette/IP plate. State what other cassette/IP plate sizes are available.
9 | State the cost of the various size cassettes/IP plates on the optional schedule of accessories.
10 | The minimum throughput must be 120 (35cmX 43cm) cassettes /imaging plates or higher per hour.
11 | The system must have the ability to do image control and system control, as well as patient data manipulation.
12 | Imaging plates must be automatically read after insertion into the plate reader.
13 | The CR Reader unit must be connected to a dedicated CR console, providing the operator with all necessary controls in respect of window/level and image transfer and formatting to the existing PACS system and/or Dry Laser Imager via the DICOM 3 standard for reporting, distribution, printing and archiving. This console must be mounted above or alongside the CR Reader. Ability to store and print.
14 | State all radiography tooling functionality of the console offered. The system should have the capability for the provision of pre – defined text and annotation.
15 | The console described in Clause T10 above must have the ability to auto-route image and patient information to either the radiologist’s or the radiographer’s workstation as well as automatic transfer of images to a RIS/ PACS system.
16 | In the event of the Hospital network failure, it must be possible to directly print images from the CR system to a Dry Laser Imager/ Printer. The Network cost between the CR System and Dry Laser Imager/Printer must be included as part of the main bid price. The network cable must be Cat 5/6 UTP Cable
17 | The bidder must state the image storage capacity on-line on the CR Reader System.
18 | State the time-frame required to retrieve an image from CR Reader system storage for image review.
19 | RANGE OF CR CASSETTES. (INCLUDING IMAGING PLATES)
19.1 | The following range and quantity of cassettes with imaging plates must be included in the offer. State price per each cassette/IP plate as well as a total price for the following requirement. (Moderate variation in cassette sizes will be considered)
19.1.1 | Size and Quantity
(i)24 X 30 cm = 2
(ii)18 x 24 cm = 2
(iii)35 X 43 cm = 2
19.2 | The cassettes must be bar-coded or radio chipped to facilitate the operation of the system.
19.3 | The expected plate life must be 30 000 cycles. State the expected plate life in terms of cycles.
19.4 | State what warranty is given on the number of cycles per IP plate and how this is digitally recorded. Supply details.
19.5 | State if IP plates alone can be replaced, or if the complete cassette must be replaced.
19.6 | The standard resolution must be 5 pixels /mm or better for general purpose radiography, 10 pixels /mm or better for high resolution radiography.
19.7 | State the standard resolution pixels /mm for general radiography, for high resolution radiography and ultrahigh resolution for mammography.
20 | POWER SUPPLY
20.1 | The bidder must install adequate electrical power supply for the optimal functionality of the equipment.
20.2 | The mains cable of the unit being quoted for must be the hospital grade type and it must be of an adequate length (minimum length of three (3) meters).
20.3 | The power input must be 220-240V, 50Hz AC. Bidders must ensure that the product quoted for is fitted appropriately with a 15 Amp SABS approved mains plug that is held together by two screws.
20.4 | A 3 phase in / out power supply must be offered. State details.
20.5 | The equipment tendered for must not overload and trip the hospital power in the area where it is installed. Resettable overcurrent breaker/s must be fitted for protection. State details.
20.6 | All available online Uninterrupted Power Supply (UPS system) and line voltage regulators must be included for the system with a support time of 50 to 60 minutes. Attach brochure.
20.7 | The lifetime of UPS must be at least 15 years.
20.8 | An Uninterrupted Power Supply (UPS system) conforming to IS-302 must be offered for the computer system.
20.9 | The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade type and it must be a minimum length of three (3) meters.
20.10. | The mains cable of the unit tendered for must be SABS colour coded. State details.
20.11 | The equipment quoted for must be protected against electromagnetic interference. State details of measures that will be taken to protect the unit against electromagnetic interference.
20.12 | Maintenance free batteries must be offered.
21 | AIR CONDITIONER
21.1 | Bidder must include any air conditioning that is considered necessary for optimal functioning of the unit.
21.2 | Any such air conditioning must be covered by the guarantee bided for on the unit and servicing of the air conditioning units for the guarantee period and must be included in the bid price.
21.3 | The air conditioner must also be included in the maintenance plan.
21.4 | Successful bidder to replace existing air conditioner with an air conditioner of 24,000 BTU. (Where applicable).
21.5 | The air conditioner must be but no less than 24,000 BTU. State details on the air conditioner to be supplied according to the room size and machine requirements.
21.6 | The Successful bidder must install an outlet-plumbing pipe for water overflow from air conditioner to be re-directed to a drainage system.
21.7 | Air-conditioning must be included in Maintenance plan

## CR Reasder Multi with Mamo
Multislot Feeder CR Reader with Mammography
ITEM | TECHNICAL SPECIFICATION
1 | The system must be able to accept patient and imaging data.
2 | The unit must be a 19 inch touch screen system.
3 | The unit must come with a auto-stitching software
4 | The system must have a configurable auto delete function that can alert when near maximum storage capacity is reached
5 | The unit must be equipped with a barcode reader or radio chipped, which will enable the radiographer to link patient data to imaging data on a specific cassette.
6 | If an exam has been pre-ordered through the RIS (Radiology Information System), the plate reader must provide the ability to acquire appropriate patient and exam identification from the RIS via a DICOM Modality work-list.
7 | The plate reader must also provide an option for the operator to enter exam and patient identification information manually, for unscheduled exams such as emergency cases
8 | The unit must accept the following radiographic cassette/IP plates sizes: 35 X 43; 24 X 30 and 18 X 24; State the throughput of the unit for each different size of cassette/IP plate. State what other cassette/IP plate sizes are available.
9 | State the cost of the various size cassettes/IP plates on the optional schedule of accessories.
10 | The minimum throughput must be 120 (35cmX 43cm) cassettes /imaging plates or higher per hour
11 | The system must have the ability to do image control and system control, as well as patient data manipulation
12 | Imaging plates must be automatically read after insertion into the plate reader.
13 | The CR Reader unit must be connected to a dedicated CR console, providing the operator with all necessary controls in respect of window/level and image transfer and formatting to the existing PACS system and/or Dry Laser Imager via the DICOM 3 standard for reporting, distribution, printing and archiving. This console must be mounted above or alongside the CR Reader. Ability to store and print.
14 | State all radiography tooling functionality of the console offered. The system should have the capability for the provision of pre – defined text and annotation.
15 | The console described in Clause T10 above must have the ability to auto-route image and patient information to either the radiologist’s or the radiographer’s workstation as well as automatic transfer of images to a RIS/ PACS system.
16 | In the event of the Hospital network failure, it must be possible to directly print images from the CR system to a Dry Laser Imager/ Printer. The Network cost between the CR System and Dry Laser Imager/Printer must be included as part of the main bid price. The network cable must be Cat 5/6 UTP Cable.
17 | The bidder must state the image storage capacity on-line on the CR Reader System.
18 | State the time-frame required to retrieve an image from CR Reader system storage for image review.
19 | RANGE OF CR CASSETTES. (INCLUDING IMAGING PLATES)
19.1 | The following range and quantity of cassettes with imaging plates must be included in the offer. State price per each cassette/IP plate as well as a total price for the following requirement. (Moderate variation in cassette sizes will be considered)
19.1.1 | Size Quantity
(i) 24 X 30 cm 2
(ii) 18 x 24 cm 2
(iii) 35 X 43 cm 2
19.1.2 | Mammography Cassettes
Size Quantity
(i) 24 x 30cm 4
(ii) 18 x 24cm 4
19.2 | The cassettes must be bar-coded or radio chipped to facilitate the operation of the system.
19.3 | The expected plate life must be 30 000 cycles. State the expected plate life in terms of cycles.
19.4 | State what warranty is given on the number of cycles per IP plate and how this is digitally recorded. Supply details.
19.5 | State if IP plates alone can be replaced, or if the complete cassette must be replaced.
19.6 | The standard resolution must be 5 pixels /mm or better for general purpose radiography, 10 pixels /mm or better for high resolution radiography and 20 pixels /mm or better for mammogarphy.
19.7 | State the standard resolution pixels /mm for general radiography, for high resolution radiography and ultrahigh resolution for mammography.
20 | POWER SUPPLY
20.1 | The bidder must install adequate electrical power supply for the optimal functionality of the equipment.
20.2 | The mains cable of the unit being quoted for must be the hospital grade type and it must be of an adequate length (minimum length of three (3) meters).
20.3 | The power input must be 220-240V, 50Hz AC. Bidders must ensure that the product quoted for is fitted appropriately with a 15 Amp SABS approved mains plug that is held together by two screws.
20.4 | A 3 phase in / out power supply must be offered. State details.
20.5 | The equipment tendered for must not overload and trip the hospital power in the area where it is installed. Resettable overcurrent breaker/s must be fitted for protection. State details.
20.6 | All available online Uninterrupted Power Supply (UPS system) and line voltage regulators must be included for the system with a support time of 50 to 60 minutes. Attach brochure.
20.7 | The lifetime of UPS must be at least 15 years.
20.8 | An Uninterrupted Power Supply (UPS system) conforming to IS-302 must be offered for the computer system.
20.9 | The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade type and it must be a minimum length of three (3) meters.
20.10. | The mains cable of the unit tendered for must be SABS colour coded. State details.
20.11 | The equipment quoted for must be protected against electromagnetic interference. State details of measures that will be taken to protect the unit against electromagnetic interference.
20.12 | Maintenance free batteries must be offered.
21 | AIR CONDITIONER
21.1 | Bidder must include any air conditioning that is considered necessary for optimal functioning of the unit.
21.2 | Any such air conditioning must be covered by the guarantee bided for on the unit and servicing of the air conditioning units for the guarantee period and must be included in the bid price.
21.3 | The air conditioner must also be included in the maintenance plan.
21.4 | Successful bidder to replace existing air conditioner with an air conditioner of 24,000 BTU. (Where applicable).
21.5 | The air conditioner must be but no less than 24,000 BTU. State details on the air conditioner to be supplied according to the room size and machine requirements.
21.6 | The Successful bidder must install an outlet-plumbing pipe for water overflow from air conditioner to be re-directed to a drainage system.
21.7 | Air-conditioning must be included in Maintenance plan

## X-Ray mobile unit
MOBILE X – RAY UNIT
ITEM | TECHNICAL SPECIFICATIONS:
The unit must be able to pass through a door of 210cm height. Spare parts must be guaranteed available for the specified life of the equipment, with a minimum of 10 years.
1 | GENERATOR:
1.1 | Compact falling load generator with an output of at least 30 kW
1.2 | The maximum mAs at various kVs must be stated
These factors will be taken into consideration in the bid adjudication. The unit must have a mA output that will make radiography with reduced exposure times possible.
State:
Maximum mAs at 125 kV
Maximum mAs at 110 kV
Maximum mAs at 100 kV
Maximum mAs at 90 kV
Maximum mAs at 80 kV
Maximum mAs at 70 kV
Maximum mAs at 60 kV
Maximum mAs at 50 kV
Maximum mAs at 40 kV
1.3 | The kV control must provide for steps of 1kV
1.4 | The bidder to state the kV range
1.5 | The cable connecting the hand switch must be at least two metres. State length
1.6 | The unit offered must incorporate automatic mains compensation
2 | X-RAY TUBE AND CABLES:
2.1 | The size of the focus must be in the range of (0.8 – 1.2)
2.2 | A pair of high voltage cables and installation cables to suit installation must be provided
2.3 | Anode heat storage capacity must not be less than 40kHU
3 | COLLIMATOR:
3.1 | A light beam diaphragm must be fitted
3.2 | The leaves must be independently adjustable in both lateral and longitudinal directions
3.3 | A rotating flange between the collimator and the tube must be supplied in order to permit diagonal collimation
3.4 | The unit is required for radiography of patients in hospital beds and a FFD of less than 100cms is frequently used. A field coverage of 43 x 43cms at a focal film distance of 80cms must be achieved.
State field coverage at 80cms. _________________
4 | TUBE STAND/MOBILE BASE:
4.1 | The unit must have a base frame with the tube parked or clipped close to the base frame for transport
4.2 | The tube stand must provide for vertical movement of the focus of the X-Ray tube up to a height of at least 1,8 metres above the floor
4.3 | The bidder to state maximum height of unit when prepared for transport
4.4 | State minimum height of tube above floor level when prepared for transport
4.5 | Angulation of the tube head must be controlled by an adequate locking device
4.6 | The unit must be suitable for radiography on a theatre table and therefore the tube when positioned to give a FFD of 100 cms must extend past the line of the base plate by an adequate distance
4.7 | If the floor to focus distance were 1800 mm (maximum) state what the maximum distance would be between the front end of the base plate of the unit and the central ray
4.8 | Bidders must supply details of the movements of the tube and tube arm/column and must state whether the tube arm can swing to the right and left (wig-wag)
4.9 | The base must be fitted with an adequate floor brake
4.1 | State the mass of the unit
5 | RADIATION PROTECTION
5.1 | The system must have the ability to display Exposure Indicator values (EI/ S-values), immediately after exposure.
5.2 | X-ray Warning Notice for display at the X-ray Room entrance (not branded).
5.3 | An advanced Retake / Reject / Repeat analysis software package that includes analysis as well as calculations in percentages, must be provided for applicable systems.
5.4 | The applicable systems must have an integrated Dose Area Product (DAP) meter that includes skin dose equivalent measurement functions.
5.5 | A radiation report must be automatically generated and transferable via DICOM.
5.6 | Should be FDA or CE approved product
5.7 | Electrical safety conforms to standards for electrical safety IEC-60601
5.8 | The unit must comply to IEC-61267
5.9 | The equipment/system must be approved and licensed by the Directorate Radiation Control of DoH: South Africa
5.10. | A copy of a valid license issued in terms of the Hazardous Substance Act, Act No 15 of 1973 must be submitted with the tender. Failure to submit such a valid license may result in a tender not being considered.
5.11 | Acceptance testing must be performed as part of the installation. (Failure to do will lead to failure to pay
5.12 | Lead Glass window to be installed as per room requirement
5.13 | Safety aspects of Radiation dosage leakage should be spelt out
5.14 | DB boards and electrical filters to be installed
5.15 | QA TOOLS
5.16 | Display Exposure index values
5.17 | Radiation protection Lead equivalence of the radiation protection screen must be 0.5m Pb and must be convex, covering the radiographer from various angles against scatter radiation.
5.18 | Patient Conventional face shield to be included.
5.19 | Patient Tomosynthesis face shield to be included, removable and retractable.
5.20. | Emergency stop controls must be supplied on both sides of the gantry and on the Acquisition Workstation
5.21 | Interlocks to stop system movement when compression is applied must be provided.
5.22 | Interlock for premature release of the X-ray switch must be available
5.23 | Interlock preventing an exposure whenever the filter or light field is not positioned correctly must be possible.
5.24 | Phantom for calibration should be provided.
6 | POWER SUPPLY
6.1 | The bidder must install adequate electrical power supply for the optimal functionality of the equipment.
6.2 | The mains cable of the unit being quoted for must be the hospital grade type and it must be of an adequate length (minimum length of three (3) meters).
6.3 | The power input must be 220-240V, 50Hz AC. Bidders must ensure that the product quoted for is fitted appropriately with a 15 Amp SABS approved mains plug that is held together by two screws.
6.4 | A 3 phase in / out power supply must be offered. State details.
6.5 | The equipment tendered for must not overload and trip the hospital power in the area where it is installed. Resettable overcurrent breaker/s must be fitted for protection. State details.
6.6 | All available online Uninterrupted Power Supply (UPS system) and line voltage regulators must be included for the system with a support time of 50 to 60 minutes. Attach brochure.
6.7 | The lifetime of UPS must be at least 15 years.
6.8 | An Uninterrupted Power Supply (UPS system) conforming to IS-302 must be offered for the computer system.
6.9 | The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade type and it must be a minimum length of three (3) meters.
6.10. | The mains cable of the unit tendered for must be SABS colour coded. State details.
6.11 | The equipment quoted for must be protected against electromagnetic interference. State details of measures that will be taken to protect the unit against electromagnetic interference.
6.12 | Maintenance free batteries must be offered.
7 | AIR CONDITIONER
7.1 | Bidder must include any air conditioning that is considered necessary for optimal functioning of the unit.
7.2 | Any such air conditioning must be covered by the guarantee bided for on the unit and servicing of the air conditioning units for the guarantee period and must be included in the bid price.
7.3 | The air conditioner must also be included in the maintenance plan.
7.4 | Successful bidder to replace existing air conditioner with an air conditioner of 24,000 BTU. (Where applicable).
7.5 | The air conditioner must be but no less than 24,000 BTU. State details on the air conditioner to be supplied according to the room size and machine requirements.
7.6 | The Successful bidder must install an outlet-plumbing pipe for water overflow from air conditioner to be re-directed to a drainage system.
7.7 | Air-conditioning must be included in Maintenance plan

## Digital mobile unit
DIGITAL MOBILE X-RAY UNIT
Item | SPECIFICATIONS
1.1 | The unit must be a light weight, motorized, compact system and easy to manoeuvre.
1.2 | The unit must have a built- in battery.
1.3 | The unit must have a lock function.
1.4 | The unit must have electro-magnetic locks.
1.5 | Tube rotation must be 180 degrees around the vertical axis.
1.6 | A manual collimator with +90 degrees and –90 degrees swivel must be included.
1.7 | “All free” lock control buttons on the collimator handle must be available for fast positioning, including the Film Focus Distance (FFD) setting, tube rotation, extension and contraction of the cross-arm.
1.8 | The unit's focus to floor distance should at least be 1500 mm.
1.9 | The unit must easily pass through a standard doorway (at least 800 mm width, height 2100mm).
1.10 | The unit must have a lead apron hanger.
1.11 | The unit offered must be able to make sharp spot turning.
1.12 | Large sturdy wheels should be incorporated.
1.13 | The unit should have a tube storage compartment in front of the control panel for safe transport.
2 | DIGITAL MOBILE FLAT PANEL DETECTOR WITH IMAGE PROCESSING
2.1 | The detectors must be of the flat panel type and wireless.
2.2 | 2 flat panel detectors must be provided (sizes at least 35x43cm for adults & 24x30cm for children).
2.3 | Detectors must be of Cesium Iodide (Csl) material.
2.4 | The pixel size (matrix size) must be not > 150µm
2.5 | The acquisition depth must at least be 12 bit.
2.6 | Weight of the detector panel must not be more than 5kg including the battery.
2.7 | The detectors must be supplied with 1 additional battery.
2.8 | The unit must have detector storage compartments for various sizes
2.9 | The charging station for the batteries must be supplied.
2.10 | A grid for the wireless detectors must be provided
2.11 | The flat panel battery:
2.11.1 | The battery should be able to acquire and transfer at least 90 images or more when fully charged. Details are stated in the supporting documents provided .
2.11.2 | The battery should be able to handle 5 hours of active use and 15 hours standby, when fully charged. Details are stated in the supporting documents provided .
2.11.3 | Charging time for each battery from fully depleted to fully charged should be less than 2 hours. Details are stated in the supporting documents provided.
2.11.4 | The batteries must be interchangeable between the 24 x30 cm and 43 x 43 cm panels.
2.11.5 | The battery charger must be light weight and should weigh less than 2 kg. Details of the weight of the battery charger are stated in the supporting documents provided .
2.11.6 | The battery must be able to operate at ambient temperature range of minumum of 5°C to maximum of 35 °C.
2.11.7 | The battery must be able to operate at ambient humidity of minimum humidity of 20 % to maximum of 85% RH.
2.11.8 | The battery must be able tolerate storage at a temperature range of minumum of -30 to to maximum of 60 degrees celsius.
2.11.9 | The battery must be able tolerate storage at a minimum humidity of 10 to maximum of 85% RH.
2.12 | The Grid
2.12.1 | An encasement with a built in grid that fits the 43cm x 43cm detector must be supplied.
2.12.2 | The grid must have at least:
2.12.3 | Grid ratio 8:1
2.12.4 | 100cm to 180cm focal distance
2.13 | Load capacity
2.13.1 | The uniform load that the detector must be able to withstand should be at least 150kg. Details are stated in the supporting documents provided .
2.13.2 | The local load (on an area of 40 mm diameter) that the detector must be able to withstand should be at least 100kg. Details are stated in the supporting documents provided .
2.14 | Reference image to be displayed on the mobile monitor within at least 3 seconds.
2.15 | Storage of not less than 2,000 images must be possible.
2.16 | DICOM format transfer to an image server and/or laser printer must be possible.
2.17 | After exposure, various types of processing must be possible to obtain the required image information selectively.
3 | X-Ray Generator (Micro-Processor Controlled High Frequency)
3.1 | The nominal kilowatt output rating may not be less than 30 kW.
3.2 | The system must be single phase( 230 V, 50Hz) system.
3.3 | The unit should have anatomical programming facilities with at least 300 programs.
3.4 | Automatic overload protection must be standard on the generator with error code readout.
3.5 | Automatic mains compensation must be standard on the system (State the tolerance of mains power fluctuation permissible on the system).
3.6 | Minimum exposure time may not be greater than 1 millisecond.
3.7 | The kV range must be from 40 kV (minimum) to 125 kV (maximum).
3.8 | At 60 kV the unit must be able to give at least 320 mA.
3.9 | The generator must have a high speed starter unit to drive the x-ray tube to at least 3000 rpm.
4 | X-Ray Tube and Cables
4.1 | The unit must have high speed anode rotation with dual focus x-ray tube.
4.2 | The anode rotation speed must be at least 3000 rpm.
4.3 | The unit must have focus sizes of at least 0,6 mm (small) and 1,2 mm (large).
4.4 | A full size multileaf collimator unit with integrated centering lamp must be offered. Double-slot collimator units are not acceptable.
4.5 | A rotation flange must be provided with the unit for full rotation of the collimator on the x-ray tube.
4.6 | Minimum anode heat storage capacity must be at least 100 kHU.
5 | Control Console
5.1 | The control console must have the following patient data entering functions:
5.1.1 | Automatically retrieve patient data via LAN and Wireless network
5.1.2 | Manual patient registration must be possible
5.2 | Touch panel LCD monitor or latest technology.
5.3 | Monitor size must be at least 15 inch
5.4 | Image processing must be possible
5.5 | The Menu Control should be able to do the following:
5.5.1 | Organ program selection
5.5.2 | Window position/width
5.5.3 | Horizontal and vertical image mirroring
5.5.4 | Image rotation
5.5.5 | Antero-Posterior/ Postero-Anterior views
5.5.6 | Configurable text annotation
5.5.7 | Filter selection
5.5.8 | Image zoom
5.5.9 | Must be able to print different layouts of selected images.
5.5.10 | Edge enhancement and noise suppression
5.5.11 | Quality management program including reject analysis
5 | RADIATION PROTECTION
5.1 | The system must have the ability to display Exposure Indicator values (EI/ S-values), immediately after exposure.
5.2 | X-ray Warning Notice for display at the X-ray Room entrance (not branded).
5.3 | An advanced Retake / Reject / Repeat analysis software package that includes analysis as well as calculations in percentages, must be provided for applicable systems.
5.4 | The applicable systems must have an integrated Dose Area Product (DAP) meter that includes skin dose equivalent measurement functions.
5.5 | A radiation report must be automatically generated and transferable via DICOM.
5.6 | Should be FDA or CE approved product
5.7 | Electrical safety conforms to standards for electrical safety IEC-60601
5.8 | The unit must comply to IEC-61267
5.9 | The equipment/system must be approved and licensed by the Directorate Radiation Control of DoH: South Africa
5.10. | A copy of a valid license issued in terms of the Hazardous Substance Act, Act No 15 of 1973 must be submitted with the tender. Failure to submit such a valid license may result in a tender not being considered.
5.11 | Acceptance testing must be performed as part of the installation. (Failure to do will lead to failure to pay
5.12 | Lead Glass window to be installed as per room requirement
5.13 | Safety aspects of Radiation dosage leakage should be spelt out
5.14 | DB boards and electrical filters to be installed
5.15 | QA TOOLS
5.16 | Display Exposure index values
5.17 | Radiation protection Lead equivalence of the radiation protection screen must be 0.5m Pb and must be convex, covering the radiographer from various angles against scatter radiation.
5.18 | Patient Conventional face shield to be included.
5.19 | Patient Tomosynthesis face shield to be included, removable and retractable.
5.20. | Emergency stop controls must be supplied on both sides of the gantry and on the Acquisition Workstation
5.21 | Interlocks to stop system movement when compression is applied must be provided.
5.22 | Interlock for premature release of the X-ray switch must be available
5.23 | Interlock preventing an exposure whenever the filter or light field is not positioned correctly must be possible.
5.24 | Phantom for calibration should be provided.
6 | POWER SUPPLY
6.1 | The bidder must install adequate electrical power supply for the optimal functionality of the equipment.
6.2 | The mains cable of the unit being quoted for must be the hospital grade type and it must be of an adequate length (minimum length of three (3) meters).
6.3 | The power input must be 220-240V, 50Hz AC. Bidders must ensure that the product quoted for is fitted appropriately with a 15 Amp SABS approved mains plug that is held together by two screws.
6.4 | A 3 phase in / out power supply must be offered. State details.
6.5 | The equipment tendered for must not overload and trip the hospital power in the area where it is installed. Resettable overcurrent breaker/s must be fitted for protection. State details.
6.6 | All available online Uninterrupted Power Supply (UPS system) and line voltage regulators must be included for the system with a support time of 50 to 60 minutes. Attach brochure.
6.7 | The lifetime of UPS must be at least 15 years.
6.8 | An Uninterrupted Power Supply (UPS system) conforming to IS-302 must be offered for the computer system.
6.9 | The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade type and it must be a minimum length of three (3) meters.
6.10. | The mains cable of the unit tendered for must be SABS colour coded. State details.
6.11 | The equipment quoted for must be protected against electromagnetic interference. State details of measures that will be taken to protect the unit against electromagnetic interference.
6.12 | Maintenance free batteries must be offered.
7 | AIR CONDITIONER
7.1 | Bidder must include any air conditioning that is considered necessary for optimal functioning of the unit.
7.2 | Any such air conditioning must be covered by the guarantee bided for on the unit and servicing of the air conditioning units for the guarantee period and must be included in the bid price.
7.3 | The air conditioner must also be included in the maintenance plan.
7.4 | Successful bidder to replace existing air conditioner with an air conditioner of 24,000 BTU. (Where applicable).
7.5 | The air conditioner must be but no less than 24,000 BTU. State details on the air conditioner to be supplied according to the room size and machine requirements.
7.6 | The Successful bidder must install an outlet-plumbing pipe for water overflow from air conditioner to be re-directed to a drainage system.
7.7 | Air-conditioning must be included in Maintenance plan

## X-Ray mobile C-arm motoriesed
Mobile C – Arm image intensifier system
ITEM | TECHNICAL SPECIFICATION
1 | X-RAY GENERATOR AND X-RAY TUBE
1.1 | The generator must be of high frequency type and have an output of at least 3kW
Radiography 40kV - 110kV at 65mA or more
Fluoroscopy 40kV - 110kV and 0mA - 15mA
State generator type
1.2 | Automatic dose control regulating both kV and mA must be included, but it must be possible to select factors manually.
1.3 | A fixed or rotating anode system may be offered. State the type of system offered.
1.4 | The anode heat capacity should be at least 40 kHU or greater.
1.5 | X-Ray Tube Focal Spots must not be more than 0.6mm x 0.6mm for Fluoroscopy and Radiography
2 | COLLIMATION
2.1 | The collimation system must include both circular and slot diaphragms that are able to rotate. Facilities for positioning the diaphragms without radiation must be present in order to reduce radiation dose to the staff
2.2 | The total filtration of the x–ray tube and collimator should be at least 3 mm Al.
State the filtration offered:
2.3 | A DAP (Dose Area Product) meter must be supplied with the system offered.
3 | IMAGE INTENSIFIER/T.V. CHAIN
3.1 | The image intensifier must be dual field with the largest field at least 31cm.
State the size of field:
State actual size coverage for imaging for each unit offered. Documentary proof of the size covered must be supplied.
3.2 | Television System must have two (2) monitors of at least 19 inches. Liquid Crystal Display (LCD) or Light Emitting Diode (LED).
3.3 | The system offered must have a Last Image Hold (LIH) facility.
Indication must be displayed on the monitor to differentiate between the live and stored image.
3.4 | The bidder must state the resolution of the system __________ Lp/cm
3.5 | The image on the monitor must be able to electronically rotate.
3.6 | A timer indicating lapsed screening time must be incorporated.
4 | MOBILE STAND
4.1 | The Mobile C-Arm and Control Cabinet must be easily maneuverable.
· C-arm mass of not more than 360 kg
· Control cabinet of not more than 200 kg
4.2 | The C-Arm should have a motor-driven height adjustment of 450mm.
4.3 | The stand must have castors with cable deflectors to facilitate movement.
4.4 | A steering handle controlling the castors must be incorporated to facilitate easy manoeuvering of the unit.
4.5 | It must be possible to rotate the castors through 90º so that the unit can be moved sideways.
A central floor lock must be provided.
4.6 | The movements of the C-Arm must be counter balanced and the locking system must have electromagnetic or mechanical brakes.
(a) Orbital 90º + 25º
(b) Angulation of C-Arm must be at least 190º
i.e. image intensifier can be used either above or below patient.
(c) Swivel of C-Arm (Wig Wag) - 10º in both directions.
(d) Movement of C-Arm away from control desk at least 200mm.
4.7 | Radius of C-Arm must be at least 600mm.
4.8 | Distance between the front of the Image Intensifier to the front of the X-Ray Collimator must be at least 700mm.
4.9 | With mobile equipment the positioning of a unit at an operating table is sometimes restricted because distance between the tube/image intensifier and the base of the unit is limited. State distance between the front of the base and a point in line with the focal spot.
Bidders must provide a diagram with markings indicating this distance.
4.10. | A grid permanently mounted on image intensifier must be provided and must be covered with perspex.
4.11 | The cables must be integrated into the “C” of the C-Arm and all surfaces must be smooth and easy to clean.
4.12 | The supplier must supply ten (10) sterile cloth covers for image intensifier and ten (10) sterile cloth for tube head.
5 | IMAGING
5.1 | Fluoroscopy must be in continuous and in pulsed mode.
5.2 | The unit must provide facilities for storage of atleast 2 TB.
5.3 | A thermal printer for producing hard copies sized 80mm x 130mm
5.4 | The printer must be integrated into the mobile cart.
5.5 | Supply details of the printer eg. make, model, type of print and size.
5.6 | A foot switch and a hand switch to control fluoroscopy must be included.
5.7 | All software for DICOM compatibility must be included as well as licensing for DICOM worklist and sending to PACS.
6 | RADIATION PROTECTION
6.1 | The system must have the ability to display Exposure Indicator values (EI/ S-values), immediately after exposure.
6.2 | X-ray Warning Notice for display at the X-ray Room entrance (not branded).
6.3 | An advanced Retake / Reject / Repeat analysis software package that includes analysis as well as calculations in percentages, must be provided for applicable systems.
6.4 | The applicable systems must have an integrated Dose Area Product (DAP) meter that includes skin dose equivalent measurement functions.
6.5 | A radiation report must be automatically generated and transferable via DICOM.
6.6 | Should be FDA or CE approved product
6.7 | Electrical safety conforms to standards for electrical safety IEC-60601
6.8 | The unit must comply to IEC-61267
6.9 | The equipment/system must be approved and licensed by the Directorate Radiation Control of DoH: South Africa
6.10. | A copy of a valid license issued in terms of the Hazardous Substance Act, Act No 15 of 1973 must be submitted with the tender. Failure to submit such a valid license may result in a tender not being considered.
6.11 | Acceptance testing must be performed as part of the installation. (Failure to do will lead to failure to pay
6.12 | Lead Glass window to be installed as per room requirement
6.13 | Safety aspects of Radiation dosage leakage should be spelt out
6.14 | DB boards and electrical filters to be installed
6.15 | QA TOOLS
6.16 | Display Exposure index values
6.17 | Radiation protection Lead equivalence of the radiation protection screen must be 0.5m Pb and must be convex, covering the radiographer from various angles against scatter radiation.
6.18 | Patient Conventional face shield to be included.
6.19 | Patient Tomosynthesis face shield to be included, removable and retractable.
6.20. | Emergency stop controls must be supplied on both sides of the gantry and on the Acquisition Workstation
6.21 | Interlocks to stop system movement when compression is applied must be provided.
6.22 | Interlock for premature release of the X-ray switch must be available
6.23 | Interlock preventing an exposure whenever the filter or light field is not positioned correctly must be possible.
6.24 | Phantom for calibration should be provided.
7 | POWER SUPPLY
7.1 | The bidder must install adequate electrical power supply for the optimal functionality of the equipment.
7.2 | The mains cable of the unit being quoted for must be the hospital grade type and it must be of an adequate length (minimum length of three (3) meters).
7.3 | The power input must be 220-240V, 50Hz AC. Bidders must ensure that the product quoted for is fitted appropriately with a 15 Amp SABS approved mains plug that is held together by two screws.
7.4 | A 3 phase in / out power supply must be offered. State details.
7.5 | The equipment tendered for must not overload and trip the hospital power in the area where it is installed. Resettable overcurrent breaker/s must be fitted for protection. State details.
7.6 | All available online Uninterrupted Power Supply (UPS system) and line voltage regulators must be included for the system with a support time of 50 to 60 minutes. Attach brochure.
7.7 | The lifetime of UPS must be at least 15 years.
7.8 | An Uninterrupted Power Supply (UPS system) conforming to IS-302 must be offered for the computer system.
7.9 | The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade type and it must be a minimum length of three (3) meters.
7.10. | The mains cable of the unit tendered for must be SABS colour coded. State details.
7.11 | The equipment quoted for must be protected against electromagnetic interference. State details of measures that will be taken to protect the unit against electromagnetic interference.
7.12 | Maintenance free batteries must be offered.
8 | AIR CONDITIONER
8.1 | Bidder must include any air conditioning that is considered necessary for optimal functioning of the unit.
8.2 | Any such air conditioning must be covered by the guarantee bided for on the unit and servicing of the air conditioning units for the guarantee period and must be included in the bid price.
8.3 | The air conditioner must also be included in the maintenance plan.
8.4 | Successful bidder to replace existing air conditioner with an air conditioner of 24,000 BTU. (Where applicable).
8.5 | The air conditioner must be but no less than 24,000 BTU. State details on the air conditioner to be supplied according to the room size and machine requirements.
8.6 | The Successful bidder must install an outlet-plumbing pipe for water overflow from air conditioner to be re-directed to a drainage system.
8.7 | Air-conditioning must be included in Maintenance plan
9 | OPTIONAL EXTRAS
Orthopaedic softwares
Vascular softwares
This must be listed separately and the cost of each software package must be supplied on the schedule of optional accessories at the end of this specification.

## 3D C-Arm Flat Panel 10kW
3D C-Arm Flat Panel (10kW)
ITEM | TECHNICAL SPECIFICATIONS
1 | GENERATOR
1.1 | Generator must be of medium/high frequency type and have an 10kW output.
a. Radiography 110kV at 16 – 20 mA
b. Fluoroscopy a minimum of 110kV and 0.5 – 3 mA
c. State generator type:
1.2 | Automatic dose control regulating both kV and mA must be included, but manual override to be available.
1.3 | Fixed or rotating anode system may be offered. Bidder to state type.
1.4 | Anode heat capacity should be at least 40 kHU.
1.5 | X-ray tube focal spots must not be more than 0.6mm x 0.6mm for fluoroscopy and radiography.
1.6 | A High mode pulsed fluoroscopy must be possible for larger patients
1.7 | Laser positioning device mounted on generator.
1.8 | Collimation
a. System to include circular and or slot diaphragms that are able to rotate. Facility to position the diagphram without radiation must be present to reduce exposure,
b. Total filtration of x-ray tube and collimator should be at least 3mm AI.
c. Dose area product (DAP) meter must be supplied.
1.9 | System must be easily maneuverable
2 | Flat Panel Detector System
2.1 | Grid requirement: Bidder to specify type.
2.2 | Bidder to state resolution of system in Lp/cm
2.3 | Panel size of minimum 30cm x 30cm
3 | User Interface (UI)
3.1 | It must be premised on an End-User oriented design and must comply with the following:
3.2 | Touch screens with intuitive icons
3.3 | Emergency stop button
3.4 | Radiation lamp
4 | C-arm mobile stand
4.1 | Visual display of radiation warning signs.
4.2 | 3D acquisition must be motorized
4.3 | Mass of no more than 500kg
4.4 | Motor driven vertical travel more than 40cm
4.5 | Movements of the C-arm must be counter balanced and the locking system must have electromagnetic or mechanical brakes.
a. Orbital 90°both directions, Bidder to state range
b. Angulation of C-arm must be at least 190°
c. Swivel of C-arm 10° both directions
d. Movement of C-arm away from control desk at least 20cm.
Colour coding of C-arm movements.
4.6 | Stand castors must have cable deflectors
4.7 | Steering handle controlling castors must be incorporated and allow for rotation through 90°
4.8 | A central floor lock/brake for stand
4.9 | Vertical free space to be at least 70cm
4.10. | Must come with at least 2 footswitch pedals
4.11 | The whole C-arm system must be able to through a standard door entrance.
5 | Monitor cart
5.1 | Monitors
a. Dual flat screen- high definition
b. Minimum size of 19 inch touch screen
c. Connectors for external monitors
5.2 | Control cabinet of no more than 200kg
5.3 | Timer indicating lapsed screening time
5.4 | Integrated thermal printer to be supplied: Bidder to state make, model and size of prints
5.5 | 3D technology hardware and software:
a. Control software for 3D visualisation and reconstruction
b. 3D image acquisition
5.6 | Digital image processing:
a. Digital rotation of image
b. Grayscale inversion
c. Last image hold
d. Orthopaedic software (anatomic/3D/patient/metal)
e. Metal artifacts reduction software
f. Continuous and pulsed mode fluoroscopy
5.7 | Digital storage:
a. Storage capacity minimum of 4TB: Bidder to state amount
b. USB port
c. DICOM networking and software sending to PACS
5.8 | It must be compatible with navigation
6 | RADIATION PROTECTION
6.1 | The system must have the ability to display Exposure Indicator values (EI/ S-values), immediately after exposure.
6.2 | X-ray Warning Notice for display at the X-ray Room entrance (not branded).
6.3 | An advanced Retake / Reject / Repeat analysis software package that includes analysis as well as calculations in percentages, must be provided for applicable systems.
6.4 | The applicable systems must have an integrated Dose Area Product (DAP) meter that includes skin dose equivalent measurement functions.
6.5 | A radiation report must be automatically generated and transferable via DICOM.
6.6 | Should be FDA or CE approved product
6.7 | Electrical safety conforms to standards for electrical safety IEC-60601
6.8 | The unit must comply to IEC-61267
6.9 | The equipment/system must be approved and licensed by the Directorate Radiation Control of DoH: South Africa
6.10. | A copy of a valid license issued in terms of the Hazardous Substance Act, Act No 15 of 1973 must be submitted with the tender. Failure to submit such a valid license may result in a tender not being considered.
6.11 | Acceptance testing must be performed as part of the installation. (Failure to do will lead to failure to pay
6.12 | Lead Glass window to be installed as per room requirement
6.13 | Safety aspects of Radiation dosage leakage should be spelt out
6.14 | DB boards and electrical filters to be installed
6.15 | QA TOOLS
6.16 | Display Exposure index values
6.17 | Radiation protection Lead equivalence of the radiation protection screen must be 0.5m Pb and must be convex, covering the radiographer from various angles against scatter radiation.
6.18 | Patient Conventional face shield to be included.
6.19 | Patient Tomosynthesis face shield to be included, removable and retractable.
6.20. | Emergency stop controls must be supplied on both sides of the gantry and on the Acquisition Workstation
6.21 | Interlocks to stop system movement when compression is applied must be provided.
6.22 | Interlock for premature release of the X-ray switch must be available
6.23 | Interlock preventing an exposure whenever the filter or light field is not positioned correctly must be possible.
6.24 | Phantom for calibration should be provided.
7 | POWER SUPPLY
7.1 | The bidder must install adequate electrical power supply for the optimal functionality of the equipment.
7.2 | The mains cable of the unit being quoted for must be the hospital grade type and it must be of an adequate length (minimum length of three (3) meters).
7.3 | The power input must be 220-240V, 50Hz AC. Bidders must ensure that the product quoted for is fitted appropriately with a 15 Amp SABS approved mains plug that is held together by two screws.
7.4 | A 3 phase in / out power supply must be offered. State details.
7.5 | The equipment tendered for must not overload and trip the hospital power in the area where it is installed. Resettable overcurrent breaker/s must be fitted for protection. State details.
7.6 | All available online Uninterrupted Power Supply (UPS system) and line voltage regulators must be included for the system with a support time of 50 to 60 minutes. Attach brochure.
7.7 | The lifetime of UPS must be at least 15 years.
7.8 | An Uninterrupted Power Supply (UPS system) conforming to IS-302 must be offered for the computer system.
7.9 | The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade type and it must be a minimum length of three (3) meters.
7.10. | The mains cable of the unit tendered for must be SABS colour coded. State details.
7.11 | The equipment quoted for must be protected against electromagnetic interference. State details of measures that will be taken to protect the unit against electromagnetic interference.
7.12 | Maintenance free batteries must be offered.
8 | AIR CONDITIONER
8.1 | Bidder must include any air conditioning that is considered necessary for optimal functioning of the unit.
8.2 | Any such air conditioning must be covered by the guarantee bided for on the unit and servicing of the air conditioning units for the guarantee period and must be included in the bid price.
8.3 | The air conditioner must also be included in the maintenance plan.
8.4 | Successful bidder to replace existing air conditioner with an air conditioner of 24,000 BTU. (Where applicable).
8.5 | The air conditioner must be but no less than 24,000 BTU. State details on the air conditioner to be supplied according to the room size and machine requirements.
8.6 | The Successful bidder must install an outlet-plumbing pipe for water overflow from air conditioner to be re-directed to a drainage system.
8.7 | Air-conditioning must be included in Maintenance plan
9 | Software Package options:
a. Orthopaedic software (anatomic/3D/patient/metal)
b. Vascular software (anatomic/3D/patient)

## 3D C-Arm Flat Panel 7kW
3D C-Arm Flat Panel (7kW)
ITEM | TECHNICAL SPECIFICATIONS
1 | GENERATOR
1.1 | Generator must be of medium/high frequency type and have an 7kW output.
a. Radiography 110kV at 16 – 20 mA
b. Fluoroscopy a minimum of 110kV and 0.5 – 3 mA
c. State generator type:
1.2 | Automatic dose control regulating both kV and mA must be included, but manual override to be available.
1.3 | Fixed or rotating anode system may be offered. Bidder to state type.
1.4 | Anode heat capacity should be at least 40 kHU.
1.5 | X-ray tube focal spots must not be more than 0.6mm x 0.6mm for fluoroscopy and radiography.
1.6 | A High mode pulsed fluoroscopy must be possible for larger patients
1.7 | Laser positioning device mounted on generator.
1.8 | Collimation
a. System to include circular and or slot diaphragms that are able to rotate. Facility to position the diagphram without radiation must be present to reduce exposure,
b. Total filtration of x-ray tube and collimator should be at least 3mm AI.
c. Dose area product (DAP) meter must be supplied.
1.9 | System must be easily maneuverable
2 | Flat Panel Detector System
2.1 | Grid requirement: Bidder to specify type.
2.2 | Bidder to state resolution of system in Lp/cm
2.3 | Panel size of minimum 30cm x 30cm
3 | User Interface (UI)
3.1 | It must be premised on an End-User oriented design and must comply with the following:
3.2 | Touch screens with intuitive icons
3.3 | Emergency stop button
3.4 | Radiation lamp
4 | C-arm mobile stand
4.1 | Visual display of radiation warning signs.
4.2 | 3D acquisition must be motorized
4.3 | Mass of no more than 500kg
4.4 | Motor driven vertical travel more than 40cm
4.5 | Movements of the C-arm must be counter balanced and the locking system must have electromagnetic or mechanical brakes.
a. Orbital 90°both directions, Bidder to state range
b. Angulation of C-arm must be at least 190°
c. Swivel of C-arm 10° both directions
d. Movement of C-arm away from control desk at least 20cm.
Colour coding of C-arm movements.
4.6 | Stand castors must have cable deflectors
4.7 | Steering handle controlling castors must be incorporated and allow for rotation through 90°
4.8 | A central floor lock/brake for stand
4.9 | Vertical free space to be at least 70cm
4.10. | Must come with at least 2 footswitch pedals
4.11 | The whole C-arm system must be able to through a standard door entrance.
5 | Monitor cart
5.1 | Monitors
a. Dual flat screen- high definition
b. Minimum size of 19 inch touch screen
c. Connectors for external monitors
5.2 | Control cabinet of no more than 200kg
5.3 | Timer indicating lapsed screening time
5.4 | Integrated thermal printer to be supplied: Bidder to state make, model and size of prints
5.5 | 3D technology hardware and software:
a. Control software for 3D visualisation and reconstruction
b. 3D image acquisition
5.6 | Digital image processing:
a. Digital rotation of image
b. Grayscale inversion
c. Last image hold
d. Orthopaedic software (anatomic/3D/patient/metal)
e. Metal artifacts reduction software
f. Continuous and pulsed mode fluoroscopy
5.7 | Digital storage:
a. Storage capacity minimum of 4TB: Bidder to state amount
b. USB port
c. DICOM networking and software sending to PACS
5.8 | It must be compatible with navigation
6 | RADIATION PROTECTION
6.1 | The system must have the ability to display Exposure Indicator values (EI/ S-values), immediately after exposure.
6.2 | X-ray Warning Notice for display at the X-ray Room entrance (not branded).
6.3 | An advanced Retake / Reject / Repeat analysis software package that includes analysis as well as calculations in percentages, must be provided for applicable systems.
6.4 | The applicable systems must have an integrated Dose Area Product (DAP) meter that includes skin dose equivalent measurement functions.
6.5 | A radiation report must be automatically generated and transferable via DICOM.
6.6 | Should be FDA or CE approved product
6.7 | Electrical safety conforms to standards for electrical safety IEC-60601
6.8 | The unit must comply to IEC-61267
6.9 | The equipment/system must be approved and licensed by the Directorate Radiation Control of DoH: South Africa
6.10. | A copy of a valid license issued in terms of the Hazardous Substance Act, Act No 15 of 1973 must be submitted with the tender. Failure to submit such a valid license may result in a tender not being considered.
6.11 | Acceptance testing must be performed as part of the installation. (Failure to do will lead to failure to pay
6.12 | Lead Glass window to be installed as per room requirement
6.13 | Safety aspects of Radiation dosage leakage should be spelt out
6.14 | DB boards and electrical filters to be installed
6.15 | QA TOOLS
6.16 | Display Exposure index values
6.17 | Radiation protection Lead equivalence of the radiation protection screen must be 0.5m Pb and must be convex, covering the radiographer from various angles against scatter radiation.
6.18 | Patient Conventional face shield to be included.
6.19 | Patient Tomosynthesis face shield to be included, removable and retractable.
6.20. | Emergency stop controls must be supplied on both sides of the gantry and on the Acquisition Workstation
6.21 | Interlocks to stop system movement when compression is applied must be provided.
6.22 | Interlock for premature release of the X-ray switch must be available
6.23 | Interlock preventing an exposure whenever the filter or light field is not positioned correctly must be possible.
6.24 | Phantom for calibration should be provided.
7 | POWER SUPPLY
7.1 | The bidder must install adequate electrical power supply for the optimal functionality of the equipment.
7.2 | The mains cable of the unit being quoted for must be the hospital grade type and it must be of an adequate length (minimum length of three (3) meters).
7.3 | The power input must be 220-240V, 50Hz AC. Bidders must ensure that the product quoted for is fitted appropriately with a 15 Amp SABS approved mains plug that is held together by two screws.
7.4 | A 3 phase in / out power supply must be offered. State details.
7.5 | The equipment tendered for must not overload and trip the hospital power in the area where it is installed. Resettable overcurrent breaker/s must be fitted for protection. State details.
7.6 | All available online Uninterrupted Power Supply (UPS system) and line voltage regulators must be included for the system with a support time of 50 to 60 minutes. Attach brochure.
7.7 | The lifetime of UPS must be at least 15 years.
7.8 | An Uninterrupted Power Supply (UPS system) conforming to IS-302 must be offered for the computer system.
7.9 | The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade type and it must be a minimum length of three (3) meters.
7.10. | The mains cable of the unit tendered for must be SABS colour coded. State details.
7.11 | The equipment quoted for must be protected against electromagnetic interference. State details of measures that will be taken to protect the unit against electromagnetic interference.
7.12 | Maintenance free batteries must be offered.
8 | AIR CONDITIONER
8.1 | Bidder must include any air conditioning that is considered necessary for optimal functioning of the unit.
8.2 | Any such air conditioning must be covered by the guarantee bided for on the unit and servicing of the air conditioning units for the guarantee period and must be included in the bid price.
8.3 | The air conditioner must also be included in the maintenance plan.
8.4 | Successful bidder to replace existing air conditioner with an air conditioner of 24,000 BTU. (Where applicable).
8.5 | The air conditioner must be but no less than 24,000 BTU. State details on the air conditioner to be supplied according to the room size and machine requirements.
8.6 | The Successful bidder must install an outlet-plumbing pipe for water overflow from air conditioner to be re-directed to a drainage system.
8.7 | Air-conditioning must be included in Maintenance plan
9 | Software Package options:
a. Orthopaedic software (anatomic/3D/patient/metal)
b. Vascular software (anatomic/3D/patient)

## Digital X-Ray-U-Arm
DIGITAL X-RAY UNIT, U-ARM TYPE
ITEM | TECHNICAL SPECIFICATIONS
1 | CEILING SUSPENDED / MOUNTED X-RAY UNIT, U-ARM TYPE, INDEPENDENT FOR A SINGLE BAY ( IN TRAUMA UNIT)
1.1 | Counter balance system.
1.2 | One cordless / wireless detector .
1.3 | One workstation with a viewing station.
1.4 | Ease of movement to be well secured with locks at the tube head, if possible, or at the arm. Details must be stated.
1.5 | Cables to be well secured, avoid dangling above operators heads.
1.6 | Horizontal / transverse travel to be 3 meters and more, as per need in area positioned.
1.7 | The vertical travel of the U-arm must reach as low as 40-60cm above the floor.
1.8 | Distance from the detector to the X-Ray source (SID) must not be less than 100cm.
1.9 | Distance from the detector to the X-Ray source (SID) must be adjustable.
1.1 | The projection direction for the recumbent patients must include all angles between +90° and -90°.
1.11 | The X-Ray system must be equipped with collision protection.
2 | ONE DETECTOR TRAY
2.1 | The tube on the one end of the U-arm and the detector must be synchronized.
2.2 | The detector tray must be integrated with the U-arm.
2.3 | A battery pack must be provided to charge the detector.
2.4 | Tether interphase cable, LAN cable, generator interphase cables must be provided.
2.5 | The detector tray must accommodate both the detector and the removable grid.
2.6 | The grid ratio must be 10:1 lines per centimeter. The details must be stated.
2.7 | 43cm x 43cm Detector tray.
3 | ONE FLAT PANEL DETECTOR
3.1 | Pixel pitch up to 160 micrometer (µm).
3.2 | The flat panel detector must be wireless/fixed
3.3 | Detectors have acquisition depth of 14bits and above grayscale. Details are stated in supporting documents provided.
3.4 | The size of the flat panel detector must be between 35cm x 43cm to 43cm x 43cm.
3.5 | Detector to have matrix in pixels of 2000 x 2000 and above
3.6 | Patient’s information/ image preview data must be transferable in not more than 3 seconds.
3.7 | The flat panel detector must be DICOM compliant and enabled.
3.8 | The weight of the the flat panel detector must be less than 5kg.
3.9 | The flat panel detector must be of Cesium Iodide material.
3.10. | DQE must be between a range of 60%-70% @ 0 line pairs per millimeter / 50%- 60 % @ 0.05 line pairs per millimeter / 40 %-50% @ 1 line pairs per millimeter (lp/mm). DQE and MTF details are stated in supporting documents.
3.11 | The wall mounted cabinet must be installed for storage of removable wireless detector (optional) , QC test Tools and Grid.
3.12 | The flat panel battery:
3.12.1 | The battery should be able to acquire and transfer at least 90 images or more when fully charged. Details are stated in the supporting documents provided .
3.12.2 | The battery should be able to handle 5 hours of active use and 15 hours standby, when fully charged. Details are stated in the supporting documents provided .
3.12.3 | Charging time for each battery from fully depleted to fully charged should be less than 2 hours. Details are stated in the supporting documents provided.
3.12.4 | The batteries must be interchangeable between the 24 x30 cm and 43 x 43 cm panels.
3.12.5 | The battery charger must be light weight and should weigh less than 2 kg. Details of the weight of the battery charger are stated in the supporting documents provided .
3.12.6 | The battery must be able to operate at ambient temperature range of minumum of 5°C to maximum of 35 °C.
3.12.7 | The battery must be able to operate at ambient humidity of minimum humidity of 20 % to maximum of 85% RH.
3.12.8 | The battery must be able tolerate storage at a temperature range of minumum of -30 to to maximum of 60 degrees celsius.
3.12.9 | The battery must be able tolerate storage at a minimum humidity of 10 to maximum of 85% RH.
3.13 | The Grid
3.13.1 | An encasement with a built in grid that fits the 43cm x 43cm detector must be supplied.
3.13.2 | The grid must have at least:
3.13.3 | Grid ratio 8:1
3.13.4 | 100cm to 180cm focal distance
3.14 | Load capacity
3.14.1 | The uniform load that the detector must be able to withstand should be at least 150kg. Details are stated in the supporting documents provided .
3.14.2 | The local load (on an area of 40 mm diameter) that the detector must be able to withstand should be at least 100kg. Details are stated in the supporting documents provided .
4 | ONE WORKSTATION
4.1 | The workstation must be connected to the PACS system, if available at the hospital.
4.2 | The successful bidder must provide the license to connect the workstation to the PACS and link to hospital network.
4.3 | The size of the touch screen LCD monitor must be 19” or above.
4.4 | Enhanced and extensive connectivity with MHL, display port, HDMI, VGA, USB ports.
4.5 | Full HD resolution and wide viewing angle 1000 x 1000 matrix.
4.6 | Stand designed uniquely for the workstation.
4.7 | Keyboard and mouse to be included.
4.8 | Repeat and reject analysis software package must be included.
4.9 | QC software package for viewing monitors must be included (SMPTE pattern or similar.)
4.1 | Browser must be compatible to Windows 10 or higher.
5 | IMAGE ACQUISITION AND PROCESSING
5.1 | Must have an organ related application program.
5.2 | Must be able to configure the organ related application program to individual parameters.
5.3 | Pre- and post-processing of image parameters must be included.
5.4 | The following must be available on image processing:
5.5 | a. Identification of patients
5.6 | b. ID marker selection Left or Right
5.7 | c. AP/PA selection
5.8 | Image functions to include:
a. PAN and zoom capability
b. Rotation
c. Inversion
d. Windowing
e. Filtration
f. Annotation
g. Subtraction
h. Measurement
i. Image manipulation
j. Filter selection for noise suppression and edge enhancement.
5.9 | The image contrast must be adjustable.
5.10. | All updated parameters to be included at no extra cost.
6 | Image documentation and archiving:
6.1 | a. DICOM Print, Transfer, Modality, Work List, Send, Storage and Storage Commit and MPPS.
6.2 | b. All DICOM conformance statements must be provided.
6.3 | c. Storage capacity for 10 000 or more images must be available on the internal hard drive.
7 | ONE X-RAY GENERATOR
7.1 | Must be a three phase, high frequency system.
7.2 | Output must be 80kW and above.
7.3 | The minimum KV must be 40KV up to 125KV at reasonable incremental values, 1 KV step. State incremental step.
7.4 | The mA must be 800mA for a 80kW generator at 100kV. The details must be stated.
7.5 | The generator must have selectable mAs.
7.6 | The settings on the operator console must be displayed digitally.
7.7 | Troubleshooting, overload protection and error codes to be incorporated into the system.
8 | ONE CEILING SUSPENDED X-RAY TUBE ON THE U-ARM
8.1 | Rotating anode.
8.2 | Dual Focal Spot (small and large focus).
8.3 | Anode Heat Storage Capacity must be 300kHU or above.
8.4 | Multi lead collimator to have rotating flange for collimation around the central beam.
8.5 | Swivel / Rotation of collimator around its own axis +/-45 degrees.
8.6 | Light localizer with electronic switch.
9 | CPU HARD-DRIVE
9.1 | Must have a high storage capacity of 1TB.
9.2 | DVD RW and readable facilities to be available.
9.3 | Random access memory must not be less than 4GB.
9.4 | Must have a i5 core Intel processor or higher.
9.5 | Integrated graphics.
9.6 | Antivirus software protector to be provided.
10 | RADIATION PROTECTION
10.1 | The system must have the ability to display Exposure Indicator values (EI/ S-values), immediately after exposure.
10.2 | X-ray Warning Notice for display at the X-ray Room entrance (not branded).
10.3 | An advanced Retake / Reject / Repeat analysis software package that includes analysis as well as calculations in percentages, must be provided for applicable systems.
10.4 | The applicable systems must have an integrated Dose Area Product (DAP) meter that includes skin dose equivalent measurement functions.
10.5 | A radiation report must be automatically generated and transferable via DICOM.
10.6 | Should be FDA or CE approved product
10.7 | Electrical safety conforms to standards for electrical safety IEC-60601
10.8 | The unit must comply to IEC-61267
10.9 | The equipment/system must be approved and licensed by the Directorate Radiation Control of DoH: South Africa
10.10. | A copy of a valid license issued in terms of the Hazardous Substance Act, Act No 15 of 1973 must be submitted with the tender. Failure to submit such a valid license may result in a tender not being considered.
10.11 | Acceptance testing must be performed as part of the installation. (Failure to do will lead to failure to pay
10.12 | Lead Glass window to be installed as per room requirement
10.13 | Safety aspects of Radiation dosage leakage should be spelt out
10.14 | DB boards and electrical filters to be installed
10.15 | QA TOOLS
10.16 | Display Exposure index values
10.17 | Radiation protection Lead equivalence of the radiation protection screen must be 0.5m Pb and must be convex, covering the radiographer from various angles against scatter radiation.
10.18 | Patient Conventional face shield to be included.
10.19 | Patient Tomosynthesis face shield to be included, removable and retractable.
10.20. | Emergency stop controls must be supplied on both sides of the gantry and on the Acquisition Workstation
10.21 | Interlocks to stop system movement when compression is applied must be provided.
10.22 | Interlock for premature release of the X-ray switch must be available
10.23 | Interlock preventing an exposure whenever the filter or light field is not positioned correctly must be possible.
10.24 | Phantom for calibration should be provided.
11 | POWER SUPPLY
11.1 | The bidder must install adequate electrical power supply for the optimal functionality of the equipment.
11.2 | The mains cable of the unit being quoted for must be the hospital grade type and it must be of an adequate length (minimum length of three (3) meters).
11.3 | The power input must be 220-240V, 50Hz AC. Bidders must ensure that the product quoted for is fitted appropriately with a 15 Amp SABS approved mains plug that is held together by two screws.
11.4 | A 3 phase in / out power supply must be offered. State details.
11.5 | The equipment tendered for must not overload and trip the hospital power in the area where it is installed. Resettable overcurrent breaker/s must be fitted for protection. State details.
11.6 | All available online Uninterrupted Power Supply (UPS system) and line voltage regulators must be included for the system with a support time of 50 to 60 minutes. Attach brochure.
11.7 | The lifetime of UPS must be at least 15 years.
11.8 | An Uninterrupted Power Supply (UPS system) conforming to IS-302 must be offered for the computer system.
11.9 | The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade type and it must be a minimum length of three (3) meters.
11.10. | The mains cable of the unit tendered for must be SABS colour coded. State details.
11.11 | The equipment quoted for must be protected against electromagnetic interference. State details of measures that will be taken to protect the unit against electromagnetic interference.
11.12 | Maintenance free batteries must be offered.
12 | AIR CONDITIONER
12.1 | Bidder must include any air conditioning that is considered necessary for optimal functioning of the unit.
12.2 | Any such air conditioning must be covered by the guarantee bided for on the unit and servicing of the air conditioning units for the guarantee period and must be included in the bid price.
12.3 | The air conditioner must also be included in the maintenance plan.
12.4 | Successful bidder to replace existing air conditioner with an air conditioner of 24,000 BTU. (Where applicable).
12.5 | The air conditioner must be but no less than 24,000 BTU. State details on the air conditioner to be supplied according to the room size and machine requirements.
12.6 | The Successful bidder must install an outlet-plumbing pipe for water overflow from air conditioner to be re-directed to a drainage system.
12.7 | Air-conditioning must be included in Maintenance plan
12 | OPTIONAL ITEMS
12.4 | One LCD monitor of 18" or above.

## X-Ray, panorex
DIGITAL PANORAMIC DENTAL X-RAY UNIT
Item | SPECIFICATIONS | Weight
1.1 | The Unit must be a Digital Panoramic Dental X - ray Unit | #
1.2 | The unit must have a software system capable of automatically removing redundant tomographic shadows and ghost images that result in geometric unsharpness, e.g. ghost images/shadows of the cervical spine and the occiput. | #
1.3 | The unit must allow for optimum selection of exposure factors to reduce the radiation dose. | #
1.4 | It must be capable of performing image segmenting to reduce radiation, e.g. Vertical segmenting to expose a smaller area of diagnostic interest. | #
1.5 | The unit must have an automatic doubleTemporo - Mandibular Joint ( TMJ) program to perform full temporo-mandibular studies on a single display or film. | #
1.6 | The unit must have a sinus program, which will show a clear view of the maxillary sinuses. | #
1.7 | The GUI must be simple and intuitive. Text and or graphic symbols should prompt the operator. | #
1.8 | All necessary information must be displayed on the main console. | #
1.9 | The unit must be equipped with a self-diagnostic test system, capable of displaying error messages/codes shall be incorporated into the unit. | #
1.1 | The unit must have a fixed bite block which is easy to clean after each patient. | #
1.11 | The unit must incorporate a fully automatic film marking system for panoramic radiography to mark the exposure factors, the chosen program, as well as patient identification information | #
1.12 | The unit must have a stable headrest for temple support | #
1.13 | The rating of the x ray tube must be at least 2.5kW.Please supply rating charts. | #
1.14 | The tube focal spot size must range from 0,5 X 0,5 mm to 0.8x0.8mm. | #
1.15 | The total tube filtration must be 2.5mm AI. | #
1.16 | The anode voltage must be from 60kVp to 90 kVp. | #
1.17 | The tube anode current to be from 4 to 10 mA or better | #
1.18 | The unit must have exposure time of 4.5 to 20 seconds continuous or split. | #
1.19 | The unit must include an electronic kV compensation over the cervical spine. | #
1.2 | Exposure factors must be pre-programmed, and also manually selectable. | #
2 | SYSTEM FUNCTIONALITY
2.1 | The Unit must have fully customizable patient work list
2.2 | The Unit must have examination related pre-configured views
2.3 | System must set automatically all generator settings
2.4 | System must set collimators automatically
2.5 | System must set pre-filters automatically
2.6 | Manual input for patient data must be possible (Please describe)
3 | WORKSTATION
3.1 | A WINDOWS Operating computer system is required with at least 20 GB hard drive, a network card and suitable graphics card. (State capacity)
3.2 | At least a 17 inch touch screen flat panel LCD monitor (supporting at least 3 Megapixel resolution) must be supplied.
3.4 | A Mono Laser jet printer with at least 1200 dpi must be provided
4 | IT Integration
4.1 | The system must have the following DICOM compatibility:
4.1.1 | DICOM print class
4.1.2 | DICOM storage class
4.1.3 | DICOM query/retrieve
4.1.4 | DICOM worklist
4.1.5 | The system must comply to the following IHE profiles (Integrating the Healthcare Enterprise):
4.1.6 | Scheduled Workflow (SWF)
4.1.7 | Simple image and Numeric Reports (SINR)
4.1.8 | Patient Information Reconciliation (PIR)
5 | RADIATION PROTECTION
5.1 | The system must have the ability to display Exposure Indicator values (EI/ S-values), immediately after exposure.
5.2 | X-ray Warning Notice for display at the X-ray Room entrance (not branded).
5.3 | An advanced Retake / Reject / Repeat analysis software package that includes analysis as well as calculations in percentages, must be provided for applicable systems.
5.4 | The applicable systems must have an integrated Dose Area Product (DAP) meter that includes skin dose equivalent measurement functions.
5.5 | A radiation report must be automatically generated and transferable via DICOM.
5.6 | Should be FDA or CE approved product
5.7 | Electrical safety conforms to standards for electrical safety IEC-60601
5.8 | The unit must comply to IEC-61267
5.9 | The equipment/system must be approved and licensed by the Directorate Radiation Control of DoH: South Africa
5.10. | A copy of a valid license issued in terms of the Hazardous Substance Act, Act No 15 of 1973 must be submitted with the tender. Failure to submit such a valid license may result in a tender not being considered.
5.11 | Acceptance testing must be performed as part of the installation. (Failure to do will lead to failure to pay
5.12 | Lead Glass window to be installed as per room requirement
5.13 | Safety aspects of Radiation dosage leakage should be spelt out
5.14 | DB boards and electrical filters to be installed
5.15 | QA TOOLS
5.16 | Display Exposure index values
5.17 | Radiation protection Lead equivalence of the radiation protection screen must be 0.5m Pb and must be convex, covering the radiographer from various angles against scatter radiation.
5.18 | Patient Conventional face shield to be included.
5.19 | Patient Tomosynthesis face shield to be included, removable and retractable.
5.20. | Emergency stop controls must be supplied on both sides of the gantry and on the Acquisition Workstation
5.21 | Interlocks to stop system movement when compression is applied must be provided.
5.22 | Interlock for premature release of the X-ray switch must be available
5.23 | Interlock preventing an exposure whenever the filter or light field is not positioned correctly must be possible.
5.24 | Phantom for calibration should be provided.
6 | POWER SUPPLY
6.1 | The bidder must install adequate electrical power supply for the optimal functionality of the equipment.
6.2 | The mains cable of the unit being quoted for must be the hospital grade type and it must be of an adequate length (minimum length of three (3) meters).
6.3 | The power input must be 220-240V, 50Hz AC. Bidders must ensure that the product quoted for is fitted appropriately with a 15 Amp SABS approved mains plug that is held together by two screws.
6.4 | A 3 phase in / out power supply must be offered. State details.
6.5 | The equipment tendered for must not overload and trip the hospital power in the area where it is installed. Resettable overcurrent breaker/s must be fitted for protection. State details.
6.6 | All available online Uninterrupted Power Supply (UPS system) and line voltage regulators must be included for the system with a support time of 50 to 60 minutes. Attach brochure.
6.7 | The lifetime of UPS must be at least 15 years.
6.8 | An Uninterrupted Power Supply (UPS system) conforming to IS-302 must be offered for the computer system.
6.9 | The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade type and it must be a minimum length of three (3) meters.
6.10. | The mains cable of the unit tendered for must be SABS colour coded. State details.
6.11 | The equipment quoted for must be protected against electromagnetic interference. State details of measures that will be taken to protect the unit against electromagnetic interference.
6.12 | Maintenance free batteries must be offered.
7 | AIR CONDITIONER
7.1 | Bidder must include any air conditioning that is considered necessary for optimal functioning of the unit.
7.2 | Any such air conditioning must be covered by the guarantee bided for on the unit and servicing of the air conditioning units for the guarantee period and must be included in the bid price.
7.3 | The air conditioner must also be included in the maintenance plan.
7.4 | Successful bidder to replace existing air conditioner with an air conditioner of 24,000 BTU. (Where applicable).
7.5 | The air conditioner must be but no less than 24,000 BTU. State details on the air conditioner to be supplied according to the room size and machine requirements.
7.6 | The Successful bidder must install an outlet-plumbing pipe for water overflow from air conditioner to be re-directed to a drainage system.
7.7 | Air-conditioning must be included in Maintenance plan

## 64 slices CT scanner
CT SCANNER (64 SLICES)
ITEM | TECHNICAL SPECIFICATIONS
1 | GANTRY
1.1 | The system must have 64 Slices or more, producing 64 Slices or more per 360 degrees rotation
1.2 | The unit must be low voltage slip-ring designed gantry technology.
1.3 | The unit must support whole body CT Scanning.
1.4 | The unit must have a minimum scan time of 0.4 second or less for 360 degrees rotation.
1.5 | The unit must have a gantry aperture size of at least 700mm or larger.
1.6 | Angulations must at least + 30degrees and - 30 degrees.
1.7 | The Unit must be able to scan either head first or feet first.
1.8 | The unit must have a minimum field of view at least 500mm or larger.
1.9 | The unit must have control panels on either side (left and right) of table on the gantry to adjust gantry movements.
1.10. | The system must have triple laser lights for accurate patient positioning at setup plane. (Accuracy of 1mm)
1.11 | Control of gantry movements from operator console must be possible.
1.12 | The system shall have an Intercom between gantry and operator console.
1.13 | Electronic patient breathing instructions in multiple languages must be possible.
1.14 | Emergency stop switches to be available on gantry as well as on the console.
2 | PATIENT SUPPORT SYSTEM (PATIENT TABLE)
2.1 | Support patients 200kg or more.
2.2 | Lowest table height must be 550mm or less
2.3 | The table length must be 2 meters or more
2.4 | Longitudinal tabletop travel must not be less than 1700mm for whole body coverage.
2.5 | Must have motorized patient support movements (longitudinal and vertical movements)
2.6 | A metal-free tabletop over scanner range allowing patient scanning without re-positioning patient must be provided.
2.7 | Control of table movements from operator console and gantry must be available.
2.8 | Motorized removal and re-indexing of table top without repositioning patient scanning for emergency access to patient must be available.
2.9 | Patient positioning accessories required to perform the procedures specified, including headrest support, axial and coronal head rest support, table leg extender, security straps, infant immobiliser, arm support, knee support and immobilizing straps must be provided. As well as the radiolucent mattress
2.10. | Emergency override to enable manual removal of patients in case of power failure or emergency must be included
3 | DETECTOR ACQUISITION SYSTEM
3.1 | Solid-state detector technology using low-dose and high resolution acquisitions.
3.2 | Multi-slice detector technology for 64-slice acquisitions per 360 degree gantry rotation.
3.3 | Multi-Slice detector with 1200 views per second sampling rates or better.
3.4 | Pre-body collimator must be available.
3.5 | The ability to select 64-slice acquisition with 0.6mm Slice thickness or thinner. State number of channels
3.6 | All calibration and service phantoms with wall brackets must be included
4 | X-RAY GENERATOR
4.1 | Modern high frequency generator allowing at least between 80kV to 140kV and 500mA or better.
4.2 | To support at least 200 scans in 100 seconds using a 0.5 second scan time or better during the spiral scan mode.
4.3 | The unit must be able to scan continuously a length of 1.7m high resolution without cooling time.
4.4 | The unit must support repeat scanning with minimal wait time.
4.5 | The unit must have at least 3 variable kV setting between 80kV (or less) and 135kV (or more).
4.6 | The unit must have variable mA settings from 20mA up to 800mA or better.
4.7 | State maximum mA allowed for each kV.
4.8 | Generator Power output shall be 60kW or higher.
State actual power output.
5 | X-RAY TUBE
5.1 | The unit must have a high performance rotating anode tube with dual focal spot.
5.2 | The unit must have a total Filtration of 3.5mm AI equivalent or greater.
5.3 | Anode heat dissipation rate not less than 1300kHU/minute.
5.4 | Cooling performance curves required.
5.5 | The unit must have a heat storage capacity of at least 7MHU.
5.6 | Expected tube life. State in years and slices to include tube replacement in warrantee
6 | OPERATOR CONSOLE
6.1 | The operator console must have a display and entry of patient data pre-planning
6.2 | The console must have comprehensive adult and paediatric anatomical programs/protocols.
6.3 | Must at least have 100 pre-programmed scan protocols required with option to add user defined protocols.
6.4 | One operator console and two independent post -processing workstations must be included in the bid.
6.5 | The operator console to include a high speed computer which will support multi-tasking operation to allow simultaneous scanning reconstruction, viewing, archiving, processing and filming operation, 2X19 inch high resolution of a minimum 3 MP or above flat touch screen LCD colour monitors and mouse.
6.6 | Image review system must be able to display, analyze, process and manage acquired imaging data such as:
6.6.1 | - Zoom
6.6.2 | Grey scale (colour display) adjustment
6.6.3 | Measurement of distance and angle
6.6.4 | - Rotation
6.6.5 | - Grid
6.6.6 | - Annotation
6.6.7 | - Cine display
6.7 | Facilities must be available to perform basic CT examination of all parts of the body, including reconstruction, review and printing.
6.8 | All post processing software must be available on all three workstations
6.9 | Compatibility: To comply to DICOM 3 standards. Please provide DICOM 3 conformance statement.
6.10. | On line storage: - To store at least 800 Tera Bytes uncompressed images of 512 x 512 images.
6.11 | Archival storage by means of an DVD recording system must be provided
6.12 | X-Ray tube heat loading display is required on console
6.13 | Print facility to laser imager must be supported from all CT workstations.
6.14 | Chairs: wheeled, height adjustable, swivel for the operators must be supplied. (Quantity 4)
6.15 | All high end post processing software must be on both the CT console and Doctors console. (The user must be able to either configure the post processing software on CT console or Doctors console depending on workflow requirements.)
7 | COMPUTER SYSTEM / RECONSTRUCTION
7.1 | At least a 512 x 512 reconstruction matrix is required.
Details of matrix sizes available to be stated
7.2 | Real time reconstruction for scout-view required.
7.3 | A 0.1 second or faster reconstruction time is required per 512 x 512 reconstructed image
7.4 | 3 - D re-construction must be possible
8 | IMAGE STORAGE
8.1 | Large image archive to store raw scan-data of at least 5 TB as well as reconstructed data required. State size
8.2 | At least 3600 scan data sets are to be stored. State storage size.
8.3 | The Unit must be able to burn images onto CD & DVD. (200 CDs must be included)
9 | IMAGE VIEWING MODES
9.1 | Comprehensive image acquisition, viewing, processing and filming modes required must be provided and supported by the systems.
9.2 | Noise reduction techniques must be available.
10 | SCAN MODES
10.1 | Scout-view mode for scan planning must be provided.
10.2 | Single scan mode must be provided.
10.3 | Scanning with serial reconstruction must be provided.
10.4 | Scanning with parallel reconstruction must be provided.
10.5 | Dynamic scan mode must be provided.
11 | SPIRAL MODE PERFORMANCE
11.1 | Tilted spiral scan mode with 64 slice and other combination for thin 0,6mm slice acquisition being acquired per gantry rotation must be provided.
11.2 | 10 scan plans per scout view required.
11.3 | The Unit must provide a variable pitch setting for spiral scanning mode.
11.4 | Single spiral sequence must be provided.
11.5 | Multiple spiral sequence must be provided.
11.6 | Combined conventional and spiral sequence must be provided.
11.7 | Tilted spiral mode must be possible.
11.8 | Real time reconstruction and display of images during the spiral mode must be provided at a viewing rate of at 16 images per second or better.
11.9 | For accurate contrast examinations, a bolus-timing / tracking package must be included.
11.10. | Isotropic volume imaging is required. State voxel sizes used for isotropic volume imaging.
11.11 | Standard mode spatial resolution of 15 lp/cm at cut-off or better is required.
State high contrast spatial resolution in line pairs per cm.
11.12 | High-resolution mode spatial resolution of 21 lp/cm at cut-off or better is required.
State high contrast spatial resolution in line pairs per cm.
11.13 | 4.0mm at 0.3% low-contrast resolution must be possible at 90mAs, 10mm slice thickness, 120kV.
State resolution and scan technique used to obtain measurement. "
11.15 | International CATPHAN performance phantom must be provided for measurements. Supply details used for measurements.
12 | DOSE REDUCTION TECHNIQUE
12.1 | Current dose reduction techniques used during patient scanning to reduce the total applied patient dose must be included.
12.2 | The applied patient dose must be quantified and displayed on the operators console for the patient examination.
12.3 | Bidders must include current dose reduction technology specifically to the following:
12.3.1 | - Gantry design.
12.3.2 | - Choice of scan-times available.
12.3.3 | - Detector design
12.3.4 | - Use of 64 x 0.6mm (or smaller) imaging techniques.
12.3.5 | - Slice profile used
12.3.6 | - Real-time exposure control during scanning
12.3.7 | - Variable pitch settings
12.3.8 | - Beam collimation and X-ray tube filtration
12.3.9 | - Adult and paediatric scanning exposure techniques/protocols.
12.3.10 | - Carbon fibre patient support tabletop attenuation
12.3.11 | - Accurate contrast timing.
12.3.12 | - Real-time spiral scanning display
12.3.13 | - Calculated dose display on operators console monitor.
12.3.14 | - Software image quality correction that allows reconstructed images to be improved without having to rescan the patient with additional scans or scans with a higher applied dose.
12.3.15 | - For shoulders and pelvis with streak reduction techniques
12.3.16 | - Posterior fossa with image fusion techniques.
12.3.17 | - Abdomen with motion detection techniques.
12.4 | Any other techniques must be detailed.
12.5 | Supporting documentation indicating applied dose reduction must be included.
13 | DICOM 3 COMPATIBILITY
13.1 | The CT Scanner must support image & data linkage via network including all additional workstations and to laser printer. All DICOM software must be included in the cost. It must have the ability to be easily linked to any PACS and RIS system. It must have ability to be linked to ADSL/ISDN line or any other tele-radiology system
13.2 | All hardware & software relating to DICOM compatibility, including licences, must be supplied.
13.3 | The Laser printer to support DICOM compatible printing on 35 x 43 cm Laser film as wel as option for paper printer.
13.4 | Links to additional workstation and Laser printer systems is required through DICOM compatibility.
13.5 | DICOM Conformance statements on the various service classes per modality to be provided.
14 | INDEPENDENT WORKSTATIONS (the number of workstations are as per user requirements)
14.1 | Latest Windows OS with minimum 32 GB RAM or better
800 GB hard drive, network card, suitable graphics card and at least a 48cm flat panel touch LCD monitor (supporting at least 3 Megapixel resolution) is required.
14.2 | Provide modern user-interface with high speed-processing and storage facilities as per operators console.
14.3 | 2D, 3D, surface rendering, volume rendering, image sculpting, Maximum Intensity Projections, Minimum Intensity Projections, MPR and curved MPR and other image processing tools must be supported by system.
14.4 | All post processing software must be available on all workstations.
14.5 | Image manipulation and arithmetic must be available
14.6 | Archival storage using TIFF, JPEG, DICOM or similar format by means of a DVD must be included. (200 discs must be included)."
14.7 | Print facility to laser imager must be supported.
16 | SOFTWARE PACKAGES AND CLINICAL APPLICATIONS
16.1 | All packages required to perform the clinical applications as listed in SECTIONS 10.11 and 12. must be included.
16.2 | All inclusive Vascular, Oncology, Orthopedic clinical applications, Neuro DSA software with total bone removal and Metal artifact reduction must be included in the package. All other clinical applications as per list below must be available, however these must be quoted as optional:
16.2.1 | · Virtual colonoscopy package with CO2 insufflator.
16.2.2 | · CT Fluoroscopy for drainages, FNA biopsies, etc. is required, including the ability to interactively change table height, longitudinal table position and gantry angulation on remote hand control with ceiling mounted LCD monitor.
16.2.3 | · Brain perfusion package.
16.2.4 | · Cardiac package including patient monitor, ECG gated scanning, ECG gated reconstruction and full ventricle and coronary evaluation software.
16.3 | Any clinical application not listed in 6.2. but is offered by the manufacturer, must be quoted separately as an option.
17 | CT CONTRAST INJECTOR with Patient Monitor
17.1 | Tree channel syringeless injector for CT Scanner (Air and occlusion detection on fluid channels. Unidirectional flow of fluid. Pressure of 300psi and higher. Adjustable flow rate of 0.5ml/sec to 6ml/sec) and will be used for all contrast enhanced examinations including cardiac and bolus chasing. (Please supply 200 units of each of the Consumables)
17.2 | The consumables for the injector mentioned above must readily available in South Africa
17.3 | Adult/pediatric patient monitor to measure heart rate, saturation and blood pressure
18 | RADIATION PROTECTION
18.1 | The system must have the ability to display Exposure Indicator values (EI/ S-values), immediately after exposure.
18.2 | X-ray Warning Notice for display at the X-ray Room entrance (not branded).
18.3 | An advanced Retake / Reject / Repeat analysis software package that includes analysis as well as calculations in percentages, must be provided for applicable systems.
18.4 | The applicable systems must have an integrated Dose Area Product (DAP) meter that includes skin dose equivalent measurement functions.
18.5 | A radiation report must be automatically generated and transferable via DICOM.
18.6 | Should be FDA or CE approved product
18.7 | Electrical safety conforms to standards for electrical safety IEC-60601
18.8 | The unit must comply to IEC-61267
18.9 | The equipment/system must be approved and licensed by the Directorate Radiation Control of DoH: South Africa
18.10. | A copy of a valid license issued in terms of the Hazardous Substance Act, Act No 15 of 1973 must be submitted with the tender. Failure to submit such a valid license may result in a tender not being considered.
18.11 | Acceptance testing must be performed as part of the installation. (Failure to do will lead to failure to pay
18.12 | Lead Glass window to be installed as per room requirement
18.13 | Safety aspects of Radiation dosage leakage should be spelt out
18.14 | DB boards and electrical filters to be installed
18.15 | QA TOOLS
18.16 | Display Exposure index values
18.17 | Radiation protection Lead equivalence of the radiation protection screen must be 0.5m Pb and must be convex, covering the radiographer from various angles against scatter radiation.
18.18 | Patient Conventional face shield to be included.
18.19 | Patient Tomosynthesis face shield to be included, removable and retractable.
18.20. | Emergency stop controls must be supplied on both sides of the gantry and on the Acquisition Workstation
18.21 | Interlocks to stop system movement when compression is applied must be provided.
18.22 | Interlock for premature release of the X-ray switch must be available
18.23 | Interlock preventing an exposure whenever the filter or light field is not positioned correctly must be possible.
18.24 | Phantom for calibration should be provided.
19 | POWER SUPPLY
19.1 | The bidder must install adequate electrical power supply for the optimal functionality of the equipment.
19.2 | The mains cable of the unit being quoted for must be the hospital grade type and it must be of an adequate length (minimum length of three (3) meters).
19.3 | The power input must be 220-240V, 50Hz AC. Bidders must ensure that the product quoted for is fitted appropriately with a 15 Amp SABS approved mains plug that is held together by two screws.
19.4 | A 3 phase in / out power supply must be offered. State details.
19.5 | The equipment tendered for must not overload and trip the hospital power in the area where it is installed. Resettable overcurrent breaker/s must be fitted for protection. State details.
19.6 | All available online Uninterrupted Power Supply (UPS system) and line voltage regulators must be included for the system with a support time of 50 to 60 minutes. Attach brochure.
19.7 | The lifetime of UPS must be at least 15 years.
19.8 | An Uninterrupted Power Supply (UPS system) conforming to IS-302 must be offered for the computer system.
19.9 | The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade type and it must be a minimum length of three (3) meters.
19.10. | The mains cable of the unit tendered for must be SABS colour coded. State details.
19.11 | The equipment quoted for must be protected against electromagnetic interference. State details of measures that will be taken to protect the unit against electromagnetic interference.
19.12 | Maintenance free batteries must be offered.
20 | AIR CONDITIONER
20.1 | Bidder must include any air conditioning that is considered necessary for optimal functioning of the unit.
20.2 | Any such air conditioning must be covered by the guarantee bided for on the unit and servicing of the air conditioning units for the guarantee period and must be included in the bid price.
20.3 | The air conditioner must also be included in the maintenance plan.
20.4 | Successful bidder to replace existing air conditioner with an air conditioner of 24,000 BTU. (Where applicable).
20.5 | The air conditioner must be but no less than 24,000 BTU. State details on the air conditioner to be supplied according to the room size and machine requirements.
20.6 | The Successful bidder must install an outlet-plumbing pipe for water overflow from air conditioner to be re-directed to a drainage system.
20.7 | Air-conditioning must be included in Maintenance plan

## Digital Mamography
DIGITAL MAMMOGRAPHY UNIT
ITEM | GENERAL
FULL FIELD DIGITAL MAMMOGRAPHY UNIT WITH C-ARM ASSEMBLY AND ERGO DYNAMICALLY ACQUISITION WORKSTATION WITH 2D IMAGING AS WELL AS 3D DIGITAL BREAST TOMOSYNTHESIS SYSTEM
COMPUTER AIDED DIAGNOSTIC SYSTEM WITH SOFTWARE
UPRIGHT STEREO-TAXY BIOPSY SYSTEM
VACCUUM ASSISTED BREAST BIOPSY AND EXCISION SYSTEM
BIOPSY CHAIR
EXTERNAL HARD DRIVE minimum 4 TB (as option)
TECHNICAL SPECIFICATIONS
1 | FULL FIELD DIGITAL MAMMOGRAPHY UNIT WITH C-ARM ASSEMBLY AND ERGO DYNAMICALLY ACQUISITION WORKSTATION WITH 2D IMAGING AS WELL AS 3D DIGITAL BREAST TOMOSYNTHESIS SYSTEM
1.1 | GENERATOR
1.1.1 | The generator must be of Constant Potential High frequency Inverter Type
1.1.2 | The generator rating must be 7.0kW, 200mA @35kV
1.1.3 | The KV must range from 20KV to 50KV (State Range)
1.1.4 | KV increments to be 1kV
1.1.5 | The mAs must range from 3 to 400mAs (State Range)
1.1.6 | Automatic exposure device to be capable of 7 manual positions from nipple to chest wall and Auto AEC to find the optimum area of the breast
1.1.7 | The mA range must be 10 to 200mA for Large focal Spot and 10 to 50mA for small focal Spot
1.2. | X-RAY TUBE
1.2.1 | Tungsten rotating biangular anode X-ray Tube must be provided.
1.2.2 | A dual focal spot X-ray Tube must be provided at most 0.13 mm small focal and 0.33 mm large focal spot
1.2.3 | Anode heat capacity must be at least 300KHU
1.2.4 | Total inherent filtration of the X-ray tube and shield (mm) to be at least 0,63 .
1.2.5 | Filtration to be included for Tomosynthesis. State Filter:
1.2.6 | Filter selection method must be automatic and manual
1.2.7 | Tube light beam intensity must be at least 160 (LUX)
1.2.8 | Tube heat Storage must at least be 500KHU or more
1.2.9 | Fan Speed controlled in tube housing via temperature sensor must be provided.
1.3 | COLLIMATOR
1.3.1 | Collimation to be manually or automatically selected.
1.3.2 | Automatic collimation adjustment is required when compression paddles are changed
1.3.3 | X-Ray field collimation for 18x24, 24x29 ,15x15 10x10cm and 5x10cm views to be included
1.3.4 | Collimation for Tomosynthesis must be included at 18 x 29cm
1.4 | C-ARM ASSEMBLY
1.4.1 | The System must incorporate a Biopsy Ready C-arm should an upright stereotactic unit be acquired.
1.4.2 | The system must provide indented hand spaces on the side of the gantry for natural positioning of the patient hand holds
1.4.3 | The C-arm must have a Tomosynthesis Face Shield for optimum safety to the patient during Tomosynthesis and must be removable and retractable.
1.4.4 | Range of breast platform vertical movement (from floor) must be at least from 71cm to 142cm motorized.
1.4.5 | A Streamlined Tube head and SID of 70cm must be offered to provide more workspace during interventional procedures
1.4.6 | Method of angulations must be motorised from +/- +195° to -155°. State range.
1.4.7 | Angle and SID must be displayed on the Gantry and the C-arm Head as well as on the Acquisition station
1.4.8 | Compression thickness and compression force information must be displayed on both sides of the Compression device and on the Acquisition Station.
1.4.9 | Automatic release of compression after exposure must be supplied.
1.4.10 | Emergency release of compression during power failure must be available or during Exposure abortion or failure.
1.4.11 | C-arm rotation indents must be pre-programmable at any angle for up to four indents
1.5 | DIGITAL IMAGE RECEPTOR
1.5.1 | Direct capture flat panel technology must be offered.
1.5.2 | Pixel pitch must be minimum 70x70 microns
1.5.3 | Detector must be direct conversion type, eliminating the need to convert X-rays to Light. No Scintillation scanning method of imaging acquisition must be offered.
1.5.4 | Active area should be minimum 24x29 cm
1.5.5 | Resolution should be 3328x4096 pixels
1.5.6 | High contrast limiting resolution(1p/mm) must be at least 7 lines per mm
1.5.7 | State detector calibration frequency period. Calibration period to be at least 1 week valid
1.5.8 | Calibration to be performed by the user
1.5.9 | The chest gap between detector cover and detector plate to be at maximum 4mm
1.5.10 | Image acquisition time to display image on the workstation to be less than 4 seconds.
1.5.11 | Imaging Modes must be Conventional Mammography, Tomosynthesis, Combined Conventional and Tomosynthesis under the same compression in CC and MLO positions.
1.5.12 | A Grid reducing scatter radiation drastically must be offered.(State Grid type and advantages)
1.5.13 | A Grid with Auto motorized retraction for Geometric views and Tomosynthesis must be offered
1.5.14 | Linear Response over 400:1 in X-ray exposure or more must be offered.
1.5.15 | Output image bit data must be 14 bit
1.5.16 | X-ray Exposure level at which image pixels are saturated is not less than 500mR
1.5.17 | Failed Pixels on the Detector must be fixed and the Detector Certified as usable after Pixel fixes
1.5.18 | The Detector must be powered from its own individual supply and controllers offering the detector on/off times to be managed individually to prolong the detector’s lifespan by more than 100%
1.6 | GANTRY
1.6.1 | Input Line Voltage to be 200/208/220/230/240VAC Single Phase @50Hz with 10% mains compensation
1.6.2 | A 15VA online UPS to be installed with a Line ripple Filter
1.6.3 | The Gantry must be licensed for Tomosynthesis and must be ready for Tomosynthesis upon Installation. No additional upgrades will be accepted.
1.6.4 | The Tomosynthesis must be used on its own as well as with a Combo Mode, where a conventional mammogram is taken and a Tomo sweep is done under the same compression without manually removing a grid. The grid must be automatically removed and motorized.
1.6.5 | The sweep angle for Tomosynthesis must be 15 degrees to compensate for patient movement and minimizing breast compression time, optimizing image quality and throughput per patient.
1.6.6 | The total Tomo sweeps must be done in less than 4 seconds and the dosage must be similar to that of a Conventional mammogram. State Tomo Dosage @ACR phantom of 4.5cm
1.6.7 | The Tomo sweep exposure shots have to be rapid scans of 15 projections and each exposure taken per degree to prevent overexposure of the Patient.
1.6.8 | Reconstruction of the Projections has to be done in a manner to construct 1mm slices of the breast for the full Compression thickness.
1.6.9 | An upper 5mm of reconstruction slices above the breast Skin line must be available to compensate for any paddle offsets.
1.6.11 | The Tomosynthesis software must be able to Synthesize 2D images from 3D tomosynthesis Studies or able to be upgradable for this: State Price as optional Extra:
1.6.12 | State if the software will be able to be upgraded to perform localisation hookwire and contrast enhanced breast mammography in both 2D and Tomosynthesis.
1.7 | COMPRESSION
1.7.1 | Maximum Manual compression that can be applied must be 300(N)
1.7.2 | 2 pairs of foot pedals for compression adjustment required as well as vertical C-arm height adjustment.
1.7.3 | Adjustable and displayed compressor force with Pre, Full, Dual compression selectable modes must be available. Pre compression mode from 70 to 134N must be available.
Full motorized compression from 89 to 178N must be available with a manual override up to 300N
1.7.4 | Motor driven compression paddles for 18x24, 24x30 and ‘spot’ views to be included.
1.7.5 | The Compression Device must be configurable for Fixed Paddle and Fully Adjustable Self Tilting paddle system conforming to the natural contours of the breast applying equal force from Nipple to Chest Wall.
1.7.6 | Manual and motorised compression movements must be provided
1.7.7 | The following Paddles must be provided:
1.7.7.1 | · 24x29cm Paddle
1.7.7.2 | · 18x24cm Paddle
1.7.7.3 | · Small Breast Paddle
1.7.7.4 | · Frameless Spot Paddle
1.7.7.5 | · 10cm contact Paddle
1.7.7.6 | · 10 cm magnification Paddle
1.7.7.7 | · 5-10 cm spot contact paddle(State Size)
1.7.7.8 | · 5-10cm spot magnification Paddle(State Size)
1.7.7.9 | · Removable Face Shield
1.7.7.10 | · Removable, Retractable Tomosynthesis Face Shield
1.7.7.11 | · Magnification Platform
1.7.7.12 | · 15cm Perforated Localization Paddle
1.7.8 | Automatic Motorized paddle Shift must be possible for the next view optimizing workflow and time
1.7.9 | Compression tilt must be standard as well as fully adjustable and user selectable
1.7.10 | Range of movement of the compression device , relative to the breast support platform must be at least 30cm or more
1.8 | MAGNIFICATION
1.8.1 | Carbon Fibre magnification platform must be provided.
1.8.2 | Area of detector used for magnification must be 18x24 cm
1.8.3 | The size of the magnification compression paddles in cm must be 18x24cm Paddle,10cm x15cm Magnification paddle and 5-10cm spot magnification Paddles.
1.8.4 | Clearance above the magnification device when fitted in cm must be between 18 and 30 cm.
1.8.5 | Compressed breast thickness must be displayed when using magnification on the Gantry and the Acquisition Station.
1.8.6 | AEC function must be activated during magnification.
1.8.7 | Focal spot size used for magnification must be 0.1mm.
1.8.8 | System for magnification views 1.5 and 1.8 times to be included.
1.9 | RADIATION PROTECTION
1.9.1 | The system must have the ability to display Exposure Indicator values (EI/ S-values), immediately after exposure.
1.9.2 | X-ray Warning Notice for display at the X-ray Room entrance (not branded).
1.9.3 | An advanced Retake / Reject / Repeat analysis software package that includes analysis as well as calculations in percentages, must be provided for applicable systems.
1.9.4 | The applicable systems must have an integrated Dose Area Product (DAP) meter that includes skin dose equivalent measurement functions.
1.9.5 | A radiation report must be automatically generated and transferable via DICOM.
1.9.6 | Should be FDA or CE approved product
1.9.7 | Electrical safety conforms to standards for electrical safety IEC-60601
1.9.8 | The unit must comply to IEC-61267
1.9.9 | The equipment/system must be approved and licensed by the Directorate Radiation Control of DoH: South Africa
1.9.10 | A copy of a valid license issued in terms of the Hazardous Substance Act, Act No 15 of 1973 must be submitted with the tender. Failure to submit such a valid license may result in a tender not being considered.
1.9.11 | Acceptance testing must be performed as part of the installation. (Failure to do will lead to failure to pay
1.9.12 | Lead Glass window to be installed as per room requirement
1.9.13 | Safety aspects of Radiation dosage leakage should be spelt out
1.9.14 | DB boards and electrical filters to be installed
1.9.15 | QA TOOLS
1.9.16 | Display Exposure index values
1.9.17 | Radiation protection Lead equivalence of the radiation protection screen must be 0.5m Pb and must be convex, covering the radiographer from various angles against scatter radiation.
1.9.18 | Patient Conventional face shield to be included.
1.9.19 | Patient Tomosynthesis face shield to be included, removable and retractable.
1.9.20 | Emergency stop controls must be supplied on both sides of the gantry and on the Acquisition Workstation
1.9.21 | Interlocks to stop system movement when compression is applied must be provided.
1.9.22 | Interlock for premature release of the X-ray switch must be available
1.9.23 | Interlock preventing an exposure whenever the filter or light field is not positioned correctly must be possible.
1.9.24 | Phantom for calibration should be provided.
1.9.25 | Geometry Phantom for Tomosynthesis calibration must be provided.
1.10. | RADIOGRAPHER’S ACQUISITION STATION
1.10.1 | The Workstation must be Multi Core Intel Based CPU Technology.
1.10.2 | Minimum RAM requirements must be 32 GB
1.10.3 | Data Disk Drive must be 1TB or more
1.10.4 | A 3MP Greyscale Medical Grade LCD Display must be provided on a left or right articulating Arm
1.10.5 | A 1.2MP screen colour LCD Display must be provided as a user interface display which is built in the ergonomic station table, optimizing workflow and throughput. It is preferred that the said screen be of touch type.
1.10.6 | Quick login and access to the application software must be available. State method.
1.10.7 | A built in Automated barcode reader must be provided if required, in liaising with Worklist operations.
1.10.8 | A keyboard, Trackball and/or mouse on the workstation must be provided.
1.10.9 | Ergonomically Designed X-ray Exposure switches must be provided, eliminating repetitive motion discomfort. The system must be configurable for one or two lever activation as well as remote X-ray activation.
1.10.10 | DICOM Services must include Worklist, Print, Storage Commitment, Query/Retrieve and Modality Performed Procedure Step.
1.10.11 | Image storage of up to 9000 conventional images and 3000 Tomosynthesis images or more. State.
1.10.12 | The Acquisition Station must have its own UPS installed on the Station.
1.10.13 | All Patient details must be ordered via completed, In Progress, Pending, Scheduled and All.
1.10.14 | Advanced search facility must be available for searching the patient database.
1.10.15 | Query/Retrieve functionality must be available from PACS to the Acquisition Station.
1.10.16 | DICOM communications from the radiographer’s Acquisition workstation to the Radiologist review Station must be possible.
1.10.17 | A full Quality Control tab must be visible to perform all the required checks as per it’s due time including Automatic SNR/CNR, Artefacts, Image Quality, etc.
1.10.18 | The AEC Exposure modes must include: Auto Time, Auto kV, Auto Filter and Manual. In addition, a Tissue exposure Control mode must be available
1.10.19 | Implant software must be available to reprocess implant views for optimized image quality should the radiographer be unaware of an implant present.
1.10.20 | A full reject and repeat analysis report must be generated per radiographer
1.10.21 | The view orders have to be programmed according to each radiographer’s preferences.
1.10.22 | Output Groups must be setup to maximize the radiographer’s workflow.
1.10.23 | Printing of images must be available to DICOM printers as well as normal Printers
1.10.24 | Importing of Other vendors Images must be possible and must be IHE compliant and importation problem free. Reprocessing of the imported images must be possible.
1.10.25 | The Images must be Grey Scale DICOM Compliant
1.10.26 | The View Order must be configurable and programmable for each Radiographer’s preferences.
1.10.27 | A CDRW/DVD Writer must be available for importing or exporting of images
1.10.28 | A USB 2.0 Port must be available for importing or exporting of images and must be visible on the station outside.
1.10.29 | QA and QC Test patterns including AAPM-TG and SMPTE patterns must be available in the application.
1.10.30 | Only licence keys required to install a stereotactic device must be offered. No additional hardware and software upgrades must be necessary to upgrade to a Stereotactic Biopsy Device.
1.10.31 | All Image optimization and Radiologists preferences settings must be adjusted via software settings
Access to prio(old) and current images also exchange of annotation
1.11 | IMAGING ON ACQUISITION WORKSTATION
1.11.1 | Window width and level with fine adjustment tool must be available. State.
1.11.2 | The following image processing must be possible on the workstation:
1.11.2.1 | Image display:
1.11.2.2 | - Freely selectable screen layout
1.11.2.3 | - Windows settings (contrast and brightness setting)
1.11.2.4 | - Magnification, stepped and dynamic zoom
1.11.2.5 | - Image inversion (black/white)
1.11.2.6 | Ability to cross reference
1.11.3 | Roaming of image
1.11.4 | Image Mirrored
1.11.5 | Edge enhancement
1.11.6 | ROI
1.11.7 | AEC sensor position displayed on image
1.11.7.1 | - SNR Automatic Calculation
1.11.7.2 | - CNR Automatic Calculation as well as Base Value Programmable
1.11.7.3 | - Patient Information as DICOM Header info must be available
1.11.8 | Zoom and Crosshair
1.11.9 | Annotation:
1.11.9.1 | - Draw Oval marking
1.11.9.2 | - Measure a distance
1.11.9.3 | - Draw freehand
1.11.9.4 | - Add a text note
1.11.9.5 | - Send a Notice for the current image to the destination selected
1.11.9.6 | - Send notices for All images in the opened study to the selected destination
1.11.10 | Image evaluation:
1.11.10.1 | - Fit to Viewport
1.11.10.2 | - True Size
1.11.10.3 | - View Actual Pixels
1.11.10.4 | - Multi up Display
1.11.10.5 | - Image tiling Advance
1.11.11 | Administration:
1.11.11.1 | - The demographic patient data should be retrieved directly from a HIS/RIS system
1.11.11.2 | - The demographic patient data can be entered manually or via Worklist.
1.11.11.3 | - Barcode scanner for liaising with Patient Worklist, ensuring the correct Patient is in the Mammo Room
1.11.11.4 | - Retrieval of images from CD, DVD or PACS or USB 2.0 Port
1.11.11.5 | - The workstation should be fully DICOM compatible and include MPPS functionally
1.11.12 | Exposure Index must be visible after the exposure to assess the Index for the Image
1.11.13 | Tomosynthesis Projections must be displayed on the workstation
1.11.14 | Tomosynthesis Construction images must be displayed on the Workstation Display and must be scrollable and with Cine Mode as well.
1.11.15 | The possibility to add a view or a whole procedure must be available on the Workstation.
1.12 | RADIOLOGIST’S REPORTING STATION
1.12.1 | Two High Contrasts 21.3 Inch, 5 Mega Pixel monitors must be provided and must be approved for Mammography.
1.12.2 | Monitor size to be at least 422mm by 337mm
Monitors must have sensors installed with Built in auto calibration and QA self -diagnosis and report writing of QA/QC diagnostics.
1.12.3 | Monitor resolution to be 2048 x 2560
1.12.4 | User interface to be by mouse and keyboard.
1.12.5 | Displayed image bit depth must be 10 bits.
1.12.6 | Removable archive media must be CD/DVD and USB 2.0
1.12.7 | A PC with Data Drive, Application Drive and an Operating System Drive to be offered in RAID1 for optimal protection of Patient Data.
1.12.8 | Storage Space to be a minimum of 4 TB
1.12.9 | PC must be Windows based system with High end Dual Quad Processor @ minimum 32 GB RAM. Latest Windows, 64 Bit Operating System
1.12.10 | A UPS must be provided for the Workstation
1.12.11 | The Reviewing Software must be dedicated mammography Based Image review Software customized for radiologist use, optimizing Workflow and ease of use.
1.12.12 | DICOM services to include Print, Storage, Query /Retrieve/Auto Fetching of Priors from PACS and media export/import
1.12.13 | A Dedicated Mammo Keypad, apart from a normal keyboard must be provided for easy workflow, imaging tools and reviewing of Images via dedicated keys and specified Workflow.
1.12.14 | User preferences for scrolling mode, cine speed, default slab thickness and default initial slice must be available for Tomosynthesis reviewing
1.12.15 | Display initial hanging snapshot for a four-view Tomo study in less than 2 seconds must be possible
1.12.16 | Registration of, and instant toggle between 2D and 3D images
1.12.17 | Automatic scrolling through a limited range of slices must be possible
1.12.18 | Generate Tomosynthesis Movies as .AVI files for presentation and possible to export movies
1.12.19 | Annotate or tag reconstructed slices and store to PACS as a GSPS or DICOM Secondary Capture Image
1.12.20 | Tag reconstructed slices to save as TIFF images for presentation and printing
1.12.21 | Efficient review with personalised report flows for any clinical situation per Radiologist must be provided.
1.12.22 | Single click navigation through each programmed step must be possible
1.12.23 | Intelligent roaming through images must be possible and be programmable for each Radiologists Preferences
1.12.24 | Automatic step-through full image resolution of breast quadrants
1.12.25 | Automatic retrieval (Auto Fetching) of prior studies from PACS upon receipt of current images must be offered. State
1.12.26 | Manually or Automatically merge patient cases if patient presents with different ID and date of Birth combination must be offered. State
1.12.27 | Receive and analyze images from remote sites and mobile units through an existing PACS or SecureXchange product must be possible.
1.12.28 | Display of standard mammography images for all FFDM vendors as well as CR mammography, MRi mammography and Ultrasound MG must be possible. State
1.12.29 | Correct comparisons of current and priors images
1.12.30 | Clarified display of CAD marks and information on 2D must be possible
1.12.31 | Consistent Patient, study and image information labelling to be displayed
1.12.32 | Digital Communication channel between the radiologist and the Radiographer, independent of PACS must be possible.
1.12.33 | Exchange markings, annotations, comments, queries and instruction must be possible and must be able to save to PACS.
2 | COMPUTER AIDED DIAGNOSTICS PC AND SOFTWARE
2.1 | CAD PC with 2.6GHz Processor or more. 32 GB RAM or More, HDD space of 4 TB or more must be offered
2.2 | CD RW/DVD Drive with 2 Ethernet ports to be offered
2.3 | CAD must support up to 4 licensed digital ports
2.4 | CAD Software using sophisticated algorithms, searches digital mammograms for potential micro calcifications and masses- characteristics commonly associated with breast cancer to be offered
2.5 | The Algorithms must also incorporate anatomic correlation technology where this technology correspond findings in the contra lateral breast and different views of the breast.
2.6 | CAD must support outputs to PACS or DICOM Workstations
2.7 | Throughput of CAD cases must be 40-60 per hour/nominal
2.8 | CAD Display able to be archived to PACS
2.9 | Additional CAD specifics with Calcifications and Masses.
2.10. | CAD specificity and sensitivity to be programmed as per Radiologist’s preferences.
CAD overlay markers to be visible on Images when requested.
3 | UPRIGHT STEREO-TAXY BIOPSY SYSTEM
3.1 | BIOPSY SYSTEM
3.1.1 | Weight of the Biopsy system shall be less than 7kg
3.1.2 | Ability to do stereo Biopsy in any angle of the c-arm rotation
3.1.3 | The Biopsy window should be a minimum of 6 x 7cm.
3.1.4 | Auxiliary paddle should be included to offer a solution to biopsy in the auxiliary region
3.1.5 | Stereo pair images shall be acquired automatically without requiring the technician to move the x-ray tube.
3.1.6 | Large SID, equal to or larger then 68cm, to allow for easy installation of the biopsy system.
3.1.7 | Detachable control unit
3.1.8 | Screen color display on the control unit must be provided. It is preferred that this should be of the touch type.
3.1.9 | Biopsy add-on shall support Tomo guided biopsy procedures.
3.1.10 | Transparent Biopsy compression paddles (5cm x 5cm axillary biopsy paddle, 5cm x 5cm standard biopsy paddle and a 6cm x 7cm standard biopsy paddle).
3.2 | OPERATOR CONSOLE
3.2.1 | Biopsy user interface and targeting shall be integrated on the mammography system acquisition workstation.
3.2.2 | Target should be possible on both stereo, or on one stereo and one scout.
3.2.3 | Ability to acquire a scout with 18x24cm collimation, to be able to view tissue around biopsy window.
3.2.4 | Possibility to target up to 6 targets.
3.2.5 | Automatic placement of target on the third image after targeting on two images.
3.2.6 | Possibility to program different biopsy devices and needles.
3.2.7 | Automatic calculation of safety margins versus the detector and the paddle according to the needle used.
3.2.8 | DICOM 3.0 Print/Storage/Storage commitment and worklist approved.
3.2.9 | X-Ray generation and detector control functions integrated in a single operator console and controlled through the graphical user interface.
3.2.10 | Control device capable of tracking actual number of exposures by display icon versions of acquired images.
3.2.11 | Grayscale monitor for previewing images, 3MP minimum.
3.2.12 | Preview and image info (KV, mas, dose, compression thickness, force) Available after every exposure.
3.2.13 | Ability to add comments to an image
3.2.14 | Ability to edit views.
4 | LATERAL ARM
4.1 | 15 cm lateral biopsy compression paddle and a support paddle for lateral arm
4.2 | Should come in a case holder with needle guides and support
5 | VACUUM ASSISTED BREAST BIOPSY AND EXCISION SYSTEM MUST BE OFFERED AS AN OPTION.
5.1 | To include the following:
5.1.1 | Breast Biopsy and Excision system with
5.1.2 | Footswitch for use in MRI as well
5.1.3 | Power Cord
5.1.4 | Vacuum line and assembly with Line Filtering
5.1.5 | Stereotactic Adapter for use on system offered
5.2 | Must be compatible with the stereotactic system offered and utilized for handheld MRI.
5.3 | Must be compatible with Ultrasound use
5.4 | Using needled lengths of 9, 12 and 14 cm.
Needle gauge of 9 and 12 gauge.
5.5 | Vacuum generated: 71 cm Hg at sea level
5.6 | Biopsy and Excision system must be Mobile on rugged, Sturdy Base for use with other Modalities.
6 | MAMMOGRAPHIC BIOPSY CHAIR
6.1 | Mammography chair with low, removable arm rest and adjustable height, allowing easy access to the unit on offer, must be included.
6.2 | Must include a Foot pedal controlled hydraulic pump
6.3 | Four caster lock system to lock the wheels.
6.4 | 300lb Maximum patient weight
6.5 | Upright Seated position: 25” W x 38.5”H
6.6 | Fully reclined: 25” W x 73”L
6.7 | Overall Height: 29” raises to 36.75”
6.8 | Seat Back: 32.5” above the seat
6.9 | Wheels: 6” with locks.
7 | POWER SUPPLY
7.1 | The bidder must install adequate electrical power supply for the optimal functionality of the equipment.
7.2 | The mains cable of the unit being quoted for must be the hospital grade type and it must be of an adequate length (minimum length of three (3) meters).
7.3 | The power input must be 220-240V, 50Hz AC. Bidders must ensure that the product quoted for is fitted appropriately with a 15 Amp SABS approved mains plug that is held together by two screws.
7.4 | A 3 phase in / out power supply must be offered. State details.
7.5 | The equipment tendered for must not overload and trip the hospital power in the area where it is installed. Resettable overcurrent breaker/s must be fitted for protection. State details.
7.6 | All available online Uninterrupted Power Supply (UPS system) and line voltage regulators must be included for the system with a support time of 50 to 60 minutes. Attach brochure.
7.7 | The lifetime of UPS must be at least 15 years.
7.8 | An Uninterrupted Power Supply (UPS system) conforming to IS-302 must be offered for the computer system.
7.9 | The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade type and it must be a minimum length of three (3) meters.
7.10. | The mains cable of the unit tendered for must be SABS colour coded. State details.
7.11 | The equipment quoted for must be protected against electromagnetic interference. State details of measures that will be taken to protect the unit against electromagnetic interference.
7.12 | Maintenance free batteries must be offered.
8 | AIR CONDITIONER
8.1 | Bidder must include any air conditioning that is considered necessary for optimal functioning of the unit.
8.2 | Any such air conditioning must be covered by the guarantee bided for on the unit and servicing of the air conditioning units for the guarantee period and must be included in the bid price.
8.3 | The air conditioner must also be included in the maintenance plan.
8.4 | Successful bidder to replace existing air conditioner with an air conditioner of 24,000 BTU. (Where applicable).
8.5 | The air conditioner must be but no less than 24,000 BTU. State details on the air conditioner to be supplied according to the room size and machine requirements.
8.6 | The Successful bidder must install an outlet-plumbing pipe for water overflow from air conditioner to be re-directed to a drainage system.
8.7 | Air-conditioning must be included in Maintenance plan

## Bone densitometer
BONE DENSITOMETER, FAN BEAM AND LINEAR SCANNING
ITEM | TECHNICAL SPECIFICATIONS
1 | Modality Details
1.1 | Functionality:
1.2 | The system must be a Linear Scanning device.
1.3 | 2D Fan beam Technology
1.4 | Motorized Table Top and Arm.
1.5 | Laser light for patient positioning
1.6 | Maximum patient weight of 200kg or above.
1.7 | Table lowest height range between 60cm - 72cm.
1.8 | A complete scan / exposure time for AP spine & femur and hip study should be completed within 40 seconds.
1.9 | Multi-element, high resolution, digital detector array. The details of the detector must be stated.
1.10. | The maximum scanning area must be a maximum range of between 60-75 cm in width.
1.11 | The maximum scanning area must be a maximum range of between of 190-210 cm in length.
1.12 | Dual Energy X-Ray Absorptiometry (DEXA) Dual pulse energy.
1.13 | The system must be able to generate a maximum kVp of 100 or above.
1.14 | Both Automatic and Manual selection of Region of Interest must be possible.
1.15 | Automatic and continouous calibration. Operator calibration not required.
1.16 | Automatic quality control program, including required phantoms.
1.17 | In addition to its normal operation, the system must be suited to the examination of paediatric patients.
1.18 | OPERATOR CONSOLE
1.18.1 | Must come with a PC with the following requirements:
1.18.1.1 | Windows 10 or newer operating system.
1.18.1.2 | i7 or better processor.
1.18.1.3 | 32 GB RAM
1.18.1.4 | 4 TB hard drive.
1.18.1.5 | 17 Inch or larger 3 MP Flat Panel Touch Monitor.
1.18.1.6 | Keyboard and mouse
1.18.1.7 | CD/DVD R/W drive enabled (it should be license activated and ready for CD/DVD writing)
1.18.1.8 | Network ready
1.18.1.9 | DICOM compliant and enabled (HIS, RIS,PACS).
2 | Software tools and functions
2.1 | The following Software Packages must be included in the offer:
2.1.1 | Bone Mineral Density (BMD), including hip and Lateral Spine BMD.
2.1.2 | Bone Mineral Content (BMC)
2.1.3 | Body Composition Analysis
2.1.4 | BMI Calculator
2.1.5 | Scan / analyze protocols for whole body, AP and lateral spine, proximal femur, dual hip, forearm, scoliotic spine for paediatric and adult patients.
2.1.6 | Digital vertebra assessment (IVA).
2.1.7 | The system must have reporting packages.
2.1.8 | The reporting packages must be DICOM compatible.
2.1.9 | Customizable reports.
2.1.10 | The system must support multiple reports per patients, including comparison against earlier studies.
2.1.11 | The reporting software must be compliant with the International Society for Clinical Densitometry (ISCD).
2.1.12 | Integrated FRAX report to be included (fracture risk assessment).
2.1.13 | The system must include a compatible Colour Printer.
4 | Operating Requirements
4.1 | Temperature: 15 - 32°Celsius
4.2 | Humidity: 20 - 80%
4.3 | Power: 220 - 240 V, 50 Hertz
5 | IT Integration
5.1 | The system must have the following DICOM compatibility:
5.1.1 | DICOM print class
5.1.2 | DICOM storage class
5.1.3 | DICOM query / retrieve
5.1.4 | DICOM work list
6 | ACCESSORIES
6.1 | forearm positioner with palm mouse positioner
6.2 | head positioner
6.3 | hip positioner fixture with a velco strap
6.4 | spine phantom
6.5 | step phantom for vertebral assessment
6.6 | positioning sponges
7 | POWER SUPPLY
7.1 | The bidder must install adequate electrical power supply for the optimal functionality of the equipment.
7.2 | The mains cable of the unit being quoted for must be the hospital grade type and it must be of an adequate length (minimum length of three (3) meters).
7.3 | The power input must be 220-240V, 50Hz AC. Bidders must ensure that the product quoted for is fitted appropriately with a 15 Amp SABS approved mains plug that is held together by two screws.
7.4 | A 3 phase in / out power supply must be offered. State details.
7.5 | The equipment tendered for must not overload and trip the hospital power in the area where it is installed. Resettable overcurrent breaker/s must be fitted for protection. State details.
7.6 | All available online Uninterrupted Power Supply (UPS system) and line voltage regulators must be included for the system with a support time of 50 to 60 minutes. Attach brochure.
7.7 | The lifetime of UPS must be at least 15 years.
7.8 | An Uninterrupted Power Supply (UPS system) conforming to IS-302 must be offered for the computer system.
7.9 | The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade type and it must be a minimum length of three (3) meters.
7.10. | The mains cable of the unit tendered for must be SABS colour coded. State details.
7.11 | The equipment quoted for must be protected against electromagnetic interference. State details of measures that will be taken to protect the unit against electromagnetic interference.
7.12 | Maintenance free batteries must be offered.
8 | RADIATION PROTECTION
8.1 | The system must have the ability to display Exposure Indicator values (EI/ S-values), immediately after exposure.
8.2 | X-ray Warning Notice for display at the X-ray Room entrance (not branded).
8.3 | An advanced Retake / Reject / Repeat analysis software package that includes analysis as well as calculations in percentages, must be provided for applicable systems.
8.4 | The applicable systems must have an integrated Dose Area Product (DAP) meter that includes skin dose equivalent measurement functions.
8.5 | A radiation report must be automatically generated and transferable via DICOM.
8.6 | Should be FDA or CE approved product
8.7 | Electrical safety conforms to standards for electrical safety IEC-60601
8.8 | The unit must comply to IEC-61267
8.9 | The equipment/system must be approved and licensed by the Directorate Radiation Control of DoH: South Africa
8.10. | A copy of a valid license issued in terms of the Hazardous Substance Act, Act No 15 of 1973 must be submitted with the tender. Failure to submit such a valid license may result in a tender not being considered.
8.11 | Acceptance testing must be performed as part of the installation. (Failure to do will lead to failure to pay
8.12 | Lead Glass window to be installed as per room requirement
8.13 | Safety aspects of Radiation dosage leakage should be spelt out
8.14 | DB boards and electrical filters to be installed
8.15 | QA TOOLS
8.16 | Display Exposure index values
8.17 | Radiation protection Lead equivalence of the radiation protection screen must be 0.5m Pb and must be convex, covering the radiographer from various angles against scatter radiation.
8.18 | Patient Conventional face shield to be included.
8.19 | Patient Tomosynthesis face shield to be included, removable and retractable.
8.20. | Emergency stop controls must be supplied on both sides of the gantry and on the Acquisition Workstation
8.21 | Interlocks to stop system movement when compression is applied must be provided.
8.22 | Interlock for premature release of the X-ray switch must be available
8.23 | Interlock preventing an exposure whenever the filter or light field is not positioned correctly must be possible.
8.24 | Phantom for calibration should be provided.
9 | AIR CONDITIONER
9.1 | Bidder must include any air conditioning that is considered necessary for optimal functioning of the unit.
9.2 | Any such air conditioning must be covered by the guarantee bided for on the unit and servicing of the air conditioning units for the guarantee period and must be included in the bid price.
9.3 | The air conditioner must also be included in the maintenance plan.
9.4 | Successful bidder to replace existing air conditioner with an air conditioner of 24,000 BTU. (Where applicable).
9.5 | The air conditioner must be but no less than 24,000 BTU. State details on the air conditioner to be supplied according to the room size and machine requirements.
9.6 | The Successful bidder must install an outlet-plumbing pipe for water overflow from air conditioner to be re-directed to a drainage system.
9.7 | Air-conditioning must be included in Maintenance plan
8 | Optional Items
8.1 | Orthopaedics software for prosthesis.
8.2 | Orthopaedic software
8.3 | Single energy femur scan software- to detect atypical femur fractures

## Fluoroscopic system
FLOUROSCOPIC SYTEM
ITEM | SPECIFICATIONS
1 | REMOTE CONTROLLED TABLE.
1.1 | A remote controlled table with easy patient transfer access. Please state confirmation and describe table top mounting.
1.2 | The table top must be able to accommodate a patient weight of at least 200 kg or above.
1.3 | The table system must be able to tilt to +90 degree's and to -15 degree's (minimum range)
1.4 | The table top to floor height should be stated. Low lower limit is preferred.
1.5 | Table top must be X-ray translucent
1.6 | Table top length approximately 220cm x 55cm
1.7 | The table system must have a motorised compression cone device. Please state compliance.
1.8 | The table must be able to perform tomography. Please state the angles offered and the speeds related to each angle.
1.9 | The tomographic device should have a variable layer height. State range.
1.10. | The table must be supplied with the following:
1.10.1 | Foot rest
1.10.2 | Hand grips
1.10.3 | Shoulder rests
1.10.4 | Band compression device (for IVP)
1.10.5 | Mattress with water resistant covers
1.10.6 | Compensation filter pad for peripheral studies
2 | X-RAY TUBE AND COLLIMATOR SYSTEM
2.1 | The tube and collimator must be able to vary in FFD (state range).
3 | HIGH FREQUANCY GENERATOR SYSTEM
3.1 | A 65 kWatt high frequency, microprocessor controlled x-ray generator is to be supplied together with the table system
3.2 | A matching x-ray tube unit with small and large focus distance must be provided.
3.3 | The x-ray generator must be delivered with a set of matching high tension cables.
3.4 | The x-ray generator and table system must have AEC built into the systems
3.5 | A full range of standard organ programs must be supplied. State number of programs.
3.6 | The x-ray generator must have an output ability of at least 80 kV and 800 mA.
3.7 | The x-ray generator and x-ray tube unit must have a rapid starter unit ensuring anode rotation of at least 8000 RPM.
3.8 | The system must be delivered with a dose measuring chamber to meet the new national radiation protection requirements. Please state confirmation of delivery and fitment. DAP
3.9 | The table should be of a smooth and easy to clean construction.
3.10. | An automatic collimator must be provided with full control from the table side control and the control room console.
3.11 | An automatic collimator light must be included for off radiation positioning of the patient.
3.12 | State radiation reduction techniques.
4 | CCD TELEVISION CAMERA SYSTEM WITH LARGE SIZE IMAGE INTENSIFIER AND TELEVISION MONITORS.
4.1 | A CCD television camera system must be supplied with this offer and must have a resolution of at least 1024 x 1024.
4.2 | A high resolution image intensifier of a at least an effective 38 cm input diameter must be supplied with the system.
4.3 | The image intensifier must have at least three zoom/magnification sizes. Please state sizes and compliance.
4.4 | The unit must be delivered with two flat screen TFT or LCD type monitors of at least 48 cm diagonal size. One in the control room and one in the examination room.
5 | DIGITAL IMAGING SYSTEM
5.1 | A full digital imaging system is to be offered with the system and must be totally integrated into the main system.
5.2 | Please indicated if the system can accommodate a full field flat panel detector system at a later stage. Please state cost of such upgrade in your offer
5.3 | The digital system must have digital pulsed fluoroscopy as standard. Please state the frame rates offered
5.4 | Digital radiography: Minimum of 2048 x 2048 (2 kilo Matrix), 12 bit, 0.5 to 7.5 frames/sec
5.5 | The following applications should form part of the digital imaging system (1K x 1K x 12 bits):
5.5.1 | Continuous digital fluoroscopy
5.5.2 | Pulsed fluoroscopy
5.5.3 | Digital radiography
5.5.4 | Overview of images from 4 to 16 on one monitor
5.5.5 | Digital zoom with magnification of at least 2X
5.5.6 | Windowing and contrast adjustment
5.5.7 | Image flip to the left and right as well as up and down
5.5.8 | Image rotation
5.5.9 | Digital edge enhancement
5.5.10 | Digital filtration
5.5.11 | Cine loop after a scene
5.5.12 | Electronic collimation
5.5.13 | Text annotation and free word insertion
5.5.14 | Graphic calculation of angles and relative distances
5.6 | The system storage capaciity should be at least 15 000images of 1K x 1K x 12 bits
6 | GENERAL ISSUES
SAFETY STANDARDS OF THE SYSTEM
6.1 | X-ray machine should be CE marked
6.2 | X-ray Compliance to IEC 601
7 | IT INTEGRATION
7.1 | The system must have the following DICOM compatibility:
7.1.1 | DICOM print class
7.1.2 | DICOM storage class
7.1.3 | DICOM query/retrieve
7.1.4 | DICOM worklist
7.1.5 | The system must comply to the following IHE profiles (Integrating the Healthcare Enterprise)
7.1.6 | Scheduled Workflow (SWF)
7.1.7 | Patient Information Reconciliation (PIR)
7.1.8 | Simple image and Numeric Reports (SINR)
8 | RADIATION PROTECTION
8.1 | The system must have the ability to display Exposure Indicator values (EI/ S-values), immediately after exposure.
8.2 | X-ray Warning Notice for display at the X-ray Room entrance (not branded).
8.3 | An advanced Retake / Reject / Repeat analysis software package that includes analysis as well as calculations in percentages, must be provided for applicable systems.
8.4 | The applicable systems must have an integrated Dose Area Product (DAP) meter that includes skin dose equivalent measurement functions.
8.5 | A radiation report must be automatically generated and transferable via DICOM.
8.6 | Should be FDA or CE approved product
8.7 | Electrical safety conforms to standards for electrical safety IEC-60601
8.8 | The unit must comply to IEC-61267
8.9 | The equipment/system must be approved and licensed by the Directorate Radiation Control of DoH: South Africa
8.10. | A copy of a valid license issued in terms of the Hazardous Substance Act, Act No 15 of 1973 must be submitted with the tender. Failure to submit such a valid license may result in a tender not being considered.
8.11 | Acceptance testing must be performed as part of the installation. (Failure to do will lead to failure to pay
8.12 | Lead Glass window to be installed as per room requirement
8.13 | Safety aspects of Radiation dosage leakage should be spelt out
8.14 | DB boards and electrical filters to be installed
8.15 | QA TOOLS
8.16 | Display Exposure index values
8.17 | Radiation protection Lead equivalence of the radiation protection screen must be 0.5m Pb and must be convex, covering the radiographer from various angles against scatter radiation.
8.18 | Patient Conventional face shield to be included.
8.19 | Patient Tomosynthesis face shield to be included, removable and retractable.
8.20. | Emergency stop controls must be supplied on both sides of the gantry and on the Acquisition Workstation
8.21 | Interlocks to stop system movement when compression is applied must be provided.
8.22 | Interlock for premature release of the X-ray switch must be available
8.23 | Interlock preventing an exposure whenever the filter or light field is not positioned correctly must be possible.
8.24 | Phantom for calibration should be provided.
9 | POWER SUPPLY
9.1 | The bidder must install adequate electrical power supply for the optimal functionality of the equipment.
9.2 | The mains cable of the unit being quoted for must be the hospital grade type and it must be of an adequate length (minimum length of three (3) meters).
9.3 | The power input must be 220-240V, 50Hz AC. Bidders must ensure that the product quoted for is fitted appropriately with a 15 Amp SABS approved mains plug that is held together by two screws.
9.4 | A 3 phase in / out power supply must be offered. State details.
9.5 | The equipment tendered for must not overload and trip the hospital power in the area where it is installed. Resettable overcurrent breaker/s must be fitted for protection. State details.
9.6 | All available online Uninterrupted Power Supply (UPS system) and line voltage regulators must be included for the system with a support time of 50 to 60 minutes. Attach brochure.
9.7 | The lifetime of UPS must be at least 15 years.
9.8 | An Uninterrupted Power Supply (UPS system) conforming to IS-302 must be offered for the computer system.
9.9 | The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade type and it must be a minimum length of three (3) meters.
9.10. | The mains cable of the unit tendered for must be SABS colour coded. State details.
9.11 | The equipment quoted for must be protected against electromagnetic interference. State details of measures that will be taken to protect the unit against electromagnetic interference.
9.12 | Maintenance free batteries must be offered.
10 | AIR CONDITIONER
10.1 | Bidder must include any air conditioning that is considered necessary for optimal functioning of the unit.
10.2 | Any such air conditioning must be covered by the guarantee bided for on the unit and servicing of the air conditioning units for the guarantee period and must be included in the bid price.
10.3 | The air conditioner must also be included in the maintenance plan.
10.4 | Successful bidder to replace existing air conditioner with an air conditioner of 24,000 BTU. (Where applicable).
10.5 | The air conditioner must be but no less than 24,000 BTU. State details on the air conditioner to be supplied according to the room size and machine requirements.
10.6 | The Successful bidder must install an outlet-plumbing pipe for water overflow from air conditioner to be re-directed to a drainage system.
10.7 | Air-conditioning must be included in Maintenance plan

## Digital Flat PanelFluoro system
MULTI FUNCTIONAL DIGITAL FLAT PANEL FLUOROSCOPIC SYSTEM
ITEM | TECHNICAL SPECIFICATIONS | Weight
1. | X-Ray Generator
1.1 | Micro-processor controlled high frequency generator | #
1.2 | Nominal rating: > 80kW | #
1.3 | Shortest exposure time: 1msec | #
1.4 | Maximum tube current: ± 1000 mA | #
1.5 | Fluoroscopic programmed radiography must be possible | #
1.6 | Auto-setting of all parameters of at least 10 different examinations should be possible | #
1.7 | Anatomical programming must be included. State # of programmes | #
1.8 | Pulsed fluoroscopy for dose reduction should be available
1.9 | Automatic mains compensation | #
1.10 | Continuous self-check of all generator functions
1.11 | An integrated operator's console is required. | #
1.12 | Can the measuring fields be graphically displayed over the LIH image?
2. | X-Ray Tube:
2.1 | A high speed X-Ray tube is essential. State speed | #
2.2 | Dual focus with ratings compatible with the offered generator | #
State focal sizes.
2.2.1 | Large focus to be ± 1.0 mm | #
2.2.2 | Small focus to be ± 0.6 mm | #
2.3 | Anode material to be tungsten rhenium molybdenum compound with a maximum heat storage capacity of at least 1 500 kHU
2.4 | State anode heat dissipation HU/min | #
2.5 | Short term rating to be at least 45 kW and 85 kW | #
2.6 | Set of high tension cables to suit the installation | #
2.7 | The table must be able to perform tomography. Please state the angles offered and the speeds related to each angle.
3 | Integrated Table and C-Arm
3.1 | Remote controlled table/detector with integrated C-arm | #
3.1.1 | Table angulation: + 90˚ and -45˚ | #
3.2 | Table top must be X-ray translucent | #
3.3 | Table top length approximately 220cm x 55cm | #
3.4 | Range of table top/detector movements: | #
3.4.1 | Transverse at least 36 cm | #
3.4.2 | Longitudinal patient coverage of approx.195cm should be possible | #
3.5 | SID should be variable to position the detector close to the patient without tabletop or C-arm elevation | #
3.6 | Table height at lowest level should be ± 70 cm | #
3.7 | AP and PA imaging must be possible | #
3.8 | Programming for at least 50 preset positions in both PA and AP must be available | #
3.9 | C-arm angulation of approx. 35° cranial and 35°caudal | #
3.10 | C-arm angulation speed variable, at least 15 ˚/sec | #
3.11 | At least 90° RAO and 45° LAO must be possible | #
3.12 | An automatic collimator system | #
3.13 | The remote control panel must be fully integrated and include controls for generator, image acquisition and tabletop/C-arm | #
3.14 | All relevant generator and system information must be presented as clearly legible displays | #
3.15 | Controls for table top/C-arm movement should be available at the table side and the remote panel | #
3.16 | Collision sensors for patient safety must be incorporated within the system | #
3.17 | Maximum weight limit for the table should be at least 165kg | #
3.18 | The following table accessories are to be included:- | #
3.18.1 | Shoulder rests
3.18.2 | Mattress
3.18.3 | Hand grips
3.18.4 | Foot rest
3.18.5 | Compression cone
3.18.6 | Compensation filter pad for peripheral studies
4 | Digital Flat Panel Detectors
4.1 | A large field digital flat panel with a maximum input field of 43 x 43 cm is required. | #
4.1.1 | Direct conversion method will be an advantage. Tenderer to state and motivate what technology/ method is used | #
4.2 | Detector matrix must be at least 2048 x 2048 or greater, 16bit. State and motivate | #
4.3 | Processor matrix must be at least 2048 x 2048 or greater, 16 bit. State and motivate | #
4.4 | State acquisition field sizes in sq.cm | #
4.5 | Resolution must not be less than 2.5 lp/mm. State | #
4.6 | State and motivate pixel size. Should not be more than 180µm | #
4.7 | State DQE – Acquisition dose (%). Preference will be given to the lowest dose system | #
4.8 | State DQE – Fluoroscopy dose (%). Preference will be given to the lowest dose system | #
5 | T.V. Monitors
5.1 | Two high resolution TFT LCD monitors of at least 48 cm to be ceiling mounted in the examination room | #
5.2 | One at least 48 cm high line-rate TFT LCD monitor to be supplied at the control desk | #
5.3 | Automatic brightness control to be included | #
5.4 | State monitor output brightness: (cd/m2)
5.4.1 | Monitors in exam room | #
5.4.2 | Monitor at control panel | #
6 | Digital Image Acquisition and Processing System
6.1 | A digital imaging system suitable for general fluorography | #
6.2 | Digital acquisition matrix:
6.2.1 | Digital radiography:Minimum of 2048 X 2048 (2 Kilo Matrix), 12 bit, 0.5 to 7.5 frames/sec | #
6.2.2 | Digital fluoroscopy should be possible at 1024 X 1024, 30 FPS | #
6.3 | At least 3 pulsed fluoroscopy modes must be available | #
6.4 | Single shot exposures must be possible | #
6.5 | Last image hold function and instant transfer of fluoroscopy image, both single image and dynamic, to digital storage on a hard drive with at least 1 TB capacity. State capacity | #
6.6 | Storage capacity of 20 000 frames in 2048 X 2048 matrix for immediate access must be available. State capacity | #
6.7 | Separate storage for at least 3000 serial and 600 spot images in 2048 X 2048 matrix on CD-R should be possible. State | #
6.8 | DICOM conformance statement to be included DICOM Print must be possible | #
6.9 | Backup recording on DVD must be possible | #
6.10 | Connection for sending images to a laser imager should utilize a digital or DICOM printing environment | #
6.11 | Functionality must also include:-
6.11.1 | Measuring of distance, angle and area | #
6.11.2 | Gray scale inversion | #
6.11.3 | Window and centre control | #
6.11.4 | Text annotation | #
6.11.5 | Up to X2 or better zoom facility required complete with scroll, zoom and cine loop display | #
6.11.6 | Horizontal (R/L) and vertical (up/down) image flip | #
6.11.7 | Gamma curve selection | #
6.11.8 | 16 & 4 on 1 multi-frame display with study/series overview | #
6.11.9 | Split display for reference image comparison on 1 monitor | #
7 | IT INTEGRATION
7.1 | The system must have the following DICOM compatibility:
7.1.1 | DICOM print class
7.1.2 | DICOM storage class
7.1.3 | DICOM query/retrieve
7.1.4 | DICOM worklist
7.1.5 | The system must comply to the following IHE profiles (Integrating the Healthcare Enterprise)
7.1.6 | Scheduled Workflow (SWF)
7.1.7 | Patient Information Reconciliation (PIR)
8 | RADIATION PROTECTION
8.1 | The system must have the ability to display Exposure Indicator values (EI/ S-values), immediately after exposure.
8.2 | X-ray Warning Notice for display at the X-ray Room entrance (not branded). | #
8.3 | An advanced Retake / Reject / Repeat analysis software package that includes analysis as well as calculations in percentages, must be provided for applicable systems. | #
8.4 | The applicable systems must have an integrated Dose Area Product (DAP) meter that includes skin dose equivalent measurement functions. | #
8.5 | A radiation report must be automatically generated and transferable via DICOM. | #
8.6 | Should be FDA or CE approved product | #
8.7 | Electrical safety conforms to standards for electrical safety IEC-60601
8.8 | The unit must comply to IEC-61267 | #
8.9 | The equipment/system must be approved and licensed by the Directorate Radiation Control of DoH: South Africa | #
8.10. | A copy of a valid license issued in terms of the Hazardous Substance Act, Act No 15 of 1973 must be submitted with the tender. Failure to submit such a valid license may result in a tender not being considered.
8.11 | Acceptance testing must be performed as part of the installation. (Failure to do will lead to failure to pay
8.12 | Lead Glass window to be installed as per room requirement
8.13 | Safety aspects of Radiation dosage leakage should be spelt out
8.14 | DB boards and electrical filters to be installed
8.15 | QA TOOLS
8.16 | Display Exposure index values
8.17 | Radiation protection Lead equivalence of the radiation protection screen must be 0.5m Pb and must be convex, covering the radiographer from various angles against scatter radiation.
8.18 | Patient Conventional face shield to be included.
8.19 | Patient Tomosynthesis face shield to be included, removable and retractable.
8.20. | Emergency stop controls must be supplied on both sides of the gantry and on the Acquisition Workstation
8.21 | Interlocks to stop system movement when compression is applied must be provided.
8.22 | Interlock for premature release of the X-ray switch must be available
8.23 | Interlock preventing an exposure whenever the filter or light field is not positioned correctly must be possible.
8.24 | Phantom for calibration should be provided.
9 | POWER SUPPLY
9.1 | The bidder must install adequate electrical power supply for the optimal functionality of the equipment.
9.2 | The mains cable of the unit being quoted for must be the hospital grade type and it must be of an adequate length (minimum length of three (3) meters).
9.3 | The power input must be 220-240V, 50Hz AC. Bidders must ensure that the product quoted for is fitted appropriately with a 15 Amp SABS approved mains plug that is held together by two screws.
9.4 | A 3 phase in / out power supply must be offered. State details.
9.5 | The equipment tendered for must not overload and trip the hospital power in the area where it is installed. Resettable overcurrent breaker/s must be fitted for protection. State details.
9.6 | All available online Uninterrupted Power Supply (UPS system) and line voltage regulators must be included for the system with a support time of 50 to 60 minutes. Attach brochure.
9.7 | The lifetime of UPS must be at least 15 years.
9.8 | An Uninterrupted Power Supply (UPS system) conforming to IS-302 must be offered for the computer system.
9.9 | The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade type and it must be a minimum length of three (3) meters.
9.10. | The mains cable of the unit tendered for must be SABS colour coded. State details.
9.11 | The equipment quoted for must be protected against electromagnetic interference. State details of measures that will be taken to protect the unit against electromagnetic interference.
9.12 | Maintenance free batteries must be offered.
10 | AIR CONDITIONER
10.1 | Bidder must include any air conditioning that is considered necessary for optimal functioning of the unit.
10.2 | Any such air conditioning must be covered by the guarantee bided for on the unit and servicing of the air conditioning units for the guarantee period and must be included in the bid price.
10.3 | The air conditioner must also be included in the maintenance plan.
10.4 | Successful bidder to replace existing air conditioner with an air conditioner of 24,000 BTU. (Where applicable).
10.5 | The air conditioner must be but no less than 24,000 BTU. State details on the air conditioner to be supplied according to the room size and machine requirements.
10.6 | The Successful bidder must install an outlet-plumbing pipe for water overflow from air conditioner to be re-directed to a drainage system.
10.7 | Air-conditioning must be included in Maintenance plan

## Fluoroscopic, floor mounted
FLUOROSCOPIC SYSTEM
ITEM | TECHNICAL SPECIFICATIONS
1. | X-Ray Generator
1.1 | Micro-processor controlled high frequency generator | #
1.2 | Nominal rating: > 80kW | #
1.3 | Shortest exposure time: 1msec | #
1.4 | Maximum tube current: ± 1000 mA | #
1.5 | Fluoroscopic programmed radiography must be possible | #
1.6 | Auto-setting of all parameters of at least 10 different examinations should be possible | #
1.7 | Anatomical programming must be included. State # of programmes | #
1.8 | Pulsed fluoroscopy for dose reduction should be available
1.9 | Automatic mains compensation | #
1.10 | Continuous self-check of all generator functions
1.11 | An integrated operator's console is required. | #
1.12 | Can the measuring fields be graphically displayed over the LIH image?
2. | X-Ray Tube:
2.1 | A high speed X-Ray tube is essential. State speed | #
2.2 | Dual focus with ratings compatible with the offered generator | #
State focal sizes.
2.2.1 | Large focus to be ± 1.0 mm | #
2.2.2 | Small focus to be ± 0.6 mm | #
2.3 | Anode material to be tungsten rhenium molybdenum compound with a maximum heat storage capacity of at least 1 500 kHU
2.4 | State anode heat dissipation HU/min | #
2.5 | Short term rating to be at least 45 kW and 85 kW | #
2.6 | Set of high tension cables to suit the installation | #
2.7 | The table must be able to perform tomography. Please state the angles offered and the speeds related to each angle.
3 | Integrated Table and C-Arm
3.1 | Remote controlled table/detector with integrated C-arm | #
3.1.1 | Table angulation: + 90˚ and -45˚ | #
3.2 | Table top must be X-ray translucent | #
3.3 | Table top length approximately 220cm x 55cm | #
3.4 | Range of table top/detector movements: | #
3.4.1 | Transverse at least 36 cm | #
3.4.2 | Longitudinal patient coverage of approx.195cm should be possible | #
3.5 | SID should be variable to position the detector close to the patient without tabletop or C-arm elevation | #
3.6 | Table height at lowest level should be ± 70 cm | #
3.7 | AP and PA imaging must be possible | #
3.8 | Programming for at least 50 preset positions in both PA and AP must be available | #
3.9 | C-arm angulation of approx. 35° cranial and 35°caudal | #
3.10 | C-arm angulation speed variable, at least 15 ˚/sec | #
3.11 | At least 90° RAO and 45° LAO must be possible | #
3.12 | An automatic collimator system | #
3.13 | The remote control panel must be fully integrated and include controls for generator, image acquisition and tabletop/C-arm | #
3.14 | All relevant generator and system information must be presented as clearly legible displays | #
3.15 | Controls for table top/C-arm movement should be available at the table side and the remote panel | #
3.16 | Collision sensors for patient safety must be incorporated within the system | #
3.17 | Maximum weight limit for the table should be at least 165kg | #
3.18 | The following table accessories are to be included:- | #
3.18.1 | Shoulder rests
3.18.2 | Mattress
3.18.3 | Hand grips
3.18.4 | Foot rest
3.18.5 | Compression cone
3.18.6 | Compensation filter pad for peripheral studies
4 | Digital Flat Panel Detectors
4.1 | A large field digital flat panel with a maximum input field of 43 x 43 cm is required. | #
4.1.1 | Direct conversion method will be an advantage. Tenderer to state and motivate what technology/ method is used | #
4.2 | Detector matrix must be at least 2048 x 2048 or greater, 16bit. State and motivate | #
4.3 | Processor matrix must be at least 2048 x 2048 or greater, 16 bit. State and motivate | #
4.4 | State acquisition field sizes in sq.cm | #
4.5 | Resolution must not be less than 2.5 lp/mm. State | #
4.6 | State and motivate pixel size. Should not be more than 180µm | #
4.7 | State DQE – Acquisition dose (%). Preference will be given to the lowest dose system | #
4.8 | State DQE – Fluoroscopy dose (%). Preference will be given to the lowest dose system | #
5 | T.V. Monitors
5.1 | Two high resolution TFT LCD monitors of at least 48 cm to be ceiling mounted in the examination room | #
5.2 | One at least 48 cm high line-rate TFT LCD monitor to be supplied at the control desk | #
5.3 | Automatic brightness control to be included | #
5.4 | State monitor output brightness: (cd/m2)
5.4.1 | Monitors in exam room | #
5.4.2 | Monitor at control panel | #
6 | DIGITAL IMAGE ACQUISITION AND PROCESSING SYSTEM
6.1 | A digital imaging system suitable for general fluorography
6.2 | Digital acquisition matrix:
6.2.1 | Digital radiography: Minimum of 2048 X 2048 (2 Kilo Matrix), 12 bit, 0.5 to 7.5 frames/sec
6.2.2 | Digital fluoroscopy should be possible at 1024X1024, 30 FPS
6.3 | At least 3 pulsed fluoroscopy modes must be available
6.4 | Single shot exposures must be possible
6.5 | Last image hold function and instant transfer of fluoroscopy image, both single image and dynamic, to digital storage on a hard disk with at least 4 GB capacity. State capacity
6.6 | Storage capacity of 20 000 frames in 1024 X 1024 matrix for immediate access must be available. State capacity
6.7 | Matrix on CD-R must be possible. State
6.8 | DICOM 3 conformance statement to be included DICOM Print must be possible
6.9 | Backup recording on CD-R should also be possible
6.1 | Connection for sending images to a laser imager should utilize a digital or DICOM printing environment
6.11 | Functionality must also include:-
6.11.1 | Measuring of distance, angle and area
6.11.2 | Gray scale inversion
6.11.3 | Window and centre control
6.11.4 | Text annotation
6.11.5 | Up to X2 zoom facility required complete with scroll, zoom and cine loop display
6.11.6 | Horizontal (R/L) and vertical (up/down) image flip
6.11.7 | Gamma curve selection
6.11.8 | 16 & 4 on 1 multi-frame display with study/series overview Split display for reference image comparison on 1 monitor
7 | IT INTEGRATION
7.1 | The system must have the following DICOM compatibility:
7.1.1 | DICOM print class
7.1.2 | DICOM storage class
7.1.3 | DICOM query/retrieve
7.1.4 | DICOM worklist
7.1.5 | The system must comply to the following IHE profiles (Integrating the Healthcare Enterprise)
7.1.6 | Scheduled Workflow (SWF)
7.1.7 | Patient Information Reconciliation (PIR)
8 | RADIATION PROTECTION
8.1 | The system must have the ability to display Exposure Indicator values (EI/ S-values), immediately after exposure.
8.2 | X-ray Warning Notice for display at the X-ray Room entrance (not branded). | #
8.3 | An advanced Retake / Reject / Repeat analysis software package that includes analysis as well as calculations in percentages, must be provided for applicable systems. | #
8.4 | The applicable systems must have an integrated Dose Area Product (DAP) meter that includes skin dose equivalent measurement functions. | #
8.5 | A radiation report must be automatically generated and transferable via DICOM. | #
8.6 | Should be FDA or CE approved product | #
8.7 | Electrical safety conforms to standards for electrical safety IEC-60601
8.8 | The unit must comply to IEC-61267 | #
8.9 | The equipment/system must be approved and licensed by the Directorate Radiation Control of DoH: South Africa | #
8.10. | A copy of a valid license issued in terms of the Hazardous Substance Act, Act No 15 of 1973 must be submitted with the tender. Failure to submit such a valid license may result in a tender not being considered.
8.11 | Acceptance testing must be performed as part of the installation. (Failure to do will lead to failure to pay
8.12 | Lead Glass window to be installed as per room requirement
8.13 | Safety aspects of Radiation dosage leakage should be spelt out
8.14 | DB boards and electrical filters to be installed
8.15 | QA TOOLS
8.16 | Display Exposure index values
8.17 | Radiation protection Lead equivalence of the radiation protection screen must be 0.5m Pb and must be convex, covering the radiographer from various angles against scatter radiation.
8.18 | Patient Conventional face shield to be included.
8.19 | Patient Tomosynthesis face shield to be included, removable and retractable.
8.20. | Emergency stop controls must be supplied on both sides of the gantry and on the Acquisition Workstation
8.21 | Interlocks to stop system movement when compression is applied must be provided.
8.22 | Interlock for premature release of the X-ray switch must be available
8.23 | Interlock preventing an exposure whenever the filter or light field is not positioned correctly must be possible.
8.24 | Phantom for calibration should be provided.
9 | POWER SUPPLY
9.1 | The bidder must install adequate electrical power supply for the optimal functionality of the equipment.
9.2 | The mains cable of the unit being quoted for must be the hospital grade type and it must be of an adequate length (minimum length of three (3) meters).
9.3 | The power input must be 220-240V, 50Hz AC. Bidders must ensure that the product quoted for is fitted appropriately with a 15 Amp SABS approved mains plug that is held together by two screws.
9.4 | A 3 phase in / out power supply must be offered. State details.
9.5 | The equipment tendered for must not overload and trip the hospital power in the area where it is installed. Resettable overcurrent breaker/s must be fitted for protection. State details.
9.6 | All available online Uninterrupted Power Supply (UPS system) and line voltage regulators must be included for the system with a support time of 50 to 60 minutes. Attach brochure.
9.7 | The lifetime of UPS must be at least 15 years.
9.8 | An Uninterrupted Power Supply (UPS system) conforming to IS-302 must be offered for the computer system.
9.9 | The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade type and it must be a minimum length of three (3) meters.
9.10. | The mains cable of the unit tendered for must be SABS colour coded. State details.
9.11 | The equipment quoted for must be protected against electromagnetic interference. State details of measures that will be taken to protect the unit against electromagnetic interference.
9.12 | Maintenance free batteries must be offered.
10 | AIR CONDITIONER
10.1 | Bidder must include any air conditioning that is considered necessary for optimal functioning of the unit.
10.2 | Any such air conditioning must be covered by the guarantee bided for on the unit and servicing of the air conditioning units for the guarantee period and must be included in the bid price.
10.3 | The air conditioner must also be included in the maintenance plan.
10.4 | Successful bidder to replace existing air conditioner with an air conditioner of 24,000 BTU. (Where applicable).
10.5 | The air conditioner must be but no less than 24,000 BTU. State details on the air conditioner to be supplied according to the room size and machine requirements.
10.6 | The Successful bidder must install an outlet-plumbing pipe for water overflow from air conditioner to be re-directed to a drainage system.
10.7 | Air-conditioning must be included in Maintenance plan

## Full Body X-ray Scanner
FULL BODY X-RAY SCANNER
ITEM
1 | CLINICAL PURPOSE & FUNCTION
1.1 | A real time-saving, full-body digital X-ray imaging device, it provides an excellent quality X-ray image up to 6 feet in length in just 13 seconds, emitting up to 10 times less harmful dose than regular X-ray systems.
1.2 | The full-body X-ray machine plays a significant role in the management of the trauma patient, and is an important advance in the trauma imaging.
1.3 | The effectiveness and high speed of imaging dramatically reduce the resuscitation time of patients with major injury, and allow imaging of a large number of patients in a very short time
1.4 | The detection rate for treatment, low radiation dose and speed at which the whole body can be evaluated are advantages in the primary survey of acute trauma patients.
2 | PRODUCT CONFIGURATION
2.1 | The unit supplied shall be a Slit Scanner
2.2 | The system must be able to produce a full body digital X- ray image of a patient to assist in early location of life threatening injuries and locate fractures and foreign bodies.
2.3 | It shall be possible to have the unit located near or inside an emergency or resuscitation room, for this reason the unit shall emit very low doses of x - rays scattered into the environment around the system.
2.4 | The primary environment around the unit must comply with the international Commission for Radiation Protection's recommendations for x- ray exposure to members of the general public (<1mSv of radiation dose per year per any one person). Ref : ICRP Publication No 60. ISBN 0146 -6453.
2.5 | The system shall be capable of image quality of at least 4 lp/mm for spatial resolution by imaging a standard radiological line pair tool. State the resolution modes possible and describe the type of images that are normally captured in the different resolution modes.
3 | TECHNICAL SPECIFICATION
3.1 | X -Ray Generator :
3.1.1 | Generator Rating: minimum 60 KW.
3.1.2 | The system must allow for a 2 scans ( e.g.. AP and Lateral scan view) immediately after each other without overheating. The system must have a general through rate of at least 1 patient per minute. State the duty cycle of the offered sytem.
3.1.3 | The system must be equipped with emergency stop switch.
3.1.4 | The system must have an External manual exposure release switch
3.1.5 | The system must have an Automatic Exposure Control mechanism.
3.2 | Control Console
3.2.1 | Touch screen or touch key console.
3.2.2 | Selectable exposure facility for the various anatomic studies must be available
3.2.3 | Display of selected KV and mA/mAs
3.3 | X - Ray Tube and Tube Assembly
3.3.1 | Rotating Anode x- ray tube
3.3.2 | Provide thermal cut off for avoiding tube over heating.
3.3.3 | Tube rating should be higher than the specifications of generator ratings.
3.3.4 | The tube assembly must have fully counterbalanced movements with electromagnetic locks.
3.3.5 | The longitudinal travel must be able to cover a full body scan for a standard adult patient. State longitudinal travel range.
3.4 | Diagnostic Viewing Station
3.4.1 | The X- Ray unit shall be supplied with a medical diagnostic viewing station.
3.4.2 | The unit shall allow for a full function radiological viewing manipulation and storage of an X-Ray image
3.4.3 | The viewing station shall feature 20'' ( at least) high quality LCD radiological standard monitor designed for examining X- ray images. The monitor shall :
3.4.3.1 | Have a contrast ratio of at least 700:1
3.4.3.2 | Have a brightness of at least 300Cd/m2
3.4.3.3 | Have a resolution of at least 1600x 1200 pixels
3.4.3.4 | The system shal have a local storage capacty of at least 3 months of patient's information, at 50 patient per day. The system should have a removable storage device, for example a DVD - writer for transfer of images. State the size of the internal hard drive for storage of images.
3.4.3.5 | The system must be able to produce removable CD/DVD for off - line reviewing of individual or groups of cases. Each CD/DVD shal have the viewing program loaded which is needed in order to view the images on a commercially available PC.
3.4.3.6 | A plain paper printer of suitable quality will be supplied in order to print copies of images for record keeping or review purposes.
3.5 | Patient Trolley
3.5.1 | The imaging trolley shall be suitable to transport patients directly from a helicopter or ambulance to the proposed x -ray machine.
3.5.2 | The imaging trolley shall also be suitable, to function as a patient treatment trolley in the Trauma and Emergency Unit so that a minimum of patient transfers kis necessary. It shall be equipped with : a drip stand, patient retention rails,patient retention straps, x - ray translucent mattress, equipment basket and a gas cylinder holder.
4 | IT NETWORKING AND CONNECTIVITY
4.1 | The system will be DICOM Version 3 and HL7 compliant or better
4.2 | The system must have the following DICOM compatibility:
4.2.1 | DICOM print class
4.2.2 | DICOM storage class
4.2.3 | DICOM query/retrieve
4.2.4 | DICOM work list
4.3 | The system must comply to the following IHE profiles (Integrating the Healthcare Enterprise)
4.3.1 | Scheduled Workflow (SWF)
4.3.2 | Patient Information Reconciliation (PIR)
4.3.3 | Simple image and Numeric Reports (SINR)
4.3.4 | A safeguard method that protects the digital modality from potential malicious software attacks within the hospital network must be included. Describe method
5 | RADIATION PROTECTION
5.1 | The system must have the ability to display Exposure Indicator values (EI/ S-values), immediately after exposure.
5.2 | X-ray Warning Notice for display at the X-ray Room entrance (not branded).
5.3 | An advanced Retake / Reject / Repeat analysis software package that includes analysis as well as calculations in percentages, must be provided for applicable systems.
5.4 | The applicable systems must have an integrated Dose Area Product (DAP) meter that includes skin dose equivalent measurement functions.
5.5 | A radiation report must be automatically generated and transferable via DICOM.
5.6 | Should be FDA or CE approved product
5.7 | Electrical safety conforms to standards for electrical safety IEC-60601
5.8 | The unit must comply to IEC-61267
5.9 | The equipment/system must be approved and licensed by the Directorate Radiation Control of DoH: South Africa
5.10. | A copy of a valid license issued in terms of the Hazardous Substance Act, Act No 15 of 1973 must be submitted with the tender. Failure to submit such a valid license may result in a tender not being considered.
5.11 | Acceptance testing must be performed as part of the installation. (Failure to do will lead to failure to pay
5.12 | Lead Glass window to be installed as per room requirement
5.13 | Safety aspects of Radiation dosage leakage should be spelt out
5.14 | DB boards and electrical filters to be installed
5.15 | QA TOOLS
5.16 | Display Exposure index values
5.17 | Radiation protection Lead equivalence of the radiation protection screen must be 0.5m Pb and must be convex, covering the radiographer from various angles against scatter radiation.
5.18 | Patient Conventional face shield to be included.
5.19 | Patient Tomosynthesis face shield to be included, removable and retractable.
5.20. | Emergency stop controls must be supplied on both sides of the gantry and on the Acquisition Workstation
5.21 | Interlocks to stop system movement when compression is applied must be provided.
5.22 | Interlock for premature release of the X-ray switch must be available
5.23 | Interlock preventing an exposure whenever the filter or light field is not positioned correctly must be possible.
5.24 | Phantom for calibration should be provided.
6 | POWER SUPPLY
6.1 | The bidder must install adequate electrical power supply for the optimal functionality of the equipment.
6.2 | The mains cable of the unit being quoted for must be the hospital grade type and it must be of an adequate length (minimum length of three (3) meters).
6.3 | The power input must be 220-240V, 50Hz AC. Bidders must ensure that the product quoted for is fitted appropriately with a 15 Amp SABS approved mains plug that is held together by two screws.
6.4 | A 3 phase in / out power supply must be offered. State details.
6.5 | The equipment tendered for must not overload and trip the hospital power in the area where it is installed. Resettable overcurrent breaker/s must be fitted for protection. State details.
6.6 | All available online Uninterrupted Power Supply (UPS system) and line voltage regulators must be included for the system with a support time of 50 to 60 minutes. Attach brochure.
6.7 | The lifetime of UPS must be at least 15 years.
6.8 | An Uninterrupted Power Supply (UPS system) conforming to IS-302 must be offered for the computer system.
6.9 | The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade type and it must be a minimum length of three (3) meters.
6.10. | The mains cable of the unit tendered for must be SABS colour coded. State details.
6.11 | The equipment quoted for must be protected against electromagnetic interference. State details of measures that will be taken to protect the unit against electromagnetic interference.
6.12 | Maintenance free batteries must be offered.
7 | AIR CONDITIONER
7.1 | Bidder must include any air conditioning that is considered necessary for optimal functioning of the unit.
7.2 | Any such air conditioning must be covered by the guarantee bided for on the unit and servicing of the air conditioning units for the guarantee period and must be included in the bid price.
7.3 | The air conditioner must also be included in the maintenance plan.
7.4 | Successful bidder to replace existing air conditioner with an air conditioner of 24,000 BTU. (Where applicable).
7.5 | The air conditioner must be but no less than 24,000 BTU. State details on the air conditioner to be supplied according to the room size and machine requirements.
7.6 | The Successful bidder must install an outlet-plumbing pipe for water overflow from air conditioner to be re-directed to a drainage system.
7.7 | Air-conditioning must be included in Maintenance plan

## MRI 1.5T
70cm Wide Bore Magnetic Resonance Imaging(MRI) Unit (1.5T) Tesla
Item | Technical Specifications
1 | MAGNET
1.1 | The Magnet Field Strength shall be at least 1.5 Tesla active shielded
1.1.1 | State available upgrade path for unit offered (without magnet replacement). Please also state if the magnet on offer is the latest Technology and Release
1.2 | The Magnet shall be short bore cylindrical type and design
1.3 | Bore length including all covers shall not exceed 180 cm
1.4 | State internal bore dimensions (L x W x H)
1.4.1 | Maximum Vision Angle
1.5 | Magnet homogeneity shall meet the following specifications using the standard deviation VRMS (volume root-meansquare) using 24 plane plot measurement method:
1.5.1 | 50 cm DSV state
1.5.2 | 48 cm DSV state
1.5.3 | 45 cm DSV state
1.5.4 | 40 cm DSV state
1.5.5 | 30 cm DSV state
1.5.6 | 20 cm DSV state
1.5.7 | 10 cm DSV state
1.6 | Magnetic field stability shall be less than 0.1 ppm/hour or better
1.7 | The 5 Gauss/0.5mT fringe field shall be contained in an area of typically 2.5 m (radial) by 4.0m (axial). State actual area
1.8 | State helium boil-off details
1.8.1 | State typical cryogen refilling interval
1.8.2 | State maximum helium capacity
1.8.3 | State the weight of magnet with cryogens, table and covers (kg) in full operation
1.8.4 | The bore diameter shall be 70cm measured at the center in an operational mode
1.9 | The tenderer shall guarantee and ensure that there is no magnet vibration of the system, on the current location, with all possible imaging sequences. Give details.
1.9.1 | Space required (Magnet, Electronics and Control room)
2 | MAGNET SAFETY
2.1 | Magnet shall be equipped with quench exhaust leading to the outside of the building in event of Magnet quench to prevent injury to staff and patient
2.2 | Magnet shall be equipped with emergency ramp down unit for fast Magnetic Field reduction.
2.3 | Any other safety regulations by Environmental Safety Law and Regulations must be applied
2.4 | Successful bidder must provide a MRI user license obtained from Radiation Directorate
2.5 | Metal detector to be build into the magnet entrance door.
2.6 | Examination room shall be marked off by warning labels/lights
3 | GRADIENT SYSTEM
3.1 | Actively shielded hi-performance non-resonant gradient coil system
3.2 | Gradient amplitude/peak strength shall be at least 30mT/m measured per real axis plateau (100% duty cycle) not effective gradients
3.3 | Gradient duty cycle shall be 100%
3.4 | Minimum rise time from 0 to max. gradient amplitude shall be 240 micro-seconds or better
3.5 | Gradient slew rate shall be 120 T/m/s or better measured per real axis plateau, not effective values
3.5.1 | Shortest TE (SSh DWI, b=1000, 128 matrix)
3.5.2 | Shortest TE / TR (2D, 128 matrix) : 2D GRE
3.5.3 | Shortest TE / TR (2D, 128 matrix) : 3D GRE
3.5.4 | Shortest TE / TR (2D, 128 matrix) : 2D GRE EPI
3.5.5 | Shortest TE / TR (2D, 256 matrix) : 2D GRE
3.5.6 | Shortest TE / TR (2D, 256 matrix) : 3D GRE
3.5.7 | Shortest TE / TR (2D, 256 matrix) : 2D GRE EPI
3.5.8 | Maximum scan matrix
3.5.9 | Maximum recon matrix
3.5.10 | Maximum scan matrix
3.5.11 | Maximum recon matrix
3.5.12 | Increment steps for matrix
3.5.13 | Maximum number of slices
3.5.14 | Minimum slice thickness (3D)
3.5.15 | Duty cycle
3.5.16 | Gradient linearity at 30 cm DSV
3.5.17 | Gradient linearity at 40 cm DSV
3.5.18 | Gradient linearity at 50 cm DSV
3.6 | Gradient upgrades to different levels should be possible with out replacing magnet.
3.6.1 | State high-end gradient specifications for optional gradient upgrade:
3.6.2 | Max slew rate and amplitude and FOV at max slew rate and amplitude
3.6.3 | Cooling for gradient coil; state type of coolant
3.7 | The output linearity of the gradient amplifiers should be no worse than + 0.1% of peak
3.8 | A Child Friendly scanning enviroment is required. Noise should be kept to a minimum. State advanced ergonomic features that characterize the configuration offered.
3.9 | Free choice of flip angle while maintaining signal to noise ratio to be supported. Specify.
4 | RADIOFREQUENCY SYSTEM
4.1 | Resonance frequency shall be 63.86MHz (1.5 Tesla)
4.1.1 | Direct Radio Frequency (RF) Transmitter System: Digital signal generation and processing.
4.1.2 | RF system shall be digital transmitter and receiver design
4.1.3 | Maximum power output of transmitter amplifier rating shall be 28kW
4.1.4 | State the bandwidth of the RF transmitter
4.1.5 | The system shall be equipped with RF fault protection limiting RF output in event of malfunction
4.1.6 | Frequency resolution of the RF synthesizer shall be 0.35 Hz or better
4.2 | The receive components shall be integrated into the magnet housing
4.2.1 | The system should have at least 18 independent RF receiver channels or better
4.2.2 | Phase resolution shall be 0.1 degree/bit or better
4.2.4 | Noise of the preamplifier to be ≤ 0.5 decibels
4.2.5 | State maximum receiver bandwidth of each receiver channel
4.3 | All of these coil elements must be seamlessly used in one study without coil changes and without repositioning of the patient
4.4 | State maximum number of simultaneously connected coil elements. Preference will be given to highest number. Indicate coil combination.
4.5 | State maximum receiver bandwidth of each receiver channel.
4.6 | Sampling rate of each ADC to be about 80 MHz
4.7 | Receiver signal resolution should be 32 bit
5 | RF COILS
5.1 | The bidder shall supply the latest Integrated Coil Technology (i.e. GE = GEM Technology, Toshiba = Atlas Technology, Siemens = TIM Technology and Philips = dStream Technology). Please state how the Coil Philosophy of the unit contributes and improve the Image Quality and Workflow .
5.2 | Coil pre-amplifiers shall be on the patient table connector and coils shall be interchangeable and be of light construction.
5.3 | Diameter of 70cm is required (larger bore will be preffered because of the patient profile at this institution e.g. obese, claustrophic, paediatrics etc.) NB: The bore to be measured at centre of the system and not on the flare!
5.4 | Give standard surface coils supplied with system
5.4.1 | Maximum coils connection to do full body scan. Give details of coils and connections (short cable coil connections)
5.5 | Connected coils must be detected automatically
5.6 | It must be possible to select active coil or elements from the main console
5.7 | Special coils for the following applications should be available with the system (state price separately). State standard coils with the system.
5.7.1 | Body (Chest, abdomen, pelvis)
5.7.2 | TMJ kit
5.7.3 | Head
5.7.4 | Head-neck (without repositioning patient)
5.7.5 | C/T/L spine ( State what coil is used for imaging of the total spine cervical, thoracic and lumbar )
5.7.6 | Whole spine
5.7.7 | Neck soft tissue
5.7.8 | Extremities and joints: wrist, hips, shoulder, knee, foot/ankle,
5.7.9 | Peripheral array coil to be included in the price
5.7.10 | Prostate, colon, cervix (Endorectal)
5.7.11 | Whole body vascular (Peripheral vascular array)
5.7.12 | Peripheral MR Angiography
5.7.13 | Neurovascular
5.7.14 | Cardiac
5.7.15 | All paediatric applications
5.7.16 | Supply a cupboard in the room for storage of all coils provided or alternatively sufficient coil cabinets to house the coils.
5.7.17 | State all other SW packages that comes standard with the system.
5.7.18 | Other features
Coils to facilitate optimal facial, calvarial and intracranial imaging in the child from 1.5-60kgs, including visualisation of the orbits, optic nerves, optic tracts, pituitary fossa, hypothalamus, temporal lobe, posterior fossa and cerebral cortex. Coils to provide optimal imaging including, but not limited to the following clinical settings:
6 | PATIENT SUPPORT/TABLE and MANAGEMENT/ PATIENT COMFORT
6.1 | 2 X Patient table shall be dockable
6.2 | Vertical table movement
6.2.1.1 | Minimum table height from floor
6.2.1.2 | Maximum table height from floor
6.2.1.3 | Maximum vertical speed
6.2.2.1 | Horizontal table movement
6.2.2.2 | Maximum horizontal range of table movement
6.2.2.3 | Maximum speed:
6.2.2.4 | Accuracy of repositioning
6.3 | Dual table control panels shall be located at either side of aperture/gantry for easy access. State if controls are available on the rear of the magnet
6.4 | Three alignment light beams for anatomical references in axial, coronal and sagittal planes
6.5 | Table movement shall be controlled from both gantry and operator console
6.6 | Patient table shall be equipped with manual override for quick removal of patient from the magnet-bore in case of emergency
6.7 | Maximum load shall be 200kg or better. State actual maximum load.
6.8 | Physiologic measurement unit essential with display of ECG, respiration and pulse at the main console
6.9 | Two way in-bore intercom system shall allow communication with patient while gradient is running
6.10 | In-bore music/stereo system shall be included
6.11 | Hand held alarm button for patient signalling
6.12 | Table restraining/immobilization straps are required
6.13 | Separate Patient transport stretcher –MR compatible Max load 200kg to be supplied.
6.14 | Integrated music for patient, including CD interface to be included
6.15 | Variable patient lighting to be included.
6.16 | Child friendly interior decoration to be included.
6.17 | State of the art technology is required to allow young patients to see outside the magnet during the examination. Give details
6.18 | MR compatible sand bags
6.19 | Two comfortable patient mattresses to be included
7 | ADDITIONAL WORKSTATION: DOCTORS CONSOLE
7.1 | In addition to the acquisition console, additional, 2 x dedicated Workstations (Doctors Workstation) will be supplied (or better solution).
7.2 | The dedicated workstation will be used for all post processing. The system must included licenses for all 3D reconstruction and post processing requirements. Please give details and options where applicable
7.3 | Please state if the additional workstation has the same user interface as the console
7.4 | Please state if the additional workstation can have the same software packages as the console
7.5 | High quality read/write CD/DVD storage device required
8 | OPERATOR USER INTERFACE
8.1 | 23 inch hi-resolution colour LCD flat-panel flicker-free monitor with undistorted image display required
8.2 | Monitor 1024 x 1024 dot resolution required
8.3 | User interface providing flexible multi-tasking in foreground or background (scanning, filming, reconstruction)
8.4 | All standard image processing feature are required in addition the following advance features must be available:
8.4.1 | Multi-planner reformatting (MPR)
8.4.2 | Multi-projection volume rendering (MPVR)
8.4.3 | Maximum intensity projection (MIP)
8.4.4 | MR Angiography processing
8.4.5 | 3D surface rendering
8.4.6 | MR Hydrography MRCP/Urography/Myelography
8.4.7 | Image add/subtract
8.4.8 | CT image display
8.4.9 | CT image integrating must be offered.
8.4.10 | State any special features to improve productivity and consistency (e.g. automated scanning procedures)
9 | ACQUISITION PARAMATERS, APPLICATION SOFTWARE AND IMAGING TECHNIQUES
9.1 | State slice and slab thickness:
9.1.1 | 2D
9.1.2 | 3D
9.2 | State number of slices:
9.2.1 | 2D
9.2.2 | 3D
9.3 | Variable Field Of View required to maximum of 48cm or better. State
9.4 | STANDARD/CONVENTIONAL AND FAST IMAGING TECHNIQUES/ SEQUENCES REQUIRED FOR ALL THE FOLLOWING APPLICATIONS WITH DEDICATED POST PROCESSING
a | State standard sequences in offered software package
b | State maximum parallel imaging factor possible for offered software package (and compatible coils)
c | State sequences that are compatible with parallel imaging
d | State optional software packages. State details and costs
9.4.1 | Neuro:
i | Brain
ii | Conventional imaging sequences
iii | Spectroscopy (Single voxel and multi voxel). State if compatible with parallel imaging 2D / 3D
iv | Diffusion weighted
v | Perfusion weighted. Sate sequences used.
vi | Fast MRI sequences. Sate sequences used.
vii | Tractography / Fibre tracking
viii | Spine
9.4.2 | Cardiac:
i | Conventional sequences
9.4.3 | Abdomen:
i | Conventional sequences
ii | Virtual endoscopy
iii | MRCP
iv | Small bowel studies
v | Dynamic liver contrast studies
9.4.4 | Chest:
i | Conventional sequences
ii | Mammography
iii | Dynamic (3D FSGRE etc.)
iv | Biopsy capability
v | Pulmonary ventilation studies
9.4.5 | Pelvis:
i | Conventional sequences
ii | Prostate imaging (including spectroscopy)
iii | Dynamic pelvic floor imaging
9.4.6 | Vascular:
i | TOF (2D/3D)
ii | PC (2D/3D)
iii | Contrast enhancement for:
iv | Neck vessel studies
v | Intracranial vessel studies
vi | Aortic arch and branches
vii | Abdominal aorta and outflow
viii | Upper limbs
ix | Pulmonary arteries
x | Renal arteries
xi | Multistep MR angiography
xii | Lower limps
9.4.7 | Head and neck:
i | Conventional sequences
ii | IAM's
iii | TMJ's
10 | PERFORMANCE EVALUATION PHANTOMS REQUIRED / QUALITY CONTROL
10.1 | Performance evaluation software and appropriate phantoms required for evaluation of image quality required as per licensing conditions
11 | SAFETY
11.1 | The Magnet system shall include an emergency ramp down unit (RDU) for fast reduction of The Magnetic field.
11.2 | The Magnet shall have quench bands that contain The fringe Field in The event of a Magnet quench.
11.3 | Real-time SAR calculation shall be performed by software to ensure that RF power levels comply with regulatory guidelines and be displayed on each image.
11.4 | The system shall have manual override of the motor drive for quick removal of the patient from the magnet bore.
11.5 | The system shall have RF fault protection hardware (RF transmit enable limit and RF power and duty cycle limit) to limit The RF output in The event of equipment malfunction.
11.6 | Walk through metal detector
Require a scavenger for anaesthetic gases to be removed from the magnet room
12 | ADDITIONAL EQUIPMENT
12.1 | MRI compatible ECG cables required
12.3 | MRI integrated pulse oximeter for adults and children
13 | MR FLUROSCOPY
13.1 | a MR Fluoroscopy mode with real-time acquisition and display techniques is required as an option. Please State the soft and hardware required if this is optional.
14 | STANDARD ACCESSORIES
List of standard accessories included but not specified in the tender.
14.1 | Incubator for Neonatal Imaging
14.2 | MRI compatible contrast injector interphased with MR unit
14.3 | 200 syringes and 200 matching syringe lines
15 | TECHNOLOGY
15.1 | Innovation. Give explanation
15.2 | Unique. Give explanation
15.3 | Clinical benefits. Give detailed explanation
16 | OPTIONAL ACCESSORIES
16.1 | The tenderer to give full description and pricing of optional accessories available for the system. Each item must be priced separately
17 | MRI COMPATIBLE ANAESTHETIC UNIT
17.1 | MRI Compatible Anaesthetic unit for General Anaesthetics of the paedeatric and adult patients is required. Please state the specifications of the unit on offer. Please also note the warranty and customer support agreement terms.
17 | LASER CAMERA
17.1 | A state of the art dry laser camera to be included. Specify: Make Model
17.2 | UPS for the laser camera to be included
B General Issues
1 | SAFETY STANDARDS OF THE SYSTEM
1.1 | MR Unit should be CE marked
1.2 | MR Unit should be manufacture according to FDA 1020.30 regulations
1.3 | MR Unit must Comply to IEC 601
1.4 | The equipment/system must be approved and licensed by Radiation Control.
1.5 | A copy of a valid license issued in terms of the Hazardous Substance Act, Act No 15 of 1973 must be submitted with the tender. Failure to submit such a valid license may result in a tender not being considered.
1.6 | Electrical safety conforms to standards for electrical safety IEC-60601 / IS-13450
2 | IT INTEGRATION
The system must have the following DICOM 3 compatibility:
2.1 | DICOM Send / Receive
2.2 | DICOM Storage Commitment (SC):
2.3 | DICOM query/retrieve
2.4 | DICOM worklist
2.5 | DICOM Basic Print
2.6 | DICOM MPPS (Modality Performed Procedure Steps)
2.7 | The system must comply to the following IHE profiles (Integrating the Healthcare Enterprise)
2.7.1 | Scheduled Workflow (SWF)
2.7.2 | Patient Information Reconciliation (PIR)
2.7.3 | Simple Image and Numeric Reports (SINR)
2.8 | System to be integrated to the existing PACS
7 | AIR CONDITIONER
7.1 | Bidder must include any air conditioning that is considered necessary for optimal functioning of the unit.
7.2 | Any such air conditioning must be covered by the guarantee bided for on the unit and servicing of the air conditioning units for the guarantee period and must be included in the bid price.
7.3 | The air conditioner must also be included in the maintenance plan.
7.4 | Successful bidder to replace existing air conditioner with an air conditioner of 24,000 BTU. (Where applicable).
7.5 | The air conditioner must be but no less than 24,000 BTU. State details on the air conditioner to be supplied according to the room size and machine requirements.
7.6 | The Successful bidder must install an outlet-plumbing pipe for water overflow from air conditioner to be re-directed to a drainage system.
7.7 | Air-conditioning must be included in Maintenance plan

## MRI 3T
MRI MACHINES (3T)
Item | Technical Specification | Weight
A | Modality Details
1 | MAGNET
1.1 | The Magnet Field Strength shall be at least 3 Tesla active shielded | #
1.1.1 | State available upgrade path for unit offered (without magnet replacement) Please also state if the magnet of offer is the latest Technology and Release | #
1.2 | The Magnet shall be short bore cylindrical type and design | #
1.3 | Bore length including all covers shall not exceed 180 cm | #
1.4 | State internal bore dimensions (L x W x H) | #
1.4.1 | Maximum Vision Angle | #
1.5 | Magnet homogeneity shall meet the following specifications using the standard deviation VRMS (volume root-meansquare) using 24 plane plot measurement method: | #
1.5.1 | 50 cm DSV state | 1
1.5.2 | 48 cm DSV state | 1
1.5.3 | 45 cm DSV state | 1
1.5.4 | 40 cm DSV state | 1
1.5.5 | 30 cm DSV state | 1
1.5.6 | 20 cm DSV state | 1
1.5.7 | 10 cm DSV state | 1
1.6 | Magnetic field stability shall be less than 0.1 ppm/hour or better | 5
1.7 | The 5 Gauss/0.5mT fringe field shall be contained in an area of typically 2.5 m (radial) by 4.6m (axial). State actual area | 5
1.8 | State helium boil-off details | 5
1.8.1 | State typical cryogen refilling interval | 5
1.8.2 | State maximum helium capacity | 5
1.8.3 | State the weight of magnet with cryogens, table and covers (kg) in full operation | 5
1.8.4 | The bore diameter shall be 70cm measured at the center in an operational mode | #
1.9 | The tenderer shall guarantee and ensure that there is no magnet vibration of the system, on the current location, with all possible imaging sequences. Give details. | 5
1.9.1 | Space required (Magnet, Electronics and Control room) | 5
2 | MAGNET SAFETY
2.1 | Magnet shall be equipped with quench exhaust leading to the outside of the building in event of Magnet quench to prevent injury to staff and patient | #
2.2 | Magnet shall be equipped with emergency ramp down unit for fast Magnetic Field reduction. | #
2.3 | Any other safety regulations by Environmental Safety Law and Regulations must be applied | #
2.4 | Successful bidder must provide a MRI user license obtained from Radiation Directorate | #
2.5 | Metal detector to be build into the magnet entrance door. | #
2.6 | Examination room shall be marked off by warning labels/lights | #
3 | GRADIENT SYSTEM
3.1 | Actively shielded hi-performance non-resonant gradient coil system | #
3.2 | Gradient amplitude/peak strength shall be at least 40mT/m measured per real axis plateau (100% duty cycle) not effective gradients | 5
3.3 | Gradient duty cycle shall be 100% | 5
3.4 | Minimum rise time from 0 to max. gradient amplitude shall be 240 micro-seconds or better | 5
3.5 | Gradient slew rate shall be 180 T/m/s or better measured per real axis plateau, not effective values | 5
3.5.1 | Shortest TE (SSh DWI, b=1000, 128 matrix) | #
3.5.2 | Shortest TE / TR (2D, 128 matrix) : 2D GRE | #
3.5.3 | Shortest TE / TR (2D, 128 matrix) : 3D GRE | #
3.5.4 | Shortest TE / TR (2D, 128 matrix) : 2D GRE EPI | #
3.5.5 | Shortest TE / TR (2D, 256 matrix) : 2D GRE | #
3.5.6 | Shortest TE / TR (2D, 256 matrix) : 3D GRE | #
3.5.7 | Shortest TE / TR (2D, 256 matrix) : 2D GRE EPI | #
3.5.8 | Maximum scan matrix | #
3.5.9 | Maximum recon matrix | #
3.5.10 | Maximum scan matrix | #
3.5.11 | Maximum recon matrix | #
3.5.12 | Increment steps for matrix | #
3.5.13 | Maximum number of slices | #
3.5.14 | Minimum slice thickness (3D) | #
3.5.15 | Duty cycle | #
3.5.16 | Gradient liniearity at 10 cm DSV | #
3.5.17 | Gradient linearity at 20 cm DSV | #
3.5.18 | Gradient linearity at 30 cm DSV | #
3.5.19 | Gradient linearity at 40 cm DSV | #
3.5.20 | Gradient linearity at 50 cm DSV | #
3.6 | Gradient upgrades to different levels should be possible without changing gradient coil. Quote optional gradient upgrades.Quote prices | #
3.6.1 | State high-end gradient specifications for optional gradient upgrade: | #
3.6.2 | Max slew rate and amplitude and FOV at max slew rate and amplitude | #
3.6.3 | Cooling for gradient coil; state type of coolant | 4
3.7 | The output linearity of the gradient amplifiers should be no worse than + 0.1% of peak | 5
3.8 | A Child Friendly scanning enviroment is required. Noise should be kept to a minimum. State advanced ergonomic features that characterize the configuration offered. | #
3.9 | Free choice of flip angle while maintaining signal to noise ratio to be supported. Specify. | 5
4 | RADIOFREQUENCY SYSTEM
4.1 | Resonance frequency shall be 128MHz (3 Tesla) | #
4.1.1 | Direct Radio Frequency (RF) Transmitter System: Digital signal generation and processing. | #
4.1.2 | The transmission system to be integrated in the magnet housing | #
4.1.3 | Maximum power output of transmitter amplifier rating shall be 35kW | 5
4.1.4 | State the bandwidth of the RF transmitter | 5
4.1.5 | The system shall be equipped with RF fault protection limiting RF output in event of malfunction | 5
4.1.6 | Frequency resolution of the RF synthesizer shall be 0.35 Hz or better | 5
4.2 | The receive components shall be integrated into the magnet housing | #
4.2.1 | The system should have at least 48 independent RF receiver channels or better | 5
4.2.2 | Phase resolution shall be 0.1 degree/bit or better | 5
4.2.3 | Noise of the preamplifier to be ≤ 0.5 decibels | 5
4.2.4 | State maximum receiver bandwidth of each receiver channel | 5
4.3 | Can all of these coil elements be seamlessly used in one study without coil changes and without repositioning of the patient. | #
4.4 | State maximum number of simultaneously connected coil elements. Preference will be given to highest number. Indicate coil combination. | 5
4.5 | State maximum receiver bandwidth of each receiver channel. | 5
4.6 | Sampling rate of each ADC to be about 80 MHz | 5
4.7 | Receiver signal resolution should be 32 bit | 5
5 | RF COILS
5.1 | The bidder shall supply the latest Integrated Coil Technology (i.e. GE = GEM Technology, Toshiba = Atlas Technology, Siemens = TIM Technology and Philips = dStream Technology). Please state how the Coil Philosophy of the unit contributes and improve the Image Quality and Workflow . | 5
5.2 | Coil pre-amplifiers shall be on the patient table connector and coils shall be interchangeable and be of light construction. | 5
5.3 | Diameter of 70cm is required (larger bore will be preferred because of the patient profile at this institution e.g. obese, claustrophic, paediatrics etc.) NB: The bore to be measured at centre of the system and not on the flare! | #
5.4 | Maximum coils connection to do full body scan. Give details of coils and connections (short cable coil connections) | 4
5.5 | Connected coils must be detected automatically | 4
5.6 | It must be possible to select active coil or elements from the main console | 4
5.7 | Special coils for the following applications should be available with the system (state price separately)
5.7.1 | Body (Chest, abdomen, pelvis) | #
5.7.2 | TMJ kit | 3
5.7.3 | Head | #
5.7.4 | Head-neck | #
5.7.5 | C/T/L spine ( State what coil is used for imaging of the total spine cervical, thoracic and lumbar ) | #
5.7.6 | Whole spine | #
5.7.7 | Neck soft tissue | #
5.7.8 | Extremities and joints: wrist, hips, shoulder, knee, foot/ankle, | #
5.7.9 | Breast . State solution for Breast MR with interventional capabilities. | #
5.7.10 | Prostate, colon, cervix | #
5.7.11 | Whole body vascular (Peripheral vascular array) | 4
5.7.12 | Peripheral vascular | #
5.7.13 | Neurovascular | #
5.7.14 | Dedicated Advanced Cardiac | #
5.7.15 | All paediatric applications | #
5.7.16 | Supply a cupboard in the room for storage of all coils provided or alternatively sufficient coil cabinets to house the coils. | #
5.7.17 | State all other SW packages that comes standard with the system. All these SW packages have to be Advanced, for use in an academic and research institution and included in the offer. | #
5.7.18 | Other features
| Coils to facilitate optimal facial, calvarial and intracranial imaging in the child from 1.5-60kgs, including visualisation of the orbits, optic nerves, optic tracts, pituitary fossa, hypothalamus, temporal lobe, posterior fossa and cerebral cortex. Coils to provide optimal imaging including, but not limited to the following clinical settings: | 5
6 | PATIENT SUPPORT/TABLE and MANAGEMENT/ PATIENT COMFORT
6.1 | Patient table shall be preferable dockable/detachable or similar solution to increase throughput and for comfort | #
6.2 | One table must be inluded in the total price, state price of a second table (or device as per 6.1 above ). Please give details.Quote prices | #
6.3 | Vertical table movement
6.3.1 | Minimum table height from floor | 3
6.3.2 | Maximum table height from floor | 3
6.3.3 | Maximum vertical speed | 3
6.3.4 | Horizontal table movement:
6.3.5 | Maximum horizontal range of table movement | 3
6.3.6 | Maximum speed: | 3
6.3.7 | Accuracy of repositioning | 3
6.4 | Dual table control panels shall be located at either side of aperture/gantry for easy access. State if controls are available on the rear of the magnet | 5
6.5 | Three alignment light beams for anatomical references in axial, coronal and sagittal planes | 5
6.6 | Table movement shall be controlled from both gantry and operator console | 5
6.7 | Patient table shall be equipped with manual override for quick removal of patient from the magnet-bore in case of emergency | 5
6.8 | Maximum load shall be 200kg or better. State actual maximum load. | 5
6.9 | Physiologic measurement unit essential with display of ECG, respiration and pulse at the main console | 5
6.10. | Two way in-bore intercom system shall allow communication with patient while gradient is running | #
6.11 | In-bore music/stereo system shall be included | #
6.12 | Hand held alarm button for patient signalling | 3
6.13 | Table restraining/immobilization straps are required | 3
6.14 | Separate Patient transport stretcher –MR compatible Max load 200kg to be supplied. | 3
6.15 | Integrated music for patient, including CD interface to be included | #
6.16 | Variable patient lighting to be included. | 5
6.17 | Child friendly interior decoration to be included. | 5
6.18 | State of the art technology is required to allow young patients to see outside the magnet during the examination. Give details | 5
6.19 | MR compatible sand bags | 3
6.2 | Two comfortable patient mattresses to be included | 5
7 | ADDITIONAL WORKSTATION: DOCTORS CONSOLE
7.1 | In addition to the acquisition console, additional, 3 X Dedicated Workstations (Doctors Workstation) will be supplied (or better solution). | #
7.2 | The dedicated workstation will be used for all post processing. The system must included licenses for all 3D reconstruction and post processing requirements. Please give details and options where applicable | 5
7.3 | Please state if the additional workstation has the same user interface as the console | #
7.4 | Please state if the additional workstation can have the same software packages as the console | #
7.5 | High quality read/write CD/DVD storage device required | 5
8 | OPERATOR USER INTERFACE
8.1 | 23 inch hi-resolution colour LCD flat-panel flicker-free monitor with undistorted image display required | 5
8.2 | Monitor 1024 x 1024 dot resolution required | 5
8.3 | User interface providing flexible multi-tasking in foreground or background (scanning, filming, reconstruction) | 5
8.4 | All standard image processing feature are required in addition the following advance features must be available:
8.4.1 | Multi-planner reformatting (MPR) | #
8.4.2 | Multi-projection volume rendering (MPVR) | #
8.4.3 | Maximum intensity projection (MIP) | #
8.4.4 | MR Angiography processing | #
8.4.5 | 3D surface rendering | 3
8.4.6 | MR Hydrography MRCP/Urography/Myelography | 3
8.4.7 | Image add/subtract | 3
8.4.8 | CT image display | 3
8.4.9 | CT image integrating must be offered. | 5
8.4.10 | State any special features to improve productivity and consistency (e.g. automated scanning procedures) | 4
9 | ACQUISITION PARAMATERS, APPLICATION SOFTWARE AND IMAGING TECHNIQUES
9.1 | State slice and slab thickness:
9.1.1 | 2D | 5
9.1.2 | 3D | 5
9.2 | State number of slices:
9.2.1 | 2D | 5
9.2.2 | 3D | 5
9.3 | Variable Field Of View required to maximum of 48cm or better. State | 5
9.4 | Standard/Conventional and Fast Imaging Techniques/ Sequences Required for all the following Applications with Dedicated Post Processing: | #
a | State standard sequences in offered software package | #
b | State maximum parallel imaging factor possible for offered software package (and compatible coils) | #
c | State sequences that are compatible with parallel imaging | #
d | State optional software packages. State details and costs | #
9.4.1 | Neuro:
i | Brain | #
ii | Conventional imaging sequences | #
iii | Spectroscopy (Single voxel and multi voxel). State if compatible with parallel imaging 2D / 3D | #
iv | Diffusion weighted | #
v | Perfusion weighted. Sate sequences used. | #
vi | Fast MRI sequences. Sate sequences used. | #
vii | Tractography / Fibre tracking | 3
viii | Spine | #
9.4.2 | Cardiac:
i | Functional evaluation | #
ii | Morphology | #
iii | Valvular analysis | #
iv | Coronary artery imaging | #
v | Viability studies | #
vi | ECG gating | #
vii | Myocardial tagging | #
9.4.3 | Abdomen:
i | Conventional sequences | #
ii | Virtual endoscopy | #
iii | MRCP | #
iv | Small bowel studies | #
v | Dynamic liver contrast studies | #
9.4.4 | Chest:
i | Conventional sequences | #
ii | Mammography | #
iii | Dynamic (3D FSGRE etc.) | #
iv | Biopsy capability | #
v | Pulmonary ventilation studies | #
9.4.5 | Pelvis:
i | Conventional sequences | #
ii | Prostate imaging (including spectroscopy) | #
iii | Dynamic pelvic floor imaging | #
9.4.6 | Vascular:
i | TOF (2D/3D) | #
ii | PC (2D/3D) | #
iii | Contrast enhancement for:
iv | Neck vessel studies | #
v | Intracranial vessel studies | #
vi | Aortic arch and branches | #
vii | Abdominal aorta and outflow | #
viii | Upper limbs | #
ix | Pulmonary arteries | #
x | Renal arteries | #
xi | Multistep MR angiography | 3
9.4.7 | Head and neck:
i | Conventional sequences | #
ii | IAM's | #
iii | TMJ's | 3
10 | PERFORMANCE EVALUATION PHANTOMS REQUIRED / QUALITY CONTROL
10.1 | Performance evaluation software and appropriate phantoms required for evaluation of image quality required as per licensing conditions | #
11 | SAFETY
11.1 | The Magnet system shall include an emergency ramp down unit (RDU) for fast reduction of The Magnetic field. | #
11.2 | The Magnet shall have quench bands that contain The fringe Field in The event of a Magnet quench. | #
11.3 | Real-time SAR calculation shall be performed by software to ensure that RF power levels comply with regulatory guidelines and be displayed on each image. | #
11.4 | The system shall have manual override of the motor drive for quick removal of the patient from the magnet bore. | #
11.5 | The system shall have RF fault protection hardware (RF transmit enable limit and RF power and duty cycle limit) to limit The RF output in The event of equipment malfunction. | #
11.6 | Walk through metal detector | 5
12 | ADDITIONAL EQUIPMENT
12.1 | MRI compatible ECG cables required | #
12.2 | MRI integrated pulse oximeter for adults and children | #
13 | MR FLUROSCOPY
13.1 | a MR Fluoroscopy mode with real-time acquisition and display techniques is required as an option. Please State the soft and hardware required if this is optional. | 5
14 | STANDARD ACCESSORIES
14.1 | List of standard accessories included but not specified in the tender.
14.1.1 | Incubator for Neonatal Imaging | #
14.1.2 | MRI compatible contrast injector interphased with MR unit | #
14.1.3 | 200 syringes and 200 matching syringe lines | 4
14.1.4 | Drug Infusion pump compatible to MRI | #
15 | TECHNOLOGY
15.1 | Innovation. Give explanation | 3
15.2 | Unique. Give explanation | 3
15.3 | Clinical benefits. Give detailed explanation | 5
16 | OPTIONAL ACCESSORIES
16.1 | The tenderer to give full description and pricing of optional accessories available for the system. Each item must be priced separately.Quote prices | 5
17 | MRI COMPATIBLE ANAESTHETIC UNIT
7.1 | MRI Compatible Anaesthetic unit for General Anaesthetics of the paedeatric patients is required. Please state the specifications of the unit on offer. Please also note the warranty and customer support agreement terms. | #
18 | LASER CAMERA
18.1 | A state of the art dry laser camera to be included. Specify: Make Model | 5
18.2 | UPS for the dry laser camera to be included | #
19 | General Issues
19.1 | IT INTEGRATION
The system must have the following DICOM 3 compatibility:
19.1.1 | DICOM Send / Receive | #
19.1.2 | DICOM Storage Commitment (SC): | #
19.1.3 | DICOM query/retrieve | #
19.1.4 | DICOM worklist | #
19.1.5 | DICOM Basic Print | #
19.1.6 | DICOM MPPS (Modality Performed Procedure Steps) | #
19.2 | The system must comply to the following IHE profiles (Integrating the Healthcare Enterprise) | #
19.2.1 | Scheduled Workflow (SWF) | 5
19.2.2 | Patient Information Reconciliation (PIR) | 5
19.2.3 | Simple Image and Numeric Reports (SINR) | 5
19.2.4 | System to be integrated to the existing PACS | #
7 | AIR CONDITIONER
7.1 | Bidder must include any air conditioning that is considered necessary for optimal functioning of the unit.
7.2 | Any such air conditioning must be covered by the guarantee bided for on the unit and servicing of the air conditioning units for the guarantee period and must be included in the bid price.
7.3 | The air conditioner must also be included in the maintenance plan.
7.4 | Successful bidder to replace existing air conditioner with an air conditioner of 24,000 BTU. (Where applicable).
7.5 | The air conditioner must be but no less than 24,000 BTU. State details on the air conditioner to be supplied according to the room size and machine requirements.
7.6 | The Successful bidder must install an outlet-plumbing pipe for water overflow from air conditioner to be re-directed to a drainage system.
7.7 | Air-conditioning must be included in Maintenance plan