## Linear accelerator
LINEAR ACCELERATOR
Item |  | Weighting
1 | GENERAL DESCRIPTION
1.1 | Dual energy photon linear accelerator (linac) with electrons. Fitted with MLC, EPID and kV on board imaging capable of doing IMRT, VMAT, IGRT,SRS, SBRT and adapative RT.
2 | TECHNICAL SPECIFICATIONS
2.1 | Compulsory Technical specifications
2.1.1 | Photons energies of 6MV, 6MV FFF, 10MV, 10MV FFF and 15MV
2.1.2 | Must have at least five standard electron energies ranging from 4 MeV to 15 MeV
2.1.3 | MLC minimum of 120 leaves (0.5 cm at isocenter) up to maximum allowed field size.
2.1.4 | EPID high end with portal and in vivo dosimetry (state detector resolution and size)
2.1.5 | On board kV imaging (retractable)
2.1.6 | Soft/virtual/Motorized/EDW wedge.
2.1.7 | Physical/motorized wedges (15 deg to 60 deg in 15deg steps)
2.1.8 | Treatment couch must take minimum of 200 kg.
2.1.9 | Treatment couch must have carbon fibre top compatible with system in use
2.1.10 | Shadow tray (if applicable)
2.1.11 | Standard electron applicators - square from 6x6 cm to 25x25 cm (including 10x10 cm) and (Optional- circular)
2.1.12 | Lasers, sagittal and 2 x side (green or red)
2.1.13 | Optional Back pointer, Isocentric check front pointer or Back pointer laser
2.1.14 | Minimum 2 high definition CCTV, to enable full vision of patient during therapy, 1 with full movements and rotation; Intercom, last man button
2.1.15 | Collision sensors/interlogs must be available
2.2 | Functional specifications
2.2.1 | 2D, 3D , IMRT, VMAT enabled ( dynamic MLC), SRS enabled with frames and immobilizations
2.2.2 | Variable gantry speed, variable dose rate and dynamic MLC
2.2.3 | At least two Treatment planning system for one linac with 2D, 3D , IMRT, VMAT enabled ( dynamic MLC), SRS enabled licenses and atleas two Oncology workstations.( one must be added for each added Linear accelerator)
2.2.4 | Remote setup - it must be possible to set the linac up from the console without the need of an operator in the treatment room (as per manufactured recommendations, state the details)
2.2.5 | Must interface with all Record & verify, oncology management system (Very important) - OMS, state number of R& V workstations (minimum 4 for one linac and 3 for each added linac).
2.2.6 | IGRT based on kV Cone beam CT, EPID based treatment verification (2D-2D imaging with MV and kV, CBCT verification and offline review.
2.2.7 | All patient immobilization devices required to cover different cancer diagnoses, for both linear accelerator and CT scanner ( Head and Neck, Thorax, Abdomen, knees and feet rest).
2.2.8 | Head and Neck frames and all supporting accessories, breast Boards, pelvis frames, Knee rest foot rest, Masks).
2.2.9 | Include all supplied accessories for immobilization with specific details.
2.2.10 | Sample mask must be 50 for each treatment site
2.2.11 | Vacuum cushion set for different patient size ( x4 in each) with pump
2.2.12 | Compulsory onsite training and 4 offsite (international training for all speciality as agreed by all parties)
2.3 | Optional extras
Specify the cost for optional items. Please note that this will form part of the total cost for the evaluation of the tender
2.3.1 | 0.25 cm MLC for SRS treatment (Build- in or removable).
3 | Dosimetry and Patient Verification equipment (separately priced)
3.1 | The following equipment's are essential for ensuring quality, compliance and safety of delivered treatment to patients. All compulsory dosimetry equipment must be provided, installed and tested for functionality on each application.
3.2 | Barometer and thermometer with measuring probe for water temperature.
3.3 | 3 absolute dosimetry ionization chambers for photons (1), electrons (1) and Small fields (1), with electrometer and calibration certificates (refer to Dosimeter specifications)
3.4 | 1-D depth adjustable water tank with chamber holders fro TRS 398 absolute dosimetry out put measurements.
3.5 | 3D Scanning tank with all accessories required for commissioning of the Linear accelerator and treatment planning systems (training must be provided before commissioning of the all units) with at least 3 licenses and software's
3.6 | Latest technology 2D array detector complete package system for advance therapy patient pre-treatment verification in three dimension (IMRT, VMAT and SRS).
3.7 | MV, kV and cone beam CT, alignment tools, phantoms for quality assurance and image quality checks with processing software and licenses.
3.8 | Daily output check system (Diodes detector system or any accepted equivalent)
3.9 | In vivo dosimetry comprehensive solution.
3.10. | Linear array detector system for relative dosimetry (Flatness and Symmetry, Energy check), with water equivalent phantom plates (2x0.1, 1x0.2, 1x0.5, and 20x1.0 cm.), water equivalent plates with insert for photon ionization chamber and electrons ionization chamber.
3.11 | Radiation survey meter
3.12 | Radiation pocket dosimeter
3.13 | SRS and SBRT patient based pre-treatment verifications, complete solution.
3.14 | Two boxes of EBT 3 Film for film dosimetry
4 | POWER SUPPLY
4.1 | The bidder must install adequate electrical power supply for the optimal functionality of the equipment.
4.2 | The mains cable of the unit being quoted for must be the hospital grade type and it must be of an adequate length (minimum length of three (3) meters).
4.3 | The power input must be 220-240V, 50Hz AC. Bidders must ensure that the product quoted for is fitted appropriately with a 15 Amp SABS approved mains plug that is held together by two screws.
4.4 | A 3 phase in / out power supply must be offered. State details.
4.5 | The equipment tendered for must not overload and trip the hospital power in the area where it is installed. Resettable overcurrent breaker/s must be fitted for protection. State details.
4.6 | All available online Uninterrupted Power Supply (UPS system) and line voltage regulators must be included for the system with a support time of 50 to 60 minutes. Attach brochure.
4.7 | The lifetime of UPS must be at least 15 years.
4.8 | An Uninterrupted Power Supply (UPS system) conforming to IS-302 must be offered for the computer system.
4.9 | The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade type and it must be a minimum length of three (3) meters.
4.10. | The mains cable of the unit tendered for must be SABS colour coded. State details.
4.11 | The equipment quoted for must be protected against electromagnetic interference. State details of measures that will be taken to protect the unit against electromagnetic interference.
4.12 | Maintenance free batteries must be offered.
5 | DOSIMETRY (optional extra, must be priced separatly)
5.1 | Therapy Electrometer
5.1.1 | Therapy and diagnostic dosimeter, for dose and dose rate measurements. Connecting system M type. Includes RS232, rechargeable batteries and mains cord for 115/230V, 50-60Hz
5.2 | Therapy Ionization Chambers
5.2.1 | Farmer type chamber 0.6 cm³, waterproof acrylic/aluminum; fully guarded therapy chamber with acrylic wall and aluminum electrode. Cable length 1.3m, connecting system M type. Includes acrylic build-up cap
5.2.2 | Advanced Markus electron chamber 0.02 cm³ wide guard ring, parallel plate, poly-ethylene therapy chamber with graphite coated acrylic electrode. Open chamber volume. Chamber is waterproof when using 0.87 mm acrylic cap. Cable length 1.05 m, connecting system M type
5.3 | Supporting accessories
5.3.1 | Extension cable 'M type 6m HD'
5.3.2 | 6 m 'heavy duty' detector to dosemeter extension cable, connecting system M type.
5.3.3 | Wall mount for M type connector, single
5.4 | Environmental conditions
5.4.1 | Barometer, digital precision barometer for measurement of absolute air pressure,
5.4.2 | Precision thermometer with probe -20°C to 50°C, relative humidity 10% to 95%, software Smart Graph 3, USB, LAN
5.5 | Radiation Beam Characteristics
5.5.1 | 3D water phantom, complete for atleast 50 x 50 x 41 cm scans. Includes lifting carriage/reservoir, dual channel electrometer, Two 3D chambers for measurements and reference, software with all options option. Connecting system M type. Laptop must be included
5.5.2 | 4D Pretreament verirfication system for IMRT/VMAT/SRS dosimetry. Automatic motorized tracing of the detector following the rotation of the gantry. Includes all software package and trolley. Laptop must be included
5.6 | Radiation Safety equipments
5.6.1 | Portable Survey meter, dose rate measuring with internal proportional counter tube for Gamma, 30 keV - 4.4 MeV H*(10). Range of indication 10nSv/h-100 mSv/h. Calibratable 0.5 μSv/h-100 mSv/h. Includes batteries, carrying strap, carrying bag and protective foil
5.6.2 | EBT3 Films
5.6.3 | Electronic Personal Dosemeter with alarm function and dose rate indication for the detection of the personal - dose when handling ionising radiation. :
5.6.4 | Accredited laboratory approval dosemeter for measuring gamma and X-rays for the measuring size Hp(10) (personal dose equivalent)
5.6.5 | Detector: energy compensated GM tube Radiation in front of the user's body preset dose and dose rate alarm thresholds Measuring Areas
5.6.6 | Dose indication range: 0.1 µSv ≤ Hp(10) ≤ 9.0 Sv
5.6.7 | Dose rate indication range: 0.1 µSv/h to 1.5 Sv/h.
5.7 | Radiation output measurement devices and QA
5.7.1 | 1D Scanning Arm, Tank, with Software
5.7.2 | Water Phantom - with vertical lines for alignment to isocenter using room lasers Fill lines for 20cm and 25cm depth measurements.
5.7.3 | Integral drain valve with included draining tube.
5.7.4 | Control Arm - Rotary encorder directly mounted to leadscrew shaft.
5.7.5 | 0.05mm accuracy of position and repeatability over 275.00mm travel distance.
5.7.6 | 0.05mm accuracy of position and repeatability over 275.00mm travel distance.
5.7.7 | Remote Handheld Controller.
5.7.8 | Internal memory.
5.7.9 | Software included for fully automated programmable PDD collection
5.8 | Daily Output check
5.8.1 | QUICKCHECK with docking station, for daily QA on LINACs.
5.8.2 | Measures CAX dose, dose rate, flatness, symmetry, radiation quality.
5.8.3 | Stores data in non-volatile memory for later download to PC.
5.8.4 | Features TCP/IP, USB and RS232 interfaces.
5.8.5 | Docking station and software included
5.9 | Imaging Qa
5.9.1 | EPID QC phantom QC phantom for tests on conventional EPIDs and Flat Panel Imagers.
5.9.2 | Checks linearity, isotropy, noise, low and high contrast resolution.
5.9.3 | Test elements are focused to 100 cm focus distance.
5.9.4 | For 4-25 MV photon beams.
5.9.5 | Incl. carrying case
5.9.6 | kV QC phantom QC phantom for tests on conventional Flat Panel Imagers.
5.9.7 | CathPhan Phantom

## Treatment planning systems
RADIOTHERAPY TREATMENT PLANNING SYSTEMS
Item | TECHNICAL SPECIFICATIONS | Weight
1 | Treatment Planning system
1.1 | The systems must be compatible with the existing unit in the purchasing institution
1.2 | 2D, 3D, IMRT and VMAT Software License
1.3 | Includes the following features:
1.3.1 | Contouring
1.3.2 | CT Simulation
1.3.3 | Fusion 4D
1.3.4 | Support Plan Review and Approval
1.3.5 | DICOM Import and Export Capabilities
1.3.6 | 3D Planning:
1.3.7 | Support for MLC shapes, blocks, apertures, wedges and electron applicators
1.3.8 | Support for static and dynamic gantry delivery techniques
1.3.9 | Collapsed Cone Algorithm
1.3.10 | Monte Carlo Algorithm for MLC based planning
2 | IMRT software license for static gantry IMRT planning including:
2.1 | Step and Shoot delivery mode, supported for all MLC Types
2.2 | Dynamic MLC (Sweeping Window) delivery mode, supported for Elekta and Varian MLC Types
2.3 | Dynamic conformal Arc with MLC shaping based on the view of the target. Supported for all MLC Types
2.4 | VMAT software license for volume-modulated arc planning including:
2.5 | Fully dynamic gantry and MLC movement
2.6 | Variable dose rate delivery, controlled by varying gantry speed, leaf speed, and dose rate
2.7 | Printer compatible with treatment planning system.
3 | Minimum PC Specification
3.1 | Hardware Package consisting of the following components:
3.2 | Please note that Supplier must provide the latest technology of computer hardware available at time of shipping.
3.2.1 | Base PC - HP Z8 G4 (Dual Processor)
3.2.2 | Processor (x2) - Xeon 14C
3.2.3 | Graphics - NVIDIA Quadro P2000 5.0GB
3.2.4 | Hard drive (x2) Int. SATA SSD 1TB
3.2.5 | Keyboard & Mouse
4 | Contouring systems
4.1 | Software License includes the following features:
4.1.1 | Contouring
4.1.2 | CT Simulation
4.1.3 | Fusion
4.1.4 | 4D Support
4.1.5 | Plan Review and Approval
4.1.6 | DICOM Import and Export Capabilities
4.1.7 | Use with Treatment Planning Systems for dose calculation
5 | Recommended hardware
5.1 | Xeon 4-Core Processor
5.2 | 8GB Memory
5.3 | NVIDIA Video Card
5.4 | 2560 x 1600 Resolution
5.5 | 500GB SATA 6Gb/s Hard Disk
6 | Record and verify system
6.1 | R & V must be Image-Enabled Electronic Charting features required for clinical modules with site licenses.
6.2 | Procedure Management: Advanced procedure management system including complete charge/code capture, audit, and export capability. Site license.
6.3 | IQ Server and SQL. License included. One server is required per department.
6.4 | Image-Guided Treatment Management (IGTM) system that includes support for automated patient set up, image registration and analysis tools to assist the user in evaluating the quality of patient setup based on quantitative comparisons between the reference image and the port film. Site license.
6.5 | Advanced automation and process improvement to deliver greater control, efficiency and customization across your entire oncology workflow. Site license.
6.6 | Import and Export
7 | Oncology Maganement server
7.1 | Server complatitible with data storage for the system used for treatment delivery including record and verify system.
7.2 | Latest model Hardware, software and calculation licences.
7.3 | Server Rack must be included.
8 | POWER SUPPLY
8.1 | The bidder must install adequate electrical power supply for the optimal functionality of the equipment.
8.2 | The mains cable of the unit being quoted for must be the hospital grade type and it must be of an adequate length (minimum length of three (3) meters).
8.3 | The power input must be 220-240V, 50Hz AC. Bidders must ensure that the product quoted for is fitted appropriately with a 15 Amp SABS approved mains plug that is held together by two screws.
8.4 | A 3 phase in / out power supply must be offered. State details.
8.5 | The equipment tendered for must not overload and trip the hospital power in the area where it is installed. Resettable overcurrent breaker/s must be fitted for protection. State details.
8.6 | All available online Uninterrupted Power Supply (UPS system) and line voltage regulators must be included for the system with a support time of 50 to 60 minutes. Attach brochure.
8.7 | The lifetime of UPS must be at least 15 years.
8.8 | An Uninterrupted Power Supply (UPS system) conforming to IS-302 must be offered for the computer system.
8.9 | The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade type and it must be a minimum length of three (3) meters.
8.10. | The mains cable of the unit tendered for must be SABS colour coded. State details.
8.11 | The equipment quoted for must be protected against electromagnetic interference. State details of measures that will be taken to protect the unit against electromagnetic interference.
8.12 | Maintenance free batteries must be offered.

## HDR Brachytherapy Unit
HIGH DOSE RATE (HDR) BRACHYTHERAPY UNIT
Item | TECHNICAL SPECIFICATIONS | Weight
1 | GENERAL DESCRIPTION
1.1 | High dose rate remote afterloader brachytherapy unit with radioactive source
2 | EQUIPMENT DETAILED SPECIFICATIONS
2.1 | Compulsory Technical specifications
2.1.1 | Remote after loading unit
2.1.2 | Iridium (Ir-192) preferred, Cobalt (Co-60), or equivalent high dose rate source
2.1.3 | Successful testing according to ISO 2919:2012-08
2.1.4 | Two dimension and three dimension (CT/MRI based ) treatment planning hardware, software's and license - latest in the market with most recent software version at the time of installation.
2.1.5 | Emergency source container
2.1.6 | Source replacement if Ir=192, to be included on warranty period.
2.2 | Compulsory Functional specifications
2.2.1 | Capable of delivering high dose rate remote afterload treatment.
2.2.2 | Two dimensional and three dimensional treatment planning unit.
2.2.3 | CT/MRI based and Dicom image compatible planning.
2.2.4 | Must interface with all Record & Verify systems for Oncology Management System (OMS) .
2.2.5 | Server based image import/export system which allows connectivity between CT/MRI simulator and HDR planning unit.
2.2.6 | In vivo dosimetry for real time patient monitoring (bladder and rectal doses).
2.3 | Compulsory Treatment delivery accessories : All applicators must be titanium (priced separetly as an option)
2.3.1 | Transfer tubes compatible with user applicators x18.
2.3.2 | Fletcher type, titanium CT/MRI shielded applicator set including intrauterine (including angulation variations) and ovoid set with varying cap diameters per set. X6
2.3.3 | Rectal retractor x6 (compatible with 2.3.2)
2.3.4 | Intrauterine titanium tandem of varying standard lengths x6- compatible with 2.3.2
2.3.5 | CT/MRI titanium vaginal ring applicator set (300, 450and 600), intrauterine tandem (varying lengths) with perineal bar x6
2.3.6 | Rectal retractor x6(compatible with 2.3.5)
2.3.7 | Intrauterine Y type titanium applicator x2
2.3.8 | Vaginal cylinder set with applicator and varying diameter caps including standard accessories x4
2.3.9 | Standard vaginal multichannel applicator with accessories x2
2.3.10 | Surface applicator slabs with minimum 5 channels.
2.3.11 | Universal applicator with x-ray marker plug: channel indicator x4
2.3.12 | Applicator Stabiliser/Jig
2.3.13 | Electronic Mobile Adjustable (x,y,z axis) and electric hand controlled Brachytherapy Radio-Translucent Couch Lithotomy positioning accessories. Minimum weight capacity of 200kg. Must include patient leg support during procedure.
2.3.14 | Oesophageal applicator set varying tube dimensions as per supplier availability x1
2.3.15 | Complete high dose rate prostate brachytherapy solution x1
2.3.16 | Consumable prostate Interstitial tube on contract
2.3.17 | Sterilization Oven or alternate system available from the appointed supplier.
2.4 | Compulsory Dosimetry and Patient Verification equipment
2.4.1 | The following equipment are essential for ensuring quality and safety of treatment and must be included upon delivery:
2.4.2 | Room dose monitoring detector unit with an external electric buzzer, 2 way Intercom system, two in- room zoomable cameras with external monitors (x2).
2.4.3 | Standard absolute dosimetry calibration kit with well ionisation chamber, electrometer, thermometer and barometer with calibration certificate.
2.4.4 | Radiation warning light system (inside and outside the room).
2.4.5 | Check source ruler for daily source position check, with transfer tube.
2.4.6 | Two Boxes of EBT 3 films for Autoradiography.
2.4.7 | Survey meter with external Geiger-Mueller counter.
2.4.8 | Radiation pocket dosimeter
2.4.9 | Independent inroom radiation monitor (ionization chamber)
2.4.10 | Laptop/PC with recent software for Calibration of source activity (Microsoft office, Invivo dosimtery).
2.4.11 | Supply and install a single LED head ceiling mounted theatre light (inclusive on the maintenance warranty) for brachytherapy procedure.
2.5 | Important additional Items to be considered
2.5.1 | Treatment delivery and treatment planning printer
2.5.2 | Data backup solution
2.5.3 | Nasopharynx applicator set.
2.5.4 | Oesophagus applicator set. Size: from smallest to a diameter of 12 mm
2.5.5 | Bronchus applicator set (2 mm diameter)
2.5.6 | Interstitial ring and IU applicator set (CT/MRI)
2.5.7 | Appropriate source insert for well chamber calibration (If CO-60 or X-ray source)
2.6 | POWER SUPPLY
2.6.1 | The bidder must install adequate electrical power supply for the optimal functionality of the equipment.
2.6.2 | The mains cable of the unit being quoted for must be the hospital grade type and it must be of an adequate length (minimum length of three (3) meters).
2.6.3 | The power input must be 220-240V, 50Hz AC. Bidders must ensure that the product quoted for is fitted appropriately with a 15 Amp SABS approved mains plug that is held together by two screws.
2.6.4 | A 3 phase in / out power supply must be offered. State details.
2.6.5 | The equipment tendered for must not overload and trip the hospital power in the area where it is installed. Resettable overcurrent breaker/s must be fitted for protection. State details.
2.6.6 | All available online Uninterrupted Power Supply (UPS system) and line voltage regulators must be included for the system with a support time of 50 to 60 minutes. Attach brochure.
2.6.7 | The lifetime of UPS must be at least 15 years.
2.6.8 | An Uninterrupted Power Supply (UPS system) conforming to IS-302 must be offered for the computer system.
2.6.9 | The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade type and it must be a minimum length of three (3) meters.
2.6.10 | The mains cable of the unit tendered for must be SABS colour coded. State details.
2.6.11 | The equipment quoted for must be protected against electromagnetic interference. State details of measures that will be taken to protect the unit against electromagnetic interference.
2.6.12 | Maintenance free batteries must be offered.
2.7 | Brachytherapy C-Arm Table (optional extra)
2.7.1 | Minimum Dimensions: 214.6 x 64.8 cm with OR accessory rails
2.7.2 | Minimum Tabletop: 134.6 x 53.3 cm; 213.4 cm with extension
2.7.3 | Minimum Lightweight Extension (For patient transfer only, not for imaging.): 67.3 x 53.3 cm
2.7.4 | Minimum Radiolucent Extension: 78.7 x 64.8 cm with OR accessory rails
2.7.5 | Minimum Radiolucent Area: 137.2 x 44.5 cm
2.7.6 | Tabletop Material: Three Fiberesin™ sections – each with individual 5cm mattress
2.7.7 | Mattress: Individual, seamless, thick (5 cm), one for each section
2.7.8 | Motions:
2.7.9 | Height Adjustable: 86.4 to 110.5 cm
2.7.10 | Tabletop X Motion (head-to-toe): atleast 50 cm
2.7.11 | Tabletop Y Motion (side-to-side): atleast 25 cm
2.7.12 | Trendelenburg: 0° to 20°
2.7.13 | Reverse Trendelenburg: 0° to 20°
2.7.14 | Lateral Roll: 0° to 20°
2.7.15 | Fowler Positioning: Adjusts from 0° to 80° with touch handle lock/unlock controls
2.7.16 | Hand Control: activates all motions
2.7.17 | Foot Control: activates all motions
2.7.18 | Optional:
2.7.19 | X-Ray Grid: Removable X-Ray grid, 103 line, 8:1 ratio
2.7.20 | Accessory Rails: Standard OR accessory rails 2.86 x .95 cm on table (rails also on table imaging extension)
2.7.21 | Patient Restraints: Two body straps, two ankle straps
2.7.22 | Finish: Powder coat
2.7.23 | Gynaecological position or Lithotomy position: 220 kg patient capacity and use with radiolucent extension.
2.7.24 | Weight tested to four times the patient load rating.
2.7.25 | Power: 115 VAC or 230 VAC and Battery. Battery automatically charges when table is plugged in. An additional battery can be charged with optional wall-mounted charger.
2.7.26 | Certifications: ETL and cETL Listed to UL 60601-1®, CAN/CSA© C22.2

## 64 Slices CT Simulator
64 Slices CT Simulator
Item | Technical Specifications | Weight
System should be compatible with or allow for inclusion of surface guided and/or image guided/adaptive radiotherapy systems.
1 | The system must be a large Bore CTSim for radiotherapy use (at least 80 cm)
1.1 | The system must have 64 Slices, producing 64 Slices per 360 degrees rotation
1.2 | The unit must be low voltage slip-ring designed gantry technology.
1.3 | The unit must support whole body CT Scanning.
1.4 | The unit must have a minimum scan time of 0.4 second or less for 360 degrees rotation
1.5 | The unit must have a gantry aperture size of at least 800mm or larger.
1.6 | Angulations must at least + 30degrees and - 30 degrees.
1.7 | The Unit must be able to scan either head first or feet first.
1.8 | The unit must have a minimum field of view at least 500mm or larger.
1.9 | The unit must have control panels on either side (left and right) of table on the gantry to adjust gantry movements.
1.1 | The system must have triple laser lights for accurate patient positioning at setup plane. (Accuracy of 1mm)
1.11 | Control of gantry movements from operator console must be possible.
1.12 | The system shall have an Intercom between gantry and operator console.
1.13 | Electronic patient breathing instructions in multiple languages must be possible.
1.14 | Emergency stop switches to be available on gantry as well as on the console.
2 | PATIENT SUPPORT SYSTEM (PATIENT TABLE)
2.1 | Support patients 220kg or more.
2.2 | Lowest table height must be 550mm or less
2.3 | The table length must be 2 meters or more
2.4 | Longitudinal table top travel must not be less than 1700mm for whole body coverage.
2.5 | Must have motorized patient support movements (longitudinal and vertical movements)
2.6 | A metal-free table top over scanner range allowing patient scanning without re-positioning patient must be provided.
2.7 | Control of table movements from operator console and gantry must be available.
2.8 | Motorized removal and re-indexing of table top without repositioning patient scanning for emergency access to patient must be available.
2.9 | Patient positioning accessories required to perform the procedures specified, including headrest support, axial and coronal head rest support, table leg extender, security straps, infant immobiliser, arm support, knee support and immobilizing straps must be provided
2.1 | Emergency override to enable manual removal of patients in case of power failure or emergency must be included
3 | DETECTOR ACQUISITION SYSTEM
3.1 | Solid-state detector technology using low-dose and high resolution acquisitions.
3.2 | Multi-slice detector technology for 64-slice acquisitions per 360 degree gantry rotation.
3.3 | Multi-Slice detector with 1200 views per second sampling rates or better.
3.4 | Pre-body collimator must be available.
3.5 | The ability to select 64-slice acquisition with 0.6mm Slice thickness or thinner. State number of channels
3.6 | All calibration and service phantoms with wall brackets must be included
4 | X-RAY GENERATOR
4.1 | Modern high frequency generator allowing at least between 80kV to 135kV and 500mA or better.
4.2 | To support at least 200 scans in 100 seconds using a 0.5 second scan time or better during the spiral scan mode.
4.3 | The unit must be able to scan continuously a length of 1.7m high resolution without cooling time.
4.4 | The unit must support repeat scanning with minimal wait time.
4.5 | The unit must have at least 3 variable kV setting between 80kV (or less) and 135kV (or more).
4.6 | The unit must have variable mA settings up to at least 500mA or better.
4.7 | State maximum mA allowed for each kV.
4.8 | Generator Power output of at least 60kW or higher. State actual power output.
5 | X-RAY TUBE
5.1 | The unit must have a high performance rotating anode tube with dual focal spot.
5.2 | The unit must have a total Filtration of 3.5mm AI equivalent or greater.
5.3 | Anode heat dissipation rate not less than 1300kHU/minute.
5.4 | Cooling performance curves required.
5.5 | The unit must have a heat storage capacity of at least 7MHU.
5.6 | Expected tube life. State in years and slices
6 | OPERATOR CONSOLE
6.1 | The operator console must have a display and entry of patient data pre-planning
6.2 | The console must have comprehensive adult and paediatric anatomical programs/protocols.
6.3 | Must at least have 100 pre-programmed scan protocols required with option to add user defined protocols.
6.4 | One operator console and two independent post -processing workstations must be included in the bid.
6.5 | The operator console to include a high speed computer which will support multi-tasking operation to allow simultaneous scanning reconstruction, viewing, archiving, processing and filming operation.
6.6 | Image review system must be able to display, analyse, process and manage acquired imaging data such as:
6.6.1 | - Zoom
6.6.2 | - Grey scale (colour display) adjustment
6.6.3 | - Measurement of distance and angle
6.6.4 | - Rotation
6.6.5 | - Grid
6.6.6 | - Annotation
6.6.7 | - Cine display
6.6.8 | Facilities must be available to perform basic CT examination of all parts of the body, including reconstruction, review and printing.
6.7 | All post processing software must be available on all three workstations
6.8 | Compatibility: To comply to DICOM 3 standards. Please provide DICOM 3 conformance statement.
6.9 | On line storage: - To store at least 800 Tera Bytes uncompressed images of 512 x 512 images.
6.1 | Archival storage by means of an DVD recording system must be provided
6.11 | X-Ray tube heat loading display is required on console
6.12 | Print facility to laser imager must be supported from all CT workstations.
3 | Chairs: wheeled, height adjustable, swivel for the operators must be supplied. (Quantity 4)
6.14 | All high end post processing software must be on both the CT console and Doctors console. (The user must be able to either configure the post processing software on CT console or Doctors console depending on workflow requirements.)
7 | COMPUTER SYSTEM / RECONSTRUCTION
7.1 | At least a 512 x 512 reconstruction matrix is required.
Details of matrix sizes available to be stated
7.2 | Real time reconstruction for scout-view required.
7.3 | A 0.1 second or faster reconstruction time is required per 512 x 512 reconstructed image
7.4 | 3 - D re-construction must be possible
8 | IMAGE STORAGE/Export
8.1 | Large image archive to store raw scan-data of at least 1000 GB as well as reconstructed data required. State size
8.2 | At least 3600 scan data sets are to be stored. State storage size.
8.3 | The Unit must be able to burn images onto CD & DVD. (200 CDs must be included)
8.4 | Dicom Images must be exported to Radiotherapy treatment planning system (¦Supplier must setup the export)
9 | IMAGE VIEWING MODES
9.1 | Comprehensive image acquisition, viewing, processing and filming modes required must be provided and supported by the systems.
9.2 | Noise reduction techniques must be available.
10 | SCAN MODES
10.1 | Scout-view mode for scan planning must be provided.
10.2 | Single scan mode must be provided.
10.3 | Scanning with serial reconstruction must be provided.
10.4 | Scanning with parallel reconstruction must be provided.
10.5 | Dynamic scan mode must be provided.
11 | SPIRAL MODE PERFORMANCE
11.1 | Tilted spiral scan mode with 64 slice and other combination for thin 0,6mm slice acquisition being acquired per gantry rotation must be provided.
11.2 | 10 scan plans per scout view required.
11.3 | The Unit must provide a variable pitch setting for spiral scanning mode.
11.4 | Single spiral sequence must be provided.
11.5 | Multiple spiral sequence must be provided.
11.6 | Combined conventional and spiral sequence must be provided.
11.7 | Tilted spiral mode must be possible.
11.8 | Real time reconstruction and display of images during the spiral mode must be provided at a viewing rate of at 16 images per second or better.
11.9 | For accurate contrast examinations, a bolus-timing / tracking package must be included.
11.1 | Isotropic volume imaging is required. State voxel sizes used for isotropic volume imaging.
11.11 | Standard mode spatial resolution of 15 lp/cm at cut-off or better is required.
State high contrast spatial resolution in line pairs per cm.
11.12 | High-resolution mode spatial resolution of 21 lp/cm at cut-off or better is required.
State high contrast spatial resolution in line pairs per cm.
11.13 | 4.0mm at 0.3% low-contrast resolution must be possible at 90mAs, 10mm slice thickness, 120kV.
State resolution and scan technique used to obtain measurement. "
11.15 | International CATPHAN performance phantom must be provided for measurements. Supply details used for measurements.
12 | DOSE REDUCTION TECHNIQUE
12.1 | Current dose reduction techniques used during patient scanning to reduce the total applied patient dose must be included.
12.2 | The applied patient dose must be quantified and displayed on the operators console for the patient examination.
12.3 | Bidders must include current dose reduction technology specifically to the following:
12.3.1 | - Gantry design.
12.3.2 | - Choice of scan-times available.
12.3.3 | - Detector design
12.3.4 | - Use of 64 x 0.6mm (or smaller) imaging techniques.
12.3.5 | - Slice profile used
12.3.6 | - Real-time exposure control during scanning
12.3.7 | - Variable pitch settings
12.3.8 | - Beam collimation and X-ray tube filtration
12.3.9 | - Adult and paediatric scanning exposure techniques/protocols.
12.3.10 | - Carbon fibre patient support table top attenuation
12.3.11 | - Accurate contrast timing.
12.3.12 | - Real-time spiral scanning display
12.3.13 | - Calculated dose display on operators console monitor.
12.3.14 | Software image quality correction that allows reconstructed images to be improved without having to rescan the patient with additional scans or scans with a higher applied dose.
12.3.15 | - For shoulders and pelvis with streak reduction techniques
12.3.16 | - Posterior fossa with image fusion techniques.
12.3.17 | - Abdomen with motion detection techniques.
12.4 | Any other techniques must be detailed.
12.5 | Supporting documentation indicating applied dose reduction must be included.
13 | DICOM 3 COMPATIBILITY
13.1 | The CT Scanner must support image & data linkage via network including all additional workstations and to laser camera. All DICOM software must be included in the cost. It must have the ability to be easily linked to any PACS and RIS system. It must have ability to be linked to ADSL/ISDN line or any other tele-radiology system
13.2 | All hardware & software relating to DICOM compatibility, including licences, must be supplied.
13.3 | The Laser camera to support DICOM compatible printing on 35 x 43 cm Laser film.
13.4 | Links to additional workstation and Laser camera systems is required through DICOM compatibility.
13.5 | DICOM Conformance statements on the various service classes per modality to be provided.
14 | THREE INDEPENDENT WORKSTATIONS ( One for Doctors reporting and two for operators control rooms)
14.1 | A WINDOWS Operating computer system is required with at least 4 GB RAM,
800 GB hard drive, network card, suitable graphics card and at least a 48cm flat panel LCD monitor (supporting at least 3 Megapixel resolution) is required.
14.2 | Provide modern user-interface with high speed-processing and storage facilities as per operators console.
14.3 | 2D, 3D, surface rendering, volume rendering, image sculpting, Maximum Intensity Projections, Minimum Intensity Projections, MPR and curved MPR and other image processing tools must be supported by system.
14.4 | All post processing software must be available on all workstations.
14.5 | Image manipulation and arithmetic must be available
14.6 | Archival storage using TIFF, JPEG, DICOM or similar format by means of a DVD must be included. (200 discs must be included)."
14.7 | Print facility to laser imager must be supported.
15 | DICOM COMPATIBLE LASER CAMERA
15.1 | DICOM Compatible networked laser camera must be provided.( 455 DPI or better, for printing on 35 x 43cm film size is required).
15.2 | 35 x 43cm laser film to be used
15.3 | Daylight loading film magazines to be used
15.4 | Linkage via network to additional workstation & CT Console must be provided.
15.5 | At least 100 films per hour printing to be possible
15.6 | Laser printing using dry printing technology required. Details of the unit to be offered.
15.7 | 35 x 43 cm film to be included in the bid price and delivered with the unit (500 sheets)
15.8 | UPS support must be included with at least 20 minutes back – up.
15.9 | All necessary hardware, software and licences must be included.
16 | SOFTWARE PACKAGES AND CLINICAL APPLICATIONS
16.1 | All packages required to perform the clinical applications as listed in SECTIONS 10.11 and 12. must be included.
16.2 | All inclusive Vascular, Oncology and orthopaedic clinical applications must be standard and included in the package. All other clinical applications as per list below must be available, however these must be quoted has optional:
16.3.1 | Virtual colonoscopy package with CO2 insufflator.
16.3.2 | CT Fluoroscopy for drainages, FNA biopsies, etc. is required, including the ability to interactively change table height, longitudinal table position and gantry angulation on remote hand control with ceiling mounted LCD monitor.
16.3.3 | Brain perfusion package.
16.3.4 | Cardiac package including patient monitor, ECG gated scanning, ECG gated reconstruction and full ventricle and coronary evaluation software.
16.3.5 | Neuro DSA software with total bone removal software.
16.4 | Any clinical application not listed in 6.2. but is offered by the manufacturer, must be quoted separately as an option.
17 | CT CONTRAST INJECTOR
17.1 | A modern dual head microprocessor controlled contrast injector is required for use with disposable syringes and will be used for all contrast enhanced examinations including cardiac and bolus chasing. (Please supply 200 units of each of the Consumables i.e. Syringes, Connecting Tubes, Y- Adaptor Tube with Dual check valve)
17.2 | The consumables for the injector mentioned above must readily available in South Africa
18 | POWER SUPPLY
18.1 | The bidder must install adequate electrical power supply for the optimal functionality of the equipment.
18.2 | The mains cable of the unit being quoted for must be the hospital grade type and it must be of an adequate length (minimum length of three (3) meters).
18.3 | The power input must be 220-240V, 50Hz AC. Bidders must ensure that the product quoted for is fitted appropriately with a 15 Amp SABS approved mains plug that is held together by two screws.
18.4 | A 3 phase in / out power supply must be offered. State details.
18.5 | The equipment tendered for must not overload and trip the hospital power in the area where it is installed. Resettable overcurrent breaker/s must be fitted for protection. State details.
18.6 | All available online Uninterrupted Power Supply (UPS system) and line voltage regulators must be included for the system with a support time of 50 to 60 minutes. Attach brochure.
18.7 | The lifetime of UPS must be at least 15 years.
18.8 | An Uninterrupted Power Supply (UPS system) conforming to IS-302 must be offered for the computer system.
18.9 | The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade type and it must be a minimum length of three (3) meters.
18.10. | The mains cable of the unit tendered for must be SABS colour coded. State details.
18.11 | The equipment quoted for must be protected against electromagnetic interference. State details of measures that will be taken to protect the unit against electromagnetic interference.
18.12 | Maintenance free batteries must be offered.
20 | SUPPORTING DOCUMENTATION
20.1 | Sales licence issued in terms of Hazardous Substance Act must be submitted.
20.2 | All supporting documentation as requested in the various sections to be included.
20.3 | All original system brochures, product specifications and application notes to be supplied.
20.4 | DICOM 3 conformance statements to be included.
21 | CUSTOMER SUPPORT DETAILS
21.1 | The application training for radiologists and radiographers by a fully trained application specialist to include at least the following:
21.2 | CV of trainer to be supplied with the bid.
21.3 | An initial period of at least one week to be included patient scanning of multiple clinical problems.
21.4 | At least 4X 1week follow up periods during the first 12 months.
21.5 | Training in the specialized packages to be provided (in addition to the above training) Training schedule to be included. (to be discussed with End User)"
21 | STANDARDS AND SAFETY
21.1 | Should be FDA or CE approved product
21.2 | Electrical safety conforms to standards for electrical safety IEC-60601 / IS-13450
21.3 | The equipment/system must be approved and licensed by the Directorate Radiation Control of DoH: South Africa
21.4 | A copy of a valid license issued in terms of the Hazardous Substance Act, Act No 15 of 1973 must be submitted with the tender. Failure to submit such a valid license may result in a tender not being considered.
21.5 | Acceptance testing must be performed as part of the installation. (Failure to do will lead to failure to pay)
21.7 | Lead Glass window to be installed as per room requirement
21.8 | Safety aspects of Radiation dosage leakage should be spelt out

## 128 Slices CT Simulator
128 Slices CT Simulator
Item | Technical Specifications | Weight
System should be compatible with or allow for inclusion of surface guided and/or image guided/adaptive radiotherapy systems.
1 | The system must be a large Bore CTSim for radiotherapy use (at least 80 cm)
1.1 | The system must have 128 Slices, producing 128 Slices per 360 degrees rotation
1.2 | The unit must be low voltage slip-ring designed gantry technology.
1.3 | The unit must support whole body CT Scanning.
1.4 | The unit must have a minimum scan time of 0.4 second or less for 360 degrees rotation
1.5 | The unit must have a gantry aperture size of at least 800mm or larger.
1.6 | Angulations must at least + 30degrees and - 30 degrees.
1.7 | The Unit must be able to scan either head first or feet first.
1.8 | The unit must have a minimum field of view at least 500mm or larger.
1.9 | The unit must have control panels on either side (left and right) of table on the gantry to adjust gantry movements.
1.1 | The system must have triple laser lights for accurate patient positioning at setup plane. (Accuracy of 1mm)
1.11 | Control of gantry movements from operator console must be possible.
1.12 | The system shall have an Intercom between gantry and operator console.
1.13 | Electronic patient breathing instructions in multiple languages must be possible.
1.14 | Emergency stop switches to be available on gantry as well as on the console.
2 | PATIENT SUPPORT SYSTEM (PATIENT TABLE)
2.1 | Support patients 220kg or more.
2.2 | Lowest table height must be 550mm or less
2.3 | The table length must be 2 meters or more
2.4 | Longitudinal table top travel must not be less than 1700mm for whole body coverage.
2.5 | Must have motorized patient support movements (longitudinal and vertical movements)
2.6 | A metal-free table top over scanner range allowing patient scanning without re-positioning patient must be provided.
2.7 | Control of table movements from operator console and gantry must be available.
2.8 | Motorized removal and re-indexing of table top without repositioning patient scanning for emergency access to patient must be available.
2.9 | Patient positioning accessories required to perform the procedures specified, including headrest support, axial and coronal head rest support, table leg extender, security straps, infant immobiliser, arm support, knee support and immobilizing straps must be provided
2.1 | Emergency override to enable manual removal of patients in case of power failure or emergency must be included
3 | DETECTOR ACQUISITION SYSTEM
3.1 | Solid-state detector technology using low-dose and high resolution acquisitions.
3.2 | Multi-slice detector technology for 64-slice acquisitions per 360 degree gantry rotation.
3.3 | Multi-Slice detector with 1200 views per second sampling rates or better.
3.4 | Pre-body collimator must be available.
3.5 | The ability to select 64-slice acquisition with 0.6mm Slice thickness or thinner. State number of channels
3.6 | All calibration and service phantoms with wall brackets must be included
4 | X-RAY GENERATOR
4.1 | Modern high frequency generator allowing at least between 80kV to 135kV and 500mA or better.
4.2 | To support at least 200 scans in 100 seconds using a 0.5 second scan time or better during the spiral scan mode.
4.3 | The unit must be able to scan continuously a length of 1.7m high resolution without cooling time.
4.4 | The unit must support repeat scanning with minimal wait time.
4.5 | The unit must have at least 3 variable kV setting between 80kV (or less) and 135kV (or more).
4.6 | The unit must have variable mA settings up to at least 500mA or better.
4.7 | State maximum mA allowed for each kV.
4.8 | Generator Power output of at least 60kW or higher. State actual power output.
5 | X-RAY TUBE
5.1 | The unit must have a high performance rotating anode tube with dual focal spot.
5.2 | The unit must have a total Filtration of 3.5mm AI equivalent or greater.
5.3 | Anode heat dissipation rate not less than 1300kHU/minute.
5.4 | Cooling performance curves required.
5.5 | The unit must have a heat storage capacity of at least 7MHU.
5.6 | Expected tube life. State in years and slices
6 | OPERATOR CONSOLE
6.1 | The operator console must have a display and entry of patient data pre-planning
6.2 | The console must have comprehensive adult and paediatric anatomical programs/protocols.
6.3 | Must at least have 100 pre-programmed scan protocols required with option to add user defined protocols.
6.4 | One operator console and two independent post -processing workstations must be included in the bid.
6.5 | The operator console to include a high speed computer which will support multi-tasking operation to allow simultaneous scanning reconstruction, viewing, archiving, processing and filming operation.
6.6 | Image review system must be able to display, analyse, process and manage acquired imaging data such as:
6.6.1 | - Zoom
6.6.2 | - Grey scale (colour display) adjustment
6.6.3 | - Measurement of distance and angle
6.6.4 | - Rotation
6.6.5 | - Grid
6.6.6 | - Annotation
6.6.7 | - Cine display
6.6.8 | Facilities must be available to perform basic CT examination of all parts of the body, including reconstruction, review and printing.
6.7 | All post processing software must be available on all three workstations
6.8 | Compatibility: To comply to DICOM 3 standards. Please provide DICOM 3 conformance statement.
6.9 | On line storage: - To store at least 800 Tera Bytes uncompressed images of 512 x 512 images.
6.1 | Archival storage by means of an DVD recording system must be provided
6.11 | X-Ray tube heat loading display is required on console
6.12 | Print facility to laser imager must be supported from all CT workstations.
3 | Chairs: wheeled, height adjustable, swivel for the operators must be supplied. (Quantity 4)
6.14 | All high end post processing software must be on both the CT console and Doctors console. (The user must be able to either configure the post processing software on CT console or Doctors console depending on workflow requirements.)
7 | COMPUTER SYSTEM / RECONSTRUCTION
7.1 | At least a 512 x 512 reconstruction matrix is required.
Details of matrix sizes available to be stated
7.2 | Real time reconstruction for scout-view required.
7.3 | A 0.1 second or faster reconstruction time is required per 512 x 512 reconstructed image
7.4 | 3 - D re-construction must be possible
8 | IMAGE STORAGE/Export
8.1 | Large image archive to store raw scan-data of at least 1000 GB as well as reconstructed data required. State size
8.2 | At least 3600 scan data sets are to be stored. State storage size.
8.3 | The Unit must be able to burn images onto CD & DVD. (200 CDs must be included)
8.4 | Dicom Images must be exported to Radiotherapy treatment planning system (¦Supplier must setup the export)
9 | IMAGE VIEWING MODES
9.1 | Comprehensive image acquisition, viewing, processing and filming modes required must be provided and supported by the systems.
9.2 | Noise reduction techniques must be available.
10 | SCAN MODES
10.1 | Scout-view mode for scan planning must be provided.
10.2 | Single scan mode must be provided.
10.3 | Scanning with serial reconstruction must be provided.
10.4 | Scanning with parallel reconstruction must be provided.
10.5 | Dynamic scan mode must be provided.
11 | SPIRAL MODE PERFORMANCE
11.1 | Tilted spiral scan mode with 64 slice and other combination for thin 0,6mm slice acquisition being acquired per gantry rotation must be provided.
11.2 | 10 scan plans per scout view required.
11.3 | The Unit must provide a variable pitch setting for spiral scanning mode.
11.4 | Single spiral sequence must be provided.
11.5 | Multiple spiral sequence must be provided.
11.6 | Combined conventional and spiral sequence must be provided.
11.7 | Tilted spiral mode must be possible.
11.8 | Real time reconstruction and display of images during the spiral mode must be provided at a viewing rate of at 16 images per second or better.
11.9 | For accurate contrast examinations, a bolus-timing / tracking package must be included.
11.1 | Isotropic volume imaging is required. State voxel sizes used for isotropic volume imaging.
11.11 | Standard mode spatial resolution of 15 lp/cm at cut-off or better is required.
State high contrast spatial resolution in line pairs per cm.
11.12 | High-resolution mode spatial resolution of 21 lp/cm at cut-off or better is required.
State high contrast spatial resolution in line pairs per cm.
11.13 | 4.0mm at 0.3% low-contrast resolution must be possible at 90mAs, 10mm slice thickness, 120kV.
State resolution and scan technique used to obtain measurement. "
11.15 | International CATPHAN performance phantom must be provided for measurements. Supply details used for measurements.
12 | DOSE REDUCTION TECHNIQUE
12.1 | Current dose reduction techniques used during patient scanning to reduce the total applied patient dose must be included.
12.2 | The applied patient dose must be quantified and displayed on the operators console for the patient examination.
12.3 | Bidders must include current dose reduction technology specifically to the following:
12.3.1 | - Gantry design.
12.3.2 | - Choice of scan-times available.
12.3.3 | - Detector design
12.3.4 | - Use of 64 x 0.6mm (or smaller) imaging techniques.
12.3.5 | - Slice profile used
12.3.6 | - Real-time exposure control during scanning
12.3.7 | - Variable pitch settings
12.3.8 | - Beam collimation and X-ray tube filtration
12.3.9 | - Adult and paediatric scanning exposure techniques/protocols.
12.3.10 | - Carbon fibre patient support table top attenuation
12.3.11 | - Accurate contrast timing.
12.3.12 | - Real-time spiral scanning display
12.3.13 | - Calculated dose display on operators console monitor.
12.3.14 | Software image quality correction that allows reconstructed images to be improved without having to rescan the patient with additional scans or scans with a higher applied dose.
12.3.15 | - For shoulders and pelvis with streak reduction techniques
12.3.16 | - Posterior fossa with image fusion techniques.
12.3.17 | - Abdomen with motion detection techniques.
12.4 | Any other techniques must be detailed.
12.5 | Supporting documentation indicating applied dose reduction must be included.
13 | DICOM 3 COMPATIBILITY
13.1 | The CT Scanner must support image & data linkage via network including all additional workstations and to laser camera. All DICOM software must be included in the cost. It must have the ability to be easily linked to any PACS and RIS system. It must have ability to be linked to ADSL/ISDN line or any other tele-radiology system
13.2 | All hardware & software relating to DICOM compatibility, including licences, must be supplied.
13.3 | The Laser camera to support DICOM compatible printing on 35 x 43 cm Laser film.
13.4 | Links to additional workstation and Laser camera systems is required through DICOM compatibility.
13.5 | DICOM Conformance statements on the various service classes per modality to be provided.
14 | THREE INDEPENDENT WORKSTATIONS ( One for Doctors reporting and two for operators control rooms)
14.1 | A WINDOWS Operating computer system is required with at least 4 GB RAM,
800 GB hard drive, network card, suitable graphics card and at least a 48cm flat panel LCD monitor (supporting at least 3 Megapixel resolution) is required.
14.2 | Provide modern user-interface with high speed-processing and storage facilities as per operators console.
14.3 | 2D, 3D, surface rendering, volume rendering, image sculpting, Maximum Intensity Projections, Minimum Intensity Projections, MPR and curved MPR and other image processing tools must be supported by system.
14.4 | All post processing software must be available on all workstations.
14.5 | Image manipulation and arithmetic must be available
14.6 | Archival storage using TIFF, JPEG, DICOM or similar format by means of a DVD must be included. (200 discs must be included)."
14.7 | Print facility to laser imager must be supported.
15 | DICOM COMPATIBLE LASER CAMERA
15.1 | DICOM Compatible networked laser camera must be provided.( 455 DPI or better, for printing on 35 x 43cm film size is required).
15.2 | 35 x 43cm laser film to be used
15.3 | Daylight loading film magazines to be used
15.4 | Linkage via network to additional workstation & CT Console must be provided.
15.5 | At least 100 films per hour printing to be possible
15.6 | Laser printing using dry printing technology required. Details of the unit to be offered.
15.7 | 35 x 43 cm film to be included in the bid price and delivered with the unit (500 sheets)
15.8 | UPS support must be included with at least 20 minutes back – up.
15.9 | All necessary hardware, software and licences must be included.
16 | SOFTWARE PACKAGES AND CLINICAL APPLICATIONS
16.1 | All packages required to perform the clinical applications as listed in SECTIONS 10.11 and 12. must be included.
16.2 | All inclusive Vascular, Oncology and orthopaedic clinical applications must be standard and included in the package. All other clinical applications as per list below must be available, however these must be quoted has optional:
16.3.1 | Virtual colonoscopy package with CO2 insufflator.
16.3.2 | CT Fluoroscopy for drainages, FNA biopsies, etc. is required, including the ability to interactively change table height, longitudinal table position and gantry angulation on remote hand control with ceiling mounted LCD monitor.
16.3.3 | Brain perfusion package.
16.3.4 | Cardiac package including patient monitor, ECG gated scanning, ECG gated reconstruction and full ventricle and coronary evaluation software.
16.3.5 | Neuro DSA software with total bone removal software.
16.4 | Any clinical application not listed in 6.2. but is offered by the manufacturer, must be quoted separately as an option.
17 | CT CONTRAST INJECTOR
17.1 | A modern dual head microprocessor controlled contrast injector is required for use with disposable syringes and will be used for all contrast enhanced examinations including cardiac and bolus chasing. (Please supply 200 units of each of the Consumables i.e. Syringes, Connecting Tubes, Y- Adaptor Tube with Dual check valve)
17.2 | The consumables for the injector mentioned above must readily available in South Africa
18 | POWER SUPPLY
18.1 | The bidder must install adequate electrical power supply for the optimal functionality of the equipment.
18.2 | The mains cable of the unit being quoted for must be the hospital grade type and it must be of an adequate length (minimum length of three (3) meters).
18.3 | The power input must be 220-240V, 50Hz AC. Bidders must ensure that the product quoted for is fitted appropriately with a 15 Amp SABS approved mains plug that is held together by two screws.
18.4 | A 3 phase in / out power supply must be offered. State details.
18.5 | The equipment tendered for must not overload and trip the hospital power in the area where it is installed. Resettable overcurrent breaker/s must be fitted for protection. State details.
18.6 | All available online Uninterrupted Power Supply (UPS system) and line voltage regulators must be included for the system with a support time of 50 to 60 minutes. Attach brochure.
18.7 | The lifetime of UPS must be at least 15 years.
18.8 | An Uninterrupted Power Supply (UPS system) conforming to IS-302 must be offered for the computer system.
18.9 | The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade type and it must be a minimum length of three (3) meters.
18.10. | The mains cable of the unit tendered for must be SABS colour coded. State details.
18.11 | The equipment quoted for must be protected against electromagnetic interference. State details of measures that will be taken to protect the unit against electromagnetic interference.
18.12 | Maintenance free batteries must be offered.
20 | SUPPORTING DOCUMENTATION
20.1 | Sales licence issued in terms of Hazardous Substance Act must be submitted.
20.2 | All supporting documentation as requested in the various sections to be included.
20.3 | All original system brochures, product specifications and application notes to be supplied.
20.4 | DICOM 3 conformance statements to be included.
21 | CUSTOMER SUPPORT DETAILS
21.1 | The application training for radiologists and radiographers by a fully trained application specialist to include at least the following:
21.2 | CV of trainer to be supplied with the bid.
21.3 | An initial period of at least one week to be included patient scanning of multiple clinical problems.
21.4 | At least 4X 1week follow up periods during the first 12 months.
21.5 | Training in the specialized packages to be provided (in addition to the above training) Training schedule to be included. (to be discussed with End User)"
21 | STANDARDS AND SAFETY
21.1 | Should be FDA or CE approved product
21.2 | Electrical safety conforms to standards for electrical safety IEC-60601 / IS-13450
21.3 | The equipment/system must be approved and licensed by the Directorate Radiation Control of DoH: South Africa
21.4 | A copy of a valid license issued in terms of the Hazardous Substance Act, Act No 15 of 1973 must be submitted with the tender. Failure to submit such a valid license may result in a tender not being considered.
21.5 | Acceptance testing must be performed as part of the installation. (Failure to do will lead to failure to pay)
21.7 | Lead Glass window to be installed as per room requirement
21.8 | Safety aspects of Radiation dosage leakage should be spelt out

## 320 Slices CT Simulator
320 Slices CT Simulator
Item | Technical Specifications | Weight
System should be compatible with or allow for inclusion of surface guided and/or image guided/adaptive radiotherapy systems.
1 | The system must be a large Bore CTSim for radiotherapy use (at least 80 cm)
1.1 | The system must have 320 Slices, producing 320 Slices per 360 degrees rotation
1.2 | The unit must be low voltage slip-ring designed gantry technology.
1.3 | The unit must support whole body CT Scanning.
1.4 | The unit must have a minimum scan time of 0.4 second or less for 360 degrees rotation
1.5 | The unit must have a gantry aperture size of at least 800mm or larger.
1.6 | Angulations must at least + 30degrees and - 30 degrees.
1.7 | The Unit must be able to scan either head first or feet first.
1.8 | The unit must have a minimum field of view at least 500mm or larger.
1.9 | The unit must have control panels on either side (left and right) of table on the gantry to adjust gantry movements.
1.1 | The system must have triple laser lights for accurate patient positioning at setup plane. (Accuracy of 1mm)
1.11 | Control of gantry movements from operator console must be possible.
1.12 | The system shall have an Intercom between gantry and operator console.
1.13 | Electronic patient breathing instructions in multiple languages must be possible.
1.14 | Emergency stop switches to be available on gantry as well as on the console.
2 | PATIENT SUPPORT SYSTEM (PATIENT TABLE)
2.1 | Support patients 220kg or more.
2.2 | Lowest table height must be 550mm or less
2.3 | The table length must be 2 meters or more
2.4 | Longitudinal table top travel must not be less than 1700mm for whole body coverage.
2.5 | Must have motorized patient support movements (longitudinal and vertical movements)
2.6 | A metal-free table top over scanner range allowing patient scanning without re-positioning patient must be provided.
2.7 | Control of table movements from operator console and gantry must be available.
2.8 | Motorized removal and re-indexing of table top without repositioning patient scanning for emergency access to patient must be available.
2.9 | Patient positioning accessories required to perform the procedures specified, including headrest support, axial and coronal head rest support, table leg extender, security straps, infant immobiliser, arm support, knee support and immobilizing straps must be provided
2.1 | Emergency override to enable manual removal of patients in case of power failure or emergency must be included
3 | DETECTOR ACQUISITION SYSTEM
3.1 | Solid-state detector technology using low-dose and high resolution acquisitions.
3.2 | Multi-slice detector technology for 64-slice acquisitions per 360 degree gantry rotation.
3.3 | Multi-Slice detector with 1200 views per second sampling rates or better.
3.4 | Pre-body collimator must be available.
3.5 | The ability to select 64-slice acquisition with 0.6mm Slice thickness or thinner. State number of channels
3.6 | All calibration and service phantoms with wall brackets must be included
4 | X-RAY GENERATOR
4.1 | Modern high frequency generator allowing at least between 80kV to 135kV and 500mA or better.
4.2 | To support at least 200 scans in 100 seconds using a 0.5 second scan time or better during the spiral scan mode.
4.3 | The unit must be able to scan continuously a length of 1.7m high resolution without cooling time.
4.4 | The unit must support repeat scanning with minimal wait time.
4.5 | The unit must have at least 3 variable kV setting between 80kV (or less) and 135kV (or more).
4.6 | The unit must have variable mA settings up to at least 500mA or better.
4.7 | State maximum mA allowed for each kV.
4.8 | Generator Power output of at least 60kW or higher. State actual power output.
5 | X-RAY TUBE
5.1 | The unit must have a high performance rotating anode tube with dual focal spot.
5.2 | The unit must have a total Filtration of 3.5mm AI equivalent or greater.
5.3 | Anode heat dissipation rate not less than 1300kHU/minute.
5.4 | Cooling performance curves required.
5.5 | The unit must have a heat storage capacity of at least 7MHU.
5.6 | Expected tube life. State in years and slices
6 | OPERATOR CONSOLE
6.1 | The operator console must have a display and entry of patient data pre-planning
6.2 | The console must have comprehensive adult and paediatric anatomical programs/protocols.
6.3 | Must at least have 100 pre-programmed scan protocols required with option to add user defined protocols.
6.4 | One operator console and two independent post -processing workstations must be included in the bid.
6.5 | The operator console to include a high speed computer which will support multi-tasking operation to allow simultaneous scanning reconstruction, viewing, archiving, processing and filming operation.
6.6 | Image review system must be able to display, analyse, process and manage acquired imaging data such as:
6.6.1 | - Zoom
6.6.2 | - Grey scale (colour display) adjustment
6.6.3 | - Measurement of distance and angle
6.6.4 | - Rotation
6.6.5 | - Grid
6.6.6 | - Annotation
6.6.7 | - Cine display
6.6.8 | Facilities must be available to perform basic CT examination of all parts of the body, including reconstruction, review and printing.
6.7 | All post processing software must be available on all three workstations
6.8 | Compatibility: To comply to DICOM 3 standards. Please provide DICOM 3 conformance statement.
6.9 | On line storage: - To store at least 800 Tera Bytes uncompressed images of 512 x 512 images.
6.1 | Archival storage by means of an DVD recording system must be provided
6.11 | X-Ray tube heat loading display is required on console
6.12 | Print facility to laser imager must be supported from all CT workstations.
3 | Chairs: wheeled, height adjustable, swivel for the operators must be supplied. (Quantity 4)
6.14 | All high end post processing software must be on both the CT console and Doctors console. (The user must be able to either configure the post processing software on CT console or Doctors console depending on workflow requirements.)
7 | COMPUTER SYSTEM / RECONSTRUCTION
7.1 | At least a 512 x 512 reconstruction matrix is required.
Details of matrix sizes available to be stated
7.2 | Real time reconstruction for scout-view required.
7.3 | A 0.1 second or faster reconstruction time is required per 512 x 512 reconstructed image
7.4 | 3 - D re-construction must be possible
8 | IMAGE STORAGE/Export
8.1 | Large image archive to store raw scan-data of at least 1000 GB as well as reconstructed data required. State size
8.2 | At least 3600 scan data sets are to be stored. State storage size.
8.3 | The Unit must be able to burn images onto CD & DVD. (200 CDs must be included)
8.4 | Dicom Images must be exported to Radiotherapy treatment planning system (¦Supplier must setup the export)
9 | IMAGE VIEWING MODES
9.1 | Comprehensive image acquisition, viewing, processing and filming modes required must be provided and supported by the systems.
9.2 | Noise reduction techniques must be available.
10 | SCAN MODES
10.1 | Scout-view mode for scan planning must be provided.
10.2 | Single scan mode must be provided.
10.3 | Scanning with serial reconstruction must be provided.
10.4 | Scanning with parallel reconstruction must be provided.
10.5 | Dynamic scan mode must be provided.
11 | SPIRAL MODE PERFORMANCE
11.1 | Tilted spiral scan mode with 64 slice and other combination for thin 0,6mm slice acquisition being acquired per gantry rotation must be provided.
11.2 | 10 scan plans per scout view required.
11.3 | The Unit must provide a variable pitch setting for spiral scanning mode.
11.4 | Single spiral sequence must be provided.
11.5 | Multiple spiral sequence must be provided.
11.6 | Combined conventional and spiral sequence must be provided.
11.7 | Tilted spiral mode must be possible.
11.8 | Real time reconstruction and display of images during the spiral mode must be provided at a viewing rate of at 16 images per second or better.
11.9 | For accurate contrast examinations, a bolus-timing / tracking package must be included.
11.1 | Isotropic volume imaging is required. State voxel sizes used for isotropic volume imaging.
11.11 | Standard mode spatial resolution of 15 lp/cm at cut-off or better is required.
State high contrast spatial resolution in line pairs per cm.
11.12 | High-resolution mode spatial resolution of 21 lp/cm at cut-off or better is required.
State high contrast spatial resolution in line pairs per cm.
11.13 | 4.0mm at 0.3% low-contrast resolution must be possible at 90mAs, 10mm slice thickness, 120kV.
State resolution and scan technique used to obtain measurement. "
11.15 | International CATPHAN performance phantom must be provided for measurements. Supply details used for measurements.
12 | DOSE REDUCTION TECHNIQUE
12.1 | Current dose reduction techniques used during patient scanning to reduce the total applied patient dose must be included.
12.2 | The applied patient dose must be quantified and displayed on the operators console for the patient examination.
12.3 | Bidders must include current dose reduction technology specifically to the following:
12.3.1 | - Gantry design.
12.3.2 | - Choice of scan-times available.
12.3.3 | - Detector design
12.3.4 | - Use of 64 x 0.6mm (or smaller) imaging techniques.
12.3.5 | - Slice profile used
12.3.6 | - Real-time exposure control during scanning
12.3.7 | - Variable pitch settings
12.3.8 | - Beam collimation and X-ray tube filtration
12.3.9 | - Adult and paediatric scanning exposure techniques/protocols.
12.3.10 | - Carbon fibre patient support table top attenuation
12.3.11 | - Accurate contrast timing.
12.3.12 | - Real-time spiral scanning display
12.3.13 | - Calculated dose display on operators console monitor.
12.3.14 | Software image quality correction that allows reconstructed images to be improved without having to rescan the patient with additional scans or scans with a higher applied dose.
12.3.15 | - For shoulders and pelvis with streak reduction techniques
12.3.16 | - Posterior fossa with image fusion techniques.
12.3.17 | - Abdomen with motion detection techniques.
12.4 | Any other techniques must be detailed.
12.5 | Supporting documentation indicating applied dose reduction must be included.
13 | DICOM 3 COMPATIBILITY
13.1 | The CT Scanner must support image & data linkage via network including all additional workstations and to laser camera. All DICOM software must be included in the cost. It must have the ability to be easily linked to any PACS and RIS system. It must have ability to be linked to ADSL/ISDN line or any other tele-radiology system
13.2 | All hardware & software relating to DICOM compatibility, including licences, must be supplied.
13.3 | The Laser camera to support DICOM compatible printing on 35 x 43 cm Laser film.
13.4 | Links to additional workstation and Laser camera systems is required through DICOM compatibility.
13.5 | DICOM Conformance statements on the various service classes per modality to be provided.
14 | THREE INDEPENDENT WORKSTATIONS ( One for Doctors reporting and two for operators control rooms)
14.1 | A WINDOWS Operating computer system is required with at least 4 GB RAM,
800 GB hard drive, network card, suitable graphics card and at least a 48cm flat panel LCD monitor (supporting at least 3 Megapixel resolution) is required.
14.2 | Provide modern user-interface with high speed-processing and storage facilities as per operators console.
14.3 | 2D, 3D, surface rendering, volume rendering, image sculpting, Maximum Intensity Projections, Minimum Intensity Projections, MPR and curved MPR and other image processing tools must be supported by system.
14.4 | All post processing software must be available on all workstations.
14.5 | Image manipulation and arithmetic must be available
14.6 | Archival storage using TIFF, JPEG, DICOM or similar format by means of a DVD must be included. (200 discs must be included)."
14.7 | Print facility to laser imager must be supported.
15 | DICOM COMPATIBLE LASER CAMERA
15.1 | DICOM Compatible networked laser camera must be provided.( 455 DPI or better, for printing on 35 x 43cm film size is required).
15.2 | 35 x 43cm laser film to be used
15.3 | Daylight loading film magazines to be used
15.4 | Linkage via network to additional workstation & CT Console must be provided.
15.5 | At least 100 films per hour printing to be possible
15.6 | Laser printing using dry printing technology required. Details of the unit to be offered.
15.7 | 35 x 43 cm film to be included in the bid price and delivered with the unit (500 sheets)
15.8 | UPS support must be included with at least 20 minutes back – up.
15.9 | All necessary hardware, software and licences must be included.
16 | SOFTWARE PACKAGES AND CLINICAL APPLICATIONS
16.1 | All packages required to perform the clinical applications as listed in SECTIONS 10.11 and 12. must be included.
16.2 | All inclusive Vascular, Oncology and orthopaedic clinical applications must be standard and included in the package. All other clinical applications as per list below must be available, however these must be quoted has optional:
16.3.1 | Virtual colonoscopy package with CO2 insufflator.
16.3.2 | CT Fluoroscopy for drainages, FNA biopsies, etc. is required, including the ability to interactively change table height, longitudinal table position and gantry angulation on remote hand control with ceiling mounted LCD monitor.
16.3.3 | Brain perfusion package.
16.3.4 | Cardiac package including patient monitor, ECG gated scanning, ECG gated reconstruction and full ventricle and coronary evaluation software.
16.3.5 | Neuro DSA software with total bone removal software.
16.4 | Any clinical application not listed in 6.2. but is offered by the manufacturer, must be quoted separately as an option.
17 | CT CONTRAST INJECTOR
17.1 | A modern dual head microprocessor controlled contrast injector is required for use with disposable syringes and will be used for all contrast enhanced examinations including cardiac and bolus chasing. (Please supply 200 units of each of the Consumables i.e. Syringes, Connecting Tubes, Y- Adaptor Tube with Dual check valve)
17.2 | The consumables for the injector mentioned above must readily available in South Africa
18 | POWER SUPPLY
18.1 | The bidder must install adequate electrical power supply for the optimal functionality of the equipment.
18.2 | The mains cable of the unit being quoted for must be the hospital grade type and it must be of an adequate length (minimum length of three (3) meters).
18.3 | The power input must be 220-240V, 50Hz AC. Bidders must ensure that the product quoted for is fitted appropriately with a 15 Amp SABS approved mains plug that is held together by two screws.
18.4 | A 3 phase in / out power supply must be offered. State details.
18.5 | The equipment tendered for must not overload and trip the hospital power in the area where it is installed. Resettable overcurrent breaker/s must be fitted for protection. State details.
18.6 | All available online Uninterrupted Power Supply (UPS system) and line voltage regulators must be included for the system with a support time of 50 to 60 minutes. Attach brochure.
18.7 | The lifetime of UPS must be at least 15 years.
18.8 | An Uninterrupted Power Supply (UPS system) conforming to IS-302 must be offered for the computer system.
18.9 | The mains cable of the unit being quoted for must be 15 amp 3 prong hospital grade type and it must be a minimum length of three (3) meters.
18.10. | The mains cable of the unit tendered for must be SABS colour coded. State details.
18.11 | The equipment quoted for must be protected against electromagnetic interference. State details of measures that will be taken to protect the unit against electromagnetic interference.
18.12 | Maintenance free batteries must be offered.
20 | SUPPORTING DOCUMENTATION
20.1 | Sales licence issued in terms of Hazardous Substance Act must be submitted.
20.2 | All supporting documentation as requested in the various sections to be included.
20.3 | All original system brochures, product specifications and application notes to be supplied.
20.4 | DICOM 3 conformance statements to be included.
21 | CUSTOMER SUPPORT DETAILS
21.1 | The application training for radiologists and radiographers by a fully trained application specialist to include at least the following:
21.2 | CV of trainer to be supplied with the bid.
21.3 | An initial period of at least one week to be included patient scanning of multiple clinical problems.
21.4 | At least 4X 1week follow up periods during the first 12 months.
21.5 | Training in the specialized packages to be provided (in addition to the above training) Training schedule to be included. (to be discussed with End User)"
21 | STANDARDS AND SAFETY
21.1 | Should be FDA or CE approved product
21.2 | Electrical safety conforms to standards for electrical safety IEC-60601 / IS-13450
21.3 | The equipment/system must be approved and licensed by the Directorate Radiation Control of DoH: South Africa
21.4 | A copy of a valid license issued in terms of the Hazardous Substance Act, Act No 15 of 1973 must be submitted with the tender. Failure to submit such a valid license may result in a tender not being considered.
21.5 | Acceptance testing must be performed as part of the installation. (Failure to do will lead to failure to pay)
21.7 | Lead Glass window to be installed as per room requirement
21.8 | Safety aspects of Radiation dosage leakage should be spelt out

## Immobilization
IMMOBILIZATION
Item | Technical Specifications | Weight
1 | All immobilization devices must be MRI/CT compatible
2 | Following considerations when designing and selecting immobilization devices for photon therapy:
2.1 | Rigid device that will maintain shape while allowing for patient comfort
2.2 | Uniform, low-density if possible, material (i.e., minimal part-to-part variability) so that it would introduce minimal range perturbation in the beam path;
2.3 | Devices with gradual slope and no sharp edge and conforms to body contours
2.4 | Devices producing minimal imaging artifacts; and
2.5 | Allows for placement of reference marks
2.6 | Avoiding the field passing part of the frame and/or table edge during treatment planning. Beam path should not be obstructed and surface dose not affected
2.7 | Device sizing: Pediatric and adult age groups where appropiate
2.8 | Device requirements may be determined by equipment and capabilities thereof e.g radiosurgery, stereotactic body radiotherapy
3 | Carbon fibre immobilization devices
3.1 | All in one Base plate
3.2 | Standard head base plate for 3 point marks
3.3 | head and shoulder base plate for 5 point markers
3.4 | breastboard
3.5 | tharacic wing board
3.6 | pelvis base plate with knee and foot rest support
4 | Head and neck immobilization devices
4.1 | Various sizes including pediatric and adult
4.2 | Machine technique determined
4.3 | Standard head base plate for 3 point marks
4.4 | Head and shoulder base plate for 5 point markers
4.5 | Specilalised systems including radiosurgery immoblisation devices, must include appropiate head support systems
4 | Head support
4.1 | 6 pieces set
4.2 | include set A, B, C, D, E, F and paediatric sizes
5 | Breast immobilisation system
5.1 | Breast Board, Wing board, prone board or similar devices as compatible
6 | Chest immobilisation thermoplastic casts
7 | Pelvic and abdominal systems
7.1 | Abdominal compression devices, belly board, prone and supine devices
8 | Limb Systems
8.1 | Knee and foot positioning devices, pelvic immobilisation devices, specific devices for limb immobilisation
8.2 | Acrylic prone/hypophysis base plate
9 | Acrylic prone/hypophysis base plate
10 | Carbon fibre Knee rest
11 | Caron fibre foot rest
12 | Table indexer bars
13 | Masks Items (quantities and sizes as per instituion requirements)
13.1 | Water tank/oven for mask preparations
13.2 | Head Masks
13.3 | Head and neck Masks
13.4 | Head and shoulder masks
13.5 | IMRT/VMAT masks for head and neck
14 | Vaccume system
14.1 | Vacuume pump
14.2 | Vaccume bags for small, medium and large body heights
15 | Skin Markers
15.1 | Isocenter field line sheets
15.2 | CT Markers