## Ultrasound entry, B&W
BLACK & WHITE ULTRASOUND UNIT (ENTRY LEVEL)
ITEM | TECHNICAL SPECIFICATION
1 | The specification calls for the supply, delivery, installation and commissioning of basic/portable type ultrasound unit, fixed to a trolley with real‑time scanning and image recording systems. State make and model offered.
2 | To be used for diagnostic investigations of patients.
3 | Depending upon the clinical requirements, trans-abdominal and/or trans-vaginal Gynaecology and Obstetric scanning, cardiac scanning, general abdominal scanning, Urological scanning, and/or Anterior fontanelle scanning may be required.
4 | The offered system shall be a real‑time black and white sector Scanner M‑Mode Cardiac Screening facilities, Thermal Printer and probe as specified.
5 | The final bid price must include the following minimum components:
5.1 | Ultrasound unit as specified. | -
5.2 | Outlets for Video Recording of Images - DICOM | -
5.2.1 | Appropriate software.
5.2.2 | 3.5 MHz convex sector probe.
5.2.3 | Probe connections X 2 | -
5.2.4 | Thermal printer. | -
5.3 | Unit fixed to mobile cart
6 | OPERATOR'S CONSOLE / TROLLEY:
6.1 | The operator's console must contain the following:
6.1.1 | Connections to the probes, thermal printer and all recordings systems.
6.1.2 | All the display monitors, keyboard and variable controls necessary for the operation of the unit.
6.1.3 | It must be possible to connect two (2) probes simultaneously.
7 | DIGITAL MEMORY AND PROCESSOR:
7.1 | The processor must provide for a minimum of 256 shades of grey imaging.
7.2 | A cine-loop memory capacity must be available
7.3 | Tissue harmonic imaging must be possible
7.4 | Supply details of the programmable features of the memory must be at least.
7.4.1 | Real time pan zoom.
7.4.2 | Magnification of frozen image.
7.5 | Image storage up to a minimum of 1000 images must be available. Measurement must be possible on stored images.
7.6 | Semi‑fixed data storage is required for patient and image identification. The number of characters available must be 40:
7.6.1 | Patient data
7.6.2 | Institution name
7.6.3 | Free text
7.7 | The unit must have a freeze‑frame facility that could be operated by a push‑button. There should be no deterioration in brightness or contrast whilst the frozen image is being viewed.
7.8 | The processor should be capable of both pre‑ and post ‑processing. A minimum of an 8 bit processor is required. Details of all such processing must be provided.
7.9 | A standard 625 line video‑output must be provided for transmitting the processed analogue signal to an image recording system.
8 | DISPLAY SYSTEMS:
8.1 | All images should be displayed for identification purposes on a sharp‑focus high-resolution non‑glare monochrome monitor. The monitor of the unit must have a diagonal size of at least 15” (inch) TFT.
8.2 | The instrument should be capable of displaying real‑time B‑Mode and M‑Mode images in grey scale.
8.3 | It must be possible to display combinations of B‑Mode and M‑Mode simultaneously.
8.4 | In addition to the real‑time image display, patient information and scan details should be provided by full alpha‑numeric keyboard, as well as automatically recording the distance between the markers.
9 | MEASURING SYSTEM:
9.1 | An electronic calliper measuring system, calibrated in centimetres range, is required with an accuracy of at least 1mm.
9.2 | The system should provide two independent calliper markers in the range of (6 - 8), superimposed on the displays and capable of being positioned at any point on the image display.
9.3 | The separation between two markers should be directly readable on the margin of the display.
9.4 | The unit must be capable of circumference, volume, ellipse and trace measurements and the answers must be displayed on the margin of the image.
9.5 | Scale markers for M‑Mode Displays should be provided 1cm apart and sub‑divisions should be stated.
9.6 | On M‑Mode there should be 1cm depth markers along the transducer axis and time markers at least every half second.
9.7 | The means of operation of the electronic calliper measuring system must be by track‑ball.
10 | SOFTWARE:
10.1 | Obstetric Software must be available:-
10.2 | To calculate foetal age and foetal weight from caliper measurements
10.3 | The software offered must be able to estimate the date of delivery from the date of the last menstrual period.
10.4 | Percentile graphs and report facilities must be obtainable from recorded data.
10.5 | Any further available obstetrical software must be stated.
10.6 | Cardiac and Vascular Studies software(In Radiology, Obstetrics, Gynaecology and Including Paediatrics)
10.7 | Software must be available for calculations pertaining to blood flow studies in:-
a) Obstetrics and Gynae: umbilical cord, uterine and ovarian pathology,
b) Abdominal and small parts.
c) Vascular: Aorta, neck, deep vein thrombosis, and superficial veins.
Software pertaining to the above-named applications must be available for obtaining the following:-
a) Display of maximum, minimum and mean mode curves with spectral Doppler.
b) Pulsatility and resistive index calculation.
c) Lumen calculations via diameter/area method.
d) Percentage stenosis based on area or diameter calculation.
e) Any other desirable features of the software package must be stated.
11 | PROBES:
11.1 | Bidders must state if exchange probes are available. If so, quote the prices offered on an exchange basis and indicate what guarantee will be offered on these probes.
11.2 | Bidders must include the one probe (3.5 MHz multi frequency curvilinear sector probe and 5MHz ) as part of the offer
11.3 | Details regarding the prices of probe and any other optional extras must be stated in the “schedule of optional extras” at the end of the specification.
12 | TRANSMITTER:
12.1 | The power output should enable tissue interfaces up to a depth of 30cm to be visualised.
13 | RECEIVER:
13.1 | The receiver should have gain controls to compensate for depths up to about 30cm.
13.2 | The bidder to state the overall system gain of the unit offered.
13.3 | Any additional features must be stated and will be considered.
14 | IMAGE RECORDING:
14.1 | Each system must be supplied with a 110 mm black and white thermal printer fixed permanently to the trolley and capable of providing good quality copies of selected frozen images displayed on the monitor screen.
15 | MOBILE CART: The unit must be attached to a sturdy trolley on castors, fitted with brakes.
16 | POWER SUPPLY:
16.1 | The equipment offered shall operate from a standard 220 to 240 Volt A.C. 50Hz mains supply. The bidder must ensure that the product being quoted for is fitted with a 15 Amp sealed plugtop S.A.N.S approved
16.2 | The unit must be supplied with an on-line uninterrupted power supply (U.P.S.) to provide minimum of 50 to 60 minutes mains backup power.
16.3 | The UPS should not adversely affect the image quality.
16.4 | This unit must be included in the bid price and be covered by the guarantee and maintenance plan.
17 | Optional accessories (bidder must list the probes available)
Abdominal

## Ultrasound entry, colour
COLOUR ULTRASOUND UNIT (ENTRY LEVEL)
ITEM | TECHNICAL SPECIFICATIONS
1 | PURPOSE:
1.1 | The system is intended to serve the hospitals for the diagnostic investigations of patients. The unit must be moderately priced and have the ability to produce high quality colour doppler images.
2 | MEDICAL USES:
2.1 | The unit must have the ability to do the following:-
2.1.1 | Produce the highest quality real-time B-Mode scans, with tissue harmonics, of all body organs and any pathology or abnormal phenomena in different shades of grey and colour.
2.1.2 | The system must have a full software package to calculate the various parameters for all ultrasound imaging as specified.
3 | SYSTEM HARDWARE:
3.1 | The system must consist of the following minimum components:
3.1.1 | Operator's console, with all controls and a full alpha-numeric keyboard.
3.1.2 | A digital memory and processor.
3.1.4 | Display monitor as specified.
3.1.5 | A high precision measuring system.
3.1.6 | A recording system as specified in line 22
3.1.7 | Pulsed wave doppler with real-time duplex mode.
4 | SYSTEM SOFTWARE:
4.1 | The unit must have software for the analysis of recorded data in respect of the following investigations:-
4.1.1 | Obstetrics.
4.1.2 | Vascular flow studies in obstetrics, and central and peripheral vascular.
4.1.3 | Paediatrics.
4.1.4 | The following software must be offered as options for future up-grading of the system:
a) Panoramic imaging
b) Free hand 3D
c) Trapezoid imaging
d) State prices of the above options under “schedule of accessories” at the end of the specification.
4.1.5 | Any other software that is presently available and not included on the offer must be stated and may be taken into consideration.
4.1.6 | The bidder must state if any other software will become available in the near future. Details of the software and cost must be listed in the schedule of optional accessories.
4.1.7 | Cardiac software and continuous wave must be quoted for and offered as an option on the schedule of optional accessories for future up-grading of the system.
4.1.8 | All version software upgrades must be supplied free of charge.
4.1.9 | The software for Dicom/PACS compatibility must be offered with worklist
5 | PROBES:
5.1 | PROBE CONNECTION:
5.1.1 | A probe holder for 3 probes and cable support must be supplied.
5.1.2 | It must be possible to connect 3 probes simultaneously
5.2 | PROBES REQUIRED:
The bid price must include the following probe. The bidder must supply details of the probe and of the cost of the probe in the schedule of accessories at the end of the specification.
5.2.1 | 3.5 MHz and 5 MHz multi-frequency convex sector probe.
6 | OPTIONAL PROBES:
6.1 | The following probes should be offered as options and the prices must be stated on the price schedule at the end of the specification.
6.1.1 | Neonatal head 5 – 7.5 MHz or higher.
6.1.2 | Linear 7.5 -10 MHz probe or higher
6.1.3 | Endo vaginal 5 – 7 MHz. or higher
6.1.4 | Cardiac Probe 5 MHz or higher
6.5 | The bidder may offer other probes as optional extras.
6.6 | Details and the cost of the probes offered must be stated on the schedule of optional accessories at the end of the specification.
6.7 | The specified frequencies of the above probes are nominal values: Variations from the central frequency of 0,5MHz each way will be acceptable.
6.8 | Mechanically driven probes, except for 3D and 4D, will not be considered.
7 | ELECTRONIC FOCUSING:
7.1 | Bidders must indicate which of the offered probes are capable of electronic dynamic focusing in which case the number of focusing zones along the depth of the field should be stated.
8 | PULSED DOPPLER:
8.1 | Transducers must be capable of switching (simultaneous) between 2-D, pulsed Doppler, continuous wave doppler and colour flow Doppler imaging in real-time (triplex mode).
9 | REPLACEMENT PROBES:
9.1 | Bidders must state if service exchange probes are available.
9.2 | If available, quote the prices offered on an exchange basis and indicate what guarantee will be offered on these probes.
10 | BIOPSY NEEDLE GUIDES:
10.1 | The bidder must indicate which of the probes quoted for can be fitted with biopsy needle guides.
10.2 | Details of such needle guides should be given as well as quotations for their supply should be stated on the schedule of optional accessories at the end of the specification.
11 | TRANSMITTER AND RECEIVER:
11.1 | The power output should enable tissue interfaces up to a depth of 30cm to be visualized.
11.2 | Variable focusing during transmission must be available.
11.3 | The range of pulse repetition rates available must be stated.
11.4 | Compensation for depth should be accomplished by slide controls on the control panel, the positions of which will indicate the Time Gain Compensation(TGC) slope, or TGC should be achieved by a control that provides for variation in near gain, slope rate and slope position.
11.5 | The frame rate will be assessed during demonstration and will be taken into consideration in the clinical evaluation of the system.
11.6 | Any additional features or offers may be stated and will be considered.
12 | OPERATOR'S CONSOLE:
12.1 | The operator's console must contain the following:-
12.2 | All the electronic equipment.
12.3 | Connections to the probes, footswitches and all recording systems.
12.4 | The facilities for an integrated writable DVD/CD drive must be available for image storage.
12.5 | The console panels must be made of rust-proof material which is resistant to marine air and is of a high quality finish. The colour must be stated.
13 | DIGITAL MEMORY AND PROCESSOR:
13.1 | The memory must use at least 8 bit technology.
13.2 | The cine loop memory capacity must be at least 1000 frames.
14 | DISPLAY:
14.1 | The instrument must be capable of displaying:-
14.2 | Real-time B-Mode scans in shades of grey - minimum 256.
14.3 | Real-time B-mode, pulsed wave, and colour Doppler images in triplex mode.
15 | MONITOR:
15.1 | The display should be on a high resolution colour monitor of at least 38cm (15”) diagonal.
15.2 | The monitor must have swivel and tilt facilities to enable optimal visualization from different angles.
16 | ANNOTATION:
16.1 | The screen must display the following information:-
16.2 | Patient's name and I.D. number.
16.3 | Date and time.
16.4 | Annotation pointer.
16.5 | All scanning parameters.
16.6 | Results of calculations.
16.7 | The number of digits available for annotation must be at least 40 or above
17 | AUDIO DOPPLER:
17.1 | The Doppler shift frequency must have an audible output that can be heard via speakers.
18 | SOFTWARE:
18.1 | Obstetric Software must be available:-
18.2 | To calculate foetal age and foetal weight from caliper measurements
18.3 | The software offered must be able to estimate the date of delivery from the date of the last menstrual period.
18.4 | Percentile graphs and report facilities must be obtainable from recorded data.
18.5 | Any further available obstetrical software must be stated.
18.6 | Cardiac and Vascular Studies software(In Radiology, Obstetrics, Gynaecology and Including Paediatrics)
18.7 | Software must be available for calculations pertaining to blood flow studies in:-
a) Obstetrics and Gynae: umbilical cord, uterine and ovarian pathology,
b) Abdominal and small parts.
c) Vascular: Aorta, neck, deep vein thrombosis, and superficial veins.
Software pertaining to the above-named applications must be available for obtaining the following:-
a) Display of maximum, minimum and mean mode curves with spectral Doppler.
b) Pulsatility and resistive index calculation.
c) Lumen calculations via diameter/area method.
d) Percentage stenosis based on area or diameter calculation.
e) Any other desirable features of the software package must be stated.
19 | IMAGE PROCESSING:
19.1 | The unit must be capable of real time and frozen pan-zoom of B-Mode and colour flow displays. Magnification factors must be at least 10 times. Any better capabilities must be stated.
20 | OPERATIONAL:
20.1 | The unit must have a freeze-frame facility that must be operated by a push-button.
20.2 | There should be no deterioration in brightness or contrast whilst the frozen image is being viewed.
20.3 | Repetitive viewing in a cine-loop with freeze frame and slow motion must be possible.
21 | MEASURING SYSTEM:
21.1 | An electronic caliper measuring system, calibrated in millimeters over the scan depth, is required with an accuracy of at least 1mm.
21.2 | The system must be capable of multiple pairs of calipers.
21.3 | The calipers must be movable to any position on the screen by the trackball. And must be in the range of (6 – 8)
21.4 | The separation between markers, area, or circumference must be displayed on the screen and be easily readable.
21.5 | Scale marker displays should be provided 1cm apart.
21.6 | Details of any additional measurement capability that is present in the system offered must be supplied.
22 | RECORDING SYSTEM:
22.1 | The system must include a black and white thermal printer with a capacity to take 110mm thrmal paper
22.2 | The printer must be integrated into the unit. The cables of the printer should be easily accessible for any changes that need to be made.
23 | POWER SUPPLY:
23.1 | The equipment offered shall operate from a standard 220 to 240 Volt A.C. 50Hz mains supply. The bidder must ensure that the product being quoted for is fitted with a 15 Amp sealed plugtop S.A.N.S approved
23.2 | The unit must be supplied with an on-line uninterrupted power supply (U.P.S.) to provide minimum of 50 to 60 minutes mains backup power.
23.3 | The UPS should not adversely affect the image quality.
23.4 | This unit must be included in the bid price and be covered by the guarantee and maintenance plan.
23.5 | The UPS should not adversely affect the image quality.

## Mobile Ultrasound
MOBILE ULTRASOUND UNIT
ITEM | TECHNICAL SPECIFICATIONS
1 | IMAGING MODALITIES
1.1 | Unit must offer excellent high image quality in B-Mode.
1.2 | Unit should be capable of offering a Maximum frame rate of no less than 394 frames per second
1.3 | The system must offer a Digital Continuous Beam Former to generate sharp defined ultrasound beams, to reduce artifacts and to provide more consistent overall image quality
1.4 | Unit should offer a 38-cm LCD colour monitor with a flexible arm to support LCD monitor
1.5 | Unit must have a rechargeable battery and sufficient battery life when operating the unit is operating on battery power only.
1.6 | Free programmable panel with changeable labels must be offered
1.7 | Unit should offer hard drive with a minimum memory capacity of 80GB. State maximum memory capacity.
1.8 | This unit must offer Tissue Harmonics on all probes
1.9 | The Tissue Doppler imaging software should be included
1.10. | Unit must be portable but also housed on a mobile cart.
1.11 | The system must have security functions
1.11.1 | Unit should offer entry of user name and
1.11.2 | passwords auto locking if there is no activity for security purpose of patients
1.11.3 | anti virus
1.12 | Panoramic software should be available in B-mode and colour
1.13 | Images not less than 40 000 must be able to be stored on hard drive
1.14 | Unit should allow emergency log on
1.15 | Advanced directional dynamic flow must be included
1.16 | The unit should have Programmable screensaver.
1.17 | Unit must offer handle for easy portability of equipment around hospital.
1.18 | Post processing must be available on frozen images.
1.19 | System Dynamic Range should not be less than 172 dB
1.20. | Unit should offer a minimum of 5 frequencies on all probes
1.21 | The unit must be able to offer scanning methods for Electronic Convex, Linear and phased array.
1.22 | Maximum depth should not be less than 30cm.
1.23 | Unit should offer depth gain in not less than 8 steps
1.24 | Image processing adjustments via a rotary dial with 5 combinations for optimising the best image quality for each patient’s physique (user programmable) should be available.
1.25 | Pan-zoom enlargement should be in real and frozen mode up to 2cm
1.26 | Sweep speed adjustable with a minimum of 4 steps. Independent M-mode , gain adjustment should be available
1.27 | B-mode image on linear probe must be electronically steerable.
1.28 | Grey scale level should not be less than 256 shades
1.29 | Oblique steerable B-mode must be offered on linear probes
1.30. | The unit must be connected to the SERVER, Printer, the PACS via the network and the HIS/RIS
2 | DOPPLER
2.1 | There must be Pulsed Wave Doppler on all imaging probes.
2.2 | The maximum detectable velocity for PW must be stated.
2.3 | Pulse Repetitions Frequency should varies between 1.3 to 12.5kHz
2.4 | High Pulse Repetitions Frequency should varies between 5.8 to 15.8 kHz
2.5 | State Max detectable velocity
2.6 | Independent angle steering left and right up to ±20° for colour and Doppler must be offered
2.7 | Independent gain control for PW
2.8 | Colour pallet offering 7 types independent of Pulse Doppler pallet to be available.
3 | COLOUR DOPPLER
3.1 | Maximum colour frame rate should not be less than 371 frames p/sec.
3.2 | Directional colour angio should be standard
3.3 | Tissue Doppler imaging should be standard
3.4 | Minimum of 20 colour maps should be available for colour display
3.5 | Colour maps must be able to be selected after freeze.
3.6 | Independent colour gain adjustment 1 – 16 dB in 0.5 Db steps
3.7 | Colour flow must operate to the full depth of minimum 24cm.
3.8 | There should be a minimum of 7 processing selections i.e. Wall filters, colour averaging etc. State the number and list the possible selection.
3.9 | Colour maps and processing selections must be able to be selected by the user and stored in the presets
3.10. | Quad signal processing to be available
3.11 | Parallel signal processing to be available. Bidder to provide details
3.12 | Dynamic flow on all transducers should be offered
4 | IMAGE MEMORY
4.1 | Image memory to operate in both dual and quad screen. 2D,2D/2D,2D/M, 2D/Doppler,2D/C/Doppler State format
4.2 | B-Loop facility up to 511 frames should be possible
4.3 | M/D cine review of a minimum of 128 sec should be possible.
4.4 | B-Mode image on linear transducer to be electronically steerable.
5 | TRANSDUCERS
5.1 | All probes required must be offered
5.2 | All imaging transducers must be multi-frequency
5.3 | Tissue Harmonics should be available on Convex probes
5.4 | The selected frequency of the imaging transducer must be displayed on the screen. There should be at least B-mode frequencies, THI frequencies, PW frequencies and colour frequencies.
5.5 | All transducers must be digitally controlled so continuous beam forming is possible.
5.6 | A minimum of 3 active probe plus pencil connectors must be available
5.7 | All probes must be included in the main offer
5.7.1 | Multi frequency Convex probe with tissue for abdominal , O&G ,vascular applications :3.0 MHZ to 6 MHz
5.7.2 | Pediatric curvilinear probe for premature to toddler abdomens with a 5MHz to 8MHz as well as neonatal neuro-imaging must be offered
5.7.3 | High frequency and multi frequency linear probe with tissue harmonies for small parts and muscular applications from 8MHz to 14 MHz with a trapezoid footprint.
5.7.4 | Phased array multi frequency probe with tissue harmonics for trans cranial and adult cardiac applications : central frequency 5 MHz
6 | MEASUREMENTS
6.1 | On screen measurements must be possible from real time examination.
6.2 | Basic cardiac measurements must be included
6.3 | End diastolic volume (simpson, bullet)/stroke volume, cardiac output must be possible
6.4 | Not less than 8 pairs of distance callipers must be offered for B-mode measurements. Trace, area, circumference ,ellipse and volume must be possible
6.5 | Valve orifice area measurement must be offered
6.6 | 16 callipers (8 channel) for measuring distance angle area volume should be offered
6.7 | 16 callipers (8 channels) hip joint should be offered
6.8 | The unit must offer storage facility for up to 2000 patient foetal growth measurement data. Must be able to be saved and be retrievable for follow-up visits
6.9 | The unit must offer trends on all Doppler measurements e.g. Resistive Index + Pulsitility Index of the umbilical cord, MCA, uterine arteries etc. Unit must be able to store this information for the follow-up visits
6.10. | 50th percentile growth graphs must be offered on all measurements for example BPD, FL, AC, Foetal Weight, and HC. This information must be retrievable for all follow-up visits, to detect abnormal growth of foetus early
6.11 | Unit must offer user programmable annotation and measurement labelling to reduce scanning time.
6.12 | Display of graphs must be available to compare measurements of twins, triplets and quads
7 | RECORDING DEVICES
7.1 | Black and white thermal printer must be offered
7.2 | CD/DVD Writer/Recorder should be integrated part of the unit
7.3 | CD/DVD Recorder must be operated through panel
7.4 | High quality laser printer and desktop printer must be compatible
7.5 | Advanced DICOM/PACS with full worklist must be offered and to be included in main offer.
8 | SOFTWARE UPGRADEABILITY:
8.1 | Bidders are to state the policy with regard to future software upgrades and the costs.
8.2 | The following option should quoted separately
8.2.1 | Spatial compounding
8.2.2 | Panoramic imaging
8.2.3 | Cardiac upgrade
9 | POWER SUPPLY:
9.1 | The equipment offered shall operate from a standard 220 to 240 Volt A.C. 50Hz mains supply. The bidder must ensure that the product being quoted for is fitted with a 15 Amp sealed plugtop S.A.N.S approved
9.2 | The unit must be supplied with an on-line uninterrupted power supply (U.P.S.) to provide minimum of 50 to 60 minutes mains backup power.
9.3 | The UPS should not adversely affect the image quality.
9.4 | This unit must be included in the bid price and be covered by the guarantee and maintenance plan.
9.5 | The UPS should not adversely affect the image quality.

## Ultrasound mid
ULTRASOUND UNIT (MID-RANGE)
ITEM | Tender Specification | Weight
1 | Imaging Modes
1.1 | Unit must be latest model with latest technology. Details are stated. | #
1.2 | Ultrasound applications for abdominal, obstetrics, gynecological, small parts, vascular / peripheral vascular, musculoskeletal, urological, transcranial, cardiac and elastography . | #
1.3 | Unit should be capable of offering a frame rate from 400 frames per second and above. | #
1.4 | The colour frame rate should be 400 or above. State maximum colour frame rate. | 2
1.5 | Unit should offer a 42 cm or better high definition LCD colour monitor. | #
1.6 | Ergonomic design keyboard. Details must be stated. | 2
1.7 | Brightness / contrast adjustment on monitor. | #
1.8 | Programmable panel with changeable labels must be offered | #
1.9 | Graphical user interphase (user friendly and exam specific) | #
1.10. | Height adjustable user interphase | #
1.11 | Unit must be housed on a ergonomic and compact mobile cart with lockable wheels, for easy maneuvaribility | #
1.13 | Post processing must be available on frozen images ( Gain, Gray Map,Chrom Color map, Baseline, Remove color) . Details are stated. | #
1.14 | The gray scale level for image display modes, should be 256 shades and above. | #
1.15 | Touch screen Control Panel 8 inch and above. | #
1.16 | Unit should offer emergency log on. | #
2 | Functionality
2.1 | 12 bit and above digital converter for dynamic image / live image (beam former). | #
2.2 | B -mode must have trapezoid function. | #
2.3 | B mode must have panoramic view function. | #
2.4 | Autocalibration for measurements on panoramic image must be standard. | #
2.5 | B mode with broad band harmonic for increased resolution from near and far field. | #
2.6 | B-Mode image on linear transducer to be electronically steerable. | #
2.7 | Image format to operate in dual and quad screen. | #
2.8 | Independent gain control for Pulsed Wave Doppler. | #
2.9 | Scanning Operating Modes: B-mode, M-mode, Pulsed Wave Doppler (Spectral), Continuous Wave Doppler, Power Doppler & Color Doppler, 2D. | #
2.1 | Pulse Repetition Frequency (PRF) / Velocity must be adjustable in increments. | #
2.11 | Pulse Repetition Frequency (PRF) / Velocity values must be displayed on the screen. | 2
2.12 | Reference Frequency (REF) must be displayed. | 2
2.13 | Wall filter must be adjustable. | 2
2.14 | The sampling gate must be adjustable: range between 0.3mm and 15mm. | #
2.15 | Directional Power Doppler | 2
2.16 | M-mode / Doppler review (cine playback or loop) of 10 seconds and above. Details are stated. | 2
2.17 | M-mode / Doppler to be steerable. | #
2.18 | Independent B-mode and Color Doppler gain adjustment must be available. | #
2.19 | Independent adjustment of each mode for image optimization of each mode during triplex mode . | #
2.2 | The Maximum scanning depth must be 30 cm and above. Details are stated. | #
2.21 | One button image optimization must be standard resolution | 2
2.22 | Internal Hard drive with minimum usable image storage space of 250GB or higher. You must be able to store videos and still images on internal hard drive. | 2
2.23 | Auto deletion of images should be configurable | #
2.24 | Notice warning be included when minimum image storage capacity is reached on the system. The actual minimum storage capacity in GB and total number of images is stated in supporting documents provided. | #
2.25 | Black and white digital thermal printer attached to the unit must be included | #
2.26 | an Active DVD Writer/Recorder attached to the unit must be included | 1
2.27 | 2x USB Port (External) should be an integral part the unit for storage/recording of real time loops (video ) and still images. 2x Internal USB for Printer & memory stick use. | 2
2.28 | Software to ensure functionality of DVD Writer/Recorder with all types of computers/laptops must be included as a standard. | 2
2.29 | The system must be DICOM 3 compliant. | 2
2.3 | Advanced DICOM/PACS including full worklist must be offered, and to be included in Main Unit. | 2
2.31 | Export of images from the system must be possible via LAN, USB, DVD, (including transfer of images to PACS) . The methods shall be stated in supporting documents. | #
2.32 | the following storage formats must be possible on unit: JPEG, Bitmap, Tiff, AVI, DICOM | #
2.33 | On screen measurements must be possible on real time examination. | #
2.34 | 50th percentile growth graphs must be offered on all measurements for example BPD, FL, AC, Fetal Weight, HC. This information must be retrievable for all follow-up visits, to detect abnormal growth of foetus early | #
2.35 | Display of graphs must be available to compare measurements of twins , triplets in single and quad display. | #
2.36 | User programmable exam preset capability | #
2.37 | Unit must offer user programmable annotation for quick labelling. | #
2.38 | Storage and review of patient data must be available for comparative exams. Details are stated on the options you offer.if there is functional PACS. | 1
3 | Transducers
3.1 | The selected frequency of the imaging transducer must be displayed on the screen. | #
3.2 | Three transducers must be included and the end user must be able to select the final transducers Clearly State all available transducers with prices per transducers (swapping of products/transducers and software to be allowed with equitable costs). Details are stated. | 2
3.3 | Additional probes and software must be available as optional items. Optional items must be listed and priced separately. | 2
3.4 | Types and frequency range of transducers:
3.4.1 | Convex Probe (Curvilinear). The MHz range are stated (within 1-6 MHz) | #
3.4.2 | Microconvex for neonatal / paediatrics.The MHz range are stated (4-8) | 2
3.4.3 | Microconvex Transvaginal / transrectal. The MHz range are stated (within 3-8 MHz) | #
3.4.4 | Linear probe (vascular) The MHz range are stated (4-12) | 2
3.4.5 | Linear probe (MSK)The lower MHz range must be 5-6 MHz and the upper range must be between 14 and 21 MHz. | 2
3.4.6 | Adult Cardiac Probe. State the MHz range (5 MHz). Must have pulse and continuous Wave Doppler software | 2
3.4.7 | Paediatric Cardiac Probe. State the MHz range (7.5-10 MHz). Must have pulse and continuous Wave Doppler software | 2
3.4.8 | Minimum 3 x standard Probe holders. | #
3.4.9 | Minimum 3 x Active probe ports. Details of total number of active probe ports are stated. | #
3.4.10 | Foot rest as a standard. | 1
3.4.11 | Chair, ergonomic, height adjustable with lockable castors for the Health professional. | 2
3.4.12 | Printer mounted on the machine. | 1
3.4.13 | Cable tray /hooks cable support system. | 1
3.4.14 | Gel warmer. | 2
3.4.15 | Easy removable air filters (by end users). | 2
3.4.16 | Ultrasound Quality Control Phantom (swapping of phantoms, products and software to be allowed with equitable costs). Details are stated. | 2
3.5 | SOFTWARE:
3.5.1 | Compound Imaging ( Cross Beam angles). | #
3.5.2 | Speckle Reduction (Minimum of 4-5 steps ). Refines images(reduces noise) | #
3.5.3 | System must have software to enhance needle visualization during biopsy | 2
3.5.4 | Measurement and Report software packages must be configurable | 1
3.5.5 | Details of tools/software provided as a standard are stated | 1
3.5.6 | The standard software packages supplied must be fully upgradable . | #
3.5.7 | The cost of updating and upgrading all standard software packages supplied , during warranty and maintenance period must be included in the offer . It must be clearly outlined as a line item in the offer documentation. | #
4 | OPTIONAL ITEMS
4.1 | Fold down monitor for easy transportation. | 1
4.2 | USB Port, additional | 1
4.3 | Any additional tools / software not included in the standard package must be clearly listed and priced. | 1
4.4 | Side storage bin | 1
4.5 | Elastography software | 1
4.6 | 3D software | 1
4.7 | 4D software | 1
4.8 | Hard Drive of 5TB for additional storage/archiving is included.
5 | POWER SUPPLY:
5.1 | The equipment offered shall operate from a standard 220 to 240 Volt A.C. 50Hz mains supply. The bidder must ensure that the product being quoted for is fitted with a 15 Amp sealed plugtop S.A.N.S approved
5.2 | The unit must be supplied with an on-line uninterrupted power supply (U.P.S.) to provide minimum of 50 to 60 minutes mains backup power.
5.3 | The UPS should not adversely affect the image quality.
5.4 | This unit must be included in the bid price and be covered by the guarantee and maintenance plan.
5.5 | The UPS should not adversely affect the image quality.

## Ultrasound high-end
ULTRASOUND UNIT (HIGH END)
TECHNICAL SPECIFICATION
ITEM | IMAGING MODALITIES
1.1 | This unit must be a fully digital high end Colour Doppler unit with pulsed wave Doppler and Continuous wave Doppler.
1.2 | The unit must have a standby mode for quick start up & shutdown.
1.3 | The unit must have 3 active probe connectors with electronic switching.
1.4 | The unit must be ergonomically designed with an elevating adjusting panel to reduce repetitive stress injuries.
1.5 | The unit must offer a 48 cm or better high resolution flat panel colour monitor which must be able to move up and down and from side-to-side.
1.6 | A cine loop function of at least 30sec must be offered for B-mode and D - mode.
1.7 | The unit must be fully DICOM/PACS compatible. DICOM verification, worklist, structured reporting, storage and print must be available as standard.
1.8 | The unit offered must have at least 29,000 channels.
1.9 | The unit must offer at least 512 line density in B-mode.
1.1 | Phase Inversion Tissue Harmonics must be offered on all probes
1.11 | Spatial compounding offering multiple lines of sight must be offered to improve contrast resolution and speckle reduction.
1.12 | The system offered must have at least a 500GB hard drive.
1.13 | The unit must offer TGC in not less than 8 steps.
1.14 | Pan-zoom enlargement must be at least 10 times read & write. Image quality must not be compromised by the zoom function.
1.15 | Easy programmable image processing must be in not less than 8 steps to enable the user to change dynamic range, persistence, contrast, enhancement, smoothing and post-processing.
1.16 | Independent M-mode gain adjustment must be available.
1.17 | B-mode frame rate of at least 498 fps
1.18 | Colour frame rate higher than 200 fps
1.19 | B-Mode image on the linear probe must be electronically steerable.
1.20. | Trapezoidal imaging must be available on linear probes.
1.21 | Sector FOV must be at least 90 deg.
1.22 | Maximum depth must not be less than 28cm.
1.23 | System dynamic range must be greater than 160dB.
1.24 | One button B-mode image optimisation must be possible.
1.25 | Panoramic software must be available in B-mode & Colour.
1.26 | Real-time 4D imaging of at least 30 volumes per second. Trans-abdominal and trans-vaginal 4D transducers and software to be included.
1.27 | Images of video clips and still images must be able to be stored on hard drive. Measurements on stored images must be possible.
1.28 | Loop / Clip storage of up to 1 minute must be possible.
1.29 | Please state any other additional features/technologies that are available on the unit offered.
2 | PULSED WAVE (PW) & CONTINUOUS WAVE DOPPLER (CW).
2.1 | PW with a one button Doppler optimisation must be possible.
2.2 | PW Doppler on all imaging probes.
2.3 | The max detectable velocity for PW must be stated.
2.4 | Maximum PRF must be equal to or greater than 40kHz.
2.5 | Independent gain adjustments for PW & CW must be available.
2.6 | Independent angle steering up to 20° for colour and PW must be offered on Linear Probes.
2.7 | Colour pallet offering 7 types independent of B-Mode, Pulse Doppler & Continuous wave doppler pallet must be available.
2.8 | Sampling with Doppler gate must be varied between 1 – 20mm.
2.9 | Display of Doppler scan modes side by side and up and down must be included .
2.10. | Post-processing of the Doppler trace must be possible in freeze mode.
2.11 | The unit must offer real time triplex mode on PW.
3 | COLOUR DOPPLER
3.1 | At least 8 colour maps must be available for colour Doppler.
3.2 | Colour maps must be able to be selected after frozen mode.
3.3 | Independent colour gain adjustment must be +15dB to -15dB in 0.2dB steps.
3.4 | Maximum colour frame rate must not be less than 200 frames p/sec.
3.5 | Colour flow must operate to the full depth of 28cm.
3.6 | No deterioration of B-mode must be noticeable in colour mode.
3.7 | There must be a minimum of 7 processing selections i.e. wall filters, colour averaging etc. State the number of selections and list the possible selections.
3.8 | The user must be able to select colour maps and processing selections. It must be possible to store these preset selections.
3.9 | The colour box must move with the Doppler gate to reduce scanning time.
3.10. | Power Doppler must be available to display the existence of blood flow with high sensitivity.
4 | TRANSDUCERS
4.1 | All probes required must be offered and included in the total bid price.
4.2 | All imaging transducers must be multi-frequency or broadband.
4.3 | Tissue Harmonics should be available on all probes.
4.4 | The selected frequency of the imaging transducer must be displayed on the screen. There should be at least B-mode frequencies, THI frequencies, PW frequencies and colour frequencies.
4.5 | All transducers must be digitally controlled so that continuous beam forming is possible.
4.6 | It must be possible to connect 3 probes simultaneously.
4.7 | The following transducers must be offered and included in the total bid price.
4.7.1 | Multi-frequency convex trans-abdominal transducer with Tissue Harmonics for Abdominal, O&G and vascular applications: 2 - 5 MHz central frequency.
4.7.2 | Wideband Linear transducer 5- 14MHz for general imaging
4.7.3 | High density Linear probe 18 MHz for breast and superficial imaging
4.7.4 | 5 – 9 MHz intercavity transducer for rectal and vaginal scanning
4.7.5 | Adult cardiac probe 5 MHz and Pediatric cardiac probe 7.5 MHz
5 | MEASUREMENTS (add for cardiac)
5.1 | 16 callipers (8 channel) for measuring distance angle area volume must be included.
5.2 | The unit must offer storage facility for up to 2000 patient fetal growth measurement data. This information must be retrievable for follow-ups at a later stage.
5.3 | 50th percentile growth graphs must be offered on all measurements for example BPD, FL, AC, Fetal Weight, HC. Information saved must be retrievable for all follow-up at a later stage.
5.4 | Display of graphs must be available to compare measurements of twins, triplets, quads.
5.5 | Unit must offer user programmable annotation and measurement labels to reduce scanning time.
5.6 | Automatic trace measurements must be possible on real-time and frozen Doppler traces.
5.7 | Supply details of any other measurement capabilities that are available on the system offered.
6 | RECORDING DEVICES:
6.1 | A black and white thermal printer must be included in the offer.
6.2 | A digital colour thermal printer must be offered as an option.
6.3 | Internal CD/DVD read / write must be offered as standard. It must be possible to write images in DICOM and PC format when storing to CD/DVD/PACS.
6.4 | A DVD writer/recorder for storage/recording of real time loops and still images must be offered as an option.
6.6 | A USB port for image transfer to a memory stick must be available.
7 | POWER SUPPLY:
7.1 | The equipment offered shall operate from a standard 220 to 240 Volt A.C. 50Hz mains supply. The bidder must ensure that the product being quoted for is fitted with a 15 Amp sealed plugtop S.A.N.S approved
7.2 | The unit must be supplied with an on-line uninterrupted power supply (U.P.S.) to provide minimum of 50 to 60 minutes mains backup power.
7.3 | The UPS should not adversely affect the image quality.
7.4 | This unit must be included in the bid price and be covered by the guarantee and maintenance plan.
7.5 | The UPS should not adversely affect the image quality.

## Ultrasound high-portable
ULTRASOUND MACHINE, PORTABLE, SMALL AND COMPACT, HIGH END
ITEM | TECHNICAL SPECIFICATIONS | Weight
1 | The machine must be configurable to perform the following ultrasound investigations:
1.1 | Adult cardiac and paeditric | #
1.2 | Adult transesophageal | 2
1.3 | Stress echo | #
1.4 | Abdominal | #
1.5 | Pediatric | #
1.6 | Vascular | #
1.7 | Obstetrical | #
1.8 | Gynecological | #
1.9 | Small parts | #
1.10 | Breast | #
1.11 | Musculoskeletal | #
1.12 | Emergency medicine | #
1.13 | Regional anesthesia | 2
1.14 | Integrated Standby Battery | #
2 | imaging modes
2.1 | • 2D | #
2.2 | • M-mode | #
2.3 | • Anatomical M-mode | #
2.4 | • Color M-mode | #
2.5 | • Color Power doppler imaging | #
2.6 | • Pulsed wave (PW) Doppler | #
2.7 | • HPRF PW Doppler | #
2.8 | • Continuous wave (CW) Doppler | 2
2.9 | • Freehand 3D imaging (available as an option) | 1
2.10 | •advanced quantification/ measurement software | 2
2.11 | • Invert and color invert | #
2.12 | • Color compare mode | #
2.13 | • Dual mode | #
2.14 | • Duplex for simultaneous 2D and Doppler | #
2.15 | • Triplex for simultaneous 2D, Doppler, and color or color power doppler | #
2.16 | • 2D and flow optimization signal processing | #
2.17 | • Tissue harmonic imaging (THI) | #
2.18 | • Reconstructed zoom with pan | #
2.19 | • High definition write zoom | #
2.20 | • Trapezoidal imaging | #
2.21 | • Contrast imaging (as an option) | 1
2.22 | • Adaptive Doppler | #
2.23 | • Adaptive color Doppler | #
2.24 | • Color tissue Doppler imaging | #
2.25 | • Pulsed wave tissue Doppler imaging | #
3 | 2D grayscale
3.1 | • Smart TGC: pre-defined TGC curves optimized for consistently | #
3.2 | • User adjustable LGC control | #
3.3 | • High definition zoom concentrates all image processing power into a user-defined area of interest | #
3.4 | • possible to combine high definition zoom with pan zoom | #
3.5 | • Cineloop image review | #
3.6 | • Selectable 2D compression settings | #
3.7 | • Sector size and steering control for sector and curved array | #
4 | image formats
4.1 | • Dual imaging with either independent cineloop buffers or split screen imaging | #
4.2 | • Live compare | #
4.3 | • 256 (10-12 bits) discrete gray levels or better | #
4.4 | • state 2D acquisition frame rate (preferably 400 up to 755 frames/sec) | #
5 | M-mode
5.1 | • Available on cardiac imaging transducers | #
5.2 | • Selectable sweeping rates | #
5.3 | • Time markers: 0.2 and 1.0 seconds | 2
5.4 | • Acquisition zoom capability | #
5.5 | • Selectable display format prospective or retrospective | #
5.6 | • Cineloop review for retrospective analysis | #
6 | Anatomical M-mode
6.1 | • Available for cardiac imaging transducers | #
6.2 | • Uses 2D image as a basis for M-mode analysis at a defined line, independent of transducer orientation | #
6.3 | • Anatomical M-mode trace can be generated or modified post Freeze | #
7 | Tissue Doppler imaging
7.1 | • Allows high frame rate acquisition of tissue motion | #
7.2 | • Color gain, TGC, and LGC | 2
7.3 | • 15 to 20 color maps | #
7.4 | • Velocity (6cm to 50cmcm/s) | #
8 | Pulsed wave Doppler
8.1 | • Available on all imaging transducers | #
8.2 | • Adjustable sample volume size:(e.g. 0.5-20 mm) -state size | #
8.3 | • Simultaneous or duplex mode of operation | #
8.4 | • Simultaneous 2D, color Doppler or CPA, pulsed Doppler | #
8.5 | • automatic optimization - adjusts scale, baseline and doppler gain | #
9 | Continuous wave Doppler
9.1 | • Steerable through 90° sector | 2
9.2 | • Maximum velocity range: 15sec (transducer dependent) | #
10 | Spectral Doppler
10.1 | • Display annotations including Doppler mode, scale (cm/sec),Nyquist limit, wall filter setting, gain, acoustic output status, sample volume size, normal/inverted, angle correction,grayscale curve | #
10.2 | • Angle correction with automatic velocity scale adjustment | #
10.3 | • Adjustable velocity display ranges | #
10.4 | • Normal/invert display around horizontal zero line | #
10.5 | • Selectable sweep speeds | #
10.6 | • Selectable low-frequency signal filtering with adjustable wall filter settings | #
10.7 | • Selectable grayscale curve for optimal display | #
10.8 | • Selectable display format prospective or retrospective - state options | #
10.9
10.10 | • 256 (10 t0 12 bits) discrete gray levels | #
10.11 | • Post-processing in PW frozen mode includes map, baseline,sweep speed, invert, compress, reject, and Chroma | #
11 | Color Doppler
11.1 | • Adaptive mode adjusts Doppler frequency and sensitivity based on color ROI placement within image available on all imaging transducers | #
11.2 | • Cineloop review with full playback control | #
11.3 | • Advanced motion suppression with intelligent algorithms; to eliminate color motion artifacts. Give details | 2
11.4 | • Trackball-controlled color region of interest: size and position | #
11.5 | • Maps, filters, color sensitivity, line density, smoothing, echo, write priority, color persistence, gain and baseline optimized automatically by exam type or is user selectable | #
11.6 | • Velocity and variance displays | #
11.7 | • Color invert in live and frozen imaging | #
11.8 | • User selectable smoothing control | #
11.9 | • User selectable persistence control | #
11.1 | • Color/2D line density control | #
12 | Tissue harmonic imaging
12.1 | • Available on all imaging transducers | #
12.2 | • Extends high performance imaging capabilities to all patient body types | #
13 | Color Power doppler
13.1 | Directional power flow available | #
13.2 | • Highly sensitive mode for small vessel visualization | #
13.3 | • Available on all imaging transducers | #
13.4 | • Cineloop review | #
13.5 | • Multiple color maps | #
13.6 | • Individual controls for gain, PRF, baseline, filters, sensitivity, | #
13.7 | • Dynamic motion differentiation | #
13.8 | • Adjustable region of interest: size and position | #
13.9 | • User-selectable persistence | #
13.10 | • User-selectable blending | #
14 | Connectivity
14.1 | • Two USB ports on control panel | 2
14.2 | • At least 80 GB internal hard drive space or more | #
14.3 | • Internal slot-load CD/DVD RW drive | 2
14.4 | • DICOM print, store, and storage commitment | #
14.5 | • DICOM structured reporting for cardiac and obstetrics | 2
14.6 | • Modality worklist | #
14.7 | • DICOM reader saved onto media | #
14.8 | • Export data as PC-compatible or DICOM files | #
14.9 | • Ethernet at 100 Mb/second | 2
14.1 | • USB to serial converter adapter | 1
14.11 | • Support for optional small B/W and color printers | 1
14.12 | • Must be PACS ready with no added cost to connect to PACS | #
15 | Probes: the following probes must be included in the price. Give prices for individual probes.
15.1 | Adult TEE (offer as optional) | 1
15.2 | Vascular linear (L12-3) | #
15.3 | Abdominal curviliar (C5-1) | #
15.4 | Hockey stick (L15-7) optional | 1
15.5 | Microconvex/ pediatric abdomen (C8-5) | #
15.6 | Sector transducer 1.5Mhz to 4.5 Mhz (adult cardiac) | 2
15.7 | Sector transducer 3 to 8Mhz for paediric cardiac | 2
16 | Power Supply
16.1 | The unit must be supplied with an on-line uninterrupted power supply (U.P.S.) to provide minimum ten 10 to 20 minutes mains backup power.
16.2 | The UPS should not adversely affect the image quality.
16.3 | This unit must be included in the bid price and be covered by the guarantee and maintenance plan.

## Ultra High-portable & compact
ULTRASOUND MACHINE, PORTABLE, HIGH END AND COMPACT
ITEM | TECHNICAL SPECIFICATIONS | Weight
1 | General
1.1 | Unit should offer at least a 17 to 19 inch LCD high resolution colour monitor with integrated keyboard (laptop type). State the size and technology. | 2
1.2 | Indicate if system has a built-in UPS | 2
1.3 | The unit must come with a trolley
2 | The machine must be configurable to perform the following ultrasound investigations:
2.1 | Adult cardiac and paediatric | #
2.2 | Adult and paediatric transesophageal | 2
2.3 | Stress echo | #
2.4 | Abdominal | #
2.5 | Paediatric | #
2.6 | Vascular | #
2.7 | Obstretrical | #
2.8 | Gynaecological | #
2.9 | Small parts | #
2.1 | Breast | #
2.11 | Musculoskeletal | #
2.12 | Emergency medicine | 2
2.13 | Regional anaesthesia | 2
2.14 | Intergrated standby battery
3 | imaging modes
3.1 | • 2D | #
3.2 | • M-mode | #
3.3 | • Anatomical M-mode | #
3.4 | • Color M-mode | #
3.5 | • Color Power doppler imaging | #
3.6 | • Pulsed wave (PW) Doppler | #
3.7 | • High Pulse Repetition Frequency (HPRF) PW Doppler | #
3.8 | • Continuous wave (CW) Doppler | #
3.9 | • Freehand 3D imaging (available as an option) | 1
3.1 | •advanced quantification/ measurement software | 2
3.11 | • Invert and color invert | #
3.12 | • Color compare mode | #
3.13 | • Dual mode | #
3.14 | • Duplex for simultaneous 2D and Doppler | #
3.15 | • Triplex for simultaneous 2D, Doppler, and color or color power doppler | #
3.16 | • 2D and flow optimization signal processing | #
3.17 | • Tissue harmonic imaging (THI) | #
3.18 | • Reconstructed zoom with pan | #
3.19 | • High definition write zoom | #
3.20. | • Trapezoidal imaging | #
3.21 | • Contrast imaging (as an option) | 1
3.22 | • Adaptive Doppler | #
3.23 | • Adaptive color Doppler | #
3.24 | • Color tissue Doppler imaging | #
3.25 | • Pulsed wave tissue Doppler imaging | #
4 | 2D grayscale
4.1 | • Time Gain Compensation (TGC): pre-defined TGC curves optimized for consistently | 2
4.2 | • User adjustable Lateral Gain Control (LGC) | #
4.3 | • High definition zoom concentrates all image processing power into a user-defined area of interest | #
4.4 | • ability to combine high definition zoom with pan zoom | #
4.5 | • Cineloop image review | #
4.6 | • Selectable 2D compression settings | #
4.7 | • Sector size and steering control for sector and curved array | #
5 | image formats
5.1 | • Dual imaging with either independent cineloop buffers or split screen imaging | #
5.2 | • Live compare | #
5.3 | • 256 Gray scale levels or more | #
5.4 | • 2D acquisition frame rate of 400 frames/sec up to 800 frames/sec. State details. | 2
6 | M-mode
6.1 | • Available on all imaging transducers | #
6.2 | • Selectable sweeping rates | #
6.3 | • Time markers: minimum of 0.2 seconds up to a maximum of 2.0 seconds | 2
6.4 | • Acquisition zoom capability | #
6.5 | • Selectable display format prospective and retrospective | #
6.6 | • Cineloop review for retrospective analysis | #
7 | Anatomical M-mode
7.1 | • Available for all imaging transducers | #
7.2 | • Uses 2D image as a basis for M-mode analysis at a defined line, independent of transducer orientation | #
7.3 | • Anatomical M-mode trace can be generated or modified post freeze | #
8 | Tissue Doppler imaging
8.1 | • Allows high frame rate acquisition of tissue motion | #
8.2 | • Color gain, TGC, and LGC | 2
8.3 | • 15 to 20 color maps | 2
8.4 | • Tissue velocity (6cm to 50cm/s) | 2
9 | Pulsed wave Doppler
9.1 | • Available on all imaging transducers | #
9.2 | • Adjustable sample volume size:(e.g. 0.5-20 mm) -state size | 2
9.3 | • Triplex / real time duplex mode of operation | #
9.4 | • Simultaneous 2D, color Doppler or CPA, pulsed Doppler | #
9.5 | • automatic optimization button / knob - adjusts scale, baseline and doppler gain | 2
10 | Continuous wave Doppler
10.1 | • Steerable through 90° sector | 2
10.2 | Maximum velocity range: 15 sec (transducer dependent)
11 | Spectral Doppler
11.1 | • Display annotations including: | #
11.1.1 | Doppler mode
11.1.2 | scale (cm/sec)
11.1.3 | Nyquist limit
11.1.4 | wall filter setting
11.1.5 | gain
11.1.6 | acoustic output status
11.1.7 | sample volume size
11.2 | • Angle correction with automatic velocity scale adjustment | #
11.3 | • Adjustable velocity display ranges | #
11.4 | • Normal/invert display around horizontal zero line | #
11.5 | • Selectable sweep speeds | #
11.6 | • Selectable low-frequency signal filtering with adjustable wall filter settings | #
11.7 | • Selectable grayscale curve for optimal display | #
11.8 | • Selectable display format of prospective and retrospective | #
11.9 | • 256 Gray scale levels or more | #
11.10. | • Post-processing in Pulsed Wave (PW) frozen mode includes map, baseline,sweep speed, invert, compress, reject, and chroma | #
12 | Color Doppler
12.1 | • Adaptive mode adjusts Doppler frequency and sensitivity based on color ROI placement within image available on all imaging transducers | #
12.2 | • Cineloop review with full playback control | #
12.3 | • Advanced motion suppression with intelligent algorithms; to eliminate color motion artifacts. State details. | 2
12.4 | • Trackball-controlled color region of interest: size and position | #
12.5 | • Maps, filters, color sensitivity, line density, smoothing, echo, color persistence, gain and baseline optimized automatically by exam type or is user selectable | #
12.6 | • Velocity and variance displays | #
12.7 | • Color invert in live and frozen imaging | #
12.8 | • User selectable smoothing control | #
12.9 | • User selectable persistence control | #
12.1 | • Color/2D line density control | #
13 | Tissue harmonic imaging
13.1 | • Available on all imaging transducers | #
13.2 | Extends high performance imaging capabilities to all patient body types
14 | Color Power doppler
14.1 | Directional power flow available | #
14.2 | • Available on all imaging transducers | #
14.3 | • Cineloop review | #
14.4 | • Multiple color maps | #
14.5 | • Individual controls for gain, Pulse Repetition Frequency (PRF), baseline, filters, sensitivity, | #
14.6 | • Dynamic motion differentiation | #
14.7 | • Adjustable region of interest: size and position | #
14.8 | • User selectable Persistence button / knob. | 2
14.9 | • User selectable Blending button / knob. | 2
14.10. | Highly sensetive mode for small vessel visualization
15 | Connectivity
15.1 | • One USB ports on control panel | 2
15.2 | • At least 80 GB hard drive space or more | 2
15.3 | • Internal slot-load CD/DVD RW drive | 2
15.4 | • DICOM print, store, and storage commitment | #
15.5 | • DICOM structured reporting capability | 2
15.6 | • Modality worklist | #
15.7 | • A DICOM images must exported with the DICOM reader. | 2
15.8 | The DICOM conformance statement must be included. | 2
15.9 | • Export data as PC compatible or DICOM files | #
15.1 | • Ethernet at 100 Mb/second | 2
15.11 | • B/W printers to be supplied. | 1
15.12 | • The ultrasound unit must be DICOM and HL7 compliant and enabled to be PACS ready without any additional costs. | 2
16 | Probes: the following probes must be included in the price. Give prices for individual probes.
16.1 | The selected frequency of the imaging transducer must be displayed on the screen. | #
16.2 | Three transducers must be included and the end user must be able to select the final transducers Clearly State all available transducers with prices per transducers (swapping of products/transducers and software to be allowed with equitable costs). Details are stated | 2
16.3 | The transducers selected by the end-user must be supplied with the relevant software. | 1
16.4 | The relevant and compatible software are the following in line with the transducers selected: MSK, TEE adult, TEE pediatric, pediatric TTE, adult TTE, small parts, vascular, breast, abdominal and obstretrics & gynaecology. | 2
16.5 | additional probes and software must be available as optional items. Optional items must be listed and priced separately. | 1
16.6 | Hockey stick optional | 1
16.7 | Convex Probe (Curvilinear). The MHz range are stated (within 1-6 MHz) | #
16.8 | Microconvex for neonatal/ paediatrics.The MHz range are stated (4-8) | 2
16.9 | Microconvex Transvaginal/transrectal. The MHz range are stated (within 3-8 MHz) | 2
16.1 | Linear probe (vascular) The MHz range are stated (4-12) | 2
16.11 | Linear probe (MSK)The lower MHz range must be 5-6 MHz and the upper range must be between 14 and 21 MHz. | 2
16.12 | Adult Cardiac Probe. State the MHz range (1-4MHz) | 2
16.13 | Paediatric Cardiac Probe. State the MHz range (2-10 MHz) | 2
16.14 | TEE Probe, adult. The lower MHz range must be 2MHz and the upper range must be 9MHz. | 2
16.15 | TEE Probe, pediatric. The lower MHz range must be 3MHz and the upper range must be 11MHz. | 2
17 | SOFTWARE
17.1 | Compound Imaging ( Cross Beam angles) | #
17.2 | Speckle Reduction (Minimum of 4-5 steps ). Refines images(reduces noise) | #
17.3 | System must have software to enhance needle visualization during biopsy | 2
17.4 | Measurement and Report software packages must be configurable | 1
17.5 | Details of tools/software provided as a standard are stated | 1
17.6 | The standard software packages supplied must be fully upgradable . | #
17.7 | The cost of updating and upgrading all standard software packages supplied , during warranty and maintenance period must be included in the offer . It must be clearly outlined as a line item in the offer documentation. | #
17.8 | A rechargeable battery with a charger must be supplied. | 1
17.9 | Fully charged battery must have an actual time use of one hour or more. | 2
17.10. | The weight of the portable ultrasound unit must be less than five kg. | 2
17.11 | A multi port (for transducer connectivity) cart with castors with two wheel lock brakes | 2
17.12 | A compact carry case with wheels for the ultrasound unit must be supplied. | 2
17.13 | Paediatric and adult ECG cables (three lead) must be supplied. | 2
18 | OPTIONAL ITEMS
18.1 | USB PORT additional | 1
18.2 | Any additional tools/software not included in the standard package must be clearly listed and priced. | 1
18.3 | Elastography software | 1
18.4 | 3D software | 1
18.5 | 4D software | 1
19 | POWER SUPPLY:
19.1 | The equipment offered shall operate from a standard 220 to 240 Volt A.C. 50Hz mains supply. The bidder must ensure that the product being quoted for is fitted with a 15 Amp sealed plugtop S.A.N.S approved
19.2 | The unit must be supplied with an on-line uninterrupted power supply (U.P.S.) to provide minimum of 50 to 60 minutes mains backup power.
19.3 | The UPS should not adversely affect the image quality.
19.4 | This unit must be included in the bid price and be covered by the guarantee and maintenance plan.
19.5 | The UPS should not adversely affect the image quality.

## Doppler ultrasound high-end
THE ART TOP END PREMIUM DIAGNOSTIC COLOUR DOPPLER ULTRASOUND UNIT WITH PULSED WAVE DOPPLER
ITEM | TECHNICAL SPECIFICATIONS
1 | IMAGING MODALITIES
1.1 | This unit must be fully digital high end Colour Doppler Ultrasound Unit
1.2 | This unit must be latest technology, please state year in which this unit was first manufactured
1.3 | Unit should have 4 active Probe connectors with electronic switching
1.4 | The unit must offer Touch command screen for easy operation
1.5 | The unit must offer a 19” high resolution,high quality LCD colour monitor
1.6 | Panel and Touch Command Screen must be fully programmable
1.7 | Panel must be able to swivel left to right and up and down for ease of work
1.8 | The unit should offer fully Advanced Dicom software with Modality worklist included in main offer
1.9 | Unit must offer Pulse Subtraction technique
1.1 | Conventional Tissue Harmonics must be offered on all probes
1.11 | Differential Tissue Harmonics to improve spatial resolution and tissue visualization must be offered
1.12 | Multiple compounding technology to increase contrast resolution should be offered
1.13 | Tissue Specific Optimization to corrects for differences in acoustic characteristics of tissue in the body to improve lateral resolution
1.14 | B-Mode must offer single , dual and quad display
1.15 | System offered not less than 2 x 350 GB hard drive
1.17 | Unit should offer depth gain in not less than 8 steps
1.18 | Pan-zoom enlargement should be up to 12 times
1.19 | Easy programmable image processing should be in not less than 8 steps using one button to enable user to change dynamic range, persistence, contrast, enhancement, smoothing and post-processing
1.2 | Independent M-mode gain adjustment should be available.
1.21 | Frame rate of 500 frames should be offered
1.22 | Unit should not offer not less than a 202 transmitting channels
1.23 | B-Mode image on linear probe must be electronically steerable not less than 10 degrees left and right
1.24 | Unit should operate on Windows 10
1.25 | Maximum depth should not be less than 40 cm
1.26 | System dynamic range should not be less than 172dB
1.27 | Pan-zoom enlargement should be in real and frozen mode up to 2cm
1.28 | Contrast Harmonics software must be available. Please quote optionally
1.29 | The unit should offer trapezoid software capabilities on all Sector and Linear Probes
1.30. | Unit should support and be upgradeable to Panoramic software and should offer not less than 400cm in length
1.31 | Images not less than 150 000 must be able to be stored on hard drive
1.32 | Elastography software ,which displays the elasticity of tissue in real time and MicroPure software to depicts small calcifications must be included
1.33 | Automatic gain and STC control as well as velocity range and base line for Spectrum dopplerr through one touch operation must be available on unit.
1.34. | Unit must be upgradeable to display volume data obtained by another modality in synchronization with the ultrasound image. - fusion
2 | DOPPLER
2.1 | PW Doppler on all imaging probes.
2.1.1 | Max detectable velocity for PW must be stated.
2.2 | Pulse Repetition Frequency should varies between 0.5 to 24kHz
2.2.1 | Unit must come with the Continuous wave doppler
2.2.2 | Max detectable velocity must be stated.
2.3 | Independent angle steering up to ±30° for colour and Doppler must be offered on Linear Probes
2.4 | Colour pallet offering 7 types independent of B-Mode and Pulse Doppler pallet to be available.
2.5 | Sampling with Pulse Doppler should be varies between 1 – 20mm
2.6 | Display of Doppler scan modes should be offered side by side and up and down
2.7 | Unit must offer real time triplex mode
3 | COLOUR DOPPLER
3.1 | Minimum of 20 colour maps should be available for colour display.
3.2 | Colour maps must be able to be selected after frozen mode.
3.3 | Independent colour gain adjustment 1 – 16 dB in 0.5 dB steps
3.4 | Maximum colour frame rate should not be less than 307 frames /sec
3.5 | Colour flow must operate to the full depth of 40cm.
3.6 | There should be a minimum of 8 processing selections i.e. wall filters, colour averaging etc. Tendered to state the Number and list the possible selections.
3.7 | Colour maps and processing selections must be able to be
Selected by the user and stored in the presets.
3.8 | Advanced Dynamic flow should be offered for vascular recognition imaging on all transducers
3.9 | Contrast Harmonic Imaging should be available
3.10. | Directional Colour Angio should be available to display the existence of blood flow and the direction of flow movement with high sensitivity
3.11 | Unit offered must be upgradable to do high resolution 3D/ 4D applications
4 | TRANSDUCERS
4.1 | All probes required must be offered and must be included in main offer.
4.2 | All imaging transducers must be multi-frequency.
4.3 | Tissue Harmonics should be available on all probes
4.4 | The selected frequency of the imaging transducer must be displayed on the screen.
4.5 | All transducers must be digitally controlled so that continuous beam forming is possible
4.6 | It must be possible to connect 4 probes & a pencil probe simultaneously, and should be electronically selectable.
4.7 | THE FOLLOWING TRANSDUCERS MUST BE OFFERED AND INCLUDED IN THE MAIN OFFER
4.7.1 | High frequency linear probe for breast imaging from 8MHz t0 18MHz
4.7.2 | Linear probe frequency range 6MHz to 12Hz
4.7.3 | Adult cardiac probe 5 MHz
4.7.4 | Pediatric cardiac probe 7.5 MHz
4.7.5 | Convex probe 3.5 MHz
5 | MEASUREMENTS
5.1 | 16 callipers (8 channel) for measuring distance angle area volume should be offered
5.2 | Basic measurements – B-Mode : Distance , Area,Angle,Volume,Stenosis ratio
5.3 | PW/CW- mode measurements – Velocity, Acceleration, Time, Heart rate, PI , RI, S/D , Flow Volume , Doppler Trace
5.4 | Unit must offer user programmable annotation for quick labelling.
5.5 | Applications measurements should automatically be labelled on touch command screen for all applications,(B-mode and Doppler). It must automatically generate a comprehensive report.
Institutions should have the opportunity to choose their specific probe requirements.
6 | RECORDING DEVICES
6.1 | Digital Black and white thermal printer must be offered
6.2 | Unit must offer DVD/CD writer/PACS for storage/recorder for storage/recording of real-time loops and still images and USB flash drive , USB-HDD
6.3 | Recorder CD/DVD-writer should be compatible with any computer/laptop, with no additional software, images must be offered.
7 | POWER SUPPLY:
7.1 | The equipment offered shall operate from a standard 220 to 240 Volt A.C. 50Hz mains supply. The bidder must ensure that the product being quoted for is fitted with a 15 Amp sealed plugtop S.A.N.S approved
7.2 | The unit must be supplied with an on-line uninterrupted power supply (U.P.S.) to provide minimum of 50 to 60 minutes mains backup power.
7.3 | The UPS should not adversely affect the image quality.
7.4 | This unit must be included in the bid price and be covered by the guarantee and maintenance plan.
7.5 | The UPS should not adversely affect the image quality.

## Doppler utrasound mid-range
MID-RANGE DIAGNOSTIC COLOUR DOPPLER ULTRASOUND UNIT WITH PULSED DOPPLER
The unit required must be a fully digital, high performance diagnostic ultrasound unit. It must be multi-purpose. It will be used for Vascular, Cardiac, Abdominal, Obstetrics & Gynaecology examinations. Unit must be a compact, mobile and hosted on a fully integrated trolley with 4 wheels.
ITEM | TECHNICAL SPECIFICATIONS
1 | IMAGING MODALITIES
1.1 | This unit must be fully digital Colour flow Doppler, Pulsed Doppler and Continouos wave Doppler
1.2 | This unit must be latest technology, please state year in which this unit was first manufactured
1.3 | Unit should have 3 active Probe connectors with electronic switching
1.5 | The unit must offer a 19 inch high resolution LCD colour monitor
1.6 | The unit must have programmable panel switches for easy operation
1.7 | The unit must be fully mobile and hosted on trolley with 4 wheels
1.8 | The unit offered should offer maximum of 40cm depth.
1.9 | Unit must have Touch Command Screen
1.1 | Tissue Harmonics must be offered on all probes for excellent image quality
1.11 | Should have a fully adjustable panel for ease of operations
1.13 | System offered should offer image memory of not less than 340 MB
1.14 | Unit should offer depth gain in not less than 8 steps
1.15 | Pan-zoom enlargement should be up to 12 times on full screen
1.16 | Easy programmable image processing should be in less than 8 steps using one button to enable user to change dynamic range, persistence, contrast, enhancement, smoothing and post-processing
1.17 | Independent M-mode gain adjustment should be available.
1.18 | Frame rate of 500 frames should be offered
1.19 | Unit should not offer not less than a 192 transmitting channels
1.2 | Post processing should be available for Gain, Dynamic range, Depth etc.
1.21 | Patient browser should offered to look at saved patients information
1.22 | Pan-zoom enlargement should be in real and frozen mode up to 2cm
1.23 | Unit should have on-board reporting system
1.24 | The unit should offer trapezoid software capabilities on all Sector and Linear Probes
1.25 | Panoramic software should be available
1.26 | Single, Dual and Quad display should be available in real time
1.27 | Please state if unit is upgradeable to 3D and 4D capabilities
1.28 | The unit should offer fully advanced Dicom/PACS software with Modality worklist included in main offer
2 | DOPPLER
2.1 | PW Doppler on all imaging probes.
2.2 | Pulse Repetition Frequency should varies between 0.5 to 24 kHz
2.2.2 | Max detectable velocity must be stated.
2.3 | Independent angle steering up to ±30° for colour and Doppler must be offered on Linear Probes
2.4 | Colour pallet offering 7 types independent of B-Mode and Pulse Doppler pallet to be available.
2.5 | Sampling with Pulse Doppler should be varies between 1 – 20mm
2.6 | Display of Doppler scan modes should be offered side by side and up and down and in quad display
2.7 | Unit must offer real time triplex mode
2.8 | Offered with continouos wave doppler
3 | COLOUR DOPPLER
3.1 | Minimum of 20 colour maps should be available for colour display.
3.2 | Colour maps must be able to be selected after frozen mode.
3.3 | Independent colour gain adjustment 1 – 16 dB in 0.5 dB steps
3.4 | Maximum colour frame rate should not be less than 387 frames p/sec
3.5 | Colour flow must operate to the full depth of 40cm.
3.6 | There should be a minimum of 8 processing selections i.e. wall filters, colour averaging etc. Tendered to state the Number and list the possible selections.
3.7 | Colour maps and processing selections must be able to be selected by the user and stored in the presets.
3.8 | Advanced Dynamic flow should be offered for vascular recognition imaging on all transducers
3.9 | Contrast Harmonic Imaging should be available
3.10. | Directional Colour Angio should be available to display the existence of blood flow and the direction of flow movement with high sensitivity
4 | TRANSDUCERS
4.1 | All probes required must be offered and must be included in main offer.
4.2 | All imaging transducers must be multi-frequency.
4.3 | Tissue Harmonics should be available on all probes
4.4 | The selected frequency of the imaging transducer must be displayed on the screen.
4.5 | All transducers must be digitally controlled so that continuous beam forming is possible.
4.6 | It must be possible to connect 3 probes & a pencil probe simultaneously, and should be electronically selectable.
4.7 | 3.75 MHz Convex probe frequency should vary between 1.9 and 6 MHz to scan large and small patients. Probe should be capable of Tissue Harmonics and Differential Tissue Harmonics to offer sharp and clear images
4.8 | Trans-Vaginal Convex Probe with frequency between 4 – 11 MHz should be offered.
4.9 | Linear probe for vascular applications, small parts with tissue harmonics from 5MHz to 11MHz
4.10. | High frequency linear probe for muscular skeletal, breast imaging from 8MHz t0 18MHz
4.11 | Multi-frequency Convex probe 6MHz with frequencies that varies between 5MHz and 9 MHz Tissue harmonics for abdominal , vascular applications for neo –natal examinations
4.12 | Pulsed Wave Doppler and Colour Doppler must be available on both probes
4.13 | Adult cardiac probe 5 MHz
4.14 | Pediatric cardiac probe 7.5 MHz
5 | MEASUREMENTS
5.1 | 16 callipers (8 channel) for measuring distance angle area volume should be offered
5.2 | 16 callipers (8 channels) hip joint should be offered
5.3 | The unit must offer storage facility for up to 2000 patient fetal growth measurement data. The information must be able to be retrievable for follow-ups.
5.4 | Unit must offer 50th percentile growth graphs of the Doppler measurements RI /PI the information saved must be retrievable for all follow-up visits.
5.5 | 50th percentile growth graphs must be offered on all measurements for example BPD, FL, AC, Fetal Weight, HC information saved must be retrievable for all follow-up visits.
5.6 | Display of graphs must be available to compare measurements of twins, triplets in single and quad display
5.7 | Unit must offer user programmable annotation for quick labelling.
6 | RECORDING DEVICES
6.1 | Digital Black and white thermal printer must be included in the offer
6.2 | Unit must offer DVD/CD writer/PACS for storage/recorder for storage/recording of real-time loops and still images
6.3 | Recorder CD/DVD-writer should be integrated part of the unit. Software should be compatible with any computer/laptop, with no additional software,
7 | POWER SUPPLY:
7.1 | The equipment offered shall operate from a standard 220 to 240 Volt A.C. 50Hz mains supply. The bidder must ensure that the product being quoted for is fitted with a 15 Amp sealed plugtop S.A.N.S approved
7.2 | The unit must be supplied with an on-line uninterrupted power supply (U.P.S.) to provide minimum of 50 to 60 minutes mains backup power.
7.3 | The UPS should not adversely affect the image quality.
7.4 | This unit must be included in the bid price and be covered by the guarantee and maintenance plan.
7.5 | The UPS should not adversely affect the image quality.

## Ultrasound high-TEE
ULTRASOUND MACHINE, HIGH END WITH TEE
ITEM | TECHNICAL SPECIFICATIONS | Weight
1 | Ultrasound high end with TEE
1.1 | Unit should offer at least a 17 to 19 inch LCD high resolution colour monitor. State the size and technology. | 2
1.2 | Unit must be latest technology. two year of which the unit was manufactured. | 2
1.3 | The Unit must be a Fully Digital high end colour doppler ultrasound unit with excellent colour and power sensitivity while maintaining excellent detail resolution. | 2
1.4 | Unit must be housed on a lightweight mobile cart with lockable wheels. | 2
1.7 | The Width of the system in cm 100 to 156cm | 2
1.8 | The Depth of the system must be 100 to 156 of the system in cm | 2
1.9 | User configurable hard keys | 2
1.1 | User interface must be height adjustable | #
1.11 | the maximum number of channels 512 | 2
1.13 | the system dynamic range must be 45 to 90db | #
1.15 | Unit should offer hard drive with image storage space of 300GB to 500GB . | 2
1.16 | CD/DVD writer with the required software for active use, must be standard and integrated in the system | #
1.17 | Recorded data on CD/DVD must be readable on any standard laptop/computer with additional software, if necessary from the supplier | #
1.18 | The system must be DICOM compliant | #
1.19 | Advanced DICOM with full worklist, transmission and print must be offered and be included in main system | #
1.2 | Export of images from the system must be possible. (including transfer of images to PACS). State the methods. | #
1.21 | The system must have a touch command screen for ease of operations | #
1.22 | Digital black and white thermal printer must be offered as part of the system. | #
2 | Imaging Modes
2.1 | Unit should be capable of offering a maximum frame rate of not less than 500 frames per second. | #
2.2 | The Maximum scanning depth of the abdominal convex probe/transducer must be 40cm and above . State the maximum scanning depth of the range of probes provided | #
2.3 | Independent gain control for Pulsed Wave | #
2.4 | The PW sampling range 0.2 to 20mm (gate sizes) | 2
2.5 | State the Pulse Repetition Frequency 10000 to 19000Hz. | 2
| Streerable continuous wave repetition frequency | 2
2.7 | State the Continuous wave repetition frequency 1.56 to 34.7Khz . | 2
2.8 | detectable velocity for CW Doppler to be 30 to 650cm per second | 2
2.9 | 2D image on linear transducers must be electronically steerable. The are up to 7 steering angles available on linear transdusers | 1
2.1 | State the degree of Independent angle steering. | 1
2.11 | Post processing must be available on frozen images | #
2.12 | Independent M-Mode gain adjustment must be available | #
2.13 | Indicate if the unit are able to support Elastography Software to display the elasticity of tissues in real time | 1
2.14 | Indicate if the unit are able to support a software that depicts small calcifications. | 1
2.15 | Indicate if the unit are able to support Panoramic Software. If so, state the maximum length on offer. | 1
2.16 | Indicate if the unit are able to offer trapezoid sofware capabilities on all sector and linear probes. | 2
2.17 | The Grey scale level must be 256 as a minimum. | 2
2.18 | maximum amount of images that can be stored on the hard drive must be 200 000 to 300 000 images. | 2
2.19 | State the maximum storage capability in GB that is available on your hard drive. | 2
2.2 | Conventional Tissue Harmonics must be offered on all probes. | #
2.21 | Differential Tissue Harmonics to improve spatial resolution and tissue visualization must be offered. | 2
2.22 | Indicate if the unit are able to offer pulse substraction technique. | 1
2.23 | 2D Loop up to 400 frames should be recorded. State the total frames if ˃ than 400. | 1
2.24 | Indicate the maximum length of the captured loop in minutes. | 2
2.25 | Real time triplex mode must be standard. | #
3 | Doppler and Colour Doppler
3.1 | State the different Doppler display modes i.e.. side by side or top to bottom | 2
3.2 | All probes must offer Pulsed Wave Doppler. | 2
3.3 | State the maximum detectable velocity for Pulsed Wave. The maximum detectable velocity for pulse wave 350 - 500 cm/per sec. | 2
3.4 | Image memory to operate in both dual and quad screen format. | 2
3.6 | State if simultaneous display of the colour image and 2D image, side by side, in real-time is possible. | 2
3.7 | Indicate the range of independent colour gain adjustment in 1 dB steps. | 2
3.8 | The colour frame rate should be 400 or above. State maximum colour frame rate. | 2
3.9 | Colour flow must be fully operational through the total depth of the image. | #
3.1 | On screen measurements must be possible from real time examination. | #
3.11 | Unit must offer user programmable annotation for quick labelling. | #
3.12 | Storage and review of patent data must be available to compare exams. | #
3.13 | Indicate the maximum number of processing selections available, for example, wall filters, colour averaging. List the processing selections available.Indicate if the user are able to store them in presets. | 2
3.14 | Indicate if Tissue Doppler Imaging (TDI) for cardiac application are standard on B-Mode and M-Mode in the unit. | 2
3.15 | State type of remote control device that links with system during difficult exams. (Please quote optionally) | 1
3.16 | Controlling the system remotely with a remote control | 1
3.17 | Controlling the system remotely with voice commands | 1
3.18 | Controlling the system remotely by any other method. Explain method if applicable. | 1
3.18.1 | Indicate if Contrast harmonics software are be available. Quote optionally | 1
3.18.2 | Indicate if Advanced Dynamic Flow are offered on all transducers for vascular recognition imaging | 2
3.19 | Indicate if Directional Colour Angio are available to display the existance of blood flow and the direction of flow movement with high sensitivity | 2
4 | Transducers/ Probes
4.1 | The selected frequency of the imaging transducer must be displayed on the screen. | #
4.2 | Five transducers must be included and the end user must be able to select the final transducers. Clearly State all available transducers with prices per transducers (swapping of products/transducers and software to be allowed with equitable costs). Details are stated. | 2
4.3 | Additional probes and software must be available as optional items. Optional items must be listed and priced separately. | 2
4.4 | Types and frequency range of Transducers
4.4.1 | Convex Probe (Curvilinear) with tissue harmonics. The MHz range are stated (within 1-6 MHz) | #
4.4.2 | Microconvex for neonatal / paediatrics with tissue harmonics.The MHz range are stated (4-8) | #
4.4.3 | Microconvex Transvaginal / transrectal with tissue harmonics. The MHz range are stated (within 3-8 MHz) | #
4.4.4 | Linear probe (vascular) with tissue harmonics. The MHz range are stated (4-12) | #
4.4.5 | Linear probe (MSK) with tissue harmonics.The lower MHz range must be 5-6 MHz and the upper range must be between 14 and 21 MHz. | #
4.4.6 | Adult Cardiac Probe with tissue harmonics. State the MHz range (1-4MHz). | #
4.4.7 | Paediatric Cardiac Probe. State the MHz range (2-10 MHz). | #
4.4.8 | Neonatal Cardiac Probe. State the MHz range (2-10 MHz). | #
4.4.9 | Mono-plane Endorectal/ (transrectal ) transducer with tissue harmonics . State the MHz range. | 2
4.4.10 | Biplane Endorectal (transrectal) transducer, linear and linear , with tissue harmonics . State the MHz range. | 2
4.4.11 | Biplane Endorectal (transrectal)transducer, linear and convex , with tissue harmonics . State the MHz range. | 2
4.4.12 | Transesophageal echocardiography (TEE)with tissue harmonics. State the MHz range. | #
4.4.13 | Paediatric Transesophageal echocardiography (TEE)with tissue harmonics. State the MHz range. | #
4.4.14 | State your advanced cardiac applications as an option. | 2
4.4.15 | State your Speckle imaging application as an option | 1
4.4.16 | State your 2D, 3D and 4D MPR applications as options | 1
4.4.17 | Tissue harmonics should be available on all offered transducers | #
4.4.18 | Indicate the footprint size of the Pediatric sector transducer | 1
4.4.19 | The selected frequency of the imaging transducer must be displayed on the screen. | #
4.4.20 | All transducers must be digitally controlled so that continuous beam forming is possible | #
4.4.21 | System must have 4 or more active imaging transducer connectors with electronic switching between probes. | 2
5 | Measurements
5.1 | 16 callipers (8channel) for measuring distance/volume/angle and area must be offered | 2
5.2 | 16 callipers (8 channels) hip joint should be offered | 2
5.3 | State the storage facility offered for patient fetal growth measurement data files. The information must be able to be retrieved for follow-up comparison. 250GB must be offered. | 2
5.4 | Basic Measurements: -B Mode: Distance , area, angle, volume, stenosis ratio; 4D mode basic measurements: Distance, Volume; M-Mode, basic measurement: Slope, Distance, Time, Heart Rate; PW/CW Mode Measurements: acceleration, velocity, time,heart rate, PI,RI, S/D,Flow Volume, Doppler Trace | #
5.5 | Unit must offer 50th percentile growth graphs from the Doppler measurements, RI/PI. The information must be able to be retrieved for follow-up comparison | 2
5.6 | 50th percentile growth graphs must be offered on all measurements for example BPD, FL, AC, Fetal weight, HC. The information must be able to be retrieved for follow-up comparison | #
5.7 | Display of graphs must be available to compare measurements of twins or triplets in single and quad display format | #
5.8 | System must offer user programmable annotation for quick labelling | #
5.9 | Applications measurements should automatically be labelled/displayed on the touch command screen for all applications, (2D mode and Doppler). Comprehensive reports must be generated automatically. | 2
5.1 | Chair, ergonomic, height adjustable with lockable castors for the Health professional. | 2
5.11 | Foot rest as a standard | 1
5.12 | Cable tray /hooks ..cable support system | 1
5.13 | easy removable air filters | 2
5.14 | Gel warmer. | 2
5.15 | USB PORT additional | 2
6 | Recording devices
6.1 | Digital colour thermal printer must be offered as a standard. Printer mounted on the machine | 1
7 | POWER SUPPLY
7.1 | The unit must be supplied with an on-line uninterrupted power supply (U.P.S.) to provide minimum ten 10 to 20 minutes mains backup power. | #
7.2 | The UPS should not adversely affect the image quality.
7.3 | This unit must be included in the bid price and be covered by the guarantee and maintenance plan.
8 | OPTIONAL ITEMS
1 | fold down monitor for easy transportation | 1
| Measurement and Report software packages | 2
| Details of tools/software provided as a standard and those as optional, are stated | 1
| side storage bin | 1
| Ultrasound Quality Control Phantom (swapping of phantoms, products and software to be allowed with equitable costs). Details are stated. | 2

## UltrasoundECO high-TEE
ULTRASOUND MACHINE, HIGH END, ECHOCARDIOGRAM UNIT WITH 4D/3D CAPABILITY WITH TEE
ITEM | TECHNICAL SPECIFICATIONS | Weight
1 | ULTRASOUND HIGH END ECHOCARDIOGRAM UNIT WITH 4D\3D CAPABILITY
1.1 | Top of the range High-End Ultrasound with Echocardiography | #
1.2 | Digital broadband beam | #
1.3 | Digital channels more than 500 000. | 2
1.4 | 2D Imaging | #
1.5 | Real time 4D (Live 3D) using matrix transducer technology for instantaneous volume rendering of cardiac anatomy | #
1.6 | Live biplane/ Triplane Imaging using Matrix transducers (infinite plane adjustments) | #
1.7 | Dynamic range of 250dB and above | #
1.8 | Real-Time Compounding (Spatial Compound) | #
1.9 | Speckle Reduction Software | #
1.10. | Speckle Trace/ Tracking Software (for wall motion abnormalities) | #
1.11 | Adaptive Image Processing for noise and artifact reduction | #
1.12 | Ergonomic transducers designs with lightweight flexible cables | #
1.13 | Color Frame Rates of 650 fps and above | #
1.14 | 2D Frame rates of 800 fps and above | #
1.15 | High-resolution interactive graphical color touch panel with adjustment for various ambient light conditions | 2
1.16 | Fully articulating flicker-free at least 19-inch to 20-inch wide format high resolution flat panel TFT with nearly infinite positioning adjustments | 2
1.17 | Fully articulating control panel, including height, swivel, and slide | #
1.18 | Must have minimum of four active transducer /probe ports and one compulsary non- imaging transducer (pencil probe) port. Thus 5 ports . | #
1.19 | Min /Max velocity for PW Doppler | #
1.2 | Min /Max velocity for CW Doppler | #
1.21 | Min /Max PRF for PW Doppler | #
1.22 | Intelligent one-button optimization in 2D and Doppler modes | #
1.23 | Integrated Standby Battery | #
1.24 | Unit must be housed on a lightweight mobile cart with lockable wheels. | 2
2 | CLINICAL APPLICATION IN SOFTWARES
2.1 | The echocardiogram machine must have the following software applications (protocols):
2.1.1 | Adult Echocardiography | #
2.1.2 | Pediatric Echocardiography (inclusive of neonatal) | #
2.1.3 | Stress Echocardiography (Treadmill and Chemical) | #
2.1.4 | 2D Transesophageal Echocardiography | #
2.1.5 | 4D / Live 3D Transthoracic Echocardiography (TTE) | #
2.1.6 | 4D / Live 3D Transesophageal Echocardiography | #
2.2 | Full Cardiac Echocardiogram Software Package | #
2.3 | Full Vascular Software Package | #
2.4 | The standard software packages supplied must be fully upgradable . | #
2.5 | The cost of updating and upgrading all standard software packages supplied , during warranty and maintenance period must be included in the offer . It must be clearly outlined as a line item in the offer documentation. | #
2.6 | Automated Intima Thickness Measurement | 2
3 | IMAGING MODES AND IMAGE DISPLAY
3.1 | The operating modes should include: 2D, M, 2D/M, 4D / Live 3D, PWD, High PRF, Color Doppler, Power Doppler, Directional Power Doppler, steerable CWD, Triplex and Harmonic with pulse inversion technology. | #
3.2 | Tissue Doppler with frame rate of 600 fps and above. | 2
3.3 | M mode color Doppler and M mode tissue Doppler | #
3.4 | Innovative Transducer crystal technology with a pure wave (or similar) capability for greater acoustic efficiency and bandwidth | 2
3.5 | Innovative Transducer crystal technology with active matrix array capability for greater acoustic efficiency | 2
3.6 | 4D / Live 3D Volume Imaging using matrix transducer | #
3.7 | 4D / Live 3D Zoom | #
3.8 | 4D / 3D Full volume | #
3.9 | 4D / Live 3D Quantification with MPR and islice. | #
3.1 | 4D / Live 3D display of dynamic 3D rendering and LV true volumes. | #
3.11 | Quantification of the mitral valve using the 4D / live 3D transesophageal transducer. | #
3.12 | Live (on screen) automatic Doppler Analysis | 2
3.13 | Automatic Doppler Analysis on still images. (autotrace) | #
3.14 | Tissue Specific Imaging Application Software | #
3.15 | Onboard quantification for :
3.15.1 | a. Automated Strain Imaging and Speckle tracking. | #
3.15.2 | b. Display of automated borders | #
3.15.3 | c. Display of Mitral valve annular motion | #
3.17 | Dual Live display with BW reference image and color Doppler image | #
3.18 | Colorization of volume imaging | #
3.19 | Trapezoid scanning providing convex view on linear probes. | #
3.2 | High Definition Write Zoom and Read Zoom with pan feature | #
3.21 | Scanning depth ranges for adult echocardiac probes of a minimum depth between 0 cm- 2cm, and a maximum depth between 30 cm-40cm | #
3.22 | Measurement tools including: distance, depth, area, and circumference | #
3.23 | Volume Flow Measurements | #
3.24 | On-board workstation-class data management with thumbnail previews and storage of images, loops, and reports | #
3.25 | Retrospective and prospective clip capture to internal drive or removable media | #
3.26 | Integrated DVD/CD burning capability for storage of DICOM 3 images or export in JPEG and avi for PC compatibility | #
3.27 | DICOM 3.0 Print and Store capability to internal drive or DVD/CD | #
3.28 | Internal HDD at least 1 TB | 2
3.29 | One Full Off-line Work station with Full Reporting Systems | #
3.3 | Chair, ergonomic, height adjustable with lockable castors for the Health professional. | 2
3.31 | Gel warmer | 2
3.32 | Foot rest | 1
3.33 | Cable tray /hooks ..cable support system | 1
3.34 | Easy removable air filters | 2
3.35 | USB PORT additional | 2
4 | TRANSDUCERS
4.1 | The following transducers below must be included as standard. Details are stated:
4.1.1 | Adult Sector transducer (1 MHz Hz to 5 Mhz) using innovating single crystal technology. | #
4.1.2 | Paediatric Sector transducer (4 Mhz to 12 MHz) using innovating single crystal | #
4.1.3 | Neonatal Sector transducer (2 Mhz to 8 MHz) using innovating single crystal | #
4.1.4 | Matrix transducer for transthoracic 4D / Live 3D Echo | #
4.1.5 | Matrix Trans-esophageal transducer (TEE) using matrix array technology for 4D / Live 3D trans-esophageal Imaging | #
4.1.6 | Matrix Trans-esophageal transducer (TEE) using Single crystal technology for 2D imaging | #
4.1.7 | Matrix Paediatric Trans-esophageal transducer (TEE) using Single crystal technology for 2D imaging | #
4.1.8 | Linear vascular Carotid transducer Probe frequency between 4-6 MHz lower range and 12 to 14 MHz upper range | #
4.1.9 | Non-Imaging Pencil Probe with 1.9 MHz and above | #
4.2 | The following transducers must be included as optional. The end user must be able to make the final selection of transducers. Clearly State all available transducers with prices per transducers (swapping of products/transducers and software to be allowed with equitable costs). Details are stated.
4.2.1 | Convex Probe (Curvilinear) with tissue harmonics. The MHz range are stated (within 1-6 MHz) | #
4.2.2 | Microconvex for neonatal / paediatrics with tissue harmonics.The MHz range are stated (4-8) | #
4.2.3 | Microconvex Transvaginal / transrectal with tissue harmonics. The MHz range are stated (within 3-8 MHz) | #
4.2.4 | Linear probe (vascular) with tissue harmonics. The MHz range are stated (4-12) | #
4.2.5 | Linear probe (MSK) with tissue harmonics.The lower MHz range must be 5-6 MHz and the upper range must be between 14 and 21 MHz. | #
4.3 | New active matrix array technnology on transthoracic and Transesophageal transducers | 1
4.4 | The selected frequency of the imaging transducer must be displayed on the screen. | #
5 | GENERAL
5.1 | Small Black White photo printer integrated into the system | #
5.2 | Measurement and Report software packages | #
5.3 | Minumum of 3-Lead ECG Cable with connecting port | #
5.4 | Integrated spyrometric function(display respiratory motion) | #
7 | Optional
7.1 | Ultrasound Quality Control Phantom (swapping of phantoms, products and software to be allowed with equitable costs). Details are stated. | 2
7.2 | side storage bin
7.3 | A second Full Off-line Work station with Full Reporting System as optional item | 2
7.4 | External Hardrive of 1 TB as optional item | 2
8 | POWER SUPPLY
8.1 | The unit must be supplied with an on-line uninterrupted power supply (U.P.S.) to provide minimum ten 10 to 20 minutes mains backup power.
8.2 | The UPS should not adversely affect the image quality.
8.3 | This unit must be included in the bid price and be covered by the guarantee and maintenance plan.

## Thermal Printer and Paper
THERMAL PRINTER AND PAPER
Item | Technical Specifications
1 | Thermal Printer
1.1 | General
1.1.1 | The unit must be compact A6 medical grade black and white digital printer that can accept analogue video as well as PC-friendly digital USB signal inputs.
1.1.2 | The unit must be able to produce hard copy prints of still images captured by ultrasound units for use in patient records and referrals.
1.1.3 | The printed still images must be storable to a connected USB flash drive.
1.1.4 | The unit must have LCD display panel and joystick for intuitive menu navigation.
1.1.5 | It must be capable of delivering high-quality black & white prints in under three seconds.
1.1.6 | The thermal printing technology should be able to enable photoquality black & white prints with a high resolution up to approximately 300 dpi and grey levels up to 240 steps or better.
1.1.7 | The unit must have a hybrid operation with analogue and digital inputs
1.1.8 | Dimensions of the unit: a minimu of 80mm (H) x 150 mm (W) x 230 mm (D) or better. Bidders must state their dimensions
1.1.9 | PC compatibility: the unit must be compatible with the latest Windows operation system
1.2 | Technical Specifications
1.2.1 | Printing Method: Direct thermal printing
1.2.2 | Resolution: 325 dp
1.2.3 | Gradations: 8 bits (256 levels) processing
1.2.4 | Picture Elements
1.2.4.1 | Digital: 4096 x 1280 dots
1.2.4.2 | Video NTSC: 720 x 504 dots
1.2.4.3 | Video PAL: 720 x 604 dots
1.2.5 | Picture Area
1.2.5.1 | Digital: 320 x 100 mm/12 5/8 x 3 7/8 inch (Max)
1.2.5.2 | STD Video NTSC: 94 x 73 mm, Video PAL: 94 x 71 mm (WIDE1) when “SIDE:OFF” is selected
1.2.5.3 | SIDE Video-NTSC: 124 x 96 mm, Video-PAL: 127 x 96 mm (WIDE1) when “SIDE:ON” is selected
1.2.6 | Paper Size Paper width of 110 mm
1.2.7 | Printing Time
1.2.7.1 | High-speed mode: Approx. 1.9 seconds/image (at standard setting)
1.2.7.2 | Normal speed mode: Approx. 3.3 seconds/image (at standard setting)
1.2.8 | Picture Memory
1.2.8.1 | Digital: 4,096 x 1,280 x 8 (bit)
1.2.8.2 | Video: 10 frame memories (850 k x 8 bits per frame)
1.2.9 | Interface
1.2.9.1 | USB terminal (type A) for USB flash drive (X1)
1.2.9.2 | Hi-Speed USB (USB 2.0) (x1)
1.2.9.3 | VIDEO INPUT: BNC type (x1) NTSC or PAL composite video signals 1.0 Vp-p, 75 ohms (NTSC/PAL automatically discriminated)
1.2.9.4 | VIDEO OUTPUT: BNC type (x1) Loop-through
1.2.9.5 | REMOTE: Stereo mini jack (x1)
1.2.10 | Printer Driver Software
1.2.10.1 | Microsoft Windows 7 32/64 bit, 8 32/64 bit, 10 32/64 bit
1.2.11 | Power Requirements AC 100 V to 240 V, 50/60 Hz
1.2.12 | Input Current 1.3 A to 0.6 A
1.2.13 | Operating Temperature 5°C to 40°C (41°F to 10 4°F)
1.2.14 | Operating Humidity 20% to 80% (no condensation allowed)
1.2.15 | Mass 2.5 kg (5 lb 8 oz)
1.2.16 | Supplied Accessories
1.2.16.1 | Thermal head cleaning sheet
1.2.16.2 | Before Using this Printer
1.2.16.3 | Service Contact List
1.2.16.4 | USB Flash
1.2.16.5 | Drive Ex. Cable
1.2.16.6 | Print media (UPP-110HG)
2 | Thermal Paper
2.1 | General
2.1.1 | It is an A6 width B&W High Density print media (type II) for use in thermal printers
2.1.2 | The paper must be usable in ultrasound
2.1.3 | The paper must be available in multiples of atleast 5 x 1 unit (master carton contains at least 100 rolls).
2.2 | Technical Specifications
2.2.1 | Roll dimensions: 110 mm (W) x 20 m
2.2.2 | Order unit: atleast 5 rolls
2.2.3 | Print quantity: Approximately 220 prints per roll (Picture size: 100x75mm)
2.3 | Optional
2.3.1 | The unit must allow paper to be cut in two directions – both up and down
2.3.2 | Wireless Printing Capability: The unit must allow a cable-free connection between the printer and ultrasound units